## Metadata+Aeris Parameters
Mineralogy and Microscopy Laboratory
Geosciences and Environmental Change Science Center
United States Geological Survey
X-ray Diffraction Report
Date of Analysis | 10/10/23 - 4/25/24
Analyst Name | J. Leidelmeijer
Project | Etosha Pan, Namibia Clays
X-Ray Diffraction Parameters
Instrument | Benchtop Malvern Panalytical AERIS Research Edition X-ray Diffractometer
Measurement
Measurement Program | GEC Clays SOP
Scan Parameters
Description | GEC Clay Scan for Phase Analysis (Cu Tube)
Scan Axis | Gonio
Start Angle | 2.0054 ° 2θ
End Angle | 29.9894 ° 2θ
Step Size | 0.011 ° 2θ
Time per step | 48.195 seconds
Scan Type | Continuous
Hardware Settings
X-ray Source, Copper (Cu)
Beta Filter | NONE
Soller Slits | 0.04 rad.
Divergence Slit | 1/16°
Beam Mask | Fixed Mask 20 mm
Stage
Sample Stage | 1 rps
Beam Knife | Low
PIXcel1D-Medipix3 Detector
Mode | Scanning Line Detector (1D)
Active Length | 5.542 ° 2θ
Anti-Scatter Slit/PPC | AS slit 9mm
Soller Slits | Large Soller Slits 0.04 rad.
Beta Filter | NICKEL
Sample Preparation
Carbonate Dissolution | 10% Acetic Acid
Clay Separation | Centrifugation
Sample Mount | Zero Background Disc
Gycolization | Vapor Treatment in 60C glycolator for 4 hours
Heating Samples | 400 C for 1/2 hour AND 550 C for 1/2 hour
Sample Mount | Zero Background Disc
X-ray Scan Mineral Identification and Quantification
X-Ray Diffraction Software | HighScore (version 5.1a (5.1.1.30138) HighScore details available at Malvern Panalytical.
Mineral Identification Resource | International Centre for Diffraction Data (ICDD)
Mineral Database | Powder Diffraction File-4/Minerals 2023
Software Version | 4.22.1.2
Database Version | 4.2211
Java Version | AdoptOpenJDK 1.80_275 (64-bit)

## SQ Mineralogy Table
SQ Mineralogy | Sample Name | Mineralogy | Actinolite | Ajoite | Analcime | Anorthite | Anthrophylite | Biotite | Carlfrancisite | Cassiterite | Celsian | Clinozoisite | Furongite | Ganophyllite | Hexahydrite | Hornblende | Illite | Leucite | Loughlinite | Montmorillonite | Microcline | Muscovite | Nontronite | Orthoclase | Palygorskite | Parnauite | Phologopite | Quartz | Rutile | Sanidine | Sepiolite | Tobermorite | Vermiculite | Zeolite | TOTAL
1 | 1 Etosha 18 RIP 0-2 cm | Regular |  |  | 39 |  |  |  | 10 | [9 empty] | 9 |  |  | 21 |  |  | 4 | [5 empty] | 18 |  |  |  | 101
|  | 400C |  |  | 39 |  | 7 |  | 11 | [9 empty] | 5 |  |  | 24 |  |  | 6 | [5 empty] | 7 |  |  |  | 99
|  | 500C |  |  | 13 |  | 12 |  | 16 | [12 empty] | 51 |  |  | 7 | [9 empty] | 99
|  | Glycolated |  |  | 36 | [13 empty] | 18 |  |  | 34 | TR |  | 6 | [5 empty] | 7 |  |  |  | 101
2 | 2 Etosha 18 SS 0-2 cm | Regular |  |  | 16 |  |  | 16 | [10 empty] | 27 | [7 empty] | 15 |  |  |  | 27 |  |  |  | 101
|  | 400C |  |  | 32 |  |  | 12 | 18 | [17 empty] | 25 |  |  |  | 5 |  |  | 8 | 100
|  | 500C |  |  | 25 |  |  | 17 | 30 | [17 empty] | 20 | [6 empty] | 8 | 100
|  | Glycolated |  |  | 20 |  |  | 14 | 11 | [9 empty] | 19 |  |  |  | TR |  |  |  | 20 | [6 empty] | 17 | 101
3 | 3 Etosha 18 SS 0-2 cm S2 | Regular |  |  | 12 |  |  |  | 46 |  |  |  |  | 5 | [10 empty] | 6 |  | 21 | 6 |  |  | 4 |  |  |  | 100
|  | 400C |  |  | 39 |  |  |  | 45 |  |  |  |  | 2 | [10 empty] | 4 |  | 9 | [7 empty] | 99
|  | 500C |  |  | 66 |  |  |  | 11 | [17 empty] | 22 | [7 empty] | 99
|  | Glycolated |  |  | 8 |  |  |  | 34 |  |  |  |  | 3 |  |  | 44 | [9 empty] | 11 | [7 empty] | 100
4 | 4 Etosha 18 SS 0-2 cm S3 | Regular |  | 7 | 23 |  |  |  | 43 | [9 empty] | 12 | [7 empty] | 15 | [7 empty] | 100
|  | 400C |  | 4 | 64 |  |  |  | TR | [9 empty] | 4 | [7 empty] | 21 | [6 empty] | 7 | 100
|  | 500C |  |  | 42 |  |  |  | 10 | [9 empty] | 12 | [7 empty] | 36 | [7 empty] | 100
|  | Glycolated |  |  | 25 |  |  |  | 55 | [9 empty] | 9 | [7 empty] | 6 |  |  |  | 5 |  |  |  | 100
5 | 5 Etosha 18 HS 0-2 cm S1 | Regular |  |  | 28 |  |  |  | TR | TR | [8 empty] | 15 | [5 empty] | 6 |  | 39 |  |  |  | 12 |  |  |  | 100
|  | 400C |  |  | 25 |  |  |  | 19 | 2 | [8 empty] | 8 | [5 empty] | 9 |  | 27 |  |  |  | 4 | 5 |  |  | 99
|  | 500C |  |  | 43 |  |  |  | 13 | [15 empty] | 8 |  | 36 | [7 empty] | 100
|  | Glycolated |  |  | 15 |  |  |  | 51 | [9 empty] | 17 | [7 empty] | 17 | [7 empty] | 100
6 | 6 Etosha 18 Ssub 2-5 cm S1 | Regular |  |  | 21 |  |  | 7 | 45 | [9 empty] | 27 | [15 empty] | 100
|  | 400C |  |  | 25 |  |  | 25 | 42 | [9 empty] | 8 | [15 empty] | 100
|  | 500C |  |  | 30 |  |  | 29 | 29 | [9 empty] | 12 | [15 empty] | 100
|  | Glycolated |  |  | 15 |  |  | 10 | 46 | [9 empty] | 29 | [15 empty] | 100
7 | 7 Etosha 22 RIP 0-1 cm S1 | Regular |  |  | 21 |  |  | 10 | 45 | [9 empty] | 14 | [5 empty] | 4 |  | 6 | [7 empty] | 100
|  | 400C |  |  | 26 |  |  | 7 | 27 | [9 empty] | 7 | [5 empty] | 9 |  | 16 |  |  |  | 7 |  |  |  | 99
|  | 500C |  |  | 32 |  |  | 35 | 16 | [15 empty] | 17 | [9 empty] | 100
|  | Glycolated |  |  | 22 |  |  |  | 53 | [9 empty] | 9 | [5 empty] | 6 |  | 7 |  |  |  | 3 |  |  |  | 100
8 | 8 Etosha 22 CRT 0-0.1 cm S1 | Regular |  |  | 12 |  |  | 13 | 26 |  |  |  |  | 23 | [11 empty] | 5 |  |  |  |  | 22 |  |  |  | 101
|  | 400C |  |  | 27 |  |  | 18 | 18 |  |  |  |  | 16 | [16 empty] | 20 |  |  |  | 99
|  | 500C |  |  | 24 |  |  | 32 | 44 | [25 empty] | 100
|  | Glycolated |  |  | 22 |  |  | 17 | 31 |  |  |  |  | 20 | [16 empty] | 10 |  |  |  | 100
9 | 9 Etosha 22 Sub 0.1-4 cm S1 | Regular |  |  | 11 |  |  |  | 41 | [9 empty] | 16 | [5 empty] | 3 | 3 | 13 |  |  |  | 13 |  |  |  | 100
|  | 400C |  |  | 23 |  |  |  | 34 | [17 empty] | 22 |  |  |  | 21 |  |  |  | 100
|  | 500C |  |  | 7 |  |  |  | 22 | [17 empty] | 70 | [7 empty] | 99
|  | Glycolated |  |  | 29 |  |  |  | TR |  |  |  |  | 36 | [12 empty] | 34 | [7 empty] | 99
10 | 10 Etosha 22 CrySub 4-10 cm S1 | Regular |  |  | 20 | [8 empty] | 27 |  |  |  |  | 25 | [7 empty] | 9 |  |  |  | 19 |  |  |  | 100
|  | 400C |  |  | 17 |  |  |  | 27 | [9 empty] | 8 | [7 empty] | 28 |  |  |  | 8 |  |  | 13 | 101
|  | 500C |  |  | 31 |  |  |  | 18 | [17 empty] | 51 | [7 empty] | 100
|  | Glycolated |  |  | 25 |  |  |  | 46 | [9 empty] | 9 | [7 empty] | 10 |  |  |  | 9 |  |  |  | 99
11 | 11 Etosha 22 Msub 10-15 cm S1 | Regular |  |  | 10 |  |  |  | 13 |  |  |  | 2 | 14 |  |  | 16 | [9 empty] | 29 |  | 2 |  | 13 |  | 2 |  | 101
|  | 400C |  |  | 38 |  |  |  | TR |  |  |  | 5 | [11 empty] | 9 |  | 27 |  |  |  | 21 |  |  |  | 100
|  | 500C |  |  | 29 |  |  |  | 10 |  |  |  | 6 | [13 empty] | 54 | [7 empty] | 99
|  | Glycolated |  |  | 15 |  |  |  | 13 |  |  |  | 3 | 16 |  |  | 9 |  |  | TR | [6 empty] | 37 |  |  |  | 6 |  |  |  | 99
12 | 12 Etosha 22 CRT 0-0.1 cm S2 | Regular |  |  | 4 |  |  |  | TR |  |  |  |  | 17 | [10 empty] | 13 |  | 66 | [7 empty] | 100
|  | 400C |  |  | 27 |  |  |  | TR |  |  |  |  | 10 | [10 empty] | 12 |  | 50 | [7 empty] | 99
|  | 500C |  |  | TR |  |  |  | TR | [15 empty] | 10 |  | 90 | [7 empty] | 100
|  | Glycolated |  |  | 23 |  |  |  | TR |  |  |  |  | 21 | [5 empty] | TR |  |  |  |  | 16 |  | 40 | [7 empty] | 100
13 | 13 Etosha 22 sub 0.1-10 cm S2 | Regular |  |  | 4 |  |  | 12 | 25 |  |  |  |  | 7 | [6 empty] | 9 |  |  |  | 12 |  | 26 |  |  |  | 4 |  |  |  | 99
|  | 400C |  |  | 19 |  |  | 20 | 26 | [17 empty] | 29 |  | 5 |  | 1 |  |  |  | 100
|  | 500C |  |  | 12 |  |  | 25 | 14 | [11 empty] | 16 | [5 empty] | 28 |  | 4 | [5 empty] | 99
|  | Glycolated |  |  | 6 |  |  | 12 | 24 |  |  |  |  | 3 | [10 empty] | 11 |  | 39 |  | 5 | [5 empty] | 100
14 | 14 Etosha 22 crt 0-0.1cm S3 | Regular |  |  | 5 |  |  |  | 47 |  |  |  |  | 11 |  |  |  |  | 15 | [5 empty] | 4 |  | 11 |  |  |  | 7 |  | 1 |  | 101
|  | 400C |  |  | 17 |  |  |  | 18 |  |  |  |  | 4 |  |  |  |  | 5 | [5 empty] | 8 |  | 42 |  |  |  | 6 |  |  |  | 100
|  | 500C | 1 |  | 7 |  |  |  | 31 | [9 empty] | 35 | [5 empty] | 5 |  | 22 | [7 empty] | 101
|  | Glycolated |  |  | 5 |  |  |  | 49 |  |  |  |  | 4 |  |  |  |  | 7 | [5 empty] | 4 |  | 29 |  |  |  | 2 |  |  |  | 100
15 | 15 Etosha 22 Sub 0.1-3 cm S3 | Regular |  |  | 22 |  |  |  | 12 |  |  |  |  | 7 |  |  |  |  | 9 |  |  | 40 |  |  | 4 | [5 empty] | 6 |  |  |  | 100
|  | 400C |  |  | 29 |  |  |  | TR |  |  |  |  | 1 |  |  |  |  | 1 |  |  | 60 |  |  | 7 | [5 empty] | 2 |  |  |  | 100
|  | 500C | 4 |  | 3 |  |  |  | 21 | [9 empty] | 11 |  |  | 45 |  |  | 9 |  | 6 | [7 empty] | 99
|  | Glycolated |  |  | 6 |  |  |  | 30 |  |  |  |  | 3 | [5 empty] | TR |  | 55 |  |  | 4 | [5 empty] | 1 |  |  |  | 99
16 | 16 Etosha 22 moisub 3-15 cm S3 | Regular |  |  | 4 |  |  |  | 31 | [9 empty] | 5 | [5 empty] | 6 |  | 28 |  | 5 | 19 | 2 |  |  |  | 100
|  | 400C | 7 |  | 23 |  |  |  | 30 | [9 empty] | 11 | [5 empty] | 5 |  | 7 |  | 5 |  | 11 |  |  |  | 99
|  | 500C | 1 |  | 10 |  |  |  | 14 | [9 empty] | TR |  | 30 | [5 empty] | 27 |  | 12 | 6 |  |  |  |  | 100
|  | Glycolated |  |  | 11 |  |  | 8 | 17 | 1 | [8 empty] | 1 |  | 11 |  |  |  | 3 |  | 11 |  | 5 | 31 |  |  |  |  | 99
17 | 17 Etosha 22 wetsub 15cm S3 | Regular |  |  | 23 | 4 |  |  | TR | [7 empty] | 17 | [7 empty] | 5 |  | 38 |  | TR |  | 13 |  |  |  | 100
|  | 400C |  |  | 37 |  |  |  | TR | [15 empty] | 11 |  | 50 |  | 1 |  | 1 |  |  |  | 100
|  | 500C | 6 |  | 14 |  |  |  | 25 | [15 empty] | 12 |  | 21 |  | 3 |  | 19 |  |  |  | 100
|  | Glycolated |  |  | 20 | 5 |  |  | 47 | [15 empty] | 11 |  | 11 |  | 3 |  | 3 |  |  |  | 100
18 | 18 Etosha 22 SSCrt 0-0.5 cm S4 | Regular |  |  | 6 |  |  | 7 | 41 | 1 |  |  |  | 14 |  |  |  |  | 15 | [7 empty] | 8 |  | 1 |  | 6 |  | 2 |  | 101
|  | 400C |  |  | 41 |  |  | 28 | TR | 1 |  |  |  | 2 |  |  |  |  | 3 | [7 empty] | 21 |  | 1 |  | 3 |  | TR |  | 100
|  | 500C | 3 |  | 23 |  |  | 14 | 22 | 1 | [16 empty] | 35 |  | 2 | [5 empty] | 100
|  | Glycolated |  |  | 10 |  |  | 14 | 35 | 2 | [8 empty] | 10 | [7 empty] | 25 |  | 4 | [5 empty] | 100
19 | 19 Etosha 22 Sssub 0.5-5cm S4 | Regular |  |  | 8 |  |  | 7 | 41 | 1 |  |  |  | 12 | [10 empty] | 7 |  | 9 |  |  |  | 11 | 3 |  |  | 99
|  | 400C |  |  | 17 |  |  | 14 | 37 | 1 |  |  |  | 8 | [10 empty] | 14 | [5 empty] | 9 |  |  |  | 100
|  | 500C |  |  | 25 |  |  | 18 | 42 | [15 empty] | 12 |  |  |  | 3 | [5 empty] | 100
|  | Glycolated |  |  | 19 |  |  | 21 | 37 |  |  |  |  | 13 | [13 empty] | 6 |  |  | 4 |  |  |  | 100
20 | 20 Etosha 22 HS crt 0-0.1 cm S4 | Regular |  |  | 7 |  |  | 11 | 43 | 2 |  |  |  | 18 | [14 empty] | 5 |  | 11 |  | 4 |  | 101
|  | 400C |  |  | 13 |  |  | 15 | 43 | 2 |  |  |  | 11 | [14 empty] | 6 |  | 10 |  | TR |  | 100
|  | 500C |  |  | 14 |  |  | 14 | 34 | [17 empty] | 34 |  | 4 | [5 empty] | 100
|  | Glycolated |  |  | 7 |  |  | 12 | 34 | 2 |  |  |  | 8 |  |  |  |  | 11 | [7 empty] | 23 |  | 3 | [5 empty] | 100
21 | 21 Etosha 22 Hssub 0.1-2.5 cm S4 | Regular |  |  | 14 |  |  | 6 | 28 | 2 |  |  |  | 5 |  |  |  | 11 | [6 empty] | 3 |  | 20 |  | 11 | [5 empty] | 100
|  | 400C |  |  | 10 |  |  | 7 | 28 | 1 |  |  |  | 3 |  |  |  | 12 | [6 empty] | 7 |  | 22 |  | 10 | [5 empty] | 100
|  | 500C |  |  | 14 |  |  | 7 | 34 | [8 empty] | 11 | [6 empty] | 6 |  | 16 |  | 12 | [5 empty] | 100
|  | Glycolated |  |  | 10 |  |  | 9 | 37 | 2 |  |  |  | 3 |  |  |  | 15 | [6 empty] | 8 |  | 11 |  | 6 | [5 empty] | 101
22 | 22 Etosha 22 MoiSub 2.5-15 cm | Regular |  |  | 9 |  |  | 15 | 31 |  | 11 | 14 |  | 4 | [13 empty] | 10 | 7 | [5 empty] | 101
|  | 400C |  |  | 18 | TR |  | 14 | 31 |  | 8 | [15 empty] | 13 | 9 | 7 | [5 empty] | 100
|  | 500C |  |  | 17 |  |  | 10 | 30 | [17 empty] | 37 |  | 6 | [5 empty] | 100
|  | Glycolated |  |  | 10 |  |  | 16 | 46 |  |  |  |  | 2 | [10 empty] | 11 |  |  | 9 | 7 | [5 empty] | 101
23 | 23 Etosha 23 CRT 0-0.1 cm S5 | Regular |  |  | 7 |  |  | 6 | 38 |  |  |  |  | 12 |  |  |  | 9 | [9 empty] | 6 | 14 |  | 5 |  | 2 |  | 99
|  | 400C |  |  | 5 |  |  | 9 | 40 |  |  |  |  | 8 |  |  |  | 7 | [6 empty] | 9 |  |  | 6 | 11 |  | 5 |  |  |  | 100
|  | 500C |  |  | 6 |  |  | 14 | 54 | [15 empty] | 15 |  |  | 2 | 10 |  | TR |  |  |  | 101
|  | Glycolated |  |  | 8 | 7 |  | 6 | 36 |  |  |  |  | 6 | [10 empty] | 6 |  | 17 | 6 | 9 | [5 empty] | 101
24 | 24 Etosha 22 sub 0.1-5 cm S5 | Regular |  |  | 6 |  |  | 16 | 23 |  |  |  |  | 28 | [12 empty] | 20 |  | 3 |  |  |  | 4 |  | 100
|  | 400C |  |  | 9 |  |  | 18 | 22 |  |  |  |  | 7 | [12 empty] | 28 |  | 3 |  |  |  | 13 |  | 100
|  | 500C |  |  | 10 |  |  | 37 | 15 | [6 empty] | 10 | [8 empty] | 28 | [9 empty] | 100
|  | Glycolated |  |  | 14 |  |  | 16 | 18 |  |  |  |  | 10 |  | 5 | [10 empty] | 35 | [5 empty] | 2 |  | 100
25 | 25 Etosha 22 sub 5-15 cm S5 | Regular |  |  | 8 |  |  | 7 | 24 |  |  |  |  | 7 | [6 empty] | 13 |  |  |  |  | 3 | 21 | 5 | 4 |  | 5 |  | 2 |  | 99
|  | 400C |  |  | 13 |  |  | 18 | 15 |  |  |  |  | 5 | [6 empty] | 17 | [6 empty] | 5 | 4 |  | 8 |  |  | 15 | 100
|  | 500C |  |  | TR |  |  | 26 | TR | [5 empty] | TR | [5 empty] | 14 |  |  |  | 9 |  | 41 |  |  | 10 |  |  |  |  | 100
|  | Glycolated |  |  | 10 |  |  |  | 18 |  |  |  |  | 2 |  |  | 37 | [6 empty] | 10 |  |  | 18 | 4 | [6 empty] | 99
26 | 26 CS2 ground | Regular |  |  | TR |  |  | 3 | 55 | [7 empty] | 16 | [7 empty] | 2 |  |  | 4 | 11 |  | 9 |  |  |  | 100
|  | 400C |  |  | 4 |  |  | 7 | 19 | [7 empty] | 25 | [7 empty] | 10 |  | 5 | 6 | 13 |  | 11 |  |  |  | 100
|  | 500C |  |  | 6 |  |  | 4 | 33 | [15 empty] | 5 |  | 11 | 1 |  | 39 |  |  |  |  | 99
|  | Glycolated |  |  | 1 |  |  |  | 20 | [7 empty] | 17 | [9 empty] | 9 | 4 | [6 empty] | 51
27 | 27 SS2_63 um | Regular |  |  | 4 | 2 |  | 9 | 61 |  |  |  |  | 5 |  |  |  |  | 5 | [5 empty] | 7 |  |  | 6 |  |  | 2 |  |  |  | 101
|  | 400C |  |  | 21 |  |  | 8 | 49 |  |  |  |  | 4 |  |  |  |  | 5 | [5 empty] | 4 |  |  | 6 |  |  | 4 |  |  |  | 101
|  | 500C |  |  | 29 |  |  | 8 | 51 | [9 empty] | TR | [5 empty] | 5 | [5 empty] | 6 |  |  |  | 99
|  | Glycolated |  |  | 25 |  |  | 9 | 34 | [9 empty] | 4 | [5 empty] | 4 | [5 empty] | 23 |  |  |  | 99
28 | 28 SS2 Ground | Regular |  |  | 13 |  |  | 5 | 32 |  |  |  |  | 2 |  |  |  | 10 | [6 empty] | 6 |  | 13 | 5 | 12 |  | 1 |  |  |  | 99
|  | 400C |  |  | 13 |  |  | 10 | 12 |  |  |  |  | 2 |  |  |  | 10 | [6 empty] | 3 |  | 25 | 7 | 15 |  | 3 |  |  |  | 100
|  | 500C |  |  | 17 |  |  | 11 | 18 | [8 empty] | 8 | [8 empty] | 29 | 2 | 15 | [5 empty] | 100
|  | Glycolated |  |  | 12 |  |  | 7 | 33 |  |  |  |  | 2 | [10 empty] | 4 |  | 21 | 8 | 12 | [5 empty] | 99
29 | 29 RS1_63 um | Regular |  |  | 6 |  |  | 5 | 36 |  |  |  |  | 3 | [10 empty] | 1 |  | 41 |  | 6 |  | 1 |  |  |  | 99
|  | 400C |  |  | 17 |  |  | 12 | 46 |  |  |  |  | 3 | [10 empty] | 9 |  |  |  | 9 |  | 4 |  |  |  | 100
|  | 500C |  |  | 19 |  |  | 10 | 66 | [15 empty] | 4 | [9 empty] | 99
|  | Glycolated |  |  | 15 |  |  | 11 | 49 |  |  |  |  | 5 | [10 empty] | 6 |  | 11 |  | 2 | [5 empty] | 99
30 | 30 RS1 ground | Regular |  |  | 4 |  |  | 4 | 27 | [7 empty] | 22 | 10 | 1 | [5 empty] | 3 |  | 15 | 5 | 10 | [5 empty] | 101
|  | 400C |  |  | 10 |  |  | 7 | 38 | [8 empty] | 10 | 1 | [7 empty] | 15 | 7 | 12 | [5 empty] | 100
|  | 500C |  |  | 7 |  |  | 7 | 32 | [8 empty] | 7 | 10 | [5 empty] | 4 |  | 21 | 3 | 9 | [5 empty] | 100
|  | Glycolated |  |  | 15 |  |  | 3 | 49 | [8 empty] | 21 | 4 | [7 empty] | 4 |  | 4 | [5 empty] | 100
1 | 1R Etosha 18 RIP 0-2 cm | Regular |  |  | 7 |  |  |  | 20 |  |  |  |  | 16 |  |  |  |  | 15 |  |  | 27 |  |  | 4 | [5 empty] | 8 |  | 2 |  | 99
|  | 400C |  |  | 8 |  |  |  | 18 |  |  |  |  | 6 |  |  |  |  | 7 |  |  | 51 |  |  | 7 | [5 empty] | 3 |  |  |  | 100
|  | 500C |  |  | 31 |  |  |  | 48 | [15 empty] | 21 | [5 empty] | TR |  |  |  | 100
|  | Glycolated |  |  | 9 |  |  |  | 56 |  |  |  |  | 8 |  |  |  |  | 16 | [5 empty] | 7 | [5 empty] | 3 |  |  |  | 99
REP | 13R Etosha 22 sub 0.1-10 cm S2 | Regular |  |  | 4 |  |  | 7 | 21 |  |  |  |  | 3 | [6 empty] | 21 |  |  |  | 6 |  | 24 |  | 4 | 10 |  |  |  |  | 100
|  | 400C |  |  | 6 |  |  | 19 | 18 |  |  |  |  | 1 | [6 empty] | 21 |  |  |  | 7 |  | 16 |  | 5 | 7 |  |  |  |  | 100
|  | 500C |  |  | 7 |  |  | 16 | 19 | [11 empty] | 17 |  |  |  | 10 |  | 21 |  | 5 | 4 |  |  |  |  | 99
|  | Glycolated |  |  | 11 |  |  | 13 | 30 |  |  |  |  | 3 | [10 empty] | 7 |  | 33 |  | 4 | [5 empty] | 101
REP | 25 Etosha 22 sub 5-15 cm S5 REPLICATE | Regular |  |  | 16 |  |  | 8 | 15 |  |  |  |  | 5 |  |  |  |  | 8 | TR | 14 |  |  |  |  | 3 | 10 |  | 3 | 10 | 4 |  | 3 |  | 99
|  | 400C |  |  | 13 |  |  |  | 19 |  |  |  |  | 3 |  |  |  |  | 3 |  | 20 | [5 empty] | 25 |  |  | 15 |  |  |  |  | 98
|  | 500C |  |  | 2 |  |  | 14 | 19 | [9 empty] | 15 |  | 17 | [5 empty] | 20 | 4 |  | 10 |  |  |  |  | 101
|  | Glycolated |  |  | 5 |  |  | 13 | 16 | [9 empty] | 3 |  | 25 |  |  |  | 4 |  | 11 | 6 | 4 | 13 |  |  |  |  | 100
0

## Anchor Scan Data
Sample ID |  | 1 Etosha 18 RIP 0-2 cm |  |  |  | 2 Etosha 18 SS 0-2 cm S1 |  |  |  | 3 Etosha 18 SS 0-2 cm S2 |  |  |  | 4 Etosha 18 SS 0-2 cm S3 |  |  |  | 5 Etosha 18 HS 0-2 cm S1 |  |  |  | 6 Etosha 18 Ssub 2-5 cm S1 |  |  |  | 7 Etosha 22 RIP 0-1 cm S1 |  |  |  | 8 Etosha 22 crt 0-0.1 cm S1 |  |  |  | 9 Etosha 22 sub 0.1-4cm S1 |  |  |  | 10 Etosha 22 CrySub 4-10 cm S1 |  |  |  | 11 Etosha 22 MoiSub 10-15 cm S1 |  |  |  | 12 Etosha 22 crt 0-0.1 cm S2 |  |  |  | 13 Etosha 22 sub 0.1-10 cm S2 |  |  |  | 14 Etosha 22 CRT 0-0.1 cm S3 |  |  |  | 15 Etosha 22 Sub 0.1 - 3 cm S3 |  |  |  | 16 Etosha 22 MoiSub 3-15 cm S3 |  |  |  | 17 Etosha 22 wetsub 15 cm S3 |  |  |  | 18 Etosha 22 SS Crt 0-0.5 cm S4 |  |  |  | 19 Etosha 22 Ssub 0.5-5 cm |  |  |  | 20 Etosha 22 HS crt 0-0.1 cm S4 |  |  |  | 21 Etosha 22 HS sub 0.1-2.5 cm S4 |  |  |  | 22 Etosha 22 msub 2.5-15 cm S4 |  |  |  | 23 Etosha 22 CRT 0-0.1 cm S5 |  |  |  | 24 Etosha 22 sub 0.1-5cm S5 |  |  |  | 25 Etosha 22 sub 5-15 cm S5 |  |  |  | 26 CS2 ground |  |  |  | 27 |  |  |  | 28 |  |  |  | 29 |  |  |  | 30 |  |  |  | 1R Etosha 18 RIP 0-2 cm REPLICATE |  |  |  | 13R Etosha 22 sub 0.1-10 cm S2 REPLICATE |  |  |  | 25R Etosha 22 sub 5-15 cm S5 REPLICATE
|  | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated | Regular | 400C | 500C | Glycolated
No. | Pos. [°2θ] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts] | Intensity [Counts]
1 | 2.0054 | 1650 | 3957 | 8805 | 4663 | 4428 | 5922 | 9379 | 5172 | 8127 | 5623 | 4419 | 7829 | 4737 | 4332 | 5072 | 4947 | 4901 | 6978 | 7129 | 7761 | 7827 | 6904 | 7057 | 6865 | 5037 | 5963 | 6996 | 5425 | 5778 | 6419 | 7702 | 5920 | 7516 | 5381 | 5288 | 5087 | 4797 | 5331 | 5667 | 4484 | 6262 | 5458 | 6101 | 6196 | 4679 | 4817 | 4811 | 4942 | 9083 | 5201 | 8698 | 7882 | 9691 | 5632 | 9654 | 9357 | 7140 | 4970 | 9066 | 8970 | 8755 | 5487 | 8625 | 8601 | 3700 | 5501 | 8911 | 8686 | 9437 | 5998 | 9352 | 8620 | 8471 | 8451 | 8549 | 8190 | 8742 | 8983 | 9004 | 8522 | 8324 | 8788 | 8750 | 8853 | 8180 | 8245 | 8802 | 8410 | 8473 | 8695 | 8717 | 8447 | 7941 | 8394 | 7758 | 7857 | 6795 | 6150 | 6322 | 4631 | 8100 | 6441 | 7624 | 8100 | 8203 | 5543 | 8767 | 8133 | 8403 | 6238 | 6955 | 7242 | 7062 | 6757 | 7591 | 7096 | 8378 | 8532 | 8371 | 7675 | 7887 | 7237 | 7637 | 7579 | 7720 | 7496 | 7842 | 7067 | 7864 | 4606 | 4786 | 7743
2 | 2.0164 | 1643.4794 | 3886.3363 | 8736.1932 | 4678.0905 | 4479.9553 | 5881.2581 | 9345.6401 | 5034.4503 | 7940.2748 | 5615.8692 | 4284.324 | 7742.9307 | 4682.9754 | 4379.8648 | 5033.4548 | 4770.7911 | 4904.3977 | 6767.3476 | 7082.5241 | 7551.5251 | 7963.6716 | 6863.0123 | 7024.648 | 6855.7876 | 4916.7498 | 5895.6067 | 7057.3933 | 5318.1106 | 5658.7665 | 6159.7296 | 7511.3889 | 5833.4671 | 7510.7252 | 5324.5128 | 5224.3328 | 5036.8604 | 4614.8033 | 5154.7866 | 5661.2345 | 4350.4022 | 6326.3363 | 5485.9307 | 6260.0949 | 6207.5374 | 4705.4267 | 4830.3609 | 4961.5988 | 4983.9184 | 9033.4749 | 5191.1888 | 8620.6313 | 8126.7104 | 9766.8402 | 5851.2424 | 9516.0492 | 9371.1072 | 7304.7296 | 5073.648 | 8857.4012 | 8937.5005 | 8673.1922 | 5390.7217 | 8602.2291 | 8577.2011 | 3675.9587 | 5512.5374 | 8655.5364 | 8634.2134 | 9346.1185 | 5980.8112 | 9256.3151 | 8451.8938 | 8210.619 | 8433.2134 | 8534.1106 | 8046.9174 | 8565.6883 | 8718.4882 | 8671.2906 | 8504.2301 | 8309.5285 | 8760.295 | 8681.2748 | 8734.4346 | 8291.1441 | 8284.3397 | 8638.8559 | 8319.6224 | 8444.2547 | 8607.2748 | 8772.8569 | 8404.4759 | 8028.2424 | 8082.3564 | 7749.9464 | 8059.3899 | 6570.1967 | 6060.9061 | 6495.296 | 4589.4056 | 8173.7586 | 6295.7006 | 7849.6893 | 8173.7586 | 8457.1529 | 5485.6067 | 8852.1273 | 8098.9966 | 8336.5206 | 6190.5418 | 6682.8569 | 7006.3397 | 6848.1844 | 6965.3206 | 7824.3653 | 7085 | 8346.0939 | 8580.2011 | 8236.6637 | 7660.9061 | 7887.9833 | 7203.5329 | 7708.5664 | 7423.6716 | 7732.0905 | 7317.7375 | 7738.5329 | 7353.8658 | 7761.6278 | 4542.8156 | 4750.4803 | 7613.762
3 | 2.0274 | 1681.2871 | 3832.7296 | 8593.3353 | 4605.0737 | 4395.1888 | 5900.5251 | 9157.4995 | 5144.5841 | 7885.3677 | 5521.9272 | 4228.4267 | 7657.7375 | 4677.8043 | 4364.4769 | 5065.821 | 4751.8859 | 4856.8604 | 6950.764 | 7200.4769 | 7430.763 | 8092.7886 | 6858.5428 | 6915.5329 | 6675.2222 | 4972.5015 | 5858.297 | 7086.9351 | 5325.0088 | 5641.8191 | 6141.5152 | 7618.944 | 5870.8525 | 7564.8456 | 5203.943 | 5012.3353 | 5106.0492 | 4598.3117 | 5218.2704 | 5439.7129 | 4223.2291 | 6270.8614 | 5398.4513 | 6107.5408 | 6088.7955 | 4578.4346 | 4781.7778 | 4843.7394 | 4898.0324 | 9151.9754 | 5123.2714 | 8590.1888 | 7781.2871 | 9745.1996 | 5707.0501 | 9514.8623 | 9218.4838 | 7280.2468 | 4963.5497 | 8966.2389 | 8818.057 | 8765.9764 | 5288.0246 | 8695.7394 | 8509.5093 | 3750.3206 | 5394.238 | 8774.8702 | 8567.9027 | 9433.1731 | 6041.4769 | 9359.1731 | 8365.7532 | 8184.0069 | 8371.8191 | 8538.0836 | 8198.0501 | 8615.8623 | 8597.295 | 8774.6804 | 8360.7463 | 8353.8682 | 8782.1072 | 8625.8761 | 8687.4759 | 8141.238 | 8389.3697 | 8784.2881 | 8444.3215 | 8299.7965 | 8543.3707 | 8678.9351 | 8281.705 | 7883.8535 | 8111.7129 | 7825.5575 | 7929.3363 | 6583.9597 | 5973.295 | 6277.5084 | 4620.4022 | 8071.7876 | 6349.1976 | 7663.5575 | 8071.7876 | 8149.3766 | 5453.1888 | 8787.7522 | 7939.057 | 8297.2134 | 5992.8525 | 6592.7207 | 7113.8525 | 6859.6814 | 6752.8269 | 7689.9027 | 7083.648 | 8434.3687 | 8516.0246 | 8289.8033 | 7576.2203 | 7957.7571 | 7166.0983 | 7510.6893 | 7557.9263 | 7662.3432 | 7375.4022 | 7701.5654 | 7101.0413 | 7570.6067 | 4525.296 | 4545.8859 | 7667.1829
4 | 2.0384 | 1653.3687 | 3822.59 | 8679.2153 | 4600.441 | 4363.0369 | 5844.3968 | 9214.1239 | 4893.7242 | 8017.2301 | 5403.9454 | 4167.2537 | 7695.4808 | 4555.7935 | 4192.6696 | 4903.3923 | 4670.323 | 4885.6534 | 6745.1386 | 7021.3319 | 7344.618 | 7886.4794 | 6642.0369 | 6892.5015 | 6566.9174 | 4795.2168 | 5634.3968 | 7005.6637 | 5239.9897 | 5711.4926 | 6273.2566 | 7451.4366 | 5863.0088 | 7401.8437 | 5160.3407 | 5101.2021 | 5103.3171 | 4567.9395 | 5222.8791 | 5468.1844 | 4274.0649 | 6176.708 | 5330.6209 | 6128.41 | 6040.41 | 4554.6726 | 4881.9336 | 4638.3555 | 4940.4012 | 9138.6844 | 5017.5442 | 8556.8201 | 7744.4823 | 9403.559 | 5591.9912 | 9311.1475 | 9196.6357 | 7169.3776 | 4902.2389 | 8820.1711 | 8863.3923 | 8641.4646 | 5284.059 | 8490.6504 | 8379.2965 | 3644.9218 | 5360.3864 | 8617.3791 | 8445.5531 | 9321.4425 | 5857.5383 | 9245.267 | 8472.6401 | 8337.3215 | 8440.9764 | 8343.0737 | 8085.351 | 8412.9956 | 8610.4381 | 8752.7817 | 8310.4469 | 8422.0265 | 8698.8879 | 8735.3171 | 8619.7729 | 8110.8156 | 8282.382 | 8633.1283 | 8221.6917 | 8134.6224 | 8325.3466 | 8599.8024 | 8271.5855 | 7763.0192 | 8136.3864 | 7805.087 | 7896.6578 | 6509.3053 | 5986.0324 | 6257.2625 | 4554.3083 | 7981.3451 | 6233.649 | 7638.5885 | 7981.3451 | 8200.2389 | 5421.4056 | 9003.6106 | 7971.9292 | 8257.4823 | 5991.9041 | 6656.7596 | 7107.705 | 6727.6681 | 6856.3717 | 7566.5206 | 7028.6726 | 8438.4543 | 8509.6254 | 8280.7463 | 7702.6726 | 7591.295 | 7171.4749 | 7419.885 | 7554.5206 | 7787.9941 | 7314.4705 | 7718.9484 | 7064.733 | 7557.5811 | 4416.3776 | 4463.7699 | 7361.0192
5 | 2.0494 | 1659.3127 | 3829.7896 | 8494.1337 | 4409.4897 | 4362.6234 | 5896.4789 | 9127.6431 | 4992.2783 | 7972.5241 | 5497.7571 | 4201.7699 | 7896.5477 | 4670.6264 | 4388.4277 | 4983.9371 | 4725.5428 | 4743.4435 | 6793.1337 | 7026.1652 | 7411.5251 | 7895.706 | 6637.9361 | 6779.0806 | 6610.6087 | 4823.411 | 5702.8014 | 7017.3953 | 5089.8338 | 5642.1652 | 6069.2793 | 7291.0619 | 5801.7571 | 7445.7719 | 5110.4612 | 4998.3127 | 4981.0983 | 4492.2773 | 5059.706 | 5463.7571 | 4319.3963 | 5942.6037 | 5171.4238 | 6101.0039 | 6015.3638 | 4516.5093 | 4620.0816 | 4727.4258 | 4970.0187 | 8822.998 | 5088.2773 | 8530.6214 | 7533.8486 | 9586.1032 | 5637.6578 | 9309.3461 | 9168.178 | 7121.7886 | 4804.1475 | 8892.8722 | 8937.8407 | 8528.6539 | 5154.7542 | 8574.5939 | 8583.1681 | 3666.8368 | 5273.3255 | 8771.2173 | 8456.411 | 9327.9852 | 5743.9007 | 9198.6234 | 8413.5074 | 8194.3097 | 8355.2645 | 8343.1318 | 8194.3294 | 8438.8525 | 8634.6273 | 8515.0482 | 8302.7847 | 8265.0659 | 8495.6037 | 8609.2635 | 8443.0944 | 8097.8879 | 8212.3284 | 8538.2439 | 8290.4759 | 8181.5723 | 8450.4966 | 8669.8535 | 8316.0973 | 7868.885 | 8030.8889 | 7723.297 | 7731.5408 | 6659.5457 | 6008.7719 | 6239.7207 | 4486.5565 | 7912.885 | 6122.3107 | 7445.2458 | 7912.885 | 8106.8181 | 5441.1937 | 8685.0639 | 7741.2439 | 8047.9794 | 6114.6411 | 6548.0649 | 7092.1829 | 6725.3422 | 6627.4602 | 7547.1662 | 6990.0177 | 8162.3117 | 8321.8673 | 8241.7561 | 7672.7198 | 7604.1131 | 7213.3432 | 7449.0472 | 7542.3933 | 7615.3736 | 7374.2311 | 7585.9184 | 7118.3579 | 7625.1967 | 4374.6391 | 4503.6578 | 7466.7375
6 | 2.0604 | 1592.4248 | 3676.0393 | 8355.6431 | 4533.3407 | 4337.2291 | 5690.8491 | 8905.2345 | 5008.8879 | 7983.6863 | 5319.0615 | 4227.2734 | 7515.0118 | 4467.6032 | 4286.2955 | 4832.2404 | 4625.9612 | 4688.7148 | 6652.2463 | 6950.2793 | 7373.9661 | 7825.2399 | 6539.324 | 6853.5752 | 6487.2458 | 4754.1844 | 5800.4966 | 6906.9105 | 5214.9941 | 5561.3628 | 6027.7035 | 7501.497 | 5617.6534 | 7367.0113 | 4992.8496 | 4932.4695 | 4978.2679 | 4578.503 | 4996.8604 | 5287.028 | 4152.5698 | 6103.1003 | 5307.3181 | 5871.9499 | 5828.324 | 4446.9105 | 4637.4356 | 4751.5585 | 4830.944 | 8934.5752 | 5172.4749 | 8608.6308 | 7745.9769 | 9304.9444 | 5517.7655 | 9196.0669 | 9291.0831 | 7019.2793 | 4804.296 | 8698.5978 | 8660.7429 | 8632.3461 | 5145.821 | 8361.8048 | 8329.4076 | 3628.8717 | 5410.2065 | 8525.7207 | 8390.1957 | 9225.1003 | 5788.262 | 9161.4911 | 8450.5693 | 8292.4523 | 8229.9164 | 8318.9838 | 8099.8825 | 8429.4243 | 8385.2797 | 8571.6647 | 8422.0777 | 8243.8658 | 8656.2847 | 8528.4577 | 8665.5693 | 7940.4469 | 8173.3014 | 8647.9381 | 8253.441 | 8255.7424 | 8179.4641 | 8612.5644 | 8367.3963 | 7878.6868 | 7810.4076 | 7605.5639 | 7746.179 | 6383.0728 | 5923.1898 | 6264.2734 | 4531.8825 | 7932.2738 | 6167.1844 | 7456.6372 | 7932.2738 | 8189.3682 | 5604.2507 | 8762.2281 | 7851.3068 | 8294.5138 | 6010.2011 | 6584.1898 | 6938.0501 | 6829.6863 | 6563.855 | 7411.1898 | 6909.497 | 8153.4248 | 8411.0221 | 8123.4582 | 7752.2006 | 7714.5536 | 7266.2006 | 7555.4911 | 7544.8992 | 7800.3289 | 7289.1288 | 7562.972 | 7271.7035 | 7622.6087 | 4383.497 | 4386.3466 | 7481.4916
7 | 2.0714 | 1658.0088 | 3749.8289 | 8560.9764 | 4502.3215 | 4348.8466 | 5660.0649 | 8874.7257 | 4882.4189 | 7718.3068 | 5317.3451 | 4097.0236 | 7717.3481 | 4430.2625 | 4088.6519 | 4898.9174 | 4715.0088 | 4613.0885 | 6799.1475 | 6909.177 | 7233.2714 | 7716.174 | 6539.9882 | 6780.2478 | 6491.1475 | 4754.8614 | 5568.528 | 6872.9351 | 5108.4425 | 5470.9882 | 5906.7522 | 7432.0708 | 5542.8584 | 7397.5133 | 5130.2861 | 5008.8466 | 4901.5752 | 4611.646 | 5008.6726 | 5326.3333 | 4142.2389 | 6007.9852 | 5189.1888 | 5918.6224 | 5825.0796 | 4419.056 | 4597.6755 | 4693.2478 | 4766.9351 | 8866.5575 | 5181.2832 | 8421.8732 | 7498.9056 | 9406.2802 | 5546.5782 | 9128.6991 | 9040.1829 | 6993.59 | 4754.9617 | 8718.8702 | 8623.5369 | 8545.5575 | 5075.1416 | 8476.5015 | 8247.8348 | 3670.6696 | 5325.9381 | 8560.5546 | 8422.4808 | 9134.7788 | 5608.4838 | 8997.3864 | 8264.2743 | 8217.9852 | 8291.5782 | 8490.2743 | 8031.2478 | 8325.1652 | 8505.6136 | 8530.2478 | 8235.7227 | 8182.0619 | 8554.9145 | 8406.0383 | 8500.2183 | 7926.8378 | 8066.4277 | 8481.8024 | 8126.9381 | 8037.4838 | 8320.0914 | 8660.8791 | 8243.0059 | 7753.7257 | 7742.2566 | 7495.0973 | 7801.4454 | 6422.9941 | 5864.0413 | 6135.1298 | 4465.3864 | 7945.9823 | 6161.4454 | 7541.997 | 7945.9823 | 8028.9912 | 5373.174 | 8459.6018 | 7686.3038 | 8222.5339 | 5978.6047 | 6526.59 | 6894.9912 | 6559.9558 | 6625.3156 | 7369.4454 | 6849.5841 | 8205.4395 | 8336.8437 | 8153.0354 | 7573.2537 | 7868.4336 | 7089.3923 | 7394.1888 | 7448.5634 | 7556.059 | 7393.1327 | 7528.0914 | 7130.1032 | 7532.3746 | 4325.2065 | 4469.9912 | 7480.8879
8 | 2.0824 | 1631.2793 | 3711.1185 | 8508.3736 | 4429.1932 | 4266.5713 | 5682.9887 | 9019.2468 | 4860.5826 | 7878.0393 | 5287.0973 | 4120.8038 | 7555.0831 | 4425.9902 | 4221.148 | 4909.4921 | 4688.8815 | 4669.4287 | 6489.1509 | 6726.5275 | 7198.2257 | 7657.2822 | 6671.117 | 6697.1593 | 6506.766 | 4718.0197 | 5455.0678 | 6842.0098 | 5116.237 | 5610.2906 | 5980.1185 | 7227.5261 | 5632.1834 | 7204.8068 | 5153.4936 | 4917.0211 | 4843.4189 | 4353.591 | 4983.4823 | 5415.8574 | 4096.0324 | 5947.1509 | 5179.3441 | 5933.2355 | 5903.2453 | 4471.2045 | 4752.7473 | 4615.2355 | 4740.5698 | 8863.4951 | 5006.4189 | 8255.1932 | 7641.0393 | 9335.2989 | 5382.3019 | 9030.117 | 8936.2355 | 6993.0846 | 4654.9464 | 8669.237 | 8777.8166 | 8528.2129 | 4992.8717 | 8412.2566 | 8288.8589 | 3465.3033 | 5186.3117 | 8641.4145 | 8224.3781 | 9130.4513 | 5596.6657 | 8954.5811 | 8129.0452 | 8143.557 | 8118.9562 | 8336.7193 | 7950.8264 | 8287.5826 | 8264.2738 | 8455.1608 | 8176.1495 | 8007.7743 | 8375.2468 | 8502.1819 | 8597.1578 | 7863.1185 | 7992.4808 | 8497.2974 | 7996.5177 | 8021.9661 | 8270.3977 | 8577.9774 | 8316.2876 | 7591.1298 | 7937.4046 | 7538.8181 | 7687.2891 | 6334.6559 | 5845.3751 | 6171.9435 | 4411.9125 | 7693.2709 | 6054.086 | 7327.7502 | 7693.2709 | 8034.0393 | 5285.2016 | 8625.2763 | 7669.56 | 8169.0197 | 5756.0792 | 6513.4371 | 6921.9789 | 6683.7955 | 6473.0551 | 7550.1027 | 6760.8264 | 8010.4302 | 8347.0084 | 8066.0777 | 7572.2227 | 7582.2468 | 7115.911 | 7392.3456 | 7215.8815 | 7488.1495 | 7329.1608 | 7473.8378 | 6938.5261 | 7621.6445 | 4271.0226 | 4485.4499 | 7385.6249
9 | 2.0934 | 1610.0501 | 3708.1524 | 8353.9833 | 4416.6863 | 4339.822 | 5635.1691 | 8973.0875 | 4963.4621 | 7678.9794 | 5306.529 | 3971.6716 | 7396.8161 | 4472.8053 | 4079.5064 | 4766.4916 | 4661.5624 | 4596.473 | 6517.8014 | 6761.2468 | 7182.4287 | 7794.644 | 6573.6145 | 6577.9685 | 6618.7817 | 4628.7158 | 5628.3176 | 6787.6234 | 5126.9223 | 5617.9036 | 5981.3805 | 7168.2989 | 5651.3491 | 7223.115 | 5081.1357 | 4897.8682 | 4816 | 4337.1445 | 4885.0354 | 5264.0462 | 4105.0334 | 5966.2635 | 5174.3471 | 5858.529 | 5792.8348 | 4466.4975 | 4454.2911 | 4551.7384 | 4731.2802 | 8865.8112 | 4972.5103 | 8254.8791 | 7452.5005 | 9230.6657 | 5442.061 | 8963.1947 | 8883.9479 | 6917.5546 | 4660.3786 | 8678.644 | 8722.9577 | 8410.8663 | 5028.7886 | 8329.3117 | 8229.7158 | 3591.0836 | 5163.0708 | 8361.4582 | 8316.1337 | 9111.1485 | 5540.6883 | 8955.648 | 8248.8387 | 8032.3245 | 8085.7345 | 8277.473 | 7866.2616 | 8382.6942 | 8528.238 | 8410.2507 | 8136.9105 | 8017.1799 | 8352.4897 | 8430.8427 | 8384.8073 | 7865.1672 | 7828.9961 | 8288.0983 | 8086.3117 | 7939.3727 | 8259.3786 | 8495.0226 | 8148.8889 | 7626.6421 | 7713.1131 | 7568.4661 | 7622.3913 | 6322.6421 | 5784.0462 | 6000.6922 | 4338.0944 | 7803.6588 | 6046.6421 | 7204.1091 | 7803.6588 | 7839.7974 | 5145.0108 | 8453.8142 | 7727.5379 | 8070.3746 | 5947.355 | 6304.2429 | 6937.1003 | 6668.8368 | 6537.1839 | 7280.4081 | 6677.7365 | 8008.9036 | 8231.9105 | 8180.8112 | 7396.1632 | 7558.1131 | 6939.8869 | 7452.1898 | 7216.4877 | 7435.6362 | 7303.5231 | 7421.8515 | 6890.3117 | 7578.9204 | 4335.295 | 4379.7719 | 7445.1839
10 | 2.1044 | 1579.0177 | 3576.0929 | 8337.4469 | 4453.7743 | 4272.1195 | 5485.6858 | 8815.3451 | 4863.9425 | 7794.5708 | 5195.6106 | 3984.1239 | 7489.4735 | 4411.7965 | 4126.5708 | 4771.5531 | 4523.5398 | 4655.2655 | 6539.6947 | 6731.354 | 7119.8584 | 7644.3053 | 6362.0575 | 6545.5354 | 6313.9602 | 4765.6283 | 5458.4381 | 6771.1283 | 5005.3407 | 5547.9513 | 5892.6283 | 7077.1504 | 5535.2522 | 7214.7788 | 4964.8407 | 4893.5398 | 4805.7124 | 4398.7655 | 4812.885 | 5349.2124 | 4084.7876 | 5936.2035 | 5111.0398 | 5724.3805 | 5818.2655 | 4322.9248 | 4559.3673 | 4532.7699 | 4698.177 | 8642.7832 | 4913.3319 | 8396.2522 | 7627.8451 | 9393.7434 | 5538.9203 | 9017.7876 | 8958.6991 | 6985.6018 | 4620.6549 | 8546.5575 | 8530.5531 | 8458.1106 | 4977.0177 | 8315.2566 | 8382.7788 | 3545.115 | 5024.9646 | 8383.6681 | 8244.5531 | 9073.385 | 5535.1637 | 8731.0354 | 8163.0487 | 8097.7965 | 8158.9602 | 8334.6062 | 7879.146 | 8235.3938 | 8290.6726 | 8281.1947 | 8310.0531 | 8047.469 | 8403.2611 | 8250.0929 | 8292.1372 | 7763.8274 | 7931.6991 | 8285.4602 | 8064.9557 | 8028.7035 | 8217.6637 | 8303.4115 | 8030.4071 | 7579.3142 | 7759.6681 | 7362.3717 | 7657.3274 | 6240.5796 | 5854.5 | 5918.3496 | 4432.4071 | 7726.8496 | 5974.7566 | 7351.0442 | 7726.8496 | 7977.5575 | 5283.4115 | 8591.9602 | 7787.0133 | 8103.4867 | 5676.3186 | 6385.5708 | 6890.4115 | 6632.7389 | 6446.7389 | 7315.8407 | 6725.385 | 7954.6593 | 8401.5133 | 7948.1593 | 7409.1416 | 7577.677 | 6921.9823 | 7213.9956 | 7329.2301 | 7530.4159 | 7310.3761 | 7423.7743 | 6885.5487 | 7510.5708 | 4335.3451 | 4335.8009 | 7352.1947
11 | 2.1154 | 1550.4415 | 3597.8151 | 8149.822 | 4249.2901 | 4253.1898 | 5462.0619 | 8793.1396 | 4767.4916 | 7759.1337 | 5287.2901 | 3965.6706 | 7432.944 | 4379.4749 | 4098.2625 | 4777.2291 | 4545.4857 | 4579.3353 | 6297.8328 | 6670.7699 | 7185.5703 | 7500.5143 | 6438.4739 | 6583.1062 | 6396.2016 | 4638.3019 | 5474.9164 | 6720.0885 | 4982.0442 | 5426.2734 | 5970.4808 | 7212.8997 | 5669.0944 | 7209.6814 | 5039.9548 | 4727.883 | 4732.4641 | 4250.4818 | 4808.8151 | 5356.3068 | 4071.056 | 5970.8604 | 5203.7414 | 5645.1455 | 5751.1849 | 4435.4012 | 4503.1622 | 4535.8151 | 4776.9272 | 8630.6647 | 4894.3677 | 8143.3353 | 7308.0728 | 9234.4739 | 5281.1229 | 9007.6205 | 8908.8496 | 6900.648 | 4631.7984 | 8531.4464 | 8533.5251 | 8286.6382 | 4940.5752 | 8239.7257 | 8363.4238 | 3537.6313 | 5076.2173 | 8362.0718 | 8241.2566 | 8965.9882 | 5448.9725 | 8871.9331 | 8073.648 | 7875.414 | 8039.8328 | 8240.296 | 7808.6254 | 8307.4572 | 8279.6313 | 8317.234 | 8005.5025 | 8174.5015 | 8326.2399 | 8296.3972 | 8439.4857 | 7680.2684 | 7879.6372 | 8261.3019 | 7955.6765 | 7764.1957 | 8191.7807 | 8178.7158 | 7924.1632 | 7766.4503 | 7753.5585 | 7296.5644 | 7468.1062 | 6211.5034 | 5750.6814 | 6066.8928 | 4305.0226 | 7782.0216 | 6007.8378 | 7264.9774 | 7782.0216 | 7761.6922 | 5138.9272 | 8452.061 | 7556.8378 | 7842.1013 | 5885.4405 | 6368.939 | 6914.6018 | 6504.704 | 6505.4464 | 7213.5536 | 6810.8761 | 7968.1229 | 8074.0334 | 7869.1121 | 7335.3176 | 7399.5438 | 7127.5909 | 7413.7817 | 7321.1455 | 7561.5969 | 7021.3068 | 7241.9056 | 6909.9223 | 7459.8928 | 4300.7257 | 4304.6932 | 7403.4749
12 | 2.1264 | 1575.8112 | 3525.4361 | 8265.7783 | 4191.3569 | 4222.0792 | 5540.5944 | 8801.1962 | 4756.6686 | 7700.1608 | 5181.2252 | 3985.615 | 7409.6785 | 4379.7247 | 4090.972 | 4780.8623 | 4499.1352 | 4585.4543 | 6375.3009 | 6547.4533 | 7232.2021 | 7548.352 | 6279.4263 | 6593.1632 | 6475.0841 | 4641.6893 | 5503.3751 | 6586.1121 | 4921.6809 | 5362.7832 | 6003.6893 | 7252.5664 | 5566.0487 | 7172.3446 | 4950.2144 | 4770.5103 | 4792.4287 | 4450.2886 | 4730.2168 | 5282.2984 | 4097.8368 | 5952.2581 | 4955.6809 | 5688.1888 | 5833.1632 | 4273.4897 | 4524.4518 | 4458.4336 | 4739.7222 | 8601.4263 | 4764.7419 | 8151.7527 | 7350.9562 | 9087.5944 | 5277.8088 | 9034.8245 | 8945.5005 | 6869.7733 | 4549.2655 | 8452.6504 | 8515.4287 | 8454.766 | 4913.8673 | 8145.2704 | 8240.8392 | 3527.9695 | 4963.6381 | 8199.4189 | 8181.4031 | 8845.9902 | 5502.9951 | 8869.0388 | 8038.3117 | 7968.8673 | 8075.118 | 8214.1047 | 7752.2144 | 8172.3265 | 8264.5615 | 8217.3776 | 7974.9843 | 7943.0438 | 8270.2935 | 8376.0462 | 8370.9621 | 7721.0792 | 7894.2812 | 8263.0669 | 8017.7733 | 7787.5408 | 7995.4995 | 8287.7247 | 8092.1765 | 7424.7321 | 7713.1814 | 7315.7807 | 7477.6224 | 6266.3274 | 5722.8088 | 5906.4617 | 4332.4518 | 7670.028 | 5959.7783 | 7224.9744 | 7670.028 | 7705.0659 | 5117.1352 | 8384.4105 | 7536.3594 | 7949.6991 | 5752.5944 | 6431.1814 | 6659.296 | 6514.9415 | 6414.2119 | 7286.8063 | 6655.8392 | 7966.4312 | 8082.4435 | 7932.2729 | 7223.3619 | 7657.7379 | 6910.5029 | 7330.2886 | 7355.9415 | 7424.5177 | 6956.8977 | 7313.855 | 6975.9366 | 7309.737 | 4222.2242 | 4378.1608 | 7397.5054
13 | 2.1374 | 1569.4956 | 3567.1003 | 8193.41 | 4248.8083 | 4199.8024 | 5507.469 | 8858.9616 | 4664.2625 | 7670.3481 | 4951.1622 | 3898.6726 | 7424.4661 | 4332.4366 | 4070.3304 | 4730.7434 | 4494.2478 | 4512.5428 | 6414.0354 | 6776.0796 | 7066.8791 | 7546.823 | 6221.7139 | 6561.1298 | 6473.1445 | 4473.7847 | 5453.3304 | 6628.118 | 4991.1799 | 5329.0413 | 5826.1976 | 7201.0118 | 5508.5369 | 7073.2006 | 4928.6549 | 4755.8437 | 4732.4956 | 4321.8083 | 4734.9027 | 5277.236 | 4085.6224 | 5782.0531 | 5014.2655 | 5679.3127 | 5822.5162 | 4291.8437 | 4439.3068 | 4479.8997 | 4636.1003 | 8553.1534 | 4895.7139 | 8094.5752 | 7459.1888 | 9079.1003 | 5247.1652 | 8860.41 | 8807.5634 | 6721.3156 | 4498.6519 | 8452.2537 | 8464.4926 | 8238.1917 | 4940.7227 | 8126.056 | 8253.1003 | 3481.9676 | 4833.1947 | 8037.4543 | 8123.3864 | 8750.1534 | 5444.7522 | 8688.1091 | 7868.9852 | 7982.0059 | 7880.7611 | 8187.2566 | 7756.1711 | 8160.177 | 8217.4484 | 8194.6106 | 8111.233 | 7821.3894 | 8220.115 | 8262.3481 | 8247.5516 | 7676.9734 | 7681.0177 | 8092.4543 | 7872.2655 | 7783.0354 | 8132.8466 | 8205.4808 | 7930.0973 | 7454.41 | 7605.0678 | 7284.4071 | 7462.4838 | 5973.7316 | 5684.7906 | 5933.3776 | 4237.5723 | 7672.233 | 5877.823 | 7224.826 | 7672.233 | 7844.9882 | 5097.7935 | 8444.3628 | 7641.295 | 7894.4454 | 5741.8879 | 6312.5959 | 6643.9174 | 6467.9145 | 6343.6932 | 7225.6401 | 6645.2389 | 7947.6401 | 8208.5959 | 7960.1003 | 7353.0678 | 7437.1888 | 6794.6814 | 7232.4867 | 7272.0413 | 7451.5664 | 7037.2773 | 7199.3068 | 6865.5044 | 7392.4956 | 4300.1888 | 4349.3805 | 7306.3068
14 | 2.1484 | 1510.9892 | 3527.4076 | 8202.3033 | 4157.325 | 4201.7793 | 5352.411 | 8654.738 | 4674.8599 | 7605.237 | 4997.9371 | 3886.0467 | 7308.9479 | 4431.9695 | 4024.8815 | 4625.8869 | 4452.6912 | 4477.0054 | 6337.4813 | 6528.6391 | 7135.7345 | 7456.0433 | 6313.3736 | 6299.1957 | 6235.2424 | 4522.7468 | 5421.9425 | 6574.5152 | 4998.4794 | 5210.797 | 5834.9263 | 7005.3033 | 5565.3412 | 7016.824 | 4932.4956 | 4883.677 | 4664.8815 | 4241.0305 | 4727.8653 | 5161.0826 | 4075.6662 | 5710.1239 | 4954.4848 | 5661.675 | 5693.7827 | 4199.384 | 4502.9857 | 4490.1922 | 4594.1706 | 8722.2227 | 4733.5261 | 8056.2173 | 7314.6337 | 9041.6445 | 5220.3663 | 8859.7163 | 8944.3019 | 6648.4794 | 4557.6912 | 8321.7109 | 8440.8923 | 8288.7704 | 4897.7847 | 8014.4272 | 8192.1563 | 3454.2694 | 4892.9676 | 8124.8348 | 7965.3501 | 8920.9803 | 5456.4292 | 8771.9803 | 7833.2891 | 7879.7576 | 8008.0128 | 8247.5334 | 7887.0433 | 7962.4361 | 8374.9607 | 8123.1598 | 8022.1563 | 7893.4794 | 8211.4848 | 8212.8958 | 8324.7129 | 7599.4361 | 7659.2783 | 8147.8781 | 7850.1937 | 7724.8564 | 7941.8869 | 8109.2011 | 7955.3501 | 7428.9498 | 7567.7163 | 7325.6499 | 7315.7989 | 6080.2424 | 5636.7773 | 5900.1922 | 4279.1795 | 7599.9336 | 5834.6588 | 7112.5619 | 7599.9336 | 7750.738 | 4993.1436 | 8398.6499 | 7615.677 | 7845.1598 | 5682.8564 | 6235.648 | 6650.5261 | 6526.0467 | 6272.1849 | 7188.352 | 6498.174 | 7909.7129 | 8090.5836 | 7914.4956 | 7305.1706 | 7531.3019 | 6919.6804 | 7281.2552 | 7103.9444 | 7313.1327 | 7051.9336 | 7270.176 | 6914.1042 | 7312.5782 | 4163.9425 | 4282.088 | 7354.9661
15 | 2.1594 | 1553 | 3493.1819 | 8224.1534 | 4132.4936 | 4117.708 | 5443.8348 | 8583.6234 | 4757.1298 | 7616.4838 | 5082.0098 | 3862.409 | 7327.9518 | 4398.7856 | 4031.2173 | 4593.1495 | 4426.7237 | 4463.3992 | 6195.7561 | 6518.7502 | 6917.117 | 7306.412 | 6124.6165 | 6326.6686 | 6207.0718 | 4550.7699 | 5464.0167 | 6556.3481 | 4969.3382 | 5348.4031 | 5936.7473 | 7057.7827 | 5483 | 7055.5978 | 4866.3215 | 4735.8702 | 4629.4031 | 4276.6234 | 4755.2311 | 5131.8997 | 3978.646 | 5715.1327 | 4932.2144 | 5664.1367 | 5760.8279 | 4393.4258 | 4410.5713 | 4371.3668 | 4561.9479 | 8584.7866 | 4725.4867 | 7990.2045 | 7308.1986 | 9048.2625 | 5181.7178 | 8871.4189 | 8728.9508 | 6634.1367 | 4512.1072 | 8349.4287 | 8376.1367 | 8122.4248 | 4796.4936 | 8043.6952 | 8056.3343 | 3345.5064 | 5007.2989 | 8151.883 | 7994.88 | 8832.0197 | 5361.412 | 8710.3796 | 7806.7758 | 7904.3147 | 7783.0393 | 8004.8967 | 7690.0914 | 8105.6401 | 8166.2379 | 8133.12 | 7862.588 | 7913.0354 | 8168.0492 | 8266.0462 | 8217.821 | 7775.8505 | 7708.3304 | 8045.5998 | 8003.8446 | 7677.3471 | 7921.1947 | 8107.237 | 7971.6303 | 7320.0551 | 7564.9646 | 7279.7729 | 7322.409 | 6042.0914 | 5683.8741 | 5756.7601 | 4327.1662 | 7559.3894 | 5883.2733 | 7108.2694 | 7559.3894 | 7707.1554 | 5043.9223 | 8349.1593 | 7463.7404 | 7848.7532 | 5624.3343 | 6350.708 | 6612.5457 | 6504.3019 | 6313.4484 | 7087.058 | 6597.1072 | 7808.3274 | 8015.175 | 7849.0098 | 7220.8092 | 7288.0747 | 6904.7502 | 7051.5133 | 7029.3048 | 7406.6106 | 7013.4543 | 7209.5487 | 6797.6303 | 7259.0354 | 4203.9518 | 4240.8476 | 7182.0226
16 | 2.1704 | 1558.531 | 3488.6209 | 8006.9381 | 4113.0059 | 4064.8879 | 5330.2006 | 8585.1121 | 4737.9144 | 7584.5015 | 5030.1608 | 3881.0059 | 7207.5649 | 4402.0664 | 4053.3112 | 4590.9956 | 4497.6195 | 4464.3392 | 6298.5369 | 6452.0501 | 6893.2183 | 7386.8879 | 6222.6873 | 6364.91 | 6137.9395 | 4439.3687 | 5309.8702 | 6397.3304 | 4812.0221 | 5194.5693 | 5816.0959 | 7006.3746 | 5468.4351 | 6939.3127 | 4824.4351 | 4709.3628 | 4504.8835 | 4250.2729 | 4581.3953 | 5043.503 | 4002.6593 | 5816.2404 | 4909.5752 | 5575.6254 | 5728.6578 | 4282.9499 | 4436.6755 | 4393.5192 | 4592.6475 | 8358.3245 | 4779.5354 | 7979.9602 | 7170.1401 | 9137.7139 | 5133.2802 | 8840.0103 | 8803.2773 | 6580.1018 | 4502.2625 | 8309.497 | 8301.0841 | 8238.3732 | 4772.028 | 7991.9071 | 8082 | 3376.6091 | 4927.4012 | 8147.7581 | 7942.3348 | 8758.0044 | 5425.4793 | 8726.7065 | 7881.6327 | 7941.1268 | 7699.3525 | 7849.4587 | 7558.0059 | 7987.2743 | 7977.8628 | 8155.2448 | 7775.8437 | 7953.149 | 8080.8171 | 8069.9218 | 8221.9454 | 7721.4041 | 7874.5737 | 8256.8879 | 7866.9381 | 7720.0487 | 7935.6976 | 8113.0221 | 7861.6032 | 7369.3024 | 7523.5605 | 7239.1003 | 7309.295 | 5900.972 | 5495.7714 | 5692.4145 | 4207.9233 | 7581.9425 | 5741.4646 | 7138.1962 | 7581.9425 | 7810.6917 | 4990.9558 | 8275.4292 | 7411.4631 | 7828.9985 | 5640.2286 | 6270.9705 | 6551.6652 | 6496.4233 | 6235.8215 | 7153.1991 | 6573.2507 | 7889.0929 | 8065.7139 | 7820.7596 | 7053.0074 | 7275.1062 | 6831.5973 | 6996.9115 | 7125.7094 | 7342.2065 | 7047.0487 | 7228.7994 | 6690.7375 | 7307.4454 | 4078.4027 | 4176.3289 | 7222.615
17 | 2.1814 | 1565.1042 | 3561.7571 | 7969.9961 | 4147.3432 | 4020.1504 | 5302.1888 | 8586.7817 | 4558.117 | 7453.6067 | 4938.2684 | 3924.7827 | 7239.7483 | 4309.4513 | 3910.0757 | 4664.1121 | 4433.6067 | 4354.6578 | 6332.7611 | 6427.0039 | 6937.9331 | 7349.2301 | 6221.6942 | 6319.7827 | 6227.3304 | 4440.2134 | 5177.0669 | 6468.297 | 4790.9833 | 5139.2802 | 5843.0531 | 7001.8319 | 5382.9577 | 6925.9499 | 4758.2301 | 4662.1042 | 4589.703 | 4178.3392 | 4490.8378 | 5059.1967 | 3984.4346 | 5819.8574 | 4852.0452 | 5525.1042 | 5571.3176 | 4267.8781 | 4384.5103 | 4343.2173 | 4551.9626 | 8398.8869 | 4663.0413 | 7916.0796 | 7201.4307 | 8990.058 | 5206.0875 | 8771.5359 | 8603.3599 | 6663.4218 | 4493.297 | 8272.5644 | 8311.4808 | 8122.4435 | 4791.4228 | 8165.5015 | 8050.7316 | 3393.4602 | 4845.4189 | 8008.3687 | 7916.1927 | 8636.5103 | 5314.6401 | 8446.6991 | 7973.8859 | 7770.7237 | 7760.6863 | 7940.6618 | 7606.1091 | 8042.2635 | 8173.0826 | 8052.117 | 7813.1052 | 7788.4523 | 8198.468 | 8046.0747 | 7853.2439 | 7361.8997 | 7714.703 | 8181.6686 | 7789.8781 | 7570.8328 | 7869.7699 | 8091 | 7847.8535 | 7435.8201 | 7672.7148 | 7214.5733 | 7193.117 | 5887.0836 | 5536.4435 | 5785.7738 | 4325.355 | 7435.3058 | 5827.2802 | 7178.7778 | 7435.3058 | 7710.8201 | 5022.8073 | 8194.356 | 7348.1337 | 7713.4818 | 5567.3048 | 6104.0039 | 6584.8869 | 6371.9666 | 6276.8161 | 6983.5152 | 6517.4395 | 7695.0462 | 7958.766 | 7791.5772 | 7227.6401 | 7268.3972 | 6774.0708 | 7087.3599 | 7121.9666 | 7309.9872 | 6912.7788 | 7257.4012 | 6712.6824 | 7155.0669 | 4090.6981 | 4117.8161 | 7228.179
18 | 2.1924 | 1508.5069 | 3521.0747 | 8038.0143 | 4153.2689 | 3996.119 | 5350.6436 | 8443.8658 | 4514.5379 | 7461.6116 | 5038.001 | 3940.0162 | 7128.2679 | 4197.2389 | 3977.3437 | 4508.2532 | 4410.8963 | 4494.3446 | 6228.6573 | 6382.5969 | 6962.149 | 7326.12 | 6159.7758 | 6338.5683 | 6174.6873 | 4424.881 | 5179.5964 | 6405.5088 | 4789.5674 | 5102.8663 | 5599.1937 | 6861.2832 | 5323.5364 | 6942.8058 | 4716.1057 | 4609.6858 | 4648.3122 | 4234.9405 | 4674 | 5021.3446 | 3893.3727 | 5751.1047 | 4727.7153 | 5454.7163 | 5510.88 | 4146.5226 | 4330.8368 | 4282.6421 | 4593.3289 | 8379.0447 | 4581.4469 | 7923.8963 | 7140.5674 | 8877.0747 | 5129.2689 | 8798.4341 | 8579.1504 | 6526.119 | 4454.9557 | 8182.2837 | 8253.1937 | 8090.2694 | 4812.2694 | 8036.6421 | 7898.1352 | 3346.6131 | 4923.8363 | 8120.149 | 7944.2694 | 8624.6264 | 5310.6416 | 8480.7625 | 7776.2532 | 7758.7606 | 7787.5821 | 8017.1362 | 7644.6578 | 8008.8215 | 7978.3427 | 8017.3137 | 7944.5369 | 7835.6431 | 8078.06 | 7955.4022 | 8124.9985 | 7422.9838 | 7744.2684 | 7994.2547 | 7695.6858 | 7627.5069 | 7855.6873 | 8065.9262 | 7829.9852 | 7381.2704 | 7564.2532 | 7216.6121 | 7188.1932 | 5861.9852 | 5520.5674 | 5706.3894 | 4195.7611 | 7478.6278 | 5798.3884 | 7060.001 | 7478.6278 | 7686.3589 | 4955.4484 | 8228.1352 | 7310.1495 | 7837.5821 | 5485.7001 | 6038.0752 | 6559.4031 | 6389.4027 | 6285.061 | 7087.6721 | 6525.2847 | 7714.6573 | 7975.7178 | 7746.9105 | 7222.3879 | 7216.6121 | 6748.4912 | 7079.6573 | 7123.3137 | 7266.0447 | 7001.2542 | 7162.0447 | 6763.5084 | 7145.4774 | 3998.3722 | 4151.7616 | 7074.4317
19 | 2.2034 | 1554.6283 | 3488.7611 | 8047.3894 | 4123.115 | 4019.2124 | 5175.2655 | 8401.2478 | 4477.0177 | 7464.8142 | 4930.8938 | 3799.5575 | 7180.4425 | 4225.6991 | 3941.646 | 4579.0531 | 4337.1947 | 4343.7965 | 6144.4159 | 6383.1858 | 6968.885 | 7270.8673 | 6171.9823 | 6194.7434 | 6106.6106 | 4398.0266 | 5256.8584 | 6303.5664 | 4774.3097 | 5048.6195 | 5640.6372 | 6834.7522 | 5370.9026 | 6959.8938 | 4580.9027 | 4686.6726 | 4723.5929 | 4196.292 | 4641.4779 | 4894.469 | 3918.115 | 5688.938 | 4822.7611 | 5470.531 | 5511.2655 | 4133.7257 | 4229.5487 | 4243.6991 | 4492.6283 | 8383.6637 | 4678.646 | 7870.1062 | 7151.1947 | 8858.2566 | 5102.4336 | 8655.7965 | 8491.4867 | 6589.2566 | 4387.1593 | 8151.4248 | 8268.0885 | 8023.4867 | 4688.1858 | 7909.2035 | 7827.2301 | 3262.2743 | 4828.8672 | 8085.9557 | 7838.3274 | 8701.0177 | 5294.6106 | 8345.9558 | 7810.9911 | 7796.5752 | 7768.7434 | 7849.0797 | 7541.4867 | 7922.7434 | 8035.5929 | 8015.8938 | 7976.5133 | 7725.2124 | 8052.646 | 8044 | 8205.2389 | 7538 | 7721.5398 | 7911.8053 | 7789.0443 | 7652.6106 | 7744.9115 | 7955.9557 | 7814.0973 | 7221.0973 | 7600.5398 | 7186.7257 | 7262.2832 | 5850.5221 | 5476.7699 | 5579.0443 | 4221.9469 | 7396.0177 | 5722.8407 | 6941.6106 | 7396.0177 | 7603.5929 | 4903.1593 | 8142.1858 | 7277.823 | 7763.4336 | 5610.3363 | 6018.6903 | 6478.7257 | 6423.4513 | 6153.4248 | 7046.7522 | 6394.8407 | 7652.7611 | 7810.4071 | 7751.6106 | 7221.5221 | 7205.7522 | 6857.5133 | 7003.823 | 7085.3894 | 7225.0885 | 6910.7434 | 7177.9469 | 6596.2035 | 7140.1327 | 4065.7168 | 4097.7345 | 7170.6903
20 | 2.2144 | 1501.6706 | 3415.6347 | 7831.6706 | 4112.9941 | 3952.5447 | 5189.0664 | 8328.263 | 4560.1077 | 7326.2817 | 4918.8264 | 3776.2817 | 7142.5742 | 4320.1863 | 3933.4793 | 4601.647 | 4350.8264 | 4305.4135 | 6094.413 | 6327.97 | 6879.0959 | 7291.0359 | 6075.5811 | 6129.5629 | 6164.3112 | 4454.2576 | 5239.6652 | 6420.324 | 4776.4194 | 5082.497 | 5680.8142 | 6645.179 | 5212.9577 | 6944.557 | 4572.5688 | 4667.9882 | 4484.3717 | 4143.1018 | 4534.0059 | 4948.2876 | 3854.0895 | 5579.0477 | 4746.8323 | 5471.5511 | 5359.8496 | 4136.234 | 4245.0418 | 4243.2212 | 4396.0413 | 8400.7124 | 4678.1858 | 7921.1322 | 7187.6716 | 8831.4671 | 5093.4852 | 8585.4199 | 8570.9469 | 6606.0964 | 4334.5211 | 8080.6347 | 8184.6888 | 8056.2817 | 4544.2694 | 7648.1308 | 7852.7193 | 3404.2394 | 4717.6401 | 7885.9577 | 7792.3648 | 8766.1259 | 5198.263 | 8353.9105 | 7752.4189 | 7660.0723 | 7695.6229 | 7951.0241 | 7574.9823 | 7855.9282 | 7963.4975 | 8044.1863 | 7912.4253 | 7767.7552 | 8016.9228 | 8015.0423 | 7955.5565 | 7462.8442 | 7543.2876 | 7941.1023 | 7786.1863 | 7555.796 | 7673.5624 | 7911.8741 | 7820.7965 | 7294.0782 | 7487.8687 | 7159.8501 | 7215.4012 | 5847.3417 | 5396.6347 | 5655.8201 | 4232.2399 | 7413.2522 | 5591.3053 | 6925.7901 | 7413.2522 | 7540.1381 | 4921.9405 | 8151.8505 | 7257.88 | 7564.9813 | 5533.3112 | 6161.2876 | 6278.8859 | 6356.1145 | 6183.4789 | 7053.6888 | 6318.3776 | 7628.6352 | 7765.5447 | 7769.144 | 7126.1381 | 7232.0718 | 6686.5093 | 6962.3953 | 7038.1976 | 7122.8137 | 6797.9636 | 7181.4675 | 6396.5084 | 7028.9282 | 3988.209 | 4068.0536 | 7014.7252
21 | 2.2254 | 1511.5919 | 3420.2625 | 7881.415 | 4083.3628 | 4007.0787 | 5074.2242 | 8434.1672 | 4552.1436 | 7286.7109 | 4888.0669 | 3721.9617 | 7269.06 | 4147.5703 | 3874.1406 | 4513.9381 | 4329.1622 | 4249.3864 | 6136.6421 | 6368.0403 | 6826.3727 | 7289.4464 | 5973.2389 | 6120.6539 | 6149.2488 | 4392.4513 | 5112.8496 | 6285.7109 | 4718.8043 | 5070.5084 | 5691.5064 | 6813.9046 | 5300.3795 | 6848.0029 | 4602.6323 | 4577.998 | 4398.3363 | 4158.1121 | 4563.6372 | 4905.8997 | 3912.5182 | 5628.3628 | 4722.8427 | 5424.766 | 5533.826 | 4039.236 | 4281.6873 | 4308.0551 | 4562.6539 | 8423.9262 | 4580.9646 | 7777.413 | 6997.6391 | 8769.0865 | 4984.3776 | 8436.1858 | 8397.1858 | 6473.5536 | 4278.3963 | 8119.6775 | 8128.9548 | 7993.6037 | 4570.9784 | 7863.8309 | 7672.9164 | 3413.7424 | 4849.4179 | 8005.5821 | 7906.1504 | 8670.1367 | 5272.7021 | 8286.3864 | 7795.764 | 7557.1622 | 7655.7375 | 7857.232 | 7515.3412 | 7855.9115 | 7823.7542 | 7858.5418 | 7852.4848 | 7614.1406 | 7905.9803 | 7918.7709 | 7996.59 | 7481.647 | 7506.3078 | 7825.8997 | 7565.881 | 7641.9764 | 7770.0383 | 7930.6755 | 7748.2488 | 7186.8761 | 7333.3058 | 7148.3038 | 7149.2035 | 5782.4484 | 5376.9262 | 5688.0315 | 4106.8594 | 7271.2478 | 5610.2793 | 6975.7021 | 7271.2478 | 7414.2507 | 4842.6155 | 8095.6372 | 7299.9213 | 7689.3156 | 5539.881 | 6101.6922 | 6356.704 | 6182.9902 | 6191.177 | 6983.4513 | 6268.6106 | 7552.4602 | 7817.1288 | 7656.1721 | 7072.768 | 7271.4199 | 6626.1858 | 6886.232 | 7002.4513 | 7266.177 | 6881.585 | 7044.9213 | 6532.943 | 7005.8043 | 4081.4661 | 4128.3176 | 6916.6677
22 | 2.2364 | 1495.5162 | 3380.8319 | 7897.177 | 4058.0959 | 4008.3171 | 5117.8053 | 8372.5103 | 4458.4558 | 7348.1224 | 4873.1608 | 3774.7743 | 7133.9292 | 4159.0929 | 3911.1239 | 4470.4499 | 4308.09 | 4302.0944 | 6097.8363 | 6434.4808 | 6765.8805 | 7221.4779 | 5912.618 | 6149.615 | 5998.6829 | 4319.7758 | 5130.9661 | 6318.472 | 4706.677 | 5062.5472 | 5557.8392 | 6790.9041 | 5255.8348 | 6671.3879 | 4655.8363 | 4431.1268 | 4462.9336 | 4149.7035 | 4558.4469 | 4920.1283 | 3819.2581 | 5558.4425 | 4666.1622 | 5467 | 5452.028 | 4126.0634 | 4291.4513 | 4237.7714 | 4523.9897 | 8321.705 | 4457.9366 | 7736.385 | 7140.6106 | 8611.9749 | 4890.2699 | 8568.1991 | 8493.8378 | 6478.0265 | 4258.1947 | 8165.441 | 8102.0664 | 8057.4263 | 4599.6091 | 7712.5398 | 7684.1283 | 3354.5133 | 4762.9012 | 8059.5767 | 7841.0575 | 8449.7714 | 5135.6755 | 8330.3186 | 7733.1991 | 7560.6106 | 7639.4159 | 7672.7729 | 7507.3481 | 7866.41 | 7790.2198 | 7882.4897 | 7754.6755 | 7666.3525 | 8024.9557 | 7908.5472 | 7864.1003 | 7435.5531 | 7500.7345 | 7785.6667 | 7450.1932 | 7576.649 | 7662.764 | 7921.7139 | 7580.6475 | 7137.326 | 7240.7434 | 6960.4926 | 6983.3643 | 5798.8879 | 5337.9425 | 5536.4174 | 4066.7788 | 7375.2522 | 5584.6091 | 6817.7684 | 7375.2522 | 7428.882 | 4881.5428 | 8069.5398 | 7291.972 | 7506.8083 | 5431.9366 | 6012.7965 | 6342.8024 | 6260.0295 | 6077.2596 | 6903.0324 | 6368.5133 | 7683.385 | 7840.2906 | 7589.292 | 7219.0265 | 7254.5988 | 6687.9926 | 6764.236 | 6889.7684 | 7149.6785 | 6777.2906 | 7005.4823 | 6485.5221 | 6976.8997 | 3883.6814 | 4123.1387 | 7038.6962
23 | 2.2474 | 1501.8673 | 3295.3619 | 7695.7847 | 4021.0157 | 3890.2812 | 5044.5673 | 8237.6539 | 4504.6126 | 7276.6372 | 4837.6893 | 3771.5919 | 7050.8722 | 4118.469 | 3853.4435 | 4415.6893 | 4209.3009 | 4218.7611 | 6042.1534 | 6337.2458 | 6820.8997 | 7127.7296 | 5997.1377 | 6131.6234 | 6068.8574 | 4344.4592 | 5115.2753 | 6174.9125 | 4711.6273 | 5015.3373 | 5546.531 | 6749.2969 | 5198.7856 | 6713.4739 | 4611.7807 | 4426.2035 | 4399.3884 | 4046.4287 | 4475.2655 | 4899.5359 | 3800.531 | 5430.8722 | 4675.3215 | 5411.2969 | 5355.7148 | 4089.531 | 4232.9194 | 4173.9902 | 4346.235 | 8232.235 | 4474.4592 | 7685.9646 | 7109.3265 | 8679.3835 | 5053.3677 | 8587.1288 | 8418.9705 | 6404.1426 | 4264.9135 | 7991.0767 | 8108.0157 | 8080.9046 | 4541 | 7577.9184 | 7701.3923 | 3307.8368 | 4716.6686 | 7966.9036 | 7749.5762 | 8438.4484 | 5106 | 8260.9695 | 7750.236 | 7506.883 | 7587.352 | 7631.179 | 7393.8879 | 7707.2094 | 7720.3363 | 7909.1848 | 7724.1888 | 7643.5202 | 7845.5713 | 7842.409 | 7898.9902 | 7430.5831 | 7416.3668 | 7665.591 | 7456.0865 | 7572.9961 | 7585.9322 | 7961.0413 | 7606.9951 | 7193.5152 | 7295.7188 | 7071.8525 | 7169.6677 | 5621.3815 | 5450.2448 | 5477.6077 | 4137.7296 | 7276.6991 | 5537.3569 | 6744.765 | 7276.6991 | 7549.0826 | 4789.4031 | 7977.9125 | 7328.5261 | 7520.3982 | 5449.5408 | 5882.4179 | 6326.8614 | 6207.5359 | 6078.0157 | 6923.3058 | 6321.8368 | 7618.9036 | 7822.6381 | 7617.0816 | 7062.94 | 7021.4798 | 6559.7502 | 6872.4346 | 6823.5251 | 7126.5211 | 6764.7758 | 7004.2547 | 6601.9843 | 6935.2498 | 3923.5565 | 4209.7257 | 6883.1623
24 | 2.2584 | 1500.5826 | 3351.3038 | 7841.6499 | 3953.9115 | 3826.3117 | 4885.1057 | 8191.5005 | 4425.5998 | 7306.5919 | 4764.8481 | 3719.119 | 7077.7134 | 4150.9287 | 3876.119 | 4352.177 | 4204.4056 | 4147.1465 | 6014.0713 | 6235.5098 | 6620.9946 | 7136.6116 | 5993.473 | 5992.1519 | 6041.2866 | 4328.0172 | 5129.4307 | 6292.2404 | 4713.0344 | 4979.2734 | 5510.9617 | 6579.7768 | 5218.0462 | 6752.0093 | 4601.502 | 4483.7979 | 4298.0846 | 4042.3963 | 4497.3407 | 4853.9287 | 3779.0964 | 5464.498 | 4658.0211 | 5260.7173 | 5330.2562 | 4028.8019 | 4109.0502 | 4108.1268 | 4334.6406 | 8190.675 | 4481.1981 | 7659.5059 | 7081.7055 | 8607.6234 | 4947.5074 | 8420.1136 | 8248.8019 | 6431.7134 | 4238.1096 | 7974.9154 | 8063.2443 | 7853.4346 | 4506.5113 | 7628.9946 | 7764.148 | 3301.7173 | 4697.2232 | 7879.6077 | 7785.4095 | 8470.3092 | 5078.296 | 8248.2193 | 7557.1942 | 7457.7173 | 7602.5905 | 7622.7753 | 7348.2537 | 7632.6844 | 7787.4135 | 7748.8732 | 7697.1229 | 7641.2709 | 7826.7055 | 7712.0211 | 7853.7635 | 7192.4267 | 7475.5575 | 7791.7384 | 7491.0885 | 7397.3461 | 7766.292 | 7856.794 | 7584.5231 | 7170.7729 | 7315.588 | 6946.8958 | 7163.4666 | 5746.0093 | 5429.8825 | 5404.675 | 4110.144 | 7286.9656 | 5585.8097 | 6774.7345 | 7286.9656 | 7427.6578 | 4751.5575 | 7974.53 | 7210.2709 | 7489.1981 | 5437.4307 | 5906.0516 | 6403.4189 | 6150.0556 | 6009.2153 | 6950.2114 | 6325.3289 | 7547.4385 | 7727.9671 | 7620.2827 | 6884.5442 | 6986.4135 | 6549.7252 | 6765.1175 | 6879.9326 | 7043.1809 | 6714.2153 | 7031.6632 | 6588.2379 | 6950.3633 | 3886.4636 | 4039.8481 | 7005.9921
25 | 2.2694 | 1460.4602 | 3333.7257 | 7781.1121 | 3954.8643 | 3884.4602 | 4873.8643 | 8290.6844 | 4407.9322 | 7199.9233 | 4749.3009 | 3593.2065 | 7046.0649 | 4087.3156 | 3811.2684 | 4353.8112 | 4158.5192 | 4268.4307 | 5993.9085 | 6303.8407 | 6782.1888 | 7117.0472 | 5970.8289 | 5993.5398 | 6021.6077 | 4302.6844 | 5087.9204 | 6244.2979 | 4612.8112 | 5040.6372 | 5513.587 | 6595.3068 | 5197.9321 | 6709.3569 | 4611.8053 | 4420.8555 | 4326.3776 | 4066.8909 | 4527.3746 | 4745.1829 | 3807.2625 | 5462.1593 | 4632.056 | 5284.4838 | 5409.5782 | 4020.9794 | 4130.1917 | 4086.5546 | 4383.1268 | 8152.2124 | 4506.4248 | 7691.5959 | 6967.8407 | 8726.8702 | 4814.3156 | 8423.5369 | 8195.8466 | 6384.1357 | 4271.0738 | 7996.1799 | 8040.6283 | 7789.1681 | 4411.3304 | 7772.1947 | 7706.8761 | 3285.3658 | 4739.8053 | 7849.7611 | 7738.5959 | 8329.59 | 5020.3894 | 8201.1681 | 7525.531 | 7503.9027 | 7580.6726 | 7706.8319 | 7295.6696 | 7611.4513 | 7746.7522 | 7792.8643 | 7628.6903 | 7529.4307 | 7748.5339 | 7745.9381 | 7863.3156 | 7163.6431 | 7410.705 | 7733.4926 | 7463.292 | 7417.8083 | 7655.0354 | 7774.7198 | 7590.3953 | 7082.8053 | 7147.3097 | 6841.233 | 6905.4189 | 5609.8466 | 5281.6637 | 5403.3805 | 4043.7699 | 7397.5133 | 5566.6932 | 6759.0089 | 7397.5133 | 7440.1917 | 4730.3628 | 7696.9941 | 7122.6195 | 7522.0944 | 5384.4159 | 5856.0118 | 6250.826 | 6104.9882 | 5977.2507 | 6887.6106 | 6342.2153 | 7496.4041 | 7663.4985 | 7606.2507 | 6900.4395 | 6982.9204 | 6605.5605 | 6834.5752 | 6817.7522 | 7013.2094 | 6770.4513 | 6947.6047 | 6404.5546 | 6888.8555 | 3933.3156 | 4040.5251 | 6938.8201
26 | 2.2804 | 1446.1617 | 3339.795 | 7677.9272 | 3959.116 | 3850.1888 | 4899.9326 | 8018.0865 | 4481.9302 | 7257.0324 | 4693.1726 | 3577.617 | 7039.0162 | 4064.3746 | 3863.5688 | 4291.3909 | 4187.8599 | 4201.2778 | 5833.5256 | 6234.9705 | 6799.7468 | 7035.353 | 5950.4828 | 5993.3963 | 5896.1509 | 4304.4233 | 5021.5285 | 6083.7198 | 4547.6254 | 4933.3127 | 5401.9435 | 6749.6146 | 5123.884 | 6787.7847 | 4588.089 | 4396.2345 | 4391.6067 | 4021.6981 | 4317.3692 | 4638.6495 | 3692.7601 | 5497.5364 | 4578.3127 | 5325.178 | 5240.7925 | 4055.0728 | 4063.5929 | 4111.3019 | 4325.1131 | 7988.2212 | 4497.3343 | 7631.8707 | 6902.0433 | 8683.8142 | 4803.6657 | 8316.4852 | 8138.3584 | 6307.674 | 4317.1726 | 8040.4823 | 7936.0806 | 7674.3368 | 4386.8545 | 7687.4425 | 7608.5231 | 3254.4528 | 4698.8815 | 7796.5418 | 7699.7655 | 8299.9651 | 5014.9705 | 8043.3181 | 7436.9031 | 7519.5152 | 7416.5177 | 7633.9784 | 7291.2616 | 7717.8628 | 7684.6332 | 7798.6524 | 7606.4744 | 7489.6254 | 7773.5742 | 7806.1214 | 7772.1593 | 7240.6362 | 7384.1888 | 7653.1696 | 7383.2075 | 7442.6175 | 7565.0192 | 7730.383 | 7539.205 | 7047.1052 | 7189.4985 | 6940.8383 | 6951.0108 | 5528.3909 | 5149.7842 | 5367.2453 | 4060.2911 | 7404.1027 | 5388.1106 | 6776.5123 | 7404.1027 | 7321.6632 | 4782.9921 | 7744.9164 | 7161.8274 | 7520.2724 | 5312.9597 | 5952.7684 | 6194.2453 | 6170.4125 | 5922.6416 | 6815.5959 | 6319.5472 | 7374.8973 | 7734.6146 | 7573.7655 | 6919.0054 | 6934.6578 | 6547.0487 | 6834.7547 | 6696.0216 | 6839.4095 | 6909.7414 | 6908.1485 | 6450.5039 | 6859.3181 | 3891.9381 | 3978.0029 | 6916.9489
27 | 2.2914 | 1429.8043 | 3382.4975 | 7671 | 3886.0108 | 3776.6608 | 4930.9764 | 7926.4091 | 4357.9784 | 7162.0954 | 4665.8348 | 3675.8368 | 6950.0039 | 4042.3323 | 3792.2281 | 4289.6283 | 4180.0197 | 4129.3913 | 5860.7522 | 6159.5152 | 6719.8653 | 7116.4415 | 5826.6264 | 5999.2458 | 5871.2802 | 4233.9282 | 4910.6391 | 6053.9784 | 4507.6716 | 4945.4523 | 5380.7827 | 6713.6106 | 5165.352 | 6662.1219 | 4529.9715 | 4331.5192 | 4268.0934 | 4001.1347 | 4316.1563 | 4685.7129 | 3710.237 | 5470.9607 | 4533.8889 | 5310.5654 | 5234.2763 | 4058.1652 | 4054.3432 | 4121.9587 | 4371.2694 | 7926.8112 | 4476.0089 | 7585.5674 | 6882.4091 | 8609.3196 | 4868.7719 | 8264.5369 | 8201.4936 | 6240.998 | 4247.0128 | 8077.117 | 7962.7689 | 7769.6588 | 4415.1829 | 7470.9892 | 7519.7414 | 3258.1347 | 4639.0413 | 7774.0629 | 7659.9823 | 8278.8938 | 4951.0197 | 8001.0737 | 7378.1111 | 7324.3697 | 7398.7375 | 7653.3825 | 7299.12 | 7577.6588 | 7739.2193 | 7700.2497 | 7549.528 | 7466.9282 | 7737.3284 | 7723.6912 | 7650.0197 | 7230.2871 | 7314.176 | 7711.5388 | 7364.7974 | 7343.2763 | 7443.3481 | 7645.9046 | 7438.3412 | 6980.1868 | 7263.9892 | 6907.3432 | 6918.2999 | 5517.3609 | 5214.0128 | 5354.1849 | 4097.2675 | 7237.4739 | 5389.8417 | 6719.8024 | 7237.4739 | 7310.0679 | 4658.0934 | 7818.6352 | 7142.1888 | 7447.8043 | 5247.8456 | 5828.237 | 6201.3176 | 6095.8781 | 5848.9607 | 6707.2173 | 6254.1888 | 7497.1062 | 7653.8712 | 7544.528 | 6864.941 | 6886.3697 | 6468.237 | 6837.589 | 6670.2262 | 6926 | 6820.0108 | 6809.7719 | 6508 | 6802.3736 | 3839.8368 | 3895.1721 | 6835.6519
28 | 2.3024 | 1437.9336 | 3274.4115 | 7571.1106 | 3878.292 | 3733.7124 | 5012.4159 | 8040.6549 | 4371.7566 | 7108.1858 | 4655.4159 | 3658.4867 | 6862.8009 | 3966.3053 | 3668.6858 | 4304.6814 | 4155.2522 | 4164.2389 | 5896.8938 | 6089.0133 | 6685.5885 | 7059.1902 | 5857.5708 | 6010.5221 | 5830.5664 | 4194.4513 | 4890.4292 | 6071.0752 | 4473.9292 | 4992.9823 | 5415.3009 | 6600.8142 | 5124.292 | 6695.4513 | 4450.6593 | 4257.7743 | 4230.4071 | 3959.4159 | 4268.9558 | 4653.9115 | 3718.9823 | 5421.2699 | 4424.8097 | 5317.8761 | 5243.2655 | 4038.7168 | 4024.6903 | 4086.8319 | 4282.1903 | 8037.9867 | 4496.4602 | 7549.7743 | 6975.0841 | 8589.7566 | 4868.531 | 8229.3717 | 8221.4292 | 6299.3673 | 4119.1991 | 7928.7876 | 7821.3938 | 7795.0354 | 4475.1327 | 7490.9867 | 7544.354 | 3225.3761 | 4627.708 | 7670.3009 | 7583.5708 | 8316.2965 | 4961.0973 | 7990.8097 | 7443.7168 | 7458.2832 | 7552.0354 | 7629.615 | 7394.469 | 7608.677 | 7681.177 | 7632.6195 | 7487.2743 | 7340.8363 | 7781.3584 | 7657.2168 | 7675.031 | 7222.9292 | 7302.9381 | 7613.6903 | 7284.1504 | 7370.8496 | 7532.3009 | 7696.1637 | 7411.6372 | 6864.6593 | 7245.4911 | 6779.4602 | 6860.1195 | 5562.4602 | 5112.0841 | 5396.5841 | 4058.9248 | 7211.4911 | 5475.8319 | 6795.5929 | 7211.4911 | 7385.6416 | 4655.9159 | 7884.9646 | 7066.5221 | 7430.3805 | 5204.1549 | 5780.8982 | 6252.0708 | 6055.9115 | 5780.354 | 6682.9867 | 6137.7035 | 7348.8628 | 7503.8894 | 7457.7168 | 6804.177 | 6972.1637 | 6462.3805 | 6710.6947 | 6732.1549 | 6864.938 | 6766.8274 | 6889.8407 | 6501.6947 | 6743.5531 | 3827.7965 | 3973.4602 | 6754.6637
29 | 2.3134 | 1443.5654 | 3241.409 | 7453.2842 | 3864.9941 | 3702.3314 | 4962.2665 | 8072.4995 | 4318.6106 | 7138.5202 | 4636.5841 | 3685.9685 | 6943.708 | 3993.7984 | 3582.9921 | 4251.5398 | 4059.9282 | 4191.8574 | 5923.3186 | 5943.0128 | 6588.5841 | 6822.6883 | 5682.8908 | 5907.7925 | 5869.4612 | 4129.6431 | 4883.824 | 6061.4346 | 4487.4159 | 4936.644 | 5343.6185 | 6503.3707 | 5104.7994 | 6648.5388 | 4467.4218 | 4311.4749 | 4230.3825 | 3905.1367 | 4323.0708 | 4661.3058 | 3697.7473 | 5444.8378 | 4409.5575 | 5194.0914 | 5137.7984 | 4006.708 | 4142.2861 | 3963.8191 | 4154.0138 | 8012.6824 | 4405.6175 | 7546.7719 | 6890.5398 | 8586.8564 | 4758.0069 | 8227.4867 | 8174.9027 | 6247.176 | 4036.5713 | 7836.1436 | 7874.4494 | 7758.0069 | 4382.3058 | 7483.5782 | 7539.1298 | 3225.6165 | 4628.2724 | 7581.1485 | 7540.2743 | 8248.4209 | 4883.2733 | 8064.2665 | 7416.8702 | 7544.2419 | 7493.6696 | 7554.2291 | 7260.0403 | 7546.527 | 7710.7021 | 7618.2085 | 7494.3835 | 7215.2124 | 7782.7335 | 7550.0521 | 7666.8505 | 7132.9086 | 7281.0983 | 7642.9213 | 7369.7532 | 7284.2144 | 7504.3255 | 7632.3245 | 7342.7483 | 6736.2724 | 7155.0718 | 6715.4612 | 6850.3953 | 5626.059 | 5108.7591 | 5343.4887 | 3938.9027 | 7140.6883 | 5398.9351 | 6766.3648 | 7140.6883 | 7187.8387 | 4703.292 | 7749.6559 | 7042.8653 | 7396.2665 | 5128.4936 | 5704.5644 | 6172.1632 | 6064.6824 | 5777.5388 | 6690.7207 | 6138.5516 | 7330.8948 | 7530.942 | 7367.5467 | 6824.7542 | 6989.4621 | 6457.5192 | 6582.4474 | 6816.4415 | 6747.1622 | 6677.1367 | 6960.2793 | 6425.6765 | 6829.7591 | 3897.2861 | 3893.9744 | 6753.9351
30 | 2.3244 | 1415.5349 | 3225.3264 | 7585.9798 | 3824.3171 | 3714.4361 | 4939.7822 | 8067.3953 | 4292.1485 | 7030.8609 | 4673.0693 | 3561.97 | 6868.6131 | 4015.5841 | 3682.2965 | 4173.6932 | 4089.9405 | 4117.732 | 5841.0093 | 5996.0797 | 6591.5639 | 6854.2974 | 5599.1986 | 5875.7129 | 5787.1376 | 4103.3663 | 4915.5044 | 6077.7026 | 4436.6731 | 4808.6534 | 5174.7227 | 6457.119 | 5048.0398 | 6619.2965 | 4445.9498 | 4277.0998 | 4270.1185 | 3895.0993 | 4361.9798 | 4597.5742 | 3659.4656 | 5330.2065 | 4495.089 | 5163.5944 | 5200.2183 | 3987.9307 | 4052.3063 | 3912.2778 | 4136.0207 | 7953.7129 | 4370.8717 | 7520.2955 | 6751.029 | 8575.9493 | 4703.0295 | 8229.1976 | 8141.5541 | 6182.703 | 4065.2576 | 7780.7527 | 7751.4061 | 7677.0093 | 4345.6637 | 7452.6534 | 7497.0983 | 3221.6529 | 4649.6731 | 7648.1878 | 7414.2085 | 8302.2861 | 4860.3471 | 7976.91 | 7388.3264 | 7337.4351 | 7378.4253 | 7445.1298 | 7134.9012 | 7444.4853 | 7634.4951 | 7579.2281 | 7430.6234 | 7338.7119 | 7718.2463 | 7517.6731 | 7668.0197 | 7178.0197 | 7169.9602 | 7647.7119 | 7365.7325 | 7280.9208 | 7441.5349 | 7619.9203 | 7279.031 | 6781.1981 | 7088.7026 | 6706.9208 | 6960.6637 | 5559.7325 | 5188.9012 | 5234.09 | 3942.2675 | 7097.4449 | 5392.1485 | 6627.0393 | 7097.4449 | 7103.3766 | 4683.5737 | 7766.1091 | 6856.9602 | 7317.4646 | 5188.9317 | 5751.1583 | 6128.763 | 5973.2768 | 5768.6136 | 6658.3559 | 6097.8604 | 7408.6131 | 7470.9602 | 7273.5747 | 6721.5147 | 6854.2571 | 6473.6731 | 6640.3761 | 6731.0383 | 6821.1593 | 6609.3559 | 6880.3653 | 6342.4263 | 6843.6032 | 3886.0797 | 3872.6736 | 6765.3269
31 | 2.3354 | 1457.1121 | 3221.9233 | 7541.9499 | 3866.531 | 3773.3687 | 4886.9528 | 7906.528 | 4228.3717 | 7004.7729 | 4629.646 | 3515.8879 | 6809.9853 | 3981.5074 | 3637.236 | 4232.4543 | 4087.1121 | 3969.2596 | 5794.2861 | 5975.4602 | 6528.944 | 6929.2802 | 5643.9705 | 5851.5575 | 5631.5015 | 4099.7935 | 4821.295 | 6006.7758 | 4428.5428 | 4821.8614 | 5136.3953 | 6449.5015 | 5001.41 | 6538.1003 | 4395.292 | 4314.5221 | 4212.6195 | 3895.9144 | 4266.6962 | 4588.3628 | 3668.8761 | 5250.1446 | 4420.7404 | 5210.2124 | 5248.7257 | 3948.0206 | 3952.5133 | 3923.7906 | 4176.4661 | 7918.4956 | 4363.6342 | 7384.3215 | 6656.0501 | 8377.6962 | 4692.5929 | 8185.2684 | 8072.0885 | 6197.7935 | 4072.9322 | 7783.7139 | 7797.0738 | 7587.1445 | 4361.3953 | 7428.2507 | 7415.1121 | 3145.8938 | 4597.6637 | 7571.233 | 7449.233 | 8154.0383 | 4925.0442 | 7880.6991 | 7317.4454 | 7261.944 | 7276.9144 | 7537.4838 | 7167.7434 | 7459.4749 | 7606.1829 | 7599.1829 | 7401.4366 | 7261.6224 | 7560.5811 | 7498.1121 | 7652.2124 | 7115.7522 | 7159.4366 | 7528.9558 | 7395.4602 | 7295.0059 | 7412.6844 | 7549.4631 | 7352.6047 | 6791.944 | 7049.8555 | 6687.8201 | 6901.9115 | 5522.351 | 5146.1534 | 5314.8142 | 3908.8525 | 7027.3245 | 5384.4189 | 6568.8761 | 7027.3245 | 7204.8348 | 4572.1209 | 7721.7787 | 6843.4867 | 7224.1622 | 5265.3776 | 5722.6195 | 6223.4808 | 5971.6785 | 5759.1298 | 6559.708 | 5984.649 | 7342.9794 | 7554.4631 | 7302.0767 | 6667.4395 | 6805.413 | 6414.6578 | 6653.236 | 6553.9381 | 6848.8731 | 6535.7817 | 6727.9882 | 6327.7728 | 6727.9823 | 3841.1239 | 3857.2714 | 6778.6195
32 | 2.3464 | 1449.6627 | 3278.8569 | 7465.4611 | 3830.6583 | 3726.7532 | 4766.6145 | 7869.2979 | 4171.4366 | 7039.8977 | 4479.9867 | 3484.854 | 6901.0978 | 3927.3781 | 3622.7021 | 4269.8992 | 4102.9051 | 3876.7183 | 5770.1298 | 5909.5841 | 6548.6716 | 6958.6657 | 5608.0044 | 5787.4248 | 5641.5826 | 4098.0934 | 4756.9007 | 5934.4203 | 4425.4218 | 4817.5634 | 5159.4425 | 6333.0831 | 4963.4454 | 6574.1037 | 4438.7591 | 4241.2935 | 4200.5737 | 3872.0556 | 4209.1475 | 4529.7532 | 3636.3343 | 5392.0393 | 4285.1839 | 5162.1809 | 5179.8963 | 3849.089 | 3985.8206 | 3876.5344 | 4113.4395 | 7889.53 | 4268.4774 | 7486.235 | 6675.3855 | 8291.1898 | 4730.5256 | 8147.6101 | 8026.5693 | 6178.1853 | 4054.5737 | 7695.8496 | 7870.1008 | 7608.0089 | 4291.473 | 7419.003 | 7383.6583 | 3161.527 | 4539.997 | 7479.5182 | 7443.2483 | 8117.3795 | 4826.8422 | 7826.413 | 7240.1868 | 7222.6131 | 7217.6205 | 7403.3722 | 7225.5797 | 7421.6116 | 7584.292 | 7518.6145 | 7338.1254 | 7340.4744 | 7530.708 | 7475.4218 | 7638.854 | 7027.4263 | 7133.0846 | 7499.5152 | 7392.6991 | 7230.4144 | 7345.2453 | 7502.0905 | 7199.0015 | 6821.9592 | 7072.8088 | 6671.1957 | 6737.6219 | 5465.2817 | 5006.3766 | 5252.6657 | 3883.2016 | 7064.4322 | 5339.2306 | 6513.854 | 7064.4322 | 7217.7473 | 4505.3825 | 7586.5226 | 6968.3038 | 7223.0934 | 5126.705 | 5685.4263 | 6182.3707 | 6022.8014 | 5798.3314 | 6513.5737 | 6012.7606 | 7365.7065 | 7602.2114 | 7260.003 | 6598.381 | 6780.2365 | 6303.8466 | 6701.1898 | 6535.381 | 6696.0949 | 6510.4744 | 6674.2802 | 6229.2935 | 6679.8073 | 3729.7739 | 3830.2498 | 6731.1386
33 | 2.3574 | 1438.2055 | 3229.409 | 7348.7325 | 3747.1672 | 3683.4317 | 4723.4602 | 7841.3068 | 4246.7198 | 6959.1209 | 4431.4366 | 3447.5605 | 6784.3736 | 3905.5398 | 3563.7709 | 4197.0079 | 4093.414 | 3943.8584 | 5678.2055 | 5957.6765 | 6555.8151 | 6841.4454 | 5643.8535 | 5715.0619 | 5758.8791 | 4061.941 | 4709.5939 | 5874.1927 | 4410.4445 | 4732.8407 | 5282.7866 | 6244.0924 | 5063.2291 | 6626.1878 | 4380.175 | 4262.409 | 4179.3087 | 3883.3963 | 4226.0079 | 4484.8584 | 3621.0079 | 5329.6195 | 4249.8122 | 5083.8997 | 5097.527 | 3795.9666 | 4043.6686 | 3961.3471 | 4183.648 | 7931.0206 | 4256.1878 | 7419.2753 | 6692.9744 | 8313.7119 | 4720.8279 | 8052.8741 | 8017.8328 | 6146.3727 | 4081.293 | 7668.3314 | 7804.5457 | 7632.9026 | 4246.8741 | 7353.7512 | 7323.3727 | 3187.6401 | 4515.5733 | 7492.9302 | 7474.0029 | 8094.3805 | 4700.5064 | 7681.1337 | 7304.3619 | 7121.2291 | 7226.0669 | 7366.5113 | 7220.9105 | 7374.0924 | 7554.2576 | 7434.9489 | 7216.2517 | 7315.9538 | 7501.7453 | 7458.8712 | 7555.9361 | 6982.6145 | 7004.3884 | 7504.6293 | 7203.1495 | 7140.6116 | 7385.0079 | 7456.4602 | 7241.4985 | 6854.468 | 6980.6509 | 6643.4503 | 6812.9508 | 5425.2802 | 4993.1721 | 5234.5605 | 3949.5349 | 7048.0157 | 5262.7738 | 6486.7866 | 7048.0157 | 7050.0619 | 4502.4808 | 7578.7532 | 7071.7561 | 7170.0285 | 5006.4553 | 5609.1495 | 6068.7945 | 5939.5939 | 5745.6018 | 6527.3137 | 6015.1593 | 7203.3913 | 7363.8535 | 7234.4759 | 6586.0206 | 6749.4808 | 6226.7325 | 6632.4631 | 6544.2596 | 6707.9459 | 6452.1288 | 6651.9587 | 6304.293 | 6709.1642 | 3765.3835 | 3912.1544 | 6617.6971
34 | 2.3684 | 1428.5723 | 3161.2729 | 7363.6711 | 3742.9513 | 3732.5664 | 4709.8879 | 7696.0856 | 4252.3776 | 6966.5044 | 4508.6077 | 3391.2227 | 6617.6726 | 3932.9027 | 3604.118 | 4174.1962 | 4013.1593 | 4011.6578 | 5697.3702 | 5899.7493 | 6470.3909 | 6655.7478 | 5631.6991 | 5716.2463 | 5760.8245 | 4075.6165 | 4717.1475 | 5834.1873 | 4415.32 | 4759.267 | 5172.8643 | 6289.441 | 4981.5103 | 6523.264 | 4293.118 | 4190.6224 | 4117.028 | 3871.2817 | 4195.0841 | 4537.2448 | 3600.2242 | 5289.2153 | 4248.4956 | 5066.6357 | 5080.3215 | 3802.6504 | 3916.0752 | 3932.9926 | 4160.0354 | 7863.1386 | 4223.028 | 7381.5251 | 6637.0973 | 8355.5664 | 4693.7625 | 8024.3496 | 8037.4617 | 6098.8584 | 4056.118 | 7668.1593 | 7631.6504 | 7540.8599 | 4181.4395 | 7254.7773 | 7335.8879 | 3129.2434 | 4553.6932 | 7558.3628 | 7355.4189 | 8119.7773 | 4739.0148 | 7705.0988 | 7315.0944 | 7167.2758 | 7196.0693 | 7363.6976 | 7149.0502 | 7357.3835 | 7360.3481 | 7412.8333 | 7192.5855 | 7196.5428 | 7536.1047 | 7512.8319 | 7458.8319 | 6970.1475 | 6994.3997 | 7588.4395 | 7171.351 | 7184.2655 | 7259.1726 | 7401.5162 | 7277.2655 | 6738.1799 | 6903.351 | 6567.7212 | 6796.3053 | 5348.236 | 4993.4675 | 5213.1047 | 3872.2021 | 7034.0295 | 5222.6637 | 6455.146 | 7034.0295 | 7075.9882 | 4431.9779 | 7492.5501 | 6963.5029 | 7047.6976 | 4932.8038 | 5573.7006 | 5999.3348 | 5887.9292 | 5725.4764 | 6445.208 | 6052.0074 | 7096.7212 | 7261.6549 | 7107.5442 | 6534.7684 | 6774.295 | 6324.0089 | 6528.6991 | 6511.9779 | 6693.6844 | 6399.7906 | 6702.9513 | 6271.4115 | 6763.9086 | 3685.9867 | 3875.5383 | 6581.9322
35 | 2.3794 | 1359.4926 | 3127.6283 | 7324.9262 | 3747.7168 | 3709.7768 | 4702.5969 | 7766.1799 | 4216.1504 | 6999.1052 | 4461.9852 | 3370.5359 | 6666.0157 | 3933.6578 | 3578.0738 | 4181.2094 | 3937.7906 | 3993.4494 | 5589.4208 | 5945.4631 | 6462.1033 | 6712.9971 | 5593.8368 | 5740.8082 | 5669.1809 | 4040.705 | 4713.0462 | 5864.175 | 4299.0157 | 4803.6726 | 4992.0305 | 6304.9705 | 4904.3009 | 6421.0747 | 4276.2675 | 4161.6578 | 4092.4189 | 3866.0639 | 4218.2085 | 4570.4169 | 3603.1052 | 5231.4916 | 4227.5959 | 5010.0305 | 5095.8201 | 3823.88 | 3852.2675 | 3878.4041 | 4173.7168 | 7785.06 | 4182.1209 | 7441.4916 | 6609.179 | 8284.7188 | 4685.0905 | 7996.8515 | 7986.7483 | 6057.0875 | 3940.644 | 7605.7453 | 7718.9076 | 7529.3432 | 4180.0875 | 7296.3589 | 7356.178 | 3055.6863 | 4507.3766 | 7478.5683 | 7381.4454 | 8042.2586 | 4756.2547 | 7721.6431 | 7163.5654 | 7176.588 | 7122.1357 | 7270.9253 | 7173.2851 | 7314.1485 | 7290.4769 | 7385.5555 | 7098.3727 | 7122.6539 | 7533.1042 | 7402.645 | 7470.2527 | 6928.4356 | 7100.2075 | 7398.6165 | 7272.7306 | 7088.5103 | 7180.4012 | 7388.3412 | 7318.1623 | 6673.8791 | 6935.7778 | 6575.3599 | 6612.6136 | 5284.4464 | 4953.2979 | 5170.4041 | 3798.6578 | 7034.6598 | 5177.3884 | 6398.2242 | 7034.6598 | 7103.7945 | 4423.2517 | 7395.5211 | 6807.0491 | 7060.6529 | 4929.9695 | 5579.883 | 5926.5978 | 5838 | 5707.4651 | 6370.9243 | 5987.4956 | 7196.7149 | 7208.0118 | 7111.9676 | 6502.3727 | 6781.1367 | 6389.6283 | 6495.0157 | 6462.7158 | 6655.2979 | 6422.0895 | 6709.3884 | 6173.8967 | 6695.3894 | 3732.4326 | 3889.059 | 6640.5693
36 | 2.3904 | 1374.9282 | 3079.4641 | 7314.0905 | 3728.5585 | 3643.853 | 4697.385 | 7772.5772 | 4157.385 | 6880.9513 | 4394.3436 | 3373.2266 | 6646.9056 | 3884.4076 | 3582.4567 | 4217.0639 | 3926.0265 | 3878.3776 | 5519.8904 | 5904.3397 | 6516.1622 | 6779.6681 | 5570.1962 | 5653.8717 | 5616.7886 | 3917.3471 | 4626.3171 | 5910.7099 | 4284.6377 | 4725.7021 | 5039.2714 | 6256.585 | 4841.4036 | 6409.2566 | 4286.1436 | 4090.468 | 4065.1849 | 3716.2075 | 4241.0113 | 4558.3776 | 3572.2793 | 5191.7173 | 4273.4302 | 5011.2414 | 5064.9022 | 3790.9361 | 3909.8943 | 3803.7473 | 4139.502 | 7769.6981 | 4149.8225 | 7349.2345 | 6633.498 | 8190.1396 | 4594.7965 | 7930.7247 | 7907 | 6038.6268 | 3868.2075 | 7615.8982 | 7682.8648 | 7483.6681 | 4239.0605 | 7224.2606 | 7274.5251 | 3058.9774 | 4400.5433 | 7387.2606 | 7378.9287 | 7848.5698 | 4678.9926 | 7677.4528 | 7221.5246 | 7011.6903 | 7084.7773 | 7300.381 | 7031.8456 | 7301.9813 | 7295.3397 | 7396.3919 | 7115.2793 | 7197.1436 | 7473.7399 | 7245.1775 | 7431.4154 | 6762.6455 | 7106.087 | 7297.0148 | 7266.4341 | 6979.3019 | 7239.0944 | 7443.468 | 7274.5477 | 6707.468 | 6928.9356 | 6564.9243 | 6578.1131 | 5260.1362 | 4921.3663 | 5136.615 | 3780.8112 | 6851.4833 | 5178.6264 | 6425.5772 | 6851.4833 | 6983.3471 | 4479.3397 | 7485.0074 | 6712.5846 | 7130.6947 | 4982.1018 | 5539.6337 | 5928.8451 | 5884.4602 | 5675.2375 | 6446.264 | 5891.5207 | 7210.8191 | 7277.7965 | 7187.0379 | 6562.7281 | 6681.9774 | 6260.589 | 6532.2601 | 6509.2301 | 6716.7168 | 6436.615 | 6618.1362 | 6153.8225 | 6673.5433 | 3749.4641 | 3855.9135 | 6512.8417
37 | 2.4014 | 1411.5044 | 3120.7611 | 7276.0885 | 3710.2832 | 3608.3717 | 4661.646 | 7746.2389 | 4147.7522 | 6801.4956 | 4354.708 | 3384.7434 | 6664.4425 | 3860.6283 | 3549.5221 | 4221.4159 | 3905.7965 | 3872.3894 | 5552.7965 | 5837.6018 | 6431.8495 | 6805.0354 | 5542.0973 | 5625.3717 | 5608.9292 | 3930.1504 | 4591.8673 | 5842.9558 | 4319.5044 | 4618.7876 | 5098.7168 | 6275.6018 | 4801.1416 | 6353.4248 | 4248.3894 | 4089.3982 | 4072.4779 | 3744.7522 | 4222.4336 | 4500.6195 | 3581.5044 | 5088.4602 | 4247.8938 | 4943.7876 | 5042.8319 | 3684.8142 | 3862.823 | 3768.9823 | 4074.7345 | 7794.9735 | 4176.4956 | 7191.9469 | 6595.6106 | 8161.5664 | 4551.7611 | 7899.4867 | 7841.0708 | 5891.9557 | 3943.2124 | 7545.9292 | 7527.3628 | 7421.4248 | 4142.3009 | 7232.6991 | 7274.6991 | 3064.2743 | 4391 | 7299.6814 | 7297.2478 | 7871.4513 | 4612.7257 | 7760.3363 | 7138.6283 | 7091.6106 | 7052.1327 | 7268.5133 | 6964.1681 | 7286.5929 | 7324.177 | 7455.7964 | 7076.2743 | 7091.2478 | 7398.3717 | 7301.6903 | 7354.7611 | 6745.5221 | 7005.469 | 7345.0354 | 7191.0088 | 6903.5841 | 7191.469 | 7353.0177 | 7122.0796 | 6597.9911 | 6944.1681 | 6589.8849 | 6568 | 5197.2655 | 4877.3186 | 5071.469 | 3826.9469 | 6810.2655 | 5127.2035 | 6378.8849 | 6810.2655 | 7029.3363 | 4451.5398 | 7507.0265 | 6800.3009 | 7030.5221 | 4941.8938 | 5499.6106 | 6002.4336 | 5865.9115 | 5682.4513 | 6411.1239 | 5860.2832 | 7112.0442 | 7221.0443 | 7079.2035 | 6531.7168 | 6542.5841 | 6204.9381 | 6546.5664 | 6502.3894 | 6654.2832 | 6425.0088 | 6548.1858 | 6162.354 | 6637.8938 | 3735.0088 | 3756.4159 | 6556.1151
38 | 2.4124 | 1400.9877 | 3137.8997 | 7268.9489 | 3661.1662 | 3586.7065 | 4564.6013 | 7584.2419 | 4162.2763 | 6813.1637 | 4334.1465 | 3426.6966 | 6602.2419 | 3834.9975 | 3565.8215 | 4094.9803 | 3947.0074 | 3877.3274 | 5528.7335 | 5704.32 | 6314.7262 | 6786.8431 | 5515.7434 | 5612.1052 | 5641.9464 | 3945.3766 | 4637.5649 | 5815.8431 | 4318.0831 | 4580.2861 | 5064.824 | 6268.7748 | 4808.0541 | 6327.8338 | 4249.4474 | 4088.9779 | 4078.4621 | 3712.8899 | 4180.3569 | 4488.0782 | 3562.3471 | 5145.3938 | 4202.7212 | 4773.1613 | 5033.7748 | 3685.7483 | 3824.7458 | 3794.4646 | 4063.1101 | 7778.973 | 4175.1883 | 7178.8043 | 6512.9071 | 8231.5969 | 4512.0536 | 7927.6406 | 7738.4449 | 5917.5674 | 3971.7089 | 7522.1003 | 7575.1785 | 7485.6652 | 4127.0344 | 7207.4911 | 7281.8087 | 3049.1662 | 4380.7532 | 7313.1151 | 7232.941 | 7845.5182 | 4626.9243 | 7753.4961 | 7034.5354 | 7094.6455 | 7029.8412 | 7142.7969 | 6983.2591 | 7270.7847 | 7379.9902 | 7297.5477 | 7065.472 | 7049.1957 | 7407.176 | 7362.7896 | 7469.675 | 6800.7188 | 6947.4449 | 7256.0831 | 7132.8972 | 6882.8584 | 7141.6308 | 7241.3422 | 7078.2468 | 6585.1298 | 6824.9803 | 6548.0098 | 6503.6357 | 5228.4449 | 4831.1416 | 4991.8289 | 3854.2788 | 6802.3348 | 5096.2104 | 6243.9951 | 6802.3348 | 7019.0708 | 4472.3717 | 7400.2394 | 6757.9902 | 7051.824 | 4840.9975 | 5438.9489 | 6015.1711 | 5734.9852 | 5621.0172 | 6279.8019 | 5843.6677 | 7085.736 | 7175.4228 | 7040.0246 | 6390.2566 | 6479.323 | 6278.1028 | 6383.0585 | 6432.7163 | 6483.7139 | 6334.1539 | 6485.8579 | 6122.2714 | 6521.0806 | 3671.8948 | 3759.3766 | 6633.2493
39 | 2.4234 | 1360.7247 | 3105.1917 | 7258.1091 | 3637.9459 | 3563.5565 | 4512.1613 | 7455.7552 | 4113.0846 | 6710.6175 | 4218.5752 | 3360.4071 | 6486.8328 | 3795.8387 | 3512.8151 | 4095.5742 | 3976.6765 | 3783.3953 | 5545.8682 | 5636.2989 | 6298.4553 | 6665.233 | 5470.4621 | 5625.5152 | 5656.8692 | 3918.1504 | 4665.3835 | 5699.8378 | 4310.6765 | 4628.88 | 5028.0905 | 6150.8564 | 4883.0187 | 6386.5506 | 4205.3776 | 4068.1613 | 4038.6175 | 3687.0177 | 4164.4671 | 4418.7011 | 3501.9164 | 5122.7856 | 4169.3176 | 4740.2812 | 4929.8682 | 3680.0187 | 3821.8928 | 3733.9951 | 4074.0364 | 7653.6834 | 4127.0128 | 7141.2635 | 6508.9341 | 8237.8997 | 4435.9459 | 7937.5329 | 7693.0482 | 5983.1436 | 3834.9892 | 7537.527 | 7607.1672 | 7458.3844 | 4104.5103 | 7047.9105 | 7099.586 | 2991.8505 | 4424.8505 | 7350.9823 | 7138.0413 | 7770.5929 | 4606.0846 | 7644.6765 | 7035.2212 | 7049.2763 | 7041.1622 | 7120.8918 | 6944.3953 | 7175.4553 | 7416.2507 | 7320.2812 | 7122.4789 | 7089.0718 | 7331.8151 | 7269.1799 | 7386.5172 | 6715.0787 | 6951.0895 | 7314.0649 | 7081.1141 | 6931.4916 | 7146.2153 | 7226.3648 | 7006.5929 | 6608.5398 | 6816.233 | 6473.9174 | 6445.4857 | 5228.353 | 4814.7355 | 5039.9518 | 3887.2153 | 6717.0541 | 5144.9823 | 6246.0118 | 6717.0541 | 6988.2153 | 4454.1681 | 7401.2026 | 6698.0364 | 6984.7316 | 4875.6342 | 5429.5103 | 5909.1327 | 5640.6175 | 5548.1032 | 6307.7237 | 5779.882 | 7042.0787 | 7181.6037 | 7078.9823 | 6343.06 | 6530.0728 | 6176.2232 | 6331.0895 | 6363.4975 | 6485.0954 | 6301.0295 | 6403.0354 | 6065.2881 | 6492.2635 | 3603.8387 | 3746.6293 | 6519.9528
40 | 2.4344 | 1335.5782 | 3087.7684 | 7163.2817 | 3654.5206 | 3536.9985 | 4560.0664 | 7509.5649 | 4113.6298 | 6715.4233 | 4251.1505 | 3348.9086 | 6529.2006 | 3803.059 | 3486.1416 | 4152.9174 | 3991.4381 | 3853.6962 | 5460.7581 | 5727.4558 | 6266.7021 | 6656.3127 | 5374.9587 | 5543.4779 | 5491.8658 | 3890.7861 | 4640.6032 | 5769.6947 | 4287.5693 | 4644.1239 | 4968.5619 | 6040.1711 | 4800.7434 | 6343.8495 | 4154.3879 | 4069.09 | 3997.4056 | 3697.0988 | 4061.4602 | 4427.2493 | 3504.885 | 5089.6962 | 4174.3643 | 4780.1327 | 4950.6947 | 3670.1254 | 3763.6445 | 3680.1416 | 4064.6372 | 7583.6032 | 4089.472 | 7148.9838 | 6492.7434 | 8078.9911 | 4474.5708 | 7871.8082 | 7618.1475 | 5962.9661 | 3756.3643 | 7538.587 | 7560.9646 | 7405.0015 | 3978.2065 | 7077.3982 | 7170.5206 | 2958.4956 | 4356.0236 | 7297.7596 | 7189.2065 | 7781.1327 | 4587.9351 | 7562.6091 | 6976.7581 | 6972.4794 | 6935.354 | 7209.1667 | 7007.0266 | 7081.3864 | 7316.2684 | 7337.1239 | 7080.5192 | 7081.6195 | 7206.3953 | 7237.2065 | 7191.6696 | 6714.8201 | 6894.2787 | 7293.1873 | 7014.7345 | 6851.9911 | 7047.6416 | 7223.4616 | 6993.6445 | 6476.1888 | 6795.3289 | 6369.6696 | 6456.7375 | 5203.7094 | 4887.8673 | 5041.5118 | 3851.9159 | 6714.3805 | 5113.3304 | 6252.8923 | 6714.3805 | 6912.6504 | 4374.09 | 7363.6578 | 6632.5531 | 6892.4794 | 4925.8348 | 5388.0516 | 5837.174 | 5610.91 | 5395.9086 | 6274.2699 | 5716.3496 | 6969.4381 | 7234.3378 | 7008.5029 | 6382.4145 | 6566.3997 | 6145.8791 | 6316.941 | 6302.0575 | 6502.9086 | 6331.8923 | 6455.7935 | 5992.4956 | 6589.1431 | 3600.5133 | 3719.4882 | 6472.2979
41 | 2.4454 | 1305.7591 | 3074.117 | 7158.6018 | 3589.5949 | 3454.826 | 4562.6421 | 7580.411 | 4097.0128 | 6763.9931 | 4248.2704 | 3334.3579 | 6561.5782 | 3827.4179 | 3552.6657 | 4089.6116 | 3990.5418 | 3932.5752 | 5434.8299 | 5718.0364 | 6216.5516 | 6604.883 | 5314.0531 | 5409.7119 | 5413.3913 | 3931.5418 | 4547.0462 | 5696.588 | 4165.8958 | 4653.7827 | 5004.3648 | 5995.9636 | 4708.2537 | 6293.1406 | 4115.7758 | 3990.6519 | 4028.9331 | 3729.1976 | 3978.9902 | 4415.1701 | 3482.3609 | 5076.1465 | 4135.1534 | 4757.2606 | 4945.4346 | 3725.5752 | 3735.4179 | 3669.9695 | 4045.9026 | 7546.0934 | 4099.9764 | 7111.3609 | 6440.0767 | 7964.0295 | 4399.1799 | 7869.8299 | 7650.0374 | 5881.2606 | 3799.7493 | 7552.4415 | 7451.6588 | 7379.5044 | 4031.0138 | 7143.9833 | 7124.2773 | 3007.9469 | 4358.589 | 7197.6617 | 7193.3412 | 7810.8663 | 4640.9833 | 7438.8427 | 6985.7365 | 6964.1504 | 6893.6991 | 7256.2006 | 6955.1563 | 7043.1976 | 7146.2743 | 7256.8122 | 7032.7325 | 7046.3343 | 7191.2311 | 7248.2242 | 7166.0767 | 6856.9469 | 6846.1976 | 7217.2812 | 6924.7856 | 6841.2576 | 6967.6559 | 7170.0669 | 6989.0167 | 6520.3117 | 6776.6057 | 6357.8427 | 6479.5113 | 5164.7158 | 4851.7424 | 5035.174 | 3816.5654 | 6675.9194 | 5076.4848 | 6189.9026 | 6675.9194 | 6778.0098 | 4366.4484 | 7320.3117 | 6541.1032 | 6896.4415 | 4907.8761 | 5426.8191 | 5833.9095 | 5656.4081 | 5430.5221 | 6219.3884 | 5784.1268 | 7039.3383 | 7192.9164 | 6925.5516 | 6408.4513 | 6574.5477 | 6130.4572 | 6377.1446 | 6294.3913 | 6523.2645 | 6329.8663 | 6555.8869 | 5968.6922 | 6623.2871 | 3656.1908 | 3739.1003 | 6560.3048
42 | 2.4564 | 1327.2911 | 3037.2679 | 7059.5123 | 3575.913 | 3411.8191 | 4498.5673 | 7428.2448 | 4021.1259 | 6693.6853 | 4184.6062 | 3351.5669 | 6452.6617 | 3793.9371 | 3473.2212 | 4032.0787 | 3941.1342 | 3892.2207 | 5487.8584 | 5583.4415 | 6212.7321 | 6613.5118 | 5388.5816 | 5400.5035 | 5482.9917 | 3870.5437 | 4455.0157 | 5591.9371 | 4119.2989 | 4577.2999 | 5021.8348 | 6057.4882 | 4640.5669 | 6251.0005 | 4031.9769 | 3968.6377 | 4023.4253 | 3705.4336 | 4004.9607 | 4296.1263 | 3459.7321 | 4958.3623 | 4105.8186 | 4722.5905 | 4853.9926 | 3699.8348 | 3724.1101 | 3733.4558 | 3950.1814 | 7492.8663 | 4069.9921 | 7020.7011 | 6388.3702 | 7968.2591 | 4372.59 | 7824.5914 | 7675.5044 | 5791.6617 | 3790.3073 | 7483.9376 | 7474.2517 | 7210.8982 | 4050.6853 | 7161.6298 | 7054.323 | 3022.6145 | 4364.6145 | 7072.5752 | 7048.2925 | 7796.8269 | 4618.8191 | 7398.8505 | 6950.8908 | 6859.7566 | 6975.4956 | 7191.087 | 6870.8658 | 7086.4567 | 7165.0393 | 7208.2439 | 6999.1106 | 7051.7006 | 7199.1971 | 7221.6617 | 7228.763 | 6829.2842 | 6859.7247 | 7223.7089 | 6846.2601 | 6894.8029 | 7021.5197 | 7183.6377 | 6929.5437 | 6628.5433 | 6809.8348 | 6372.7793 | 6389.7797 | 5124.0477 | 4832.4646 | 5032.8269 | 3849.0472 | 6657.795 | 4997.6145 | 6135.5039 | 6657.795 | 6734.4017 | 4396.9764 | 7307.5993 | 6559.6686 | 6841.0398 | 4867.8663 | 5363.8977 | 5810.5118 | 5646.882 | 5472.3623 | 6218.969 | 5777.2439 | 7012.237 | 7144.5831 | 6906.6141 | 6261.5678 | 6491.8265 | 6013.2517 | 6320.1745 | 6326.7006 | 6534.2365 | 6276.5201 | 6553.2134 | 5995.2989 | 6513.2522 | 3667.441 | 3731.4882 | 6508.0246
43 | 2.4674 | 1381.5516 | 3007.4838 | 7047.5575 | 3603.2212 | 3477.2743 | 4466.1976 | 7392.3658 | 3996.2183 | 6697.8791 | 4139.1504 | 3298.5693 | 6506.5339 | 3797.3569 | 3417.8466 | 4000.1858 | 3888.8053 | 3835.8348 | 5454.5133 | 5596.1121 | 6217.2212 | 6589.1888 | 5394.3953 | 5453.1593 | 5505.2507 | 3837.7788 | 4469.9056 | 5576.7729 | 4167.3245 | 4517.9086 | 4957.1209 | 6026.059 | 4613.8967 | 6207.6165 | 4009.8791 | 3951.9292 | 3955.6667 | 3664.9735 | 4012.3894 | 4311.9794 | 3475.0619 | 4923.4484 | 4099.0236 | 4703.0265 | 4853.4956 | 3651.1239 | 3730.2566 | 3812.7021 | 3899.1593 | 7427.4749 | 4011.6991 | 6968.5457 | 6373.5487 | 7981.9203 | 4412.413 | 7775.5605 | 7494.7846 | 5757.2596 | 3768.5811 | 7436.2301 | 7449.7316 | 7149.9705 | 3942.3422 | 7171.8053 | 7070.2006 | 2972.4513 | 4297.4159 | 7076.8171 | 6935.233 | 7644.9115 | 4522.1799 | 7425.944 | 6877.5929 | 6834.6578 | 6965.2124 | 7082.2153 | 6874.6696 | 7130.0413 | 7175.1622 | 7225.0619 | 6973.2301 | 6924.7138 | 7171.8407 | 7190.3569 | 7301.2537 | 6576.3894 | 6730.531 | 7106.469 | 6829.9381 | 6828.7227 | 6989.9292 | 7143.1829 | 6898.8437 | 6567.4071 | 6685.7551 | 6365.1534 | 6335.1563 | 5161.469 | 4805.2861 | 5013.0649 | 3848.2891 | 6678.4808 | 4991.3304 | 6104.5103 | 6678.4808 | 6744.7139 | 4297.9823 | 7171.8024 | 6601.6696 | 6765.4484 | 4827.7699 | 5333.9646 | 5797.5457 | 5628.3864 | 5508.1711 | 6160.2596 | 5775.4779 | 6880.8967 | 7069.2861 | 6849.2448 | 6233.3068 | 6439.0472 | 6045.5162 | 6197.3953 | 6317.413 | 6529.3923 | 6204.6431 | 6396.6991 | 6018.0295 | 6541.3658 | 3637.8024 | 3708.8348 | 6297.2419
44 | 2.4784 | 1356.7758 | 3015.4562 | 7120.4258 | 3546.175 | 3456.9395 | 4451.5895 | 7432.6731 | 3971.2129 | 6661.8368 | 4120.4676 | 3288.0379 | 6509.2665 | 3813.8137 | 3429.9887 | 3965.7566 | 3882.4715 | 3800.6731 | 5447.3992 | 5617.9774 | 6123.7075 | 6479.3805 | 5263.5438 | 5420.8024 | 5525.0797 | 3859.943 | 4437.0762 | 5601.1824 | 4086.0954 | 4578.2965 | 4945.5969 | 6028.1254 | 4624.8668 | 6222.3004 | 4035.1445 | 3880.3614 | 3866.118 | 3645.3766 | 4035.6273 | 4272.1676 | 3469.3992 | 4856.791 | 4050.2434 | 4687.2129 | 4857.5059 | 3639.7984 | 3696.0457 | 3702.1219 | 3843.0192 | 7316.9926 | 3952.1863 | 6913.7758 | 6386.0988 | 7801.9395 | 4327.1485 | 7728.6544 | 7426.8097 | 5706.8633 | 3742.9887 | 7349.1258 | 7390.1293 | 7106.3918 | 3896.498 | 7052.8977 | 7092.2586 | 2954.0875 | 4244.061 | 7055.9508 | 6988.384 | 7593.2925 | 4429.7453 | 7527.6652 | 6981.2242 | 6864.118 | 6815.6086 | 7043.7984 | 6803.2891 | 7148.5172 | 7093.3235 | 7104.7114 | 6988.1293 | 6852.8211 | 7175.1293 | 7097.411 | 7256.2738 | 6552.4292 | 6648.3422 | 6982.9051 | 6783.2778 | 6763.7679 | 6966.9848 | 7130.7527 | 6927.6996 | 6392.9548 | 6652.9504 | 6326.5855 | 6265.0762 | 5115.7571 | 4789.8137 | 5009.7148 | 3855.2129 | 6682.0266 | 4987.1332 | 6035.3614 | 6682.0266 | 6804.0531 | 4253.4523 | 7148.3574 | 6493.8215 | 6820.8781 | 4792.7871 | 5286.2891 | 5736.9621 | 5611.7871 | 5428.5137 | 6078.6809 | 5715.4031 | 6854.8707 | 7015.2321 | 6793.6273 | 6287.4031 | 6410.1937 | 6061.8555 | 6227.9012 | 6259.6313 | 6397.4454 | 6207.2129 | 6350.3992 | 5978.1027 | 6472.2704 | 3606.9582 | 3672.616 | 6285.2925
45 | 2.4894 | 1318.9184 | 2982.9803 | 7050.4405 | 3556.5605 | 3417.999 | 4484.8968 | 7334.6342 | 3911.6136 | 6649.2439 | 4121.0924 | 3270.5113 | 6443.9794 | 3764.3786 | 3449.3461 | 3954.3166 | 3838.1947 | 3857.3048 | 5406.409 | 5542.5821 | 6068.5713 | 6469.5192 | 5241.939 | 5331.7768 | 5541.9292 | 3797.8889 | 4404.5713 | 5591.1436 | 4043.9381 | 4534.4503 | 4911.9705 | 5943.0531 | 4608.8377 | 6228.7345 | 4015.0511 | 3855.2753 | 3852.3058 | 3616.6018 | 4039.5103 | 4195.8977 | 3389.471 | 4801.7247 | 4050.3265 | 4701.704 | 4732.2871 | 3644.8879 | 3711.4789 | 3657.469 | 3814.4385 | 7400.6913 | 3984.8564 | 6878.0816 | 6305.6853 | 7807.886 | 4302.1731 | 7612.531 | 7483.9282 | 5690.6932 | 3691.5005 | 7230.3677 | 7351.0108 | 7150.7227 | 3971.4071 | 6881.6952 | 7082.1632 | 2968.7444 | 4186.0413 | 7065.7945 | 7031.1219 | 7676.5506 | 4412.2448 | 7477.7876 | 6883.88 | 6910.3353 | 6851.1613 | 7018.6027 | 6804.2035 | 7049.3687 | 6948.94 | 7122.0285 | 6969.5615 | 6887.8781 | 7099.236 | 7008.7856 | 7204.0924 | 6635.0924 | 6711.8663 | 7009.2232 | 6854.3855 | 6704.3078 | 6891.7158 | 7168 | 6899.0924 | 6378.4474 | 6686.2045 | 6228.4916 | 6298.0089 | 4955.4395 | 4766.8781 | 4926.0423 | 3814.5418 | 6598.9508 | 4820.9105 | 6085.4474 | 6598.9508 | 6808.2556 | 4283.1839 | 7166.5506 | 6407.5821 | 6780.6244 | 4803.7139 | 5216.8987 | 5689.5919 | 5602.1632 | 5321.1121 | 5983.3166 | 5699.8161 | 6896.6932 | 7089.294 | 6855.3766 | 6243.1534 | 6447.3166 | 5975.41 | 6188.7965 | 6245.6529 | 6374.4366 | 6095.6961 | 6354.8368 | 5953.9184 | 6344.7758 | 3575.8269 | 3700.7237 | 6376.1219
46 | 2.5004 | 1313.0221 | 2936.2788 | 7061.708 | 3568.8761 | 3392.7566 | 4441.5487 | 7297.292 | 3816.1549 | 6687.0442 | 4029.8407 | 3191.792 | 6424.5796 | 3692.5575 | 3471.5133 | 3919.9823 | 3760.6239 | 3801.4026 | 5278.0575 | 5532.0619 | 6148.562 | 6539.2743 | 5232.7655 | 5334.9469 | 5486.4513 | 3742.1283 | 4423.7212 | 5536.323 | 4049.8982 | 4486.3805 | 4854.3628 | 5961.385 | 4505.5841 | 6125.8274 | 3986.2168 | 3864.5531 | 3852.0664 | 3619.5973 | 4083.7788 | 4223.146 | 3450.9558 | 4840.5664 | 4023.0265 | 4644.8186 | 4715.5664 | 3616.9823 | 3660.9292 | 3688.3805 | 3863.7168 | 7525.4248 | 4000.4071 | 6833.0929 | 6196.9203 | 7825.1814 | 4295.5575 | 7579.4071 | 7596.792 | 5703.3009 | 3620.3274 | 7226.3805 | 7233.3982 | 7126.3584 | 3933.6106 | 6800.4292 | 7056.3451 | 2923.7566 | 4232.5841 | 7087.8584 | 7077.1239 | 7597.208 | 4433.8053 | 7316.8849 | 6804.292 | 6861.5088 | 6824.6372 | 7047.3938 | 6794.708 | 6936.2212 | 6941.2257 | 7178.7301 | 6824.0442 | 6823.6239 | 7011.2389 | 6960.0088 | 7060.2964 | 6584.2168 | 6679.1416 | 7145.9867 | 6885.3318 | 6633.4248 | 6842.5708 | 7061.8053 | 6815.5841 | 6441.5973 | 6644.4292 | 6256.2788 | 6275.3761 | 4937.5929 | 4768.4425 | 4941.9646 | 3794.8894 | 6508.469 | 4823.0664 | 6087.7566 | 6508.469 | 6722.7124 | 4231.9646 | 7073.3805 | 6425.2699 | 6682.5973 | 4763.2655 | 5195.8274 | 5628.5841 | 5559.1991 | 5413.177 | 6103.8496 | 5692.4071 | 6830.761 | 7055.3186 | 6767.9159 | 6100.7876 | 6472.7168 | 5908.2522 | 6177.2743 | 6256.3186 | 6369.9115 | 6035.2478 | 6319.8142 | 5953 | 6342.615 | 3604.5664 | 3722.0664 | 6395.5752
47 | 2.5114 | 1353.6657 | 2950.8771 | 7001.0029 | 3526.1229 | 3339.3923 | 4416.8427 | 7226.057 | 3910.6893 | 6703.3117 | 3961.6037 | 3184.2655 | 6470.1082 | 3707.0541 | 3436.3845 | 3938.0541 | 3755.1386 | 3700.6499 | 5208.1308 | 5420.5723 | 6128.8456 | 6456.1455 | 5192.5959 | 5353.3382 | 5385.2802 | 3742.3923 | 4337.884 | 5420.7532 | 4059.0964 | 4422.2379 | 4768.1494 | 5856.8102 | 4560.6391 | 6059.1809 | 3965.8692 | 3890.4425 | 3745.5801 | 3687.1967 | 4007.53 | 4201.9459 | 3401.6912 | 4858.646 | 4027.9656 | 4524.6303 | 4686.7188 | 3558.7266 | 3586.7807 | 3696.6077 | 3885.3884 | 7449.2802 | 3965.5152 | 6827.6618 | 6233.1927 | 7814.4041 | 4290.2655 | 7537.7306 | 7576.057 | 5718.9656 | 3684.4661 | 7240.2616 | 7241.8506 | 6970.3264 | 3922.9312 | 6903.4277 | 6918.6607 | 2822.8997 | 4206.5841 | 6954.057 | 7020.9341 | 7617.943 | 4392.4612 | 7318.6273 | 6736.4493 | 6788.7886 | 6779.3196 | 6951.1298 | 6728.5536 | 6892.6077 | 6996.2655 | 7131.7345 | 6723.4808 | 6802.3078 | 7016.2694 | 6944.5733 | 6958.6499 | 6522.0305 | 6606.7807 | 7077.8299 | 6821.0423 | 6620.0383 | 6822.6844 | 6988.4385 | 6696.4995 | 6393.8574 | 6607.2114 | 6232.2261 | 6216.6116 | 4952.1308 | 4744.3343 | 5004.4769 | 3802.5615 | 6463.8574 | 4875.8151 | 5977.8535 | 6463.8574 | 6631.5959 | 4149.3766 | 7075.0275 | 6467.177 | 6688.0039 | 4679.9459 | 5113.7109 | 5617.4887 | 5536.5693 | 5357.0491 | 6153.1996 | 5543.8722 | 6749.6077 | 6976.7886 | 6764.9007 | 6082.2429 | 6397.707 | 5929.7847 | 6131.4612 | 6128.9606 | 6305.9194 | 6020.115 | 6309.3963 | 5891.938 | 6314.6844 | 3616.9921 | 3705.4769 | 6304.3648
48 | 2.5224 | 1366.0482 | 2913.3874 | 6886.4464 | 3539.3501 | 3403.234 | 4345.1971 | 7193.326 | 3944.3436 | 6587.4007 | 4026.4091 | 3177.5143 | 6530.9322 | 3762.884 | 3435.175 | 3920.5187 | 3693.3918 | 3657.4508 | 5274.8555 | 5482.8181 | 6092.2625 | 6435.4135 | 5190.617 | 5296.0723 | 5288.1839 | 3753.0875 | 4255.9125 | 5462.1136 | 4092.2429 | 4425.8117 | 4760.8796 | 5738.3501 | 4573.2999 | 6049.8491 | 3959.2232 | 3828.5472 | 3776.9823 | 3594.8033 | 3931.8643 | 4157.6696 | 3367.8073 | 4804.2955 | 4052.1554 | 4611.1092 | 4693.3456 | 3583.0349 | 3512.1008 | 3677.2473 | 3849.116 | 7398.9759 | 3854.8382 | 6823.116 | 6277.5339 | 7838.9125 | 4250.5865 | 7425.882 | 7453.7856 | 5703.2955 | 3740.3151 | 7295.1465 | 7180.9543 | 6982.4331 | 3911.232 | 6848.114 | 6857.7286 | 2853.7001 | 4179.2385 | 6916.1268 | 6884.8579 | 7625.8338 | 4362.2866 | 7292.6347 | 6647.9803 | 6789.558 | 6762.5428 | 6817.6259 | 6690.3348 | 6843.7856 | 6978.5383 | 7008.5078 | 6746.175 | 6709.9061 | 7022.4027 | 6931.4946 | 6927.5821 | 6606.2689 | 6644.1706 | 6973.5339 | 6730.7704 | 6456.2256 | 6672.1096 | 6936.1883 | 6663.1072 | 6262.4838 | 6556.7615 | 6198.3019 | 6140.2365 | 4925.2714 | 4663.7856 | 4946.1204 | 3714.4356 | 6437.6214 | 4787.3436 | 5907.5383 | 6437.6214 | 6605.9322 | 4111.4223 | 7090.1204 | 6405.1008 | 6674.3151 | 4543.8186 | 5153.4922 | 5634.7571 | 5448.7945 | 5291.0393 | 6072.7615 | 5589.619 | 6744.1116 | 6859.7114 | 6777.5669 | 6144.6893 | 6305.6018 | 5928.6455 | 6123.2385 | 6038.2429 | 6265.3874 | 5900.7025 | 6252.1008 | 5784.9213 | 6191.8382 | 3564.0723 | 3593.9739 | 6161.324
49 | 2.5334 | 1318.7316 | 2914.4779 | 6829.6224 | 3489.8436 | 3426.7611 | 4325.1269 | 7168.3628 | 3856.9321 | 6494.3097 | 4045.2419 | 3178.7493 | 6355.469 | 3748.6018 | 3367.354 | 3859.6814 | 3650.1091 | 3619.3805 | 5197.9262 | 5484.4956 | 6062.2124 | 6451.1711 | 5225.9794 | 5263.1504 | 5306.2478 | 3746.0206 | 4259.4395 | 5427.236 | 4044.7227 | 4366.5752 | 4774.4513 | 5783.8289 | 4500.0413 | 5977.4926 | 3930.8112 | 3843.3363 | 3833.531 | 3533.3776 | 3926.0708 | 4206.9469 | 3405.4513 | 4832.6372 | 3974.5339 | 4566.9174 | 4681.1327 | 3495.9764 | 3484.5988 | 3683.2891 | 3802.6313 | 7400.5103 | 3818.4484 | 6822.059 | 6222.4926 | 7826.5103 | 4259.0177 | 7353.5103 | 7418.2301 | 5693.8791 | 3711.2006 | 7159.9587 | 7193.0649 | 6933.6549 | 3856.6106 | 6705.7817 | 6871.5103 | 2894.5103 | 4177.2714 | 6998.8673 | 6740.8643 | 7522.3127 | 4376.1298 | 7264.9587 | 6625.7906 | 6721.7729 | 6707.2507 | 6837.3982 | 6610.8436 | 6800.5605 | 6909.2124 | 6938.7994 | 6749.7404 | 6598.8319 | 7005.1003 | 6854.4926 | 6976.8466 | 6597.6047 | 6610.1032 | 6939.1416 | 6705.9381 | 6596.1475 | 6649.6667 | 6997.1947 | 6718.9381 | 6203.7493 | 6522.7404 | 6191.3923 | 6075.4602 | 4837.7434 | 4663.0384 | 4882.4012 | 3756.1151 | 6527.7257 | 4773.6785 | 5955.6165 | 6527.7257 | 6530.174 | 4121.0295 | 7123.7463 | 6354 | 6723.6667 | 4504.708 | 5122.9056 | 5547.764 | 5329.8024 | 5240.4336 | 6016.9115 | 5615.2242 | 6717.2212 | 6901.6431 | 6645.8141 | 5994.9174 | 6248.4425 | 5806.5959 | 6114.1917 | 6133.4956 | 6246.4189 | 5964.7139 | 6258.8171 | 5726.2507 | 6222.1151 | 3544.1298 | 3551.5782 | 6228.7139
50 | 2.5444 | 1317.4769 | 2919.9095 | 6832.738 | 3408.8869 | 3405.3751 | 4325.0349 | 7145.2611 | 3845.6701 | 6427.5354 | 4002.1367 | 3123.9557 | 6217.3525 | 3686.6598 | 3321.6362 | 3845.4204 | 3657.7842 | 3720.4071 | 5123.0452 | 5271.4911 | 6029.6023 | 6419.7389 | 5190.4326 | 5253.9887 | 5254.4902 | 3680.9897 | 4264.7955 | 5339.1819 | 4017.2384 | 4253.0462 | 4719.8643 | 5757.4439 | 4550.6578 | 5965.6927 | 3894.7276 | 3779.3658 | 3782.6598 | 3551.3412 | 3941.4995 | 4113.8889 | 3309.9105 | 4862.3638 | 3967.5787 | 4538.6804 | 4633.3638 | 3444.3402 | 3547.3741 | 3661.6711 | 3814.5108 | 7216.526 | 3860.8746 | 6819.7389 | 6205.2724 | 7705.1612 | 4208.4902 | 7371.2837 | 7405.557 | 5632.9218 | 3683.1254 | 7059.4769 | 7206.1829 | 6909.9651 | 3879.3176 | 6741.6239 | 6924.2149 | 2858.6824 | 4168.8407 | 6952.6834 | 6760.4985 | 7392.854 | 4380.8068 | 7212.5354 | 6628.2384 | 6609.7286 | 6696.3977 | 6799.3648 | 6635.2827 | 6850.6126 | 6857.9322 | 7001.8058 | 6680.7625 | 6535.6598 | 6898.0133 | 6878.8284 | 6923.4223 | 6519.9548 | 6534.2847 | 6815.2178 | 6628.8087 | 6573.0604 | 6725.5108 | 6948.7409 | 6633.9793 | 6273.6239 | 6491.705 | 6066.854 | 6103.7153 | 4874.7144 | 4707.6814 | 4797.9444 | 3827.409 | 6468.1499 | 4793.7955 | 5929.3196 | 6468.1499 | 6410.7625 | 4100.8982 | 7065.9906 | 6334.7276 | 6597.7532 | 4599.851 | 5116.6691 | 5460.0801 | 5357.2262 | 5264.6804 | 5914.6947 | 5610.6013 | 6676.5118 | 7034.738 | 6670.5098 | 6022.0885 | 6340.4749 | 5805.1008 | 6030.3761 | 6032.3894 | 6190.9331 | 6000.0123 | 6173.1273 | 5727.443 | 6222.9444 | 3494.001 | 3689.0207 | 6177.2866
51 | 2.5554 | 1303.3451 | 2890.4798 | 6737.7768 | 3330.0029 | 3329.7119 | 4284.3235 | 7068.119 | 3787.3264 | 6494.0777 | 4014.9676 | 3093.9056 | 6213.3019 | 3600.5447 | 3281.294 | 3917.8299 | 3623.6578 | 3749.5959 | 5150.0511 | 5263.121 | 5969.7738 | 6272.5693 | 5035.0167 | 5256.9164 | 5227.6549 | 3706.0698 | 4252.3235 | 5340.3019 | 3907.235 | 4291.7571 | 4699.0728 | 5721.3963 | 4497.6716 | 5887.3589 | 3830.9921 | 3742.3314 | 3767.2183 | 3612.2881 | 3943.6362 | 4138.9214 | 3318.6116 | 4782.8682 | 3913.5772 | 4597.2183 | 4637.4474 | 3386.6499 | 3566.3235 | 3563.41 | 3767.6283 | 7184.0374 | 3904.1131 | 6730.9518 | 6216.1563 | 7685.768 | 4261.5467 | 7264.9331 | 7302.7443 | 5671.881 | 3666.3127 | 7001.7119 | 7094.2537 | 6839.4425 | 3759.882 | 6779.7709 | 6807.1327 | 2892.6952 | 4117.0029 | 6838.6175 | 6785.9597 | 7406.9754 | 4305.7955 | 7266.7866 | 6668.1858 | 6578.5929 | 6593.9843 | 6719.8977 | 6565.7768 | 6817.3264 | 6837.3963 | 7010.9597 | 6571.3992 | 6562.0079 | 6818.3235 | 6828.0462 | 6975.6087 | 6497.2507 | 6505.3963 | 6715.4287 | 6658.3284 | 6546.2449 | 6695.2212 | 6805.3156 | 6658.943 | 6234.9518 | 6457.178 | 6021.3638 | 6165.175 | 4931.8437 | 4745.9676 | 4878.9833 | 3866.4877 | 6397.6037 | 4812.0511 | 5915.0295 | 6397.6037 | 6469.1721 | 4024.0029 | 6927.6283 | 6261.7738 | 6511.0403 | 4612.6578 | 5148.0216 | 5440.2291 | 5375.4907 | 5231.6499 | 5928.6008 | 5542.1406 | 6689.884 | 6971.7335 | 6618.0678 | 5985.0895 | 6234.7178 | 5881.353 | 6099.1613 | 6020.2557 | 6214.059 | 5952.8653 | 6159.4445 | 5705.5015 | 6150.9597 | 3432.5742 | 3677.5772 | 6131.2478
52 | 2.5664 | 1342.4159 | 2893.2832 | 6741.9823 | 3418.9956 | 3292.5221 | 4344.6991 | 7064.4012 | 3777.4764 | 6502.3156 | 3954.1519 | 3083.1947 | 6171.4941 | 3577.6711 | 3304.6549 | 3864.4661 | 3657.2817 | 3676.6136 | 5184.5664 | 5301.4631 | 5957.0738 | 6208.9395 | 5004.8938 | 5199.823 | 5305.8038 | 3644.705 | 4159.7492 | 5293.4793 | 3905.4602 | 4317.3289 | 4660.2552 | 5798.1165 | 4426.5973 | 5863.1033 | 3833.4764 | 3701.1814 | 3726.6136 | 3620.6431 | 3849.3761 | 4146.4351 | 3386.7153 | 4729.4779 | 3979.1003 | 4509.6003 | 4598.2404 | 3413.9823 | 3506.3451 | 3542.3274 | 3761.7301 | 7236.1195 | 3938.0443 | 6633.733 | 6076.4675 | 7755.8496 | 4265.0619 | 7259.8333 | 7279.9838 | 5551.4837 | 3671.5221 | 6958.2537 | 6960.2419 | 6916.1593 | 3757.2508 | 6675.5708 | 6814.0561 | 2860.5546 | 4081.2537 | 6935.5177 | 6689.3908 | 7332.3923 | 4315.0428 | 7163.2596 | 6699.0295 | 6629.118 | 6529.3732 | 6683.4631 | 6527.9838 | 6676.4661 | 6770.5693 | 6940.7183 | 6606.1018 | 6547.6136 | 6807.2389 | 6710.823 | 6910.8407 | 6340.0649 | 6536.3569 | 6830.3983 | 6663.5841 | 6534.5251 | 6659.6121 | 6771.0148 | 6727.2684 | 6160.2994 | 6449.7611 | 5997.1504 | 6100.1077 | 5011.9823 | 4799.2375 | 4828.2433 | 3831.7035 | 6312.1077 | 4759.4351 | 5892.4941 | 6312.1077 | 6584.6549 | 4053.4454 | 6913.5811 | 6280.4159 | 6489.7625 | 4569.1799 | 5058.3908 | 5430.3437 | 5388.3274 | 5173.1047 | 5948.4336 | 5442.5236 | 6575.1091 | 6795.3923 | 6490.823 | 6007.0413 | 6199.3702 | 5776.8997 | 6054.1681 | 6050.5383 | 6281.8201 | 5930.2094 | 6124.0177 | 5695.149 | 6085.1062 | 3446.7153 | 3642.8053 | 6110.1062
53 | 2.5774 | 1300.5821 | 2894.2478 | 6679.7079 | 3373.2035 | 3273.9911 | 4251.4208 | 6988.9036 | 3712.3559 | 6399.1691 | 3868.0954 | 3040.0433 | 6191.6568 | 3599.1436 | 3305.3215 | 3805.7217 | 3625.7257 | 3713.5526 | 5102.1298 | 5216.6647 | 5891.469 | 6142.7473 | 4973.9213 | 5110.9518 | 5270.4169 | 3628.4612 | 4186.9744 | 5278.061 | 3943.5261 | 4288.826 | 4618.1042 | 5730.0383 | 4423.5565 | 5938.3265 | 3845.0521 | 3755.3137 | 3680.4651 | 3552.3432 | 3876.9744 | 4086.2694 | 3396.6912 | 4688.5693 | 4003.9823 | 4436.7434 | 4540.3215 | 3356.1819 | 3452.6303 | 3541.3825 | 3674.3471 | 7179.3953 | 3745.0334 | 6585.1868 | 6077.7267 | 7609.1858 | 4200.7955 | 7312.2262 | 7287.939 | 5447 | 3658.1042 | 6962.1957 | 7064.588 | 6945.3127 | 3728.1081 | 6635.7306 | 6825.5084 | 2878.5221 | 4021.2822 | 6802.8072 | 6649.0305 | 7389.1052 | 4267.8299 | 7122.8004 | 6573.9774 | 6660.0305 | 6566.0698 | 6606.9734 | 6512.9302 | 6689.944 | 6702.6824 | 6845.7257 | 6648.8043 | 6630.5703 | 6827.3697 | 6747.8614 | 6831.1957 | 6342.118 | 6490.2517 | 6864.3432 | 6637.6696 | 6328.5772 | 6635.2694 | 6777.1957 | 6594.2252 | 6029.9557 | 6428.9911 | 5948.8781 | 5951.408 | 4875.5683 | 4686.5034 | 4797.0177 | 3837.9046 | 6328.3048 | 4704.3825 | 5758.6254 | 6328.3048 | 6564.8692 | 4007.3559 | 6914.0216 | 6209.1819 | 6502.7915 | 4501.9518 | 4994.3825 | 5435.5044 | 5395.6912 | 5099.9734 | 5877.4779 | 5419.4435 | 6488.9695 | 6689.765 | 6530.4179 | 6068.5565 | 6151.8515 | 5788.4484 | 6065.6136 | 6000.4041 | 6179.6775 | 5950.7611 | 6059.1042 | 5745.3786 | 6125.0266 | 3491.2045 | 3612.4562 | 6104.0089
54 | 2.5884 | 1270.3771 | 2836.9317 | 6676.6283 | 3303.0457 | 3233.6401 | 4275.6682 | 6938.7714 | 3738.2625 | 6324.8687 | 3834.5772 | 3010.2743 | 6093.5777 | 3606.6514 | 3202.6632 | 3813.4685 | 3640.1598 | 3599.5206 | 5036.2743 | 5287.651 | 5853.8171 | 6177.474 | 4979.737 | 5183.9253 | 5144.0118 | 3594.3717 | 4170.4174 | 5257.9887 | 3935.1372 | 4242.823 | 4575.4287 | 5667.2571 | 4434.12 | 5982.6799 | 3815.2458 | 3728.526 | 3593.7832 | 3539.6799 | 4001.7768 | 4103.3142 | 3377.0172 | 4757.2453 | 3953.8972 | 4334.4975 | 4541.8913 | 3345.8569 | 3471.2684 | 3468.1603 | 3710.0625 | 7126.5029 | 3732.6224 | 6626.7655 | 6025.3318 | 7514.3486 | 4145.6116 | 7331.6799 | 7365.3368 | 5455.4685 | 3648 | 6890.3889 | 6983.6293 | 6844.646 | 3703.8456 | 6596.5374 | 6703.8461 | 2975.4912 | 3973.9715 | 6767.6854 | 6623.0344 | 7316.0806 | 4187.4744 | 7017.2463 | 6545.5998 | 6584.526 | 6559.6401 | 6668.971 | 6503.7257 | 6699.383 | 6664.1544 | 6820.4228 | 6424.8638 | 6464.3495 | 6786.5831 | 6858.0511 | 6789.3545 | 6442.5084 | 6487.177 | 6785.8859 | 6628.3201 | 6291.4169 | 6647.1028 | 6774.44 | 6574.3255 | 6118.7365 | 6348.9031 | 5908.2458 | 5886.2743 | 4784.3771 | 4761.7822 | 4907.6854 | 3872.7257 | 6215.5663 | 4648.7488 | 5754.0226 | 6215.5663 | 6382.2635 | 3973.9828 | 6868.5202 | 6045.1204 | 6500.8515 | 4445.2915 | 4998.0285 | 5356.0118 | 5328.4631 | 5067.5143 | 5820.2458 | 5335.2689 | 6514.9656 | 6703.5713 | 6526.0973 | 5954.6976 | 6100.2571 | 5686.4749 | 6076.7773 | 5935.7488 | 6121.1657 | 5920.0973 | 6137.5939 | 5736.4744 | 5964.6637 | 3466.263 | 3533.6288 | 6038.1775
55 | 2.5994 | 1247.1327 | 2820.6106 | 6671.4956 | 3264.8319 | 3286.0354 | 4297.8053 | 6928.7345 | 3781.9823 | 6279.7788 | 3850.2389 | 2995.3717 | 6125.2301 | 3545.938 | 3214.7434 | 3775.5221 | 3678.2832 | 3627.1859 | 5049.9115 | 5200.3274 | 5849.3717 | 6173.5398 | 5038.8319 | 5118.8672 | 5137.823 | 3549.7257 | 4192.4248 | 5197.2035 | 3928.6814 | 4230.6372 | 4567.292 | 5611.9203 | 4443.646 | 5944.2035 | 3811.9558 | 3731.1328 | 3592.823 | 3612.0443 | 3919.5575 | 4099.0885 | 3350.4159 | 4757.8761 | 3883.2566 | 4411.1062 | 4521.469 | 3399.8761 | 3484.6814 | 3435.3274 | 3627.9292 | 7158.8053 | 3743.1947 | 6560.0885 | 5901.646 | 7538.5575 | 4200.9912 | 7224.8407 | 7294.1858 | 5521.7257 | 3642.0265 | 6953.8761 | 6862.469 | 6855.4513 | 3726.9469 | 6537.1416 | 6670.5929 | 2887.2124 | 4083.8142 | 6740.6549 | 6575.8053 | 7291.7611 | 4168.3363 | 7023.9735 | 6430.6194 | 6510.8053 | 6478.6283 | 6627.4248 | 6423.0531 | 6728.7965 | 6722.6903 | 6800.115 | 6555.5576 | 6440.8938 | 6712.8672 | 6721.1062 | 6822.4159 | 6470.9912 | 6545.823 | 6765.3009 | 6572.5487 | 6329.9558 | 6573.7168 | 6771.3363 | 6526.0177 | 6104.0973 | 6281.8053 | 5942.1239 | 5948.3628 | 4836.7965 | 4736.1416 | 4855.4336 | 3803.3363 | 6223.6018 | 4562.2743 | 5762.7611 | 6223.6018 | 6339.823 | 3983.3274 | 6806.5221 | 6084.5752 | 6480.3717 | 4434.3186 | 4982.4336 | 5367.0797 | 5296.9912 | 5006.6195 | 5806.3363 | 5364.0089 | 6528.3451 | 6679.8496 | 6418.0354 | 5817.3363 | 6092.0442 | 5676.0266 | 5976.0354 | 5911 | 6127.6549 | 5900.9735 | 6060.2212 | 5690.1239 | 6027.2478 | 3471.5487 | 3552.4513 | 6142.6814
56 | 2.6104 | 1249.5762 | 2911.501 | 6669.2242 | 3244.648 | 3289.5796 | 4208.9911 | 6894.6357 | 3770.4798 | 6321.7566 | 3906.2925 | 2981.5359 | 6109.4272 | 3471.0595 | 3277.1244 | 3815.5167 | 3554.8825 | 3614.3992 | 4991.7232 | 5281.8338 | 5965.3294 | 6133.0998 | 4924.8264 | 4987.0192 | 5160.704 | 3585.5885 | 4130.7669 | 5182.7881 | 3919.5639 | 4240.0841 | 4588.2242 | 5614.0964 | 4398.0034 | 5856.1032 | 3808.2119 | 3623.5305 | 3681.2768 | 3576.7232 | 3846.5359 | 4028.1593 | 3356.5482 | 4749.5482 | 3859.4081 | 4405.3712 | 4503.296 | 3275.4184 | 3467.4518 | 3510.3609 | 3569.4641 | 7091.7669 | 3769.6726 | 6571.4887 | 5851.1278 | 7540.1839 | 4125.7547 | 7189.3083 | 7405.354 | 5442.8107 | 3643.056 | 6898.5147 | 6896.3486 | 6777.9228 | 3733.676 | 6566.2925 | 6616.1313 | 2837.1681 | 4098.3835 | 6735.0964 | 6628.9528 | 7339.9562 | 4136.9036 | 6931.8264 | 6369.4641 | 6472.6917 | 6536.9371 | 6721.6377 | 6503.3855 | 6691.8353 | 6697.8353 | 6766.0438 | 6471.7615 | 6536.2925 | 6667.0157 | 6542.4149 | 6782.0472 | 6307.0541 | 6363.1976 | 6555.3815 | 6529.7478 | 6418.149 | 6524.5919 | 6640.5305 | 6501.5482 | 5993.4833 | 6304.5325 | 5915.9754 | 6011.3923 | 4793.7512 | 4693.4921 | 4803.972 | 3833.9528 | 6153.4026 | 4553.3083 | 5783.7724 | 6153.4026 | 6199.7861 | 3962.3677 | 6759.9877 | 6132.324 | 6419.2153 | 4474.5325 | 4895.2871 | 5271.6145 | 5275.0157 | 5003.0683 | 5752.5954 | 5443.5885 | 6464.8073 | 6661.704 | 6401.6883 | 5801.9843 | 6078.8879 | 5775.2208 | 5902.2276 | 6009.6972 | 6173.6165 | 5891.1558 | 6007.1401 | 5665.8599 | 6070.4115 | 3453.9036 | 3554.0157 | 6063.5462
57 | 2.6214 | 1234.5978 | 2890.0531 | 6556.5477 | 3299.7522 | 3231.6372 | 4196.0256 | 6718.8731 | 3711.6893 | 6254.3913 | 3795.1583 | 3003.1554 | 6108.9735 | 3568.3874 | 3309.5192 | 3773.4169 | 3555.6352 | 3552.4031 | 5019.5703 | 5228.9764 | 5766.2645 | 6019.002 | 4834.1563 | 4930.9744 | 5159.2468 | 3597.7404 | 4094.1947 | 5123.6765 | 3843.7286 | 4196.1957 | 4551.9489 | 5583.5723 | 4296.5339 | 5757.5467 | 3733.7935 | 3699.3481 | 3550.1848 | 3556.3894 | 3929.414 | 3989.3638 | 3343.351 | 4671.6509 | 3798.6765 | 4303.8584 | 4421.469 | 3245.4798 | 3442.1042 | 3492.3127 | 3516.8712 | 7034.5457 | 3760.6499 | 6517.4041 | 5821.2861 | 7546.0128 | 4037.9616 | 7071.2881 | 7231.6411 | 5425.7915 | 3592.001 | 6725.1711 | 6797.4828 | 6689.767 | 3706.468 | 6650.6352 | 6577.0521 | 2783.118 | 4013.6765 | 6738.2989 | 6698.4798 | 7276.352 | 4142.649 | 6751.6775 | 6443.5182 | 6406.2478 | 6547.9616 | 6565.6145 | 6408.1327 | 6687.7788 | 6641.468 | 6678.9105 | 6572.6716 | 6474.0275 | 6587.1701 | 6620.1406 | 6721.468 | 6291.3235 | 6253.6106 | 6622.3216 | 6496.2212 | 6363.0147 | 6555.7267 | 6507.1052 | 6517.2596 | 6047.2842 | 6236.9626 | 5766.0413 | 5947.9105 | 4700.0531 | 4661.1436 | 4873.8948 | 3868.3766 | 6060.1436 | 4625.9469 | 5691.6254 | 6060.1436 | 6266.1662 | 3934.1042 | 6679.8584 | 6110.6627 | 6252.2753 | 4432.4297 | 4874.8309 | 5353.7896 | 5224.767 | 4953.6765 | 5658.8584 | 5350.4307 | 6449.4543 | 6551.4956 | 6369.9489 | 5872.882 | 6045.4031 | 5695.1839 | 5871.8574 | 5943.4307 | 6064.9115 | 5897.3894 | 5999.6106 | 5581.1573 | 6008.7021 | 3389.4425 | 3533.3638 | 6177.9056
58 | 2.6324 | 1287.0531 | 2821.3658 | 6446.6534 | 3377.5443 | 3254.736 | 4186.3835 | 6777.4115 | 3545.0796 | 6190.9823 | 3758.8201 | 2972.264 | 6096.3776 | 3563.851 | 3237.1327 | 3772.5266 | 3478.2109 | 3455.6298 | 5077.826 | 5147.8805 | 5730.5679 | 5959.7906 | 4881.5502 | 4869.8569 | 5152.3953 | 3596.1018 | 4216.1003 | 5086.5929 | 3806.0959 | 4189.0118 | 4470.3481 | 5398.3613 | 4214.7611 | 5694.7729 | 3653.6593 | 3656.2109 | 3543.2566 | 3574.6106 | 3940.9764 | 4027.736 | 3344.5089 | 4590.9646 | 3818.3407 | 4278.6047 | 4448.0531 | 3285.1195 | 3386.3599 | 3346.0973 | 3466.4793 | 7022.3953 | 3669.6357 | 6384.0192 | 5806.0959 | 7347.1283 | 4030.0059 | 7041.0236 | 7188.1608 | 5369.8392 | 3567.2994 | 6777.1903 | 6795.2448 | 6585.8274 | 3676.4852 | 6474.4277 | 6528.8628 | 2791.0295 | 3916.3422 | 6655.2345 | 6562.972 | 7210.0723 | 4123.2758 | 6942.8053 | 6452.5811 | 6326.354 | 6521.09 | 6513.0413 | 6349.2094 | 6546.2714 | 6544.9292 | 6548.0192 | 6399.8775 | 6406.5752 | 6588.2271 | 6587.4262 | 6562.5767 | 6233.1268 | 6252.2035 | 6575.6961 | 6371.5044 | 6275.8628 | 6465.9115 | 6566.7773 | 6470.5575 | 6043.6652 | 6120.2345 | 5803.5797 | 5855.5516 | 4704.9277 | 4618.7788 | 4901.0959 | 3876.0059 | 6118.6637 | 4601.4557 | 5620.5811 | 6118.6637 | 6225.2109 | 3816.0074 | 6702.9454 | 6105.9041 | 6297.5915 | 4300.6121 | 4774.8097 | 5334.6298 | 5212.6047 | 4920.7965 | 5707.8673 | 5212.8274 | 6448.2935 | 6540.233 | 6360.5089 | 5790.7079 | 6093.6519 | 5595.2581 | 5925.9513 | 5877.9882 | 5916.3245 | 5627.6681 | 5930.9469 | 5617.8614 | 6020.1254 | 3352.2876 | 3580.8437 | 6043.0147
59 | 2.6434 | 1215.6224 | 2758.3756 | 6460.7827 | 3330.8505 | 3182.4641 | 4206.5152 | 6776.2566 | 3547.4464 | 6232.6146 | 3728.3658 | 2965.4061 | 6049.3559 | 3571.7925 | 3343.4376 | 3605.6411 | 3568.4474 | 3485.5851 | 4881.4129 | 5050.3156 | 5785.6736 | 5992.0698 | 4924.1091 | 4890.7729 | 5067.3058 | 3611.9705 | 4187.1573 | 5015.2763 | 3838.2183 | 4078.0875 | 4452.5546 | 5390.0207 | 4253.4563 | 5683.8515 | 3675.7729 | 3596.0295 | 3618.3373 | 3536.5143 | 3935.4159 | 3957.6027 | 3269.8702 | 4508.8407 | 3765.5732 | 4272.4454 | 4407.9892 | 3280.7424 | 3465.0905 | 3374.5556 | 3486.6539 | 6891.6814 | 3617.7522 | 6344.4258 | 5776.9105 | 7271.4061 | 3983.0983 | 7072.9607 | 7107.0777 | 5180.7404 | 3534.2075 | 6739.5929 | 6752.4149 | 6653.0806 | 3596.999 | 6388.4258 | 6505.0098 | 2734.7119 | 3887.7129 | 6565.6824 | 6527.2281 | 7024.1956 | 4070.3658 | 6890.1662 | 6441.5742 | 6327.238 | 6493.1878 | 6548.9705 | 6420.2979 | 6489.5546 | 6502.5841 | 6636.0315 | 6372.5457 | 6279.939 | 6446.5919 | 6480.1376 | 6650.4179 | 6121.1475 | 6254 | 6557.9312 | 6348.6637 | 6132.9587 | 6383.9302 | 6515.4641 | 6258.6607 | 6025.2871 | 6048.5044 | 5779.1278 | 5937.9322 | 4691.1976 | 4647.9213 | 4876.9007 | 3823.3953 | 6190.9213 | 4484.8997 | 5593.178 | 6190.9213 | 6222.06 | 3835.2979 | 6517.8387 | 6138.6539 | 6242.058 | 4280.9508 | 4802.4268 | 5298.2871 | 5071.7109 | 4901.178 | 5610.117 | 5158.8909 | 6446.4258 | 6472.88 | 6297.9892 | 5841.6647 | 6109.3068 | 5459.2261 | 5922.7719 | 5874.1386 | 5897.2478 | 5719.5871 | 5889.0295 | 5589.118 | 5973.8309 | 3387.3471 | 3560.94 | 5891.9695
60 | 2.6544 | 1196.3048 | 2803.8152 | 6326.6145 | 3264.5659 | 3252.6003 | 4173.0393 | 6631.1711 | 3562.1224 | 6165.9793 | 3789.5934 | 2970 | 5970.4597 | 3534.2379 | 3287.5983 | 3754.9027 | 3525.8269 | 3559.1362 | 4884.7458 | 4933.0555 | 5746.1386 | 5959.1386 | 4798.2935 | 4889.1455 | 4977.4366 | 3528.8038 | 4103.5128 | 4958.2679 | 3849.8245 | 4007.7945 | 4436.4204 | 5367.8638 | 4265.3972 | 5711.8314 | 3590.7507 | 3598.7021 | 3569.559 | 3548.0069 | 3892.3142 | 4001.0438 | 3243.8014 | 4494.6028 | 3766.4572 | 4318.3117 | 4359.97 | 3280.0762 | 3455.6883 | 3411.351 | 3534.5334 | 6849.4272 | 3617.9769 | 6320.3441 | 5735.5428 | 7225.9631 | 3972.4272 | 7033.1386 | 6925.6676 | 5106.1224 | 3490.5428 | 6524.707 | 6737.6028 | 6473.8225 | 3618.0369 | 6416.5565 | 6563.9145 | 2795.9214 | 3927.1593 | 6521.2448 | 6418.6214 | 7112.3624 | 4045.2979 | 6773.3373 | 6253.1804 | 6360.9307 | 6485.4503 | 6525.9931 | 6331.6976 | 6407.4297 | 6445.6121 | 6586.4759 | 6309.1018 | 6287.6627 | 6477.5472 | 6404.8638 | 6576.118 | 6086.649 | 6279.381 | 6558.2217 | 6330.7183 | 6090.9769 | 6278.3535 | 6485.5265 | 6280.6234 | 5962.6352 | 6053.3048 | 5763.8776 | 5942.5659 | 4677.0993 | 4702.4341 | 4751.9725 | 3701.6514 | 6103.3004 | 4536.1962 | 5436.6145 | 6103.3004 | 6108.469 | 3806.1617 | 6519.5703 | 6021.9194 | 6214.0531 | 4267.7945 | 4690.5914 | 5197.2011 | 5013.2748 | 4914.3048 | 5558.4503 | 5151.2748 | 6239.6008 | 6328.7738 | 6243.4435 | 5759.8962 | 5887.4483 | 5496.3948 | 5807.3004 | 5772.0256 | 5868.7414 | 5751.0624 | 5922.1593 | 5577.3279 | 5921.7183 | 3396.0231 | 3623.7689 | 5957.471
61 | 2.6654 | 1263.7345 | 2868.5605 | 6251.5634 | 3238.2006 | 3159.5192 | 4052.7551 | 6646.2094 | 3550.6254 | 6018.6932 | 3727.5546 | 2880.0295 | 5794.0826 | 3415.4425 | 3215.5015 | 3766.7286 | 3473.5133 | 3562.9882 | 4836.879 | 4983.8584 | 5650.5664 | 5994.8348 | 4767.174 | 4854.7139 | 4946.413 | 3586.4454 | 4005.8289 | 4886.528 | 3713.3068 | 4038.823 | 4360.4159 | 5340.413 | 4215.2773 | 5626.8672 | 3560.8496 | 3586.5133 | 3538.1003 | 3538.351 | 3894.9115 | 3981.59 | 3250.4248 | 4515.2242 | 3832.5723 | 4273.5782 | 4239.9056 | 3298.9617 | 3406.0265 | 3377.6018 | 3497.115 | 6729.1062 | 3608.4631 | 6288.1386 | 5724.7375 | 7178.2389 | 3925.115 | 6944.941 | 6758.9557 | 5054.0649 | 3515.8732 | 6478.1239 | 6647.6047 | 6583.2154 | 3635.5251 | 6374.6401 | 6356.174 | 2862.4484 | 3953.7994 | 6500.4631 | 6379.6372 | 6907.2507 | 3998.4897 | 6730.1386 | 6192.3245 | 6341.2389 | 6345.882 | 6290.8112 | 6241.5398 | 6394.9115 | 6363.0531 | 6568.649 | 6264.6254 | 6312.4248 | 6479.0767 | 6403.2242 | 6461.5782 | 5983.2301 | 6091.3599 | 6363.4719 | 6233.7168 | 6143.41 | 6256.5369 | 6347.882 | 6272.2773 | 5767.1711 | 6131.7965 | 5731.8643 | 5796.7315 | 4643.5516 | 4665.3776 | 4851.8053 | 3772.4071 | 5985.8289 | 4474.1298 | 5314.3658 | 5985.8289 | 6112.8348 | 3829.9853 | 6379.4336 | 5914.0029 | 6137.1416 | 4257.5251 | 4716.646 | 5035.8436 | 5032.9115 | 4805.4808 | 5526.5516 | 5086.8407 | 6184.5988 | 6332.7227 | 6059.6961 | 5710.3127 | 5773.2891 | 5405.5811 | 5759.1504 | 5713.0885 | 5871.9617 | 5669.7906 | 5852.9292 | 5521.528 | 5847.5782 | 3449.0472 | 3463.7345 | 5811.7964
62 | 2.6764 | 1241.5143 | 2803.526 | 6197.7655 | 3210.7576 | 3099.2542 | 4095.474 | 6515.7999 | 3523.2581 | 5930.0157 | 3618.5251 | 2855.4975 | 5767.7409 | 3379.2502 | 3290.472 | 3737.2512 | 3458.2532 | 3499.5192 | 4712.0315 | 5018.998 | 5625.2493 | 6008.002 | 4839.2286 | 4746.0324 | 4829.0373 | 3526.7802 | 4040.2296 | 4942.9744 | 3692.7409 | 3927.791 | 4431.9685 | 5319.7557 | 4190.0029 | 5572.0069 | 3617.9823 | 3489.53 | 3506.0098 | 3487.5143 | 3846.5182 | 3946.7557 | 3204.0157 | 4467.0177 | 3756.7812 | 4233.7557 | 4294.469 | 3212.7802 | 3386.0029 | 3404.737 | 3466.2542 | 6635.5162 | 3656.4828 | 6202.7753 | 5730.998 | 7021.0482 | 3875.5103 | 6863.0187 | 6715.001 | 5066.9902 | 3528.2502 | 6581.9685 | 6462.5496 | 6420.3259 | 3560.0246 | 6096.585 | 6289.7448 | 2816.2689 | 3960.2493 | 6397.0324 | 6248.5437 | 6774.0128 | 4040.7311 | 6681.5133 | 6106.0295 | 6129.3161 | 6152.0472 | 6268.7311 | 6060.0541 | 6276.0413 | 6429.471 | 6428.2984 | 6199.0206 | 6200.7871 | 6402.5211 | 6239.5555 | 6364.7743 | 5899.0157 | 5911.5359 | 6282.5044 | 6095.0354 | 6152.0029 | 6259.5005 | 6364.7281 | 6229.7586 | 5819.9627 | 6085.7733 | 5664.2679 | 5657.7802 | 4707.7252 | 4617.7596 | 4855.0157 | 3809.5005 | 5903.2679 | 4414.5084 | 5324.7389 | 5903.2679 | 6007.5408 | 3773.2738 | 6297.0069 | 5734.8018 | 6003.7842 | 4234.5074 | 4648.2817 | 5049.7311 | 4968.0236 | 4817.2326 | 5406.2856 | 5150.2217 | 6142.0167 | 6258.2797 | 6072.2257 | 5647.5201 | 5767.998 | 5359.001 | 5535.5722 | 5720.2453 | 5678.5654 | 5536.2866 | 5586.8274 | 5426.5241 | 5723.7832 | 3414.2679 | 3470.2257 | 5738.2532
63 | 2.6874 | 1215.5693 | 2783.0964 | 6085.9823 | 3167.9459 | 3099.5723 | 4097.6145 | 6398.7021 | 3524.1445 | 5810.3618 | 3656.2498 | 2845.8073 | 5732.8014 | 3384.3343 | 3296.1386 | 3713.1386 | 3508.3432 | 3475.4277 | 4762.3009 | 4889.9282 | 5577.2291 | 5763.6696 | 4806.8948 | 4724.4307 | 4832.9125 | 3556.2016 | 3927.6421 | 4872.2202 | 3737.8161 | 3940.8673 | 4407.6086 | 5186.7384 | 4071.9312 | 5440.7866 | 3547.6027 | 3449.8555 | 3465.5182 | 3505.0059 | 3844.765 | 3992.8191 | 3201.7168 | 4346.8859 | 3738.7168 | 4186.0875 | 4187.3077 | 3140.0875 | 3369.2832 | 3334.9371 | 3461.238 | 6555.7502 | 3510.1603 | 6070.3137 | 5648.413 | 6885.1268 | 4031.1259 | 6712.1662 | 6594.3589 | 5073.1898 | 3506.6627 | 6404.4375 | 6387.6145 | 6297.2773 | 3589.8673 | 6093.0629 | 6291.9518 | 2797 | 3925.1839 | 6338.0905 | 6220.765 | 6741.9489 | 3965.5064 | 6498.8259 | 6038.6598 | 6084.8614 | 6142.8643 | 6280.3795 | 6072.822 | 6170.7502 | 6381.2714 | 6318.7984 | 6053.0246 | 6121.8496 | 6227.8289 | 6139.5634 | 6304.9459 | 5837.2773 | 5937.4877 | 6229.9911 | 6090.2439 | 5969.9193 | 6105.0668 | 6247.5457 | 6089.1662 | 5747.5034 | 5881.3953 | 5452.4405 | 5627 | 4661.3677 | 4536.3677 | 4839.762 | 3841.7679 | 5811.0816 | 4361.0393 | 5218.645 | 5811.0816 | 5924.0875 | 3767.2891 | 6172.645 | 5730.7709 | 5933.9459 | 4265.8643 | 4593.9007 | 4992.7955 | 5007.2498 | 4720.6932 | 5351.7109 | 5012.1121 | 6141.765 | 6069.4946 | 5988.0275 | 5533.2684 | 5607.6843 | 5194.1603 | 5440.6627 | 5584.0216 | 5650.1023 | 5452.5152 | 5589.5389 | 5264.0698 | 5753.2498 | 3400.5723 | 3460.708 | 5731.9518
64 | 2.6984 | 1138.438 | 2720.7301 | 6007.2168 | 3119.0575 | 3087.8363 | 4198.4071 | 6389.615 | 3491.3805 | 5613.8407 | 3590.823 | 2799.4115 | 5692.6726 | 3319.4071 | 3164.3318 | 3677.3496 | 3534.292 | 3357.0752 | 4647.0708 | 4745.2389 | 5461.3363 | 5581.1416 | 4676.1725 | 4787.6239 | 4734.4956 | 3523.8938 | 3921.6195 | 4800.9602 | 3672.3407 | 3933.2212 | 4265.3318 | 5172.5531 | 4084.0089 | 5300.5619 | 3559.4912 | 3431.5133 | 3524.1416 | 3619.5354 | 3916.5929 | 3897.4956 | 3222.2611 | 4371.5266 | 3746.0973 | 4135.0575 | 4099.8628 | 3162.6195 | 3363.4513 | 3340.7478 | 3390.5354 | 6426.0442 | 3489.1991 | 6019.1903 | 5533.7257 | 6822.3805 | 3992.4026 | 6593.5664 | 6484.9867 | 5020.8938 | 3560.3982 | 6274.7301 | 6328.8938 | 6203.2434 | 3572.9292 | 6052.4071 | 5970.4247 | 2749.7655 | 3804.0442 | 6084.5354 | 6181.8628 | 6615.9779 | 3966.1991 | 6466.1327 | 6109.0797 | 5992.8849 | 6098.4734 | 6113.0663 | 6022.7611 | 6106.0885 | 6216.8451 | 6205.531 | 6003.1593 | 5929.354 | 6134.3142 | 6085.7345 | 6002.8495 | 5844.6814 | 5929.1239 | 6033.0309 | 5921.1283 | 5884.5752 | 6098.0398 | 6193.9336 | 6021.6726 | 5523.2876 | 5794.0265 | 5440.7832 | 5489.9425 | 4515.1062 | 4599.177 | 4771.0841 | 3813.2832 | 5712.115 | 4230.8495 | 5171.9336 | 5712.115 | 5756.9071 | 3708.7301 | 6175.4912 | 5533.8318 | 5820.0133 | 4165.7212 | 4576.1593 | 4906.2788 | 4817.1548 | 4603.6947 | 5298.7655 | 4833.0088 | 6074.9867 | 6040.1327 | 5940.0619 | 5394.8186 | 5449.3672 | 5123.6416 | 5416.9646 | 5440.8849 | 5458.3495 | 5299.3938 | 5529.823 | 5170.6681 | 5676.9823 | 3403.5487 | 3411.2832 | 5638.854
65 | 2.7094 | 1113.0669 | 2700.4267 | 5872.646 | 3123.7325 | 3150.8633 | 4108.8594 | 6237.2478 | 3459.7611 | 5587.2389 | 3543.1878 | 2905.7807 | 5540.2605 | 3518.6824 | 3232.6215 | 3790.9273 | 3425.2989 | 3330.8663 | 4583.9744 | 4656.0029 | 5235.3431 | 5523.2547 | 4681.5447 | 4917.4877 | 4695.707 | 3475.9616 | 3992.5703 | 4740.2291 | 3657.8663 | 3848.1239 | 4210.1209 | 5054.9262 | 3991.0187 | 5306.3982 | 3506.4965 | 3438.0128 | 3427.5791 | 3527.2861 | 3847.8977 | 3897.2517 | 3188.4553 | 4321.9233 | 3742.2802 | 4143.6657 | 4215.2203 | 3180.8663 | 3218.2605 | 3332.6942 | 3336.9616 | 6357.0285 | 3449.2291 | 5858.2478 | 5278.3048 | 6540.6135 | 3867.6716 | 6438.8466 | 6416.0157 | 4862.8338 | 3486.7512 | 6223.7483 | 6162.9675 | 6121.0029 | 3514.2291 | 5975.5762 | 6066.5772 | 2774.5447 | 3894.7139 | 6079.0767 | 6090.3373 | 6648.65 | 3930.1622 | 6340.6332 | 5948.1425 | 5911.296 | 5983.6716 | 6024.0954 | 5962.6686 | 5968.7129 | 6109.1367 | 5999.4356 | 5896.9518 | 5946.2232 | 6070.6558 | 5951.2861 | 6094.9912 | 5785.7099 | 5682.6519 | 5988.6529 | 5932.7168 | 5820.9361 | 5904.6902 | 6040.7542 | 5925.8977 | 5394.057 | 5587.5565 | 5302.3943 | 5372.1239 | 4516.7581 | 4545.3373 | 4716.3884 | 3762.6686 | 5581.4454 | 4152.9489 | 5134.3884 | 5581.4454 | 5644.35 | 3675.0541 | 5895.1612 | 5485.3727 | 5679.8722 | 4015.7925 | 4453.4326 | 4865.1878 | 4674.9774 | 4531.2291 | 5185.0983 | 4817.2389 | 5958.8181 | 5990.1495 | 5767.6617 | 5537.6313 | 5450.2517 | 5174.6372 | 5341.2035 | 5333.3373 | 5413.1593 | 5179.5506 | 5403.5251 | 5112.7355 | 5557.7512 | 3322.1907 | 3449.5575 | 5569.5221
66 | 2.7204 | 1190.294 | 2727.1632 | 5637.0747 | 3086.4562 | 3125.0595 | 4179.8594 | 6024.0314 | 3354.1352 | 5550.6573 | 3541.7935 | 2793.3958 | 5522.5708 | 3416.6185 | 3197.8854 | 3746.5049 | 3411.7773 | 3423.2778 | 4559.6249 | 4611.2389 | 5138.0595 | 5365.8033 | 4731.9459 | 4733.0909 | 4519.8358 | 3418.266 | 3957.0703 | 4650.3038 | 3721.5167 | 3838.9838 | 4139.4887 | 5029.7881 | 3910.8692 | 5237.706 | 3440.266 | 3436.2065 | 3401.5924 | 3474.826 | 3828.7935 | 3877.2389 | 3175.6032 | 4218.5157 | 3719.4292 | 4035.5536 | 4178.0973 | 3185.3968 | 3291.6475 | 3260.0919 | 3367.147 | 6124.1017 | 3474.3481 | 5774.4562 | 5241.5654 | 6426.2552 | 3875.3535 | 6310.3255 | 6371.0433 | 4737.7168 | 3605.0064 | 6073.5914 | 6066.0757 | 5860.5309 | 3542.5438 | 5833.7709 | 5801.6283 | 2785 | 3892.8476 | 5845.3407 | 5868.4715 | 6372.1991 | 3890.8368 | 6093.2866 | 5778.9611 | 5796.3363 | 5776.4444 | 5894.7601 | 5879.2984 | 5799.1839 | 6025.4887 | 5840.7492 | 5658.0855 | 5712.5418 | 5944.3525 | 5818.1352 | 5960.0477 | 5614.6076 | 5558.88 | 5863.972 | 5817.3795 | 5638.2271 | 5810.6519 | 5684.4272 | 5835.6898 | 5347.6357 | 5562.7611 | 5241.027 | 5334.2227 | 4508.6249 | 4596.7016 | 4704.6032 | 3827.0983 | 5332.2974 | 4164.1686 | 5098.6357 | 5332.2974 | 5527.733 | 3693.3697 | 5933.441 | 5473.0216 | 5684.3535 | 4109.4302 | 4389.0433 | 4755.9395 | 4633.027 | 4473.2606 | 4901.5472 | 4749.4995 | 5685.3353 | 5918.2768 | 5705.2227 | 5322.1558 | 5425.4454 | 5007.8466 | 5268.4725 | 5208.1406 | 5246.0314 | 5191.5546 | 5184.2596 | 4914.852 | 5435.5319 | 3369.5059 | 3410.2714 | 5408.6022
67 | 2.7314 | 1203.5103 | 2710.4749 | 5572.2094 | 3219.5841 | 3118.056 | 4140.292 | 5830.9262 | 3309.531 | 5322.5958 | 3547.8673 | 2783.7699 | 5255.4336 | 3382.4897 | 3139.8171 | 3739.9794 | 3282.9616 | 3410.3068 | 4426.0177 | 4510.7109 | 4970.4366 | 5284.1947 | 4571.646 | 4730.1062 | 4493.8466 | 3448.9941 | 3922.9086 | 4552.3127 | 3672.1032 | 3720.3186 | 4161.9381 | 4953.3127 | 3939.3717 | 5038.7227 | 3477.4838 | 3411.4749 | 3334.8584 | 3440.7965 | 3800.7965 | 3807.5369 | 3295.8702 | 4141.8377 | 3674.2507 | 3992.3982 | 4031.8289 | 3134.8938 | 3216.1652 | 3259.2242 | 3366.4543 | 6036.3628 | 3515.6372 | 5579.0442 | 5286.1268 | 6317.0324 | 3879.1327 | 6186.41 | 6234.8289 | 4705.4543 | 3556.6342 | 5831.9528 | 5897.8702 | 5702.4661 | 3479.1799 | 5752.5162 | 5761.7493 | 2710.3687 | 3800.4454 | 5887.4572 | 5693.0737 | 6099.2537 | 3849.118 | 5933.7315 | 5599.7374 | 5606.3245 | 5742.3569 | 5795.236 | 5754.7315 | 5704.9705 | 5764.8259 | 5670.2271 | 5722.6254 | 5649 | 5757.0029 | 5569.9174 | 5750.9469 | 5506.236 | 5513.9852 | 5609.7286 | 5664.7522 | 5561.9292 | 5678.5074 | 5627.7139 | 5551.3008 | 5217.6961 | 5479.1445 | 5016.9734 | 5290.2507 | 4598.9115 | 4627.9794 | 4817.5693 | 3796.5162 | 5212.9144 | 4208.9381 | 4950.0383 | 5212.9144 | 5506.2448 | 3699.8112 | 5711.525 | 5280.7994 | 5478.0295 | 4051.7669 | 4298.2566 | 4665.9144 | 4570.4041 | 4480.6578 | 4888.9174 | 4774.8053 | 5589.7404 | 5746.6253 | 5568.7728 | 5088.8908 | 5191.4837 | 4970.1121 | 4946.7404 | 5137.3835 | 5155.9498 | 4988.2743 | 5024.8643 | 4739.8289 | 5291.7522 | 3335.0826 | 3319.1239 | 5237.9616
68 | 2.7424 | 1145.1785 | 2681.942 | 5452.0629 | 3117.1224 | 3144.5285 | 4027.0034 | 5773.7655 | 3338.7635 | 5178.9444 | 3566.2935 | 2784.2355 | 5104.7089 | 3397.176 | 3116.824 | 3762.764 | 3340.9381 | 3305.6504 | 4434.7045 | 4461.001 | 4972.7045 | 5079.6539 | 4569.0575 | 4703.2955 | 4433.767 | 3449.9415 | 3897.4125 | 4519.177 | 3708.8805 | 3756.233 | 4151.236 | 4862.5319 | 3872.4144 | 4929.4729 | 3443.7075 | 3404.353 | 3219.3564 | 3451.2935 | 3729.767 | 3894.2906 | 3290.2365 | 4094.413 | 3701.234 | 3995.588 | 3984.7065 | 3109.0005 | 3244.704 | 3152.6509 | 3320.3545 | 5802.5373 | 3476.884 | 5438.0034 | 5156.6519 | 6141.6524 | 3952.9386 | 5857.4813 | 6015.5949 | 4649.4725 | 3427.4744 | 5656.3574 | 5717.8874 | 5589.4149 | 3466.941 | 5428.0698 | 5616.1819 | 2776.1736 | 3853.4681 | 5459.0747 | 5640.1185 | 5956.6504 | 3898.645 | 5924.6465 | 5538.648 | 5654.4676 | 5652.7679 | 5677.7094 | 5451.6568 | 5569.1219 | 5580.3574 | 5716.8796 | 5495.5383 | 5311.3647 | 5672.237 | 5594.5856 | 5647.9435 | 5213.3038 | 5342.2999 | 5485.5319 | 5479.7114 | 5335.7723 | 5510.8284 | 5586.884 | 5492.294 | 5079.4744 | 5338.5924 | 4888.5909 | 5164.0629 | 4481.8284 | 4646.6465 | 4846.4705 | 3710.0614 | 5187.2945 | 4086.2399 | 4959.822 | 5187.2945 | 5334.8884 | 3673.648 | 5592.061 | 5156.2969 | 5190.8318 | 3906.6514 | 4264.177 | 4615.5895 | 4554.5295 | 4485 | 4868.06 | 4659.0044 | 5418.9469 | 5467.8908 | 5508.354 | 5026.942 | 4980.0644 | 4696.9508 | 4778.6504 | 4922.7718 | 5019.0044 | 4841.8269 | 4871.3574 | 4648.2375 | 5040.0078 | 3360.881 | 3375.9386 | 5083.4749
69 | 2.7534 | 1134.6716 | 2716.2822 | 5386.0462 | 3128.118 | 3144.1495 | 4056.1475 | 5583.4542 | 3337.6421 | 4967.6037 | 3503.1367 | 2787.179 | 4975.6755 | 3394.9852 | 3120.6863 | 3644.1386 | 3293.1367 | 3246.5536 | 4303.4228 | 4438.2842 | 4680.9508 | 5061.4022 | 4583.6863 | 4621.4503 | 4309.6755 | 3419.5831 | 3882.9705 | 4384.7807 | 3586.2733 | 3734.061 | 4098.5241 | 4613.8141 | 3824.2399 | 4630.8308 | 3503.3707 | 3415.7011 | 3321.6972 | 3559.474 | 3654.1514 | 3831.585 | 3259.5978 | 4110.4022 | 3622.0777 | 3994.4926 | 3996.88 | 3078.4189 | 3225.9115 | 3209.236 | 3388.7286 | 5525.8289 | 3538.8781 | 5266.5575 | 4985.7512 | 5915.0963 | 3858.6755 | 5768.5978 | 5798.648 | 4620.4484 | 3416.5369 | 5572.1662 | 5608.4356 | 5442.1239 | 3417.0167 | 5242.3176 | 5424.1996 | 2759.7325 | 3854.7915 | 5447.2508 | 5451.1258 | 5687.7699 | 3902.6421 | 5554.4011 | 5466.9725 | 5389.6666 | 5398.1425 | 5596.6588 | 5330.1514 | 5448.5555 | 5575.5664 | 5539.2605 | 5439.3284 | 5203.8967 | 5350.2497 | 5418.7679 | 5530.3618 | 4983.4985 | 5178.2881 | 5449.1052 | 5255.4395 | 5146.0796 | 5380.705 | 5434.0796 | 5288.231 | 5061.8505 | 5207.6902 | 4785.9135 | 5017.8102 | 4345.0187 | 4628.2694 | 4824.4336 | 3849.4573 | 5083.0334 | 4021.2252 | 4712.412 | 5083.0334 | 5107.4395 | 3608.6588 | 5477.8692 | 5080.4946 | 5201.8486 | 3997.0108 | 4210.8525 | 4602.1495 | 4403.5575 | 4321.1848 | 4748.1681 | 4504.1239 | 5224.3638 | 5364.1809 | 5327.3195 | 5012.8063 | 4808.0944 | 4602.061 | 4701.1662 | 4771.1829 | 4911.764 | 4675.5722 | 4694.3933 | 4610.762 | 4848.2438 | 3262.0187 | 3461.3707 | 4996.823
70 | 2.7644 | 1142.6327 | 2739.2655 | 5108.3097 | 3197.91 | 3021.7242 | 4087.2021 | 5516.469 | 3323.2788 | 4803.6357 | 3419.6091 | 2723.5457 | 4786.9277 | 3385.5192 | 3080.4439 | 3592.8614 | 3316.4676 | 3322.9602 | 4320.9484 | 4352.736 | 4633.0251 | 4880.2699 | 4462.8761 | 4643.4676 | 4289.6962 | 3425.0251 | 3959.8717 | 4297.938 | 3593.9233 | 3643.8879 | 4093.2404 | 4529.8363 | 3741.6726 | 4545.3171 | 3494.9498 | 3367.9631 | 3361.2655 | 3522.3304 | 3835.3629 | 3795.4941 | 3205.8112 | 4115.8481 | 3574.469 | 3945.6593 | 3937.2419 | 3063.6711 | 3147.913 | 3197.0133 | 3281.8761 | 5491.177 | 3467.3436 | 5232.7596 | 4808.7374 | 5743.851 | 3870.0752 | 5664.7743 | 5540.9542 | 4689.087 | 3348.6091 | 5414.2817 | 5448.8259 | 5148.1578 | 3342.5457 | 5103.6607 | 5292.0782 | 2721.6209 | 3774.8245 | 5268.0929 | 5309.8377 | 5559.2802 | 3789.0531 | 5571.7449 | 5172.4365 | 5218.6607 | 5379.1386 | 5456.0914 | 5252.6976 | 5342.6475 | 5186.4896 | 5297.4601 | 5224.4351 | 5044.3451 | 5232.1785 | 5244.4587 | 5288.6504 | 4992.4041 | 4946.0678 | 5263.5103 | 5000.8908 | 5050.1534 | 5165.0929 | 5161.7522 | 4988.4867 | 4895.0796 | 5064.028 | 4778.3289 | 4828.6239 | 4405.4159 | 4724.9218 | 4732.6475 | 3765.5088 | 4862.0044 | 3999.5826 | 4564.4705 | 4862.0044 | 5073.3923 | 3517.2802 | 5308.7374 | 4943.4838 | 5108.8377 | 3902.1165 | 4117.8879 | 4487.3569 | 4350.2537 | 4233.8746 | 4774.3407 | 4445.9248 | 5140.4189 | 5221.6991 | 5136.1047 | 4804.0678 | 4697.8879 | 4484.7492 | 4539.8894 | 4689.7861 | 4661.9292 | 4507.9144 | 4465.7006 | 4351.5752 | 4767.4189 | 3293.5811 | 3361.3569 | 4795.8776
71 | 2.7754 | 1173.2527 | 2731.6755 | 5042.7925 | 3141.0305 | 3030.9735 | 4132.2979 | 5301.0688 | 3284.0039 | 4725.0757 | 3474.1092 | 2628.5437 | 4609.8761 | 3379.6981 | 3107.415 | 3681.7807 | 3329.883 | 3313.0492 | 4314.3736 | 4281.2153 | 4457.763 | 4750.6381 | 4550.2675 | 4557.94 | 4078.3481 | 3308.2684 | 3745.1406 | 4388.5241 | 3581.0718 | 3577.3776 | 4012.2645 | 4550.532 | 3711.6529 | 4458.2645 | 3406.6834 | 3327.1662 | 3351.9548 | 3419.0796 | 3855.3736 | 3715.4267 | 3169.468 | 3966.295 | 3611.1583 | 3944.7434 | 3956.9735 | 3002.6342 | 3185.7621 | 3216.0905 | 3276.3471 | 5334.6912 | 3375.4267 | 5025.5103 | 4729.5172 | 5498.8839 | 3862.234 | 5494.7138 | 5379.6873 | 4366.4542 | 3497.7503 | 5119.5634 | 5231.997 | 5106.883 | 3416.177 | 4903.1819 | 5059.7669 | 2778.8073 | 3766.1622 | 5103.0836 | 5051.8161 | 5394.4346 | 3830.9469 | 5464.9666 | 4944.763 | 4990.6499 | 5202.4611 | 5143.054 | 5126.5703 | 5108.4651 | 5056.6342 | 5045.9066 | 5168.7925 | 4911.174 | 5095.5251 | 5211.3963 | 5088.1101 | 4774.5821 | 4801.4533 | 5068.3894 | 5027.3884 | 4890.0157 | 5066.3323 | 5050.8908 | 4991.0413 | 4616.352 | 4777.0049 | 4699.1249 | 4981.3992 | 4273.7365 | 4682.7738 | 4690.3284 | 3752.7434 | 4827.7207 | 3935.0531 | 4512.585 | 4827.7207 | 4922.5742 | 3525.3697 | 5109.9252 | 4856.472 | 4958.3628 | 3837.9813 | 4179.5733 | 4405.9135 | 4276.3097 | 4129.3589 | 4483.5899 | 4324.7099 | 5031.1966 | 5063.7817 | 4851.6086 | 4649.1288 | 4639.9813 | 4292.0649 | 4357.2723 | 4635.1209 | 4566.6568 | 4356.5251 | 4364.6116 | 4176.0836 | 4647.8722 | 3258.2606 | 3436.5015 | 4595.0875
72 | 2.7864 | 1191.9626 | 2703.8201 | 4949.2429 | 3061.9321 | 3056.944 | 4049.1155 | 5294.9253 | 3198.5875 | 4577.2463 | 3471.0374 | 2780.8063 | 4427.175 | 3344.9661 | 3163.6146 | 3689.146 | 3226.2802 | 3272.3481 | 4156.6657 | 4226.6214 | 4327.7709 | 4447.0206 | 4510.4041 | 4594.4872 | 4042.2547 | 3382.3751 | 3725.4533 | 4249.6657 | 3576.2547 | 3603.9066 | 4003.3633 | 4288.8372 | 3689.8014 | 4387.4041 | 3297.0069 | 3360.3973 | 3320.3294 | 3410.1087 | 3814.6027 | 3767.5772 | 3144.056 | 3890.5128 | 3642.4346 | 3906.7301 | 3852.0442 | 3102.9558 | 3147.7488 | 3132.7335 | 3237.9848 | 5129.231 | 3413.6146 | 4870.6996 | 4695.6927 | 5281.4857 | 3726.4823 | 5278.0137 | 5073.7659 | 4346.9066 | 3402.6996 | 5071.6367 | 5137.2955 | 4977.1003 | 3469.2699 | 4640.9985 | 4850.467 | 2816.0526 | 3771.3633 | 4925.793 | 4892.5723 | 5246.9543 | 3790.3854 | 5165.7473 | 4850.5501 | 4840.0629 | 4898.1106 | 4894.4144 | 5057.3854 | 4995.3328 | 4991.1495 | 4874.8643 | 4845.732 | 4815.4602 | 4922.2837 | 4979.6165 | 4895.321 | 4729.2734 | 4709.8235 | 4865.7217 | 4816.3431 | 4762.1529 | 4824.7438 | 4905.6096 | 4850.647 | 4467.7896 | 4599.6096 | 4531.5383 | 4805.9238 | 4330.2852 | 4628.4941 | 4708.6893 | 3719.0934 | 4747.7148 | 3881.3668 | 4411.8795 | 4747.7148 | 4658.0172 | 3549.944 | 5074.146 | 4647.9051 | 4703.9985 | 3751.5875 | 4079.4262 | 4192.6691 | 4241.438 | 4106.146 | 4526.6298 | 4320.8353 | 4891.0816 | 4964.3328 | 4873.7045 | 4486.0103 | 4415.7065 | 4120.0103 | 4224.6622 | 4294.642 | 4442.8082 | 4214.6996 | 4184.703 | 4019.737 | 4602.9474 | 3341.7571 | 3390.0408 | 4462.4076
73 | 2.7974 | 1120.5487 | 2655.7522 | 4837.7699 | 3055.5398 | 3008.6726 | 3906.7964 | 5259.177 | 3217.9735 | 4395.938 | 3475.1947 | 2654.5575 | 4257.0973 | 3288.6726 | 3059.4956 | 3605.6991 | 3232.2389 | 3245.7611 | 4055.2389 | 4138.5841 | 4221.8495 | 4235.9734 | 4432.0885 | 4467.4867 | 3842 | 3398.6549 | 3733.3097 | 4232.3097 | 3483.9646 | 3586.8761 | 3763.292 | 4269.469 | 3617.0885 | 4212.6283 | 3407.7788 | 3312.9026 | 3314.2301 | 3454.3628 | 3864.7522 | 3850.4513 | 3214.4513 | 3884.7699 | 3562.8141 | 3854.0973 | 3775.823 | 3008.3805 | 3115.3363 | 3091.6814 | 3345.2744 | 4823.3539 | 3489.1416 | 4711.4071 | 4576.0708 | 5174.469 | 3728.2389 | 5104.8672 | 5022.2655 | 4136.9557 | 3359.3363 | 4755.0707 | 4882.8672 | 4716.407 | 3387.3894 | 4550.3628 | 4732.9646 | 2755.3982 | 3648.1062 | 4775.5664 | 4606.8584 | 5080.3274 | 3877.6903 | 5063.0531 | 4854.2743 | 4646.0088 | 4898.708 | 4891.5487 | 4858.8495 | 4727.3628 | 4791.8495 | 4822.2566 | 4752.9823 | 4597.8053 | 4892.0354 | 4669.0088 | 4723.2124 | 4472.2477 | 4511.3894 | 4686.0177 | 4577.8053 | 4500.9823 | 4624.6283 | 4694.8495 | 4682.9823 | 4525.4513 | 4526.0177 | 4411.6549 | 4697.2389 | 4338.4602 | 4588.3717 | 4819.9646 | 3690.1062 | 4652.8141 | 3825.4425 | 4302.4248 | 4652.8141 | 4538.0442 | 3530.8761 | 4749.9911 | 4525.1947 | 4725.7522 | 3739.115 | 4054.1504 | 4111.6726 | 3954.8141 | 4018.2743 | 4364.7787 | 4171.6725 | 4731.4071 | 4692.823 | 4738.7522 | 4387.1239 | 4294.6195 | 3916.6991 | 4124.9292 | 4219.4513 | 4206.5309 | 4001.4247 | 4148.4956 | 3796.4601 | 4377.761 | 3321.3717 | 3348.2566 | 4403.6372
74 | 2.8084 | 1148.6844 | 2772.0256 | 4625.8112 | 3024.5964 | 2966.1101 | 3974.3412 | 4929.1966 | 3122.4056 | 4252.9744 | 3359.3579 | 2632.3083 | 4159.0423 | 3304.4312 | 3053.0752 | 3606.6642 | 3302.2935 | 3234.0275 | 4035.1303 | 4157.6091 | 4052.9818 | 4166.4459 | 4479.5266 | 4613.2989 | 3847.4036 | 3320.0423 | 3667.6037 | 4137.8918 | 3456.7468 | 3557.9798 | 3838.7724 | 4225.3432 | 3606.2055 | 4094.0973 | 3346.8368 | 3289.8496 | 3256.5762 | 3462.6917 | 3828.9321 | 3760.4056 | 3213.233 | 3762.0772 | 3532.2606 | 3768.145 | 3691.734 | 2993.1028 | 3095.8496 | 3105.9449 | 3254.1249 | 4813.9174 | 3457.6239 | 4649.9321 | 4397.7011 | 4875.9837 | 3765.9926 | 4864.38 | 4873.8515 | 4160.1431 | 3359.589 | 4688.8771 | 4678.7286 | 4632.9597 | 3383.178 | 4441.6111 | 4479.6406 | 2692.2679 | 3738.7724 | 4530.202 | 4650.0605 | 4951.5359 | 3811.9395 | 4820.8185 | 4603.8736 | 4584.4184 | 4543.3067 | 4709.2699 | 4795.4184 | 4527.1121 | 4670.9946 | 4512.936 | 4588.1873 | 4411.29 | 4682.8387 | 4588.1377 | 4680.6716 | 4194.0516 | 4325.4956 | 4508.0845 | 4463.967 | 4338.8992 | 4454.9543 | 4640.3909 | 4512.3653 | 4232.6278 | 4424.7891 | 4349.6514 | 4510.2424 | 4320.9523 | 4548.1101 | 4655.4479 | 3769.5467 | 4428.4552 | 3737.3505 | 4200.2753 | 4428.4552 | 4447.145 | 3579.1431 | 4731.3284 | 4484.2881 | 4550.8589 | 3687.8569 | 3889.7011 | 4086.1578 | 4064.2513 | 3908.3028 | 4155.3726 | 4068.556 | 4503.0442 | 4512.7487 | 4517.2222 | 4344.6991 | 4186.0973 | 3803.2753 | 3930.0644 | 4044.6258 | 4120.5762 | 3833.3377 | 4002.5432 | 3860.6219 | 4256.5836 | 3245.2202 | 3385.2734 | 4286.7414
75 | 2.8194 | 1098.4277 | 2721.4228 | 4488.3982 | 3064.6578 | 2997.472 | 3937.6136 | 4922.0148 | 3178.3058 | 4158.0511 | 3389.6578 | 2760.2449 | 4029.941 | 3258.7945 | 3077.648 | 3623.4621 | 3305.4523 | 3285.7532 | 3919.3982 | 4008.1976 | 3967.6942 | 4080.0511 | 4254.9774 | 4495.5693 | 3855.5477 | 3260.9754 | 3685.4671 | 4190.6677 | 3514.12 | 3469.5034 | 3730.941 | 4082.6598 | 3572.3471 | 4002.9607 | 3336.1809 | 3265.2616 | 3280.9243 | 3362.9508 | 3811.2665 | 3737.176 | 3205.6332 | 3647.9508 | 3660.0639 | 3710.7945 | 3703.0098 | 2992.0905 | 3021.237 | 3258.0738 | 3242 | 4503.2094 | 3370.232 | 4461.2733 | 4294.413 | 4832.9951 | 3760.9046 | 4695.293 | 4550.4857 | 4100.5133 | 3439.1298 | 4437.6961 | 4589.8751 | 4491.4788 | 3306.6893 | 4345.413 | 4447.4523 | 2669.6283 | 3761.0049 | 4430.3274 | 4446.1681 | 4717.0727 | 3672.0265 | 4692.1317 | 4510.4228 | 4339.882 | 4471.0757 | 4452.4297 | 4542.6961 | 4445.2419 | 4522.6598 | 4382.4985 | 4475.408 | 4369.352 | 4390.6863 | 4455.6647 | 4361.4808 | 4206.0197 | 4156.0167 | 4384.4936 | 4391.2419 | 4221.2222 | 4383.4277 | 4526.9459 | 4255.7915 | 4127.9656 | 4228.4542 | 4144.8112 | 4498.1858 | 4342.648 | 4629.4966 | 4768.7021 | 3756.2616 | 4380.352 | 3687.0708 | 4119.056 | 4380.352 | 4405.2566 | 3411.2782 | 4690.3422 | 4339.1121 | 4389.0216 | 3653.8043 | 3859.176 | 3956.6647 | 3861.8013 | 3840.6086 | 4102.1662 | 3992.6037 | 4322.5595 | 4311.3638 | 4362.117 | 4177.2832 | 4004.0069 | 3720.237 | 3687.9823 | 3770.3294 | 3888.4346 | 3894.3107 | 3751.3392 | 3639.0678 | 4065.0019 | 3340.6873 | 3395.4572 | 4138.4788
76 | 2.8304 | 1134.4823 | 2714.0782 | 4341.4911 | 2971.2861 | 2946.6121 | 3917.3289 | 4846.6445 | 3058.7094 | 3993.2891 | 3465.9027 | 2733.3614 | 3882.4911 | 3233.7198 | 3106.733 | 3491.3348 | 3243.2375 | 3194.208 | 3943.9513 | 3927.5501 | 3798.7581 | 3969.677 | 4287.3097 | 4480.1563 | 3797.0029 | 3241.2507 | 3631.7684 | 4151.8304 | 3511.7035 | 3356.8333 | 3700.5634 | 4040.4233 | 3420.5855 | 3841.9764 | 3349.8274 | 3282.3584 | 3165.0841 | 3445.1829 | 3795.2507 | 3642.8953 | 3185.6416 | 3554.1136 | 3543.0221 | 3865.4764 | 3629.41 | 2984.6254 | 3067.4661 | 3114.4454 | 3070.5398 | 4399.1755 | 3409.0133 | 4301.0383 | 4238.9867 | 4575.7315 | 3725.8923 | 4615.5501 | 4496.0324 | 3997.208 | 3392.6121 | 4279.3348 | 4414.3053 | 4308.3835 | 3346.9351 | 4086.2625 | 4273.6961 | 2656.9381 | 3664.2861 | 4323.4425 | 4250.7743 | 4398.171 | 3702.326 | 4432.1843 | 4397.9897 | 4274.5177 | 4356.4587 | 4281.5074 | 4444.2537 | 4339.4425 | 4296.0575 | 4197.8525 | 4206.0442 | 4197.3835 | 4446.9646 | 4415.3451 | 4271.0029 | 3938.9823 | 4026.6932 | 4344.5015 | 4149.0118 | 4023.8687 | 4329.1431 | 4333.4778 | 4194.205 | 3901.6047 | 3909.4837 | 4016.2242 | 4524.8112 | 4124.6829 | 4669.3422 | 4792.1401 | 3725.4233 | 4212.9926 | 3720.795 | 3981.1003 | 4212.9926 | 4191.823 | 3471.4336 | 4451.8554 | 4315.7817 | 4207.2271 | 3609.5177 | 3821.1268 | 3950.4528 | 3838.1563 | 3912.4336 | 4063.9705 | 3916.8628 | 4229.5664 | 4354.7625 | 4181.2271 | 4064.208 | 3927.0811 | 3585.7876 | 3596.8628 | 3753.2965 | 3814.3776 | 3571.0634 | 3570.3672 | 3565.2994 | 3961.2699 | 3327.9543 | 3326.8628 | 3989.4911
77 | 2.8414 | 1050.1927 | 2765.3766 | 4206.6961 | 3033.059 | 2918.647 | 3831.9292 | 4605.5359 | 3082.5624 | 3777.2959 | 3379.5221 | 2687.0305 | 3721.5408 | 3195.5693 | 3130.4848 | 3516.4307 | 3153.9292 | 3200.9459 | 3812.4867 | 3883.7011 | 3707.468 | 3785.3854 | 4193.9174 | 4385.6539 | 3626.6607 | 3270.826 | 3656.8918 | 4171.353 | 3398.9056 | 3305.9646 | 3521.5172 | 3962.3363 | 3448.1672 | 3655.2537 | 3189.3972 | 3220.4798 | 3057.8515 | 3501.5742 | 3798.6706 | 3705.059 | 3211.0895 | 3608.6519 | 3503.0423 | 3823.8869 | 3574.4258 | 2948.5034 | 3111.233 | 3051.4916 | 3096.1672 | 4246.3431 | 3330.7974 | 4076.9547 | 4042.5713 | 4392.3972 | 3707.12 | 4380.2065 | 4309.0442 | 3970.3176 | 3459.9272 | 4049.4159 | 4350.4916 | 4305.2094 | 3280.941 | 3967.8515 | 4140.3667 | 2645.7247 | 3661.6047 | 4273.8987 | 4024.8889 | 4391.6824 | 3701.3294 | 4437.1554 | 4283.4444 | 4100.9901 | 4323.647 | 4257.7906 | 4421.1858 | 4058.2369 | 4099.5831 | 4154.9764 | 4145.7011 | 3986.8348 | 4296.1337 | 4106.6086 | 4249.1858 | 3868.3599 | 3923.588 | 4111.2723 | 4037.4562 | 4073.3884 | 4067.1829 | 4157.3313 | 4162.9764 | 3946.5202 | 3979.2803 | 4049.9695 | 4370.0678 | 4275.0757 | 4619.6893 | 4827.5624 | 3724.8682 | 4014.9783 | 3709.0541 | 3874.7197 | 4014.9783 | 4164.6588 | 3509.7601 | 4314.3009 | 4058.117 | 4112.6657 | 3505.115 | 3690.0255 | 3852.9174 | 3806.2399 | 3666.7876 | 3919.0137 | 3720.4395 | 4017.4936 | 4169.4395 | 4238.9272 | 3873.8466 | 3683.7994 | 3543.2399 | 3466.762 | 3607.3431 | 3642.7925 | 3570.0895 | 3522.4376 | 3362.9665 | 3816.6843 | 3252.2045 | 3396.6637 | 3832.2773
78 | 2.8524 | 1093.2134 | 2655.4307 | 4182.8574 | 3041.7856 | 2936.8746 | 3984.9971 | 4412.0211 | 2931.6814 | 3753.5359 | 3392.8746 | 2708.6067 | 3666.5187 | 3288.5875 | 3041.1445 | 3656.2296 | 3078.912 | 3169.822 | 3698.8948 | 3842.1971 | 3496.6288 | 3641.4488 | 4450.1023 | 4281.5197 | 3501.4307 | 3298.5531 | 3694.0172 | 3915.5762 | 3537.6937 | 3371.1416 | 3618.7837 | 3849.1091 | 3368.6622 | 3525.645 | 3248.9808 | 3417.5856 | 3183.4976 | 3424.5551 | 3800.8928 | 3719.4641 | 3276.4095 | 3417.0929 | 3579.5521 | 3737.6622 | 3500.0192 | 2875.0192 | 3128.1962 | 3046.1072 | 3124.5531 | 4032.5216 | 3354.3387 | 4037.3933 | 3903.1273 | 4139.6652 | 3619.8766 | 4381.0354 | 4139.8422 | 3722.9331 | 3336.2881 | 3949.7694 | 4138.7895 | 4046.6833 | 3273.1608 | 3913.6794 | 4132.9464 | 2669.6067 | 3640.2325 | 4122.4134 | 4032.7139 | 4176.2718 | 3702.8928 | 4413.4469 | 4086.7178 | 3965.0201 | 4130.8249 | 4217.2507 | 4061.3102 | 3981.2689 | 4040.6259 | 4039.4307 | 4067.2871 | 3902.9656 | 3993.5054 | 4088.7856 | 4044.5216 | 3788.3407 | 3774.0919 | 3936.0029 | 3967.5905 | 3943.6096 | 3766.9695 | 4075.9656 | 3960.3608 | 3788.2173 | 3753.1829 | 3926.1809 | 4260.5732 | 4220.7512 | 4505.2699 | 4906.6593 | 3712.6967 | 3900.734 | 3527.2355 | 3777.1445 | 3900.734 | 4060.8412 | 3344.4852 | 4117.6642 | 3960.5015 | 3899.736 | 3428.912 | 3635.5723 | 3790.1617 | 3623.2355 | 3746.4803 | 3791.645 | 3790.2842 | 3903.6268 | 3979.9675 | 3887.203 | 3746.4843 | 3804.2296 | 3295.8795 | 3399.3225 | 3531.6976 | 3662.2316 | 3505.7512 | 3499.179 | 3316.822 | 3681.1273 | 3309.4631 | 3330.0728 | 3762.3761
79 | 2.8634 | 1030.8082 | 2715.1505 | 4048.7404 | 2928.8731 | 2929.3717 | 3977.7021 | 4416.3805 | 3010.2419 | 3490.3156 | 3370.9911 | 2694.6608 | 3476.0944 | 3181.8318 | 3063.8024 | 3554.7079 | 3140.1918 | 3050.0796 | 3800.3746 | 3797.8171 | 3445.6932 | 3531.4189 | 4243.5398 | 4171.5015 | 3476.7847 | 3267.4041 | 3736.2655 | 3899.0826 | 3527.7847 | 3236.9469 | 3453.3923 | 3821.9086 | 3260.4602 | 3608.1593 | 3289.4307 | 3245.8879 | 3209.2153 | 3426.7109 | 3719.5103 | 3777.5221 | 3153.4513 | 3514.3333 | 3512.0973 | 3629.4602 | 3429.8495 | 2842.2389 | 3042.7168 | 2989.4366 | 3187.3569 | 3896.3687 | 3393.3894 | 3909.4513 | 3884.4956 | 4105.8673 | 3697.4484 | 4277.5015 | 4089.7758 | 3734.885 | 3372.1416 | 3957.4631 | 4129.8761 | 3915.0383 | 3336.2743 | 3818.0059 | 3842.5132 | 2699.7611 | 3561.3451 | 3997.997 | 3862.3923 | 4076.7581 | 3666.5693 | 4203.0029 | 3913.0206 | 3819.9469 | 4116.1652 | 3990.6637 | 4068.8024 | 3820.7728 | 3852.0118 | 3828.7964 | 3934.5339 | 3777.2448 | 3931.941 | 3953.823 | 3856.5988 | 3634.4425 | 3737.2389 | 3701.705 | 3755.0029 | 3794.1121 | 3997.3865 | 3948.3274 | 3736.1681 | 3696.5516 | 3700.6932 | 3776.1534 | 4255 | 4108.6667 | 4531.5959 | 4871.6932 | 3788.6962 | 3866.2389 | 3527.5457 | 3754.7434 | 3866.2389 | 3822.9941 | 3355.1327 | 4041.8082 | 3825.6165 | 3876.4956 | 3421.1652 | 3665.5959 | 3675.7905 | 3540.1386 | 3630.2448 | 3685.295 | 3535.1917 | 3876.9086 | 3866.4602 | 3920.6873 | 3791.7699 | 3426.6017 | 3184.2124 | 3154.407 | 3398.5752 | 3357.2153 | 3165.4955 | 3373.3687 | 3156.8141 | 3577.1711 | 3186.4513 | 3364.3481 | 3608.4838
80 | 2.8744 | 1141.6062 | 2575.264 | 3978.9144 | 2928.855 | 2936.7679 | 3868.2488 | 4261.6411 | 3046.0428 | 3435.3053 | 3346.6962 | 2589.1327 | 3425.2763 | 3249.8245 | 3094.1298 | 3525.7252 | 3084.7404 | 3164.4322 | 3679.7419 | 3782.4061 | 3372.1465 | 3565.8397 | 4140.7846 | 4323.3299 | 3381.8137 | 3211.6244 | 3713.7252 | 3850.4503 | 3399.7999 | 3264.0138 | 3518.8245 | 3702.5388 | 3295.8107 | 3299.3402 | 3205.5526 | 3212.7542 | 3202.2321 | 3370.6824 | 3832.5192 | 3732.4213 | 3119.8412 | 3298.5845 | 3436.1755 | 3758.0261 | 3448.3619 | 2888.5787 | 2900.148 | 3033.6077 | 3105.5236 | 3823.0305 | 3345.4793 | 3870.9862 | 3716.9892 | 4006.9297 | 3591.2778 | 4058.4105 | 3954.9739 | 3681.6244 | 3369.145 | 3862.8427 | 3883.3677 | 3721.6273 | 3257.4366 | 3685.8289 | 3695.0319 | 2687.4061 | 3399.7281 | 3820.1632 | 3788.9725 | 3863.2655 | 3587.3205 | 3945.4257 | 3871.0442 | 3813.029 | 3854.8028 | 4012.0718 | 3897.1632 | 3844.1008 | 3876.0428 | 3700.5678 | 3825.1037 | 3609.8441 | 3752.3082 | 3855.7556 | 3809.1893 | 3480.409 | 3586.467 | 3762.8825 | 3695.5084 | 3734.5954 | 3684.6592 | 3808.0319 | 3567.7571 | 3478.4105 | 3566.9449 | 3688.4366 | 4019.0486 | 4069.0442 | 4586.4046 | 4728.235 | 3658.0029 | 3688.2792 | 3431.8717 | 3653.9877 | 3688.2792 | 3914.9543 | 3273.4656 | 3880.7281 | 3720.9587 | 3742.0029 | 3361.7694 | 3516.496 | 3633.1313 | 3556.4199 | 3505.5678 | 3578.0747 | 3486.1313 | 3668.2074 | 3647.4105 | 3895.8122 | 3693.0167 | 3481.8535 | 3184.855 | 3109.0442 | 3443.4916 | 3435.2734 | 3266.5482 | 3152.4395 | 3034.4518 | 3555.5221 | 3248.9695 | 3317.4503 | 3509.7556
81 | 2.8854 | 1031.9223 | 2665.3953 | 3829.5074 | 2905.4012 | 2889.8024 | 3847.3009 | 4284.532 | 2918.1898 | 3300.2733 | 3242.7551 | 2651.88 | 3283.3903 | 3244.1337 | 3037.9695 | 3558.4317 | 3149.8633 | 3124.7355 | 3520.6411 | 3659.0895 | 3327.7188 | 3417.5408 | 4108.5015 | 4334.883 | 3376.0501 | 3214.9164 | 3671.7021 | 3823.4346 | 3449.0974 | 3155.8554 | 3497.4012 | 3636.0698 | 3351.0806 | 3266.4012 | 3116.471 | 3203.1504 | 3118.4208 | 3389.649 | 3773.0698 | 3629.7217 | 3203.5625 | 3256.6018 | 3470.3815 | 3514.2124 | 3423.4346 | 2913.5988 | 2979.5792 | 2879.6411 | 3021.3874 | 3666.6244 | 3199.4572 | 3805.1032 | 3601.8722 | 3812.2596 | 3646.0138 | 3985.1367 | 3845.7886 | 3673.117 | 3317.8692 | 3655.4769 | 3846.5015 | 3728.7827 | 3237.3009 | 3619.0698 | 3687.0669 | 2711.5821 | 3435.3147 | 3642.9252 | 3741.7689 | 3855.0334 | 3527.9695 | 4094.3422 | 3764.7886 | 3670.4238 | 3852.8997 | 3736.6971 | 3823.1701 | 3684.7247 | 3671.4936 | 3675.3677 | 3908.529 | 3447.6745 | 3644.7384 | 3720.2566 | 3613.3097 | 3447.4848 | 3534.8024 | 3639.1396 | 3607.471 | 3666.1367 | 3643.5349 | 3737.1367 | 3739.9912 | 3532.9636 | 3458.7719 | 3599.471 | 3942.1868 | 4101.4317 | 4510.3205 | 4729.8997 | 3663.8328 | 3679.0669 | 3464.3982 | 3534.9725 | 3679.0669 | 3667.2625 | 3338.8466 | 3794.3874 | 3691.4513 | 3553.1425 | 3378.6824 | 3451.0364 | 3484.5074 | 3400.6578 | 3240.4464 | 3471.7551 | 3377.822 | 3562.6883 | 3720.646 | 3674.7443 | 3494.3264 | 3312.9085 | 3018.825 | 3022.4543 | 3282.7581 | 3359.3373 | 3060.5604 | 3200.0806 | 3115.5625 | 3448.8053 | 3147.7551 | 3284.5349 | 3334.9252
82 | 2.8964 | 1058.8717 | 2690.8938 | 3965.385 | 3007.5443 | 2930.8142 | 3788.2611 | 4234.2257 | 2960.8009 | 3183.5088 | 3329.6062 | 2577.3363 | 3236.1991 | 3153.4026 | 3055.9159 | 3583.8894 | 2937.9469 | 3102.0973 | 3442.3363 | 3597.2655 | 3127.8849 | 3297.5221 | 4024.3717 | 4233.4513 | 3292.3717 | 3087.5664 | 3696.885 | 3696.5619 | 3389.2699 | 3226.677 | 3424.323 | 3541.4159 | 3346.0265 | 3199.2965 | 3245.4204 | 3177.115 | 3077.177 | 3255.5973 | 3720.2301 | 3694.7035 | 3167.1681 | 3382.4381 | 3451.0885 | 3534.8894 | 3274.6593 | 2858.2478 | 3037.7478 | 3003.4381 | 3052.854 | 3574.3982 | 3339.3673 | 3508.3141 | 3524.3363 | 3930.4602 | 3570.3407 | 3814.7522 | 3682.7168 | 3427.0929 | 3183.5929 | 3477.7743 | 3723.5442 | 3621.4779 | 3275.8274 | 3586.1416 | 3679.0354 | 2757.7965 | 3616.1991 | 3682.8097 | 3515.0044 | 3687.7433 | 3557.8628 | 3910.8274 | 3588.7743 | 3665.0133 | 3798.2434 | 3671.2699 | 3798.1062 | 3546.6017 | 3537.5796 | 3598.3407 | 3703.9159 | 3563.4735 | 3593.2212 | 3583.5973 | 3451.7035 | 3370.3407 | 3552.9159 | 3500.6062 | 3390.9557 | 3460.907 | 3558.3761 | 3580.6902 | 3629.5044 | 3329.907 | 3519.7212 | 3575.1018 | 4034.5841 | 4108.969 | 4687.208 | 4782.7655 | 3799.3983 | 3556.5442 | 3425.177 | 3471.2743 | 3556.5442 | 3568.4203 | 3257.3672 | 3726.2965 | 3469.9823 | 3499.2257 | 3312.2965 | 3368.3628 | 3497.9292 | 3324.3274 | 3385.3363 | 3456.0619 | 3408.854 | 3490.3141 | 3479.0796 | 3580.4026 | 3418.323 | 3286.1062 | 3003.0575 | 2994.1195 | 3216.2832 | 3275.3672 | 3123.7035 | 3070.5796 | 2920.8717 | 3307.6194 | 3249.5398 | 3273.0487 | 3342.9513
83 | 2.9074 | 1080.7325 | 2659.3864 | 3645.8348 | 2935.8594 | 2904.1495 | 3894.0324 | 3964.883 | 3043.2999 | 3248.702 | 3325.0796 | 2585.9685 | 3098.1957 | 3233.9616 | 2948.0708 | 3521.3383 | 2999.8043 | 3186.7168 | 3461.6775 | 3548.0639 | 3119.1898 | 3281.8751 | 4076.5271 | 4229.5909 | 3185.0875 | 3155.1819 | 3639.4415 | 3657.1268 | 3427.9371 | 3026.4169 | 3349.9528 | 3564.6391 | 3137.0226 | 3141.1819 | 3153.0944 | 3182.1337 | 3106.3078 | 3357.6067 | 3636.2675 | 3637.5192 | 3208.3619 | 3306.2615 | 3546.0324 | 3478.3383 | 3208.9528 | 2677.534 | 2956.236 | 2944.5192 | 2883.0541 | 3426.763 | 3249.5664 | 3674.3874 | 3524.7807 | 3576.4494 | 3617.1731 | 3769.056 | 3605.1583 | 3541.7886 | 3257.9676 | 3517.3235 | 3650.292 | 3476.6784 | 3219.1337 | 3397.8358 | 3373.472 | 2667.0865 | 3469.4563 | 3491.2055 | 3488.7964 | 3709.2301 | 3603.4887 | 3800.1976 | 3632.5447 | 3565.5831 | 3728.8663 | 3707.764 | 3732.4258 | 3400.5192 | 3553.6539 | 3541.0482 | 3561.0236 | 3373.0954 | 3464.9921 | 3545.2448 | 3442.7168 | 3399.7797 | 3316.2429 | 3421.6539 | 3468.0088 | 3506.4346 | 3587.7895 | 3557.5526 | 3411.6529 | 3279.8112 | 3295.2203 | 3408.236 | 3994.3726 | 4218.9538 | 4441.7699 | 4857.0315 | 3691.0148 | 3558.0885 | 3101.1721 | 3386.1426 | 3558.0885 | 3413.3304 | 3154.5585 | 3638.5752 | 3502.528 | 3553.1435 | 3223.4405 | 3311.5909 | 3354.3628 | 3313.4818 | 3319.4887 | 3392.1898 | 3328.174 | 3408.8358 | 3333.9292 | 3452.0718 | 3433.7886 | 3280.1583 | 2861.5437 | 2922.8663 | 3109.6539 | 3224.4572 | 2863.9213 | 2996.5123 | 2929.9213 | 3271.3235 | 3276.0393 | 3339.7247 | 3284.0157
84 | 2.9184 | 1047.262 | 2583.5118 | 3622.2586 | 2795.0841 | 2923.4031 | 3773.8948 | 3822.3471 | 2952.4198 | 3083.2866 | 3208.4557 | 2582.5241 | 3244.9282 | 3101.7502 | 2957.0403 | 3562.851 | 2849.8869 | 3149.6091 | 3419.9921 | 3429.1534 | 3015.3269 | 3137.5206 | 4141.3672 | 4175.4803 | 3069.4759 | 3055.205 | 3568.9798 | 3548.3707 | 3419.234 | 3078.4759 | 3344.2419 | 3392.7178 | 3265.234 | 3036.1052 | 3163.085 | 3197.9115 | 3021.0157 | 3309.5118 | 3670.7861 | 3697.9203 | 3253.5767 | 3087.7055 | 3423.738 | 3521.5846 | 3200.9036 | 2868.4803 | 2985.3146 | 3008.3068 | 3015.2139 | 3518.3235 | 3318.702 | 3468.7827 | 3371.9676 | 3508.532 | 3640.6328 | 3649.0123 | 3497.6091 | 3392.1043 | 3379.7581 | 3430.4076 | 3557.3717 | 3312.1165 | 3234.206 | 3253.6608 | 3429.6942 | 2677.4154 | 3526.0123 | 3390.2945 | 3335.7301 | 3483.6013 | 3539.6652 | 3695.4926 | 3448.1735 | 3507.4365 | 3709.1175 | 3674.2709 | 3657.2384 | 3463.8028 | 3506.8353 | 3420.3953 | 3561.057 | 3258.8387 | 3464.3225 | 3448.6293 | 3406.4557 | 3369.7424 | 3469.8554 | 3375.9843 | 3340.6617 | 3308.5364 | 3309.6249 | 3375.0202 | 3489.8879 | 3257.028 | 3322.1534 | 3436.972 | 3664.8948 | 4087.766 | 4662.5408 | 4859.7099 | 3817.4479 | 3318.4513 | 3248.0123 | 3279.532 | 3318.4513 | 3453.2665 | 3226.4312 | 3580.9483 | 3439.5329 | 3318.0929 | 3277.2261 | 3259.8387 | 3272.6696 | 3123.6328 | 3252.2183 | 3290.5846 | 3222.4479 | 3263.6529 | 3322.1376 | 3335.5644 | 3282.8746 | 3172.5083 | 2803.4154 | 2901.0614 | 3063.3795 | 3152.412 | 2853.4961 | 2934.3751 | 2917.6013 | 3184.8309 | 3278.8432 | 3304.6976 | 3285.4557
85 | 2.9294 | 1060.6814 | 2654.3717 | 3485.82 | 2944.059 | 2939.4956 | 3864.3127 | 3741.6755 | 2829.2389 | 2996.9174 | 3181.8053 | 2606.6372 | 3046.8702 | 3043.3038 | 2958.2861 | 3507.4867 | 2894.7522 | 3079.9882 | 3366.9764 | 3385.8908 | 2935.0059 | 3016.1888 | 4061.9351 | 4053.6165 | 3041.0472 | 3208.7286 | 3565.6637 | 3609.7758 | 3379.7611 | 3048.9115 | 3351.5546 | 3373.5486 | 3223.3923 | 2943.0265 | 3071.2655 | 3144.3186 | 3117.7021 | 3382.059 | 3708.1239 | 3460.7021 | 3230.3068 | 3118.0088 | 3416.7109 | 3546.5929 | 3099.4218 | 2744.3451 | 2948.4631 | 2925.6755 | 2975.9941 | 3288.6313 | 3205.9941 | 3363.2714 | 3414.9115 | 3577.1593 | 3713.8672 | 3601.8112 | 3431.7522 | 3537.1268 | 3368.8437 | 3347.7758 | 3380.0265 | 3407.764 | 3144.6962 | 3376.5752 | 3269.1622 | 2649.2596 | 3477.5457 | 3353.3717 | 3276.9853 | 3364.8908 | 3568.6578 | 3527.0767 | 3404.9823 | 3242.6785 | 3568.7404 | 3492.8702 | 3524.0384 | 3353.6578 | 3347.5162 | 3387.9823 | 3415.9557 | 3250.6106 | 3473.997 | 3381.708 | 3374.7109 | 3257.9557 | 3314.0147 | 3276.9941 | 3223.528 | 3285.0265 | 3332.6991 | 3374.2802 | 3291.0059 | 3299.3746 | 3276.295 | 3381.0944 | 3753.528 | 4120.6519 | 4539.5929 | 4793.3333 | 3745.6991 | 3383.1917 | 3199.059 | 3349.3274 | 3383.1917 | 3365.4631 | 3287.9852 | 3384.1947 | 3270.3835 | 3270.4307 | 3185.4837 | 3247.9764 | 3206.5074 | 3203.7463 | 3118.7522 | 3317.767 | 3148.7847 | 3338.0619 | 3183.5841 | 3360.2743 | 3365.174 | 3094.3864 | 2773.9616 | 2659.7345 | 2935.5457 | 2978.1947 | 2827.3127 | 2900.9174 | 2799.3127 | 3125.9233 | 3174.7139 | 3288.8967 | 3256.7286
86 | 2.9404 | 1080.6868 | 2579.9371 | 3392.5211 | 2912.2866 | 2838.4897 | 3816.1922 | 3724.9311 | 2893.0472 | 2904.1947 | 3169.7812 | 2560.7207 | 3038.7945 | 3109.4051 | 2936.5079 | 3522.2237 | 2948.5659 | 3020.4027 | 3248.7001 | 3323.4474 | 2957.1706 | 2996.6082 | 4057.8972 | 4075.5344 | 2953.6711 | 3140.6421 | 3499.8923 | 3454.2684 | 3316.3107 | 2930.4897 | 3214.526 | 3211.9843 | 3085.5079 | 2907.3785 | 3107.2999 | 3122.7473 | 3073.1839 | 3355.5211 | 3682.0895 | 3735.6386 | 3148.1209 | 3000.5659 | 3588.208 | 3444.1052 | 3109.5841 | 2694.6421 | 2844.5211 | 2913.8078 | 2944.6578 | 3219.1632 | 3176.0447 | 3424.6022 | 3289.4208 | 3404.8343 | 3534.4897 | 3621.9287 | 3477.8259 | 3361.1499 | 3383.7994 | 3196.7871 | 3347.2419 | 3193.5211 | 3078.237 | 3186.8343 | 3268.7207 | 2654.0919 | 3462.0315 | 3279.6082 | 3360.4523 | 3301.0157 | 3469.5816 | 3587.971 | 3313.1657 | 3408.1971 | 3560.5079 | 3481.8864 | 3548.468 | 3251.9867 | 3189.0108 | 3291.0654 | 3349.1391 | 3232.1101 | 3295.0895 | 3366.3785 | 3219.7291 | 3192.4813 | 3221.7787 | 3246.0605 | 3119.5973 | 3249.7812 | 3118.6686 | 3321.1657 | 3317.1971 | 3071.8658 | 3130.0944 | 3209.474 | 3650.5816 | 3991.8948 | 4440.8633 | 4684.0447 | 3848.5285 | 3210.5369 | 3168.4498 | 3209.8658 | 3210.5369 | 3260.7606 | 3054.7714 | 3382.5285 | 3253.7896 | 3165.7896 | 3027.7109 | 3129.0738 | 3157.1681 | 3218.8864 | 3166.884 | 3276.3918 | 3195.7365 | 3156.5708 | 3199.6578 | 3272.5128 | 3104.058 | 2929.6844 | 2746.266 | 2626.2055 | 2922.5369 | 2981.0024 | 2774.0919 | 2810.3422 | 2747.971 | 3059.5708 | 3198.7739 | 3254.0762 | 3187.9262
87 | 2.9514 | 1071.6047 | 2660.4641 | 3423.118 | 2832.2586 | 2912.6578 | 3852.4493 | 3583.0679 | 2826.4641 | 2884.1485 | 3039.4523 | 2454.6726 | 2944.0383 | 3057.1937 | 2920.8102 | 3524.7188 | 2974.3009 | 3033.3343 | 3291.7866 | 3326.8407 | 2796.7296 | 3085.4759 | 4005.0727 | 4161.3225 | 2969.7679 | 3134.4828 | 3575.3009 | 3456.9774 | 3384.0344 | 3005.1711 | 3121.2242 | 3225.472 | 3074.7719 | 2767.5398 | 3072.7109 | 3090.9813 | 2953.6765 | 3384.4641 | 3679.3805 | 3433.1632 | 3106.175 | 3029.5516 | 3604.4759 | 3420.4375 | 3062.9046 | 2693.6156 | 2879.5093 | 2863.5782 | 2915.0875 | 3235.5595 | 3178.825 | 3285.1485 | 3211.6883 | 3434.5899 | 3598.9056 | 3522.7384 | 3338.8712 | 3338.6086 | 3171.646 | 3289.5329 | 3266.7148 | 3222.7493 | 3129.8211 | 3218.9852 | 3146.9892 | 2595.764 | 3416.293 | 3370.1298 | 3240.5742 | 3279.943 | 3486.6765 | 3559.7266 | 3269.0954 | 3323.0079 | 3539.0501 | 3331.1976 | 3403.5821 | 3200.6735 | 3291.8938 | 3144.939 | 3228.5054 | 3084.2016 | 3291.6391 | 3221.9774 | 3272.5821 | 3097.0266 | 3149.8633 | 3223.4336 | 3106.8368 | 3122.646 | 3202.4179 | 3279.4336 | 3239.9921 | 3065.0649 | 3008.7335 | 3297.7001 | 3659.35 | 4016.9882 | 4505.3648 | 4680.2665 | 3698.9804 | 3215.6696 | 3114.9695 | 3192.3431 | 3215.6696 | 3269.1868 | 3177.7719 | 3219.8063 | 3285.2704 | 3197.4415 | 3166.8368 | 3150.175 | 3058.7335 | 3062.5516 | 3073.6391 | 3172.2399 | 3062.5133 | 3055.8407 | 3166.4946 | 3124.5329 | 3149.5969 | 2984.1485 | 2802.883 | 2715.4867 | 2926.4149 | 2906.6204 | 2713.2547 | 2831.643 | 2719.4415 | 2960.0383 | 3256.0039 | 3220.2429 | 3161.293
88 | 2.9624 | 1034.7035 | 2602.9705 | 3383.4734 | 2824.09 | 2902.3554 | 3684.5856 | 3722.2551 | 2767.2375 | 2839.4956 | 3085.6844 | 2565.2773 | 2853.4734 | 3108.4749 | 2842.4174 | 3489.0103 | 2851.2979 | 3098.2758 | 3389.3038 | 3227.0797 | 2899.8894 | 2926.1976 | 3881.2095 | 4312.9365 | 2995.6032 | 3158.1681 | 3447.792 | 3405.8909 | 3331.3953 | 2998.6416 | 3148.2256 | 3109.9425 | 3069.1283 | 2790.0192 | 3100.4454 | 3012.1593 | 3029.059 | 3335.1505 | 3614.4381 | 3501.1681 | 3075.1681 | 2950.177 | 3439.9292 | 3347.8422 | 3138.4543 | 2796.3451 | 2913.9292 | 2807.3466 | 2898.6829 | 3099.9513 | 3207.3363 | 3232.177 | 3179.3245 | 3274.0782 | 3430.1283 | 3325.823 | 3228.2095 | 3255.1858 | 3265.9012 | 3130.9897 | 3260.4749 | 3163.3156 | 3186.969 | 3109.7242 | 3005.5561 | 2650.0973 | 3345.4336 | 3247.0708 | 3220.4661 | 3272.0678 | 3558.6622 | 3456.4159 | 3182.5929 | 3249.7212 | 3359.5251 | 3350.7227 | 3308.7316 | 3067.7832 | 3222.6829 | 3177 | 3287.8879 | 3066.5634 | 3203.9705 | 3171.792 | 3184.6637 | 3109.3554 | 3087.4631 | 3170.0472 | 3038.7625 | 3209.8997 | 3283.3348 | 3236.1681 | 3285.6887 | 3003.5354 | 3075.146 | 3171.0383 | 3593.8009 | 4106.4336 | 4565.3673 | 4610.2493 | 3857.6224 | 3106.91 | 3125.2463 | 2964.4882 | 3106.91 | 3138.059 | 3167.3451 | 3171.1593 | 3167.6224 | 3084.4941 | 3099.8024 | 3116.8805 | 3120.848 | 2995.9322 | 2985 | 3077.8319 | 3062.82 | 2962.5959 | 3090.6829 | 2973.8923 | 3164.9779 | 2936.2375 | 2587.8024 | 2699.9292 | 2800.3761 | 2740.823 | 2724.1563 | 2705.2861 | 2608.9203 | 2881.674 | 3225.5649 | 3332.5826 | 3099.2375
89 | 2.9734 | 1068.5516 | 2627.589 | 3269.3265 | 2875.001 | 2838.1229 | 3774.1426 | 3571.7158 | 2762.06 | 2872.6735 | 3155.2261 | 2566.0619 | 2742.1829 | 3086.6126 | 2941.6735 | 3512.0826 | 2912.999 | 3019.9971 | 3192.3265 | 3246.4307 | 2785.5319 | 2895.3078 | 3972.7768 | 4047.0432 | 2936.3707 | 3139.3284 | 3445.6529 | 3393.2035 | 3288.5497 | 2869.4907 | 3021.1623 | 3219.1455 | 3037.6342 | 2742.1652 | 3094.061 | 3069.469 | 2995.41 | 3449.7571 | 3762.2055 | 3487.6332 | 3062.8574 | 2867.9194 | 3341.2458 | 3371.4729 | 3067.1013 | 2705.7758 | 2818.2655 | 2817.1416 | 2886.06 | 3148.6745 | 3180.4297 | 3155.9587 | 3035.6735 | 3204.0826 | 3414.4671 | 3396.9813 | 3309.7384 | 3114.9784 | 3278.6529 | 3098.6952 | 3226.2074 | 3087.8161 | 3101.4297 | 3188.3294 | 3097.4513 | 2546.9381 | 3226.1406 | 3328.7787 | 3249.5349 | 3173.8584 | 3487.6529 | 3466.2439 | 3109.9184 | 3154.9381 | 3374.7768 | 3335.3294 | 3233.0216 | 3121.3903 | 3194.0639 | 3175.7758 | 3093.0797 | 3010.9174 | 3205.3294 | 3168.2655 | 3192.7994 | 3037.6145 | 2948.6932 | 3016.1445 | 3047.6126 | 3126.1229 | 3157.7778 | 3217.5732 | 3014.469 | 3045.5319 | 2919.4277 | 3092.5506 | 3519.6745 | 3990.9764 | 4610.3707 | 4716.5929 | 3766.8584 | 3049.7148 | 3047.1839 | 3142.2694 | 3049.7148 | 3104.7365 | 3080.2852 | 3222.1848 | 3073.349 | 3032.9774 | 3123.3471 | 3059.4936 | 2915.5074 | 3076.4316 | 2976.5123 | 3075.7787 | 2953.0816 | 3072.0423 | 3072.1839 | 3050.47 | 3012.4297 | 2926.8161 | 2606.4503 | 2604.1023 | 2788.2458 | 2811.3284 | 2584.5732 | 2640.1229 | 2577.7355 | 2927.6539 | 3256.409 | 3235.6116 | 3094.5497
90 | 2.9844 | 1037.056 | 2741.1912 | 3263.7974 | 2819.295 | 2811.7414 | 3790.1927 | 3501.3756 | 2821.2537 | 2857.7315 | 3079.0939 | 2523.4218 | 2767.7645 | 3063.4877 | 2937.6244 | 3471.8501 | 2954.1499 | 3134.1534 | 3194.4975 | 3151.972 | 2725.5964 | 2823.2866 | 3911.1829 | 3947.9587 | 2788.9159 | 3058.1647 | 3440.7925 | 3400.44 | 3347.056 | 2833.0049 | 3070.5354 | 3098.0329 | 2966.6573 | 2660.7611 | 3096.1401 | 3084.2158 | 2909.7183 | 3363.2916 | 3695.028 | 3463.8171 | 3060.1416 | 2825.2866 | 3322.4759 | 3212.7891 | 3096.4597 | 2689.089 | 2800.8913 | 2865.8913 | 2934.5487 | 3083.1912 | 3133.5074 | 3170.1121 | 3027.8682 | 3150.3279 | 3517.6195 | 3309.7365 | 3132.4597 | 3170.9061 | 3285.6244 | 3036.8402 | 3075.9769 | 3095.2158 | 3043.4267 | 3047.5157 | 2987.6524 | 2586.9489 | 3323.4597 | 3166.6657 | 3053.3756 | 3109.4267 | 3489.6475 | 3493.0595 | 3117.9209 | 3189.6903 | 3308.5816 | 3204.4695 | 3182.6342 | 3003.3147 | 3078.5123 | 3141.529 | 3183.7183 | 3053.9671 | 3078.883 | 3143.0973 | 3037.1318 | 2929.8402 | 2980.5025 | 2946.4184 | 3017.1632 | 3094.2011 | 3051.1829 | 3126.9572 | 3025.9292 | 2976.6755 | 2957.7183 | 3100.4218 | 3473.0659 | 4119.44 | 4464.3245 | 4672.6342 | 3709.7512 | 3014.4218 | 3022.2242 | 2953.061 | 3014.4218 | 3025.1598 | 3090.8815 | 3156.0742 | 2979.1927 | 3107.8947 | 3108.0182 | 2899.4612 | 3031.9474 | 2936.6903 | 2890.3904 | 2922.6804 | 2954.5634 | 2993.0742 | 3052.295 | 3006.412 | 2961.3756 | 2918.6278 | 2550.1977 | 2593.6244 | 2739.9902 | 2729.1318 | 2505.883 | 2611.206 | 2539.8732 | 2886.7134 | 3213.5585 | 3260.6804 | 3146.8584
91 | 2.9954 | 999.4248 | 2582.1858 | 3180.0974 | 2730.6549 | 2762.6991 | 3814.9469 | 3501.177 | 2828.5575 | 2744.5398 | 3070.5487 | 2434.5044 | 2670.1681 | 3002.2567 | 2931.2389 | 3499.6549 | 2908.5044 | 2943.6195 | 3167.5752 | 3160.6991 | 2744.292 | 2846.5044 | 3881.9823 | 4059.3009 | 2828.6902 | 3104.3894 | 3340.3098 | 3243.8496 | 3309.3186 | 2861.469 | 3093.4071 | 3106.8319 | 2961.1062 | 2713.8938 | 3001.531 | 3070.6106 | 2904.7699 | 3419.4336 | 3674.8319 | 3486.4336 | 3146.3186 | 2862.4159 | 3532.3274 | 3255.6018 | 3035.5664 | 2645.4956 | 2717.885 | 2776.4071 | 2856.292 | 2922.1681 | 3171.6903 | 3083.4867 | 2984.3894 | 3057.4248 | 3481.4956 | 3226.708 | 3078.469 | 3152.7964 | 3270.4248 | 3017.1681 | 3119.115 | 3066.3186 | 2994.0177 | 2947.823 | 2965.6372 | 2642.7433 | 3275.5221 | 3041.823 | 3046.8053 | 3052.2655 | 3429.1858 | 3298.1151 | 3189.0265 | 3190.8141 | 3233.8142 | 3213.1416 | 3229.8849 | 3015.2566 | 3144.5221 | 2955.5752 | 3178.2389 | 3030.7256 | 3116.2301 | 2981.3805 | 3065 | 2909 | 2985.9734 | 3000.6902 | 2873.3628 | 2996.2655 | 3018.5841 | 3063.885 | 2972.646 | 2910.885 | 2956.2743 | 3009.5929 | 3486.2035 | 3956.1593 | 4525.9734 | 4708.3097 | 3736.0089 | 2926.9912 | 2873.0797 | 2938.5841 | 2926.9912 | 3034.2124 | 2953.8142 | 3024.115 | 2994.7079 | 2862.3717 | 3024.5221 | 3032.3097 | 2911.5221 | 2941.8495 | 2974.3097 | 2886.0974 | 2945.8761 | 2866.1062 | 2938.0619 | 2854.9381 | 2958.8407 | 2751.531 | 2594.4956 | 2580.0177 | 2686.7788 | 2757 | 2542.8053 | 2528.9911 | 2520.0796 | 2854.0442 | 3192.9735 | 3213.5841 | 3046.5575
92 | 3.0064 | 1004.7773 | 2609.9479 | 3208.9002 | 2852.1042 | 2777.1185 | 3763.5261 | 3456.6775 | 2711.621 | 2699.1799 | 2989.3412 | 2473.4789 | 2670.2842 | 3088.469 | 2905.8673 | 3423.2512 | 2833.5639 | 3031.2227 | 3076.7444 | 3083.5974 | 2750.791 | 2770.0236 | 3845.6868 | 3937.1278 | 2685.9621 | 3067.0093 | 3469.2935 | 3335.5305 | 3298.4926 | 2851.3127 | 3075.2606 | 3090.0093 | 2848.6775 | 2622.5452 | 3008.1327 | 3077.0658 | 2877.469 | 3373.1799 | 3677.915 | 3444.9194 | 3054.6352 | 2900.5354 | 3389.8673 | 3234.5831 | 2973.3555 | 2525.4268 | 2752.9572 | 2757.6681 | 2877.4833 | 2927.9715 | 3117.4026 | 3002.1092 | 2944.0472 | 3055.1942 | 3451.4076 | 3196.7488 | 3073.763 | 3021.5074 | 3165.4169 | 3011.1229 | 3075.0379 | 2964.8009 | 3096.3599 | 2880.3033 | 2967.2085 | 2551.7915 | 3327.4504 | 3075.1849 | 2971.4268 | 3091.0187 | 3357.4882 | 3353.616 | 3064.8958 | 3063.237 | 3205.4597 | 3232.7153 | 3174.384 | 2996.4405 | 3072.4641 | 3048.853 | 3148.8387 | 2912.1278 | 3033.763 | 2990.5118 | 2966.1327 | 2905.3269 | 2911.3033 | 2856.8958 | 2863.0236 | 3042.1278 | 2988.7586 | 3090.3077 | 2934.9336 | 2929.2507 | 2964.5118 | 3013.0472 | 3422.2463 | 4002.7866 | 4462.915 | 4565.0998 | 3876.2129 | 2957.5069 | 2896.7488 | 2896.0806 | 2957.5069 | 2927.2984 | 3025.3933 | 2887.4504 | 2868.6067 | 2913.3933 | 2965.853 | 2966.2842 | 2949.706 | 2854.2798 | 2911.7202 | 2923.1942 | 2822.379 | 2772.6352 | 2855.3604 | 2850.5025 | 2900.7202 | 2762.3982 | 2467.0806 | 2500.7394 | 2729.8437 | 2660.2655 | 2465.057 | 2463.3648 | 2439.9051 | 2839.0093 | 3215.8673 | 3175.2276 | 3017.8245
93 | 3.0174 | 1018.6539 | 2605.6224 | 3094.1917 | 2823.6145 | 2790.8161 | 3752.938 | 3352.9685 | 2725.0757 | 2636.6844 | 2963.7915 | 2530.53 | 2603.1003 | 3022.646 | 2881.6303 | 3428.6224 | 2809.7001 | 3018.9154 | 3164.4838 | 3145.9764 | 2603.4622 | 2735.354 | 3722.2704 | 3876.1917 | 2702.6381 | 3027.293 | 3337.6765 | 3221.5005 | 3203.1003 | 2802.1996 | 3025.353 | 3025.9685 | 2765.3697 | 2686.4838 | 3021.646 | 2902.1386 | 2826.8161 | 3307.4149 | 3757.6529 | 3423.4385 | 3104.1308 | 2835.7148 | 3362.4916 | 3286.8151 | 2962.53 | 2593.8456 | 2710.8771 | 2710.0226 | 2839.3314 | 2889.7611 | 3041.8151 | 3043.884 | 2864.5074 | 3015.4002 | 3445.5615 | 3176.3382 | 3063.1072 | 3111.7994 | 3123.2547 | 2891.1691 | 3040.707 | 2941.7847 | 2973.4612 | 2894.1849 | 2935.7306 | 2593.6008 | 3349.9528 | 3071.3146 | 3040.1072 | 2983.469 | 3382.2773 | 3315.6381 | 3069.3225 | 3071.4081 | 3196.7463 | 3106.707 | 3217.0698 | 2906.9233 | 2933.2379 | 3013.8151 | 3113.8545 | 2854.5919 | 3024.5457 | 2961.9302 | 2981.6077 | 2857.0757 | 2910.1455 | 2867.469 | 2828.6155 | 2981.5841 | 3084.9449 | 2971.061 | 2915.6224 | 2761.0462 | 2933.707 | 2934.1613 | 3370.0383 | 3953.6686 | 4549.061 | 4681.7699 | 3728.8692 | 2868.3382 | 2882.4848 | 2909.061 | 2868.3382 | 2859.7994 | 2996.4533 | 2934.6696 | 2884.9921 | 2882.7532 | 2931.5998 | 2901.5851 | 2868.3835 | 2882.5604 | 2864.2153 | 2864.3068 | 2790.8771 | 2816.6381 | 2909.3068 | 2874.9764 | 2851.2537 | 2804.2458 | 2480.3225 | 2528.938 | 2738.707 | 2677.0914 | 2518.7227 | 2409.292 | 2411.8083 | 2783.7306 | 3188.6224 | 3225.9921 | 2947.5841
94 | 3.0284 | 1022.1416 | 2535.2935 | 3014.1209 | 2722.8496 | 2859.261 | 3697.8451 | 3267.1342 | 2603.3112 | 2611.4336 | 2897.1741 | 2439.9985 | 2628.1298 | 3004.4248 | 2820.5634 | 3366.0192 | 2826.5797 | 3004.6843 | 3062.5634 | 3076.2714 | 2663.9867 | 2737.4085 | 3843.5516 | 3829.7419 | 2664.7109 | 3094.6932 | 3307.885 | 3253.1062 | 3242.0074 | 2744.6962 | 2939.3024 | 2933.1239 | 2824.7198 | 2597.4395 | 3003.1386 | 2931.441 | 2905.4189 | 3271.736 | 3644.851 | 3432.4277 | 3106.8599 | 2756.2596 | 3290.292 | 3245.6947 | 2868.851 | 2527.9926 | 2721.559 | 2659.003 | 2899.8451 | 2844.8717 | 3025.8746 | 2954.8348 | 2815.5531 | 3044.5811 | 3419.5472 | 3146.7065 | 2999.9941 | 3031.5531 | 3213.5649 | 2885.4381 | 2962.2906 | 2936.1268 | 2948.0266 | 2913.8451 | 2918.292 | 2665.6873 | 3225.708 | 2988.5737 | 2974.8481 | 2951.8451 | 3390.2359 | 3264.1195 | 3001.2802 | 3097.261 | 3202.9823 | 3054.7581 | 3273.413 | 2919.6947 | 2879.4351 | 2981.851 | 3149.6607 | 2820.5841 | 2965.2714 | 2884.413 | 2965.9808 | 2721.7198 | 2839.5737 | 2839.1357 | 2831.8363 | 2916.7109 | 2934.1268 | 2896.4159 | 2847.2965 | 2764.1372 | 2817.2345 | 2907.5797 | 3353.5664 | 3864.6843 | 4477.5634 | 4736.2507 | 3728.2861 | 2831.1239 | 2905.5811 | 2841.2832 | 2831.1239 | 2807.3083 | 2925.2743 | 2955.6799 | 2865.2861 | 2811.7198 | 2889.1135 | 2935.8333 | 2738.7153 | 2756.8525 | 2853.1342 | 2768.0118 | 2809.7139 | 2794.8776 | 2807.5752 | 2816.7168 | 2829.5664 | 2750.5487 | 2418.2891 | 2467.1209 | 2644.2669 | 2592.736 | 2465.5546 | 2381.2906 | 2470.705 | 2749.41 | 3111.4263 | 3191.4071 | 2889.8643
95 | 3.0394 | 1013.4267 | 2563.5821 | 2947.061 | 2701.7876 | 2756.2035 | 3661.2734 | 3240.5447 | 2681.2104 | 2617.2409 | 2991.2192 | 2437.2625 | 2570.9607 | 2981.8535 | 2794.2144 | 3431.4159 | 2844.7748 | 2895.4287 | 3035.9036 | 3018.9322 | 2596.4838 | 2661.2822 | 3733.7925 | 3938.5821 | 2666.8466 | 3031.5271 | 3404.8427 | 3117.3609 | 3271.2232 | 2673.175 | 3021.7591 | 2879.6146 | 2837.9892 | 2641.2232 | 2970.9302 | 2976.5624 | 2871.1268 | 3350.9892 | 3611.766 | 3434.7483 | 3109.1445 | 2669.1092 | 3297.0462 | 3169.4179 | 2848.9715 | 2507.0108 | 2670.3786 | 2675.9715 | 2842.2999 | 2905.8643 | 3104.768 | 2870.3697 | 2752.0305 | 3070.3284 | 3325.3501 | 3114.5339 | 2968.232 | 2962.9036 | 3177.2124 | 2911.2606 | 2968.4838 | 2861.2124 | 3045.1249 | 2864.3786 | 2924.4248 | 2556.1554 | 3216.1573 | 3004.3697 | 2924.173 | 2908.1268 | 3207.5752 | 3191.9931 | 2974.6499 | 2990.0088 | 3132.5536 | 3198.9764 | 3203.4749 | 2841.3982 | 2896.2124 | 2942.5516 | 2932 | 2879.3019 | 2908.7089 | 2830.0108 | 2884.2714 | 2727.998 | 2834.3962 | 2802.6981 | 2735.9695 | 2899.2714 | 2887.4877 | 2859.7591 | 2887.2714 | 2743.4857 | 2634.6342 | 2932.9213 | 3338.4376 | 3754.6519 | 4470.7896 | 4612.4877 | 3736.2144 | 2773.4199 | 2908.1799 | 2844.0502 | 2773.4199 | 2884.2409 | 2898.1092 | 2823.6126 | 2854.2606 | 2830.8928 | 2766.7375 | 2833.0964 | 2750.9233 | 2734.116 | 2808.2124 | 2811.1839 | 2792.4641 | 2856.7158 | 2812.2802 | 2821.7178 | 2805.7375 | 2656.1357 | 2423.1445 | 2381.5733 | 2568.6214 | 2663.2891 | 2398.2429 | 2391.057 | 2423.6785 | 2687.2055 | 3118.6234 | 3115.6519 | 2925.6804
96 | 3.0504 | 1006.5738 | 2555.8973 | 3053.4203 | 2743.5998 | 2728.5939 | 3725.4852 | 3283.647 | 2551.8973 | 2535.8264 | 2840.6623 | 2455.503 | 2552.6883 | 2934.9528 | 2823.6411 | 3367.5266 | 2745.9056 | 2952.7733 | 3057.1027 | 3012.4321 | 2532.5561 | 2686.647 | 3588.2444 | 3911.9528 | 2728.6907 | 3133.2114 | 3326.7237 | 3202.2385 | 3251.2409 | 2690.3648 | 3070.3141 | 3009.5089 | 2781.791 | 2614.8702 | 2905.5974 | 2937.6116 | 2884.5821 | 3301.1617 | 3580.5797 | 3456.1086 | 3059.6588 | 2763.9469 | 3217.471 | 3196.0472 | 2863.5939 | 2564.7178 | 2679.2385 | 2708.8678 | 2797.8501 | 2884.1465 | 3131.7438 | 2897.9587 | 2801.1027 | 3001.1972 | 3343.6352 | 3095.4026 | 2916.7792 | 2986.2822 | 3138.9209 | 2837.9764 | 2889.6352 | 2798.7792 | 2974.7674 | 2863.5148 | 2812.8028 | 2552.4734 | 3277.6789 | 3015.6966 | 2854.8678 | 2920.0732 | 3319.0295 | 3298.0555 | 2933.3967 | 2943.2822 | 3139.9941 | 3027.0767 | 3143.0201 | 2841.1204 | 2841.9941 | 2872.2149 | 2937.7001 | 2894.6907 | 2932.4179 | 2871.6234 | 2864.3353 | 2665.8122 | 2757.0118 | 2771.2149 | 2679.1854 | 2854.8584 | 2952.1229 | 2935.4734 | 2864.6529 | 2768.4002 | 2771.4852 | 2882.1263 | 3353.4144 | 3834.0944 | 4600.941 | 4705.791 | 3762.2881 | 2847.3141 | 2717.8914 | 2890.8584 | 2847.3141 | 2833.5064 | 2973.1642 | 2865.503 | 2777.1204 | 2811.3176 | 2748.3294 | 2791.7792 | 2794.503 | 2729.0944 | 2760.3648 | 2774.1027 | 2710.8501 | 2739.6588 | 2739.0791 | 2785.1204 | 2781.6293 | 2644.1558 | 2377.8501 | 2393.1204 | 2567.3058 | 2560.9646 | 2392.7615 | 2339.8501 | 2379.3707 | 2750.4498 | 3189.0496 | 3180.0059 | 2955.4616
97 | 3.0614 | 1064.8614 | 2598.1239 | 2986.3599 | 2700.3156 | 2687.5133 | 3698.4248 | 3268.4395 | 2605.2802 | 2524.4218 | 2969.0029 | 2436.0472 | 2552.646 | 2892.6608 | 2820.8938 | 3393.5988 | 2692.4514 | 2902.3599 | 2995.4425 | 2955.8997 | 2580.5398 | 2653.8053 | 3668.3953 | 3875.4189 | 2613.3982 | 3014.705 | 3361.1445 | 3246.4454 | 3172.1003 | 2697.2065 | 2908.8614 | 2999.1858 | 2749.4189 | 2490.9646 | 2976.8436 | 2894.2389 | 2816.059 | 3287.1032 | 3637.7434 | 3396.7434 | 3040.4041 | 2723.8584 | 3327.1622 | 3224.3038 | 2805.9587 | 2586.9469 | 2713.5486 | 2718.7463 | 2800.1239 | 2729.059 | 3033.5074 | 2856.882 | 2760.1357 | 3000.9085 | 3302.7846 | 3023.2566 | 2883.9558 | 2944.0796 | 3200.764 | 2808.6431 | 2847.5251 | 2757.1386 | 2913.7817 | 2865.9233 | 2788.2596 | 2527.4808 | 3164.6342 | 2889.2035 | 2873.5634 | 2850.3569 | 3344.1799 | 3191.764 | 2852.118 | 2913.4956 | 3125.6431 | 3061.9498 | 3034.8348 | 2792.1032 | 2838.0796 | 2906.882 | 2957.3628 | 2775.7994 | 2977.9085 | 2830.1416 | 2846.2065 | 2763.0885 | 2746.4867 | 2800.3038 | 2697.1593 | 2755.6932 | 2777.4189 | 2885.1091 | 2840.764 | 2805.0265 | 2757.7404 | 2874.6313 | 3291.0767 | 3891.5074 | 4536.6047 | 4659.5634 | 3727.059 | 2680.6431 | 2768.4395 | 2776.7581 | 2680.6431 | 2949.5899 | 2839.5369 | 2859.9026 | 2736.0206 | 2767.5899 | 2748.059 | 2753.3776 | 2785.3038 | 2779.1298 | 2779.8024 | 2719.4602 | 2705.3864 | 2714.2861 | 2674.177 | 2734.3038 | 2752.8791 | 2650.4749 | 2371.3068 | 2365.6962 | 2545.2596 | 2534.7788 | 2353.7729 | 2358.6785 | 2382.9292 | 2708.1799 | 3091.2596 | 3166.2861 | 2909.5044
98 | 3.0724 | 1043.501 | 2557.6485 | 2996.0738 | 2754.2719 | 2763.6991 | 3644.1106 | 3262.8894 | 2604.3068 | 2475.5477 | 2916.4661 | 2341.9592 | 2458.8117 | 2890.6505 | 2888.9302 | 3390.5845 | 2814.3456 | 2903.2311 | 2937.3554 | 2965.767 | 2572.6195 | 2570.5806 | 3716.1224 | 3877.1593 | 2602.6293 | 2952.4971 | 3312.9941 | 3149.8545 | 3176.6505 | 2609.1146 | 2894.6136 | 2905.6971 | 2760.3864 | 2422.1942 | 2911.7748 | 2910.5379 | 2828.2699 | 3239.501 | 3604.4563 | 3359.0349 | 3018.563 | 2740.0408 | 3294.4621 | 3161.0039 | 2808.6214 | 2496.5747 | 2605.6873 | 2633.1146 | 2740.233 | 2726.468 | 2966.563 | 2838.2271 | 2788.0894 | 2910.4504 | 3264.1436 | 3061.5418 | 2930.6971 | 2923.8176 | 3142.3087 | 2772.3437 | 2791.8019 | 2765.6195 | 2895.1903 | 2812.3107 | 2764.0118 | 2509.6563 | 3230.1981 | 2865.5669 | 2822.9243 | 2882.0428 | 3326.9302 | 3146.1883 | 2903.0679 | 2975.5767 | 3093.1883 | 2937.3029 | 3088.7379 | 2751.998 | 2841.6544 | 2883.998 | 2906.8505 | 2767.0836 | 2859.7689 | 2802.8097 | 2765.9961 | 2661.7242 | 2644.9262 | 2730.2758 | 2687.2001 | 2829.8156 | 2789.7321 | 2768.9572 | 2783.6544 | 2710.5398 | 2730.0855 | 2957.5536 | 3324.5049 | 3804.1863 | 4461.7729 | 4634.4912 | 3736.1942 | 2654.532 | 2765.9651 | 2791.8505 | 2654.532 | 2800.8117 | 2861.47 | 2833.5728 | 2736.9961 | 2646.3009 | 2747.3904 | 2772.967 | 2733.1534 | 2646.2679 | 2718.3574 | 2700.2758 | 2605.235 | 2622.8486 | 2716.5767 | 2688.6893 | 2759.2719 | 2703.5069 | 2267.1962 | 2418.8078 | 2531.969 | 2547.002 | 2437.0447 | 2375.002 | 2264.8038 | 2700.5418 | 3093.9243 | 3100.031 | 2852.2291
99 | 3.0834 | 1034.6362 | 2587.7984 | 2999.526 | 2725.7955 | 2709.6116 | 3682.115 | 3221.8171 | 2597.6381 | 2405.6333 | 2861.6598 | 2346.9567 | 2431.3392 | 2923.1141 | 2797.8653 | 3325.1573 | 2805.2743 | 2898.3638 | 2854.9754 | 2973.8427 | 2531.8633 | 2541.8397 | 3634.8643 | 3799.3658 | 2477.5202 | 2958.6391 | 3353.2979 | 3081.4523 | 3192.1386 | 2656.2459 | 2907.8643 | 2835.3884 | 2726.6834 | 2401.591 | 2859.7719 | 2916.591 | 2802.7974 | 3243.5212 | 3668.1583 | 3393.9322 | 3153.6834 | 2705.1829 | 3292.4995 | 3115.1141 | 2750.5447 | 2499.7522 | 2645.2969 | 2644.2498 | 2733.7935 | 2660.2724 | 3092.2753 | 2856.7522 | 2660.9538 | 2983.9823 | 3338.0039 | 3017.8191 | 2853.0255 | 2840.0885 | 3167.2006 | 2795.7276 | 2834.1612 | 2717.705 | 2922.9312 | 2771.2055 | 2649.6578 | 2496.0895 | 3160.0236 | 2964.41 | 2812.4543 | 2816.3658 | 3235.0472 | 3157.9125 | 2991.41 | 2952.0255 | 3102.0491 | 2986.7502 | 3038.588 | 2781.3855 | 2819.0246 | 2895.5919 | 2870.6352 | 2696 | 2845.002 | 2780.1819 | 2797.1603 | 2723.8191 | 2601.1829 | 2665.7267 | 2620.9754 | 2729.4562 | 2763.9331 | 2829.5212 | 2757.9567 | 2701.7267 | 2669.4966 | 2811.5438 | 3246.6391 | 3851.6381 | 4423 | 4706.5024 | 3731.3392 | 2712.3422 | 2719.6145 | 2743.3638 | 2712.3422 | 2735.0934 | 2795.7473 | 2864.5015 | 2758.7522 | 2686.4808 | 2693.1583 | 2732.4071 | 2699.9587 | 2654.2733 | 2602.9086 | 2639.2035 | 2558.7729 | 2573.0708 | 2707.1376 | 2712.2055 | 2714.0246 | 2623.4553 | 2246.0433 | 2403.2743 | 2462.2262 | 2514.3422 | 2346.0708 | 2373.6342 | 2292.5015 | 2660.0914 | 3062.3884 | 3163.4336 | 2852.411
100 | 3.0944 | 1032.3274 | 2512.354 | 2884.469 | 2715.8938 | 2757.9336 | 3638.2876 | 3245.0398 | 2539.854 | 2481.2256 | 2835.6991 | 2272.469 | 2466.2832 | 2917.9735 | 2761.7655 | 3367.4026 | 2743.0133 | 2888.9646 | 2893.6814 | 2915.5531 | 2526.0885 | 2590.0708 | 3618.1549 | 3719.0973 | 2552.0885 | 2875.8274 | 3251.7743 | 3130.8628 | 3092.1372 | 2751.8805 | 2886.6018 | 2773.4071 | 2661.0265 | 2401.0841 | 2890 | 2900.0398 | 2749.562 | 3294.146 | 3670.1991 | 3380.6991 | 3103.4204 | 2673.7212 | 3298.5088 | 3080.0531 | 2774.9159 | 2416.0133 | 2621.2699 | 2644.9779 | 2770.0442 | 2675.385 | 2996.9646 | 2784.2876 | 2687.7832 | 2808.0354 | 3217.7301 | 2971.6903 | 2780.1859 | 2896.2964 | 3268.9734 | 2745.0575 | 2748.3274 | 2701.8673 | 2915.4292 | 2730.7124 | 2682.4248 | 2534.208 | 3155.2345 | 2935.1195 | 2820.0487 | 2747.3717 | 3185.1416 | 3052.177 | 2935.0177 | 2863.5752 | 2991.031 | 2995.7788 | 3124.0044 | 2805.9779 | 2759.7212 | 2832.9823 | 2904.4115 | 2703.8053 | 2781.2876 | 2797.1637 | 2758.0354 | 2666.6327 | 2579.9115 | 2683.6637 | 2663.8009 | 2670.0841 | 2725.2389 | 2873.4248 | 2686.7212 | 2682.3716 | 2772.2699 | 2859.8584 | 3153.6372 | 3773.2566 | 4420.8319 | 4625.9381 | 3794.0354 | 2674.8053 | 2669.6549 | 2731.5796 | 2674.8053 | 2662.4779 | 2886.5796 | 2793.1372 | 2684.7699 | 2584.7168 | 2728.8053 | 2777.9823 | 2550.9159 | 2620.7257 | 2620.4823 | 2666.3274 | 2560.7212 | 2507.5354 | 2651.4602 | 2671.7124 | 2663.177 | 2609.9248 | 2305.354 | 2344.9159 | 2487.4734 | 2468.3496 | 2243.2434 | 2410.1062 | 2242.8186 | 2623.6549 | 3024.6903 | 3097.2655 | 2764.9203
101 | 3.1054 | 998.0216 | 2623.6037 | 2887.8358 | 2679.7925 | 2693.9105 | 3633.9656 | 3166.4857 | 2524.114 | 2364.3717 | 2817.941 | 2287.2911 | 2385.4769 | 2945.0649 | 2827.2753 | 3306.7463 | 2725.2291 | 2889.6844 | 2812.8702 | 2901.9872 | 2474.6224 | 2591.8387 | 3636.4395 | 3703.6254 | 2522.3874 | 2929.3343 | 3224.1613 | 3013.294 | 3051.2635 | 2556.1052 | 2887.5231 | 2796.5664 | 2625.2291 | 2399.5015 | 2864.1111 | 2840.5015 | 2776.3431 | 3265.4395 | 3571.2724 | 3376.2291 | 3117.9597 | 2700.0619 | 3232.7148 | 3021.7463 | 2765.9597 | 2385.8574 | 2699.7365 | 2627.9135 | 2666.3501 | 2687.6657 | 2972.0246 | 2790.7335 | 2669.1298 | 2802.9626 | 3286.4798 | 2918.1672 | 2883.6037 | 2775.7119 | 3130.5202 | 2646.0315 | 2748.6165 | 2705 | 2812.7896 | 2695.8574 | 2641.1239 | 2492.353 | 3220.2163 | 2922.2782 | 2795.8171 | 2770.3343 | 3246.412 | 3090.2753 | 2852.9164 | 2916.3314 | 3054.3992 | 3036.8604 | 3035.4021 | 2738.2812 | 2757.4238 | 2738.0836 | 2861.6499 | 2723.9813 | 2804.9164 | 2829.6775 | 2756.6283 | 2572.4798 | 2644.8948 | 2640.0491 | 2576.1052 | 2697.827 | 2751.1052 | 2799.82 | 2712.3746 | 2559.0344 | 2703.2911 | 2837.3068 | 3177.8948 | 3717.7089 | 4421.9351 | 4662.2262 | 3794.232 | 2663.3648 | 2612.7276 | 2710.8545 | 2663.3648 | 2711.232 | 2847.0678 | 2698.4238 | 2678.6312 | 2651.5477 | 2748.294 | 2675.294 | 2559.1426 | 2565.3373 | 2605.8358 | 2635.9164 | 2590.4985 | 2604.7089 | 2631.4336 | 2662.5172 | 2709.7522 | 2606.0334 | 2237.5418 | 2332.7522 | 2436.1455 | 2465.5418 | 2218.9971 | 2264.8545 | 2250.2507 | 2568.7896 | 3138.0865 | 3101.9971 | 2716.6195
102 | 3.1164 | 995.2419 | 2605.5845 | 2833.5845 | 2669 | 2755.792 | 3568.588 | 3076.0339 | 2404.3451 | 2392.9646 | 2793.5531 | 2303.9317 | 2411.8274 | 2913.793 | 2762.0683 | 3332.4813 | 2726.7242 | 2902.6878 | 2879.7591 | 2825.8638 | 2480.205 | 2524.4828 | 3607.1377 | 3710.8284 | 2592.7561 | 2914.1032 | 3257.4803 | 2969.0334 | 3138.8963 | 2531.413 | 2835.1018 | 2756.9317 | 2627.207 | 2381.9656 | 2789.3112 | 2820.3102 | 2681.2434 | 3242.414 | 3571.0703 | 3355.0329 | 3092.5511 | 2684.7915 | 3175.8638 | 3030.1008 | 2742.7237 | 2359.001 | 2608.2414 | 2650.1367 | 2691.9292 | 2664.0319 | 2934.1731 | 2724.8958 | 2626.3805 | 2774.1052 | 3267.3786 | 2919.4833 | 2852.3077 | 2750.0703 | 3169.8987 | 2743.1013 | 2818.3077 | 2672.1701 | 2839.4828 | 2650.8963 | 2661.0349 | 2496.8613 | 3115.8963 | 2842.2768 | 2743.5521 | 2754.3791 | 3251.8928 | 3023.8613 | 2856 | 2916.3776 | 2983.2738 | 2978.0703 | 2957.207 | 2639.8643 | 2663.4135 | 2725.5855 | 2788.9985 | 2686.62 | 2779.5152 | 2693.0344 | 2699.792 | 2576.0698 | 2648.5855 | 2557.8284 | 2552.62 | 2749.2738 | 2728.5855 | 2714.8973 | 2702.0334 | 2634.5506 | 2697.2404 | 2872.7905 | 3194.3796 | 3685.1362 | 4452.4828 | 4673.4464 | 3784.3122 | 2560.416 | 2638.7591 | 2666.9282 | 2560.416 | 2678.3796 | 2784.1037 | 2630.7257 | 2613.7596 | 2593.1391 | 2703.3092 | 2669.8982 | 2595.8963 | 2538.4499 | 2588.2768 | 2615.5845 | 2595.2404 | 2577.9297 | 2642.3422 | 2698.1027 | 2652.5496 | 2476.7262 | 2229.9651 | 2299.6893 | 2421.6539 | 2473.4145 | 2229.205 | 2274.0354 | 2236.1721 | 2596.4484 | 3079.4125 | 3114.1362 | 2767.6539
103 | 3.1274 | 1021.0443 | 2540.8791 | 2759.2389 | 2643.1239 | 2665.8879 | 3603.8436 | 3047.9557 | 2421.2979 | 2367.4749 | 2832.9292 | 2340.118 | 2400.1799 | 2907.7168 | 2743.3864 | 3288.8024 | 2722.9115 | 2801.9882 | 2851.5309 | 2876.6017 | 2407.1623 | 2512.5546 | 3601.3422 | 3717.5575 | 2506.8614 | 2888.5103 | 3128.9587 | 2932.6549 | 3126.5074 | 2478.764 | 2773.4454 | 2760.8997 | 2611.7197 | 2380.4336 | 2822.6726 | 2822.8643 | 2750.7197 | 3231.7345 | 3639.6254 | 3312.1298 | 3058.3894 | 2639.8614 | 3227.8702 | 2930.5339 | 2723.5074 | 2412.1328 | 2606.4631 | 2577.0029 | 2611.0148 | 2557.5959 | 2920.6342 | 2702.4307 | 2653.8554 | 2820.2743 | 3247.4454 | 2921.6902 | 2714.4218 | 2787.9292 | 3245.4366 | 2630.5251 | 2729.8909 | 2558.3451 | 2854.1918 | 2649.4159 | 2689.8938 | 2485.5487 | 3113.9086 | 2867.5103 | 2736.8997 | 2743.7611 | 3098.4454 | 2990.9852 | 2842.0472 | 2836.4631 | 2918.6224 | 3004.2596 | 2969.118 | 2725.0324 | 2676.9056 | 2718.8319 | 2747.1003 | 2644.3599 | 2714.6077 | 2690.7611 | 2665.4454 | 2597.5103 | 2615.9557 | 2584.8525 | 2539.5339 | 2641.7492 | 2706.8643 | 2708.9616 | 2676.0561 | 2578.1357 | 2632.4366 | 2756.2094 | 3167.2006 | 3626.7876 | 4470.9971 | 4580.8672 | 3841.3776 | 2614.6489 | 2657.2094 | 2544.5516 | 2614.6489 | 2652.469 | 2792.6873 | 2673.8849 | 2651.1652 | 2633.5546 | 2646.6431 | 2735.3599 | 2549.2596 | 2585.4071 | 2613.0029 | 2557.4749 | 2550.7316 | 2551.2331 | 2531.1033 | 2653.8082 | 2556.1947 | 2586.2803 | 2220.8938 | 2282.4926 | 2376.9381 | 2478.646 | 2182.2921 | 2298.0029 | 2217.4484 | 2597.4484 | 3028.1033 | 3072.0236 | 2693.764
104 | 3.1384 | 1048.8845 | 2549.4454 | 2774.2886 | 2565.4041 | 2581.8968 | 3506.6411 | 3044.3196 | 2408.6578 | 2382.1465 | 2796.9917 | 2339.53 | 2358.8702 | 2924.4061 | 2753.7114 | 3330.3176 | 2699.4145 | 2758.0433 | 2755.206 | 2798.3938 | 2428.8289 | 2465.5526 | 3592.6042 | 3638.3732 | 2564.5732 | 2830.1588 | 3114.9504 | 2964.7979 | 3108.0516 | 2457.8638 | 2766.9857 | 2694.2557 | 2558.1917 | 2354.233 | 2811.2557 | 2849.1917 | 2674.3176 | 3210.1834 | 3652.233 | 3379.1588 | 3032.0103 | 2675.231 | 3149.4041 | 2989.5959 | 2731.6495 | 2397.7733 | 2571.6042 | 2554.3958 | 2620.9154 | 2559.7114 | 2819.1401 | 2730.2227 | 2579.8515 | 2732.2453 | 3269.9774 | 2854.7443 | 2655.1814 | 2740.8206 | 3139.7094 | 2589.2557 | 2752.8785 | 2516.3609 | 2877.0413 | 2672.3609 | 2676.4434 | 2534.8289 | 3134.2969 | 2810.7547 | 2674.2453 | 2711.1568 | 3135.0742 | 2981.0846 | 2770.4125 | 2819.4248 | 3005.1711 | 2920.6927 | 3011.3525 | 2679.1382 | 2730.2743 | 2766.3176 | 2810.0206 | 2593.3918 | 2746.5713 | 2668.796 | 2716.4371 | 2535.7011 | 2583.3279 | 2517.763 | 2569.8702 | 2597.3835 | 2701.2743 | 2626.0659 | 2723.1052 | 2549.1485 | 2585.2247 | 2734.7011 | 3054.4041 | 3630.7404 | 4497.0619 | 4561.4474 | 3768.1278 | 2500.9794 | 2636.233 | 2549.4145 | 2500.9794 | 2575.8535 | 2749.4125 | 2608.7773 | 2634.7979 | 2599.8515 | 2603.1898 | 2669.231 | 2469.6824 | 2519.4577 | 2574.4371 | 2559.3795 | 2521.732 | 2642.6165 | 2540.5629 | 2601.7011 | 2534.0949 | 2641.9361 | 2250.266 | 2273.5526 | 2402.499 | 2393.7011 | 2276.0413 | 2240.5526 | 2222.9587 | 2572.6372 | 3047.1382 | 3139.6991 | 2650.5216
105 | 3.1494 | 1075.1209 | 2517.6962 | 2773.8594 | 2585.766 | 2680.5909 | 3550.7374 | 3063.174 | 2350.061 | 2273.0688 | 2838.3205 | 2263.6519 | 2282.3304 | 2959.5034 | 2701.2694 | 3281.0089 | 2675.9302 | 2787.4523 | 2858.999 | 2758.8072 | 2419.9833 | 2438.5998 | 3513.939 | 3626.3726 | 2530.3481 | 2820.8093 | 3216.7384 | 2944.7384 | 3069.9744 | 2520.3127 | 2668.94 | 2697.0954 | 2579.8004 | 2309.8171 | 2806.5821 | 2808.7463 | 2631.1308 | 3294.5477 | 3625.3481 | 3368.4179 | 3041.3137 | 2577.9567 | 3163.3658 | 2960.3294 | 2736.0865 | 2330.471 | 2502.6785 | 2612.1907 | 2620.2429 | 2536.766 | 2900.8093 | 2689.0256 | 2583.4779 | 2728.8427 | 3224.6696 | 2820.4602 | 2663.3127 | 2804.7483 | 3047.766 | 2600.1652 | 2616.6441 | 2535.4189 | 2816.0256 | 2620.1819 | 2560.1131 | 2475.3382 | 3088.6431 | 2793.8437 | 2731.5054 | 2630.5831 | 3131.7306 | 2940.5732 | 2706.8702 | 2792.1652 | 3011.7905 | 2911.1288 | 3033.06 | 2655.3559 | 2636.2517 | 2696.1386 | 2779.9302 | 2546.0266 | 2653.6441 | 2599.4956 | 2726.5732 | 2505.6785 | 2524.4002 | 2458.6883 | 2528.3924 | 2600.8082 | 2621.4435 | 2683.2517 | 2666.4169 | 2532.7217 | 2618.0344 | 2728.8879 | 3061.4091 | 3491.296 | 4416.53 | 4560.295 | 3770.3825 | 2492.6686 | 2604.8958 | 2515.6342 | 2492.6686 | 2624.5467 | 2671.1219 | 2618.53 | 2579.9547 | 2607.0954 | 2587.5231 | 2583.3658 | 2523.1652 | 2513.1386 | 2584.295 | 2453.47 | 2449.2173 | 2593.2773 | 2505.6608 | 2561.1042 | 2512.9046 | 2614.234 | 2208.703 | 2222.6686 | 2368.2783 | 2382.8791 | 2238.4002 | 2195.234 | 2263.939 | 2530.4956 | 3061.2006 | 3045.4091 | 2663.8348
106 | 3.1604 | 1007.1283 | 2536.5929 | 2629.823 | 2626.4646 | 2646.6416 | 3446.6106 | 3063.8304 | 2383.2788 | 2239.7655 | 2725.2153 | 2222.823 | 2293.3997 | 2904.2109 | 2695.1018 | 3289.351 | 2643.3776 | 2807.1357 | 2772.5133 | 2669.7994 | 2456.3702 | 2434.2035 | 3514.7168 | 3561.6785 | 2515.3894 | 2871.0678 | 3161.0339 | 2901.528 | 3072.4498 | 2498.7168 | 2712.6194 | 2704.7994 | 2581.0383 | 2270.0192 | 2808.1652 | 2692.0413 | 2635.8835 | 3264.1136 | 3653.1018 | 3396.1814 | 3075.1136 | 2586.7021 | 3183.6873 | 2927.916 | 2661.9366 | 2419.6194 | 2495.0118 | 2548.5664 | 2579.7507 | 2578.3363 | 2944.385 | 2633.3776 | 2528.6268 | 2659.5472 | 3209.2419 | 2769.8451 | 2656.5354 | 2803.1932 | 3127.1092 | 2582.3097 | 2639.7212 | 2599.4011 | 2769.6755 | 2554.8982 | 2561.8378 | 2423.9543 | 3049.6416 | 2797.6681 | 2764.6312 | 2667.8038 | 3126.2758 | 2891.3923 | 2727.0044 | 2803.3658 | 2944.0605 | 2795.792 | 2966.7168 | 2608.4528 | 2627.3466 | 2645.7021 | 2749.1209 | 2572.7965 | 2687.7625 | 2583.4115 | 2635.8422 | 2511.3776 | 2530.5398 | 2571.4189 | 2575.1932 | 2559.9322 | 2608.9587 | 2613.6194 | 2632.9735 | 2535.1962 | 2566.9513 | 2785.174 | 3047.1253 | 3546.1844 | 4361.5664 | 4543.8068 | 3778.7242 | 2468.916 | 2601.0531 | 2458.9926 | 2468.916 | 2536.1018 | 2681.4793 | 2573.4454 | 2459.3127 | 2578.1888 | 2657.6725 | 2592.0649 | 2508.7965 | 2470.8378 | 2511.1475 | 2482.8643 | 2448.6829 | 2513.3245 | 2482.7625 | 2548.9661 | 2514.4985 | 2516.8761 | 2114.1092 | 2137.8982 | 2386.1431 | 2418.3289 | 2220.1327 | 2131 | 2244.3776 | 2510.6342 | 3097.9248 | 3073.2566 | 2664.5472
107 | 3.1714 | 998.6401 | 2485.4798 | 2603.5939 | 2584.1426 | 2642.6627 | 3503.292 | 3022.0403 | 2422.8918 | 2315.1485 | 2723.6795 | 2181.1485 | 2320.0973 | 2910.2694 | 2719.0856 | 3248.0177 | 2647.708 | 2787.9321 | 2756.4287 | 2715.0177 | 2408.12 | 2442.8004 | 3518.1317 | 3617.0865 | 2416.0059 | 2877.3658 | 3159.6401 | 2911.8633 | 2999.7709 | 2489.7148 | 2647.5084 | 2738.8122 | 2525.411 | 2264.3137 | 2834.8357 | 2740.8053 | 2604.8624 | 3203.766 | 3663.1544 | 3359.0914 | 3042.3255 | 2562.4857 | 3104.5831 | 3008.3952 | 2609.5821 | 2339.6686 | 2471.2508 | 2502.3658 | 2592.9774 | 2506.9489 | 2898.0236 | 2613.177 | 2493.0797 | 2657.8171 | 3137.0747 | 2790.5143 | 2666.1494 | 2669.6578 | 3140.6804 | 2563.0797 | 2581.6342 | 2487.1593 | 2737.6519 | 2533.0344 | 2562.9371 | 2497.2625 | 3080.7935 | 2762.4287 | 2638.2055 | 2643.0403 | 3093.297 | 2882.469 | 2664.5939 | 2782.0924 | 2941.3884 | 2855.3599 | 2976.4916 | 2611.0914 | 2662.0521 | 2609.9774 | 2677.7551 | 2502.115 | 2626.6519 | 2593.9194 | 2565.7483 | 2456.6519 | 2478.9548 | 2544.2576 | 2452.2576 | 2556.3599 | 2582.9086 | 2546.0856 | 2619.412 | 2503.3264 | 2536.474 | 2734.8741 | 2935.4572 | 3589.2576 | 4354.7365 | 4602.235 | 3823.8171 | 2592.4002 | 2527.3264 | 2440.3314 | 2592.4002 | 2549.7886 | 2671.6342 | 2531.6804 | 2490.2635 | 2544.9203 | 2555.177 | 2553.3304 | 2486.2684 | 2455.5831 | 2485.1426 | 2418.1779 | 2442.589 | 2489.9262 | 2451.0285 | 2514.0177 | 2478.0511 | 2411.1662 | 2143.7316 | 2149.0452 | 2418.9144 | 2325.8574 | 2119.2517 | 2143.7198 | 2201.4631 | 2498.2911 | 3055.0973 | 3074.1485 | 2655.8515
108 | 3.1824 | 982.7394 | 2508.7817 | 2658.8786 | 2630.6603 | 2632.1583 | 3444.3874 | 2968.5034 | 2372.3151 | 2321.6426 | 2788.8486 | 2182.2851 | 2284.6003 | 2813.1937 | 2744.412 | 3200.7449 | 2678.5649 | 2733.6111 | 2730.4911 | 2658.6057 | 2411.7394 | 2371.94 | 3520.8048 | 3515.6849 | 2426.09 | 2861.4366 | 3114.7463 | 2858.412 | 3062.2183 | 2437.9631 | 2685.6603 | 2641.794 | 2558.4066 | 2304.4434 | 2651.7817 | 2762.6372 | 2641.8977 | 3206.9263 | 3583.9292 | 3327.0614 | 3055.9823 | 2540.7217 | 3074.1091 | 2896.03 | 2672.4312 | 2319.0791 | 2505.6317 | 2471.94 | 2595.4597 | 2483.8663 | 2824.9086 | 2622.2306 | 2554.648 | 2639.94 | 3099.9631 | 2778.8731 | 2560.0969 | 2607.4297 | 3057.0614 | 2572.0192 | 2580.2306 | 2543.4857 | 2695.9086 | 2505.97 | 2549.94 | 2483.7409 | 3116.1774 | 2738.6971 | 2641.3097 | 2595.4911 | 3089.8977 | 2835.7694 | 2681.7463 | 2660.8363 | 2944.7217 | 2895.8417 | 2940.2306 | 2577.8009 | 2588.9754 | 2601.1146 | 2626.356 | 2436.0477 | 2583.648 | 2623.7832 | 2588.4489 | 2382.1337 | 2420.5757 | 2488.4666 | 2382.2851 | 2565.3274 | 2567.0369 | 2564.4789 | 2535.764 | 2435.6357 | 2525.7409 | 2756.7817 | 3005.3614 | 3544.8117 | 4398.8677 | 4559.9563 | 3828.6303 | 2536.8009 | 2431.588 | 2450.8363 | 2536.8009 | 2525.8854 | 2633.0437 | 2497.854 | 2433.5266 | 2499.7394 | 2549.1037 | 2625.644 | 2484.0914 | 2428.8909 | 2508.97 | 2355.9686 | 2405.8117 | 2431.5089 | 2443.5034 | 2433.3397 | 2479.3889 | 2377.3206 | 2129.206 | 2204.2729 | 2375.6249 | 2283.1214 | 2091.1937 | 2149.0491 | 2160.7517 | 2525.146 | 3033.7817 | 3052.91 | 2631.7094
109 | 3.1934 | 982.9469 | 2500.3009 | 2696.3009 | 2624.8938 | 2589.4248 | 3450.8584 | 2961.2301 | 2376.2743 | 2331.2832 | 2759.0796 | 2196.3275 | 2219.761 | 2816.2655 | 2707.177 | 3229.0088 | 2581.3451 | 2787.7522 | 2681.2212 | 2707.2832 | 2368.0708 | 2358.708 | 3512.8583 | 3551.3452 | 2485.2301 | 2816.5133 | 3076.177 | 2855.8849 | 3035.0088 | 2469.1416 | 2695.6106 | 2599.6725 | 2489.8407 | 2254.469 | 2670.8053 | 2702.3097 | 2617.3806 | 3255.2035 | 3491.0089 | 3284.5929 | 3053.3806 | 2511.4779 | 3109.9469 | 2891.5133 | 2581.115 | 2298.2035 | 2388.7611 | 2430.8761 | 2609.3539 | 2518.9115 | 2837.5133 | 2601.9823 | 2548.7964 | 2608.3717 | 3114.115 | 2720.1327 | 2564.0885 | 2640.2478 | 3034.239 | 2509.7522 | 2553.1062 | 2484.823 | 2721.6106 | 2491.4956 | 2508.8761 | 2393.2124 | 3016.7434 | 2719.1062 | 2595.2743 | 2607.7788 | 3025.4336 | 2872.7168 | 2647.4336 | 2729.8673 | 2889.2832 | 2936.938 | 2877.4159 | 2531.9381 | 2608.7788 | 2617.531 | 2740.5221 | 2464.5841 | 2604.9734 | 2510.1593 | 2562.9292 | 2355.2743 | 2426.6195 | 2459.3274 | 2362.2743 | 2549.7257 | 2518.8319 | 2553.1062 | 2505.6903 | 2437.5575 | 2431.6106 | 2784.646 | 3116.1504 | 3575.0354 | 4390.6549 | 4631.8761 | 3831.3097 | 2467.3628 | 2424.292 | 2432.7699 | 2467.3628 | 2493.7522 | 2550.646 | 2498.761 | 2501.4425 | 2479.646 | 2496.9646 | 2507.5487 | 2441.3628 | 2403.2035 | 2490.5664 | 2444.6726 | 2433.0885 | 2438.8672 | 2423.4602 | 2418.4779 | 2380.0619 | 2406.7168 | 2108.7876 | 2165.1858 | 2299.823 | 2293.177 | 2123.0797 | 2092.4159 | 2101.1239 | 2460.7168 | 3056.4071 | 3021.2744 | 2599.4336
110 | 3.2044 | 971.3343 | 2483.6657 | 2712.1809 | 2543.674 | 2625.4931 | 3409.3451 | 2964.7974 | 2349.3672 | 2318.2659 | 2703.5069 | 2227.8166 | 2166.293 | 2793.8137 | 2617.7945 | 3197.8137 | 2568.4494 | 2748.1534 | 2632.2684 | 2712.2714 | 2279.2439 | 2409.9562 | 3372.3422 | 3611.7316 | 2429.1042 | 2800.4656 | 3108.559 | 2802.4631 | 2963.3864 | 2432.9562 | 2659.2439 | 2553.5752 | 2466.9917 | 2262.0029 | 2680.0905 | 2715.1917 | 2645.8741 | 3188.5644 | 3582.5973 | 3319.7891 | 3072.0221 | 2529.263 | 3159.7261 | 2839.4548 | 2559.5561 | 2315.0383 | 2376.9808 | 2440.2439 | 2489.9425 | 2489.9316 | 2799.3835 | 2594.148 | 2480.7616 | 2632.8274 | 3079.6465 | 2682.6273 | 2557.3972 | 2621.5123 | 3128.8137 | 2506.0659 | 2521.6357 | 2486.9946 | 2687.2822 | 2490.6165 | 2494.1234 | 2445.6794 | 3055.4631 | 2696.1288 | 2619.6657 | 2605.7728 | 3003.8083 | 2860.8358 | 2721.5123 | 2798.529 | 2879.7616 | 2867.9508 | 2842.9587 | 2553.2193 | 2578.674 | 2546.1288 | 2682.3097 | 2428.915 | 2566.263 | 2479.0905 | 2541.7837 | 2383.7424 | 2432.6794 | 2379.5836 | 2358.8083 | 2519.2193 | 2514.6082 | 2539.9808 | 2490.0905 | 2384.9204 | 2485.5123 | 2762.4906 | 2969.3727 | 3568.6219 | 4512.0659 | 4512.2468 | 3848.6082 | 2471.4739 | 2433.0767 | 2412.2163 | 2471.4739 | 2505.3535 | 2586.2439 | 2456.1642 | 2495.4056 | 2449.7753 | 2490.8083 | 2501.6357 | 2439.3589 | 2365.7807 | 2495.5452 | 2421.2384 | 2446.441 | 2375.233 | 2439.8658 | 2438.3014 | 2380.1834 | 2355.1863 | 2079.4218 | 2099.885 | 2261.326 | 2356.4739 | 2112.5919 | 2083.3397 | 2014.0575 | 2439.0821 | 3057.4631 | 3086.3343 | 2604.3068
111 | 3.2154 | 951.4327 | 2508.4956 | 2667.9843 | 2532.529 | 2621.2311 | 3433.7355 | 2902.8093 | 2319.9056 | 2222.7847 | 2637.3658 | 2136.8319 | 2222.9036 | 2813.2242 | 2620.7915 | 3214.4081 | 2531.5673 | 2761.8879 | 2672.3206 | 2610.9508 | 2274.5762 | 2357.9754 | 3447.9891 | 3556.5428 | 2381.2085 | 2869.6234 | 3055.6077 | 2752.2802 | 2934.1455 | 2399.6323 | 2663.7286 | 2574.2782 | 2512.119 | 2185.0492 | 2650.7758 | 2729.7198 | 2523.8702 | 3176.944 | 3551.2085 | 3274.7669 | 3004.7935 | 2498.7669 | 3093.1278 | 2851.648 | 2565.6057 | 2307.7355 | 2371.8319 | 2414.4553 | 2460.7669 | 2418.2242 | 2842.5202 | 2577.8161 | 2491.2065 | 2588.9508 | 3084.1907 | 2685.8948 | 2564.704 | 2624.1593 | 3118.6549 | 2471.8093 | 2469.5673 | 2447.7758 | 2676.7198 | 2474.9371 | 2451.7286 | 2452.2802 | 2990.4956 | 2668.6637 | 2607.1819 | 2496.9528 | 3012.8653 | 2792.2488 | 2654.4601 | 2674.3117 | 2884.2222 | 2797.8722 | 2883.5654 | 2514.3923 | 2563.3048 | 2487.3343 | 2605.3677 | 2490.7601 | 2544.9194 | 2482.5133 | 2496.9685 | 2395.3746 | 2410.4149 | 2385.4327 | 2356.9528 | 2490.2488 | 2527.1593 | 2507.7266 | 2441.4149 | 2422.2694 | 2468.0944 | 2649.8004 | 2942.1681 | 3594.2734 | 4419.0226 | 4455.9705 | 3810.821 | 2459.5988 | 2448.0403 | 2459.8476 | 2459.5988 | 2451.175 | 2534.0541 | 2488.4936 | 2457.9371 | 2471.5605 | 2465.0157 | 2456.9597 | 2403.6303 | 2417.9843 | 2449.9105 | 2384.3923 | 2447.646 | 2318.881 | 2401.0069 | 2435.704 | 2443.7109 | 2355.4887 | 2121.944 | 2107.7129 | 2290.0718 | 2357.6234 | 2044.0964 | 2088.9282 | 2029.1524 | 2484.4641 | 3010.0964 | 3074.177 | 2615.5831
112 | 3.2264 | 968.8923 | 2468.7773 | 2640.3835 | 2562.7065 | 2569.1563 | 3453.0738 | 2929.5914 | 2344.8333 | 2209.5619 | 2576.9661 | 2123.7139 | 2194.5206 | 2796.5914 | 2618.0295 | 3208.9852 | 2534.5089 | 2726.3053 | 2642.5398 | 2588.1372 | 2284.4572 | 2366.7876 | 3391.382 | 3502.7212 | 2387.6136 | 2805.6548 | 3056.5723 | 2753.2566 | 2991.1711 | 2398.7065 | 2668.5988 | 2536.5162 | 2457.9779 | 2207.4985 | 2661.944 | 2758.9808 | 2486.4307 | 3204.8732 | 3551.4203 | 3245.1151 | 3016.7256 | 2488.059 | 3034.6327 | 2843.3496 | 2498.323 | 2293.1932 | 2379.8437 | 2369.792 | 2436.2153 | 2430.3717 | 2840.9852 | 2576.8068 | 2437.3791 | 2540.6357 | 3054.8584 | 2654.5797 | 2541.9557 | 2606.2935 | 3072.7286 | 2471.2684 | 2422.2345 | 2411.5841 | 2639.3127 | 2485.4838 | 2458.056 | 2450.7065 | 2991.944 | 2646.4351 | 2523.9145 | 2531.8215 | 3071.4646 | 2772.997 | 2538.531 | 2669.2861 | 2792.4203 | 2787.8776 | 2800.4086 | 2503.5914 | 2541.261 | 2441.3156 | 2622.323 | 2481.9779 | 2498.9631 | 2524.7994 | 2492.4056 | 2331.7139 | 2346.2345 | 2388.32 | 2383.4425 | 2531.472 | 2483.792 | 2497.9027 | 2423.1372 | 2320.0074 | 2403.0929 | 2642.6844 | 2964.0295 | 3584.0604 | 4398.795 | 4517.1047 | 3755.6682 | 2371.6548 | 2446.0634 | 2434.6209 | 2371.6548 | 2393.149 | 2497.9882 | 2455.9396 | 2460.8333 | 2484.0782 | 2488.2449 | 2437.0782 | 2333 | 2414.792 | 2385.3569 | 2432.9735 | 2366.7286 | 2405.4455 | 2354.6357 | 2422.792 | 2385.1563 | 2391.2861 | 2126.5575 | 2129.792 | 2239.8333 | 2288.1902 | 2050.6209 | 2085.2493 | 2048.5501 | 2482.0782 | 3061.4307 | 3085.1268 | 2549.2197
113 | 3.2374 | 980.1151 | 2439.7915 | 2592.9597 | 2571.2449 | 2564.5831 | 3500.4041 | 2873.0659 | 2274.1298 | 2126.9489 | 2625.351 | 2143.764 | 2172.1298 | 2772.4533 | 2633.9489 | 3204.7915 | 2573.9489 | 2709.413 | 2565.3107 | 2610.5339 | 2264.0275 | 2395.3785 | 3495.7532 | 3477.0403 | 2340.7679 | 2727.9616 | 3070.3255 | 2731.9744 | 2976.1042 | 2417.0364 | 2597.2232 | 2587.0914 | 2420.4297 | 2174.2979 | 2684.8319 | 2803.2743 | 2507.821 | 3225.7424 | 3519 | 3242.3913 | 2951.5959 | 2526.6765 | 2994.3343 | 2866.9066 | 2531.6893 | 2282.1573 | 2376.1317 | 2368.7257 | 2454.8191 | 2457.9597 | 2827 | 2550.1829 | 2434.6765 | 2505.7296 | 3087.351 | 2676.4553 | 2512.4277 | 2634.6618 | 3059.6893 | 2391.5447 | 2446.6323 | 2396.2704 | 2584.5978 | 2427.7532 | 2470.945 | 2422.7148 | 3020.4808 | 2666.413 | 2513.9361 | 2489.055 | 3072.7148 | 2703.5831 | 2604.9744 | 2659.4041 | 2764.5192 | 2789.2576 | 2854.6745 | 2477.5703 | 2468.1023 | 2487.2724 | 2654.9361 | 2457.234 | 2502.7896 | 2488.1701 | 2456.5447 | 2332.9341 | 2324.3107 | 2373.3619 | 2352.4169 | 2473.057 | 2430.3785 | 2540.3402 | 2468.883 | 2317.2085 | 2392.1957 | 2596.351 | 2972.4169 | 3472.3894 | 4404.6273 | 4442.4936 | 3865.9744 | 2317.3107 | 2398.5339 | 2417.7404 | 2317.3107 | 2415.1003 | 2535.3402 | 2493.4208 | 2400.2979 | 2426.8082 | 2534.0659 | 2441.5703 | 2335.2596 | 2393.766 | 2382.3894 | 2439.0167 | 2382.6873 | 2419.3549 | 2325.7551 | 2390.8446 | 2386.2321 | 2401.7404 | 2104.1062 | 2094.0787 | 2119.4189 | 2209.2468 | 2027.7148 | 2087.1023 | 2060.532 | 2459.8319 | 3059.8485 | 3000.236 | 2498.2468
114 | 3.2484 | 999.7222 | 2472.7222 | 2620.5556 | 2582.6111 | 2589.2222 | 3424.8889 | 2841.8333 | 2233.2222 | 2130.7222 | 2624.9444 | 2170.3889 | 2193.8333 | 2778.1111 | 2624.5 | 3176.2778 | 2505 | 2746.9444 | 2589.3889 | 2565.6667 | 2243.5 | 2340.2222 | 3414.9444 | 3394.6111 | 2329.3333 | 2746.8334 | 3040.1111 | 2723.9444 | 2959.0556 | 2432.1666 | 2556.6667 | 2542.5 | 2415.5556 | 2173.1667 | 2645.4444 | 2746.8889 | 2466 | 3189.5556 | 3512.7778 | 3242 | 2969.1667 | 2465.3333 | 3007.9444 | 2860.7222 | 2501.3889 | 2211 | 2347.8889 | 2362.8889 | 2428.7222 | 2450.5555 | 2813 | 2538.3889 | 2364 | 2468.0556 | 3077.2222 | 2618.5555 | 2502.2778 | 2584.0555 | 3049.2222 | 2405.1111 | 2532.6111 | 2402.8333 | 2586.0556 | 2409.5555 | 2480.5 | 2421.8333 | 3002.8889 | 2664.6666 | 2520.1111 | 2416.1111 | 3027.8889 | 2754.2778 | 2557.2778 | 2593.6111 | 2769.6667 | 2767.1111 | 2834.7222 | 2464.5555 | 2522 | 2502.6111 | 2571.6666 | 2430.7778 | 2486.5 | 2459.3333 | 2475.1667 | 2359.5 | 2350.3333 | 2383.1666 | 2329.7222 | 2388.2778 | 2429.2778 | 2430.5555 | 2440.5 | 2322 | 2364.3889 | 2596.3333 | 2964.8889 | 3512.6111 | 4304.1111 | 4425.2222 | 3835.3889 | 2340.6111 | 2373.8889 | 2387.5555 | 2340.6111 | 2445.8333 | 2473.3889 | 2410.5 | 2419.7778 | 2329.0555 | 2493.5 | 2437.8889 | 2319.9444 | 2361.8889 | 2342.5 | 2300.5555 | 2347.5 | 2308.1667 | 2285.6111 | 2342.0555 | 2351 | 2364.9444 | 2062 | 2041.3889 | 2097.7223 | 2239.1111 | 2051.3334 | 2064 | 2034.1111 | 2470.6111 | 2930.2222 | 2958.0556 | 2496.2222
115 | 3.2594 | 965.2153 | 2483.7552 | 2612.4012 | 2531.8142 | 2536.882 | 3355.3717 | 2817.767 | 2205.3157 | 2129.0855 | 2610.5162 | 2143.4484 | 2187.2891 | 2711.3186 | 2593.6607 | 3122.0472 | 2451.6696 | 2663.0737 | 2565.528 | 2557.0177 | 2254.5516 | 2300.7197 | 3335.7463 | 3373.1652 | 2377 | 2794.7847 | 3020.5929 | 2719.3776 | 2975.4897 | 2333.3304 | 2596.4307 | 2479.3599 | 2459.0236 | 2146.0118 | 2614.0118 | 2711.174 | 2503.1799 | 3148.6076 | 3483.7434 | 3287.941 | 2932.7847 | 2407.1799 | 2957.5664 | 2757.8083 | 2455.6961 | 2174.0531 | 2336.6843 | 2314.0649 | 2405.1504 | 2386.1976 | 2810.6578 | 2519.6784 | 2310.8378 | 2525.2803 | 3055.5398 | 2571.4484 | 2465.6342 | 2528.5369 | 3037.5044 | 2391.7788 | 2475.0973 | 2366.9263 | 2610.6165 | 2363.7552 | 2361.9587 | 2420.7375 | 2960.9233 | 2552.7729 | 2535.702 | 2405.354 | 2961.3481 | 2727.236 | 2476.7375 | 2587.1121 | 2770.413 | 2697.9381 | 2849.2242 | 2405.0118 | 2512.7079 | 2462.8525 | 2465.9587 | 2360.8879 | 2340.3422 | 2449.3599 | 2445.7079 | 2328.4926 | 2361.6224 | 2315.1563 | 2279.2065 | 2385.41 | 2457.0678 | 2371.2448 | 2360.1091 | 2315.2832 | 2354.4012 | 2620.2271 | 2970.3009 | 3517.2537 | 4320.0649 | 4438.9086 | 3811.0236 | 2311.115 | 2355.295 | 2327.1445 | 2311.115 | 2367.233 | 2501.2803 | 2445.4926 | 2381.2832 | 2302.6283 | 2436.3127 | 2405.8496 | 2310.4602 | 2274.3186 | 2287.8496 | 2263.3098 | 2285.6932 | 2258.882 | 2311.2448 | 2285.2035 | 2244.4425 | 2334.2625 | 2039.8702 | 1984.9941 | 2192.4071 | 2267.1386 | 2072.2891 | 2008.7286 | 1971.7965 | 2438.177 | 2897.2035 | 3009.4956 | 2535.9204
116 | 3.2704 | 972.1731 | 2484.1544 | 2547.6716 | 2509.6435 | 2516.2798 | 3374.0777 | 2779.4995 | 2290.0487 | 2137.5689 | 2575.4425 | 2200.6642 | 2121.4902 | 2672.0674 | 2538.2984 | 3140.2114 | 2472.115 | 2664.7419 | 2570.7606 | 2539.1244 | 2220.9036 | 2253.3845 | 3418.6829 | 3487.5969 | 2388.5772 | 2803.5388 | 3021.5482 | 2729.5005 | 3020.5865 | 2314.5388 | 2573.5472 | 2446.1824 | 2417.6622 | 2152 | 2618.3461 | 2706.2694 | 2465.2974 | 3112.1937 | 3510.8186 | 3206.3161 | 2964.2798 | 2396.1731 | 3053.9907 | 2794.6642 | 2427.8083 | 2210.9233 | 2316.0394 | 2321.7783 | 2406.8756 | 2360.1917 | 2804.028 | 2455.9326 | 2374.6549 | 2546.5669 | 3033.2694 | 2576.5669 | 2450.3358 | 2589.3171 | 2990.2404 | 2367.7689 | 2401.086 | 2369.3928 | 2616.5192 | 2336.8653 | 2357.1254 | 2413.8083 | 2917.8653 | 2571.2508 | 2492.6155 | 2396.058 | 2966.4238 | 2656.8176 | 2509.088 | 2644 | 2785.2694 | 2741.5378 | 2872.8633 | 2397.7689 | 2462.2788 | 2403.7119 | 2502.059 | 2299.0383 | 2361.03 | 2411.2207 | 2394.8653 | 2278.1347 | 2345.7109 | 2307.058 | 2226.0684 | 2300.9907 | 2363.1244 | 2403.0964 | 2331.7793 | 2286.1834 | 2375.2601 | 2592.6529 | 2891.4705 | 3442.1731 | 4383.1834 | 4497.7212 | 3782.8933 | 2246.4995 | 2352.857 | 2284.115 | 2246.4995 | 2369.115 | 2515.5378 | 2349.6529 | 2343.087 | 2312.8653 | 2403.5772 | 2403.7109 | 2343.8943 | 2252.0197 | 2276.2011 | 2295.0187 | 2300.7492 | 2203.9617 | 2282.7296 | 2247.7202 | 2281.2984 | 2336.6435 | 1945.884 | 2004.3274 | 2172.8456 | 2197.5472 | 2021.3751 | 2045.884 | 1998.5295 | 2347.7409 | 2937.0394 | 3025.3358 | 2594.1057
117 | 3.2814 | 976.9607 | 2485.0678 | 2469.9017 | 2463.7424 | 2567.3166 | 3402.2665 | 2750.703 | 2314.6529 | 2180.4828 | 2574.9302 | 2203.1465 | 2103.3166 | 2665.6332 | 2554.0393 | 3129.0197 | 2481.2291 | 2697.116 | 2611.6922 | 2513.3471 | 2138.175 | 2255.0698 | 3376.9282 | 3510.5634 | 2370.939 | 2789.9587 | 2999.4739 | 2731.4739 | 2952.3038 | 2293.7207 | 2524.9607 | 2442.9017 | 2385.4278 | 2134.9695 | 2571.7207 | 2686.0678 | 2460.9125 | 3216.882 | 3541.8889 | 3157.6942 | 2982.1072 | 2386.059 | 3016.0767 | 2754.1445 | 2469.8122 | 2210.0983 | 2375.9607 | 2291.5349 | 2422.8604 | 2340.3274 | 2773.8624 | 2478.9715 | 2373.7099 | 2511.3363 | 3052.6637 | 2551.1288 | 2406.1288 | 2584.238 | 2997.9607 | 2369.0108 | 2371.2773 | 2322.0217 | 2613 | 2345.4759 | 2307.1947 | 2442.6136 | 2958.4081 | 2568.0373 | 2493.4848 | 2391.7316 | 2985.88 | 2616.2448 | 2550.1858 | 2613.3451 | 2805.941 | 2681.0088 | 2757.7729 | 2373.0393 | 2458.703 | 2447.7139 | 2543.8299 | 2332.0217 | 2437.1091 | 2360.4543 | 2381.762 | 2271.6529 | 2324.0305 | 2337.6441 | 2307.4563 | 2308.2075 | 2314.7817 | 2439.4169 | 2352.3166 | 2290.2556 | 2389.5241 | 2537.8319 | 2825.0285 | 3503.1613 | 4379.5919 | 4461.7011 | 3789.0128 | 2214.7227 | 2330.9065 | 2271.0895 | 2214.7227 | 2326.7119 | 2475.4454 | 2328.8928 | 2342.7709 | 2307.4258 | 2362.9479 | 2378.1986 | 2336.5044 | 2297.7335 | 2231.238 | 2240.1376 | 2274.1593 | 2224.6244 | 2215.9607 | 2217.5457 | 2264.6814 | 2329.9125 | 1948.1505 | 2015.1376 | 2109.3451 | 2154.4956 | 1983.0177 | 2020.0502 | 1985.354 | 2393.6244 | 2986.3166 | 3013.9125 | 2549.9783
118 | 3.2924 | 960.469 | 2456.8938 | 2517.1593 | 2475.9602 | 2551.3628 | 3361.6681 | 2780.5974 | 2237.1106 | 2079.4602 | 2549.9248 | 2150.5708 | 2043.9955 | 2667.8938 | 2574.9115 | 3117.115 | 2471.1858 | 2607.5531 | 2552.8318 | 2473.4911 | 2068.531 | 2227.0708 | 3293.9248 | 3458.9071 | 2284.261 | 2749.1991 | 2950.0487 | 2677.6681 | 2864.9425 | 2244.9513 | 2517.2301 | 2482.7124 | 2371.2832 | 2119.0089 | 2548.3584 | 2663.5885 | 2469.8894 | 3185.3672 | 3496.7124 | 3176.1062 | 2944.5797 | 2409.4823 | 2949.0089 | 2677.3009 | 2452.7168 | 2163.2787 | 2310.8451 | 2285.8584 | 2400.4469 | 2284.7787 | 2813.3894 | 2490.7832 | 2305.3982 | 2478.6593 | 3098.9204 | 2508.4425 | 2434.8451 | 2564.2124 | 2973.5841 | 2368.4026 | 2366.4292 | 2339.6416 | 2614.7522 | 2365 | 2223.0177 | 2391.5088 | 3006.1062 | 2523.0266 | 2471.4867 | 2379.5885 | 2921.3363 | 2537.8761 | 2447.0796 | 2587.3319 | 2795.6593 | 2681.5398 | 2778.5044 | 2395.6637 | 2458.8097 | 2446.7743 | 2470.4115 | 2370.2301 | 2389.2433 | 2339.5797 | 2403.1018 | 2259.3805 | 2304.0708 | 2299.8717 | 2291.5752 | 2309.0177 | 2316.7699 | 2416.6814 | 2375.7876 | 2278.739 | 2355.9823 | 2543.8761 | 2835.9956 | 3542.7566 | 4324.4469 | 4421.7699 | 3873.1239 | 2216.9425 | 2260.5089 | 2263.6018 | 2216.9425 | 2285.2965 | 2426.4425 | 2345.3938 | 2265.6416 | 2297.4867 | 2327.177 | 2349.5663 | 2299.9734 | 2287.5177 | 2248.4469 | 2253.177 | 2292.115 | 2199.6327 | 2249.8451 | 2206.0044 | 2208.1814 | 2325.7478 | 1985.2876 | 1956.3363 | 2103.4823 | 2166.9336 | 1898.6593 | 2017.9734 | 1913.8761 | 2360.3097 | 2948.8982 | 3047.4203 | 2502.5708
119 | 3.3034 | 964.1269 | 2471.9469 | 2538.94 | 2477.8338 | 2534.9469 | 3369.0207 | 2782.6745 | 2221.5979 | 2088.0875 | 2557.2655 | 2094.6008 | 2009.3255 | 2686.5497 | 2584.0786 | 3090.3373 | 2411.4041 | 2601.6028 | 2550.4317 | 2423.9676 | 2120.2655 | 2181.0069 | 3250.3766 | 3432.0531 | 2249.5634 | 2790.7227 | 2951.3117 | 2662.5034 | 2869.5497 | 2240.0855 | 2492.4966 | 2463.4159 | 2312.2262 | 2142.411 | 2547.7138 | 2705.3117 | 2453.8751 | 3155.4317 | 3503.0855 | 3157.9007 | 2959.7483 | 2393.7532 | 2949.4435 | 2777.6372 | 2487.4523 | 2141.6096 | 2242.6696 | 2264.9469 | 2402.7945 | 2231.1337 | 2739.2822 | 2492.0738 | 2332.6028 | 2428.8338 | 3081.5014 | 2482.4248 | 2426.5683 | 2489.9725 | 2935.1524 | 2327.5959 | 2371.9607 | 2299.7552 | 2615.649 | 2336.178 | 2216.2311 | 2329.3441 | 2982.4966 | 2544.5959 | 2491.7758 | 2338.4483 | 2931.9882 | 2621.413 | 2441.9813 | 2574.1249 | 2732.3235 | 2717.9469 | 2728.5083 | 2412.9331 | 2433.9331 | 2371.5221 | 2457.06 | 2299.1111 | 2328.4572 | 2339.8869 | 2384.0049 | 2269.9538 | 2220.6283 | 2222.1917 | 2228.8407 | 2358.9744 | 2313.5821 | 2314.5752 | 2357.707 | 2325.5959 | 2362.3186 | 2548.1524 | 2906.2586 | 3530.5841 | 4300.2242 | 4487.0344 | 3820.2104 | 2196.7807 | 2268.6008 | 2267.0207 | 2196.7807 | 2293.3186 | 2385.5634 | 2333.4248 | 2239.2655 | 2251.5752 | 2345.7738 | 2298.1062 | 2254.0855 | 2222.1848 | 2239.5359 | 2259.1642 | 2260.3835 | 2155.709 | 2258.2311 | 2174.471 | 2192.0924 | 2266.3048 | 1946.0393 | 1953.0275 | 2131.8869 | 2147.6883 | 1917.2862 | 1989.1986 | 1897.4366 | 2324.6234 | 2909.4503 | 2999.8269 | 2513.2517
120 | 3.3144 | 960.2738 | 2446.644 | 2504.0319 | 2492.7119 | 2509.644 | 3372.3913 | 2714.2419 | 2189.7433 | 2081.3805 | 2542.2134 | 2059.0251 | 2078.2208 | 2649.4518 | 2557.8648 | 3073.7011 | 2442.057 | 2618.4912 | 2607.9395 | 2479.2955 | 2197.7404 | 2216.9823 | 3234.7439 | 3445.7901 | 2221.498 | 2757.1421 | 2943.1671 | 2644.4056 | 2895.3702 | 2261.413 | 2512.9076 | 2387.7188 | 2393.7866 | 2066.5476 | 2514.5197 | 2631.0103 | 2429.6937 | 3182.8682 | 3495.267 | 3149.2847 | 2969.0285 | 2337.0428 | 2984.8009 | 2766.6794 | 2488.6509 | 2170.8791 | 2297.1318 | 2299.324 | 2405.1033 | 2276.2099 | 2675.8756 | 2464.7684 | 2321.733 | 2455.7404 | 3015.1033 | 2523.8505 | 2356.1067 | 2467.7439 | 2922.8078 | 2344.5944 | 2374.6583 | 2257.5727 | 2552.4695 | 2271.854 | 2187.3879 | 2320.4626 | 2890.4872 | 2538 | 2502.4661 | 2307.8152 | 2935.1706 | 2677.1244 | 2451.2881 | 2522.4341 | 2650.676 | 2674.1386 | 2680.9681 | 2381.3736 | 2436.6087 | 2314.0462 | 2466.5231 | 2271.235 | 2290.498 | 2350.9218 | 2327.267 | 2206.6047 | 2161.4803 | 2229.8186 | 2226.0285 | 2369.7792 | 2270.1957 | 2270.8363 | 2277.4341 | 2271.8112 | 2389.6868 | 2555.7011 | 2903.3451 | 3490.1597 | 4237.2561 | 4442.5334 | 3788.9218 | 2169.7011 | 2181.6726 | 2263.4518 | 2169.7011 | 2302.644 | 2354.2596 | 2314.2453 | 2270.9395 | 2254.0998 | 2324.356 | 2278.854 | 2220.3594 | 2185.9715 | 2187.3913 | 2186.733 | 2190.0496 | 2170.9253 | 2237.2527 | 2171.4518 | 2190.6406 | 2279.089 | 1898.4661 | 1984.0462 | 2121.178 | 2116.0462 | 1985.6406 | 1962.943 | 1867.2099 | 2336.5551 | 2923.8791 | 2947.2704 | 2522.9395
121 | 3.3254 | 952.5428 | 2392.1298 | 2486.7493 | 2495.8318 | 2439.9321 | 3293.8318 | 2699.5723 | 2075.8732 | 2005.8997 | 2492.2006 | 2142.7198 | 2048.3097 | 2618.3982 | 2530.4307 | 3077.5251 | 2440.0796 | 2603.0796 | 2545.7109 | 2481.5546 | 2155.466 | 2213.705 | 3267.5988 | 3411.8318 | 2210.0679 | 2677.8997 | 2898.7021 | 2610.826 | 2930.7493 | 2285.4543 | 2493.6076 | 2391.0207 | 2356.8879 | 2021.1652 | 2546.7994 | 2581.8909 | 2309.882 | 3142.7021 | 3503.5015 | 3185.0177 | 2970.177 | 2359.4248 | 2955.3628 | 2713.0207 | 2426.2449 | 2146.1563 | 2266.5339 | 2247.6401 | 2384.2802 | 2227.056 | 2695.705 | 2401.3481 | 2295.4513 | 2470.4572 | 2983.3569 | 2518.8761 | 2298.8053 | 2467.8289 | 2968.59 | 2315.4837 | 2293.8849 | 2209.9233 | 2501.4218 | 2261.7463 | 2170.0029 | 2329.0738 | 2851.7404 | 2514.7286 | 2426.6312 | 2332.3805 | 2892.6962 | 2574.6578 | 2403.5251 | 2575.9646 | 2659.6165 | 2596.2301 | 2706.3245 | 2342.7551 | 2421.6076 | 2364.5399 | 2453.3009 | 2242.6106 | 2255.3215 | 2385.1504 | 2287.059 | 2139.2979 | 2161.4956 | 2243.5516 | 2199.6312 | 2325.6018 | 2292.4189 | 2283.0974 | 2300.0944 | 2163.3333 | 2339.9823 | 2539.5782 | 2830.3333 | 3465.9617 | 4247.6195 | 4338.8732 | 3787.056 | 2170.2006 | 2177.2242 | 2233.823 | 2170.2006 | 2291.823 | 2351.8909 | 2265.9469 | 2254.3038 | 2255.0324 | 2325.2035 | 2221.7286 | 2207.5251 | 2138.7551 | 2173.8968 | 2157.6018 | 2173.1947 | 2132.1062 | 2229.9528 | 2123.7758 | 2205.8496 | 2277.8761 | 1885.7965 | 1949.528 | 2107.354 | 2096.2507 | 1931.5103 | 1947.823 | 1862.2242 | 2319.823 | 2939.5251 | 2923.7758 | 2485.4071
122 | 3.3364 | 971.3564 | 2334.6136 | 2498.3338 | 2434.3323 | 2401.1804 | 3294.6962 | 2723.8722 | 2132.3624 | 2031.8737 | 2475.8196 | 2139.5609 | 2016.0285 | 2673.8751 | 2482.2045 | 3038.5098 | 2416.7444 | 2528.5865 | 2425.6121 | 2429.0241 | 2078.5157 | 2168.7429 | 3288.6677 | 3326.9459 | 2195.2256 | 2710.2075 | 2893.2571 | 2597.179 | 2888.2527 | 2283.7458 | 2421.4602 | 2407.8707 | 2243.5624 | 2059.8181 | 2482.7895 | 2616.3053 | 2291.6691 | 3135.7955 | 3536.3624 | 3127.8633 | 2927.2542 | 2351.7429 | 2921.9789 | 2687.97 | 2420.6121 | 2111.7188 | 2204.9489 | 2158.027 | 2339.2045 | 2204.4631 | 2641.4857 | 2373.0993 | 2253.8436 | 2354.6873 | 2941.7955 | 2469.0993 | 2294.618 | 2420.5113 | 2969.3535 | 2236.9218 | 2276.1067 | 2194.7158 | 2507.4346 | 2237.8166 | 2178.2331 | 2338.9744 | 2885.8948 | 2478.7684 | 2348.8692 | 2305.1804 | 2943.8737 | 2485.382 | 2389.3579 | 2538.9144 | 2651.3023 | 2613.1067 | 2725.1037 | 2314.7699 | 2350.9218 | 2410.203 | 2484.0285 | 2198.1819 | 2287.3083 | 2343.5098 | 2236.6421 | 2149.615 | 2126.7384 | 2201.3053 | 2165.2316 | 2298.3835 | 2330.6406 | 2268.8692 | 2313.203 | 2175.621 | 2326.4902 | 2525.2827 | 2796.8225 | 3467.9415 | 4272.3835 | 4392.9263 | 3817.0826 | 2145.0496 | 2198.1037 | 2230.9474 | 2145.0496 | 2278.2286 | 2379.9715 | 2228.5595 | 2229.9233 | 2160.5865 | 2294.6406 | 2216.4631 | 2159.2527 | 2141.9518 | 2182.8707 | 2141.5895 | 2142.7895 | 2085.0767 | 2152.9955 | 2135.0526 | 2196.3323 | 2259.7684 | 1928.0526 | 1854.4302 | 2069.2556 | 2106.1037 | 1875.5398 | 1938.1263 | 1852.8963 | 2298.9203 | 2962.3864 | 2954.5413 | 2454.4887
123 | 3.3474 | 917.0197 | 2380.4307 | 2482.5221 | 2421.4995 | 2441.9725 | 3377.9695 | 2675.0531 | 2160.5172 | 2047.0098 | 2438.0246 | 2072.5241 | 1995.059 | 2671.0472 | 2474.999 | 3007.9902 | 2379.0472 | 2521.4641 | 2484.3983 | 2386.0216 | 2055.0619 | 2153.0059 | 3237.5634 | 3330.9587 | 2118.5074 | 2697.5437 | 2896.0069 | 2603.4867 | 2837.001 | 2320.4936 | 2443.4602 | 2474.4199 | 2219.002 | 2034.9921 | 2442.469 | 2586.9971 | 2299.0256 | 3111.0324 | 3490.5899 | 3069.944 | 2871.5251 | 2315.526 | 2949.0334 | 2660.9685 | 2360.0157 | 2081.0413 | 2257.4504 | 2172.5044 | 2334.9961 | 2236.4897 | 2600.5211 | 2352.9764 | 2238.9813 | 2352.4356 | 2954.4956 | 2445.4799 | 2340.9892 | 2395.5044 | 2838.5555 | 2186.0334 | 2336.4956 | 2159.0069 | 2500.9892 | 2148.0364 | 2224.9843 | 2315.0246 | 2865.4848 | 2486.9813 | 2375.9391 | 2303.0147 | 2888.5791 | 2514.9361 | 2364.0108 | 2498.4248 | 2614.4651 | 2623.5447 | 2729.0118 | 2325.4985 | 2347.9862 | 2341.0403 | 2498.5241 | 2210.0206 | 2271.0236 | 2250.0521 | 2274.9764 | 2143.5015 | 2161.3973 | 2161.5074 | 2157.0187 | 2243.0393 | 2228.0953 | 2270.4641 | 2273.0118 | 2211.5388 | 2347.0197 | 2500.5329 | 2757.0659 | 3372.9971 | 4245.0118 | 4396.0393 | 3924.9695 | 2106.5162 | 2222.9912 | 2233.9794 | 2106.5162 | 2248.5005 | 2341 | 2226.4405 | 2215.0177 | 2121.4956 | 2268.0206 | 2209.0285 | 2170.4396 | 2157.0265 | 2168 | 2144.4936 | 2091.4799 | 2101.9803 | 2095.9971 | 2122.53 | 2178.0049 | 2245.5034 | 1937.0088 | 1839.4386 | 2059.4995 | 2106.5152 | 1891.0029 | 1931.9813 | 1820.0167 | 2295.9666 | 2959.527 | 2987.0059 | 2437.0373
124 | 3.3584 | 933.3378 | 2490.9204 | 2503.503 | 2430.7006 | 2410.6312 | 3339.3952 | 2607.6932 | 2129.1814 | 2053.8894 | 2421.7006 | 2017.292 | 1954.4484 | 2648.59 | 2492.8894 | 3010.8348 | 2354.031 | 2585.1254 | 2491.2699 | 2384.472 | 2020.149 | 2148.0236 | 3185.795 | 3276.7817 | 2155.0148 | 2647.3702 | 2885.4174 | 2581.174 | 2833.9528 | 2296.292 | 2433.8436 | 2416.7905 | 2209.8348 | 2029.6932 | 2455.0634 | 2596.1416 | 2265.323 | 3076.9764 | 3450.18 | 3134.677 | 2914.5812 | 2329.944 | 2883.2684 | 2712.9602 | 2317.8982 | 2047.1888 | 2277.292 | 2216.621 | 2334.9056 | 2253.1416 | 2612.7788 | 2391.3304 | 2239.0634 | 2384.733 | 2910.8908 | 2473.1652 | 2328.969 | 2371.0398 | 2908.4145 | 2167.354 | 2364.031 | 2169.913 | 2482.3156 | 2164.739 | 2208.2448 | 2282.323 | 2846.1976 | 2451.7492 | 2359.1829 | 2278.7876 | 2818.236 | 2502.7183 | 2347.882 | 2569.882 | 2599.3318 | 2530.4026 | 2711.3702 | 2292.7094 | 2306.2138 | 2329.6844 | 2458.1342 | 2211.0074 | 2274.1254 | 2249.2036 | 2334.826 | 2167.59 | 2238.3628 | 2143.2522 | 2153.8658 | 2168.1976 | 2208.2817 | 2291.1342 | 2257.4174 | 2166.3702 | 2311.1342 | 2521.7626 | 2727.3614 | 3373.441 | 4221.2286 | 4376.9838 | 3881.2773 | 2098.7006 | 2164.9542 | 2223.7802 | 2098.7006 | 2211.1504 | 2293.9144 | 2233.261 | 2197 | 2190.9986 | 2222.9454 | 2184.0944 | 2199.2448 | 2099.292 | 2172.6062 | 2128.9926 | 2126.3304 | 2060.584 | 2143.5266 | 2090.9764 | 2204.7316 | 2210.2684 | 1883.9852 | 1881.528 | 2014.9616 | 2114.5428 | 1904.3776 | 1887.025 | 1817.8422 | 2305.9454 | 2902.3156 | 2945.174 | 2355.497
125 | 3.3694 | 932.3618 | 2466.1091 | 2450.6637 | 2404.41 | 2379.6804 | 3247.3196 | 2601.2891 | 2114.9518 | 2034.8859 | 2453.5841 | 2037.2645 | 1945.2232 | 2635.2655 | 2485.94 | 2973.6027 | 2380.7168 | 2573.2655 | 2429.3973 | 2385.1327 | 2066.9154 | 2119.1268 | 3222.1082 | 3244.2222 | 2184.9459 | 2663.4877 | 2810.1996 | 2549.0964 | 2817.3255 | 2227.1691 | 2375.5182 | 2296.5782 | 2227.1082 | 1983.4582 | 2439.6205 | 2610.9095 | 2241.7532 | 3116.8791 | 3485.8977 | 3125.7532 | 2988.9095 | 2330.2173 | 2908.0295 | 2713.1327 | 2273.0846 | 2045.0423 | 2238.5664 | 2220.5487 | 2264.9646 | 2230.175 | 2656.0059 | 2378.2714 | 2159.6814 | 2310.0246 | 2879.6745 | 2428.6391 | 2289.1809 | 2354.0964 | 2979.2114 | 2155.2714 | 2322.585 | 2195.9095 | 2463.9577 | 2252.6863 | 2189.0541 | 2302.2527 | 2834.4877 | 2390.5664 | 2319.6382 | 2270 | 2781.88 | 2455.7286 | 2410.0836 | 2570.2409 | 2560.9577 | 2522.2104 | 2738.4936 | 2247.4218 | 2261.3855 | 2313.4523 | 2460.295 | 2226.235 | 2263.3137 | 2296.3795 | 2323.9282 | 2123.8682 | 2142.2959 | 2133 | 2164.8073 | 2165.9214 | 2214.9164 | 2219.3982 | 2207.88 | 2134.2714 | 2297.5723 | 2535.2173 | 2779.7709 | 3335.8859 | 4232.0118 | 4413.7709 | 3791.6568 | 2118.53 | 2092.0905 | 2151.0246 | 2118.53 | 2168.6627 | 2294.0718 | 2183.0482 | 2212.7227 | 2203.2114 | 2186.8977 | 2182.235 | 2206.7827 | 2060.5664 | 2167.0423 | 2126.8673 | 2122.1809 | 2038.1623 | 2100.5732 | 2073.7532 | 2150.6214 | 2200.9636 | 1838.3314 | 1882.4395 | 1975.6686 | 2091.4041 | 1901.1327 | 1833.3373 | 1843.0059 | 2234.2596 | 2866.6627 | 2957.12 | 2356.4513
126 | 3.3804 | 941.3014 | 2396.1268 | 2361.2453 | 2421.2744 | 2430.8919 | 3255.1101 | 2645.0831 | 2117.7547 | 1962.235 | 2426.1893 | 2053.5718 | 1930.7276 | 2602.6091 | 2455.3451 | 3019.4843 | 2373.736 | 2457.47 | 2411.4076 | 2327.7714 | 2055.499 | 2123.1101 | 3210.208 | 3292.6362 | 2135.3805 | 2667.4449 | 2815.8024 | 2569.237 | 2821.1642 | 2163.6445 | 2365.4636 | 2335.8919 | 2255.973 | 1941.1996 | 2425.9085 | 2605.0541 | 2243.6278 | 3162.0457 | 3543.7296 | 3108.3181 | 2911.1789 | 2261.208 | 2945.1268 | 2665.9626 | 2280.7463 | 2022.5905 | 2224.6362 | 2184.1455 | 2231.2827 | 2283.9209 | 2632.7443 | 2319.1352 | 2160.3201 | 2262.4179 | 2919.8191 | 2372.8815 | 2270.8545 | 2392.474 | 2900.7443 | 2147.6278 | 2285.9189 | 2133.734 | 2451.0083 | 2210.3702 | 2172.6445 | 2343.6278 | 2887.792 | 2416.8648 | 2336.9813 | 2230.8441 | 2819.9209 | 2454.2537 | 2381.9606 | 2515.4616 | 2551.763 | 2558.6819 | 2670.3785 | 2211.8628 | 2266.8545 | 2230.6258 | 2438.9897 | 2185.5988 | 2204.2537 | 2277.9085 | 2308.2099 | 2062.5364 | 2115.474 | 2136.7547 | 2105.7527 | 2159.6716 | 2150.7817 | 2191.6008 | 2185.8461 | 2097.1268 | 2266.8171 | 2504.8274 | 2733.8525 | 3449.8958 | 4284.1829 | 4423.0998 | 3760.5177 | 2088.7714 | 2044.3078 | 2061.0437 | 2088.7714 | 2175.4095 | 2264.734 | 2189.3638 | 2164.2433 | 2127.9626 | 2151.3992 | 2171.0998 | 2154.5698 | 2082.5738 | 2129.0354 | 2162.237 | 2093.6632 | 2055.9813 | 2051.9646 | 2043.9813 | 2055.2266 | 2227.5093 | 1837.1455 | 1879.5447 | 1966.1268 | 2009.7714 | 1792.2787 | 1859.3097 | 1809.5531 | 2254.3658 | 2864.8274 | 2968.4262 | 2415.3825
127 | 3.3914 | 951.7611 | 2380.0619 | 2376.6549 | 2433.3717 | 2414.9469 | 3346.7434 | 2686.7788 | 2114.4159 | 1982.9646 | 2338.4602 | 2017.4779 | 1906.531 | 2610.5841 | 2416.823 | 3075.8673 | 2361.2566 | 2442.416 | 2368.708 | 2310.0089 | 2026.708 | 2129.1504 | 3211.177 | 3259.469 | 2123.3982 | 2647.9646 | 2888.5221 | 2582.8053 | 2802.531 | 2164.6903 | 2387.4956 | 2333.115 | 2202.1947 | 1929.354 | 2438.3363 | 2516.1858 | 2254.1593 | 3119.6194 | 3537.6106 | 3131.7876 | 2822.1858 | 2245.1681 | 2916.9911 | 2649.3274 | 2302.4513 | 1988.5044 | 2161.8407 | 2168.9911 | 2247.4159 | 2260.2743 | 2583.1504 | 2282.4159 | 2187.0531 | 2280.7876 | 2924.3451 | 2374.3274 | 2245.4071 | 2382.8495 | 2959.239 | 2175.1681 | 2288.6372 | 2144.8761 | 2440.2301 | 2114.823 | 2156.1327 | 2316.3009 | 2852.8053 | 2372.3805 | 2345.9469 | 2225.531 | 2820.9646 | 2447.3628 | 2330.7876 | 2400.354 | 2593.062 | 2506.1327 | 2585.4513 | 2215.1416 | 2235.3717 | 2237.3452 | 2398.2212 | 2139.0885 | 2223.6991 | 2185.5398 | 2258.1239 | 2065.7434 | 2102.5575 | 2129.0265 | 2032.7611 | 2151.9469 | 2166.6549 | 2217.4867 | 2185.9911 | 2078.3186 | 2239.5487 | 2552.8938 | 2748.4868 | 3427.8583 | 4237.9292 | 4380.1416 | 3756.7788 | 2082.531 | 2064.8673 | 2056.5044 | 2082.531 | 2180.4159 | 2251.7345 | 2177.5398 | 2126.4602 | 2098.1593 | 2114.5398 | 2148.4425 | 2072.9646 | 2074.9823 | 2071.0796 | 2108.3186 | 2095.6195 | 2050.6814 | 2076.4513 | 2022.2212 | 2090.5752 | 2213.6991 | 1841.1947 | 1837.4071 | 1955.5487 | 2012.3098 | 1734.9469 | 1861.1858 | 1790.6549 | 2265.5575 | 2938.6814 | 2956.8142 | 2402.2743
128 | 3.4024 | 943.3412 | 2318.7817 | 2420.0482 | 2368.0496 | 2389.2925 | 3337.172 | 2648.5619 | 2086.5614 | 1977.9523 | 2336.9017 | 1989.4882 | 1922.1942 | 2586.8299 | 2404.3412 | 3015.0255 | 2384.8043 | 2499.2193 | 2355.9027 | 2359.3407 | 2026.2439 | 2154.0723 | 3209.6347 | 3219.732 | 2143.2439 | 2630.6588 | 2820.587 | 2509.1966 | 2789.5851 | 2167.0492 | 2380.1957 | 2342.9749 | 2124.3908 | 1948.1947 | 2466.1706 | 2495.6333 | 2313.828 | 3105.9263 | 3435.3181 | 3109.8058 | 2829.6337 | 2261.0487 | 2896.8294 | 2613.8058 | 2282.9272 | 1976.7316 | 2151.0236 | 2134.8053 | 2217.8786 | 2184.561 | 2602.4627 | 2214.172 | 2119.0742 | 2303.6829 | 2909 | 2375.8294 | 2251.8043 | 2350.6829 | 2928.9046 | 2145.6106 | 2240.6593 | 2164.7325 | 2422.8786 | 2092.8048 | 2175.4631 | 2234.3181 | 2787.0978 | 2324.4631 | 2306.2207 | 2223.7566 | 2781.268 | 2438.6834 | 2313.8545 | 2368.1947 | 2572.1234 | 2518.7793 | 2640.5354 | 2223.2198 | 2215.4877 | 2176.5142 | 2409.5118 | 2100.7566 | 2212.001 | 2180.4622 | 2185.5369 | 2056.7807 | 2016.6352 | 2120 | 2047.9017 | 2192.9995 | 2171.3417 | 2206.854 | 2137.7812 | 2116.8771 | 2256.8289 | 2564.8786 | 2769.7812 | 3383.413 | 4219.6333 | 4389.3894 | 3751.6588 | 2077.3908 | 2071.2935 | 2058.147 | 2077.3908 | 2126.1961 | 2228.733 | 2090.5137 | 2166.9263 | 2098.9267 | 2135.2188 | 2139.6096 | 2063.5123 | 2021.1957 | 2042.8781 | 2075.1214 | 2066.7325 | 2053.0733 | 2105.7311 | 2017.2439 | 2111.1224 | 2205.4145 | 1825.0492 | 1792.7566 | 1944.2198 | 2011.586 | 1781.0482 | 1844.5123 | 1748.2207 | 2245.8289 | 2923.0747 | 2903.3672 | 2342.2202
129 | 3.4134 | 970.1976 | 2332.472 | 2417.9233 | 2330.2675 | 2445.4464 | 3234.6873 | 2600.9616 | 2064.414 | 1972.9685 | 2367.293 | 2007.1023 | 1921.6047 | 2561.707 | 2432.3442 | 3004.1023 | 2360.8594 | 2463.2861 | 2337.1082 | 2381.8407 | 2010.8024 | 2097.5722 | 3148.8722 | 3251.1406 | 2109.4582 | 2614.1209 | 2773.7581 | 2519.4976 | 2751.7001 | 2107.7129 | 2356.5605 | 2287.1711 | 2103.7198 | 1950.2419 | 2446.9233 | 2568.5546 | 2317.7129 | 3081.7129 | 3414.6244 | 3038.1652 | 2896.5929 | 2230.0315 | 2873.0826 | 2561.6558 | 2257.6814 | 1943.5919 | 2119.235 | 2090.3373 | 2209.3186 | 2125.9547 | 2602.3373 | 2261.8801 | 2051.9105 | 2266.0187 | 2828.7443 | 2334.8849 | 2251.9361 | 2342.8279 | 2867.3442 | 2097.3756 | 2223.3186 | 2116.2291 | 2449.2871 | 2060.6047 | 2114.9164 | 2244.0069 | 2767.2547 | 2371.1151 | 2290.8093 | 2187.9489 | 2715.4641 | 2424.293 | 2367.6952 | 2373.6814 | 2580.9115 | 2530.178 | 2657.6303 | 2187.6303 | 2256.2616 | 2152.0197 | 2464.1278 | 2093.6116 | 2152.4965 | 2158.7827 | 2176.7198 | 2096.1534 | 2025.0069 | 2081.0187 | 2052.4965 | 2138.1966 | 2155.5477 | 2151.4326 | 2115.586 | 2128.4965 | 2210.9547 | 2488.1013 | 2737.2675 | 3363.7443 | 4216.8849 | 4374.5664 | 3818.8162 | 2055 | 2046.7325 | 2078.7139 | 2055 | 2088.5733 | 2208.6627 | 2101.2173 | 2188.5733 | 2077.9233 | 2187.2232 | 2116.4838 | 2052.5477 | 2039.9813 | 2031.1082 | 2105.4651 | 2045.7453 | 2027.6431 | 2097.471 | 2025.3058 | 2106.6116 | 2168.4326 | 1828.7709 | 1793.0639 | 1918.9489 | 1974.9616 | 1774.3373 | 1896.6057 | 1805.6126 | 2270.4651 | 2812.6873 | 2878.3697 | 2338.5035
130 | 3.4244 | 935.5708 | 2412.3038 | 2402.3422 | 2328.3422 | 2456.0649 | 3199.6844 | 2572.0914 | 2051.0457 | 1940.4454 | 2376.0266 | 1985.913 | 1893.6578 | 2542.4336 | 2403.8407 | 2890.9822 | 2330.9277 | 2461.0339 | 2359.4484 | 2352.236 | 1983.3614 | 2051.4941 | 3165.8437 | 3281.0723 | 2098.2242 | 2606.0723 | 2732.472 | 2530.9395 | 2706.8024 | 2087.4676 | 2336.4336 | 2166.1032 | 2125.4941 | 1961.0797 | 2471.7449 | 2576.9203 | 2255.3082 | 3055.6121 | 3476.7714 | 3066.798 | 2891.0841 | 2288.3186 | 2895.8894 | 2576.6918 | 2262.882 | 1935.4941 | 2063.4602 | 2076.6121 | 2193.5782 | 2115.9617 | 2566.9469 | 2266.6622 | 2064.3761 | 2205.9661 | 2827.5546 | 2346.0605 | 2217.0914 | 2287.3997 | 2841.1829 | 2106.6652 | 2192.354 | 2097.1445 | 2438.4985 | 2081.8437 | 2103.5738 | 2298.9086 | 2778.1519 | 2407.3687 | 2341.6918 | 2178.9277 | 2742.4897 | 2398.4793 | 2297.0811 | 2385.9086 | 2602.4071 | 2506.7566 | 2610.4793 | 2151.0192 | 2230.1298 | 2143.1563 | 2433.6888 | 2053.8141 | 2135.5855 | 2140.7906 | 2134.6696 | 2085.1563 | 2071.1254 | 2043.2168 | 2075.9882 | 2125.5738 | 2130.677 | 2078.0118 | 2098.5059 | 2018.4837 | 2205.646 | 2420.4071 | 2740.6121 | 3270.6239 | 4190 | 4440.1284 | 3877.7566 | 2063.7832 | 2034.0192 | 2055.202 | 2063.7832 | 2067.3348 | 2251.6652 | 2116.2817 | 2156.2094 | 2029.5516 | 2210.6844 | 2086.0457 | 2048.1711 | 2053.1903 | 2059.8894 | 2085.8938 | 2006.1563 | 1995.7758 | 2027.1563 | 2009.5059 | 2037.5368 | 2214.4632 | 1801.5973 | 1836.6918 | 1945.0605 | 2015.1859 | 1769.9808 | 1879.4454 | 1757.2787 | 2253.8215 | 2797.5929 | 2930.7906 | 2351.3156
131 | 3.4354 | 937.2724 | 2390.7709 | 2343.6185 | 2416.6864 | 2466.294 | 3153.9774 | 2545.4454 | 2041.6087 | 1875.0865 | 2331.9774 | 1964.1632 | 1860.3146 | 2562.881 | 2434.3815 | 2854.3815 | 2285.173 | 2466.8584 | 2330.7817 | 2344.1632 | 1985.5546 | 2072.5762 | 3176.2822 | 3225.1947 | 2164.001 | 2655.414 | 2801.2193 | 2463.3471 | 2793.0669 | 2120.9135 | 2302.4779 | 2211.4799 | 2153.7503 | 1947.9017 | 2404.5526 | 2487.2704 | 2199.7276 | 3087.3048 | 3518.5762 | 3075.9125 | 2802.0098 | 2271.4769 | 2856.2379 | 2598.7935 | 2278.5762 | 1906.3579 | 2104.4248 | 2139.9902 | 2229.0118 | 2111.9017 | 2583.3048 | 2214.8151 | 2110.7178 | 2223.4572 | 2915.6313 | 2331.1514 | 2166.5319 | 2265.5546 | 2863.8161 | 2130.9676 | 2136.9017 | 2132.4572 | 2356.1839 | 2110.8151 | 2076.2055 | 2292.0757 | 2771.8692 | 2357.8033 | 2319.5536 | 2170.2389 | 2754.5752 | 2379 | 2272.9469 | 2389.8368 | 2573.0433 | 2542.0767 | 2647.7945 | 2109.1081 | 2214.12 | 2130.3265 | 2343.173 | 2020.7935 | 2143.7719 | 2142.4779 | 2115.9567 | 1982.0531 | 1986.9331 | 2013.0865 | 2034.4552 | 2125.706 | 2068.5752 | 2113.5664 | 2060.6844 | 1933.2714 | 2191.7492 | 2420.9351 | 2732.2497 | 3236.2389 | 4218.0983 | 4373.4975 | 3780.3569 | 2034.7276 | 2065.0443 | 1986.3038 | 2034.7276 | 2043.0324 | 2233.5211 | 2099.9676 | 2102.5103 | 1998.4346 | 2123.1288 | 2055.0865 | 1995.7925 | 1997.7601 | 2057.9017 | 2020.5644 | 2007.6745 | 1971.8476 | 2020.3048 | 1989.6195 | 1987.9784 | 2151.0304 | 1724.5752 | 1841.4562 | 1939.119 | 1972.879 | 1734.173 | 1808.1081 | 1658.2389 | 2183.9882 | 2833.5329 | 2897.2055 | 2282.2379
132 | 3.4464 | 915.2404 | 2323.6126 | 2268.8328 | 2376.7694 | 2409.1239 | 3160.0482 | 2481.4783 | 2065.0153 | 1894.9976 | 2278.8658 | 2010.6583 | 1833.5088 | 2534.9695 | 2449.3746 | 2972.4863 | 2228.5949 | 2454.1696 | 2226.2251 | 2336.3569 | 1942.1391 | 2026.9695 | 3099.3771 | 3178.0177 | 2161.3417 | 2677.7974 | 2853.6962 | 2456.7444 | 2780.5644 | 2119.9164 | 2295.9317 | 2231.9518 | 2138.5949 | 1885.0708 | 2410.385 | 2504.7242 | 2143.6632 | 3067.7822 | 3434.5314 | 3044.6278 | 2894.1775 | 2211.4582 | 2804.6785 | 2564.0024 | 2248.8328 | 1895.914 | 2129.0683 | 2101.5137 | 2182.853 | 2135.709 | 2574.1544 | 2194.3063 | 2096.7139 | 2239.0177 | 2835.9567 | 2312.1519 | 2142.6101 | 2297.795 | 2912.0329 | 2089.41 | 2113.7291 | 2110.7999 | 2360.4381 | 2073.2227 | 2013.5949 | 2256.6455 | 2714.4607 | 2326.1013 | 2227.5467 | 2112.3594 | 2741.5241 | 2356.4253 | 2241.2935 | 2327.4783 | 2556.4582 | 2500.9366 | 2732.3697 | 2133.5393 | 2241.2026 | 2151.9317 | 2306.119 | 2073.8609 | 2100.7999 | 2106.0354 | 2166.709 | 1960.6254 | 1982.9622 | 1963.443 | 1931.4783 | 2063.291 | 2106.7749 | 2162.4912 | 2060.2532 | 1979.7926 | 2170.7114 | 2451.914 | 2691.2733 | 3280.8609 | 4263.8456 | 4343.2331 | 3721.3215 | 2025.0329 | 2011.0226 | 1998.7267 | 2025.0329 | 2047.3722 | 2217.1013 | 2065.1214 | 2076.5088 | 1994.4405 | 2066.2709 | 2024.3569 | 2049.0787 | 1913.0885 | 1982.8682 | 1972.3063 | 2001.7139 | 1948.3569 | 2010.5442 | 1926.2074 | 1998.1013 | 2143.6052 | 1720.676 | 1782.7999 | 1892.5949 | 1967.9799 | 1730.676 | 1801.4228 | 1683.3368 | 2144.7291 | 2856.9823 | 2887.0659 | 2279.4887
133 | 3.4574 | 942.6401 | 2343.3245 | 2306.1298 | 2322.7139 | 2380.5398 | 3186.8673 | 2535.0059 | 2063.5723 | 1853.8908 | 2305.2124 | 1978.7374 | 1801.3687 | 2517.3156 | 2418.0619 | 2911.8318 | 2203.0413 | 2459.2242 | 2278.7198 | 2319.531 | 1951.3245 | 1980.0619 | 3074.3156 | 3182.826 | 2076.2478 | 2613.9823 | 2771.6666 | 2543.4455 | 2718.6844 | 2071.0324 | 2291.6637 | 2285.4248 | 2128.6018 | 1879.4071 | 2374.4631 | 2485.1357 | 2159.1003 | 3054.5811 | 3358.0826 | 3052.3156 | 2923.2271 | 2206.6224 | 2790.0413 | 2571.4985 | 2198.3687 | 1900.531 | 2082.2979 | 2071.9174 | 2145.9381 | 2082.0855 | 2580.6519 | 2213.4985 | 2076.7552 | 2190.9203 | 2788.0295 | 2308.6637 | 2086.3599 | 2314.2448 | 2807.7197 | 2041.5723 | 2122.2035 | 2081.8673 | 2322.1357 | 2057.2242 | 1988.6637 | 2196.941 | 2700.9381 | 2291.41 | 2219.3658 | 2092.4484 | 2770.1829 | 2352.5811 | 2247.6401 | 2268.3481 | 2496.2271 | 2492.3658 | 2664.2802 | 2148.9292 | 2247.7552 | 2151.3982 | 2293.6637 | 2064.6755 | 2139.7109 | 2060.879 | 2141.7168 | 1978.8673 | 2026.2242 | 1975.6106 | 1968.2006 | 2066.4985 | 2056.7699 | 2108.7374 | 2067.5103 | 2000.5723 | 2187.4484 | 2460.2655 | 2666.6431 | 3241.177 | 4242.41 | 4340.6755 | 3813.7611 | 2011.5929 | 1975.4277 | 1975.4307 | 2011.5929 | 2062.7139 | 2199.2153 | 2075.7227 | 2008.2684 | 1954.2979 | 2014.2271 | 2019.9794 | 2006.4129 | 1930.2832 | 1920.8584 | 1932.3687 | 1976.1534 | 1896.6755 | 1969.9705 | 1941.528 | 1982.705 | 2186.2655 | 1776.9676 | 1773.1623 | 1912.4779 | 1955.6755 | 1765.1829 | 1809.3776 | 1699.6431 | 2135.1534 | 2839.7257 | 2885.0826 | 2276.1239
134 | 3.4684 | 967.4793 | 2287.1224 | 2361.676 | 2302.6116 | 2396.8083 | 3191.8083 | 2538.2232 | 2025.0728 | 1834.0413 | 2310.5934 | 1861.4877 | 1764.9587 | 2532.9818 | 2427.4066 | 2910.3751 | 2210.3289 | 2477.5983 | 2257.3918 | 2294.0546 | 1961.1322 | 1960.8446 | 3037.9356 | 3148.7438 | 2107.914 | 2566.6711 | 2758.1967 | 2489.6199 | 2746.1785 | 2029.5934 | 2277.1141 | 2273.972 | 2071.7025 | 1860.6298 | 2312.3058 | 2514.914 | 2113.7207 | 3113.6809 | 3458.9784 | 3024.8215 | 2845.5703 | 2226.0413 | 2792.2512 | 2591.2694 | 2184.8083 | 1883.3058 | 2008.762 | 2078.2876 | 2184.1967 | 2053.9454 | 2539.7802 | 2202.1686 | 1998.5654 | 2258.033 | 2754.3967 | 2221.7571 | 2058.3471 | 2216.6199 | 2802.471 | 2015.767 | 2116.3058 | 2052.5339 | 2308.7537 | 1972.1406 | 1951.8992 | 2207.0826 | 2703.0182 | 2303.5438 | 2175.529 | 2044.7802 | 2720.972 | 2353.9223 | 2283.9041 | 2333.3388 | 2529.6446 | 2444.086 | 2589.1322 | 2154.9636 | 2227.9587 | 2097.2232 | 2275.9405 | 1999.0133 | 2140.1091 | 2076.2743 | 2113.8083 | 1920.972 | 1972.2414 | 1974.4199 | 1979.6298 | 2071.0364 | 2035.8132 | 2039.7257 | 2060.9223 | 1992.5388 | 2175.9587 | 2466.8859 | 2657.0182 | 3240.1603 | 4236.1372 | 4346.7355 | 3811.7802 | 1933.8943 | 1949.8215 | 1914.0909 | 1933.8943 | 2043.2463 | 2182.1917 | 2089.7488 | 1969.6347 | 1970.8909 | 2011.4661 | 2002.0546 | 1979.0049 | 1979.1372 | 1886.5934 | 1935.9454 | 2001.8314 | 1895.119 | 2005.5074 | 1951.1504 | 2008.4479 | 2088.1588 | 1716.2183 | 1787.2694 | 1907.9769 | 1907.1504 | 1703.2778 | 1762.0314 | 1688.7488 | 2165.9091 | 2818.0182 | 2827.9356 | 2277.676
135 | 3.4794 | 909.5929 | 2275.5851 | 2278.5378 | 2284.118 | 2355.6509 | 3115.7129 | 2490.237 | 2022.587 | 1850.823 | 2268.9439 | 1854.7611 | 1754.647 | 2540.882 | 2359.0659 | 2934.413 | 2183.7089 | 2460.12 | 2278.2881 | 2208.8889 | 1942.942 | 1978.939 | 3065.7011 | 3138.587 | 2085.5349 | 2612.7011 | 2731.0039 | 2455.645 | 2713.9459 | 2011.236 | 2284.94 | 2193.4159 | 2085.5251 | 1832.942 | 2352.8781 | 2519.7089 | 2071.765 | 3101.9449 | 3426.9498 | 3034.4091 | 2805.119 | 2196.8859 | 2857.1711 | 2595.882 | 2176.648 | 1853.355 | 2026.9371 | 2095.704 | 2137.5939 | 2037.5319 | 2533.88 | 2180.471 | 2032.9341 | 2246.9469 | 2791.7001 | 2217.7021 | 2098.8791 | 2211.7601 | 2813.12 | 2044.5261 | 2151.7021 | 2028.53 | 2326.882 | 1944.292 | 1999.9941 | 2208.296 | 2662.1209 | 2253.4769 | 2197.8771 | 2049.8791 | 2666.296 | 2343.883 | 2250.0619 | 2304.5949 | 2529.8269 | 2415.352 | 2572.236 | 2159.765 | 2202.766 | 2101.6441 | 2224.3569 | 1993.469 | 2084.2969 | 2018.3599 | 2076.5329 | 1903.175 | 1948.058 | 2002.1141 | 1956.471 | 2011.2399 | 2082.88 | 2045.88 | 2045.06 | 2019.233 | 2184.41 | 2450.649 | 2664.646 | 3235.2379 | 4190.3579 | 4344.296 | 3802.645 | 1907.999 | 1907.649 | 1953.5241 | 1907.999 | 2005.708 | 2116.5939 | 2071.413 | 2012.7021 | 1980.119 | 2021.883 | 2010.057 | 1981.178 | 1910.7129 | 1910.9381 | 1929.884 | 1967.2399 | 1941.9971 | 1992.5919 | 1877.1819 | 1973.7109 | 2054.998 | 1681.587 | 1730.4169 | 1869.061 | 1881.707 | 1641.296 | 1777.7021 | 1662.649 | 2169.825 | 2806.883 | 2873.8171 | 2256.6499
136 | 3.4904 | 891.5487 | 2296.9779 | 2258.646 | 2276.0664 | 2355.7345 | 3124.4248 | 2420.6637 | 2047.8009 | 1828.9469 | 2196.6194 | 1849.3097 | 1753.0221 | 2444.7655 | 2302.8142 | 2860.4115 | 2228.2921 | 2434.792 | 2316.3186 | 2247.4558 | 1883.938 | 1947.8274 | 3119.3938 | 3143.3628 | 2040.1327 | 2583.9823 | 2668.761 | 2412.6194 | 2745.6549 | 2032.7124 | 2299.2522 | 2198.3274 | 2080.354 | 1847.7788 | 2336.8053 | 2499.6814 | 2136.5885 | 3081.6372 | 3410.2611 | 3035.177 | 2843.8098 | 2200.0531 | 2828.6858 | 2553.1681 | 2126.1681 | 1830.4513 | 2070.4602 | 2068.761 | 2153.7655 | 2091.885 | 2574.3717 | 2191.8451 | 2042.3761 | 2167.8274 | 2841.5044 | 2221.2212 | 2138.1195 | 2228.1327 | 2826.6903 | 1993.7964 | 2104.4557 | 2069.1726 | 2332.2965 | 1973.4823 | 2039.2965 | 2220.7345 | 2674.3053 | 2288.8407 | 2205.6062 | 2068.0221 | 2675.4381 | 2313.6283 | 2234.2965 | 2303.5575 | 2416.0398 | 2405.8363 | 2622.7213 | 2083.5 | 2198.2965 | 2132.5266 | 2284.3894 | 1975.3097 | 2057.7699 | 1991.1637 | 2127.5222 | 1904.7478 | 1937.8363 | 2030.5929 | 1930.2655 | 2069.3673 | 2029.0929 | 1974.292 | 2055.8009 | 2029.0442 | 2144.8938 | 2419.2212 | 2636.0575 | 3216.3628 | 4110.5531 | 4282.2345 | 3836.7566 | 1895.5398 | 1878.8584 | 1931.2787 | 1895.5398 | 1987.7257 | 2140.677 | 2023.1239 | 2043.2522 | 1952.2212 | 2082.6106 | 1949.7964 | 1993.0974 | 1903.6018 | 1914.177 | 1895.6062 | 1927.1327 | 1911.2124 | 1906.0486 | 1907.9514 | 1957.5044 | 2059.0885 | 1711.4159 | 1743.8319 | 1771.6106 | 1880.9336 | 1699.2257 | 1802.9779 | 1653.6372 | 2120.1239 | 2816.8451 | 2845.0929 | 2206.6947
137 | 3.5014 | 893.6421 | 2310.7601 | 2227.9892 | 2238.0777 | 2412.9961 | 3137.5683 | 2477.9774 | 2032.1514 | 1761.6017 | 2213.6677 | 1790.6755 | 1745.9557 | 2464.5163 | 2293.3137 | 2879.7267 | 2254.9557 | 2426.9852 | 2282.7492 | 2243.6755 | 1848.6126 | 1927.6863 | 3147.4464 | 3121.8525 | 2007.8082 | 2565.4169 | 2646.3137 | 2414.2213 | 2752.1514 | 1995.6312 | 2229.7974 | 2136.0963 | 2119.0698 | 1815.0187 | 2298.2104 | 2433.8564 | 2084.3657 | 3127.9961 | 3396.764 | 3036.4022 | 2800.1701 | 2183.4503 | 2832.5939 | 2511.1957 | 2100.9852 | 1802.8377 | 2061.8672 | 1958.3362 | 2106.7974 | 2046.7384 | 2532.2586 | 2185.5683 | 2036.4169 | 2174.5497 | 2889.295 | 2204.2842 | 2148.6716 | 2242.8191 | 2818.8377 | 1971.7748 | 2029.0482 | 2044.9892 | 2302.4356 | 1914.1701 | 1991.9744 | 2138.3067 | 2636.6017 | 2318.941 | 2217.8486 | 2028.7197 | 2706.1475 | 2273.8082 | 2215.8525 | 2373.0256 | 2496.944 | 2390.6126 | 2639.5831 | 2062.4612 | 2152.9892 | 2033.5614 | 2310.5388 | 1946.8967 | 2007.3323 | 2076.9037 | 2091.7974 | 1933.236 | 1865.528 | 1909.0855 | 1885.0924 | 1990.4287 | 2021.5349 | 1925.2989 | 2048.2104 | 1998.7935 | 2090.1219 | 2470.7857 | 2553.528 | 3261.3393 | 4064.1809 | 4268.0885 | 3844.6421 | 1928.9813 | 1882.0885 | 1849.9449 | 1928.9813 | 1989.0295 | 2139.4356 | 2032.8928 | 2003.3471 | 1925.2104 | 2081.0777 | 1960.6824 | 1963.705 | 1883.7197 | 1991.2911 | 1846.7345 | 1947.1918 | 1884.6568 | 1841.4356 | 1916.1514 | 1907.9154 | 2046.941 | 1723.6716 | 1730.7345 | 1783.7414 | 1871.2399 | 1767.9076 | 1726.0963 | 1630.823 | 2141.9813 | 2787.7345 | 2851.6382 | 2242.3835
138 | 3.5124 | 898.8387 | 2280.7547 | 2217.9936 | 2263.4002 | 2380.29 | 3065.3481 | 2523.0516 | 2011.7935 | 1752.7807 | 2181.29 | 1840.8068 | 1721.7999 | 2521.5679 | 2328.2453 | 2831.4577 | 2231.4838 | 2463.6131 | 2206.5029 | 2172.761 | 1775.9287 | 1922.1613 | 2991.554 | 3178.3358 | 2031.2453 | 2619.5551 | 2765.3619 | 2423.529 | 2675.0319 | 1983.9936 | 2252.8132 | 2172.5423 | 2037.9478 | 1820.6711 | 2290.6839 | 2368.4902 | 2115.1746 | 3157.5226 | 3440.4971 | 3034.8451 | 2723.232 | 2101.9803 | 2846.5806 | 2502.6839 | 2084.2709 | 1832.3486 | 1991.4513 | 1961.8712 | 2104.5098 | 2045.9808 | 2474.2257 | 2135.232 | 2033.1096 | 2137.7094 | 2814.6253 | 2185.2709 | 2130.5354 | 2241.3741 | 2843.5615 | 2028.9228 | 1996.3225 | 1978.232 | 2332.8196 | 1939.8132 | 1941.3289 | 2126.8776 | 2620.2581 | 2246.6126 | 2193.0708 | 1999.058 | 2642.1416 | 2269.2065 | 2145.1352 | 2378.5354 | 2494.3481 | 2372.6386 | 2586.8127 | 2034.0708 | 2085.5482 | 2053.8068 | 2267.8063 | 1922.2193 | 2017.1421 | 2076.0192 | 2058.0516 | 1942.8451 | 1922.2198 | 1885.6131 | 1860.1096 | 1995.6519 | 2012.4321 | 1999.5939 | 2021.5354 | 1974.058 | 2140.9681 | 2413.4641 | 2550.7227 | 3279.8968 | 4114.1809 | 4330.7616 | 3884.7163 | 1906.232 | 1866.7999 | 1828.7871 | 1906.232 | 1989.6323 | 2110.3225 | 2000.1288 | 1982.6839 | 1915.6323 | 2000.3997 | 2036.6583 | 1904.3353 | 1835.5934 | 1914.0059 | 1836.5226 | 1911.4449 | 1881.2645 | 1802.7483 | 1904 | 1874.529 | 2012.2257 | 1725.3677 | 1722.5742 | 1794.587 | 1902.6131 | 1750.4902 | 1733.8196 | 1629.2581 | 2080.2512 | 2825.3938 | 2902.0905 | 2211.2384
139 | 3.5234 | 931.2389 | 2257.0383 | 2217.0383 | 2250.2802 | 2332.7227 | 3050.3156 | 2420.9675 | 1987.8112 | 1786.062 | 2135.8112 | 1860.7729 | 1710.6077 | 2509.0265 | 2294.1799 | 2792.823 | 2186.4956 | 2407.9026 | 2188.7463 | 2145.823 | 1799.6932 | 1983.0472 | 2933.4041 | 3082.2979 | 2027.115 | 2564.0295 | 2732.587 | 2451.6313 | 2661.0118 | 2012.2773 | 2234.7492 | 2188.2035 | 1990.0885 | 1799.5841 | 2279.1652 | 2316.1032 | 2124.6342 | 3099.7374 | 3479.1386 | 2963.0796 | 2755.9086 | 2119.7316 | 2795.5723 | 2468.8879 | 2102.062 | 1824.5457 | 2014.5162 | 1934.5723 | 2103.292 | 2046.5959 | 2457.2153 | 2153.9853 | 1957.9911 | 2155.7699 | 2771 | 2153.6342 | 2044.0295 | 2171.9528 | 2878.7965 | 2044.3422 | 2025.5044 | 1930.8495 | 2292.7374 | 1925.2301 | 1917.8614 | 2164.5811 | 2628.4808 | 2283.9204 | 2178.0885 | 1969.6342 | 2623.3156 | 2282.7463 | 2123.2655 | 2337.9764 | 2500.8968 | 2391.7581 | 2635.3127 | 2079.6549 | 2091.4543 | 2045.9262 | 2265.2271 | 1848.5604 | 1988.8377 | 2040.5959 | 2065.2655 | 1897.5339 | 1890.5988 | 1902.3038 | 1890.1121 | 1987.7109 | 1992.9498 | 2010.6342 | 1994.8997 | 1946.0265 | 2143.8495 | 2392.4926 | 2628.0472 | 3243.1917 | 4138.5192 | 4263.0177 | 3752.9321 | 1853.9764 | 1852.823 | 1843.531 | 1853.9764 | 1960.4572 | 2078.938 | 1933.4454 | 1977.4307 | 1931.708 | 2013.4661 | 1985.8141 | 1888.0885 | 1832.4425 | 1829.0649 | 1860.1888 | 1886.2153 | 1853.5457 | 1837.2508 | 1915.8348 | 1865.7345 | 2028.6313 | 1696.7611 | 1685.2301 | 1768.2035 | 1858.0413 | 1698.7994 | 1690.9528 | 1606.1534 | 2051.7463 | 2896.2508 | 2879.1032 | 2228.1623
140 | 3.5344 | 955.3353 | 2263.2099 | 2182.8279 | 2224.6647 | 2341.4199 | 3092.4228 | 2451.8048 | 1915.2389 | 1806.2099 | 2122.7493 | 1811.0757 | 1679.5364 | 2390.6062 | 2281.2099 | 2756.2856 | 2175.9592 | 2385.2099 | 2183.7901 | 2203.5103 | 1819.0408 | 1977.7025 | 3009.7203 | 3112.0148 | 1994.6209 | 2593.7611 | 2684.0816 | 2425.3264 | 2661.6239 | 2057.4228 | 2202.2478 | 2164.414 | 1991.7085 | 1796.3353 | 2298.796 | 2391.3909 | 2121.793 | 3136.7611 | 3420.9474 | 3000.3034 | 2878.1839 | 2123.997 | 2737.1603 | 2399.4488 | 2122.2099 | 1822.9184 | 1975.3643 | 1988.8048 | 2094.5831 | 2039 | 2419.1603 | 2132.7433 | 1911.414 | 2114.4897 | 2731.3264 | 2225.8894 | 2025.586 | 2161.7114 | 2825.3643 | 2050.9184 | 2065.0875 | 1919.7085 | 2264.5831 | 1921.1691 | 1899.5393 | 2130.0757 | 2668.2567 | 2262.2827 | 2155.6209 | 1997.2567 | 2668.2567 | 2230.3235 | 2190.4695 | 2314.0816 | 2419.9415 | 2369.7025 | 2574.6091 | 2123.2945 | 2119.2945 | 1995.411 | 2205.1135 | 1886.5949 | 1979.9184 | 1988.3672 | 2020.5334 | 1851.3702 | 1871.7522 | 1916.8776 | 1871.7025 | 1952.5801 | 1977.6647 | 2025.5044 | 1972.8309 | 1923.2478 | 2151.8368 | 2418.003 | 2582.3205 | 3230.1254 | 4055.5275 | 4345.2714 | 3707.5452 | 1836.4577 | 1863.7522 | 1875.2129 | 1836.4577 | 1940.207 | 2070.8338 | 1884.9533 | 1955.1632 | 1889.408 | 2040.5015 | 1988.4228 | 1862.7871 | 1772.5305 | 1854.9651 | 1824.3672 | 1880.7493 | 1849.793 | 1839.2886 | 1917.4577 | 1808.4488 | 2011.5364 | 1544.4341 | 1642.9125 | 1732.3702 | 1910.3471 | 1692.9184 | 1669.3761 | 1587.6239 | 2057.4169 | 2770.7664 | 2787.9036 | 2175.4459
141 | 3.5454 | 889.3667 | 2223.7551 | 2224.0561 | 2233.6942 | 2326.6568 | 3036.3667 | 2442.6381 | 1881.7925 | 1768.9174 | 2073.4307 | 1887.4179 | 1628.8004 | 2390.6863 | 2299.415 | 2715.3323 | 2237.1013 | 2372.4228 | 2191.532 | 2168.8987 | 1837.1809 | 1915.7099 | 2968.8082 | 3089.3402 | 1967.7021 | 2555.1327 | 2672.7925 | 2450.0826 | 2637.3776 | 1940.8239 | 2170.3323 | 2078.7178 | 2060.471 | 1738.295 | 2301.0718 | 2423.234 | 2110.585 | 3043.3559 | 3435.3619 | 3010.7021 | 2944.9469 | 2101.1436 | 2716.9548 | 2464.9312 | 2064.7365 | 1806.8191 | 1903.1514 | 1896.0314 | 2091.8378 | 1977.0157 | 2424.1809 | 2152.3166 | 1896.351 | 2075.585 | 2711.3963 | 2141.7974 | 2021.1888 | 2133.6834 | 2775.4946 | 2014.6381 | 2018.0609 | 1894.7738 | 2279.1357 | 1875.1062 | 1833.4572 | 2157.4071 | 2680.3963 | 2232.351 | 2120.7738 | 1966.1514 | 2642.9174 | 2217.8112 | 2172.3137 | 2221.3481 | 2417.4789 | 2345.1888 | 2495.1966 | 2020.4464 | 2119.5133 | 1981.2793 | 2174.3963 | 1892.3776 | 2025.0826 | 1985.8564 | 1944.0609 | 1918.5615 | 1850.3323 | 1954.7129 | 1844.3058 | 1916.1701 | 2001.8299 | 1968.9705 | 1968.4415 | 1881.2419 | 2070.6273 | 2393.2871 | 2575.7847 | 3193.5211 | 4151.0138 | 4315.2763 | 3785.9125 | 1844.3697 | 1875.4867 | 1852.1249 | 1844.3697 | 1909.6116 | 2087.8564 | 1923.4523 | 1944.8378 | 1927.3806 | 1954.6273 | 1937.412 | 1812.3589 | 1768.3432 | 1850.8191 | 1836.2714 | 1860.9813 | 1873.9017 | 1789.4759 | 1849.6912 | 1780.9548 | 1995.3963 | 1601.5989 | 1707.237 | 1743.3432 | 1860.114 | 1654.4759 | 1728.2183 | 1596.532 | 2086.1092 | 2788.12 | 2842.0482 | 2235.3727
142 | 3.5564 | 866.3982 | 2256.1136 | 2239.8687 | 2271.9867 | 2317.267 | 2990.674 | 2365.4248 | 1882.6652 | 1701.6873 | 2067.4617 | 1794.2522 | 1696.1003 | 2461.5118 | 2329.1903 | 2667.615 | 2143.4469 | 2330.8141 | 2149.8186 | 2124.7419 | 1816.4115 | 1921.9233 | 3007.2404 | 3082.3938 | 1963.1991 | 2560.0546 | 2641.4115 | 2398.5604 | 2705.236 | 1941.3127 | 2173.1947 | 2093.4484 | 2066.4749 | 1747.7198 | 2230.5604 | 2403.4071 | 2102.8009 | 3076.9012 | 3426.0767 | 2987.1401 | 2913.9498 | 2033.4248 | 2761.1814 | 2399.3067 | 2095.2449 | 1771.1445 | 1920.9189 | 1897.8466 | 2098.3304 | 1943.4705 | 2429.8009 | 2113.9543 | 1858.5472 | 2099.3215 | 2694.6062 | 2162.7065 | 1998.674 | 2147.1224 | 2754.2714 | 1972.6106 | 2041.4484 | 1924.5841 | 2268.1401 | 1909.9779 | 1779.9454 | 2204.3215 | 2597.6327 | 2265.6519 | 2108.0678 | 1980.2537 | 2596.2124 | 2237.3938 | 2163.7965 | 2287.4263 | 2384.3525 | 2309.4793 | 2594.0871 | 2045.4396 | 2086.2124 | 2088.7567 | 2261.6977 | 1911.8555 | 1954.5014 | 1990.5339 | 1923.4661 | 1852.9719 | 1811.4115 | 1914.9543 | 1865.3894 | 1883.7419 | 1901.5781 | 1899.7507 | 1941.1445 | 1930.7744 | 2045.5295 | 2355.3436 | 2573.2714 | 3209.7876 | 4074.1843 | 4306.7242 | 3839.3894 | 1775.4292 | 1864.0723 | 1841.9322 | 1775.4292 | 1915.8599 | 2011.8997 | 1964.2581 | 1958.6608 | 1895.7551 | 1968.3805 | 1947.1859 | 1843.1814 | 1799.7242 | 1796.0855 | 1782.0265 | 1888.4528 | 1851.2124 | 1797.9233 | 1829.8643 | 1810.5207 | 2057.0413 | 1601.2846 | 1672.9587 | 1743.2035 | 1846.326 | 1675.385 | 1690.6239 | 1603.9322 | 2075.0088 | 2751.2935 | 2767.9764 | 2215.0221
143 | 3.5674 | 921.6637 | 2279.2173 | 2166.4484 | 2198.2871 | 2319.9813 | 3027.2291 | 2287.1003 | 1902.6332 | 1718.7758 | 2083.1799 | 1758.1799 | 1677.5536 | 2428.8269 | 2285.8269 | 2771.0168 | 2153.1731 | 2265.8269 | 2130.764 | 2102.0934 | 1791.8387 | 1916.0934 | 3000.4415 | 3079.056 | 1972.1986 | 2521.8082 | 2606.4041 | 2268.9085 | 2611.5113 | 1910.8082 | 2192.1613 | 2110 | 1980.7404 | 1748.0747 | 2246.9931 | 2434.2665 | 2092.3107 | 3024.7591 | 3381.2242 | 3000.3786 | 2909.4346 | 2098.2036 | 2732.5349 | 2384.6146 | 2111.2734 | 1801.9931 | 1911.6401 | 1890.6588 | 2048.0816 | 1980.9931 | 2438.6893 | 2092.764 | 1870.5959 | 2048.883 | 2739.6824 | 2160.1681 | 2001.6893 | 2132.4041 | 2730.6214 | 1922.7522 | 1963.3431 | 1883.8456 | 2201.6283 | 1868.1376 | 1830.117 | 2102.4552 | 2592.5959 | 2126.2074 | 2047.3736 | 1928.647 | 2556.6775 | 2179.647 | 2158.056 | 2283.7522 | 2410.9744 | 2366.5703 | 2601.7148 | 2041.6775 | 2083.944 | 2069.119 | 2179.2192 | 1866.8456 | 1982.6264 | 2031.7385 | 1989.589 | 1797.8761 | 1836.0305 | 1886.3107 | 1816.119 | 1925.9371 | 1915.4091 | 1960.0236 | 1947.3973 | 1898.1003 | 2036.3294 | 2378.0492 | 2624.8997 | 3177.0256 | 4072.7198 | 4241.2124 | 3830.3668 | 1793.7198 | 1835.1495 | 1803.4975 | 1793.7198 | 1882.4975 | 1981.2547 | 1907.1563 | 1896.175 | 1808.9203 | 1921.9016 | 1910.8082 | 1873.4159 | 1790.8574 | 1786.4533 | 1766.1986 | 1866.7148 | 1734.9085 | 1789.1121 | 1816.3481 | 1738.5231 | 1978.6352 | 1630.4651 | 1637.3294 | 1764.5959 | 1802.3176 | 1633.8456 | 1596.703 | 1590.8387 | 2042.4041 | 2683.0256 | 2770.7385 | 2202.4533
144 | 3.5784 | 878.6327 | 2292.3643 | 2173.0624 | 2172.7881 | 2282.3613 | 2998.058 | 2332.3978 | 1905.9695 | 1720.3933 | 2076.3933 | 1770.3338 | 1596.6022 | 2377.6647 | 2298.3048 | 2716.2974 | 2239.9435 | 2270.7286 | 2076.2089 | 2125.7896 | 1802.91 | 1861.7547 | 2980.9086 | 3053.2714 | 1986.3643 | 2539.8506 | 2603.3033 | 2292.852 | 2616.2149 | 1878.09 | 2154.5732 | 2085.7561 | 1909.3004 | 1731.8781 | 2228.4223 | 2367.4498 | 1946.5673 | 2995.4848 | 3434.913 | 2972.9071 | 2811.4788 | 2140.8796 | 2704.1204 | 2357.1494 | 2066.8156 | 1744.026 | 1960.8216 | 1885.0305 | 2027.7271 | 1979.09 | 2414.7257 | 2080.6357 | 1873.6967 | 2080.2758 | 2722.9071 | 2145.6967 | 1996.6662 | 2116.4238 | 2767.1844 | 1923.3948 | 2005.2468 | 1843.029 | 2299.9465 | 1839.4238 | 1837.3628 | 2097.3353 | 2516.9041 | 2192.1888 | 2062.6082 | 1919.1219 | 2573.5167 | 2172.3948 | 2097.4208 | 2252.5747 | 2422.9695 | 2414.2439 | 2538.4818 | 1969.8751 | 2113.5167 | 1977.026 | 2181.6691 | 1810.8766 | 1972.6952 | 1989.9361 | 1955.1784 | 1804.6372 | 1786.6327 | 1836.0275 | 1757.3613 | 1947.5457 | 1977.1844 | 1921.6022 | 1881.7837 | 1807.026 | 2090.3368 | 2299.6008 | 2602.3309 | 3214.2758 | 4082.6967 | 4199.3323 | 3773.2394 | 1751.4208 | 1770.1784 | 1807.4253 | 1751.4208 | 1823.5428 | 2026.4572 | 1881.9695 | 1848.7561 | 1788.2729 | 1937.1234 | 1951.5792 | 1888.6067 | 1797.8796 | 1806.4253 | 1780.3643 | 1820.9071 | 1770.2163 | 1754.2104 | 1732.2074 | 1778.9435 | 1972.032 | 1608.8766 | 1560.2684 | 1762.3933 | 1743.3309 | 1641.5167 | 1625.0334 | 1602.6677 | 2052.9405 | 2771.0364 | 2843.3368 | 2182.2714
145 | 3.5894 | 892.1062 | 2296.7699 | 2116.7876 | 2207.8761 | 2281.5487 | 3090.2478 | 2324.5133 | 1909.6195 | 1652.3274 | 2140.7522 | 1679.8407 | 1549.823 | 2374.6991 | 2285.6726 | 2722.9558 | 2194.2478 | 2283.3894 | 2125.2124 | 2149.6283 | 1786.4425 | 1820.5929 | 2854.8141 | 3013.6637 | 1990.7699 | 2419.3451 | 2613.469 | 2282.7434 | 2618.6814 | 1942.2921 | 2154.8584 | 2072.6106 | 1924.2566 | 1750.3274 | 2298.8407 | 2319.6726 | 1924.0974 | 2983.4513 | 3452.1504 | 3068.2478 | 2826.7345 | 2118.823 | 2656.115 | 2336.7611 | 2026.0531 | 1698.6726 | 1888.1504 | 1929.885 | 2060.1062 | 1968.6106 | 2405.2301 | 2033.8849 | 1895.8673 | 2112.3982 | 2644.2478 | 2099.3451 | 1941.646 | 2140.177 | 2754.0531 | 1925.6903 | 1950.469 | 1856.7876 | 2199.4159 | 1871.646 | 1810.6726 | 2121.0885 | 2605.4868 | 2150.6991 | 2040.1239 | 1926.0796 | 2499.6194 | 2159.292 | 2081.3097 | 2194.646 | 2345.7787 | 2395.823 | 2564.3451 | 1911.5044 | 1997.2743 | 1975.5575 | 2184.9911 | 1884.4602 | 1921.4248 | 1944.2832 | 1989.354 | 1888.3009 | 1741.5929 | 1831.0089 | 1768.0974 | 1872.0088 | 1889.3716 | 1922.5575 | 1925.9823 | 1871.7699 | 2048.177 | 2341.062 | 2545.115 | 3192.5044 | 4119.4956 | 4133.9469 | 3800.1151 | 1788.4336 | 1813.5929 | 1771.4248 | 1788.4336 | 1823.6283 | 2021.7522 | 1907.177 | 1827.1416 | 1842.8938 | 1903.5752 | 1911.4867 | 1768.0442 | 1719.6194 | 1835.8673 | 1755.5133 | 1766.115 | 1675.2035 | 1748.8496 | 1719.3186 | 1724.7787 | 2026.4248 | 1622.4071 | 1593.2036 | 1683.5486 | 1793.1328 | 1655.3894 | 1638.9204 | 1551.7964 | 2078.2478 | 2795.531 | 2782.6991 | 2073.3716
146 | 3.6004 | 862.8899 | 2339.239 | 2189.2999 | 2263.3678 | 2220.8594 | 3042.9975 | 2359.0212 | 1907.089 | 1694.794 | 2115.6322 | 1744.2999 | 1574.8525 | 2420.0212 | 2296.9907 | 2658.1263 | 2157.5737 | 2274.0492 | 2096.5034 | 2157.911 | 1719.7704 | 1884.9627 | 2913.4287 | 3007.911 | 1982.3958 | 2410.5059 | 2552.6416 | 2264.6136 | 2666.6254 | 2008.3678 | 2156.6534 | 2092.1195 | 1998.705 | 1719.4543 | 2285.4543 | 2386.5271 | 1934.1288 | 3048.8338 | 3379.0373 | 2940.4306 | 2780.7704 | 1957.6882 | 2750.1618 | 2323.8314 | 2017.2178 | 1735.0212 | 1850.8314 | 1928.9203 | 2025.4144 | 1938.0585 | 2399.5246 | 2035.0797 | 1840.4238 | 2077.1077 | 2695.1805 | 2108.6839 | 2016.7355 | 2123.7517 | 2720.9695 | 1878.85 | 1949.6839 | 1888.8127 | 2181.4661 | 1781.0958 | 1765.6721 | 2078.85 | 2650.1686 | 2147.3373 | 2001.8407 | 1934.0398 | 2525.8432 | 2140.7026 | 2057.7517 | 2189.1686 | 2348.0305 | 2333.2458 | 2494.8687 | 2004.033 | 2078.1126 | 1967.9203 | 2129.5924 | 1894.3559 | 1949.6745 | 1930.5246 | 1965.1077 | 1866.6322 | 1758.2483 | 1847.2974 | 1764.0492 | 1858.3864 | 1863.6932 | 1884.4144 | 1866.2153 | 1862.0187 | 2016.7026 | 2370.6042 | 2493.6721 | 3139.5433 | 3990.9857 | 4161.3678 | 3830.2085 | 1758.85 | 1745.1755 | 1757.4356 | 1758.85 | 1792.0187 | 1969.0678 | 1844.6017 | 1852.7237 | 1800.3348 | 1911.1195 | 1850.1966 | 1798.2694 | 1710.6441 | 1783.5526 | 1720.4051 | 1762.3864 | 1739.3398 | 1745.3068 | 1668.2458 | 1752.061 | 1935.1848 | 1553.3441 | 1593.5737 | 1628.9297 | 1768.9788 | 1621.8009 | 1609.8407 | 1507.6229 | 2129.3678 | 2839.9322 | 2845.8245 | 2094.6254
147 | 3.6114 | 902.8476 | 2180.9783 | 2122.8387 | 2200.5231 | 2221.2734 | 3025.767 | 2316.1799 | 1880.5762 | 1619.8387 | 2015.6057 | 1793.0403 | 1593.5339 | 2460.6018 | 2254.3304 | 2668.7798 | 2089.4405 | 2259.8771 | 2008.3176 | 2070.0167 | 1750.6696 | 1848.5909 | 2914.5762 | 2992.3156 | 1878.3048 | 2441.2458 | 2551.6844 | 2294.5742 | 2671.4661 | 1951.0845 | 2071.7836 | 2017.7561 | 1969.3304 | 1706.6165 | 2209.057 | 2351.3854 | 1980.2862 | 2948.8259 | 3263.6607 | 2895.2055 | 2812.4641 | 2011.1338 | 2677.5103 | 2373.4366 | 2054.3412 | 1671.5231 | 1939.3668 | 1849.5231 | 1997.9046 | 1910.9321 | 2384.4936 | 2062.0403 | 1819.822 | 2001.646 | 2620.4552 | 2087.1927 | 1970.4956 | 2070.1249 | 2771.9981 | 1915.2586 | 1941.6991 | 1855.4808 | 2165.7542 | 1744.2606 | 1826.5044 | 2131.943 | 2558.4552 | 2087.057 | 2056.765 | 1945.5339 | 2524.1101 | 2113.1652 | 2073.2458 | 2269.0659 | 2324.8092 | 2309.2606 | 2478.356 | 1999.6313 | 2091.9321 | 1967.178 | 2227.8181 | 1802.6057 | 1926.8368 | 1922.644 | 2068.4071 | 1830.9597 | 1811.0806 | 1870.8623 | 1741.1377 | 1853.8496 | 1827.6588 | 1916.4917 | 1851.9449 | 1910.7925 | 2028.0954 | 2350.7542 | 2520.3412 | 3130.3835 | 4042.5723 | 4134.2753 | 3857.2458 | 1762.6844 | 1778.9981 | 1755.411 | 1762.6844 | 1816.1632 | 1973.4789 | 1841.6844 | 1863.3835 | 1797.5064 | 1884.4258 | 1815.7542 | 1840.0403 | 1694.3432 | 1750.3156 | 1738.4238 | 1764.2055 | 1773.356 | 1730.4936 | 1760.2567 | 1673.8387 | 1944.9302 | 1556.6293 | 1561.9872 | 1659.7247 | 1786.3963 | 1563.5359 | 1636.1632 | 1482.5211 | 2076.8515 | 2761.6332 | 2819.1524 | 2235.3245
148 | 3.6224 | 872.4336 | 2150.7153 | 2111.9086 | 2206.0103 | 2258.2714 | 2902.7551 | 2293.1254 | 1895.7832 | 1582.2847 | 2041.4071 | 1701.379 | 1582.2847 | 2375.2197 | 2200.854 | 2709.2714 | 2161.7463 | 2183.5826 | 2105.177 | 2096.2139 | 1758.9661 | 1827.5457 | 2908.8068 | 2912.1622 | 1914.1224 | 2449.4204 | 2607.6342 | 2245.7183 | 2622.2065 | 1923.3186 | 2075.0103 | 2025.2375 | 1914.2404 | 1697.1254 | 2165.3761 | 2328.1254 | 1998.4543 | 2967.2139 | 3406.0045 | 2924.851 | 2810.4779 | 2071.4985 | 2601.7861 | 2367.6372 | 2045.4439 | 1703.6342 | 1872.6401 | 1834.0339 | 1998.0339 | 1882.6032 | 2443.1224 | 2005.0708 | 1835.9425 | 2014.851 | 2673.4514 | 2128.4071 | 2001.2139 | 2057.1932 | 2725.6843 | 1895.7286 | 2012.2242 | 1853.0678 | 2090.6165 | 1731.5118 | 1772.3894 | 2148.1932 | 2598.5428 | 2152.2921 | 2092.3628 | 1886.6843 | 2584.5428 | 2081.9897 | 2068.5118 | 2276.0914 | 2314.5796 | 2272.3082 | 2525.18 | 2022.4646 | 2132.9867 | 1897.8097 | 2210.2065 | 1774.3525 | 1906.705 | 1943.59 | 1977.789 | 1829.0339 | 1773.7861 | 1847.7625 | 1756.1696 | 1780.1961 | 1826.0678 | 1818.2772 | 1866.9086 | 1809.823 | 2049.7493 | 2379.6918 | 2536.4543 | 3178.6003 | 4082.1357 | 4174.8275 | 3882.3628 | 1757.3186 | 1725.4233 | 1730.7965 | 1757.3186 | 1749.1622 | 1940.8879 | 1855.6814 | 1801.8776 | 1821.9764 | 1860.6711 | 1788.2168 | 1794.3658 | 1741.441 | 1723.41 | 1732.0914 | 1774.4882 | 1681.4129 | 1734.2611 | 1790.4204 | 1771.8245 | 1981.6239 | 1550.8982 | 1577.59 | 1667.5796 | 1812.5561 | 1601.0207 | 1681.3053 | 1523.9867 | 2074.5221 | 2786.4071 | 2768.0811 | 2168.7655
149 | 3.6334 | 811.2173 | 2209.6972 | 2169.8781 | 2200.5329 | 2234.056 | 2929.5162 | 2217.2832 | 1856.7699 | 1668.2793 | 2048.0855 | 1673.0855 | 1597.0197 | 2360.3717 | 2222.1249 | 2680.1416 | 2158.4277 | 2237.2596 | 2083.0364 | 2118.648 | 1756.3618 | 1784.2271 | 2901.6283 | 2866.1514 | 1934.7335 | 2377.1022 | 2635.5624 | 2265.5821 | 2732.2232 | 1935.7434 | 2045.7797 | 2031.1809 | 1940.0393 | 1701.5526 | 2199.1347 | 2330.9145 | 1961.5034 | 3050.2596 | 3389.4838 | 2929.5428 | 2826.6578 | 2061.9803 | 2718.5949 | 2331.8751 | 1959.9115 | 1688.1514 | 1855.9145 | 1835 | 1986.7139 | 1828.6745 | 2427.6942 | 1986.7237 | 1834.2665 | 2023.6382 | 2595.349 | 2070.3785 | 1958.1219 | 2053.7237 | 2748.6775 | 1892.0954 | 1921.0629 | 1853.3618 | 2097.7532 | 1732.2763 | 1828.7168 | 2135.7139 | 2549.293 | 2075.349 | 2023.2832 | 1931.7925 | 2496.7964 | 2088.2861 | 2007.8456 | 2200.9213 | 2341.0295 | 2322.1544 | 2558.1151 | 1958.7404 | 2017.0432 | 1919.4966 | 2130.1976 | 1831.9735 | 1876.1514 | 1892.4838 | 1949.8191 | 1813.6185 | 1784.1151 | 1819.704 | 1662.7109 | 1817.059 | 1818.9902 | 1778.0757 | 1834.7797 | 1811.4867 | 1995.8456 | 2282.3392 | 2604.3452 | 3081.0629 | 3986.8918 | 4128.0265 | 3823.1121 | 1766.4671 | 1687.704 | 1720.0855 | 1766.4671 | 1813.7267 | 1878.7601 | 1817.8653 | 1792.7335 | 1743.4542 | 1837.3422 | 1829.6873 | 1788.5526 | 1713.0462 | 1726.1908 | 1651.1878 | 1745.056 | 1683.5723 | 1674.5693 | 1679.6057 | 1755.9508 | 1954.4179 | 1517.6942 | 1556.7896 | 1657.1711 | 1739.1878 | 1575.3225 | 1669.5231 | 1480.1219 | 2028.0364 | 2744.7601 | 2804.8683 | 2199.8781
150 | 3.6444 | 849.882 | 2243.6155 | 2140.4316 | 2202.9882 | 2325.1013 | 3004.3525 | 2234.2547 | 1832.0354 | 1655.5767 | 2022.059 | 1667.0118 | 1553.6121 | 2362.8314 | 2197.4081 | 2772.5954 | 2042.4282 | 2322.8348 | 2086.1332 | 2061.9729 | 1726.0708 | 1813.2429 | 2845.9611 | 2909.882 | 1988.3761 | 2449.1249 | 2529.2596 | 2274.6627 | 2624.3185 | 1822.6086 | 2067.9292 | 2029.3373 | 1982.086 | 1656.4434 | 2133.177 | 2315.3727 | 1906.4434 | 2995.3338 | 3296.3574 | 2906.5531 | 2702.1381 | 1984.0083 | 2550.7974 | 2361.5801 | 2015.9917 | 1680.3373 | 1773.8633 | 1789.2748 | 1950.7453 | 1824.6509 | 2386.0826 | 1978.5177 | 1764.8279 | 1969.1298 | 2590.6273 | 1985.9965 | 1837.9223 | 2048.0118 | 2731.2276 | 1792.4046 | 1845.4552 | 1836.3727 | 2097.1804 | 1749.6273 | 1766.8515 | 2048.2006 | 2551.6273 | 2093.0743 | 2051.0625 | 1900.1062 | 2531.3643 | 2044.4434 | 2102.2463 | 2129.2984 | 2333.6981 | 2290.9375 | 2547.204 | 1964.6155 | 1993.9882 | 1952.1096 | 2122.4823 | 1769.3456 | 1875.8196 | 1852.5649 | 1931.7099 | 1714.3928 | 1685.9105 | 1771.7689 | 1701.6898 | 1806.5531 | 1849.4233 | 1859.7994 | 1844.4587 | 1807.1922 | 2021.2429 | 2292.2547 | 2594.059 | 3108.5448 | 3896.8279 | 4112.5059 | 3922.0659 | 1709.1416 | 1668.0354 | 1689.9022 | 1709.1416 | 1793.7571 | 1924.3643 | 1809.8314 | 1794 | 1632.7183 | 1792.263 | 1762.6829 | 1689.4046 | 1631.6711 | 1717.1922 | 1747.8973 | 1656.526 | 1672.3727 | 1650.1922 | 1645.0236 | 1729.3727 | 1968.6155 | 1546.7488 | 1580.0978 | 1609.2748 | 1782.6898 | 1578.6391 | 1594.3338 | 1468.8314 | 2001.204 | 2836.9209 | 2819.1568 | 2128.0319
151 | 3.6554 | 902.1829 | 2191.4749 | 2105.8466 | 2216.6549 | 2283.413 | 2960.4749 | 2166.7581 | 1869.4956 | 1618.6283 | 2060.0266 | 1748.6785 | 1502.6608 | 2400.9646 | 2133.6312 | 2752.5044 | 2089.18 | 2203.702 | 2042.5988 | 2054.2183 | 1694.6903 | 1824.3746 | 2833.3127 | 2893.0029 | 1872.4513 | 2456.8761 | 2464.8171 | 2275.8437 | 2580.1268 | 1864.8053 | 2043.6283 | 2029.8437 | 1926.5693 | 1679.9027 | 2213.3334 | 2348.2773 | 1919.7788 | 3023.1829 | 3365.2095 | 2910.4366 | 2852.3245 | 1974.9056 | 2567.5251 | 2349.5959 | 2008.2212 | 1602.3835 | 1826.18 | 1847.4926 | 1917.0029 | 1812.5015 | 2457.4602 | 1962.6578 | 1742.5634 | 1984.1534 | 2673.3039 | 1999.8407 | 1900.3658 | 2088.3392 | 2745.4041 | 1823.4631 | 1898.2744 | 1862.528 | 2094.3127 | 1761.4366 | 1829.4602 | 2063.059 | 2484.9764 | 2059.6608 | 1987.4454 | 1835.3186 | 2439.2301 | 2027.4071 | 2060.7256 | 2096.9085 | 2328.469 | 2360.4602 | 2448.6666 | 1906.5693 | 1899.6666 | 1886.9764 | 2190.649 | 1786.9027 | 1911.59 | 1852.0619 | 1962.59 | 1688.2507 | 1767.9263 | 1739.3776 | 1708.0944 | 1800.3127 | 1869.4985 | 1820.4749 | 1829.4395 | 1778.0029 | 1942.9469 | 2289.7198 | 2517.2891 | 3119 | 4011.236 | 4082.472 | 3938.2507 | 1680.1268 | 1721.5251 | 1656.1593 | 1680.1268 | 1803.7168 | 1979.0266 | 1742.5074 | 1840.4012 | 1775.1387 | 1735.9115 | 1746.1563 | 1767.708 | 1690.3983 | 1705.7198 | 1695.0708 | 1595.2861 | 1702.7463 | 1679.2153 | 1554.4159 | 1716.7198 | 1932.5044 | 1539.6578 | 1545.5044 | 1674.2419 | 1749.4425 | 1559.0649 | 1630.6195 | 1479.1239 | 1997.1563 | 2824.2537 | 2812.4071 | 2119.4366
152 | 3.6664 | 898.0167 | 2210.1023 | 2147.2213 | 2194.1775 | 2234.3382 | 2983.382 | 2216.3402 | 1841.0501 | 1616.4238 | 2068.8559 | 1672.4134 | 1516.3987 | 2360.0668 | 2048.5324 | 2635.4385 | 2075.4656 | 2246.62 | 2018.5929 | 2034.4572 | 1730.2296 | 1795.0418 | 2844.9833 | 2864.0334 | 1854.9916 | 2412.2025 | 2510.0688 | 2204.9562 | 2584.7035 | 1852.1775 | 2022.5929 | 2029.1441 | 1896.1691 | 1680.0084 | 2147.9729 | 2322.4739 | 1888.0501 | 3063.5261 | 3402.3987 | 2898.3048 | 2814.8122 | 2032.4926 | 2618.2296 | 2334.0167 | 2018.5511 | 1670.1711 | 1786.7871 | 1828.3299 | 1930.8309 | 1861.3569 | 2401.8122 | 1924.9061 | 1798.3486 | 1947.0585 | 2668.3132 | 2017.8392 | 1869.7787 | 2076.3132 | 2675.3884 | 1829.5762 | 1900.5846 | 1778.9813 | 2126.5261 | 1709.9312 | 1796.9228 | 2039.0418 | 2468.5762 | 2035.737 | 1950.8894 | 1778.7787 | 2442.2547 | 2101.3235 | 2066.9749 | 2117.8225 | 2280.9228 | 2231.2005 | 2424.1524 | 1915.5428 | 1978.7223 | 1913.3737 | 2120.8289 | 1752.7704 | 1857.2276 | 1812.7704 | 1928.7704 | 1698.1273 | 1717.6681 | 1771.8058 | 1667.6263 | 1711.8372 | 1798.3884 | 1789.4572 | 1873.0772 | 1803.8142 | 1982.6367 | 2318.8142 | 2540.8058 | 3107.0167 | 4013.6013 | 4106.1023 | 3853.117 | 1671.0084 | 1676.3633 | 1662.9833 | 1671.0084 | 1763.3466 | 1883.4302 | 1790.7724 | 1752.2777 | 1734.8122 | 1779.2129 | 1758.8392 | 1709.8205 | 1701.4322 | 1659.4906 | 1701.5428 | 1695.7055 | 1631.9646 | 1725.2296 | 1619.3152 | 1701.3048 | 1870.2193 | 1509.6096 | 1611.6116 | 1673.4489 | 1757.263 | 1516.4823 | 1615.1775 | 1467.3048 | 1981.5929 | 2792.3215 | 2706.5742 | 2080.7704
153 | 3.6774 | 919.4425 | 2150.7502 | 2082.9361 | 2202.1554 | 2260.1249 | 3042.8309 | 2214.8446 | 1855.8378 | 1623.9459 | 2039.4798 | 1635.294 | 1611.1652 | 2293.8279 | 2121.9558 | 2676.9086 | 2082.2871 | 2144.2605 | 2016.8682 | 2078.7532 | 1706.6116 | 1860.3275 | 2866.9695 | 2914.6991 | 1852.527 | 2349.4326 | 2529.4189 | 2254.6451 | 2574.7129 | 1886.2026 | 1984.4022 | 1993.4022 | 1888.6588 | 1706.7837 | 2171.3343 | 2272.7197 | 1875.0541 | 3036.0069 | 3260.1356 | 2836.9597 | 2786.5034 | 1945.6253 | 2549.4631 | 2309.4258 | 2017.527 | 1614.5408 | 1784.0777 | 1883.7768 | 1952.233 | 1931.6451 | 2380.4494 | 1932.8918 | 1793.2399 | 1975.5978 | 2599.0137 | 1980.9292 | 1967.0571 | 2025.2468 | 2712.4897 | 1849.8378 | 1895.527 | 1750.9764 | 2097.2704 | 1717.4966 | 1684.8112 | 2049.6283 | 2447.1622 | 2112.0099 | 2017.3275 | 1806.5978 | 2415.6116 | 2113.1318 | 2047.4258 | 2125.473 | 2273.8682 | 2196.5811 | 2442.1013 | 1935.4966 | 1969.821 | 1900.0305 | 2141.7296 | 1769.4425 | 1823.5575 | 1808.6047 | 1840.3854 | 1643.5644 | 1666.061 | 1749.3481 | 1710.4897 | 1767.3275 | 1753.8751 | 1732.1691 | 1851.4798 | 1779.4798 | 1976.3717 | 2313.5811 | 2572.179 | 3058.9833 | 3881.5005 | 4109.2635 | 3872.8614 | 1614.4326 | 1660.9223 | 1657.1858 | 1614.4326 | 1782.9154 | 1848.6352 | 1743.6961 | 1757.706 | 1678.7197 | 1729.8279 | 1809.6215 | 1764.7227 | 1617.6489 | 1658.0777 | 1730.179 | 1707.9223 | 1612.9223 | 1617.3923 | 1672.0708 | 1627.8043 | 1961.7159 | 1502.3953 | 1558.882 | 1595.4631 | 1696.6185 | 1575.118 | 1579.7434 | 1437.2163 | 1979 | 2752.4022 | 2745.4356 | 2098.3107
154 | 3.6884 | 908.031 | 2112.823 | 2080.8053 | 2111.5442 | 2287.8938 | 2995.5265 | 2207.7168 | 1827.3009 | 1529.0221 | 1959.2743 | 1727.9779 | 1547.5132 | 2359.8452 | 2179.9071 | 2636.6902 | 2097.208 | 2166.6239 | 1982.6593 | 2027.8849 | 1620.9911 | 1802.3186 | 2797.7964 | 2911.5531 | 1878.4159 | 2401.7567 | 2465.8407 | 2209.4071 | 2572.1195 | 1766.9646 | 2013.3407 | 1956.3009 | 1842.2124 | 1671.3761 | 2118.0044 | 2345.2478 | 1884.5664 | 2948.2301 | 3362.1549 | 2798.0619 | 2760.2257 | 1962.6991 | 2632.2168 | 2246.1239 | 1982.6593 | 1530.9159 | 1664.3672 | 1800.5442 | 1895.1239 | 1833.5752 | 2423.9071 | 2011.6062 | 1811.2522 | 1933.4513 | 2554.6593 | 1954.7478 | 1908.1548 | 2030.3274 | 2702.3894 | 1871.4912 | 1886.1947 | 1725.2257 | 2065.5841 | 1761.385 | 1676.9248 | 2029.9867 | 2510.9204 | 2007.7699 | 2005.1062 | 1797.031 | 2522.8275 | 2031.9911 | 2029.9115 | 2092.5398 | 2269.5973 | 2253.1151 | 2442.3584 | 1898.3761 | 2007.385 | 1890.8363 | 2139.0752 | 1721.0487 | 1836.7301 | 1810.7611 | 1846.4912 | 1718.6505 | 1702.2655 | 1743.2389 | 1679.2566 | 1699.6327 | 1810.6372 | 1808.531 | 1854.0443 | 1738.0177 | 1931.7345 | 2242.438 | 2539.8982 | 3103.9513 | 3898.9823 | 4078.1814 | 3892.0885 | 1624.4912 | 1601.8849 | 1621.6593 | 1624.4912 | 1720.9602 | 1789.646 | 1793.1151 | 1740.7478 | 1690.4912 | 1693.7035 | 1852.2213 | 1720.4071 | 1629.6549 | 1644.031 | 1607.2035 | 1677.9425 | 1594.3894 | 1618.4912 | 1664.792 | 1712.9779 | 1897.7522 | 1527.5354 | 1543.7168 | 1566.3894 | 1723.9823 | 1546.2566 | 1577.6416 | 1448.8496 | 1987.8053 | 2777.8186 | 2714.3761 | 2094.0752
155 | 3.6994 | 883.6794 | 2110.7856 | 2073.0718 | 2095.4287 | 2169.5762 | 2863.0049 | 2214.1426 | 1816.7502 | 1520.4641 | 1932.9292 | 1641.8967 | 1528.0718 | 2303.9311 | 2107.4316 | 2561.2527 | 2115.0708 | 2172 | 2012.8201 | 2027.1426 | 1661.7837 | 1793.8928 | 2846.9263 | 2910 | 1884.6785 | 2444.8201 | 2543.6038 | 2260.9263 | 2519.3235 | 1879.4592 | 2042.1062 | 1921.644 | 1913.6755 | 1616.002 | 2081.4297 | 2236.3618 | 1866.3579 | 3041.4602 | 3239.7197 | 2927.9233 | 2814.1406 | 1930.3943 | 2469.114 | 2279.0698 | 1963.7148 | 1586.176 | 1802.4583 | 1764.7866 | 1920.9272 | 1884.176 | 2342.5034 | 1974.6804 | 1838.1062 | 1951.1062 | 2576.7847 | 2005.5693 | 1873.001 | 2002.5369 | 2731.2488 | 1793.646 | 1878.7502 | 1739.8564 | 2115.5693 | 1664.3254 | 1639.6086 | 2014.5005 | 2436.0747 | 2055.3904 | 1984.1436 | 1811.0708 | 2471.9311 | 2063.8555 | 2039.6067 | 2149.6047 | 2241.3225 | 2232.4297 | 2492.8909 | 1905.4995 | 1946.9311 | 1831.0737 | 2170.1416 | 1615.7895 | 1835.4287 | 1731.5388 | 1847.2144 | 1656.7886 | 1671.3943 | 1683.1809 | 1659.7148 | 1717.5703 | 1813.6431 | 1802.1436 | 1837.1436 | 1750.8564 | 1969.7483 | 2245.4995 | 2488.5732 | 3033.3245 | 3991.4966 | 4096.3206 | 3885.0718 | 1620.7502 | 1658.2832 | 1621.4641 | 1620.7502 | 1696.822 | 1856.9971 | 1683.1475 | 1724.6077 | 1544.7915 | 1801.6028 | 1786.467 | 1679.9656 | 1553.5388 | 1701.8191 | 1582.8574 | 1579.3608 | 1593.7856 | 1621.8928 | 1614.7876 | 1623.3254 | 1877.179 | 1458.6814 | 1526.0364 | 1590.7847 | 1675.4307 | 1516.8938 | 1536.0374 | 1389.3953 | 1991.6785 | 2679.3254 | 2739.356 | 2122.4631
156 | 3.7104 | 866.5182 | 2186.1244 | 2106.3943 | 2133.0148 | 2111.5546 | 2914.0733 | 2215 | 1849.4892 | 1585.263 | 1942.7665 | 1628.1898 | 1426.5324 | 2330.3943 | 2088.0438 | 2617.263 | 2004.7659 | 2276.9931 | 2068.5915 | 1929.3426 | 1659.854 | 1821.1534 | 2837.708 | 2878.321 | 1792.6784 | 2441.7591 | 2460.7733 | 2148.525 | 2628.8545 | 1828.0216 | 2023.2773 | 1920.7153 | 1834.2989 | 1640.7739 | 2049.0364 | 2214.7591 | 1879.3869 | 2987.0216 | 3274.5915 | 2986.5403 | 2750.4012 | 1992.0733 | 2581.9568 | 2259.5619 | 1866.9631 | 1643.7812 | 1776.6494 | 1743.8977 | 1841.7296 | 1929.1096 | 2411.65 | 1835.0432 | 1750.4887 | 1906.9341 | 2637.0221 | 1921.4887 | 1806.6421 | 2018.0074 | 2647.204 | 1795.7665 | 1875.2847 | 1713.9415 | 2046.8756 | 1654.7153 | 1686.5915 | 2070.9273 | 2474.8761 | 2069.3358 | 1940.3648 | 1765.029 | 2434.1313 | 2017.4085 | 2001.175 | 2131.6568 | 2266.0585 | 2226.2847 | 2428.4012 | 1877.1313 | 1946.2409 | 1870.1608 | 2070.8171 | 1703.5625 | 1747.204 | 1769.0659 | 1803.5546 | 1735.796 | 1546.6199 | 1752.0295 | 1684.2483 | 1759.9199 | 1673.6125 | 1739.3574 | 1816.8977 | 1799.1608 | 1963.2335 | 2243.3795 | 2542.7375 | 3185.1102 | 3874.379 | 4089.9489 | 3807.0654 | 1612.7591 | 1610.7881 | 1593.0364 | 1612.7591 | 1680.5182 | 1880.4818 | 1678.9562 | 1684.0801 | 1583.117 | 1785.1313 | 1730.9341 | 1609.0216 | 1579.2994 | 1632.0654 | 1537.6494 | 1661.2267 | 1595.8102 | 1583.0801 | 1564.6494 | 1566.6494 | 1907.679 | 1531.4602 | 1422.6273 | 1571.3722 | 1693.4381 | 1511.2847 | 1569.0148 | 1373.854 | 2002.8614 | 2807.2926 | 2822.65 | 2070.7881
157 | 3.7214 | 877.8437 | 2183.9086 | 2066.882 | 2163.6047 | 2212.5015 | 2898.9823 | 2202.1563 | 1800.7079 | 1524.1121 | 1930.2389 | 1650.0177 | 1418.7522 | 2220.2419 | 2054.8908 | 2642.3481 | 2113.18 | 2244.882 | 1992.351 | 1966.705 | 1656.2478 | 1775.3776 | 2841.5044 | 2889.6785 | 1848.1357 | 2361.1858 | 2460.174 | 2157.1003 | 2574.0295 | 1752.1947 | 2009.0737 | 1879.7168 | 1835.0914 | 1642.1652 | 2096.4572 | 2146.1121 | 1817.6991 | 3047.8879 | 3366.236 | 2857.9764 | 2782.4484 | 1976.9823 | 2576.7345 | 2312.0443 | 1956.0797 | 1625.1475 | 1773.8348 | 1781.4484 | 1872.4484 | 1819.2419 | 2385.1386 | 1911.823 | 1729.9911 | 1893.7434 | 2582.4513 | 1920.2566 | 1844.0354 | 1987.8908 | 2682.9528 | 1824.5221 | 1821.7905 | 1655.0383 | 2127.9882 | 1588.9469 | 1697.5044 | 2038.469 | 2436.882 | 2065.413 | 2016.7316 | 1792.1092 | 2456.6873 | 1980.8082 | 1989.6608 | 2130.8348 | 2113.0412 | 2261.7788 | 2449.4395 | 1929.9617 | 1960.7611 | 1837.3864 | 2078.3481 | 1738.7699 | 1789.3746 | 1816.2006 | 1930.1888 | 1630.3244 | 1669.7759 | 1676.5162 | 1647.5516 | 1717.2124 | 1773.6756 | 1696.3038 | 1704.9115 | 1704.9203 | 1930.8082 | 2194.3776 | 2532.4867 | 3093.7463 | 3809.6254 | 4076.1563 | 3901.7493 | 1579.8908 | 1610.5044 | 1673.4838 | 1579.8908 | 1711.1092 | 1882.9174 | 1692.8437 | 1647.8908 | 1624.7788 | 1744.3864 | 1764.531 | 1670.8053 | 1563.7345 | 1523.0855 | 1556.8525 | 1656.3215 | 1631.7788 | 1573.4956 | 1582.9351 | 1583.1003 | 1879.8082 | 1469.3599 | 1473.4661 | 1592.8525 | 1704.2566 | 1480.7257 | 1571.1652 | 1363.826 | 1966.3864 | 2679.8023 | 2799.8082 | 2098.0266
158 | 3.7324 | 902.2424 | 2114.2025 | 2043.4764 | 2171.5782 | 2162.4995 | 2903.5079 | 2242.8363 | 1802.5079 | 1504.0546 | 1895.531 | 1566.468 | 1464.5551 | 2326.2119 | 2169.1416 | 2622.617 | 1999.4208 | 2237.2812 | 1984.0703 | 2026.7817 | 1661.5782 | 1784.0157 | 2840.1406 | 2950.5005 | 1825.1794 | 2470.2119 | 2548.3914 | 2184.1018 | 2552.1952 | 1808.2896 | 1993.1249 | 1864.8437 | 1862.0315 | 1707.0787 | 2196.7586 | 2193.9223 | 1823.1563 | 3004.7261 | 3367.4921 | 2883.1878 | 2691.1711 | 1926.6558 | 2526.6558 | 2295.617 | 1852.5068 | 1566.2886 | 1681.9601 | 1793.3673 | 1856.4764 | 1873.149 | 2332.7261 | 1883.7419 | 1738.297 | 1881.8437 | 2560.2655 | 1842.9056 | 1808.9528 | 1918.0619 | 2611.6947 | 1747.0462 | 1795.617 | 1792.2434 | 2046.7492 | 1612.7503 | 1658.9528 | 2026.6327 | 2479.4848 | 1974.8195 | 1957.0619 | 1791.2812 | 2524.0787 | 1949.1794 | 2004.8048 | 2107.7576 | 2231.9465 | 2216.6558 | 2460.3673 | 1842.9601 | 1869.0305 | 1901.8678 | 2053.9685 | 1661.9759 | 1673.1396 | 1806.0546 | 1889.9371 | 1665.5551 | 1659.6873 | 1677.3673 | 1657.9454 | 1739.0315 | 1683.5929 | 1764.5866 | 1797.2665 | 1760.149 | 1966.1175 | 2210.4533 | 2536.5782 | 3113.7503 | 3933.298 | 3936.4754 | 4043.0168 | 1629.0629 | 1592.4061 | 1685.6485 | 1629.0629 | 1712.1406 | 1792.8195 | 1665.1794 | 1650.5782 | 1625.2109 | 1714.9685 | 1675.9601 | 1605.4134 | 1559.2109 | 1669.2591 | 1558.0703 | 1648.4921 | 1557.8353 | 1633.4302 | 1566.4061 | 1541.3043 | 1883.5079 | 1431.3904 | 1368.3662 | 1555.0231 | 1683.617 | 1438.0231 | 1544.9685 | 1362.0703 | 1951.8437 | 2631.6716 | 2685.4365 | 2030.3431
159 | 3.7434 | 885.1603 | 2152.3314 | 2042 | 2172 | 2191.9115 | 2865.3786 | 2178.177 | 1846.3314 | 1483.2183 | 1930.6795 | 1599.5634 | 1454.8122 | 2324.6381 | 2079.0914 | 2588.0305 | 2022.9695 | 2231.3481 | 1914.293 | 2028.174 | 1597.9449 | 1729.4228 | 2772.177 | 2817.47 | 1827.7099 | 2339.586 | 2479.6411 | 2204.8397 | 2507.6686 | 1779.0885 | 1982.5801 | 1939.3589 | 1883.7817 | 1554.6883 | 2199.2901 | 2165.0305 | 1869.2154 | 3015.1298 | 3206.8623 | 2815.351 | 2726.3894 | 1908.8702 | 2466.4808 | 2351.4612 | 1920.3009 | 1523.3786 | 1693.8397 | 1762.9144 | 1874.7817 | 1780.9754 | 2311.1023 | 1921.5054 | 1668.351 | 1946.9391 | 2493.0609 | 1950.0384 | 1717.4533 | 1978.9391 | 2653.0993 | 1770.2016 | 1708.2792 | 1730.467 | 2054.1298 | 1594.2763 | 1644.6962 | 2112.6628 | 2416.119 | 2032.9971 | 1990.6214 | 1731.6548 | 2335.8928 | 1905.5526 | 1962.6686 | 2089.9862 | 2254.2458 | 2296.8368 | 2404.4808 | 1867.0856 | 1891.2016 | 1823.1632 | 2016.2045 | 1654.116 | 1708.2734 | 1782.3923 | 1940.6352 | 1629.4366 | 1585.467 | 1668.5801 | 1594.0029 | 1735.3481 | 1667.5221 | 1728.5526 | 1776.7404 | 1691.5251 | 1905.8869 | 2278.881 | 2514.3923 | 3022.7433 | 3963.7237 | 4063.53 | 3956.9174 | 1590.6106 | 1551.4366 | 1562.5998 | 1590.6106 | 1714.826 | 1816.1436 | 1628.1465 | 1724.4749 | 1634.4199 | 1635.7571 | 1738.707 | 1611.3038 | 1612.6932 | 1549.1219 | 1577.7817 | 1553.8013 | 1550.116 | 1561.6991 | 1547.1023 | 1541.7679 | 1835.0167 | 1386.6106 | 1501.298 | 1538.8122 | 1667.8702 | 1517.8948 | 1550.5359 | 1340.3343 | 1949.116 | 2670.4474 | 2711.8535 | 2055.1436
160 | 3.7544 | 876.3658 | 2144.5029 | 2076.354 | 2115.6976 | 2179.6858 | 2858.2286 | 2075.7094 | 1714.4469 | 1450.5088 | 1862.3835 | 1543.7891 | 1409.149 | 2237.1607 | 2057.7316 | 2574.5487 | 1990.646 | 2235.7714 | 1898.5487 | 1946.8864 | 1600.6799 | 1759.3997 | 2663.9837 | 2824.2684 | 1786.9661 | 2434.062 | 2447.326 | 2207.9086 | 2551.7596 | 1863.8393 | 1972.4572 | 1946.8628 | 1825.8348 | 1558.3599 | 2005.2817 | 2268.9248 | 1831.9203 | 2988.326 | 3363.9986 | 2864.4395 | 2727.0914 | 1974.7994 | 2521.2109 | 2249.0752 | 1902.0973 | 1568.7139 | 1671.1431 | 1729.5088 | 1811.1091 | 1846.5708 | 2240.3377 | 1921.1372 | 1679.2684 | 1912.8289 | 2605.4219 | 1853.938 | 1856.0959 | 1998.2684 | 2558.6637 | 1747.1888 | 1797.4735 | 1694.5088 | 2062.6342 | 1566.2802 | 1711.7537 | 2041.6578 | 2447.3083 | 2096.1652 | 1956.7375 | 1731.8171 | 2381.2567 | 1967.8909 | 1926.6003 | 2060.3717 | 2255.9543 | 2135.9557 | 2497.382 | 1889.9941 | 1850.0575 | 1797.0059 | 2062.6682 | 1642.5487 | 1808.3363 | 1662.6696 | 1791.3156 | 1535.3894 | 1612.4912 | 1676.5885 | 1639.6224 | 1649.115 | 1659.3658 | 1711.6858 | 1717.7433 | 1744.2567 | 1964.9823 | 2303.9941 | 2499.6401 | 3000.7773 | 3979.3142 | 4009.0177 | 3908.0575 | 1614.7198 | 1540.4115 | 1555.9543 | 1614.7198 | 1645.3377 | 1741.6578 | 1645.0856 | 1666.972 | 1589.1947 | 1720.7021 | 1666.6578 | 1617.7257 | 1604.5487 | 1581.0797 | 1568.5029 | 1582.354 | 1549.0457 | 1570.3599 | 1563.177 | 1555.3599 | 1864.4454 | 1448.8909 | 1386.8982 | 1515.0973 | 1658.4115 | 1478.149 | 1507.1032 | 1389.5767 | 1932.7773 | 2689.2227 | 2682.5088 | 2082.7655
161 | 3.7654 | 923.3619 | 2183.6293 | 2056.7355 | 2094.6352 | 2182.8663 | 2804.8387 | 2158.9577 | 1785.3255 | 1516.6598 | 1871.5654 | 1508.1367 | 1395.4012 | 2280.3619 | 2031.8692 | 2409.6833 | 2030.5625 | 2142.2428 | 1958.8938 | 1910.1367 | 1594.234 | 1750.3677 | 2691.8633 | 2797.9361 | 1810.1308 | 2458.3314 | 2361.9085 | 2148.0727 | 2546.2675 | 1770.3766 | 1948.8024 | 1963.4317 | 1857.7965 | 1642.1249 | 2077.1918 | 2205.3736 | 1853.1976 | 3082.6903 | 3293.7079 | 2895.9971 | 2839.122 | 1932.5713 | 2519.1672 | 2195.6716 | 1909.6991 | 1559.4012 | 1660.3343 | 1685.471 | 1784.8024 | 1740.7777 | 2246.6657 | 1936.4651 | 1711.8968 | 1874.2704 | 2508.5437 | 1910.8604 | 1744.1121 | 1951.6381 | 2611.0275 | 1723.7689 | 1706.3097 | 1707.4985 | 1976.9754 | 1621.061 | 1602.878 | 1971.3068 | 2379.4071 | 1969.4798 | 1928.9026 | 1764.8633 | 2304.7414 | 1986.4651 | 1910.5015 | 1999.0727 | 2188.3402 | 2125.1003 | 2423.7079 | 1817.5408 | 1851.8663 | 1816.2979 | 2163.59 | 1598.5044 | 1728.9085 | 1698.5625 | 1784 | 1562.8968 | 1585.0029 | 1562.9449 | 1623.6352 | 1635.3343 | 1640.6352 | 1708.1003 | 1670.5378 | 1669.6745 | 1865.5437 | 2270.2035 | 2520.2645 | 2923.6411 | 3832.115 | 3977.0029 | 3893.4346 | 1585.0698 | 1570.9302 | 1512.5044 | 1585.0698 | 1711.5929 | 1697.4041 | 1631.5015 | 1620.5044 | 1595.6991 | 1641.8751 | 1573.1091 | 1557.1062 | 1482.2124 | 1560.7689 | 1511.939 | 1571.4346 | 1529.6686 | 1542.0029 | 1510.8387 | 1570.4985 | 1885.3314 | 1460.6991 | 1446.7267 | 1488.8692 | 1636.7355 | 1432.5044 | 1552.3619 | 1389.1003 | 2002.5595 | 2702.5654 | 2702.2645 | 1952.5467
162 | 3.7764 | 871.1627 | 2086.1352 | 2093.1967 | 2115.5349 | 2090.0083 | 2777.5496 | 2114.0359 | 1794.852 | 1446.5644 | 1882.7675 | 1610.7591 | 1416.5349 | 2239.909 | 2033.9366 | 2566.4907 | 2042.7675 | 2155.6406 | 1963.9366 | 1782.7271 | 1563.6553 | 1715.7399 | 2714.5349 | 2758.8245 | 1778.6976 | 2323.9174 | 2413.8161 | 2173.4081 | 2514.6765 | 1775.8097 | 1889.2684 | 1817.171 | 1833.5496 | 1575.501 | 2049.6553 | 2192.2325 | 1844.2114 | 2926.4459 | 3376.1608 | 2896.0211 | 2812.6553 | 1915.3382 | 2549.3447 | 2205.7463 | 1899.2325 | 1557.0423 | 1671.7463 | 1652.6765 | 1848.6047 | 1789.8584 | 2326.2876 | 1899.8033 | 1655.2473 | 1900.4081 | 2608.7591 | 1963.9007 | 1831.034 | 1945.1057 | 2663.9007 | 1673.0782 | 1729.4292 | 1676.6765 | 1964.2114 | 1570.1416 | 1619.5561 | 1995.4292 | 2500.3363 | 1950.3171 | 1778.2345 | 1800.2601 | 2336.2601 | 1937.0782 | 1870.8456 | 1964.6976 | 2238.8584 | 2086.8245 | 2520.8437 | 1816.9577 | 1896.0487 | 1737.7124 | 2065.1986 | 1639.091 | 1741.6618 | 1733.2812 | 1763.4439 | 1554.2114 | 1621.1967 | 1592.2812 | 1669.0064 | 1643.8097 | 1712.4356 | 1759.8796 | 1693.4715 | 1709.091 | 1950.0974 | 2192.6489 | 2535.6829 | 2895.5496 | 3862.2389 | 4028.8584 | 3879.296 | 1476.3254 | 1575.9154 | 1476.7399 | 1476.3254 | 1609.262 | 1764.5202 | 1709.3088 | 1664.0275 | 1607.7463 | 1690.8796 | 1726.6175 | 1586.3235 | 1554.3511 | 1581.5349 | 1495.3171 | 1622.8796 | 1490.8245 | 1548.8309 | 1486.4862 | 1544.5708 | 1902.6406 | 1401.2896 | 1461.7252 | 1476.2537 | 1593.909 | 1446.6618 | 1515.8245 | 1361.5919 | 1943.3318 | 2610.9872 | 2676.5708 | 2021.4145
163 | 3.7874 | 923.5221 | 2168.2478 | 2042.4602 | 2177.4513 | 2130.6195 | 2820.6106 | 2112.0088 | 1793.0177 | 1509.4248 | 1865.1593 | 1581.2832 | 1356.5398 | 2163.6725 | 2029.0442 | 2489.7168 | 1965.6902 | 2239.2567 | 1921.3805 | 1933.6284 | 1566.9646 | 1695.177 | 2715 | 2815.4867 | 1914.7788 | 2337.8496 | 2459.6018 | 2114.531 | 2544.7257 | 1786.9027 | 1916.7434 | 1803.0973 | 1784.4336 | 1598.7788 | 1992.5044 | 2221.7345 | 1844 | 2927.9558 | 3342.3186 | 2872.2124 | 2773.3451 | 1977.4425 | 2501.4336 | 2184.1947 | 1875.2124 | 1550.0619 | 1633.3451 | 1728.3186 | 1800.4425 | 1781.0885 | 2221.9469 | 1881.1593 | 1692.6549 | 1806.8407 | 2533.6902 | 1809.3893 | 1663.5132 | 1934.0973 | 2594.6283 | 1710.6549 | 1698.2832 | 1678.9735 | 1974.9027 | 1545.2124 | 1614.0531 | 1991.0442 | 2429.6548 | 1891.5221 | 1931.6018 | 1703.8672 | 2412.3275 | 1941.9469 | 1924.5133 | 2111.6815 | 2137.9115 | 2186.1062 | 2494.2566 | 1813.0354 | 1792.9292 | 1710.2301 | 2122.469 | 1557.7345 | 1727.1327 | 1733.0088 | 1778.8584 | 1485.6106 | 1508.0177 | 1572.1858 | 1554.0354 | 1705.4513 | 1658.4956 | 1740.1858 | 1721.7522 | 1682.2478 | 1895.5044 | 2244.5221 | 2500.3186 | 2893.0177 | 3919.4956 | 4025.0442 | 3944.4159 | 1615.7346 | 1517.5221 | 1529.5221 | 1615.7346 | 1603.0354 | 1769.9646 | 1588.0973 | 1631.3009 | 1507.8938 | 1631.5398 | 1662.6017 | 1561.2301 | 1524.2832 | 1530.4602 | 1525.7257 | 1590.3009 | 1475.1327 | 1493.4956 | 1491.9469 | 1489.4867 | 1842.5398 | 1389.0973 | 1418.3894 | 1520.6018 | 1574.1681 | 1451.9558 | 1467.4248 | 1353.0708 | 1992.5487 | 2598.0973 | 2617.5221 | 2047.7788
164 | 3.7984 | 840.2547 | 2155.6971 | 1979.6765 | 2025.6195 | 2122.1893 | 2803.0103 | 2056.2375 | 1689.1239 | 1427.1342 | 1879.8038 | 1535.031 | 1427.6593 | 2247.0069 | 2009.7247 | 2466.2031 | 1950.4784 | 2163.8899 | 1811.3889 | 1908.0069 | 1560.2168 | 1778.5904 | 2687.0998 | 2743.0447 | 1851.8486 | 2275.3235 | 2328.6471 | 2031.3786 | 2571.0792 | 1737.9002 | 1914.0585 | 1848.2994 | 1847.8176 | 1523.0654 | 2085.6834 | 2198.062 | 1785.9036 | 3040.6971 | 3253.0207 | 2807.1342 | 2758.9862 | 2015.525 | 2503.1101 | 2134.2685 | 1812.1205 | 1455.2031 | 1670.7109 | 1674.7763 | 1877.9002 | 1798.8761 | 2294.7247 | 1939.7935 | 1656.883 | 1858.8658 | 2464.1893 | 1906.7074 | 1747.8314 | 1930.1205 | 2631.8038 | 1620.3683 | 1744.9897 | 1665.7419 | 1986.8658 | 1555.9553 | 1579.0688 | 1976.8623 | 2466.6283 | 2010.7487 | 1866.0688 | 1694.8382 | 2282.1067 | 1945.7556 | 1863.2581 | 2032.2203 | 2228.0585 | 2235.8038 | 2451.1755 | 1868.9553 | 1859.9484 | 1878.5319 | 2150.7797 | 1615.086 | 1677.2478 | 1714.8486 | 1765.9897 | 1620.6593 | 1533.1239 | 1603.6524 | 1529.8967 | 1608.2134 | 1624.0034 | 1648.7935 | 1738.5663 | 1742.7281 | 1899.296 | 2243.1101 | 2547.9277 | 3023.9277 | 3841.969 | 3991.8382 | 3902.6058 | 1509.9759 | 1523.0138 | 1451.1583 | 1509.9759 | 1694.5319 | 1733.8864 | 1552.1686 | 1640.8899 | 1552.9415 | 1602.9449 | 1592.1549 | 1543.1136 | 1527.8623 | 1455.1721 | 1478.2994 | 1581.5491 | 1531.0482 | 1536.9621 | 1516.8382 | 1531.7901 | 1846.0757 | 1423.0413 | 1359.3476 | 1548.5973 | 1594.0585 | 1483.7625 | 1448.0792 | 1385.7143 | 1871.2134 | 2741.852 | 2704.7006 | 2061.8038
165 | 3.8094 | 913.056 | 2067.8584 | 1888.1554 | 2204.1032 | 2175.647 | 2806 | 2124.3707 | 1725.2203 | 1465.3707 | 1821.6784 | 1543.3707 | 1329.9597 | 2250.0698 | 1931.0384 | 2524.0875 | 2087.2851 | 2133.3845 | 1923.8899 | 1909.0246 | 1623.236 | 1659.3864 | 2716.4405 | 2756.0787 | 1808.2517 | 2369.4405 | 2513.7974 | 2141.645 | 2594.4563 | 1709.1731 | 1930.5437 | 1932.6086 | 1793.5998 | 1542.5821 | 1993.292 | 2216.5978 | 1758.0787 | 2953.4248 | 3259.3619 | 2846 | 2756.0069 | 1877.7325 | 2534.6224 | 2210.6155 | 1846.4965 | 1513.4651 | 1587.472 | 1609.8741 | 1847.9459 | 1763.1101 | 2216.2744 | 1879.119 | 1622.4965 | 1820.6922 | 2518.8741 | 1821.7483 | 1700.5418 | 1964.1662 | 2574.8741 | 1726.1032 | 1742.8653 | 1591.1573 | 1947.8427 | 1542.4336 | 1598.6224 | 1937.6765 | 2359.1101 | 1937.2291 | 1885.1504 | 1647.0875 | 2314.0856 | 1876.9509 | 1938.3776 | 2095.0403 | 2243.82 | 2178.4651 | 2501.0403 | 1854.7256 | 1846.1642 | 1743.3461 | 2085.9774 | 1640.0787 | 1748.8741 | 1671.1259 | 1762.3864 | 1521.9056 | 1567.7256 | 1503.6608 | 1500.4877 | 1669.6696 | 1624.764 | 1590.0226 | 1614.18 | 1664.4563 | 1983.3638 | 2275.472 | 2527.5978 | 3004.3688 | 3831.1593 | 3946.0856 | 4074.9321 | 1502.8269 | 1525.9459 | 1496.355 | 1502.8269 | 1559.4563 | 1701.3461 | 1601.0944 | 1608.7168 | 1535.2291 | 1588.1957 | 1636.6224 | 1574.3638 | 1488.5978 | 1505.2675 | 1562.939 | 1452.7778 | 1543.4739 | 1526.1416 | 1470.7866 | 1471.2045 | 1866.8043 | 1434.292 | 1459.4965 | 1431.6853 | 1611.3147 | 1414.8112 | 1470.2606 | 1302.8584 | 1933.9125 | 2699.0629 | 2617.7325 | 2005.1573
166 | 3.8204 | 853.972 | 2062.8673 | 2021.7109 | 2147.9159 | 2093.7183 | 2804.2625 | 2082.674 | 1739.3083 | 1427.0929 | 1742.0546 | 1503.1313 | 1390.9248 | 2315.9631 | 1995.944 | 2468.2891 | 1976.497 | 2220 | 1915.3186 | 1848.2242 | 1637.8879 | 1623.7183 | 2744 | 2761.3643 | 1824 | 2395.0944 | 2376.5325 | 2180.9631 | 2620.3938 | 1718.9351 | 1901.6917 | 1786.6431 | 1708.354 | 1576.7301 | 1953.5118 | 2253.4675 | 1763.3643 | 2977.702 | 3280.0944 | 2843.3658 | 2817.0103 | 1862.9336 | 2509.1224 | 2126.6062 | 1814.7758 | 1480.0841 | 1615.1224 | 1601.5516 | 1781.4675 | 1768.9911 | 2295.3835 | 1826.6062 | 1591.7566 | 1864.1962 | 2508.9823 | 1806.6165 | 1795.9071 | 1911.2994 | 2567.8407 | 1668.795 | 1793.4129 | 1600.6917 | 1939.7095 | 1508.5974 | 1574.4395 | 1977.7861 | 2365.8879 | 1890.3171 | 1835.2227 | 1593.5871 | 2378.6371 | 1997.2065 | 1960.7389 | 2035.3923 | 2168.6268 | 2145.4204 | 2441.4204 | 1773.879 | 1918.2714 | 1762.674 | 2024.0738 | 1642.9823 | 1738.0295 | 1677.795 | 1725.6254 | 1515.7758 | 1488.3923 | 1545.6726 | 1531.8968 | 1648.5236 | 1609.4115 | 1653.7758 | 1687.5236 | 1692.2434 | 1880.0752 | 2304.6917 | 2563.0944 | 3088.7109 | 3716.9248 | 4012.5988 | 3881.7389 | 1493.8216 | 1461.8216 | 1456.1681 | 1493.8216 | 1587.4675 | 1723.7286 | 1605.7109 | 1591.5044 | 1486.4956 | 1661.7669 | 1612.7198 | 1473.5133 | 1523.0295 | 1520.9631 | 1435.6814 | 1562.2905 | 1446.851 | 1534.3274 | 1516.7773 | 1483.0192 | 1792.3643 | 1390.5044 | 1426.5147 | 1412.1121 | 1629.4026 | 1403.5044 | 1423.3835 | 1294.6726 | 1989.0664 | 2700.7301 | 2605.2316 | 2014.4395
167 | 3.8314 | 853.3225 | 2190.8318 | 1978.2094 | 2147.3383 | 2185.2586 | 2747.6687 | 2004.8624 | 1749.355 | 1462.6617 | 1848.1209 | 1474.6932 | 1421.7502 | 2281.2498 | 1997.7493 | 2512.7581 | 1858.4671 | 2216.2104 | 1856.6775 | 1856.4838 | 1632.1455 | 1716.589 | 2742.9115 | 2772.1858 | 1823.5565 | 2330.9754 | 2395.0806 | 2143.9115 | 2532.177 | 1683.7739 | 1890.8791 | 1906.9528 | 1790.4838 | 1500.7414 | 2105.8476 | 2230.5654 | 1775.4356 | 2901.0236 | 3214.645 | 2757.2409 | 2746.9518 | 1927.7424 | 2443.1937 | 2211.88 | 1808.8545 | 1483.6047 | 1550.8869 | 1587.6932 | 1763.8466 | 1736.298 | 2271.1376 | 1913.6942 | 1616.468 | 1870.6372 | 2440.9351 | 1826.3471 | 1752.3156 | 1888.3707 | 2597.4031 | 1628.7335 | 1673.5153 | 1592.177 | 1997.6293 | 1560.0894 | 1553.7502 | 1901.8466 | 2363.8948 | 1933.0973 | 1944.7424 | 1718.7257 | 2259.9439 | 1932.823 | 1916.9597 | 2047.9518 | 2123.0561 | 2156.4523 | 2454.8309 | 1738.3865 | 1889.2507 | 1684.7414 | 2096.0806 | 1570.4198 | 1637.2173 | 1637.0079 | 1778.2507 | 1509.9754 | 1505.4268 | 1569.4435 | 1560.057 | 1630.3874 | 1590.7739 | 1645.4356 | 1671.6047 | 1665.1534 | 1851.3865 | 2288.9606 | 2495.0649 | 3041.5005 | 3751.3786 | 3961.702 | 3840.298 | 1553.5732 | 1522.2586 | 1461.9597 | 1553.5732 | 1568.1131 | 1748.8791 | 1563.177 | 1608.7581 | 1470.7739 | 1617.3471 | 1601.1455 | 1525.8791 | 1419.3953 | 1481.0491 | 1491.057 | 1503.0403 | 1408.4513 | 1507.7345 | 1406.1691 | 1447.9518 | 1806.1691 | 1407.7581 | 1377.2576 | 1418.177 | 1575.5723 | 1419.9115 | 1402.2016 | 1316.7906 | 1922.6047 | 2625.3058 | 2750.8309 | 1994.7984
168 | 3.8424 | 862.5256 | 2042.4744 | 2005.3166 | 2198.6863 | 2116.6332 | 2829.6342 | 2070.5275 | 1733.0531 | 1404.9484 | 1846.7915 | 1513.527 | 1390.3692 | 2309.5796 | 2083.8962 | 2500.7399 | 1942.473 | 2126.6839 | 1850.4204 | 1867.1583 | 1583.8967 | 1684.473 | 2713.6858 | 2731.0516 | 1808.4233 | 2368.8466 | 2395.5796 | 2112.8972 | 2582.1062 | 1726.4735 | 1887.2635 | 1766.4194 | 1801.7896 | 1560.792 | 2005.4744 | 2128.1038 | 1759.5275 | 2972.8422 | 3228.3692 | 2875.0034 | 2723.0521 | 1874 | 2425.8412 | 2093.2616 | 1853.0526 | 1492.7394 | 1570.1062 | 1625.0541 | 1787.209 | 1789.8953 | 2205.0516 | 1835.1578 | 1582.1568 | 1861.8948 | 2481.9498 | 1787.1052 | 1621.3137 | 1917.8436 | 2600.2664 | 1670.1578 | 1776.0024 | 1641.1593 | 1962.3171 | 1471.1033 | 1523.1057 | 1920.8427 | 2403.6868 | 1863.5265 | 1891.7891 | 1728.6878 | 2210.6318 | 1877.8962 | 1916.4735 | 2017.8446 | 2166.0555 | 2074.7881 | 2426.0005 | 1815.9489 | 1802.3673 | 1747.1593 | 2099.9464 | 1507.5796 | 1683.5821 | 1659.1062 | 1687.001 | 1553.791 | 1584.1593 | 1580.1578 | 1489.0005 | 1593.8412 | 1632.1057 | 1629.1593 | 1685.1057 | 1647.1042 | 1927.2129 | 2166.4223 | 2450.2099 | 2983.6834 | 3806.4218 | 3912.208 | 3897.5777 | 1470.527 | 1456.1583 | 1455.3712 | 1470.527 | 1546.8422 | 1741.0024 | 1610.055 | 1601.1062 | 1510.8958 | 1577.4739 | 1559.1568 | 1519.2124 | 1403.7901 | 1390.6298 | 1426.5236 | 1502.6848 | 1441.6322 | 1429.6844 | 1435.0521 | 1425.4725 | 1833.0531 | 1402.9469 | 1427.1598 | 1469.1057 | 1611.001 | 1391.7905 | 1412.0502 | 1290.7881 | 1934.2119 | 2704.263 | 2719.7384 | 1988.9479
169 | 3.8534 | 888.5074 | 2020.1269 | 1966.3687 | 2115.6135 | 2060.8673 | 2731.767 | 2028.4012 | 1713.5988 | 1369.6873 | 1770.2861 | 1489.4012 | 1359.6372 | 2281.5339 | 2018.7846 | 2391.0502 | 1973.2832 | 2138.7404 | 1871.8466 | 1849.9233 | 1510.8437 | 1707.6254 | 2646.407 | 2771.764 | 1726.9735 | 2213.9941 | 2365.4572 | 2028.7228 | 2541.1858 | 1735.3894 | 1869.4572 | 1823.6519 | 1790.7315 | 1471.0236 | 1981.3098 | 2183.8053 | 1710.7817 | 2971.6873 | 3198.2861 | 2751.7257 | 2744.9853 | 1868.2891 | 2453.7433 | 2147.587 | 1856.531 | 1400.7257 | 1533.9882 | 1575.2449 | 1842.5988 | 1771.0384 | 2239.1445 | 1834.5605 | 1620.7404 | 1854.0295 | 2396.3216 | 1785.0413 | 1700.1033 | 1865.0885 | 2486.7876 | 1672.2212 | 1684.3127 | 1586.1062 | 1915.6873 | 1543.5074 | 1505.5723 | 1900.5841 | 2313.9056 | 1857.0767 | 1899.1209 | 1601.0885 | 2208.2183 | 1821.6755 | 1928.2449 | 1932.3983 | 2068.6697 | 2123.6372 | 2412.7375 | 1766.3216 | 1842.0147 | 1698.8555 | 2133.1829 | 1489.4632 | 1569.6224 | 1622.8584 | 1653.4012 | 1500.413 | 1531.4425 | 1575.5368 | 1471.5192 | 1625.8967 | 1618.7906 | 1575.469 | 1671.2714 | 1684.351 | 1841.7788 | 2127.593 | 2480.5664 | 3016.7906 | 3780.6608 | 3992.5663 | 3954.6431 | 1444.649 | 1438.2743 | 1358.0236 | 1444.649 | 1545.3628 | 1651.0649 | 1519.5723 | 1567.5192 | 1473.8437 | 1567.4159 | 1595.3392 | 1453.0767 | 1386.4779 | 1452.7551 | 1519.056 | 1480.1681 | 1415.5074 | 1423.9351 | 1451.4041 | 1466.5605 | 1813.3392 | 1421.233 | 1357.0265 | 1455.0767 | 1572.8584 | 1354.9853 | 1499.3009 | 1338.5339 | 1889.7758 | 2699.3952 | 2661 | 1976.5546
170 | 3.8644 | 870.8594 | 2095.0806 | 1851.2626 | 2039.8525 | 2123.204 | 2710.5265 | 2063.5403 | 1674.9862 | 1399.4032 | 1759.4597 | 1523.2281 | 1340.6175 | 2241.4459 | 1885.1509 | 2428.5162 | 2028.2074 | 2129.8284 | 1816.3997 | 1817.7615 | 1520.8387 | 1657.4562 | 2548.4238 | 2811.643 | 1735.7512 | 2355.7615 | 2308.3363 | 2129.4252 | 2452.1854 | 1736.5231 | 1804.3859 | 1756.6775 | 1686.4631 | 1527.4562 | 2020.0462 | 2179.379 | 1651.5438 | 2986.5059 | 3197.2626 | 2786.8318 | 2789.5438 | 1785.6141 | 2376.5334 | 2083.0737 | 1887.2881 | 1438.6106 | 1512.4101 | 1616.1337 | 1804.9897 | 1784.4562 | 2196.6878 | 1779.1337 | 1620.705 | 1763.909 | 2461.1902 | 1729.4666 | 1708.1268 | 1923.2281 | 2592.9793 | 1654.9021 | 1657.6844 | 1534.1682 | 1935.6003 | 1519.0669 | 1485.0772 | 1905.1406 | 2394.3687 | 1885.4562 | 1803.705 | 1657.2212 | 2270.0034 | 1818.9297 | 1970.2143 | 2034.7547 | 2206.8422 | 2102.5334 | 2473.9759 | 1836.825 | 1802.3191 | 1729.8112 | 2103.9125 | 1591.5968 | 1589.295 | 1598.2522 | 1687.4597 | 1556.3584 | 1506.2984 | 1469.0034 | 1475.9897 | 1608.6878 | 1653.674 | 1536.53 | 1705.3122 | 1672.3756 | 1857.0138 | 2234.4528 | 2599.8318 | 3048.2006 | 3815.6809 | 3807.3363 | 4095.8815 | 1455.235 | 1396.53 | 1417.5438 | 1455.235 | 1562.6072 | 1655.4735 | 1504.326 | 1572.4528 | 1484.765 | 1556.0669 | 1553.2178 | 1484.1578 | 1445.7512 | 1403.6072 | 1420.9297 | 1494.144 | 1394.9263 | 1420.8387 | 1418.4769 | 1414.6844 | 1773.6878 | 1433.8284 | 1341.4562 | 1476.8952 | 1539.4528 | 1397.6141 | 1441.3687 | 1297.674 | 1971.3019 | 2571.1116 | 2663.5472 | 2061.7374
171 | 3.8754 | 908.6784 | 2100.4837 | 1959.5841 | 2058.5998 | 2035.4582 | 2676.4268 | 2067.763 | 1633.8899 | 1457.7365 | 1758.1003 | 1512.7473 | 1348.0895 | 2193.2635 | 1887.1692 | 2367.9116 | 1955.411 | 2074.2635 | 1862.0206 | 1787.5103 | 1492.4946 | 1640.0738 | 2674.0835 | 2645.7267 | 1682.0374 | 2321.3048 | 2416.2321 | 2005.9528 | 2558.5998 | 1682.4631 | 1816.4268 | 1712.8742 | 1696.5634 | 1509.8053 | 1915.1475 | 2156.3314 | 1672.7739 | 2885.706 | 3302.3628 | 2738.5211 | 2801.7365 | 1783.5054 | 2360.6106 | 2071.2842 | 1821.8633 | 1419.763 | 1601.5369 | 1547.4317 | 1779.7375 | 1755.0835 | 2184.3107 | 1715.9794 | 1655.4257 | 1714.3736 | 2323.5792 | 1748.3422 | 1620.2635 | 1918.8319 | 2515.9892 | 1583.1632 | 1633.9371 | 1580.6893 | 1880.5841 | 1484.6264 | 1483.763 | 1875.8221 | 2340.4317 | 1871.0266 | 1848.8948 | 1623.1947 | 2278.5998 | 1845.2213 | 1905.6627 | 1982.1209 | 2191.5683 | 2085.8053 | 2380.8731 | 1765.8633 | 1850.6795 | 1620.5271 | 2070.4002 | 1521.2733 | 1555.2901 | 1661.8151 | 1682.4317 | 1470.1003 | 1473.5526 | 1481.3579 | 1441.6106 | 1594.6106 | 1587.4051 | 1513.2635 | 1712.3205 | 1640.2842 | 1896.3205 | 2204.4474 | 2549.9105 | 2959.1957 | 3775.9105 | 3899.7522 | 3952.0738 | 1461.7679 | 1373.1741 | 1428.6627 | 1461.7679 | 1535.7896 | 1688.3205 | 1563.4366 | 1545.4897 | 1472.4159 | 1520.1947 | 1512.7424 | 1492.7207 | 1383.9843 | 1379.2055 | 1432.4572 | 1468.769 | 1408.058 | 1398.1318 | 1459.7414 | 1388.8791 | 1755.6735 | 1376.6529 | 1326.8476 | 1382.7739 | 1514.6372 | 1391.9262 | 1384.0315 | 1238.2055 | 1949.3894 | 2693.7001 | 2693.6421 | 1955.8319
172 | 3.8864 | 825.0177 | 2030.8186 | 1927.9779 | 2026.3407 | 2078.584 | 2754.4425 | 2041.1327 | 1710.5133 | 1385.7566 | 1799.1106 | 1492.7743 | 1348.5929 | 2247.5354 | 1997.9734 | 2519.3406 | 2035.0973 | 2143.0885 | 1785.6062 | 1754.4425 | 1472.2921 | 1651 | 2575.5177 | 2710.7168 | 1763.3141 | 2249.3098 | 2478.1194 | 2041.8009 | 2529.1903 | 1669.9779 | 1878.6681 | 1786.0796 | 1743.9159 | 1512.3009 | 1933.6416 | 2117.8805 | 1701.9159 | 3003.6017 | 3231.9734 | 2754.7124 | 2733.8274 | 1816.531 | 2380.8673 | 2066.646 | 1847.031 | 1400.4602 | 1595.1947 | 1564.646 | 1894.9027 | 1649.5974 | 2217.0841 | 1792.2921 | 1559.6283 | 1714.1062 | 2356.2345 | 1770.2434 | 1687.7655 | 1957.5177 | 2468.2566 | 1598.4602 | 1676.9513 | 1545.5752 | 1862.5133 | 1491.1018 | 1466.2921 | 2036.4159 | 2353.9823 | 1911.2876 | 1849.7566 | 1612.2522 | 2244 | 1871.9956 | 1838.4292 | 1959.9469 | 2179.4425 | 2093.8982 | 2302.7567 | 1782.177 | 1924.8539 | 1743.4203 | 2103.3717 | 1465.9513 | 1659.6947 | 1561.3319 | 1698.9336 | 1522.5089 | 1475.0044 | 1483.7788 | 1455.6593 | 1605.7079 | 1557.438 | 1578.7301 | 1629.9027 | 1628.6239 | 1816.5487 | 2216.6328 | 2535.3628 | 3101.5088 | 3777.2389 | 3824.3717 | 3963.6106 | 1377.3894 | 1433.7611 | 1351.5575 | 1377.3894 | 1543.7212 | 1598.5752 | 1515.7788 | 1571.4336 | 1503.5133 | 1519.1239 | 1551.4336 | 1441.8496 | 1391.4557 | 1414.7743 | 1328.9425 | 1544.3628 | 1420.8584 | 1422.0619 | 1335.4071 | 1403.4203 | 1733.2212 | 1416.9779 | 1363.8628 | 1411.9159 | 1567.9292 | 1379.2699 | 1367.969 | 1276.6239 | 1934.2212 | 2709.1681 | 2681.3495 | 2005.2035
173 | 3.8974 | 838.7906 | 2080.2782 | 1866.4562 | 2045.8594 | 1985.0826 | 2766.0767 | 2001.2969 | 1757.9626 | 1315.0531 | 1707.4484 | 1430.939 | 1337.2438 | 2163.4297 | 1889.9577 | 2341.6765 | 1992.9233 | 2178.761 | 1744.2734 | 1769.6126 | 1515.35 | 1649.4031 | 2675.7345 | 2759.8485 | 1671.0531 | 2315.2655 | 2278.2704 | 2093 | 2573.4031 | 1618.3796 | 1767.8437 | 1728.0797 | 1741.5064 | 1464.7375 | 1914.2655 | 2110.2812 | 1712.2812 | 2923.1888 | 3137.531 | 2753.0157 | 2707.8358 | 1875.0157 | 2401.4267 | 2096.0452 | 1815.764 | 1428.1485 | 1493.6234 | 1572.3736 | 1929.6627 | 1699.9125 | 2214.0187 | 1824.058 | 1576.6234 | 1768.2281 | 2330.1859 | 1834.1032 | 1709.9361 | 1865.4562 | 2506.353 | 1623.0688 | 1693.9892 | 1533.6844 | 1914.8672 | 1485.2704 | 1506.0423 | 1942.7719 | 2326.9783 | 1896.5595 | 1779.2812 | 1592.7906 | 2247.65 | 1825.3874 | 1904.3422 | 1941.6687 | 2193.0157 | 2090.6047 | 2432.7719 | 1678.0108 | 1804.7906 | 1694.7876 | 2080.5595 | 1484.0157 | 1641.8299 | 1669.9892 | 1709.3844 | 1516.9626 | 1524.0079 | 1474.825 | 1429.8672 | 1562.9921 | 1524.9734 | 1586.1298 | 1676.0688 | 1588.5172 | 1891.4297 | 2288.5625 | 2431.6529 | 2918.2842 | 3795.4031 | 3811.2547 | 3973.2203 | 1374.3953 | 1410.35 | 1343.5733 | 1374.3953 | 1458.589 | 1591.4749 | 1496.2173 | 1487.7984 | 1544.9174 | 1577.5516 | 1464.1485 | 1465.7453 | 1368.8859 | 1360.1249 | 1444.0157 | 1426.3953 | 1346.7188 | 1349.4907 | 1345.7296 | 1360.6047 | 1751.528 | 1352.7188 | 1349.6765 | 1413.6126 | 1527.4375 | 1369.4749 | 1405.2812 | 1239.1406 | 1940.4375 | 2681.9361 | 2600.7797 | 2011.7906
174 | 3.9084 | 869.6824 | 2004.4548 | 1876.0649 | 2036.6209 | 2106.4621 | 2692.9169 | 1955.4838 | 1685.5595 | 1325.3141 | 1728.8594 | 1451.8122 | 1270.0506 | 2164.5595 | 1942.6101 | 2373.704 | 2055.4695 | 2113.1913 | 1788.325 | 1716.4621 | 1482.8304 | 1635.4838 | 2599.6642 | 2682.5162 | 1703.4258 | 2374.6573 | 2238.1298 | 2082.6391 | 2548.3648 | 1706.4548 | 1826.5595 | 1753.3899 | 1660.2527 | 1474.5019 | 1988.3141 | 2121.0506 | 1722.7979 | 3072.7979 | 3287.9026 | 2697.5703 | 2777.4695 | 1813.2527 | 2303.7075 | 2136.6824 | 1837.7979 | 1400.8196 | 1507.6824 | 1456.8122 | 1788.4297 | 1697.9026 | 2247.3141 | 1717.6028 | 1597.5054 | 1755.7222 | 2386.5054 | 1768.0325 | 1626.2743 | 1865.0831 | 2568.2601 | 1584.5811 | 1625.0108 | 1535.7148 | 1870.1986 | 1441.4656 | 1456.296 | 1929.6283 | 2202.5413 | 1828.5993 | 1795.7365 | 1630.2527 | 2269.4297 | 1874.4946 | 1908.8953 | 1994.5958 | 2134.0325 | 2069.3466 | 2422.5344 | 1748 | 1829.4911 | 1639.8447 | 2010.8304 | 1420.3574 | 1534.9927 | 1599.2419 | 1664.0649 | 1446.5811 | 1492.8736 | 1491.9061 | 1398.3393 | 1583.5128 | 1515.9892 | 1532.7075 | 1624.3141 | 1588.8771 | 1883.5885 | 2305.4262 | 2559.7365 | 3019.5634 | 3781.6101 | 3830.2237 | 3998.5452 | 1396.0325 | 1378.9675 | 1437.3756 | 1396.0325 | 1495.177 | 1634.6318 | 1440.0035 | 1500.9494 | 1427.7075 | 1550.8338 | 1450.2129 | 1455.4946 | 1367.4946 | 1425.8196 | 1389.7075 | 1434.6391 | 1337.0865 | 1335.0541 | 1391.0506 | 1342.2527 | 1798.7546 | 1338.9169 | 1380.3141 | 1374.8663 | 1495.4115 | 1397.8485 | 1425.4007 | 1246.2311 | 1909.1696 | 2601.5595 | 2574.3176 | 2044.1986
175 | 3.9194 | 834.5959 | 2055.4454 | 1901.7787 | 2035.6637 | 2078.3953 | 2707.938 | 1953.18 | 1717.3451 | 1294.3186 | 1642.2979 | 1490.7522 | 1234.3393 | 2203.5546 | 1924.9587 | 2330.5516 | 1947.7905 | 2140.056 | 1762.5162 | 1708.4986 | 1461.4867 | 1625.1386 | 2634.2891 | 2700.2095 | 1805.7699 | 2232.8909 | 2309.1062 | 2015.0207 | 2474.5871 | 1751.761 | 1858.0678 | 1773.3983 | 1669.9027 | 1402.4867 | 1954.4071 | 2078.823 | 1675.3569 | 3013.7109 | 3225.1386 | 2745.7729 | 2686.2891 | 1811.118 | 2364.764 | 2091.1976 | 1771.6401 | 1360.5634 | 1465.1091 | 1497.5546 | 1735.6519 | 1630.5634 | 2209.6637 | 1785.1091 | 1603.4956 | 1677.1033 | 2390.5546 | 1775.9675 | 1627.8584 | 1843.1298 | 2489.9793 | 1631.5929 | 1640.0413 | 1505.236 | 1829.0973 | 1480.3126 | 1480.4543 | 1856.1298 | 2298.354 | 1813.6637 | 1796.2891 | 1631.4454 | 2210.5516 | 1864.6224 | 1921.3068 | 1907.6165 | 2150.1475 | 2048.9203 | 2374.9528 | 1745.7817 | 1750.0708 | 1692.2655 | 1995.0325 | 1414.1298 | 1610.3068 | 1583.3245 | 1696.1563 | 1491.0973 | 1489.0885 | 1500.9085 | 1450.4543 | 1514.882 | 1507.059 | 1573.7994 | 1586.8024 | 1661.9882 | 1853.8997 | 2186.8437 | 2552.2478 | 2915.4986 | 3757.4424 | 3881.7994 | 3941.177 | 1415.3363 | 1353.8407 | 1351.5929 | 1415.3363 | 1451.7434 | 1628.4631 | 1507.056 | 1541.6519 | 1466.0265 | 1599.7316 | 1502.7021 | 1444.236 | 1357.3451 | 1386.9498 | 1436.4838 | 1368.2684 | 1380.1062 | 1377.764 | 1343.1386 | 1353.0118 | 1731.5369 | 1359.2065 | 1341.2773 | 1446.6726 | 1528.8437 | 1305.4071 | 1437.5634 | 1231.7817 | 1928.8495 | 2632.3746 | 2621.7434 | 2005.8702
176 | 3.9304 | 890.415 | 2042.0383 | 1811.4145 | 2042.3776 | 2064.9051 | 2625.6042 | 2006.7168 | 1671.0192 | 1341.7738 | 1661.7551 | 1483.7547 | 1297.7178 | 2199.6042 | 1873.3397 | 2259.3766 | 1917.0369 | 2042.8112 | 1807.3771 | 1798.0196 | 1503.0192 | 1616.793 | 2674.1897 | 2721.4542 | 1736.6037 | 2240.1878 | 2243.4897 | 2045.2827 | 2564.5846 | 1698.2084 | 1809.7173 | 1824.5668 | 1707.0762 | 1444.1701 | 1981.208 | 2109.2645 | 1675.9631 | 2955.5472 | 3226.9056 | 2755.6042 | 2743.0944 | 1784.2832 | 2396.85 | 2086.4902 | 1683.4332 | 1345.9804 | 1474.264 | 1500.2453 | 1696.943 | 1618.6972 | 2213.8117 | 1781.5855 | 1601.0388 | 1806.0752 | 2379.5103 | 1693.0379 | 1543.3206 | 1854.6239 | 2460.9248 | 1604.6411 | 1681.2458 | 1464.0939 | 1789.3579 | 1357.2453 | 1444.9435 | 1882.7173 | 2232.1883 | 1830.0374 | 1842.0762 | 1614.114 | 2143.7537 | 1821.6794 | 1907.6421 | 1990.8299 | 2117.585 | 2025.7365 | 2449.0374 | 1728.9813 | 1785.8495 | 1634.5285 | 2082.943 | 1445.3963 | 1599.208 | 1530.735 | 1636.3584 | 1428.265 | 1455.1519 | 1447.5462 | 1406.6421 | 1531.6042 | 1496.9621 | 1505.0939 | 1596.6416 | 1609.1696 | 1819.1696 | 2242.2645 | 2583.6967 | 3019.6603 | 3651.1519 | 3810.9808 | 3925.0743 | 1392.6234 | 1329.8112 | 1323.1318 | 1392.6234 | 1476.6603 | 1579.15 | 1419.1691 | 1504.6033 | 1377.528 | 1553.114 | 1482.0374 | 1389.8117 | 1321.2827 | 1364.3776 | 1389.208 | 1396.7738 | 1321.4327 | 1415.2837 | 1337.208 | 1416.3958 | 1802.4336 | 1319.5659 | 1344.3584 | 1420.5664 | 1583.208 | 1337.7925 | 1400.6421 | 1218 | 1879.3958 | 2593.1505 | 2632.3215 | 1967.8117
177 | 3.9414 | 896.4916 | 1998.2419 | 1869.3471 | 2029.0187 | 2111.5723 | 2600.4858 | 2010.9007 | 1649.6922 | 1316.2045 | 1616.9194 | 1480.8181 | 1228.3864 | 2156.2045 | 1883.4238 | 2314.6362 | 1983.9213 | 2054.3058 | 1820.0845 | 1725.1632 | 1478.1013 | 1562.0619 | 2570.2458 | 2588.6735 | 1727.5949 | 2302.3716 | 2308.5142 | 2063.2478 | 2594.6942 | 1616.7955 | 1778.2045 | 1738.5536 | 1644.0806 | 1422.5497 | 1949.2026 | 2085.4877 | 1603.3658 | 2946.7974 | 3217.3097 | 2716.7748 | 2740.7748 | 1764.9194 | 2308.1219 | 2110.7581 | 1767.4877 | 1406.7581 | 1486.8368 | 1488.3284 | 1741.06 | 1694.5949 | 2177.2045 | 1731.9774 | 1508.9774 | 1808.5968 | 2309.0993 | 1693.9155 | 1564.1219 | 1750.4464 | 2445.1819 | 1630.7374 | 1621.7168 | 1490.1258 | 1842.6529 | 1360.8348 | 1447.1613 | 1854.7955 | 2262.1032 | 1854.2065 | 1785.9361 | 1538.7748 | 2242.7335 | 1791.9007 | 1843.7374 | 2008.6961 | 2099.9194 | 1988.4032 | 2477.1239 | 1723.589 | 1744.0806 | 1621.6922 | 2111.2271 | 1428.2891 | 1559.3677 | 1600.6549 | 1641.8781 | 1354.1013 | 1376.4464 | 1520.6362 | 1367.3845 | 1499.7935 | 1500.3284 | 1530.6362 | 1582.9607 | 1627.9794 | 1844.8348 | 2196.2891 | 2705.9007 | 3015.6155 | 3568.0393 | 3827.0845 | 4038.4897 | 1324.1426 | 1334.6136 | 1340.0845 | 1324.1426 | 1486.6529 | 1654.9833 | 1479.7148 | 1552.0206 | 1400.3884 | 1484.4671 | 1510.94 | 1374.4238 | 1350.8368 | 1343.3471 | 1335.4936 | 1368.1032 | 1419.6381 | 1352.7148 | 1308.4671 | 1440.8613 | 1815.0452 | 1325.8387 | 1350.8574 | 1387.5516 | 1538.6342 | 1330.3471 | 1353.2232 | 1212.94 | 1921.1632 | 2633.94 | 2554.6342 | 1943.6097
178 | 3.9524 | 857.1224 | 2046.7832 | 1870.2124 | 2056.3997 | 2098.7389 | 2673.2404 | 1953.2979 | 1681.1917 | 1304.3673 | 1607.0575 | 1437.5059 | 1245.82 | 2132.9587 | 1966.5014 | 2245.2124 | 1928.1431 | 2068.0207 | 1756.1917 | 1737.9351 | 1481.3628 | 1556.6534 | 2540.9956 | 2588.2286 | 1667.1593 | 2222.531 | 2218.559 | 2004.2286 | 2577.6445 | 1633.5472 | 1763.7345 | 1696.2331 | 1668.1593 | 1447.8407 | 1969.5914 | 2124.4366 | 1627.0487 | 2920.09 | 3135.3304 | 2725.0324 | 2761.1431 | 1745.9307 | 2307.7669 | 2047.2979 | 1806.4366 | 1337.5059 | 1488.8083 | 1447.4204 | 1765.2448 | 1635.3879 | 2162.559 | 1765.5383 | 1554.8245 | 1713.1269 | 2355.9956 | 1743.7669 | 1573.7714 | 1811.5472 | 2484.9838 | 1587.2168 | 1577.7758 | 1430.7596 | 1833.5383 | 1403.1062 | 1485.8407 | 1865.2286 | 2245.9513 | 1803.2979 | 1839.1917 | 1560.0207 | 2258.3997 | 1746.0575 | 1773.7138 | 1962.2286 | 2069.5221 | 2092.2979 | 2448.5796 | 1765.7994 | 1751.6607 | 1639.5014 | 2058.9838 | 1375.3761 | 1552.4041 | 1567.9145 | 1629.4204 | 1368.5959 | 1430.5265 | 1452.792 | 1433.5428 | 1536.3097 | 1467.4941 | 1459.4572 | 1571.8334 | 1622.469 | 1879.91 | 2118.4528 | 2642.1549 | 2954.3717 | 3628.3113 | 3761.4366 | 4047.8451 | 1337.0207 | 1309.6327 | 1280.2286 | 1337.0207 | 1473.3259 | 1578 | 1468.351 | 1542.4366 | 1386.2979 | 1517.2035 | 1478.1549 | 1458.2035 | 1347.8938 | 1336.1386 | 1277.7641 | 1350.8982 | 1313.6003 | 1349.7758 | 1357.351 | 1351.2448 | 1723 | 1286.8083 | 1345.8776 | 1376.0531 | 1507.0531 | 1320.1549 | 1376.4897 | 1187.8776 | 1912.8731 | 2600.6283 | 2521.1224 | 1988.3466
179 | 3.9634 | 856.0049 | 1989.9538 | 1892.4415 | 2031.588 | 2119.8732 | 2604.6657 | 1945.6804 | 1643.7728 | 1289.7463 | 1651.4877 | 1407.1878 | 1207.0148 | 2086.3442 | 1888.6244 | 2268.5683 | 1943.8535 | 2088.9803 | 1721.9655 | 1779.1632 | 1434.2075 | 1579.6509 | 2503.2026 | 2600.1977 | 1691.9538 | 2272.939 | 2223.237 | 2008.3097 | 2534.2586 | 1666.1829 | 1749.1416 | 1733.7876 | 1680.5585 | 1442.2026 | 1905.295 | 2098.8682 | 1587.2537 | 2954.8928 | 3109.2586 | 2716.1564 | 2769.7168 | 1767.0197 | 2260.0708 | 2050.1977 | 1779.3658 | 1303.6195 | 1485.2851 | 1448.6391 | 1769.7709 | 1655.061 | 2144.5123 | 1700.4267 | 1540.29 | 1763.1367 | 2320.6441 | 1692.3147 | 1554.29 | 1820.7069 | 2490.3609 | 1638.2242 | 1608.4828 | 1495.9538 | 1781.0098 | 1416.7562 | 1458.0413 | 1861.6096 | 2286.1219 | 1810.705 | 1814.7315 | 1596.3805 | 2200.1583 | 1800.6293 | 1737.5979 | 1952.7876 | 2039.8368 | 2069.0924 | 2451.8682 | 1688.7562 | 1694.6342 | 1577.4002 | 2066.8023 | 1447.4218 | 1561.0659 | 1581.6096 | 1630.6391 | 1363.6903 | 1433.939 | 1483.7562 | 1408.6657 | 1472.5634 | 1497.061 | 1458.4169 | 1644.1534 | 1602.8614 | 1865.5782 | 2090.8879 | 2614.6952 | 2980.9243 | 3721.3018 | 3744.1318 | 4078.4975 | 1358.9194 | 1335.8978 | 1283.5585 | 1358.9194 | 1398.5683 | 1576.4975 | 1452.1583 | 1512.173 | 1379.2438 | 1437.5536 | 1448.6293 | 1382.3097 | 1327.8584 | 1311.9902 | 1370.9852 | 1371.6607 | 1347.9902 | 1381.0462 | 1356.3707 | 1335.1121 | 1723.7512 | 1298.1219 | 1338.7315 | 1374.6145 | 1502.4218 | 1284.9951 | 1358.4975 | 1190.4975 | 1872.2537 | 2568.0659 | 2524.7001 | 1948.6342
180 | 3.9744 | 867.1991 | 2018.3928 | 1883.7905 | 2093.7974 | 2101.7935 | 2559.0059 | 1934.2036 | 1555.8083 | 1286.4032 | 1636.2085 | 1405.5988 | 1249.8019 | 2156.9936 | 1869.7915 | 2289.8029 | 1891.8024 | 2010.3869 | 1640.2045 | 1720.8029 | 1487.6003 | 1497.8073 | 2541.7994 | 2622.5978 | 1726.5959 | 2282.2026 | 2344.7955 | 1981.9906 | 2566.6047 | 1639.793 | 1803.205 | 1682.0069 | 1664.5973 | 1458.7886 | 1878.0094 | 2069.3943 | 1615.4091 | 2981.1897 | 3133.6008 | 2783.9926 | 2726.622 | 1807.5944 | 2295.6067 | 2082.2026 | 1750.5978 | 1291.6013 | 1405.9965 | 1483.1986 | 1815.9956 | 1648.1986 | 2174.5988 | 1675.8019 | 1490.4046 | 1783.5949 | 2359.5944 | 1689.5959 | 1515.21 | 1737.6151 | 2439.3933 | 1573.8019 | 1580.6087 | 1523.5924 | 1822.4076 | 1416.5919 | 1406.4022 | 1822.6008 | 2271.997 | 1856.1996 | 1788.211 | 1617.003 | 2126.8019 | 1802.3953 | 1831.8004 | 1901.6072 | 2090.3933 | 1985.6013 | 2439.8029 | 1706.6008 | 1718.7999 | 1632.7945 | 2048.8043 | 1406 | 1564.1971 | 1540.5998 | 1658.1952 | 1374.4017 | 1432.7974 | 1482.1912 | 1355.8024 | 1472.7989 | 1502.0044 | 1405.8009 | 1550.5983 | 1712.5904 | 1872.3829 | 2125.8004 | 2631.5983 | 2950.8024 | 3538.205 | 3749.5939 | 4069.3987 | 1311.4002 | 1371.588 | 1288.2075 | 1311.4002 | 1410.1981 | 1572.2011 | 1377.001 | 1478.4022 | 1366.2016 | 1425.2055 | 1460.2001 | 1380.6028 | 1301.0084 | 1297.8058 | 1334.001 | 1295.9975 | 1312.5954 | 1327.795 | 1342.5988 | 1301 | 1727.2001 | 1297.2036 | 1290.6003 | 1342.998 | 1470.4046 | 1270.3992 | 1352.4002 | 1183.7999 | 1885.4017 | 2584.6037 | 2563.4027 | 1981.0039
181 | 3.9854 | 854.4779 | 1983.5753 | 1787.4601 | 2066.6283 | 2042.8407 | 2596.5929 | 1944.3097 | 1628.3097 | 1315.6106 | 1676.9646 | 1388.9911 | 1261.177 | 2105.8496 | 1805.885 | 2304.7168 | 1928.6283 | 1929.062 | 1686.7699 | 1746.1239 | 1491.0619 | 1552.5398 | 2515.9734 | 2617.3186 | 1695.3363 | 2305.1239 | 2297.3805 | 1931.062 | 2585.6725 | 1578.7522 | 1821.4867 | 1728.2301 | 1650.646 | 1379.8053 | 1957.3982 | 2031.9204 | 1669.6725 | 2896.7257 | 3135.5398 | 2725.0619 | 2877.3893 | 1762.4248 | 2323.0265 | 2105.3363 | 1747.6549 | 1303.6991 | 1404.7611 | 1458.1947 | 1769.5929 | 1626.5929 | 2173.0708 | 1697.5841 | 1506.5486 | 1729.0442 | 2314.4248 | 1665.7699 | 1591.5398 | 1844.0708 | 2389.4602 | 1580.9292 | 1634.2212 | 1484.9381 | 1859.7876 | 1352.6991 | 1416.5664 | 1816.2566 | 2256.2832 | 1841.9557 | 1866.8053 | 1626.7699 | 2157.0797 | 1760.177 | 1835.6372 | 1953.8407 | 2057.6638 | 1985.5929 | 2473.4071 | 1691.5486 | 1724.1593 | 1595.2213 | 2094.6195 | 1416.1947 | 1534.2035 | 1536.8761 | 1618.1504 | 1397.7699 | 1404.9911 | 1413.5487 | 1373.7257 | 1467.3982 | 1518.6725 | 1420.646 | 1545.0708 | 1627.9911 | 1742.708 | 2144.7168 | 2603.7699 | 2965.5398 | 3549.4778 | 3704.6726 | 4071.3186 | 1306.8407 | 1284.6195 | 1325.6371 | 1306.8407 | 1405.8319 | 1573.1416 | 1397.531 | 1486.2301 | 1362.4336 | 1456.4071 | 1451.2301 | 1406.6106 | 1346.9115 | 1336.5044 | 1336.9026 | 1295.3275 | 1275.6372 | 1291.6726 | 1324.5044 | 1300.1504 | 1716.7433 | 1314.3186 | 1295.3363 | 1338.0443 | 1488.4601 | 1263.7876 | 1351.0265 | 1170.0442 | 1881.1593 | 2609.0531 | 2565.0354 | 2011.3009
182 | 3.9964 | 822.766 | 2070.1013 | 1726.4587 | 2037.383 | 2010.5634 | 2573.2783 | 1890.9523 | 1656.5379 | 1307.6553 | 1584.4587 | 1363.674 | 1222.7345 | 2109.8987 | 1814.114 | 2254.4528 | 1964.3421 | 2012.3387 | 1699.4085 | 1756.2468 | 1478.823 | 1552.6264 | 2442.4936 | 2645.4749 | 1700.8545 | 2304.4621 | 2245.1932 | 2013.44 | 2515.5728 | 1569.7404 | 1737.7753 | 1690.8766 | 1675.4051 | 1430.5506 | 1934.7251 | 2065.2247 | 1597.7753 | 2898.9651 | 3124.617 | 2716.4179 | 2765.114 | 1777.2562 | 2257.6008 | 2100.7438 | 1788.5192 | 1308.2026 | 1504.8987 | 1422.0536 | 1725.383 | 1604.6264 | 2180.3072 | 1722.2247 | 1451.4902 | 1698.7498 | 2316.2247 | 1697.057 | 1577.6962 | 1794.6613 | 2434.7591 | 1556.8417 | 1573.7974 | 1547.0826 | 1766.1804 | 1357.6519 | 1377.0919 | 1782.4808 | 2271.7217 | 1801.8987 | 1819.114 | 1571.7532 | 2203.5634 | 1759.0315 | 1830.8102 | 1928.7532 | 2127.6647 | 1957.1106 | 2498.5192 | 1635.3668 | 1733.9685 | 1618.6298 | 2128.383 | 1444.114 | 1493.2723 | 1516.1932 | 1619.4936 | 1429.9557 | 1363.0477 | 1417.383 | 1340.0162 | 1461.1268 | 1464.0757 | 1447.6519 | 1589.6042 | 1592.114 | 1801.2876 | 2202.5506 | 2579.3098 | 2937.3166 | 3481.5659 | 3729.4145 | 4103.7438 | 1288.2689 | 1306.6996 | 1248.886 | 1288.2689 | 1432.5192 | 1545.3481 | 1432.766 | 1443.0349 | 1310.886 | 1405.1013 | 1422.8638 | 1398.88 | 1266.3481 | 1310.8417 | 1309.4621 | 1351.9464 | 1269.2375 | 1303.8323 | 1290.5285 | 1292.5064 | 1662.8672 | 1265.383 | 1300.4051 | 1365.7438 | 1428.3574 | 1266.3702 | 1347.4145 | 1123.3481 | 1821.2911 | 2581.6077 | 2509.4238 | 1989.7217
183 | 4.0074 | 789.9243 | 2039.7482 | 1783.7729 | 2034.1141 | 1988.6637 | 2610.8338 | 1932.9666 | 1657.0894 | 1232.3412 | 1619.0236 | 1362.4454 | 1156.5359 | 2114.8958 | 1801.5644 | 2208.3275 | 1864.8437 | 1988.2419 | 1636.9577 | 1728.2173 | 1433.1898 | 1591.7532 | 2473.4032 | 2589.2655 | 1707.062 | 2243.4827 | 2205.5025 | 1960.412 | 2524.5546 | 1582.3363 | 1727.7581 | 1680.8387 | 1659.2085 | 1410.9901 | 1839.0669 | 2032.35 | 1587.5211 | 2942.6687 | 3161.7729 | 2691.2468 | 2738.8201 | 1811.1514 | 2343.8239 | 2045.7581 | 1745.7915 | 1291.8102 | 1496.1514 | 1473.763 | 1726.3796 | 1639.0137 | 2112.4268 | 1718.2606 | 1511.8289 | 1739.8574 | 2317 | 1646.8239 | 1559.5113 | 1832.9764 | 2511.9616 | 1553.3028 | 1579.0521 | 1446.1986 | 1750.3933 | 1338.7817 | 1339.7915 | 1828.0521 | 2211.1799 | 1807.3697 | 1785 | 1555.469 | 2175.9764 | 1780.7768 | 1815.9528 | 1907.7296 | 2080.3884 | 1999.3461 | 2443.943 | 1711.2606 | 1695.8722 | 1606.0521 | 2095.2556 | 1405.7345 | 1462.2232 | 1476.9764 | 1631.6824 | 1401.0708 | 1310.3697 | 1396.5782 | 1259.9056 | 1412.588 | 1506.6489 | 1437.7866 | 1609.4306 | 1566.057 | 1878.7296 | 2200.5261 | 2683.8722 | 2876.6204 | 3622.3933 | 3749.3127 | 4054.2085 | 1308.6873 | 1256.4789 | 1267.3599 | 1308.6873 | 1391.6873 | 1528.9095 | 1389.4975 | 1404.7817 | 1349.7394 | 1411.2419 | 1436.7296 | 1335.7158 | 1277.6352 | 1320.8102 | 1273.8387 | 1309.0049 | 1249.9007 | 1273.6775 | 1284.8437 | 1266.0993 | 1637.588 | 1268.0384 | 1276.8623 | 1309.0993 | 1449.9095 | 1276.8151 | 1368.9715 | 1121.3648 | 1856.2704 | 2566.0806 | 2564.706 | 1932.0236
184 | 4.0184 | 809.7478 | 1966.7065 | 1821.8068 | 2066.8068 | 2015.5221 | 2573.9646 | 1880.7802 | 1579.205 | 1219.2729 | 1596.0206 | 1364.7095 | 1195.5516 | 2109.5339 | 1845.3259 | 2272.0147 | 1893.4425 | 2039.1873 | 1687.0324 | 1639.1726 | 1478.2139 | 1530.7566 | 2490.5958 | 2575.5811 | 1749.7861 | 2161.941 | 2231.9646 | 1931.677 | 2517.3053 | 1558.469 | 1705.7271 | 1654.5192 | 1648.7448 | 1343.2787 | 1927.4867 | 1939.4543 | 1562.733 | 2986.3082 | 3217.0059 | 2724.8245 | 2760.5103 | 1763.7773 | 2267.1608 | 2016.2493 | 1762.736 | 1262.7389 | 1455.5074 | 1473 | 1769.8156 | 1572.4779 | 2127.8038 | 1703.2493 | 1442.4749 | 1692.6977 | 2316.0029 | 1563.4307 | 1485.9558 | 1830.056 | 2421.5575 | 1484.2286 | 1583.4779 | 1386.7183 | 1749.4779 | 1305.9617 | 1369.9469 | 1812.2876 | 2184.5044 | 1792.5663 | 1769.795 | 1522.267 | 2187.0147 | 1713.9764 | 1847.2522 | 1877.7124 | 2038.0147 | 1998.5722 | 2439.6859 | 1690.7655 | 1701.0501 | 1612.9705 | 2050.5015 | 1377.2935 | 1512.8215 | 1499.2316 | 1616.5133 | 1356.9882 | 1318.4985 | 1376.5457 | 1325.7448 | 1462.2611 | 1444.2787 | 1411.7773 | 1523.9617 | 1579.7861 | 1816.059 | 2176.8156 | 2681.5722 | 2856.6151 | 3634.6888 | 3726.0147 | 4039.7566 | 1305.5634 | 1286.2522 | 1226.7213 | 1305.5634 | 1377.8451 | 1455.5162 | 1373.5044 | 1379.1903 | 1363.7537 | 1441.4985 | 1382.2846 | 1360.5192 | 1284.792 | 1276.7832 | 1269.4543 | 1351.4602 | 1277.5221 | 1240.4808 | 1293.7507 | 1241.9676 | 1682.5251 | 1329.2787 | 1257.4513 | 1280.9794 | 1396.4572 | 1275.9912 | 1356.9971 | 1116.2021 | 1881.0531 | 2568.5221 | 2544.4808 | 1943.4189
185 | 4.0294 | 817.2311 | 1973.8593 | 1757.1996 | 2015.0157 | 1996.9233 | 2586.9833 | 1845.292 | 1621.3382 | 1216.4012 | 1562.707 | 1381.1524 | 1155.0315 | 2083.3697 | 1765.3845 | 2236.0757 | 1947.8151 | 2070.8289 | 1628.8908 | 1702.5683 | 1505.3225 | 1542.8319 | 2410.001 | 2503.4622 | 1721.3697 | 2210.7532 | 2240.7984 | 1985.5988 | 2485.5861 | 1590.2616 | 1717.2301 | 1627.6765 | 1645.1072 | 1325.7699 | 1904.9056 | 1970.3068 | 1578.1691 | 2947.9243 | 3210.7384 | 2650.5693 | 2752.0924 | 1739.1386 | 2326.8751 | 2024.3235 | 1755.291 | 1268.7689 | 1420.0137 | 1438.998 | 1713.4622 | 1588.4612 | 2068.0305 | 1690.9223 | 1449.6912 | 1737.3687 | 2304.6303 | 1633.7384 | 1515.0452 | 1762.7994 | 2380.001 | 1498.938 | 1590.8299 | 1416.646 | 1785.0777 | 1334.7994 | 1394.9675 | 1776.8771 | 2182.354 | 1731.5388 | 1722.4307 | 1499.5074 | 2161.0147 | 1738.7847 | 1852.5851 | 1906.8456 | 2034.7699 | 1935.9853 | 2486.6912 | 1671.8614 | 1666.7395 | 1636.1996 | 2071.7552 | 1341.3697 | 1435.0305 | 1499.6912 | 1625.4936 | 1371.0924 | 1321.8309 | 1339.1544 | 1325.5074 | 1448.646 | 1418.4002 | 1387.4002 | 1561.6922 | 1551.8928 | 1759.9233 | 2129.6165 | 2603.0305 | 2983.5378 | 3684.9685 | 3728.5241 | 4077.4642 | 1243.8771 | 1274.8456 | 1248.938 | 1243.8771 | 1295.647 | 1451.8161 | 1357.2301 | 1429.7994 | 1354.5074 | 1418.6755 | 1368.1859 | 1350.939 | 1249.8466 | 1245.6155 | 1310.7689 | 1388.8614 | 1241.9223 | 1258.9538 | 1278.8761 | 1264.2763 | 1688.3865 | 1293.9843 | 1290.7837 | 1293.8918 | 1435.0305 | 1271.6607 | 1353.4612 | 1184.355 | 1830.9233 | 2553.8466 | 2562.1691 | 2028.0924
186 | 4.0404 | 824.6062 | 1898.5753 | 1735.5354 | 2048.7045 | 1998.1416 | 2510.2247 | 1805.1632 | 1612.6431 | 1276.1908 | 1557.8491 | 1311.8215 | 1197.763 | 2114.3938 | 1776.8953 | 2190.4184 | 1920.9385 | 1981.3692 | 1579.6338 | 1654.0708 | 1463.2124 | 1586.6893 | 2434.4493 | 2517.9385 | 1722.5447 | 2184.2031 | 2182.0738 | 1920.3323 | 2544.5939 | 1603.323 | 1677.4277 | 1623.1569 | 1620 | 1355.44 | 1817.2399 | 1940.5138 | 1582.5846 | 2959.6248 | 3175.7694 | 2645.4154 | 2742.117 | 1749.9169 | 2266.3845 | 2026.6524 | 1686.88 | 1238.3908 | 1348.3938 | 1356.4184 | 1729.5816 | 1562.3323 | 2058.3815 | 1666.0492 | 1410.323 | 1708.0708 | 2275.5231 | 1630.5354 | 1494.0984 | 1750.7045 | 2417.32 | 1482.56 | 1563.3353 | 1407.5354 | 1779.554 | 1281.6278 | 1336.003 | 1768.8431 | 2168.7355 | 1746.7139 | 1726.4892 | 1473.5939 | 2130.2738 | 1728.677 | 1860.0339 | 1873.88 | 2039.5201 | 1946.4616 | 2454.7168 | 1652.0462 | 1710.7139 | 1612.3446 | 2125.0462 | 1361.9877 | 1413.88 | 1463.0615 | 1667.9139 | 1376.1785 | 1328.3692 | 1348.0216 | 1304.5231 | 1439.8092 | 1418.1416 | 1385.7724 | 1574.1569 | 1587.68 | 1791.8122 | 2177.6062 | 2566.2709 | 2945.7847 | 3633.2001 | 3774.1662 | 4163.5663 | 1265.1446 | 1238.5939 | 1220.5108 | 1265.1446 | 1332.6185 | 1491.8338 | 1343.7168 | 1393.6062 | 1325.3078 | 1366.1569 | 1424.9877 | 1381.7384 | 1243.2463 | 1245.3323 | 1275.1878 | 1334.8156 | 1215.9538 | 1248.88 | 1230.7847 | 1246.0492 | 1740.0309 | 1265.2247 | 1234.5939 | 1266.8923 | 1426.7507 | 1237.2861 | 1344.5846 | 1206.6863 | 1827.7601 | 2583.12 | 2566.0369 | 2029.0708
187 | 4.0514 | 810.7227 | 1958.4012 | 1729.2861 | 2067.5339 | 1990.3127 | 2496.1682 | 1788.0797 | 1625.2979 | 1271.6991 | 1592.6489 | 1306.8466 | 1199.2537 | 2127.705 | 1821.5634 | 2232.8673 | 1883.2891 | 1983.8555 | 1647.3009 | 1635.8496 | 1423.4277 | 1541.1121 | 2376.8643 | 2520.9204 | 1658.1357 | 2214.882 | 2217.9823 | 1926.6107 | 2544.1829 | 1622.8377 | 1656.8525 | 1650.5487 | 1631.1593 | 1360.8791 | 1847.2743 | 1976.1386 | 1614.5811 | 2879.2095 | 3096.1593 | 2642.0325 | 2725.7522 | 1747.2743 | 2329.5575 | 1976.8555 | 1673.7227 | 1200.7198 | 1371.7198 | 1396.5782 | 1692.2832 | 1544.003 | 2068.5693 | 1679.6873 | 1446.1475 | 1682.41 | 2235.5339 | 1603.6843 | 1566.0118 | 1789.5634 | 2395.295 | 1505.1623 | 1591.6962 | 1386.6932 | 1674.7257 | 1228.4454 | 1380.9764 | 1776.9322 | 2121.8673 | 1754.3098 | 1750.1445 | 1458.59 | 2125.8496 | 1692.5664 | 1833.8732 | 1889.0502 | 1998.6225 | 1910.5988 | 2504.997 | 1618.4218 | 1700.5693 | 1593.1327 | 2089.1327 | 1362.2979 | 1407.8761 | 1478.1062 | 1604.2743 | 1351.6814 | 1332.8584 | 1310.4395 | 1307.4543 | 1440.0089 | 1391.1416 | 1350.5634 | 1586.0973 | 1604.5546 | 1822.9646 | 2173.7375 | 2561.7965 | 2935.1593 | 3539.9321 | 3720.5073 | 4070.1593 | 1291.4159 | 1228.7404 | 1191.1327 | 1291.4159 | 1322.4307 | 1487.6932 | 1377.4012 | 1384.0148 | 1309.8879 | 1403.2773 | 1386.3834 | 1358.6696 | 1166.1416 | 1235.587 | 1228.2979 | 1238.177 | 1247.2891 | 1211.115 | 1217.2979 | 1257.3982 | 1664.6047 | 1236.003 | 1216.7286 | 1266.0148 | 1425.9764 | 1266.3157 | 1355.1209 | 1135.3068 | 1824.6225 | 2607.4366 | 2553.3746 | 1986.8141
188 | 4.0624 | 818.4346 | 1913.1937 | 1756.8584 | 2024.5915 | 2030.9267 | 2521.9897 | 1900.618 | 1654.6544 | 1237.0472 | 1488.8638 | 1297.9528 | 1148.5811 | 2097.6755 | 1795.4661 | 2215.9793 | 1911.0836 | 1971.7488 | 1657.7802 | 1628.9685 | 1426.178 | 1505.4135 | 2403.6229 | 2603.2306 | 1654.6544 | 2261.7748 | 2201.7276 | 1949.4296 | 2609.3142 | 1590 | 1647.8063 | 1619.1623 | 1665.267 | 1388.7488 | 1830.3874 | 1953.1834 | 1588.6234 | 2938.2831 | 3123.7016 | 2677.9007 | 2780.1519 | 1733.3717 | 2283.0211 | 1972.8063 | 1699.6387 | 1242.2301 | 1364.2881 | 1399.644 | 1680.2409 | 1541.8377 | 2054.2566 | 1653.9946 | 1473.6962 | 1655.4764 | 2197.3663 | 1573.8899 | 1560.3353 | 1761.6755 | 2416.0733 | 1515.9164 | 1548.7645 | 1374.5757 | 1702.6544 | 1287.1411 | 1344.2827 | 1868.4872 | 2142.8796 | 1788.733 | 1738.6755 | 1500.7434 | 2115.9843 | 1665.1991 | 1857.4189 | 1935.8013 | 2061.7227 | 1976.3978 | 2499.1047 | 1626.9946 | 1677.9007 | 1574.2094 | 2082.7434 | 1357.0368 | 1444.178 | 1421.0733 | 1595.7434 | 1296.2984 | 1312.0211 | 1333.7748 | 1323.2463 | 1450.2252 | 1383.6283 | 1370.1254 | 1523.9214 | 1552.1258 | 1768.9685 | 2181.0683 | 2677.7227 | 2941.8117 | 3472.7483 | 3639.4553 | 4086.0629 | 1268.5079 | 1259.2724 | 1188.0261 | 1268.5079 | 1312.5708 | 1462.7434 | 1328.0157 | 1425.1151 | 1339.3195 | 1390.1937 | 1329.9946 | 1289.8638 | 1174.4449 | 1240.8117 | 1240.3456 | 1271.2306 | 1235.3928 | 1159.2252 | 1212.9321 | 1237.9317 | 1703.6912 | 1260.3978 | 1250.5497 | 1292.1937 | 1387.4921 | 1288.0526 | 1337.9946 | 1165.9267 | 1904.1362 | 2587.5762 | 2516.9106 | 1932.8009
189 | 4.0734 | 819.4681 | 1932.1573 | 1784.5781 | 2046.2782 | 2015.5821 | 2485.4179 | 1841.7856 | 1661.0177 | 1218.4857 | 1533.5861 | 1312.7856 | 1153.8928 | 2068.0393 | 1757.6106 | 2149.6716 | 1955.8535 | 1946.8928 | 1612.9892 | 1626.2212 | 1443.4572 | 1519.1425 | 2487.0935 | 2522.6283 | 1666.6106 | 2238.821 | 2230.8928 | 1837.9892 | 2529.6705 | 1552.7138 | 1661.7788 | 1624.1927 | 1648.6175 | 1351.1534 | 1833.825 | 1950.179 | 1509.5074 | 2845.8683 | 3158.1111 | 2629.5821 | 2774.0895 | 1739.1642 | 2241.4287 | 1981.4966 | 1722.3392 | 1295.3176 | 1328.7286 | 1380 | 1638.0747 | 1560.7326 | 2038.0108 | 1677.3854 | 1497.0069 | 1652.9498 | 2230.2782 | 1635.118 | 1517.4002 | 1713.3962 | 2378.2851 | 1475.4966 | 1536.4681 | 1399.5496 | 1700.939 | 1350.232 | 1343.9213 | 1785 | 2152 | 1769.0285 | 1737.6254 | 1533.4572 | 2135.118 | 1625.0393 | 1823.5211 | 1864.0217 | 2010.4179 | 1987.5103 | 2474.939 | 1656.9105 | 1655.6824 | 1578.0787 | 2098.057 | 1319.3039 | 1444.3677 | 1425.232 | 1627.2144 | 1279.5998 | 1264.8358 | 1371.7218 | 1250.8535 | 1439.3107 | 1381.2537 | 1424.6961 | 1560.2429 | 1521.4789 | 1756.589 | 2127.4396 | 2654.3825 | 2905.9823 | 3543.6745 | 3680.6961 | 4108.3363 | 1269.3501 | 1223.5821 | 1207.3284 | 1269.3501 | 1311.1397 | 1481.1642 | 1330.3215 | 1431.9282 | 1292.8358 | 1388.0285 | 1363.6391 | 1285.8033 | 1191.5428 | 1194.1072 | 1216.5821 | 1228.5359 | 1203.0108 | 1156.3894 | 1162.0787 | 1297.1534 | 1683.1967 | 1277.1032 | 1261.5034 | 1270.4533 | 1393.4857 | 1228.0787 | 1368.5319 | 1148.0964 | 1884.8928 | 2498.7463 | 2635.7463 | 1944.3746
190 | 4.0844 | 860.1991 | 1949.2655 | 1732.1637 | 1988.9381 | 1999.9248 | 2489.4469 | 1840.5974 | 1626.4469 | 1193.6371 | 1573.4203 | 1311.5974 | 1149.9071 | 2056.0044 | 1750.3053 | 2159.0044 | 1897.6858 | 1950.6593 | 1594.292 | 1600.6283 | 1424.792 | 1482.4115 | 2556.2699 | 2498.8407 | 1657.6903 | 2172.0443 | 2244.3496 | 1853.854 | 2435.4779 | 1495.8407 | 1675.4203 | 1577.9734 | 1608.0708 | 1323.1062 | 1818.0177 | 2006.2964 | 1552.0929 | 2907.7566 | 3192.3938 | 2660.4381 | 2771.3008 | 1697.1902 | 2258.0177 | 2001.1858 | 1702.6327 | 1233.0487 | 1322.4646 | 1375.4779 | 1628.6062 | 1629.9867 | 2043.1416 | 1632.4292 | 1435.6637 | 1625.9381 | 2197.1637 | 1662.323 | 1547.6018 | 1697.1505 | 2318.3894 | 1511.0133 | 1553.9425 | 1342.0133 | 1672.3496 | 1318.7876 | 1329.2301 | 1777.8938 | 2149.4159 | 1738.2964 | 1809.031 | 1531.9115 | 2141.9734 | 1652.7345 | 1777.3363 | 1897.8628 | 2021.5531 | 1893.5575 | 2448.2876 | 1627.0752 | 1698.5708 | 1587.9248 | 2073.0929 | 1355.1725 | 1409.4336 | 1409.615 | 1626.1416 | 1265.2257 | 1277.4159 | 1380.9336 | 1247.5974 | 1423.5265 | 1373.6371 | 1380.1062 | 1549.7566 | 1568.9557 | 1721.9027 | 2154.9734 | 2686.2079 | 2942.1681 | 3466.2124 | 3597.3451 | 4194.4823 | 1263.4956 | 1199.2035 | 1178.292 | 1263.4956 | 1342.146 | 1431.761 | 1340.9867 | 1359.4469 | 1312.5221 | 1385.7212 | 1330.7699 | 1253.2743 | 1190.1814 | 1207.7832 | 1194.1416 | 1249.7566 | 1223.646 | 1192.4425 | 1162.2345 | 1253.9248 | 1724.1239 | 1204.9248 | 1222.4336 | 1240.2257 | 1399 | 1239.2168 | 1335.8761 | 1119.1725 | 1875.4159 | 2503.4558 | 2608.4779 | 1895.2567
191 | 4.0954 | 852.5703 | 1922.0708 | 1781.6352 | 2009.4297 | 1982.6647 | 2520.2242 | 1834.4759 | 1624.0295 | 1173.9243 | 1535.7355 | 1340.0118 | 1136.8879 | 2017.7001 | 1748.1819 | 2117.6824 | 1822.0767 | 1944.8348 | 1615.1062 | 1625.3353 | 1383.5241 | 1504.235 | 2475.3471 | 2522.4877 | 1658.5526 | 2198.7119 | 2260.176 | 1913.2409 | 2610.1829 | 1524.5172 | 1633.0413 | 1568.1888 | 1657.5998 | 1351.412 | 1814.4474 | 1984.4474 | 1553.7119 | 2892.4061 | 3088.2586 | 2709.8171 | 2675.941 | 1670.1711 | 2244.3884 | 1934.0108 | 1687.3589 | 1205.6529 | 1319.4946 | 1353.2586 | 1698.5408 | 1610.6765 | 2023.7355 | 1635.5939 | 1408.2999 | 1610.5526 | 2241.941 | 1635.5054 | 1503.176 | 1693.1111 | 2407.5703 | 1537.3648 | 1553.9007 | 1366.0177 | 1720.7001 | 1289 | 1314.2055 | 1781.7768 | 2176.5241 | 1687.9469 | 1799.9587 | 1486.9587 | 2072.5054 | 1680.9351 | 1799 | 1869.9115 | 2010.2232 | 1880.8063 | 2504.3884 | 1633.941 | 1664.5585 | 1556.2468 | 2080.0767 | 1326.6942 | 1402.7945 | 1423.059 | 1606.0649 | 1293.1996 | 1263.2822 | 1290.4179 | 1277.353 | 1398.0531 | 1326.0944 | 1371.529 | 1535.0767 | 1558.3825 | 1790.4415 | 2110.3176 | 2604.3176 | 2935.2291 | 3505.0531 | 3522.9528 | 4162.7945 | 1249.7768 | 1168.3589 | 1176.6411 | 1249.7768 | 1271.6942 | 1368.3825 | 1298.5408 | 1364.4061 | 1287.1288 | 1375.8348 | 1328.0236 | 1243.4582 | 1195.176 | 1211.941 | 1156.294 | 1247.1475 | 1208.9056 | 1231.6824 | 1143.4415 | 1228.9528 | 1696.5526 | 1201.412 | 1205.4002 | 1268.1996 | 1383.2409 | 1211.9233 | 1322.5408 | 1107.1239 | 1856.2527 | 2553.6234 | 2608.0531 | 1917.176
192 | 4.1064 | 832.1657 | 1934.1819 | 1759.1819 | 2038.1765 | 1988.2566 | 2558.8077 | 1760.2674 | 1629.1337 | 1233.1927 | 1538.1445 | 1311.674 | 1109.4277 | 2056.4597 | 1698.738 | 2097.9734 | 1905.7002 | 1904.6416 | 1641.2674 | 1646.1603 | 1424.7114 | 1480.385 | 2453.2355 | 2478.6844 | 1628.6898 | 2192.7969 | 2217.4277 | 1898.6844 | 2645.9307 | 1497.9838 | 1662.9892 | 1618.4439 | 1639.5187 | 1359.2301 | 1806.4228 | 2019.1765 | 1556.4867 | 2944.8073 | 3085.2458 | 2666.4223 | 2692.4971 | 1718.615 | 2242.4921 | 1856.2832 | 1691.8289 | 1204.1283 | 1359.4012 | 1322.0909 | 1676.1603 | 1592.7969 | 2041.0909 | 1600.9199 | 1390.0482 | 1589.3796 | 2204.6204 | 1568.1391 | 1460.4277 | 1652.4115 | 2345.4547 | 1540.353 | 1601.3692 | 1366.4385 | 1708.7114 | 1289.6952 | 1284.9037 | 1810.3102 | 2222.2301 | 1705.8019 | 1749.1553 | 1463.615 | 2047.5453 | 1673.9464 | 1776.7542 | 1851.4705 | 1951.7969 | 1944.3368 | 2479.5511 | 1624.7753 | 1643.9892 | 1507.8825 | 2077.4975 | 1338.4331 | 1437.3102 | 1454.4705 | 1608.5349 | 1255.9946 | 1238.7006 | 1311.353 | 1276.7006 | 1395.8397 | 1367.385 | 1338.2085 | 1572.4705 | 1592.3796 | 1795.3692 | 2103.6096 | 2671.299 | 2926.4169 | 3501.4493 | 3528.5507 | 4202.1765 | 1253.6312 | 1189.8609 | 1157.5723 | 1253.6312 | 1300.4651 | 1428.4489 | 1250.7861 | 1377.1819 | 1245.8073 | 1305.0536 | 1336.6952 | 1187.9464 | 1159.032 | 1197.2139 | 1163.1603 | 1276.8397 | 1148.8451 | 1193.2035 | 1175.1337 | 1217.5187 | 1668.7753 | 1234.6042 | 1216.8771 | 1236.5561 | 1349.5669 | 1197.8181 | 1278.6096 | 1121.1927 | 1835.9518 | 2510.8451 | 2592.9788 | 1882.4223
193 | 4.1174 | 835.2065 | 1890.7669 | 1754.6342 | 2025.9616 | 1987.4041 | 2485.4012 | 1750.6372 | 1607.4041 | 1179.9204 | 1503.6785 | 1275.1121 | 1140.5162 | 2080.2891 | 1724.6726 | 2088.7876 | 1980.1711 | 1947.3156 | 1621.2006 | 1639.3746 | 1375.2419 | 1491.1622 | 2382.8466 | 2494.8849 | 1639.6018 | 2168.4749 | 2215.0472 | 1902.6489 | 2557.7728 | 1573.9499 | 1680.0442 | 1569.6076 | 1615.6106 | 1300.9675 | 1755.3481 | 1968.0944 | 1519.6844 | 2955.6135 | 3141.0354 | 2696.6814 | 2736.8879 | 1724.3924 | 2193.9233 | 1929.0354 | 1676.0826 | 1203.882 | 1303.2094 | 1303.4425 | 1713.6401 | 1542.5634 | 2058.7906 | 1597.8348 | 1383.4779 | 1620.2006 | 2170.2006 | 1529.6047 | 1458.0472 | 1699.7256 | 2297.5575 | 1521.6814 | 1518.9675 | 1370.7552 | 1677.5959 | 1274.0855 | 1301.2832 | 1765.2094 | 2157.4896 | 1751.4071 | 1701.9528 | 1454.9971 | 2094.0029 | 1684.0059 | 1764.5104 | 1901.4307 | 1944.0295 | 1963.1239 | 2389.0029 | 1626.646 | 1661.0442 | 1515.5458 | 2043.7935 | 1316.7286 | 1400.4867 | 1443.9705 | 1554.3658 | 1264.1622 | 1225 | 1281.8849 | 1251.1239 | 1387.5104 | 1397.236 | 1353.6431 | 1561.6106 | 1599.4484 | 1729.5221 | 2124.3924 | 2763.0944 | 2967.7227 | 3499.8673 | 3577.6372 | 4189.2419 | 1191 | 1173.6932 | 1184.4248 | 1191 | 1325.8436 | 1456.7374 | 1248.9174 | 1369.3953 | 1276.4071 | 1283.4779 | 1304.5309 | 1227.8761 | 1181.2832 | 1190.4454 | 1173.6489 | 1251.236 | 1125.5104 | 1130.0738 | 1158.4425 | 1200.4484 | 1692.9587 | 1251.3186 | 1211.8348 | 1242.2065 | 1371.6313 | 1195.528 | 1262.6844 | 1070.4395 | 1858.7168 | 2544.7316 | 2529.2006 | 1936.0738
194 | 4.1284 | 811.1652 | 1860.4435 | 1756.2261 | 2048.2626 | 1935.2261 | 2456.2974 | 1771.2104 | 1589.0939 | 1170.9287 | 1499.7183 | 1265.1514 | 1152.0678 | 2017.2296 | 1744.9026 | 2148.1844 | 1914.0452 | 1970.6278 | 1632.5148 | 1555.5339 | 1359.4174 | 1480.1165 | 2347.7443 | 2471.3496 | 1640.793 | 2198.9774 | 2184.9322 | 1851.767 | 2494.6991 | 1595.8766 | 1631.6922 | 1550.7026 | 1574.0035 | 1310.6156 | 1776.8279 | 1900.4209 | 1497.4887 | 2918.8452 | 3115.0122 | 2654.3879 | 2733.0782 | 1747.7635 | 2159.2522 | 1936.5078 | 1659.767 | 1183.3235 | 1260.1878 | 1265.8087 | 1693.647 | 1531.4435 | 2053.0575 | 1592.5044 | 1371.1809 | 1614.3948 | 2187.8417 | 1541.9617 | 1415.2783 | 1701.3791 | 2332.5861 | 1499.1131 | 1490.2817 | 1364.9931 | 1701.4887 | 1232.8348 | 1267.6244 | 1740.4243 | 2128.8191 | 1703.3235 | 1713.4592 | 1481.0487 | 2061.9965 | 1646.4174 | 1814.767 | 1906.2227 | 1976.8087 | 1934.9287 | 2436.6417 | 1583.9513 | 1627.9513 | 1589.2783 | 2086.4435 | 1328.8034 | 1377.3982 | 1387.3722 | 1545.3304 | 1244.5565 | 1227.6052 | 1265.0487 | 1228.1391 | 1414.32 | 1384.76 | 1322.4852 | 1512.1878 | 1550.2783 | 1724.6991 | 2166 | 2689.4661 | 2936.2522 | 3552.6244 | 3600.0713 | 4226.266 | 1189.8835 | 1121.7252 | 1236.4818 | 1189.8835 | 1263.9444 | 1379.12 | 1290.5408 | 1361.5044 | 1235.3948 | 1296.1165 | 1243.1391 | 1278.0295 | 1170.6991 | 1166.5826 | 1146.5861 | 1252.1583 | 1147.854 | 1151.3461 | 1137.5565 | 1167.6539 | 1667.9931 | 1250.44 | 1214.32 | 1232.3078 | 1403.7409 | 1172.4695 | 1261.7026 | 1072.2557 | 1845.854 | 2638.7079 | 2534.9322 | 1974.8383
195 | 4.1394 | 839.9961 | 1859.9282 | 1718.3029 | 2051.7885 | 1925.7532 | 2403.3904 | 1767.6971 | 1575.6234 | 1173.6096 | 1487.8427 | 1225.2115 | 1105.4739 | 2015.9145 | 1687.9321 | 2094.2468 | 1932.8348 | 1965.3864 | 1636.2389 | 1550.3068 | 1381.5261 | 1472.7729 | 2358.9961 | 2477.3845 | 1673.9941 | 2216.3029 | 2203.8446 | 1862.6096 | 2530.9902 | 1537.6932 | 1588.6155 | 1602.4563 | 1615.6096 | 1383.3127 | 1730.1436 | 1906.0816 | 1501.1534 | 2922.5614 | 3043.9941 | 2603.6893 | 2774.7787 | 1691.2409 | 2132.1534 | 1912.3825 | 1677.53 | 1165.8466 | 1273.6155 | 1244.2232 | 1680.6077 | 1527.8525 | 1956.2311 | 1562.9125 | 1376.3668 | 1602.3107 | 2229.5359 | 1586.7768 | 1441.6795 | 1741.3904 | 2306.7168 | 1462.8505 | 1509.0777 | 1347.1397 | 1684.3923 | 1221.6893 | 1242.3068 | 1772.3127 | 2150.5457 | 1685.8466 | 1711.9823 | 1477.6234 | 2032.3087 | 1661.2213 | 1817.5359 | 1830.9223 | 1944.0738 | 1892.6273 | 2535.5418 | 1561.233 | 1627.9164 | 1547.6293 | 2102.4602 | 1364.9478 | 1384.6214 | 1388.3845 | 1596.4563 | 1225.0796 | 1228.9243 | 1273.9262 | 1215.7729 | 1362.1573 | 1353.528 | 1307.9961 | 1539.0718 | 1547.4602 | 1736 | 2166 | 2675.0855 | 2894.5437 | 3518.0796 | 3587.0098 | 4249.8682 | 1184.5398 | 1165.765 | 1162.3166 | 1184.5398 | 1202.6116 | 1385.3166 | 1276.6391 | 1297.3107 | 1243.6873 | 1290.3107 | 1254.6096 | 1291.2448 | 1156.2409 | 1159.468 | 1189.5359 | 1203.5437 | 1097.3845 | 1118.0816 | 1122.3087 | 1173.6214 | 1634.7611 | 1215.2389 | 1184.8407 | 1253.8545 | 1368.0039 | 1182.7787 | 1300.3845 | 1068.3087 | 1813.8388 | 2564.6814 | 2509.2468 | 1976.8486
196 | 4.1504 | 863.1416 | 1849.3156 | 1728.4779 | 1968.2389 | 2016.3924 | 2399.441 | 1735.7006 | 1567.5988 | 1189.8584 | 1482.1622 | 1240.5191 | 1065.7124 | 2051 | 1700.9912 | 2065.1048 | 1920.6769 | 1946.4941 | 1578.4012 | 1570.6637 | 1413.7492 | 1472.2832 | 2401.1859 | 2459.3525 | 1686.0973 | 2224.4941 | 2198.0973 | 1871.7168 | 2575.7611 | 1540.3481 | 1632.8334 | 1577.1828 | 1586.6283 | 1336.8141 | 1767.2389 | 1917.5502 | 1493.4779 | 2831.7773 | 3102.9145 | 2611.6357 | 2739.9676 | 1652.3805 | 2137.5708 | 1930.9189 | 1698.9071 | 1153.2876 | 1279.7448 | 1271.4292 | 1682.6401 | 1479.7817 | 1972.4572 | 1586.9557 | 1387.8997 | 1574.5265 | 2204.292 | 1543.3038 | 1479.7168 | 1688.6401 | 2247.0723 | 1433.0973 | 1510.3643 | 1383.972 | 1634.3525 | 1233.6032 | 1270.5988 | 1763.1092 | 2136.4735 | 1693.1254 | 1751.0693 | 1438.6519 | 2055.7729 | 1671.8304 | 1827.6032 | 1825.2714 | 1941.1298 | 1849.9513 | 2507.7168 | 1566.68 | 1628.0206 | 1479.8908 | 2126.6475 | 1279.0767 | 1363.5708 | 1367.1784 | 1592.6076 | 1231.854 | 1218.8422 | 1297.8983 | 1192.6681 | 1350.3805 | 1386.8746 | 1320.826 | 1542.146 | 1558.1578 | 1772.5015 | 2151.32 | 2705.0679 | 2909.1048 | 3471.292 | 3562.4735 | 4230.6225 | 1192.0929 | 1155.6239 | 1114.4941 | 1192.0929 | 1218.3968 | 1349.7935 | 1182.3805 | 1287.7773 | 1231.0693 | 1283.1578 | 1279.1903 | 1233.6032 | 1134.8378 | 1133 | 1188.0973 | 1181.2065 | 1099.9838 | 1106.1416 | 1109.6681 | 1142.6519 | 1641.8141 | 1176.4454 | 1193.1136 | 1209.6844 | 1340.4941 | 1164.2345 | 1293.0649 | 1050.1136 | 1864.2345 | 2548.6932 | 2462.2508 | 1949.9395
197 | 4.1614 | 867.4553 | 1897.3147 | 1708.6824 | 1955.2733 | 2017.7768 | 2442.4985 | 1727.6352 | 1603.3638 | 1160.1839 | 1478.588 | 1166.0885 | 1088.9538 | 2046.9095 | 1711.1435 | 2107.7286 | 1808.6843 | 1942.0895 | 1563.0885 | 1592.7276 | 1355.3215 | 1484.1829 | 2450.4081 | 2416.6834 | 1695.8152 | 2203.7276 | 2159.9567 | 1875.3609 | 2596.9076 | 1541.41 | 1672.7758 | 1495.3648 | 1567.1298 | 1305.59 | 1755.5005 | 1904.6882 | 1504.7699 | 2867.0875 | 3116.3225 | 2630.7699 | 2726.7748 | 1654.3638 | 2163.8614 | 1893.642 | 1682.6814 | 1155.8161 | 1298.1809 | 1256.5919 | 1626.5467 | 1448.8633 | 2009.6352 | 1557.591 | 1330.9961 | 1587.4051 | 2178.2242 | 1524.409 | 1449.41 | 1691.0404 | 2311.2724 | 1439.9518 | 1501.8653 | 1361.8171 | 1648.9499 | 1210.5024 | 1256.8692 | 1810.2694 | 2148.1376 | 1692.6834 | 1692.8191 | 1451.4976 | 2067.7306 | 1628.591 | 1848.2704 | 1825.4985 | 1933.7709 | 1877.4071 | 2476.1829 | 1593.4995 | 1585.7729 | 1505.1318 | 2121.4582 | 1303.2271 | 1355.9105 | 1342.999 | 1558.0905 | 1252.8643 | 1223.6804 | 1313.5044 | 1193.9076 | 1369.5438 | 1370.9548 | 1291.4572 | 1504.3648 | 1567.2271 | 1797.3215 | 2131.8633 | 2786.411 | 2983.2252 | 3433.7709 | 3541.0049 | 4301.1858 | 1200.9557 | 1117.6814 | 1090.0462 | 1200.9557 | 1232.9076 | 1360.8171 | 1203.589 | 1286.6824 | 1187.9538 | 1294.2724 | 1266.6834 | 1219.5015 | 1167.4543 | 1131.7729 | 1175.7267 | 1210.2694 | 1111.1809 | 1153.8142 | 1096.591 | 1144.0443 | 1580.3205 | 1187.6342 | 1203.0875 | 1202.999 | 1325.8633 | 1205.2262 | 1290.4081 | 1037.999 | 1904.4081 | 2484.4061 | 2533.3589 | 1948.5428
198 | 4.1724 | 830.6131 | 1919.9838 | 1712.3761 | 1998.8245 | 1982.5389 | 2466.2134 | 1753.9548 | 1582.0398 | 1133 | 1484.557 | 1216.1175 | 1124.7016 | 2010.6637 | 1624.0054 | 2063.4557 | 1769.9439 | 1926.4503 | 1584.4666 | 1568.1971 | 1316.0561 | 1437.1357 | 2413.4395 | 2380.4105 | 1722.9385 | 2164.8717 | 2139.0054 | 1880.3997 | 2614.8933 | 1519.3092 | 1608.4051 | 1503.1736 | 1611.0236 | 1304.5172 | 1783.5497 | 1835.8535 | 1500.7089 | 2887.9838 | 3125.15 | 2628.8156 | 2772.3127 | 1671.8029 | 2176.3038 | 1837.8029 | 1695.5841 | 1166.5678 | 1301.5732 | 1247.2134 | 1607.8933 | 1479.9656 | 2021.208 | 1578.0669 | 1333.7035 | 1618.0904 | 2166.7541 | 1565.4484 | 1443.7414 | 1720.6583 | 2299.9892 | 1462.0959 | 1469.4666 | 1373.472 | 1689.0904 | 1196.9602 | 1214.6003 | 1856.1067 | 2091.8156 | 1690.1175 | 1719.7631 | 1471.7812 | 2003.3545 | 1639.3761 | 1834.3382 | 1835.7306 | 1932.7197 | 1893.8879 | 2443.8825 | 1620.6455 | 1591.5841 | 1496.4793 | 2081.8481 | 1302.8933 | 1342.8987 | 1329.0344 | 1583.3815 | 1230.9385 | 1225.5678 | 1270.1682 | 1203.4666 | 1345.2315 | 1334.293 | 1256.8427 | 1505.4322 | 1567.3146 | 1777.4376 | 2134.3146 | 2814.0885 | 2902.2949 | 3466.0054 | 3464.6239 | 4306.9873 | 1166.8825 | 1132.6908 | 1071.5732 | 1166.8825 | 1226.0344 | 1369.9439 | 1225.208 | 1305.5443 | 1192.1013 | 1270.7704 | 1238.8373 | 1181.4611 | 1149.9331 | 1134.2134 | 1163.7758 | 1227.4557 | 1138.1682 | 1169.9223 | 1123.3869 | 1145.2532 | 1557.0054 | 1206.7306 | 1236.6239 | 1233.6962 | 1358.651 | 1220.8427 | 1274.085 | 1068.6962 | 1870.3888 | 2549.7016 | 2564.821 | 1980.208
199 | 4.1834 | 796.0354 | 1845.3716 | 1695.4778 | 2096.885 | 1994.9911 | 2419.9646 | 1759.4602 | 1556.8673 | 1116.5221 | 1459.5133 | 1232.7611 | 1149.1593 | 1964.6637 | 1652.7788 | 2033.9381 | 1770.3717 | 1909.7876 | 1550.0531 | 1551.8407 | 1345.9823 | 1425.823 | 2347.469 | 2392.3894 | 1694.1593 | 2122.3363 | 2152.6018 | 1846.9735 | 2615.469 | 1499.9292 | 1578.1681 | 1545.4071 | 1569.7788 | 1299.4425 | 1785.469 | 1836.0177 | 1463.9469 | 2841.5575 | 3157.4601 | 2614.8938 | 2844.8141 | 1649.1327 | 2156.3363 | 1855.4602 | 1668.2743 | 1145.3009 | 1282.6283 | 1241.7257 | 1650.531 | 1558.1947 | 2025.5398 | 1615.2389 | 1296.7876 | 1573.2566 | 2116.6814 | 1551.6194 | 1449.9646 | 1664.531 | 2230.7345 | 1424.354 | 1469.7434 | 1342.6814 | 1644.2566 | 1229.9381 | 1258.8673 | 1796.1238 | 2034.9646 | 1721.0708 | 1770.8319 | 1478.8496 | 1994.9469 | 1596.8938 | 1791.2124 | 1813.0265 | 1967.8142 | 1888.6726 | 2422.9027 | 1635.9911 | 1641.8938 | 1467.7788 | 2088.3628 | 1286.1239 | 1333.6549 | 1298.0708 | 1599.1593 | 1239.0177 | 1194.3274 | 1273.2212 | 1195.9381 | 1327.8319 | 1291.469 | 1252.3628 | 1536.6726 | 1576.2743 | 1786.2832 | 2105.9823 | 2883.0973 | 2817.177 | 3436.2389 | 3425.1504 | 4349.8407 | 1133.7611 | 1106.1062 | 1089.0531 | 1133.7611 | 1223.6903 | 1371.2389 | 1232.5752 | 1350.6726 | 1206.1151 | 1215.0885 | 1246.7788 | 1172.9115 | 1124.9027 | 1104.4424 | 1117.7876 | 1193.1681 | 1138.1858 | 1126.8053 | 1157.0973 | 1121.0973 | 1587.0797 | 1160.177 | 1218.3982 | 1220.2743 | 1371.5487 | 1190.3451 | 1202.646 | 1057.8761 | 1808.469 | 2588.9027 | 2509.8584 | 1932.0708
200 | 4.1944 | 806.2429 | 1797.6888 | 1685.5521 | 2082.057 | 1972.203 | 2381.4307 | 1747.8918 | 1570.9317 | 1082.5142 | 1440.8539 | 1217.9867 | 1100.6642 | 1943.3928 | 1630.9621 | 2025.6091 | 1811.2144 | 1891.5845 | 1532.4725 | 1573.1347 | 1363.0683 | 1468.2163 | 2397.9297 | 2419.0551 | 1685.797 | 2126.2163 | 2113.1917 | 1850.1898 | 2629.3643 | 1486.203 | 1590.3643 | 1515.2999 | 1585.8653 | 1317 | 1776.3889 | 1834.0929 | 1481.5275 | 2835.5939 | 3046.2183 | 2653.6224 | 2797.2979 | 1643.3073 | 2156.649 | 1872.8766 | 1622.5674 | 1137.1898 | 1280.5806 | 1242.6072 | 1670.3131 | 1584.353 | 2007.6775 | 1608.6091 | 1348.5939 | 1556.6755 | 2176.6603 | 1475.7154 | 1429.501 | 1662.2296 | 2227.4155 | 1395.5408 | 1487.6622 | 1311.9051 | 1628.5674 | 1246.1347 | 1266.2429 | 1754.5787 | 2065.7817 | 1697.7856 | 1738.1234 | 1454.706 | 2014.7571 | 1570.4971 | 1811.4326 | 1810.7021 | 2007.7192 | 1906.6888 | 2446.2827 | 1626.1347 | 1648.5162 | 1538.0551 | 2096.1613 | 1239.4611 | 1313.6755 | 1329.1878 | 1554.6775 | 1220.3815 | 1193.7021 | 1238.852 | 1203.2163 | 1371.6091 | 1268.7989 | 1285.5787 | 1567.6205 | 1569.1499 | 1780.7552 | 2129.0777 | 2851.6907 | 2933.5521 | 3378.9867 | 3408.7154 | 4317.8368 | 1149.3093 | 1056.2296 | 1120.3112 | 1149.3093 | 1240.0152 | 1355.2448 | 1246.7704 | 1385.1878 | 1190.408 | 1178.2846 | 1224.9621 | 1190.5408 | 1102.353 | 1099.9165 | 1070.3663 | 1180.3928 | 1085.1367 | 1126.2979 | 1142.6357 | 1102.6224 | 1589.704 | 1153.4573 | 1203.852 | 1182.5939 | 1321.3928 | 1175.3093 | 1186.6337 | 1013.7837 | 1848.6736 | 2555.6509 | 2495.3378 | 1894.7286
201 | 4.2054 | 809.8761 | 1823.8702 | 1692.0924 | 1982.4267 | 1990.2016 | 2432.2045 | 1758.1642 | 1576.0029 | 1087.1613 | 1399.7463 | 1192.8387 | 1063.9567 | 1948.2852 | 1613.4897 | 1982.4641 | 1840.8358 | 1821.7493 | 1509.5506 | 1552.6745 | 1325.0059 | 1456.3776 | 2415.5939 | 2408.9971 | 1662.2999 | 2158.3658 | 2107.1957 | 1831.585 | 2622.6716 | 1508.7837 | 1552.055 | 1470.6224 | 1560.0462 | 1292.7148 | 1809.2163 | 1857.0865 | 1477.415 | 2875.6598 | 3010.3314 | 2593.0118 | 2745.5103 | 1662.4729 | 2133.5044 | 1862.3461 | 1614.1701 | 1141.9538 | 1256.0924 | 1259.2507 | 1621.4582 | 1551.8328 | 1973.2075 | 1591.1239 | 1353.8358 | 1536.173 | 2219.6657 | 1513.2449 | 1436.4926 | 1658.1268 | 2343.53 | 1398.2478 | 1498.8299 | 1324.8732 | 1621.0865 | 1213.3028 | 1251.0895 | 1759.1789 | 2086.0895 | 1622.8013 | 1695.9597 | 1419.4494 | 2006.0836 | 1628.7925 | 1820.2881 | 1860.1987 | 1943.5418 | 1914.0836 | 2481.8299 | 1607.9656 | 1569.5477 | 1502.7955 | 2122.3314 | 1233.3255 | 1342.2016 | 1324.5506 | 1542.6598 | 1174.2449 | 1170.8013 | 1231.415 | 1173.5074 | 1340.8358 | 1240.4985 | 1324.5447 | 1555.8013 | 1550.706 | 1820.8732 | 2114.3146 | 2827.2478 | 2939.6342 | 3380.8732 | 3441.6627 | 4323.9626 | 1171.4612 | 1079.1613 | 1086.3402 | 1171.4612 | 1190.7119 | 1316.9597 | 1232.5447 | 1377.8013 | 1121.7119 | 1186.118 | 1204.7404 | 1177.6686 | 1075.7896 | 1121.2999 | 1053.3717 | 1160.5418 | 1058.4985 | 1091.9223 | 1114.5044 | 1089.2075 | 1608.7375 | 1195.8761 | 1154.7178 | 1231.3658 | 1301.5418 | 1215.3314 | 1224.2478 | 1032.9538 | 1851.3028 | 2539.1583 | 2466.7581 | 1920.7493
202 | 4.2164 | 801.236 | 1821.9454 | 1681.6416 | 1967.1711 | 1959.7758 | 2429.7094 | 1761.472 | 1554.531 | 1088.8864 | 1408.9425 | 1184.9971 | 1092.6431 | 1971.8245 | 1674.9971 | 1972.9971 | 1826.7065 | 1802.4159 | 1493.5251 | 1541.9956 | 1338.9927 | 1428.5324 | 2346.8864 | 2423.4115 | 1638.7035 | 2188.649 | 2078.3097 | 1831.0575 | 2564.7714 | 1482.8348 | 1504.292 | 1436.8318 | 1527.767 | 1225.8894 | 1765.0605 | 1840.177 | 1486.4115 | 2881.5339 | 3034.0561 | 2555.5856 | 2777.1623 | 1626.2301 | 2106.0015 | 1862.4602 | 1656.4012 | 1125.7714 | 1209.8835 | 1260.2345 | 1606.4734 | 1511.1843 | 1983.1165 | 1597.882 | 1336.413 | 1490.5914 | 2204.2389 | 1509.1814 | 1468.059 | 1652.3525 | 2318.7787 | 1403.7655 | 1480.5944 | 1312.7684 | 1634.2891 | 1183.2906 | 1208.9439 | 1704.5295 | 2069.8215 | 1618.1106 | 1712.1136 | 1449.59 | 1988.0619 | 1627.823 | 1828.9425 | 1835.7182 | 1926.5324 | 1883.3584 | 2537.4632 | 1587.7625 | 1564.6431 | 1452.5944 | 2178.0502 | 1257.1785 | 1366.354 | 1290.646 | 1558.4144 | 1197.4705 | 1148.7021 | 1244.1755 | 1156.0561 | 1323.413 | 1259.233 | 1278.0649 | 1527.1151 | 1560.5295 | 1781.4793 | 2055.8157 | 2847.8215 | 2870.7699 | 3374.4144 | 3423.7124 | 4315.8215 | 1126.0649 | 1062.5368 | 1055.059 | 1126.0649 | 1206.5236 | 1300.649 | 1228.8805 | 1321.8259 | 1124.3496 | 1258.8747 | 1225.3569 | 1175.1165 | 1092.528 | 1066.1799 | 1077.41 | 1163.764 | 1079.5856 | 1054.2965 | 1089.8245 | 1066.1814 | 1635.7699 | 1186.295 | 1134.6431 | 1293.6431 | 1299.5885 | 1180.2433 | 1241.9985 | 1045.0015 | 1834.4646 | 2514.3643 | 2427.236 | 1931.9985
203 | 4.2274 | 825.9676 | 1801.3107 | 1686.5152 | 2023.1741 | 1934.003 | 2356.824 | 1724.998 | 1566.7257 | 1085.7257 | 1411.6214 | 1177.412 | 1093.8259 | 1964.5516 | 1637.9616 | 2015.6224 | 1860.6578 | 1831.5556 | 1542.2085 | 1519.7551 | 1368.5506 | 1434.1396 | 2369.7965 | 2485.3491 | 1657.3451 | 2156.2399 | 2010.0364 | 1826.4474 | 2548.6224 | 1465.6588 | 1562.8653 | 1424.6234 | 1532.8289 | 1201.8289 | 1725.5496 | 1884.7965 | 1481.7237 | 2869.4494 | 3089.6234 | 2570.0688 | 2845.6539 | 1653.6549 | 2105.2075 | 1894.6185 | 1710.3786 | 1098.0354 | 1212.6559 | 1273.2075 | 1612.1741 | 1448.999 | 1990 | 1575.3776 | 1326.586 | 1512.5202 | 2157.8604 | 1485.7591 | 1458.7925 | 1623.6185 | 2252.7945 | 1422.6569 | 1443.4828 | 1299.6559 | 1658.2409 | 1156.2379 | 1193.4484 | 1723.7945 | 2092.1387 | 1675.4838 | 1688.7207 | 1392.1691 | 1971.414 | 1634.3796 | 1853.3471 | 1751.3097 | 1884.3431 | 1883.3127 | 2534.7207 | 1580.2399 | 1602.5182 | 1452.6578 | 2133.9586 | 1258.5182 | 1337.3087 | 1349.6244 | 1533.3097 | 1198 | 1178.7591 | 1251.9312 | 1168.6549 | 1303.9302 | 1235.8249 | 1243.414 | 1540.1032 | 1568.2075 | 1732.8279 | 2057.6519 | 2864.8279 | 2946.18 | 3282.2015 | 3406.2085 | 4324.6204 | 1125.2095 | 1020.8623 | 1055.9312 | 1125.2095 | 1223.1711 | 1332.9342 | 1210.5496 | 1328.2429 | 1139.2065 | 1270.2733 | 1193.2753 | 1144.8249 | 1095.6893 | 1079.7611 | 1093.3107 | 1159.3097 | 1088.7581 | 1036.6204 | 1089.7935 | 1074.1406 | 1635.3815 | 1165.5506 | 1196.1751 | 1261.996 | 1337.589 | 1103.4808 | 1214.3077 | 1046.1387 | 1868.586 | 2467.6913 | 2462.1062 | 1919.2743
204 | 4.2384 | 828.2851 | 1852.5477 | 1631.2458 | 2054.4361 | 1961.177 | 2344.0821 | 1721.7345 | 1581.0885 | 1097.118 | 1389.8294 | 1185.151 | 1040.2851 | 2001.6165 | 1583.4558 | 2050.4164 | 1866.8392 | 1905.8427 | 1526.8884 | 1482.8427 | 1364.8722 | 1422.8687 | 2398.7212 | 2510.8884 | 1665.4656 | 2117.6622 | 2075.1411 | 1820.9213 | 2603.7149 | 1480.4031 | 1563.4818 | 1499.6657 | 1555.9705 | 1230.9115 | 1714.2984 | 1896.3933 | 1485.3968 | 2883.5246 | 3042.38 | 2547.7016 | 2817.0885 | 1666.3279 | 2150.6362 | 1802.5801 | 1668.7999 | 1090.7114 | 1216.0688 | 1222.0948 | 1618.6819 | 1477.5674 | 1981.0885 | 1549.9902 | 1299.7409 | 1587.6952 | 2132.4951 | 1485.0787 | 1467.823 | 1594.4558 | 2256.6721 | 1455.9213 | 1462.3181 | 1282.7704 | 1623.2753 | 1131.2296 | 1202.4263 | 1706.3835 | 2090.6229 | 1653.8687 | 1663.2394 | 1339.0295 | 1941.8195 | 1642.4558 | 1851.8589 | 1793.4887 | 1857.0197 | 1853.6489 | 2485.9705 | 1572.8525 | 1574.9179 | 1494.9803 | 2097.8265 | 1252.2065 | 1310.0295 | 1343.6194 | 1533.9017 | 1191.5639 | 1184.0197 | 1229.2458 | 1133.7999 | 1307.3378 | 1235.6067 | 1241.5442 | 1536.0098 | 1542.3441 | 1728.0983 | 2097.1151 | 2888.8132 | 2994.1637 | 3261.4002 | 3458.7149 | 4340.3083 | 1155.0885 | 1073.5182 | 1063.4459 | 1155.0885 | 1164.3835 | 1357.7212 | 1188.4066 | 1310.8982 | 1150.3279 | 1244.8722 | 1156.7802 | 1123.3181 | 1084.059 | 1112.0492 | 1063.849 | 1117.8195 | 1074.0688 | 1064.1903 | 1104.9115 | 1099.1573 | 1653.1573 | 1148.0098 | 1227.5836 | 1194.5737 | 1341.3835 | 1109.151 | 1167.0787 | 1023.8294 | 1875.9017 | 2499.4263 | 2488.1573 | 1894.5147
205 | 4.2494 | 802.5516 | 1845.7256 | 1631.8761 | 2037.1534 | 1887.1887 | 2390.0443 | 1753.059 | 1572 | 1036.292 | 1388.472 | 1182.9469 | 1031.8024 | 2032.882 | 1601.1918 | 2011.8141 | 1827.4336 | 1859.0855 | 1456.0501 | 1500.0148 | 1364.8082 | 1411.3835 | 2379.0443 | 2416.7109 | 1652.2566 | 2123.3245 | 2139.3983 | 1768.115 | 2650.7581 | 1470.9646 | 1530.236 | 1458.9085 | 1575.7434 | 1232.3894 | 1710.1386 | 1873.2508 | 1477.7345 | 2881 | 3001.2154 | 2497.0649 | 2830.3245 | 1645.9469 | 2171.764 | 1832.298 | 1659.3098 | 1083.2065 | 1161.3067 | 1188.3983 | 1590.8672 | 1457.5634 | 1966.4189 | 1544.9115 | 1309.5015 | 1567.7551 | 2105.0944 | 1489.177 | 1467.7345 | 1626.3983 | 2235.9557 | 1434.0649 | 1437.8731 | 1267.6844 | 1596.5221 | 1171.177 | 1205.4336 | 1686.7876 | 2094.7434 | 1633.2508 | 1693.5516 | 1401.4514 | 1933.354 | 1623.6607 | 1832.2508 | 1847.0148 | 1854.4926 | 1818.4573 | 2497.118 | 1559.587 | 1573.6048 | 1485.3304 | 2130.469 | 1262.4573 | 1314.1033 | 1295.7286 | 1534.3599 | 1176.3746 | 1192.1475 | 1212.6254 | 1125.3245 | 1317.9115 | 1261.7935 | 1226.8525 | 1556.3098 | 1483.6607 | 1733.7286 | 2162.6312 | 2959.2449 | 2923.4631 | 3352.4278 | 3425.5929 | 4375.7286 | 1115.0501 | 1109.3009 | 1059.3599 | 1115.0501 | 1141.236 | 1312.6755 | 1198.5074 | 1272.0295 | 1117.2625 | 1233.1386 | 1136.6254 | 1124.7492 | 1067.8525 | 1057.7846 | 1051.3393 | 1089.5664 | 1063.9557 | 1052.8731 | 1096.7492 | 1073.8672 | 1595.3952 | 1132.2861 | 1203.2714 | 1209.3245 | 1302 | 1149.059 | 1214.8113 | 1011.3098 | 1904.2655 | 2503.9557 | 2498.3835 | 1932.0679
206 | 4.2604 | 805.4803 | 1777.8904 | 1661.1573 | 2010.882 | 1909.3909 | 2339.4744 | 1756.4803 | 1571.4803 | 1064.8712 | 1384.6627 | 1136.646 | 1056.677 | 1990.1824 | 1633.7473 | 1965.4803 | 1825.5983 | 1754.3481 | 1466.0143 | 1487.4887 | 1341.6067 | 1431.2837 | 2366.41 | 2393.1013 | 1638.5983 | 2166.826 | 2160.2837 | 1739.6603 | 2599.4041 | 1488.4887 | 1489.3564 | 1389.6377 | 1576.9297 | 1236.9523 | 1668.3314 | 1845.8284 | 1441.292 | 2900.2276 | 3019.379 | 2505.9356 | 2843.1657 | 1609.5197 | 2120.143 | 1861.8367 | 1643.1037 | 1080.8904 | 1166.9666 | 1173.9464 | 1616.298 | 1484.9022 | 1960.4803 | 1553.031 | 1353.913 | 1495.1347 | 2141.3993 | 1502.0393 | 1451.559 | 1607.2694 | 2216.3707 | 1361.3481 | 1442.2503 | 1283.118 | 1564.2217 | 1135.7104 | 1203.0787 | 1668.0644 | 2077.2217 | 1671.826 | 1717.7866 | 1422.3874 | 1961.1573 | 1593.7247 | 1818.8594 | 1863.59 | 1844.646 | 1845.3623 | 2555.9273 | 1590.2026 | 1575.025 | 1436.584 | 2125.1406 | 1259.2217 | 1291.8987 | 1302.1657 | 1525 | 1160.7247 | 1197.8427 | 1211.646 | 1145.9607 | 1296.41 | 1250.5983 | 1202.969 | 1532.8284 | 1470.8737 | 1727.174 | 2117.2134 | 2945.0811 | 2909.5673 | 3358.1347 | 3386.826 | 4408.6686 | 1065.7724 | 1064.82 | 1027.1347 | 1065.7724 | 1167.913 | 1284.764 | 1153.2694 | 1282.8343 | 1105.9356 | 1204.8427 | 1147.0787 | 1129.3063 | 1026.3004 | 1006.882 | 1045.938 | 1053.9857 | 1056.8427 | 1025.5506 | 1088.3147 | 1074.354 | 1599.1716 | 1125.795 | 1180.7247 | 1228.4017 | 1336.8176 | 1123.8987 | 1241.4577 | 1007.5757 | 1850.2861 | 2520.7473 | 2548.944 | 1960.1263
207 | 4.2714 | 790.6372 | 1799.705 | 1686.2783 | 2026.6185 | 1931.5919 | 2330.4287 | 1676.2064 | 1549.7217 | 1107.3382 | 1372.7463 | 1158.3216 | 1055.6804 | 1996.3009 | 1621.1465 | 1949.0413 | 1859.9174 | 1783.6864 | 1514.7031 | 1496.2783 | 1275.6125 | 1387.6991 | 2355.1917 | 2400.8495 | 1623.4454 | 2180.0846 | 2097.0167 | 1748.3176 | 2573.237 | 1487.1711 | 1495.5153 | 1431.8329 | 1556.0846 | 1220.8702 | 1662.1072 | 1848.0885 | 1413.8928 | 2855.761 | 3011.1505 | 2572.9174 | 2828.0846 | 1611.5959 | 2136.472 | 1804.2124 | 1622.2989 | 1084.0865 | 1190.53 | 1156.7463 | 1576.1858 | 1512.1445 | 1929.1465 | 1561.2124 | 1385.2783 | 1505.9174 | 2185.1052 | 1477.998 | 1449.8083 | 1583.6185 | 2224.9381 | 1373.6638 | 1459.4661 | 1259.9341 | 1551.0865 | 1079.765 | 1162.9754 | 1642.0433 | 2028.9774 | 1667.53 | 1740.7237 | 1373.4248 | 1935.2104 | 1579.0206 | 1781.4248 | 1856.7217 | 1816.3176 | 1817.53 | 2550.3362 | 1561.5466 | 1544.3196 | 1384.4454 | 2157.2557 | 1229.0639 | 1301.705 | 1271.1258 | 1532.4474 | 1152.4041 | 1196.6598 | 1205.1278 | 1140.2124 | 1263.3815 | 1262.0433 | 1246.9194 | 1476.5732 | 1531.5998 | 1724.5546 | 2090.5526 | 2987.0719 | 2912.8309 | 3380.6205 | 3437.8535 | 4408.5093 | 1070.8742 | 1054.0865 | 984.7217 | 1070.8742 | 1173.7443 | 1261.4454 | 1126.9587 | 1265.4661 | 1102.1671 | 1217.4907 | 1143.1485 | 1126.6578 | 1040.6205 | 1005.0639 | 1022.6598 | 1105.3049 | 1053 | 1026.233 | 1063.5506 | 1040.9754 | 1626.7011 | 1105.3382 | 1154.8495 | 1230.7876 | 1317.9321 | 1135.3009 | 1201.6372 | 1002.0226 | 1814.3216 | 2514.53 | 2532.6352 | 1921.7424
208 | 4.2824 | 776.3673 | 1827.1018 | 1696.292 | 2028.8495 | 1923.2876 | 2291.438 | 1657.7036 | 1540.2522 | 1049.2566 | 1367.5973 | 1132.7522 | 1024.8761 | 1971.7654 | 1571.1416 | 1901.2566 | 1850.9911 | 1839.0133 | 1520.646 | 1497.0398 | 1292.3983 | 1357.8097 | 2337.8628 | 2328.4867 | 1578.0664 | 2152.2301 | 2046 | 1787.0576 | 2642.4646 | 1494.2788 | 1518.1548 | 1445.8628 | 1556.531 | 1205.1637 | 1696.0354 | 1862.0531 | 1425.1372 | 2840.1859 | 3028.5442 | 2634.1991 | 2853.5664 | 1630.4159 | 2113.6017 | 1762.9911 | 1639.9867 | 1067.8894 | 1169.5442 | 1136.3584 | 1576.1328 | 1460.2788 | 1955.3363 | 1537.9513 | 1348.4867 | 1523.115 | 2164.3805 | 1457.0885 | 1429.0398 | 1619.0487 | 2215.5221 | 1381.3584 | 1474.0089 | 1252.7832 | 1558.8363 | 1091.8584 | 1145.0487 | 1691.3629 | 1997.1283 | 1630.2168 | 1731.4823 | 1373.1637 | 1953.6283 | 1558.3938 | 1774.6327 | 1787.9292 | 1893.1549 | 1807.9735 | 2508.8982 | 1563.0178 | 1575.6682 | 1391.8584 | 2147.938 | 1188.0929 | 1292.0929 | 1270.0487 | 1542.8097 | 1122.1548 | 1148.9292 | 1200.469 | 1131.1018 | 1249.177 | 1243.8761 | 1240.3982 | 1501.6416 | 1566.115 | 1741.6327 | 2094.885 | 3049.3318 | 2951.1239 | 3335.73 | 3432.3451 | 4425.5664 | 1054.5354 | 1029.1814 | 978.5177 | 1054.5354 | 1172.177 | 1243.2788 | 1157.3274 | 1231.0265 | 1105.5531 | 1174.3097 | 1177.1859 | 1141.8319 | 1065.938 | 997.8363 | 1010.3053 | 1126.3982 | 1047.5575 | 1018.2655 | 998.5354 | 1019.7832 | 1613.9115 | 1132.7036 | 1153.7832 | 1189.2832 | 1252.8628 | 1157.0265 | 1199.3407 | 1035.5044 | 1812.6991 | 2512.5929 | 2517.0487 | 1910.1505
209 | 4.2934 | 752.7384 | 1856.9489 | 1685.6608 | 2028.1308 | 1946.4346 | 2242.9135 | 1687.0865 | 1532.4169 | 1079.4366 | 1345.2173 | 1089.5565 | 1003.5565 | 1896.2861 | 1564.2439 | 1875.7827 | 1791.06 | 1786.1376 | 1500.3038 | 1472.5388 | 1339.6519 | 1363.7827 | 2298.1819 | 2320.9577 | 1584.7227 | 2136.2262 | 2054.3481 | 1812.295 | 2630.1465 | 1468.6254 | 1500.1219 | 1436.7915 | 1548.0865 | 1225.3569 | 1689.7296 | 1856.47 | 1430.3127 | 2866.3215 | 3066.2881 | 2592.9891 | 2912.8358 | 1618.8427 | 2041.8604 | 1789.3235 | 1627.7473 | 1058.2527 | 1137.7207 | 1098.4346 | 1623.9911 | 1458 | 1977.1819 | 1539.2527 | 1301.4435 | 1527.9135 | 2188.6097 | 1475.8093 | 1390.6254 | 1636.1131 | 2223.9489 | 1344.5477 | 1459.3638 | 1263.7738 | 1540.3127 | 1111 | 1120.9292 | 1673.1288 | 1965.1907 | 1623.6785 | 1673.1376 | 1403.3746 | 1967.6254 | 1590.4101 | 1791.0266 | 1751.5654 | 1875.7187 | 1832.47 | 2522.4612 | 1553.763 | 1560.7119 | 1403.765 | 2137 | 1167.47 | 1288.7227 | 1311.0443 | 1525.408 | 1113.5831 | 1141.5054 | 1199.4523 | 1110.8604 | 1259.3481 | 1222 | 1217.4877 | 1547.9135 | 1561.4523 | 1775.8358 | 2083.9646 | 2975.7984 | 2957.9734 | 3288.8181 | 3398.8781 | 4443.0954 | 1030.6785 | 1025.6097 | 1000.1308 | 1030.6785 | 1163.9823 | 1242.6696 | 1134.2684 | 1234.3569 | 1108.06 | 1159.9489 | 1164.3638 | 1117.0157 | 1010.4769 | 977.0865 | 969.06 | 1084.0089 | 1021.8515 | 1001.5211 | 992.4877 | 1028.5477 | 1560.1288 | 1166.5388 | 1164.2173 | 1153.2262 | 1264.3835 | 1094.1288 | 1224.1396 | 1028.1819 | 1834.0797 | 2506.1819 | 2418.9095 | 1967.6362
210 | 4.3044 | 784.4882 | 1767.3898 | 1684.7065 | 2011.4026 | 1921.5221 | 2290.974 | 1635.6131 | 1537.0649 | 1081.6312 | 1328.7792 | 1107.0649 | 1021.6337 | 1934.7272 | 1615.5246 | 1874.0182 | 1810.0934 | 1804.3142 | 1493.826 | 1490.0285 | 1304.5585 | 1309.6858 | 2285.1195 | 2327.5403 | 1615.6883 | 2116.3844 | 2039.3844 | 1758.026 | 2596.5688 | 1440.8987 | 1544.8181 | 1465.1844 | 1554.6337 | 1230.8805 | 1694.735 | 1843.5403 | 1443.1844 | 2843.3663 | 3057.8884 | 2567.5792 | 2929.3299 | 1603.6883 | 2001.1273 | 1819.6416 | 1615.7065 | 1049.2468 | 1124.5143 | 1145.1947 | 1633.0831 | 1408.5118 | 1970.9454 | 1533.6598 | 1278.2468 | 1565.735 | 2119.143 | 1513.322 | 1418.0753 | 1622.7792 | 2220.2832 | 1334.7429 | 1416.7429 | 1258.0182 | 1530.4312 | 1084.8338 | 1122.561 | 1682.7636 | 1947.587 | 1619.8987 | 1637.2753 | 1417.4494 | 1927.9533 | 1588.026 | 1852.4882 | 1740.6416 | 1795.0182 | 1774.0339 | 2534.5506 | 1546.4882 | 1565.6986 | 1402.5506 | 2146.6701 | 1210.2857 | 1294.6598 | 1290.6131 | 1491.9351 | 1149.322 | 1136.5767 | 1157.4936 | 1112.7896 | 1262.0182 | 1202.0909 | 1176.1352 | 1517.4754 | 1545.0909 | 1744.5481 | 2080.5143 | 3008.9479 | 2909.026 | 3299.0364 | 3388.8805 | 4482.0753 | 1070.1298 | 1046.587 | 995.9169 | 1070.1298 | 1145.1922 | 1236.6052 | 1103.8623 | 1246.7065 | 1052.5767 | 1155.1455 | 1121.174 | 1095.0467 | 987.5143 | 989.8909 | 1009.709 | 1112.4597 | 1004.4312 | 975.7611 | 971.6131 | 1042.0649 | 1550.4233 | 1119.1819 | 1180.2026 | 1135.6598 | 1290.9818 | 1054.2753 | 1200.026 | 1015.6883 | 1831.3195 | 2508.4779 | 2378.6259 | 1938.0157
211 | 4.3154 | 796.4307 | 1725.4897 | 1673.7965 | 1969.3363 | 1879.2684 | 2311.4454 | 1653.2331 | 1554.649 | 1045.6755 | 1347.0561 | 1143.2449 | 1022.2153 | 1924.7935 | 1568.5486 | 1955.2626 | 1846.0678 | 1797.6047 | 1517.5959 | 1494.472 | 1258.0118 | 1338.4779 | 2284.9322 | 2313.649 | 1638.5959 | 2105.2301 | 2025.3245 | 1740.4602 | 2676.6136 | 1411.3363 | 1546.5251 | 1450.0383 | 1558.1209 | 1257.7994 | 1686.5251 | 1831.9735 | 1464.0678 | 2869.5163 | 3046.4336 | 2575.472 | 2886.5516 | 1610.9056 | 2006.1505 | 1827.7581 | 1633.8525 | 1058.7316 | 1147.3658 | 1132.8613 | 1594.7669 | 1368.0944 | 1913.171 | 1499.823 | 1302.2596 | 1577.2832 | 2106.1121 | 1492.646 | 1421.6194 | 1595.7286 | 2194.6755 | 1332.3245 | 1412 | 1251.5133 | 1517.7965 | 1048.7286 | 1131.7286 | 1695.2684 | 2005.4366 | 1572.9026 | 1643.649 | 1408.1888 | 1939.3127 | 1568.7168 | 1876.6342 | 1762.705 | 1797.8112 | 1796.5458 | 2542.0678 | 1575.2832 | 1562.3628 | 1379.2684 | 2189.4484 | 1195.646 | 1238.7404 | 1219.9026 | 1486.5546 | 1165.2566 | 1100.3776 | 1124.1622 | 1092.0383 | 1240.4041 | 1185.8378 | 1174.7463 | 1476.7021 | 1518.9587 | 1707.3923 | 2090 | 3044.6608 | 2936.7876 | 3297.0678 | 3403.5959 | 4553.6106 | 1064.6194 | 1027.1475 | 983.0944 | 1064.6194 | 1153.8525 | 1217.0531 | 1125.6637 | 1268.9204 | 1049.5015 | 1136.8112 | 1118.3245 | 1095.3097 | 995.8378 | 1021.9469 | 1061.4867 | 1110.0796 | 997.6077 | 959.8643 | 972.3806 | 1003.2802 | 1579.0678 | 1092.4867 | 1181.7965 | 1151.7994 | 1280.6342 | 1096.6785 | 1133.0649 | 986.8761 | 1839.1918 | 2519.9322 | 2412.9322 | 1902.528
212 | 4.3264 | 812.4179 | 1751.8791 | 1660.2527 | 2027.1765 | 1852.2856 | 2296.5054 | 1649.4282 | 1529.5162 | 1060.0221 | 1328.8348 | 1088.1971 | 982.736 | 1904.5826 | 1518.5275 | 1962.5929 | 1811.0216 | 1761.4066 | 1491.2414 | 1470.9449 | 1258.3958 | 1362.703 | 2276.6593 | 2276.0545 | 1653.1868 | 2094.7581 | 2082.1874 | 1752.7473 | 2713.824 | 1390.8461 | 1489.8018 | 1427 | 1561.6814 | 1234.9007 | 1700.4179 | 1826.3186 | 1468.1868 | 2864.1755 | 3009.7468 | 2540.6701 | 2879.9892 | 1566.7138 | 2053.4838 | 1782.6588 | 1618.703 | 1026.3731 | 1166.4946 | 1122.1101 | 1598.5718 | 1360.6593 | 1875.088 | 1477.0988 | 1305.8899 | 1507.5049 | 2115.3623 | 1459.0659 | 1403.1539 | 1564.4174 | 2143.1534 | 1348.8353 | 1441.6706 | 1216.3623 | 1509 | 1065.4838 | 1153.8245 | 1685.4946 | 2030.0108 | 1553.4946 | 1693.2862 | 1356.527 | 1938.2306 | 1542.3623 | 1856.912 | 1761.4833 | 1802 | 1783.351 | 2553.9012 | 1524.4941 | 1551.6814 | 1389.6706 | 2219.7473 | 1211.9125 | 1211.9671 | 1213.5497 | 1496.1539 | 1145.0108 | 1121.6706 | 1170.0993 | 1089.176 | 1224.0329 | 1166.6042 | 1182.6701 | 1484.2419 | 1519.9012 | 1697.1647 | 2110.176 | 3076.8796 | 2921.8126 | 3223.4498 | 3307.8121 | 4518.2851 | 1060.3958 | 1007.8791 | 984.5605 | 1060.3958 | 1122.6809 | 1199.8791 | 1144.3736 | 1250.1755 | 1095.8574 | 1162.3628 | 1106.3515 | 1088.7473 | 976.0108 | 1005.2964 | 983.2959 | 1094.4066 | 955.0216 | 996.7257 | 991.4395 | 1006.9455 | 1602.7694 | 1122.297 | 1154.6042 | 1149.6701 | 1250.2306 | 1145.5826 | 1122.209 | 953.5054 | 1811.4061 | 2480.8018 | 2419.4946 | 1946.8245
213 | 4.3374 | 814.1947 | 1711.2635 | 1650.943 | 2067.0973 | 1867.5939 | 2208.8357 | 1595.7433 | 1558.4563 | 1046.7719 | 1325.6224 | 1092.3019 | 951.703 | 1889.5319 | 1535.7198 | 1932.057 | 1765.0118 | 1774.3255 | 1424.0236 | 1488.0452 | 1336.416 | 1420.3186 | 2291.7768 | 2292.2164 | 1655.8171 | 2123.765 | 2070.3323 | 1731.7719 | 2691.5202 | 1350.4464 | 1467.4228 | 1365.824 | 1521.7551 | 1210.057 | 1676.7551 | 1768.7266 | 1417.5152 | 2873.708 | 2977.2114 | 2531.2684 | 2894.8171 | 1571.9479 | 2038.4749 | 1778.9479 | 1621.1996 | 1012.8053 | 1129.9951 | 1109.1091 | 1592.8977 | 1385.4681 | 1862.6175 | 1484.9312 | 1286.7886 | 1485.8053 | 2090.5772 | 1428.9548 | 1381.7719 | 1570.3825 | 2117.4798 | 1359.6057 | 1460.7886 | 1184.3087 | 1519.297 | 1075.3087 | 1136.4179 | 1663.5319 | 1984.8977 | 1539.4061 | 1683.2065 | 1328.2399 | 1945.1996 | 1536.9597 | 1903.3019 | 1751.8741 | 1881.3471 | 1738.8171 | 2540.6175 | 1521.2567 | 1557.1829 | 1463.3019 | 2254.4966 | 1184.4582 | 1257.0049 | 1281.0905 | 1538.9764 | 1127.7601 | 1112.3205 | 1166.4749 | 1128.4395 | 1228.057 | 1177.6224 | 1202.9479 | 1537.0049 | 1518.1259 | 1700.9312 | 2039.8072 | 3096.2802 | 2836.1494 | 3145.5605 | 3333.5468 | 4557.1643 | 1040.8977 | 1002.3943 | 976.0973 | 1040.8977 | 1101.4798 | 1218.6224 | 1084.824 | 1172.4867 | 1081.2635 | 1156.4061 | 1123.651 | 1095.0285 | 963.6057 | 977.9144 | 957.0738 | 1033.4297 | 964.7365 | 966.0354 | 986.5772 | 1018.0403 | 1547.3441 | 1162.1996 | 1169.2567 | 1139.057 | 1252.354 | 1140.2232 | 1157.0167 | 946.3943 | 1818.2164 | 2468.7198 | 2448.411 | 1988.1711
214 | 4.3484 | 773.4262 | 1718.8997 | 1585.2006 | 2061.854 | 1870.4941 | 2100.9985 | 1610.1357 | 1606.0679 | 1020.4941 | 1307.0103 | 1142.4499 | 941.292 | 1854.9321 | 1511.8643 | 1898.9203 | 1760.0339 | 1845.4174 | 1496.8791 | 1481.5162 | 1312.2566 | 1420.5604 | 2290.4041 | 2271.5604 | 1637.8422 | 2149 | 1979.2905 | 1739.4838 | 2590.2109 | 1384.0915 | 1441.1902 | 1393.3614 | 1536.7316 | 1188.0442 | 1650.944 | 1753.9218 | 1414.8997 | 2872.1342 | 2941.5383 | 2577.4056 | 2936.7876 | 1568.1902 | 2049.3717 | 1771.4823 | 1637.472 | 1002.2802 | 1100.146 | 1104.4499 | 1591.1578 | 1431.8098 | 1873.4838 | 1466.8082 | 1313.8098 | 1522.2478 | 2081.5841 | 1423.1239 | 1379.5959 | 1578.674 | 2112.4941 | 1365.562 | 1443.9321 | 1221.2596 | 1505.6062 | 1075.7979 | 1125.4499 | 1647.472 | 1981.9218 | 1572.3496 | 1708.5295 | 1341.5059 | 1946.6401 | 1543.056 | 1861.9867 | 1745.2581 | 1839.6828 | 1751.7419 | 2521.2021 | 1572.9897 | 1512.7404 | 1383.053 | 2224.3245 | 1119.1902 | 1284.382 | 1288.4262 | 1545.8422 | 1110.7861 | 1103.5059 | 1145.236 | 1110.3923 | 1228.91 | 1224.5177 | 1195.4823 | 1545.2242 | 1496.9321 | 1706.292 | 2031.9676 | 3160.9233 | 2881.6313 | 3156.1918 | 3285.289 | 4629.7758 | 1032.3599 | 934.6386 | 960.9661 | 1032.3599 | 1088.4602 | 1229.056 | 1064.1578 | 1200.0354 | 1045.0221 | 1091.0885 | 1094.6504 | 1068.5723 | 994.8997 | 984.9779 | 979.708 | 1033.5516 | 970.6962 | 949.8658 | 934.3245 | 986.6062 | 1529.8098 | 1097.5147 | 1175.8761 | 1085.5265 | 1220.3584 | 1123.056 | 1155.3141 | 957.1239 | 1856.2699 | 2417.0545 | 2422.8304 | 2001.854
215 | 4.3594 | 775.9086 | 1731.6971 | 1529.352 | 2007.5349 | 1831.2252 | 2167.2154 | 1634.412 | 1581.9646 | 971.1406 | 1303.0177 | 1071.3235 | 964.2114 | 1892.2714 | 1502.7149 | 1839.4434 | 1768.1514 | 1776.8515 | 1467.2389 | 1439.2674 | 1278.5634 | 1389.3697 | 2257.4189 | 2338.064 | 1640.4966 | 2164.1269 | 2016.7434 | 1798.7503 | 2655.7965 | 1434.4966 | 1441.0177 | 1428.0669 | 1563.7394 | 1204.7748 | 1664.7149 | 1776.3451 | 1364.0383 | 2895.1868 | 2944.2114 | 2474.4217 | 2864.3057 | 1598.2714 | 2011.3834 | 1724.8977 | 1589.8377 | 983.8486 | 1096.4789 | 1099.5034 | 1579.6794 | 1474.2114 | 1903.7571 | 1414.0737 | 1305.1583 | 1556.1691 | 2035.4012 | 1431.2606 | 1359.6372 | 1547.0069 | 2114.824 | 1350.7217 | 1471.8171 | 1217.7463 | 1510.0177 | 1078.7394 | 1117.352 | 1618.7463 | 1950.1406 | 1614.236 | 1709.4857 | 1364.3874 | 1969.884 | 1601.1377 | 1846.5143 | 1736.1091 | 1821.0777 | 1762.7817 | 2552.296 | 1600.3029 | 1545.8771 | 1306.9577 | 2210.0177 | 1157.4651 | 1261.3097 | 1261.9154 | 1530.8486 | 1109.5634 | 1105.588 | 1136.5457 | 1086.4789 | 1257.884 | 1191.6863 | 1169.6971 | 1513.4611 | 1486.3697 | 1731.7326 | 2064.7817 | 3153.056 | 2984.4651 | 3182.4543 | 3212.4789 | 4609.2252 | 1020.4611 | 938.3806 | 937.7217 | 1020.4611 | 1112.9263 | 1212.764 | 1060.8731 | 1201.1829 | 1039.5211 | 1093.6657 | 1128.3983 | 1090.708 | 1004.5457 | 968.4189 | 952.3766 | 1039.0914 | 962.5211 | 951.3943 | 962.2468 | 978.412 | 1550.1514 | 1107.6166 | 1128.6617 | 1125.6343 | 1198.412 | 1077.1406 | 1160.1269 | 936.2674 | 1883.4297 | 2391.9754 | 2378.4611 | 2004.8909
216 | 4.3704 | 770.293 | 1796.1131 | 1571.0487 | 2003.9159 | 1842.8437 | 2283.1893 | 1626.7188 | 1569.0683 | 979.414 | 1280.8712 | 1031.7109 | 1013.0566 | 1930.2812 | 1529.5625 | 1870.6913 | 1780.9238 | 1724.2812 | 1404.3614 | 1433.9238 | 1288.1406 | 1335.0236 | 2290.5507 | 2326.8869 | 1662.4941 | 2126.7345 | 2104.9804 | 1755.6017 | 2709.7227 | 1407.6583 | 1481.8437 | 1371.9513 | 1565.5703 | 1225.2852 | 1609.9552 | 1752.7989 | 1349.3456 | 2904.791 | 2970.5664 | 2483.8201 | 2873.1854 | 1627.2852 | 2002.0605 | 1778.3983 | 1579.6347 | 991.6386 | 1090.502 | 1102.5664 | 1553.791 | 1441.4493 | 1948.3417 | 1422.5585 | 1320.8437 | 1494.7463 | 2040.5585 | 1399.2286 | 1382.1972 | 1565.9843 | 2135.4218 | 1356.0644 | 1459.3009 | 1197.4297 | 1496.2247 | 1048.4415 | 1123.707 | 1575.2969 | 1957.1288 | 1606.5782 | 1717.9882 | 1350.8672 | 1939.8068 | 1539.3284 | 1855.3614 | 1765.2011 | 1813.7989 | 1734.1563 | 2621.1932 | 1605.9277 | 1528.3087 | 1300.7871 | 2202.0079 | 1144.3771 | 1218.4415 | 1218.088 | 1494.3009 | 1142.2734 | 1103.6426 | 1097.586 | 1087.5703 | 1268.2891 | 1172.5664 | 1176.9238 | 1490.4336 | 1536.121 | 1756.3535 | 2010.4533 | 3158.6269 | 2916.1155 | 3200.209 | 3225.7793 | 4616.3456 | 1030.8476 | 950.8633 | 910.293 | 1030.8476 | 1089.0196 | 1195.3614 | 1077.7031 | 1195.707 | 1026.8633 | 1122.9961 | 1141.2247 | 1077.3731 | 955.9474 | 946.9317 | 949.1367 | 1031.8555 | 946.6504 | 967.7753 | 937.2404 | 977.7871 | 1537.8633 | 1115.6544 | 1104.3574 | 1174.498 | 1220.2773 | 1061.6386 | 1144.5821 | 949.5585 | 1829.8146 | 2453.619 | 2319.4493 | 2016.2812
217 | 4.3814 | 775.4071 | 1739.0442 | 1600.6549 | 2050.1239 | 1825.292 | 2277.6725 | 1592.8407 | 1537.6371 | 944.9557 | 1247.2832 | 1053.6814 | 961.2743 | 1926.8319 | 1503.6725 | 1939.9469 | 1784.3805 | 1760.7788 | 1387.2124 | 1361.4159 | 1281.8319 | 1348.8761 | 2233.8318 | 2338.1151 | 1613.0265 | 2119.9912 | 2071.7168 | 1708.4867 | 2688.5221 | 1388.6549 | 1502.9292 | 1359.3717 | 1522.469 | 1257.4336 | 1608.0885 | 1728.7965 | 1375.1328 | 2856.0885 | 2931.1593 | 2469.4424 | 2877.9115 | 1597.2212 | 2045.469 | 1739.0707 | 1597.2035 | 979.9734 | 1097.823 | 1091.3186 | 1532.5929 | 1378.3274 | 1971.2743 | 1494.6638 | 1314.4248 | 1476.2655 | 2038.6637 | 1384.2743 | 1339.823 | 1506.761 | 2105.3628 | 1355.177 | 1464.2301 | 1197.3097 | 1483.6549 | 1032.6549 | 1122.5221 | 1593.0885 | 1957.9734 | 1561.3274 | 1648.5221 | 1303.292 | 1894.6637 | 1536.0797 | 1909.593 | 1783.3097 | 1793.7257 | 1693.4248 | 2628.354 | 1556.5752 | 1461.0177 | 1320.6106 | 2226.708 | 1115.0708 | 1211.1681 | 1222.5398 | 1509.708 | 1109.5398 | 1102.0354 | 1107.1593 | 1076.5575 | 1260.7257 | 1213.0531 | 1135.8495 | 1486.6195 | 1535.1858 | 1764.2743 | 2016.3275 | 3163.2832 | 2914.354 | 3277.0708 | 3283.7699 | 4689.9381 | 1082.2921 | 973.1239 | 879.3894 | 1082.2921 | 1079.062 | 1176.2478 | 1006.2123 | 1198.4513 | 1039.3009 | 1126.6903 | 1121.4513 | 1030.1504 | 921.5221 | 926.6283 | 955.6903 | 1046.4425 | 941.5841 | 939.7433 | 919.062 | 992.3717 | 1566.6726 | 1110.823 | 1077.1504 | 1166.5221 | 1265.708 | 1054.5575 | 1130.3451 | 953.5575 | 1829.0177 | 2407.3805 | 2393.8496 | 2016.1062
218 | 4.3924 | 749.6421 | 1709.0148 | 1577.6494 | 1994.614 | 1886.3919 | 2173.2708 | 1586.0059 | 1535.0118 | 975.71 | 1247.6731 | 1054.9941 | 942.6804 | 1929.3387 | 1512.0207 | 1894.2841 | 1837.707 | 1777.6672 | 1403.0162 | 1391.386 | 1279.0015 | 1377.3373 | 2212.0162 | 2282.6066 | 1584.3343 | 2101.9793 | 2030.3255 | 1723.3476 | 2749.0487 | 1374.3225 | 1464.3033 | 1393.352 | 1527.6878 | 1191.2723 | 1645.6819 | 1763.028 | 1386.3373 | 2854.6819 | 2912.3373 | 2492.3742 | 2895.6923 | 1555.648 | 2064.3358 | 1707.3402 | 1622.3491 | 1016.3668 | 1069.6421 | 1052.645 | 1541.0118 | 1399.6967 | 1929.6347 | 1490.6421 | 1316.6775 | 1477.3255 | 2012.9882 | 1409.3506 | 1328.6834 | 1494.355 | 2089.003 | 1319.9779 | 1422.293 | 1172.648 | 1447.3063 | 1089.7085 | 1062.293 | 1636.0192 | 1934.9897 | 1519.3166 | 1706.7424 | 1331.034 | 1866.9867 | 1603.0295 | 1916.3151 | 1769.9897 | 1828.6937 | 1712.3565 | 2642.6849 | 1544.6775 | 1544.4096 | 1342.6745 | 2263.6804 | 1151.028 | 1183.9764 | 1229.3284 | 1486.6362 | 1065.3196 | 1086.3225 | 1095.9808 | 1042.6465 | 1205.6273 | 1201.645 | 1148.3609 | 1538.0369 | 1531.3417 | 1755.9926 | 2033.3314 | 3149.6627 | 2902.3033 | 3203.2517 | 3272.6391 | 4706.3225 | 1042.2886 | 932.2915 | 896.0207 | 1042.2886 | 1137.0369 | 1158.9926 | 1031.7188 | 1205.0074 | 1016.6421 | 1142.0118 | 1068.6303 | 987.3122 | 914.6657 | 960.031 | 957.6686 | 1046.9956 | 958.6775 | 946.6878 | 924.3314 | 972.9793 | 1576.6598 | 1170.0399 | 1080.3461 | 1142.3196 | 1243.9705 | 1050.0015 | 1138.0059 | 932.3225 | 1849.3343 | 2331.6406 | 2438.3269 | 1936.2767
219 | 4.4034 | 766.9725 | 1719.6794 | 1585.704 | 2014.0354 | 1839.7099 | 2119.4494 | 1567.5782 | 1546.3972 | 975.1042 | 1277.1396 | 1052.3589 | 928.2193 | 1941.8328 | 1474.3618 | 1898.3697 | 1822.8741 | 1781.3972 | 1394.296 | 1433.7945 | 1293.589 | 1375.9233 | 2231.1534 | 2234.5261 | 1640.7149 | 2115.178 | 2046.0629 | 1691.5536 | 2686.877 | 1356.1288 | 1451.4739 | 1407.3589 | 1567.1396 | 1207.0738 | 1647.8466 | 1760.9371 | 1378.8083 | 2878.589 | 2907.5261 | 2539.0246 | 2867.4769 | 1557.2684 | 1981.0688 | 1709.9617 | 1690.0738 | 1017.2576 | 1048.2439 | 1062.4602 | 1544.9233 | 1362.9262 | 1904.6028 | 1472.6411 | 1300.8987 | 1495.7807 | 2011.115 | 1431.8712 | 1376.0108 | 1521.2301 | 2132.4464 | 1329.4602 | 1433.0492 | 1137.5782 | 1443.9479 | 1071.7974 | 1001.0275 | 1627.2576 | 1961.9725 | 1482.1809 | 1729.4631 | 1339.7316 | 1861.3972 | 1539.4277 | 1847.6057 | 1761.6411 | 1832.0521 | 1724.6028 | 2580.2605 | 1535.6932 | 1535.529 | 1350.0383 | 2251.9371 | 1134.6686 | 1196.5369 | 1167.5673 | 1529.1888 | 1067.5506 | 1097.3451 | 1080.2822 | 1056.178 | 1264.6107 | 1191.0767 | 1147.9754 | 1535.822 | 1511.8604 | 1721.1042 | 2025.2055 | 3265.9037 | 2956.7395 | 3174.4985 | 3163.6332 | 4702.3972 | 1007.2439 | 922.3451 | 918.8328 | 1007.2439 | 1091.3372 | 1198.8437 | 1049.3343 | 1165.8358 | 997.9617 | 1094.9646 | 1110.5861 | 1011.0492 | 933.704 | 929.3618 | 945.411 | 1057.589 | 944.2576 | 929.0659 | 933.8712 | 975.9862 | 1554.3343 | 1144.0029 | 1080.1672 | 1139.7945 | 1262.0492 | 1067.9862 | 1131.1288 | 933.8328 | 1863.589 | 2409.8024 | 2400.707 | 1946.2409
220 | 4.4144 | 785.767 | 1754.587 | 1592.9248 | 1993.6534 | 1820.444 | 2137.5192 | 1539.0147 | 1574.9779 | 988.6696 | 1282.7743 | 1040.5487 | 930.8348 | 1877.5177 | 1489.4292 | 1908.3909 | 1807.1504 | 1715.2094 | 1361.4808 | 1388.5781 | 1247.4277 | 1371.9248 | 2211.4808 | 2256.5944 | 1612.3525 | 2087.1873 | 1999.969 | 1725.9558 | 2637.8496 | 1355.8422 | 1450.8496 | 1414.2257 | 1496.6828 | 1193.879 | 1638.4661 | 1737.1652 | 1351.8643 | 2864.2478 | 2926.1357 | 2513.413 | 2914.4823 | 1575.7596 | 1980.4292 | 1772.6402 | 1649.9011 | 951.6607 | 1059.1431 | 1054.6313 | 1592.1136 | 1351.2183 | 1862.1873 | 1509.354 | 1281.0074 | 1482.7891 | 2064.4971 | 1351.9159 | 1376.3982 | 1516.6239 | 2134.7065 | 1300.879 | 1444.0826 | 1174.0384 | 1475.9027 | 1075.5944 | 1056.2493 | 1611.4661 | 1926.1194 | 1531.5649 | 1725.6844 | 1294.9543 | 1851.9321 | 1455.0368 | 1869.4219 | 1737.4808 | 1739.2241 | 1740.9853 | 2650.8363 | 1519.6991 | 1457.1947 | 1348.233 | 2286.2714 | 1117.3083 | 1215.5339 | 1140 | 1563.8422 | 1058.4734 | 1082.4808 | 1103.9027 | 1060.6991 | 1264.1873 | 1179.8569 | 1145.2257 | 1534.7596 | 1552.5723 | 1709.767 | 2077.3393 | 3283.2625 | 2947.0368 | 3231.9632 | 3224.7611 | 4627.9085 | 982.1726 | 948.2935 | 939.6018 | 982.1726 | 1071.1431 | 1161.202 | 1029.9321 | 1216.941 | 1040.5797 | 1044.7891 | 1085.2773 | 1006.1726 | 925.7065 | 902.9248 | 945.2257 | 1010.0442 | 942.9926 | 938.2109 | 897.4956 | 934.8864 | 1532.8569 | 1089.2551 | 1116.121 | 1136.0826 | 1223.969 | 1042.7964 | 1128.0752 | 980.1962 | 1818.1194 | 2446.0074 | 2415.4366 | 1993.9027
221 | 4.4254 | 765.175 | 1729.8702 | 1595.5202 | 1979.3363 | 1839.1524 | 2152.2242 | 1536.6323 | 1576.4395 | 987.6234 | 1208.1887 | 1085.6903 | 897.5742 | 1870.2331 | 1504.1839 | 1800.8033 | 1789.2871 | 1741.3225 | 1403.8024 | 1360.6637 | 1269.058 | 1367.4798 | 2204.5202 | 2261.6637 | 1601.3766 | 2097.8653 | 2011.6097 | 1703.5339 | 2688.8024 | 1355.1839 | 1455.3363 | 1343.0767 | 1512.0984 | 1158.3992 | 1607.1347 | 1685.942 | 1347.6323 | 2926.8113 | 2955.0089 | 2512.7847 | 2978.7935 | 1553.5428 | 1957.1662 | 1764.3903 | 1598.879 | 969.9459 | 1105.8338 | 1061.4887 | 1589.3589 | 1380.529 | 1855.9685 | 1437.292 | 1248.1347 | 1483.2645 | 1965.9065 | 1381.5871 | 1398.5693 | 1539.4572 | 2105.879 | 1306.2331 | 1429.0315 | 1188.6637 | 1427.8613 | 1072.2153 | 1088 | 1651.2419 | 1924.121 | 1545.6234 | 1727.1839 | 1263.9194 | 1827.9911 | 1536.3806 | 1903.6726 | 1705.175 | 1759.4346 | 1752.5202 | 2674.0629 | 1514.5919 | 1503.2911 | 1355.4484 | 2292.6234 | 1086.0226 | 1196.879 | 1189.9863 | 1554.7355 | 1043.5516 | 1078.3363 | 1125.5605 | 1104.8742 | 1255.1524 | 1209.9056 | 1167.0089 | 1568.1613 | 1555.5831 | 1681.4307 | 2059.7581 | 3308.9459 | 2903.0629 | 3137.9468 | 3235.0944 | 4749.591 | 996.6411 | 938.3992 | 925.175 | 996.6411 | 1064.3274 | 1144.8564 | 963.7492 | 1167.3323 | 1043.8073 | 1079.7621 | 1010.0137 | 1079.7798 | 898.2871 | 905.5202 | 916.3186 | 956.4307 | 942.0718 | 907.4621 | 909.2734 | 930.6008 | 1521.3677 | 1071.7758 | 1072.0452 | 1108.3589 | 1183.5113 | 1061.7935 | 1111.3992 | 944.0855 | 1866.8113 | 2472.2734 | 2326.2113 | 2014.1524
222 | 4.4364 | 727.3471 | 1730.9081 | 1585.5388 | 2019.2488 | 1848.7306 | 2127.4961 | 1563.355 | 1539.0491 | 943.3186 | 1152.9287 | 1058.29 | 901.057 | 1862.0919 | 1482.2266 | 1810.0079 | 1784.5816 | 1771.4469 | 1371.5595 | 1394.5325 | 1256.4818 | 1411.8447 | 2104.9985 | 2188.084 | 1629.206 | 2073.4675 | 1988.3186 | 1694.7448 | 2735.355 | 1367.6101 | 1423.6165 | 1365.0079 | 1526.8368 | 1181.8019 | 1600.0143 | 1699.2345 | 1348.5674 | 2941.2409 | 2926.4675 | 2478.1696 | 2958.6165 | 1533.4041 | 1939.2837 | 1701.7227 | 1610.7734 | 965.794 | 1047.7861 | 1030.0349 | 1573.2694 | 1335.5595 | 1838.0777 | 1389.2552 | 1234.5674 | 1492.2979 | 1937.206 | 1317.5943 | 1364.439 | 1572.4897 | 2019.6371 | 1288.7797 | 1424 | 1196.7448 | 1445.2631 | 1007.2615 | 1080.8368 | 1613.1268 | 1918.6735 | 1514.2124 | 1668.6942 | 1262.2979 | 1892.3486 | 1523.1268 | 1876.7655 | 1752.5895 | 1790.4469 | 1730.3613 | 2701.2488 | 1530.4754 | 1471.1126 | 1338.943 | 2315.6529 | 1103.8019 | 1200.1775 | 1229.6244 | 1537.1062 | 1080.1141 | 1046.6165 | 1099.1981 | 1064.1126 | 1208.1411 | 1212.1062 | 1136.3186 | 1580.7163 | 1567.7734 | 1697.6529 | 2114.3486 | 3316.2552 | 2948.8732 | 3091.4326 | 3175.588 | 4809.9857 | 955.4247 | 903.0634 | 903.1062 | 955.4247 | 1033.2124 | 1203.4405 | 980.9794 | 1144.4612 | 986.8574 | 1070.3471 | 1021.2345 | 1018.2964 | 918.6529 | 919.8938 | 915.8938 | 1019.2139 | 932.8368 | 915.0713 | 904.0919 | 949.4612 | 1563.412 | 1115.8447 | 1081.0998 | 1126.6529 | 1246.4833 | 1073.851 | 1102.5674 | 930.3122 | 1849.7227 | 2373.2536 | 2353.916 | 2055.2488
223 | 4.4474 | 737.8584 | 1743.1003 | 1537.2832 | 2004.8554 | 1809.469 | 2074.7404 | 1567.8997 | 1536.2006 | 961.5162 | 1202.7463 | 1034.6283 | 955.6313 | 1889.9735 | 1503.1446 | 1775.6253 | 1801.3009 | 1751.413 | 1371.8584 | 1355.8112 | 1250.5575 | 1379.7256 | 2148.7906 | 2184.0443 | 1610.3127 | 2068.472 | 1991.2153 | 1691.7139 | 2768.3156 | 1334.3835 | 1418.7434 | 1368.5133 | 1507.9705 | 1172.1416 | 1583.056 | 1699.6136 | 1345.8142 | 2816.9645 | 2959.3599 | 2510.8171 | 2996.7316 | 1588.2891 | 1960.8584 | 1658.9557 | 1642.3156 | 951.8997 | 1039.8584 | 967.6401 | 1578.2861 | 1383.2183 | 1836.2861 | 1412.6873 | 1258.7729 | 1554.475 | 1986.0738 | 1346.5753 | 1350.3717 | 1564.9705 | 2050.2036 | 1288.2006 | 1358.4247 | 1144.8967 | 1494.7876 | 1018.9735 | 1046.6696 | 1566.9557 | 1927.5723 | 1518.5723 | 1705.7463 | 1316.4602 | 1829.9233 | 1565.118 | 1874.2006 | 1788.8437 | 1717.9528 | 1707.1563 | 2677.3835 | 1528.9852 | 1477.6018 | 1349.944 | 2312.4277 | 1123.2861 | 1182.413 | 1227.528 | 1512.3274 | 1044.997 | 1086.1888 | 1098.6578 | 1080.8024 | 1204.7581 | 1151.6253 | 1078.0118 | 1567.8849 | 1550.4985 | 1718.472 | 2107.2979 | 3414.7346 | 2899.5958 | 3143.003 | 3155.0855 | 4727.8525 | 1019.5192 | 915.3156 | 906.7434 | 1019.5192 | 1030.2861 | 1157.9675 | 1008.2006 | 1155.5575 | 983.7581 | 1096.1593 | 1045.6578 | 984.3717 | 881.1829 | 911.1563 | 922.4572 | 1032.9705 | 932.9145 | 885.6548 | 940.7168 | 942.3422 | 1540.9262 | 1128.1711 | 1111.7581 | 1126.3422 | 1252.8702 | 1083.9292 | 1081.5988 | 931.0855 | 1879.0885 | 2379.2744 | 2377.6283 | 1999.3244
224 | 4.4584 | 735.6273 | 1730.1819 | 1580.2724 | 2025.118 | 1828.0816 | 2103.4543 | 1519.323 | 1511.8456 | 961.4046 | 1214.6273 | 1017.441 | 943.3638 | 1868.1775 | 1454.7276 | 1756.4818 | 1784.7728 | 1817.8314 | 1371.85 | 1330.8137 | 1295.2316 | 1292.0958 | 2135.5457 | 2238.1952 | 1639.268 | 2006.8004 | 1953.5457 | 1628.9867 | 2775.5182 | 1389.6451 | 1422 | 1336.8092 | 1572.386 | 1192.1544 | 1665.6495 | 1716.4862 | 1331.6637 | 2785.5954 | 2838.4508 | 2475.6548 | 2921.9051 | 1574.4366 | 1946.0319 | 1618.882 | 1678.4135 | 969.2635 | 1070.9681 | 987.7862 | 1623.2316 | 1387.5138 | 1856.7817 | 1400.7728 | 1276.8909 | 1496.2414 | 2018.0408 | 1368.7773 | 1392.7498 | 1559 | 2072.2591 | 1218.6504 | 1361.4135 | 1166.9725 | 1422.5368 | 1017.1819 | 1069.8225 | 1535.0319 | 1940.4091 | 1538.0772 | 1715.4046 | 1318.2183 | 1832.3043 | 1506.7595 | 1923.9361 | 1768.0683 | 1732.2316 | 1687.9228 | 2694.118 | 1566.7498 | 1474.9228 | 1361.4454 | 2253.6548 | 1133.4091 | 1186.5954 | 1185.4729 | 1516.1136 | 974.1366 | 1048.2094 | 1050.3957 | 1041.9139 | 1185.0683 | 1121.3682 | 1093.8589 | 1511.1366 | 1504.6957 | 1710.1455 | 2058.4366 | 3396.6912 | 2867.0728 | 3141.9956 | 3176.2272 | 4849.8855 | 1004.323 | 877.323 | 894.441 | 1004.323 | 1026.3319 | 1136.7045 | 1008.2955 | 1154.2955 | 972.2183 | 1040.3549 | 1050 | 1017.8589 | 839.291 | 892.6637 | 912.6273 | 974.3957 | 907.3274 | 900.6362 | 927.3638 | 928.3682 | 1532.6681 | 1065.2458 | 1071.323 | 1092.3638 | 1185.7276 | 1118.3771 | 1108.8225 | 931.4818 | 1908.1136 | 2388.3682 | 2350.8092 | 2003.3771
225 | 4.4694 | 735.766 | 1720.4808 | 1588.9489 | 2113.5792 | 1812.1317 | 2112.2212 | 1499.234 | 1566.7749 | 915.5742 | 1228.3235 | 1037.6087 | 937.5192 | 1848.468 | 1460.3619 | 1753.9872 | 1776.2468 | 1731.0206 | 1307.9911 | 1363.647 | 1267.0678 | 1319.2173 | 2147.1151 | 2251.7404 | 1594.2507 | 2028.4386 | 1948.5447 | 1653.945 | 2756.1701 | 1359.7699 | 1394.8849 | 1362.647 | 1515.6548 | 1156.8082 | 1620.9272 | 1749.6087 | 1331.5192 | 2837.4514 | 2883.3963 | 2514.9961 | 2929.6726 | 1516.8082 | 1910.8849 | 1624.5703 | 1666.1573 | 954.9105 | 1033.8082 | 988.1957 | 1590.5486 | 1453.06 | 1919.4897 | 1401.2852 | 1311.6342 | 1431.0934 | 1987.8466 | 1366.4936 | 1375.3658 | 1507.7827 | 2053.9233 | 1234.4258 | 1406.4002 | 1180.4936 | 1427.647 | 1014.2596 | 1121.1662 | 1563.1534 | 1864.6676 | 1544.5064 | 1713.2596 | 1342.1278 | 1848.766 | 1435.0678 | 1811.9567 | 1770.2979 | 1743.9872 | 1664.1701 | 2753.9577 | 1526.7699 | 1433.0934 | 1297.4719 | 2313.3324 | 1111.8977 | 1122.4592 | 1141.8722 | 1553.1406 | 1023.7876 | 1009.9233 | 1046.3107 | 1050.8683 | 1197.0895 | 1144.3619 | 1072.6126 | 1550.2429 | 1543.9705 | 1728.5959 | 2042.234 | 3379.2596 | 2787.4208 | 3142.532 | 3165.9233 | 4911.2852 | 966.9105 | 812.0423 | 888.2468 | 966.9105 | 1058.647 | 1119.9233 | 1033.3619 | 1108.5614 | 996.881 | 994.9105 | 1013.0934 | 998.8722 | 844.5575 | 906.0767 | 870.587 | 951.4808 | 873.1445 | 912.5064 | 886.119 | 922.7404 | 1542.532 | 1085.4258 | 1079.1023 | 1124.179 | 1186.8427 | 1037.1101 | 1083.3402 | 902.6254 | 1907.4808 | 2390.2724 | 2393.2173 | 2003.4553
226 | 4.4804 | 759.1991 | 1679.9602 | 1536.0088 | 2052.5 | 1722.6239 | 2159.7567 | 1486.5177 | 1496.1415 | 924.9646 | 1256.1991 | 1049.0619 | 885.4159 | 1854.4204 | 1472.8274 | 1732.3318 | 1821.9292 | 1698.5929 | 1326.2744 | 1324.4778 | 1216.2566 | 1309.9734 | 2208.8186 | 2144.5973 | 1570.4071 | 2043.7655 | 1988.7434 | 1672.8274 | 2743.1106 | 1346.5929 | 1399.7788 | 1331.8982 | 1465.2699 | 1142 | 1629.6814 | 1709.7743 | 1319.2212 | 2810.8982 | 2891.8009 | 2458.9823 | 3006.5354 | 1520.3717 | 1876.7389 | 1628.469 | 1626.8495 | 905.1327 | 1086.3629 | 1005.4336 | 1606.385 | 1359.8805 | 1853.0309 | 1442.2744 | 1313.5796 | 1449.6195 | 1986.7168 | 1346.0973 | 1397.1505 | 1503.2478 | 2034.9867 | 1284.7434 | 1466.2257 | 1112.6371 | 1458.9027 | 995.0973 | 1081.1814 | 1524.4778 | 1884.9292 | 1518.208 | 1719.3584 | 1319.8495 | 1849.2345 | 1445.4956 | 1906.2567 | 1685.969 | 1776.6814 | 1714.646 | 2659.7699 | 1502.8761 | 1473.0531 | 1262.9867 | 2341.2965 | 1072.615 | 1162.0045 | 1136.5929 | 1554.8142 | 1061.3097 | 1018.5443 | 1073.9646 | 1028.615 | 1201.7035 | 1139.2212 | 1105.6328 | 1549.8628 | 1551.8142 | 1722.2212 | 2066.2611 | 3378.469 | 2954.1373 | 3146.531 | 3174.5443 | 4868.0708 | 946.2212 | 835.8053 | 916.323 | 946.2212 | 1051.6283 | 1104.0487 | 1003.7256 | 1096.2965 | 962.6017 | 1004.0752 | 971.9115 | 999.7168 | 877.323 | 890.5664 | 884.0266 | 959.9513 | 890.2611 | 861.7124 | 924.2257 | 906.6903 | 1554.1858 | 1041.5885 | 1039.5398 | 1054.5619 | 1195.531 | 1046.0398 | 1116.1637 | 918.2611 | 1930.4956 | 2308.5619 | 2400.0487 | 2075.0177
227 | 4.4914 | 755.8889 | 1673.4444 | 1560.6667 | 2001.8889 | 1780.4445 | 2189 | 1494.6667 | 1566.0001 | 890.7778 | 1250.5555 | 999.4444 | 894.6667 | 1918.4445 | 1412.5555 | 1734.5556 | 1790.8889 | 1739.6667 | 1351.2222 | 1318.3333 | 1259.4445 | 1318.3333 | 2205.5555 | 2138.4445 | 1643.6667 | 2098.4445 | 1967.3333 | 1667.7778 | 2783.7778 | 1348 | 1344.1111 | 1363.3334 | 1508.5556 | 1134.2222 | 1554.4444 | 1705.1111 | 1301.7778 | 2845.4445 | 2847.6666 | 2503.6667 | 3000.3333 | 1449 | 1883 | 1686.3334 | 1632.5556 | 913 | 1051 | 1026.7778 | 1535.1111 | 1377.6667 | 1802.8889 | 1414.3333 | 1272.1111 | 1429.4444 | 1930.1111 | 1294.1111 | 1375.8889 | 1543.5556 | 2034.1111 | 1272.6667 | 1424.2222 | 1088.1111 | 1450.2222 | 952.4444 | 1041.6667 | 1597.6667 | 1911.4445 | 1495.2222 | 1687.5555 | 1357.4445 | 1830.3333 | 1491.8889 | 1895.8889 | 1735.2223 | 1825.1111 | 1700.1111 | 2775.3334 | 1525.8889 | 1531 | 1304.1111 | 2365.3333 | 1060.6667 | 1119.5555 | 1116.2222 | 1593.2222 | 1016.2222 | 1025.3333 | 1082.5556 | 988.6667 | 1190.1111 | 1114 | 1088.6667 | 1498.8889 | 1603.4445 | 1647.2222 | 1973.6666 | 3470.5556 | 2908.7777 | 3094.1111 | 3066.2222 | 4917 | 920.2222 | 838.6667 | 858.8889 | 920.2222 | 992.3333 | 1080 | 993.5556 | 1157.6667 | 928 | 1019.1111 | 996.4445 | 989.7778 | 913.2222 | 894.8889 | 872.5555 | 979.3333 | 866.1111 | 870.8889 | 838.6666 | 901.2222 | 1567.8889 | 1065.2222 | 1034.3333 | 1066.3334 | 1206.1111 | 1053.3333 | 1074.7777 | 919 | 1898.3333 | 2370.1112 | 2234.4444 | 2050.5555
228 | 4.5024 | 764.061 | 1686.851 | 1534.0767 | 2020.2822 | 1763.7065 | 2128.2143 | 1513.0113 | 1514.203 | 890.061 | 1211.4469 | 992.9007 | 894.0993 | 1840.4016 | 1412.3815 | 1794.6751 | 1824.6141 | 1793.6637 | 1289.4243 | 1336.8014 | 1312.1333 | 1351.4425 | 2182.6185 | 2191.0339 | 1633.3972 | 2050.7945 | 1942.9887 | 1680.2212 | 2727.264 | 1319.5575 | 1369.6637 | 1304.8938 | 1495.3476 | 1099.6072 | 1565.5531 | 1683.5079 | 1279.0383 | 2840.0383 | 2882.0339 | 2462.9547 | 2971.6681 | 1502.0566 | 1911.4921 | 1710.9007 | 1573.0044 | 896.088 | 969.4243 | 971.7448 | 1608.3884 | 1313.203 | 1822.0723 | 1477.2665 | 1231.7674 | 1492.8963 | 1947.763 | 1364.7473 | 1384.3815 | 1492.3746 | 1996.2866 | 1279.4312 | 1341.6342 | 1081.4695 | 1446 | 1002.0453 | 1057.1219 | 1577.9661 | 1903.939 | 1457.8171 | 1715.1333 | 1350.4582 | 1861.3432 | 1518.0113 | 1862.8284 | 1715.9164 | 1709.5889 | 1664.1376 | 2755.5845 | 1499.027 | 1514.5575 | 1256.1647 | 2308.0044 | 1108.1942 | 1076.7178 | 1171.6254 | 1541.7945 | 1009.1106 | 1048.1219 | 1094.851 | 983.9503 | 1160.1376 | 1098.5192 | 1089.4808 | 1580.8083 | 1522.6116 | 1699.6864 | 2048.4991 | 3504.1106 | 2912.973 | 3074.2596 | 3161.9912 | 4989.5148 | 925.061 | 872.3432 | 846.3206 | 925.061 | 983.3206 | 1077.1603 | 1032.7134 | 1116.5349 | 942.5418 | 988.027 | 1010.7404 | 1008.912 | 888.6072 | 885.3589 | 883.5915 | 958.7178 | 912.3048 | 846.3972 | 825.8014 | 885.1986 | 1554.4085 | 1072.2596 | 1052.8014 | 1152.8466 | 1195.5688 | 1038.5688 | 1064.5305 | 917 | 1849.0767 | 2405.2709 | 2376.0365 | 2087.5644
229 | 4.5134 | 750.9528 | 1618.5899 | 1511.1622 | 1993.3009 | 1778.6755 | 2046.0088 | 1530.6401 | 1513.059 | 925.8968 | 1220.4543 | 957.2891 | 887.9882 | 1823.0236 | 1402.5575 | 1727.3569 | 1804.1032 | 1785.7434 | 1319.3393 | 1321.7434 | 1278.6725 | 1314.8702 | 2139.4867 | 2165.2891 | 1698.0148 | 2058.8732 | 1955.0472 | 1643.882 | 2672.0914 | 1380.2006 | 1418.7109 | 1298.0354 | 1556.2006 | 1131.9204 | 1535.6843 | 1728.3511 | 1305.2921 | 2848.6401 | 2897.2094 | 2455.826 | 2991.6755 | 1538.8496 | 1828.938 | 1587.0294 | 1608.1534 | 918.885 | 983.2921 | 988.8968 | 1597.0914 | 1429.2036 | 1780.056 | 1480.9528 | 1215.9764 | 1427.7522 | 1932.1268 | 1307.5663 | 1365.7316 | 1496.8614 | 1964.3363 | 1315.8968 | 1383.3747 | 1118.5133 | 1428.3481 | 986.7079 | 992.3923 | 1553.9646 | 1880.941 | 1485.9086 | 1716.174 | 1302.4749 | 1776.3215 | 1558.1062 | 1914.9676 | 1717.8378 | 1765.8644 | 1656.1976 | 2786.3393 | 1458.7079 | 1466.882 | 1228.767 | 2297.4189 | 1052.7994 | 1118.1416 | 1126.2183 | 1447.1711 | 1011.8024 | 1039.5575 | 1040.2301 | 1023.5133 | 1195.5251 | 1124.885 | 1055.0914 | 1586.7434 | 1577.236 | 1706.9646 | 1994.9616 | 3512.4189 | 2946.0708 | 3068.9882 | 3021.0855 | 4838.7138 | 896.7079 | 860.3363 | 855.2212 | 896.7079 | 983.3953 | 1113.3039 | 1003.6843 | 1134.6991 | 980.0944 | 1055.2242 | 963.056 | 1003.767 | 851.115 | 896.4425 | 871.9528 | 959.2212 | 925.3953 | 893.3629 | 846.8496 | 860.941 | 1546.7906 | 1088.6519 | 1063.4189 | 1122.4513 | 1179.1622 | 1057.4661 | 1049.9528 | 904.9646 | 1877.5133 | 2261.1681 | 2397.9174 | 2060.882
230 | 4.5244 | 772.2542 | 1646.6205 | 1533.2542 | 2030.6205 | 1786.6293 | 2038.1465 | 1509.5605 | 1533.4784 | 880.233 | 1184.2674 | 1001.767 | 902.4784 | 1824.1853 | 1426.6981 | 1740.4007 | 1814.737 | 1743.526 | 1322.8535 | 1334.1077 | 1217.4916 | 1321.1077 | 2139.7758 | 2182.8835 | 1669.5649 | 2150.5384 | 1931.3019 | 1677.1765 | 2689.8835 | 1355.4528 | 1349.7935 | 1287.1509 | 1491.5349 | 1171.3274 | 1597.6893 | 1702.7114 | 1320.3319 | 2815.3407 | 2941.1765 | 2441.707 | 2967.0432 | 1537.2974 | 1903.9435 | 1658.6461 | 1608.7802 | 914.5949 | 1005.6637 | 982.3363 | 1605.1077 | 1382.3495 | 1768.8147 | 1407.7546 | 1224.5905 | 1503.9435 | 1940.6637 | 1361.3186 | 1335.5605 | 1538.6593 | 1987.5128 | 1320.7414 | 1425.6981 | 1097.8579 | 1428.0733 | 996.9223 | 1033.6549 | 1625.1333 | 1852.3751 | 1459.97 | 1745.7714 | 1304.2198 | 1748.9267 | 1516.4916 | 1913.5216 | 1708.7802 | 1760.3019 | 1658.444 | 2732.0132 | 1555.2754 | 1443.1509 | 1250.4395 | 2320.9956 | 989.5649 | 1118.5949 | 1144.4007 | 1499.6893 | 1009.556 | 1027.8535 | 1041.4784 | 1008.5605 | 1170.97 | 1116.5216 | 1077.7714 | 1555.9312 | 1566.8579 | 1705.6293 | 1988.8491 | 3536.8102 | 2959.3319 | 3121.767 | 3088.7926 | 4998.4042 | 963.1333 | 874.5516 | 841.5216 | 963.1333 | 1046.5428 | 1079.6381 | 981.7802 | 1156.0344 | 928.5305 | 1003.1253 | 948.556 | 1002.6293 | 850.5172 | 818.5349 | 840.3019 | 930.8579 | 847.6081 | 866.2674 | 859.9612 | 908.1377 | 1554.1465 | 1080.7802 | 1079.6637 | 1139.3274 | 1219.9912 | 1053.2242 | 1109.5428 | 924.8102 | 1850.3407 | 2280.3963 | 2296.5004 | 2044.4828
231 | 4.5354 | 808.5969 | 1614.8259 | 1561.3274 | 2053.4041 | 1763.8456 | 2038.3461 | 1502.3461 | 1544.4425 | 915.1357 | 1119.5949 | 968.5182 | 883.6726 | 1819.8653 | 1495.0207 | 1739.7306 | 1805.7689 | 1742.4425 | 1383.8476 | 1290.8643 | 1242.3855 | 1350.6352 | 2189.1357 | 2160.057 | 1609.3835 | 2123.6912 | 1886.9606 | 1692.7886 | 2796.887 | 1325.3264 | 1346.058 | 1298.885 | 1506.9046 | 1177.1731 | 1523.8633 | 1680.7689 | 1302.1534 | 2864.059 | 2791.4001 | 2504.1554 | 2992.2508 | 1492.1721 | 1843.3058 | 1642.3068 | 1571.2684 | 911.1731 | 1009.8269 | 972.3845 | 1624.7119 | 1356.5575 | 1795.4622 | 1353.1337 | 1198.057 | 1448.2674 | 1996.0974 | 1326.7296 | 1385.4051 | 1548.6539 | 1975.8073 | 1267.0757 | 1407.1917 | 1100.0964 | 1350.4405 | 952.2104 | 1025.4995 | 1617.1917 | 1820.5378 | 1494.1741 | 1742.2497 | 1335.943 | 1802.5782 | 1492.5959 | 1883.5378 | 1811.8486 | 1749.8653 | 1654.0383 | 2769.0974 | 1517.8063 | 1440 | 1297.3471 | 2328.4808 | 1041.8476 | 1087.057 | 1118.9223 | 1490.3648 | 953.5949 | 1010.2881 | 1050.7886 | 994.7306 | 1228.3284 | 1098.942 | 1040.3451 | 1598.7896 | 1549.2497 | 1677.884 | 2059.8093 | 3504.7876 | 2904.7679 | 3150.5388 | 3104.0767 | 5005.1337 | 894.8249 | 823.5949 | 818.1534 | 894.8249 | 996.4985 | 1061.7306 | 938.4798 | 1137.4995 | 902.5959 | 981.2114 | 967.8466 | 1000.5192 | 864.4041 | 830.5969 | 846.4041 | 990.1554 | 881.2704 | 834.6726 | 857.8653 | 860.4601 | 1436.9203 | 1082.9617 | 1103.6736 | 1160.885 | 1136.5555 | 1036.115 | 1078.5182 | 920.9036 | 1806.4415 | 2260.7492 | 2387.4642 | 2037.2114
232 | 4.5464 | 788.1445 | 1585.8584 | 1535.7876 | 1954.4528 | 1766.5531 | 2011.9838 | 1533.0516 | 1596.354 | 911.5723 | 1155.8363 | 972.4277 | 856.3407 | 1811.4469 | 1433.3082 | 1795.0679 | 1797.9646 | 1686.0929 | 1352.5206 | 1333.0325 | 1250.8392 | 1273.7551 | 2171.3053 | 2205.1932 | 1589.7684 | 2065.3245 | 1941.1033 | 1670.5015 | 2786.9484 | 1329.3923 | 1366.6239 | 1291.9292 | 1565.1932 | 1146.7876 | 1598.9941 | 1624.2891 | 1264.4307 | 2901.177 | 2800.8009 | 2452.5398 | 2975.5015 | 1488.0354 | 1881.3894 | 1620.5723 | 1640.3702 | 853.0929 | 998.2507 | 959.2507 | 1594.4307 | 1356.6785 | 1763.2345 | 1353.3923 | 1213.8201 | 1466.4086 | 1942.379 | 1293.3407 | 1381.8938 | 1494.4498 | 2034.1033 | 1255.1608 | 1412.91 | 1109.5531 | 1393.5856 | 941.1608 | 1013.4469 | 1601.9292 | 1883.8171 | 1412.3628 | 1682.0929 | 1319.3407 | 1840.0162 | 1484.1254 | 1929.1932 | 1705.5973 | 1670.9513 | 1661.3923 | 2814.4086 | 1503.3215 | 1489.5148 | 1265.7168 | 2356.6947 | 1044.9292 | 1122.9617 | 1080.9484 | 1467.3761 | 959.6239 | 975.1091 | 1037.733 | 999.7139 | 1232.7684 | 1081.733 | 1050.3024 | 1627.4631 | 1608.264 | 1700.5339 | 2038.7522 | 3587.6372 | 2895.5177 | 3033.9896 | 3002.9705 | 5082.6018 | 939.9425 | 826.267 | 769.5206 | 939.9425 | 962.0192 | 1088.5339 | 948.4631 | 1128.8038 | 913.4277 | 995.1062 | 982.91 | 1025.177 | 861.1254 | 810.1445 | 864.9454 | 946.9321 | 855.0737 | 856.0162 | 802.4498 | 855.5531 | 1572.2936 | 1037.8422 | 1024.9159 | 1116.1637 | 1174.0679 | 1037.1962 | 1080.7493 | 932.4277 | 1836.7655 | 2300.1578 | 2266.6518 | 2040.1962
233 | 4.5574 | 740.4238 | 1627.1308 | 1482.3146 | 2049.9558 | 1739.0482 | 1983.1573 | 1487.9085 | 1522.5506 | 946.4671 | 1159.0393 | 929.0177 | 868.179 | 1838.1003 | 1487.9735 | 1697.0049 | 1836.1701 | 1705.6549 | 1320.1573 | 1360.882 | 1258.6637 | 1292.9125 | 2114.057 | 2170.7207 | 1629.7247 | 1984.4719 | 1987.9125 | 1674.5152 | 2744.3844 | 1331.8515 | 1297.6204 | 1324.9125 | 1563.0787 | 1080.5418 | 1575.2753 | 1594.4149 | 1347.4101 | 2821.6991 | 2777.4936 | 2418.7119 | 3033.1613 | 1446.4238 | 1911.5851 | 1606.0787 | 1614.1268 | 891.3884 | 984.9302 | 889.7296 | 1627.1701 | 1369.7729 | 1838.5979 | 1401.8338 | 1207.7817 | 1477.2576 | 1916.3756 | 1275.3756 | 1310.0265 | 1515.5064 | 1939.5594 | 1216.6814 | 1363.9085 | 1100.7817 | 1428.8427 | 933.3363 | 976.236 | 1559.9783 | 1936.1701 | 1458.8643 | 1718.6854 | 1293.7119 | 1808.3844 | 1429.354 | 1868.3235 | 1774.6677 | 1727.5281 | 1664.703 | 2727.3323 | 1504.703 | 1442.7994 | 1237.5634 | 2382.4366 | 1039.5939 | 1154.8427 | 1144.825 | 1464 | 1003.8732 | 944.3058 | 1043.5152 | 990.7817 | 1230.1485 | 1115.6156 | 1072.7335 | 1579.9085 | 1578.2753 | 1723.0305 | 2016.6726 | 3594.7817 | 2930.061 | 2934.3628 | 3035.0915 | 5014.6991 | 945.1878 | 804.3058 | 807.8338 | 945.1878 | 992.9125 | 1059.939 | 969.0305 | 1119.4848 | 916.8515 | 1003.1091 | 970.5241 | 1000.1966 | 821.6814 | 805 | 813.5024 | 916.4149 | 859.0698 | 878.0305 | 812.4366 | 898.7247 | 1579.1573 | 1013.6726 | 1020.9213 | 1090.821 | 1106.6597 | 1003.0875 | 1025.0963 | 883.9085 | 1827.821 | 2356.6854 | 2300.3098 | 2021.4149
234 | 4.5684 | 720.0428 | 1640.1829 | 1548.2881 | 1990.529 | 1730.5447 | 2078.3388 | 1482.4553 | 1548.5993 | 921.9026 | 1145.0428 | 956.4164 | 821.7625 | 1822.2684 | 1426.5757 | 1655.5408 | 1766.1204 | 1723.0934 | 1347.0974 | 1308.7158 | 1291.0974 | 1312.821 | 2085.1322 | 2147.4513 | 1647.3658 | 2037.65 | 1915.3928 | 1619.852 | 2748.3658 | 1332 | 1353.2414 | 1272.2605 | 1586.9145 | 1120.9651 | 1501.4862 | 1714.5256 | 1302.5329 | 2742.2138 | 2818.646 | 2397.7237 | 3056.9572 | 1519.3348 | 1852.8053 | 1610.0467 | 1653.646 | 892.2257 | 976.0895 | 889.2296 | 1588.2178 | 1319.3073 | 1819.3112 | 1394.4513 | 1320.8875 | 1449.4941 | 1922.6382 | 1297.1868 | 1361.0585 | 1460.5329 | 1978.0118 | 1273.0585 | 1429.2881 | 1103.3191 | 1344.027 | 966.5526 | 1052.9263 | 1535.7237 | 1890.3073 | 1419.8987 | 1727.8638 | 1318.5059 | 1754.4901 | 1413.0895 | 1910.4204 | 1655.381 | 1701.9921 | 1681.4125 | 2779.2414 | 1485.1322 | 1445.0934 | 1261.6421 | 2418.825 | 1064.9145 | 1107.1711 | 1103.4434 | 1483.8678 | 942.7586 | 984.0546 | 1037.2257 | 1008.0039 | 1205.8132 | 1016.6145 | 1090.6849 | 1580.502 | 1441.4277 | 1700.0855 | 2083.1052 | 3683.7006 | 2911.8132 | 3025.5217 | 3115.6967 | 5068.2414 | 929.3151 | 808.6382 | 802.179 | 929.3151 | 973.0855 | 1066.0934 | 875.2527 | 1120.5914 | 868.6263 | 984.9961 | 930.8559 | 960.4474 | 821.0467 | 798.9533 | 814.3658 | 909.5447 | 832.4941 | 798.9375 | 841.0507 | 876.6303 | 1554.2684 | 1055.9184 | 1017.8171 | 1127.7355 | 1142.2375 | 1026.3697 | 1024.502 | 860.3151 | 1872.5099 | 2276.8908 | 2216.8441 | 2090.5605
235 | 4.5794 | 740.7788 | 1605.9557 | 1494.292 | 2034.0708 | 1751.9027 | 2100.1327 | 1474.2389 | 1606.6903 | 882.8318 | 1109.7079 | 966.3805 | 893.7257 | 1779.5752 | 1416.7522 | 1668.646 | 1841.6107 | 1733.7611 | 1316.9557 | 1342.0531 | 1289.8673 | 1339.531 | 2087.2566 | 2187.0531 | 1630.6018 | 2019.9646 | 1898.6195 | 1707.3717 | 2813.9558 | 1276.8053 | 1358.7434 | 1255.9911 | 1542.8141 | 1105.3451 | 1582.9912 | 1706.8407 | 1346.4425 | 2717.9823 | 2835.6372 | 2386.2389 | 2995.9203 | 1477.6814 | 1807.0796 | 1618.0089 | 1656.6195 | 860.3363 | 968.8673 | 929.5487 | 1615.1682 | 1374.0797 | 1790.3451 | 1315.7787 | 1254.0177 | 1461.646 | 1906.4779 | 1257.1947 | 1353.4779 | 1509.6991 | 1958.0885 | 1268.1062 | 1363.9026 | 1042.6725 | 1383.4425 | 966.7434 | 1060.7434 | 1552.2743 | 1880.2478 | 1506.6195 | 1720.115 | 1281.823 | 1783.7965 | 1391.8495 | 1894.469 | 1721.6107 | 1746.8142 | 1638.4425 | 2787.3717 | 1538.0708 | 1461.8938 | 1267.7522 | 2448.531 | 1059.3628 | 1059.5752 | 1077.4779 | 1474.7345 | 954.7788 | 960.3363 | 943.1327 | 999.6106 | 1177.469 | 1109.5133 | 1055.3274 | 1638.9469 | 1586.5842 | 1771.3452 | 2013.3982 | 3648.938 | 2942.1682 | 2974.5398 | 3042.0177 | 5105.2832 | 983.9469 | 735.5309 | 806.1327 | 983.9469 | 998.7876 | 1048.7257 | 922.2036 | 1094.7168 | 924.0797 | 1018.7965 | 987.2036 | 903.3097 | 817.6195 | 817.2743 | 774.823 | 908.1239 | 839.3894 | 804.3982 | 765.2743 | 836.0796 | 1562.3894 | 1102.4248 | 1070.4425 | 1110.3451 | 1152.6283 | 1000.3363 | 1073.3098 | 931.593 | 1820.4247 | 2312.062 | 2306.2567 | 2085.354
236 | 4.5904 | 765.3215 | 1622.3215 | 1512.5403 | 1882.0884 | 1757.6652 | 2041.4774 | 1509.4243 | 1515.3879 | 904.6519 | 1096.116 | 980.326 | 853.4729 | 1768.6696 | 1369.9149 | 1637.2409 | 1772.2674 | 1639.9414 | 1288.0133 | 1326.6785 | 1240.1381 | 1321.6785 | 2147.5226 | 2188.558 | 1707.9558 | 2014.2232 | 1936.2011 | 1622.3879 | 2831.884 | 1301.8707 | 1294.0354 | 1292.3127 | 1511.1249 | 1098.4464 | 1478.7315 | 1652.3613 | 1310.4685 | 2778.1878 | 2823.674 | 2430.3083 | 3002.1028 | 1515.3083 | 1784.232 | 1601.232 | 1608.9149 | 895.9779 | 978.6608 | 964.3171 | 1636.8796 | 1331.4729 | 1872.2862 | 1297.8928 | 1234.8928 | 1387.5978 | 1888.8972 | 1249.2232 | 1321.6829 | 1462.0309 | 2019.7414 | 1218.9149 | 1409.8575 | 1023.3392 | 1295.9414 | 956.2276 | 1009.3613 | 1568.326 | 1886.2188 | 1416.0575 | 1682.5757 | 1299.9867 | 1808.5447 | 1466.2906 | 1963.8486 | 1695.7989 | 1785.9823 | 1631.1116 | 2757.6829 | 1530.0088 | 1467.5536 | 1270.6652 | 2380.0398 | 1043.7856 | 1156.1657 | 1106.0939 | 1512.9779 | 915.3569 | 958.6652 | 969.9779 | 961.5757 | 1153.5669 | 1093.9061 | 1056.2188 | 1652.3304 | 1634.8751 | 1710.0398 | 2037.5359 | 3793.0276 | 2951.442 | 2982.768 | 2968.7006 | 5093.8972 | 910.2674 | 758.7591 | 780.3481 | 910.2674 | 975.4597 | 1069.0983 | 905.2365 | 1077.674 | 927.6696 | 963.589 | 937.3658 | 928.8707 | 792.1249 | 839.5447 | 828.5226 | 897.3392 | 801.9105 | 798.116 | 816.1878 | 835.442 | 1532.0177 | 1036.9282 | 1070.0044 | 1119.4376 | 1166.326 | 1015.5447 | 1004.4862 | 838.0575 | 1900.8398 | 2293.2365 | 2275.6917 | 2071.4508
237 | 4.6014 | 762.5959 | 1643.9145 | 1477.2704 | 2002.5074 | 1694.0511 | 2079.9351 | 1463.5614 | 1647.1102 | 930.5172 | 1117.5172 | 855.9036 | 885.7296 | 1762.5959 | 1399.1269 | 1592.469 | 1807.0344 | 1653.2193 | 1303.9145 | 1318.5959 | 1286.2399 | 1248.7463 | 2000.9822 | 2162.7807 | 1591.9036 | 2058.1337 | 1953.6096 | 1587.4828 | 2948.2882 | 1287.9862 | 1304.9145 | 1313.2124 | 1577.353 | 1119.6165 | 1517.1406 | 1533.4139 | 1266.2704 | 2795.0069 | 2860.7296 | 2372.9518 | 3033.7227 | 1471.4621 | 1827.0344 | 1572.88 | 1640.8289 | 806.4277 | 909.7463 | 958.0924 | 1592.1642 | 1312.3904 | 1792.0304 | 1378.9627 | 1233.0993 | 1386.1062 | 1923.0275 | 1284.1269 | 1379.2468 | 1482.0207 | 1886.5919 | 1226.5103 | 1458.4317 | 1049.9214 | 1345.2468 | 908.1642 | 1042.6303 | 1568.1986 | 1836.5614 | 1475.9558 | 1689.709 | 1275.0786 | 1739.8456 | 1394.0373 | 1984.709 | 1727.0275 | 1622.4719 | 1635.5034 | 2771.1131 | 1541.5103 | 1418.4621 | 1290.8152 | 2412.4317 | 1081.6303 | 1113.1573 | 1096.4897 | 1464.8594 | 957.7365 | 916.6676 | 968.6952 | 955.1986 | 1141.9931 | 1076.4897 | 1038.0855 | 1526.0029 | 1589.3628 | 1687.4897 | 2062.7158 | 3840.822 | 2955.6028 | 3000.3117 | 3062.7778 | 5146.942 | 880.1848 | 766.6028 | 758.9862 | 880.1848 | 949.4828 | 1039.5752 | 891.1917 | 1054.3835 | 900.88 | 936.1848 | 951.6165 | 964.5241 | 761.9793 | 798.7669 | 767.5477 | 892.3972 | 808.709 | 801.5034 | 778.0649 | 816.1848 | 1540.6096 | 1026.1986 | 1059.9931 | 1120.0993 | 1170.7021 | 1036.9145 | 1061.651 | 911.4661 | 1917.4041 | 2296.3048 | 2344.7572 | 2081.709
238 | 4.6124 | 739.1755 | 1632.3053 | 1498.5634 | 1983.3068 | 1730.823 | 2001.0059 | 1393.1784 | 1584.4837 | 876.3953 | 1134.6077 | 904.0826 | 804.1799 | 1779.3466 | 1389.3053 | 1586.174 | 1832.9985 | 1635.0015 | 1306 | 1327.1298 | 1271.0015 | 1307.6475 | 2087.1667 | 2145.3923 | 1658.3422 | 2065.7389 | 1866.441 | 1618.7803 | 2856.9203 | 1288.7389 | 1248.5693 | 1206.3554 | 1553.0457 | 1101.9144 | 1527.5649 | 1686.2493 | 1301.2581 | 2834.3452 | 2827.5678 | 2369.6519 | 3148.3835 | 1484.2596 | 1756.2227 | 1636.6475 | 1619.4366 | 851.6918 | 961.3437 | 931.4366 | 1536.7861 | 1338.3894 | 1786.5649 | 1411.7375 | 1210.1755 | 1403.4336 | 1868.5693 | 1274.3053 | 1340.351 | 1481.3481 | 1951.3422 | 1213.3923 | 1443.8274 | 1125.1254 | 1367.5207 | 977.8643 | 1013.2197 | 1598.1283 | 1834.6519 | 1471.2183 | 1756.7345 | 1237.3068 | 1784.9528 | 1479.9499 | 1956.7846 | 1783.9528 | 1675.2124 | 1644.2168 | 2835.9956 | 1544.826 | 1450.4322 | 1269.2625 | 2471.7788 | 1035.7861 | 1148.0841 | 1100.4779 | 1512.8658 | 983.9985 | 942.1283 | 957.0442 | 990.6062 | 1199.4307 | 1046.6976 | 1028.6962 | 1651.6859 | 1582.174 | 1756.2994 | 1971.8761 | 3810.2197 | 2857.3982 | 3091.472 | 3025.6121 | 5247.9499 | 851.5236 | 825.4307 | 774.3466 | 851.5236 | 894.0472 | 1033.2611 | 951.1696 | 1109.5177 | 883.3923 | 944.9558 | 923.8717 | 947.0015 | 790.0413 | 813.1726 | 790.1283 | 963.2124 | 815.4779 | 791.9572 | 856.0812 | 794.8274 | 1497.264 | 1006.8274 | 993.2655 | 1089.8717 | 1118.0472 | 983.3082 | 1011.4823 | 872.6121 | 1889.6976 | 2297 | 2256.2241 | 2103.9985
239 | 4.6234 | 770.3098 | 1643.9538 | 1478.7935 | 2017.9361 | 1754.4307 | 2018.9076 | 1508.2105 | 1541.8397 | 859.8967 | 1175.7866 | 908.1495 | 907.0325 | 1751.3913 | 1323.231 | 1713.6874 | 1805.3913 | 1646.8791 | 1291.1957 | 1268.7827 | 1271.7758 | 1247.3805 | 2014.6509 | 2083.8574 | 1573.4729 | 2074.3274 | 1908.8899 | 1642.356 | 2801.6617 | 1265.8682 | 1374.3885 | 1198.7758 | 1569.4307 | 1092.5978 | 1508.5693 | 1639.0816 | 1258.7365 | 2789.8859 | 2938.7336 | 2394.9823 | 3144.1957 | 1405.5014 | 1764.7296 | 1637.4484 | 1729.883 | 813.5123 | 942.943 | 939.178 | 1579.1888 | 1314.1671 | 1767.5693 | 1339.0531 | 1224.6549 | 1458.6657 | 1875.9538 | 1250.7935 | 1370.235 | 1460.644 | 2025.6195 | 1232.356 | 1453.1033 | 1061.6794 | 1365.2989 | 910.9783 | 935.0531 | 1515.8751 | 1841.4769 | 1455.1957 | 1741.7188 | 1266.3845 | 1784.4484 | 1490.8505 | 2015.0679 | 1714.0531 | 1715.3098 | 1595.9606 | 2838.2989 | 1516.3166 | 1448.4484 | 1269.7758 | 2372.4444 | 1010.644 | 1070.5762 | 1123.2065 | 1523.5516 | 926.7827 | 937.5978 | 977.2812 | 965.3166 | 1155.9606 | 1084.7119 | 1058.5339 | 1636.3166 | 1551.4661 | 1751.8967 | 1955.5978 | 3746.8574 | 2953.4809 | 2996.3697 | 2996.0177 | 5300.4877 | 861.9538 | 784.7365 | 768.5978 | 861.9538 | 943.591 | 1086.6657 | 894.0816 | 1054.8574 | 908.8329 | 917.3166 | 966.3383 | 900.587 | 768.0177 | 751.1563 | 754.1386 | 896.9036 | 827.1033 | 791 | 837.1671 | 754.1386 | 1610.5094 | 1076.2459 | 996.2527 | 1063.8682 | 1199.0502 | 982.6264 | 1042.0856 | 908.7119 | 1863.8682 | 2317.356 | 2180.3805 | 2152.2635
240 | 4.6344 | 771.1249 | 1603.7473 | 1470.4371 | 2090.3796 | 1764.4381 | 1992.8731 | 1444.8702 | 1488.3717 | 874.9386 | 1114.3082 | 881.7483 | 910.3746 | 1862.445 | 1360.1903 | 1632.1189 | 1744.8712 | 1697.6908 | 1199.0565 | 1260.2498 | 1251.3736 | 1280.44 | 2069.8791 | 2088.3756 | 1656.0059 | 2057.1863 | 1976.7542 | 1567.1199 | 2890.8186 | 1296.8147 | 1385 | 1288.0684 | 1524.1219 | 1055.4351 | 1503.2498 | 1557.1199 | 1265.3132 | 2783.1873 | 2759.5501 | 2371.6858 | 3099.8722 | 1483.3806 | 1873.8196 | 1555.4321 | 1589.9277 | 876.5669 | 938.1873 | 930.6244 | 1657.0674 | 1350.44 | 1829.7542 | 1356.439 | 1212.8741 | 1382.3697 | 1834.8097 | 1330.5679 | 1330.8097 | 1454.3122 | 1921.3053 | 1248.0634 | 1395.8712 | 1060.6878 | 1310.6214 | 902.1873 | 985.2537 | 1490.2488 | 1777.3082 | 1463.3756 | 1708.1229 | 1316.8157 | 1680.8682 | 1359.7413 | 1960.9336 | 1738.002 | 1688.9356 | 1570.4361 | 2886.7532 | 1469.9346 | 1438.6244 | 1282.1883 | 2410.8781 | 992.1239 | 1079.001 | 1088.4351 | 1476.1843 | 884.4975 | 892.9346 | 922.7463 | 956.4371 | 1127.6234 | 1113.3142 | 1042.2488 | 1576.8087 | 1601.5035 | 1663.8067 | 1970.001 | 3791.6908 | 2965.7502 | 2876.4301 | 2882.4301 | 5181.1794 | 861.1249 | 766.9366 | 737.998 | 861.1249 | 947.0624 | 1028.1834 | 873.0614 | 1085.6274 | 895.0614 | 920.6254 | 863.1179 | 905.4381 | 766 | 874.446 | 753.8127 | 935.0654 | 775.4965 | 742.2468 | 762.8077 | 833.5054 | 1484.9911 | 1089.5005 | 1025.1269 | 1103.2527 | 1096.3053 | 942.6848 | 1046.8751 | 865.1219 | 1877.9386 | 2321.8127 | 2240.6293 | 2095.0585
241 | 4.6454 | 778.5988 | 1578.2035 | 1440.5546 | 2019.0118 | 1704.2094 | 1920.7109 | 1364.9616 | 1488.7994 | 855.1032 | 1080.5546 | 899.9469 | 821.6637 | 1756.9675 | 1357.3009 | 1647.8466 | 1734.351 | 1655.4071 | 1280.3865 | 1243.8525 | 1270.8968 | 1246.9056 | 2045.5044 | 2051.0059 | 1611.6578 | 2060.7493 | 1866.0678 | 1592.3953 | 2877.0531 | 1269.5546 | 1277.6666 | 1287.351 | 1492.5044 | 1074.8466 | 1502.0501 | 1555.7994 | 1263.1504 | 2874.1859 | 2784.9941 | 2314.9085 | 3119.7965 | 1468.1032 | 1786.9646 | 1587.9941 | 1592.3481 | 858.2035 | 919.9528 | 911.9528 | 1611.6578 | 1281.761 | 1750.4129 | 1415.941 | 1217.6991 | 1486.2331 | 1800.6549 | 1280.9085 | 1314.7021 | 1468.2478 | 1893.1032 | 1228.1032 | 1419.5457 | 1025.8554 | 1289.9026 | 911.4985 | 1013.6962 | 1595.3835 | 1856.5871 | 1454.5015 | 1758.2419 | 1294.354 | 1783.2331 | 1456.4337 | 1921.8554 | 1717.2035 | 1611.0118 | 1585.1475 | 2843.4572 | 1481.2478 | 1486.4425 | 1229.8082 | 2498.5369 | 1020.4454 | 1084.7493 | 1076.5015 | 1454.9528 | 987.4337 | 945.0915 | 963.6431 | 959.7994 | 1182.9912 | 1056.1091 | 1078.0443 | 1617.6431 | 1628.7463 | 1671.2979 | 1978.5988 | 3832.0443 | 3016.3422 | 2994.3806 | 2965.236 | 5336.3747 | 892.3451 | 791.646 | 731.2507 | 892.3451 | 951.7493 | 982.2065 | 896.6962 | 1140.2419 | 864.5546 | 921.9499 | 904.4425 | 895.5516 | 804.944 | 827.9085 | 786.295 | 915.2035 | 791.9469 | 770.3451 | 750.4012 | 772.5103 | 1508.295 | 998.8141 | 1037.4484 | 1101.2507 | 1144.0915 | 968.4956 | 1029.9026 | 908.5428 | 1909.3953 | 2360.944 | 2245 | 2057.7551
242 | 4.6564 | 764.5678 | 1615.4858 | 1478.4479 | 2032.3186 | 1731.7886 | 2059.3973 | 1430.2744 | 1569.8201 | 888.6372 | 1074.1893 | 937.5615 | 898.7822 | 1685.5742 | 1307.9307 | 1649.9621 | 1773.4479 | 1615.4764 | 1321.5615 | 1226.6436 | 1260.4543 | 1226.7193 | 2096.9179 | 2144.2523 | 1533.9528 | 1982.1042 | 1871.5457 | 1619.0157 | 2950.0851 | 1302.6372 | 1297.9779 | 1227.3471 | 1523.7129 | 1124.1072 | 1523.1672 | 1559.8486 | 1248.5678 | 2747.1863 | 2773.53 | 2274.4764 | 3079.6278 | 1463.1514 | 1767.5678 | 1578.4543 | 1644.9558 | 876.2429 | 923.8107 | 895.6057 | 1570.5142 | 1231.8171 | 1796.0315 | 1296.8077 | 1244.94 | 1454.4764 | 1903.8737 | 1233.7414 | 1283.2114 | 1504.5993 | 1958.3879 | 1219.3028 | 1386.3628 | 995.1357 | 1378.4794 | 873.5142 | 988.06 | 1530.7635 | 1743.6184 | 1450.1514 | 1726.3628 | 1232.3849 | 1799.5836 | 1416.7792 | 1940.205 | 1706.3786 | 1769.6023 | 1621.5993 | 2846.7729 | 1525.2965 | 1471.6814 | 1201.9843 | 2492.4164 | 1047.0157 | 1041.7035 | 1022.1199 | 1532.8958 | 945.8549 | 895.0821 | 954.4543 | 914.7792 | 1160.0221 | 1027.9843 | 976.0884 | 1678.6908 | 1496.262 | 1714.3343 | 2056.9337 | 3856.1293 | 2928.5585 | 2998.1136 | 2924.7792 | 5262.2556 | 854.5521 | 748.7414 | 712.7193 | 854.5521 | 904.8549 | 981.9243 | 848.9307 | 1068.9149 | 875.5079 | 914.3028 | 858.8928 | 903.6593 | 798.3407 | 778.8171 | 772.6057 | 896.6814 | 802.6214 | 709.4606 | 764.4322 | 781.5079 | 1537.9022 | 1019.0157 | 1032.2271 | 1094.265 | 1098.8928 | 993.0915 | 978.0064 | 903.3028 | 1982.1987 | 2371.6593 | 2223.8328 | 2105.0693
243 | 4.6674 | 737.6903 | 1642.3353 | 1450.7669 | 1998.8948 | 1711.5624 | 1928.5791 | 1390.1249 | 1505.764 | 872.4602 | 1069.1278 | 872.7384 | 813.3264 | 1711.2586 | 1376.1593 | 1588.6106 | 1779.8722 | 1729.9578 | 1250.8918 | 1245.4887 | 1220.0482 | 1286.415 | 2103.8722 | 2095.3805 | 1606.0315 | 2011.1564 | 1862.2556 | 1595.6391 | 2897.4572 | 1256.2527 | 1247.3549 | 1166.5339 | 1511.3579 | 1126.9744 | 1497.6903 | 1562.9233 | 1207.0737 | 2754.6677 | 2672.6332 | 2379.2134 | 3061.4346 | 1512.6962 | 1725.0993 | 1538.0482 | 1631.409 | 832.176 | 869.4316 | 893.0511 | 1558.2812 | 1234.8722 | 1709.3264 | 1353.3894 | 1144.6588 | 1408.176 | 1864.1504 | 1253.4376 | 1379.4435 | 1436.8151 | 1951.2556 | 1205.3579 | 1437.6195 | 1057.3383 | 1246.5535 | 885.6421 | 965.5624 | 1543.591 | 1740.9489 | 1364.2497 | 1775.6706 | 1271.9007 | 1729.8407 | 1484.1849 | 1964.3864 | 1737.1819 | 1603.525 | 1575.2527 | 2842.1278 | 1443.1986 | 1379.4031 | 1271.1593 | 2529.0285 | 1057.7444 | 1074.1052 | 1063.8496 | 1494.0993 | 875.7895 | 863.7669 | 914.0482 | 961.7218 | 1215.5172 | 1044.5398 | 990.5143 | 1570.8888 | 1564.0826 | 1744.2586 | 1964.5054 | 3935.9292 | 2952.2842 | 2901.5309 | 2933.7188 | 5279.4631 | 862.7444 | 728.4857 | 796.7758 | 862.7444 | 878.6391 | 1016.1052 | 852.8466 | 1057.2301 | 816.5595 | 878.9203 | 870.6421 | 906.9233 | 758.0482 | 781.8722 | 743.7414 | 884.3068 | 701.6588 | 743.0541 | 780.591 | 762.5113 | 1564.3353 | 1010.2556 | 991.0737 | 1075.4857 | 1113.5654 | 956.9715 | 998.412 | 879.6136 | 1896.3264 | 2222.9114 | 2229.821 | 2177.1593
244 | 4.6784 | 734.0398 | 1552.2212 | 1532.885 | 2021.6814 | 1799.8053 | 1938.8142 | 1515.3009 | 1491.1726 | 844.3717 | 1086.7611 | 946.9779 | 851.4558 | 1849.1549 | 1357.2788 | 1688.6328 | 1795.7876 | 1630.3805 | 1296.3629 | 1219.3584 | 1271.2965 | 1264.3186 | 2106.9602 | 2113.7345 | 1533.0088 | 2021.8673 | 1809.7035 | 1580.1991 | 2902.9071 | 1236.2522 | 1295.3363 | 1146.2522 | 1562.3098 | 1114.1726 | 1464.438 | 1591.6151 | 1280.0045 | 2727.3717 | 2745.9779 | 2405.6814 | 3100.469 | 1531.7611 | 1722.0265 | 1532.0664 | 1560.9424 | 773.7699 | 819.6504 | 893 | 1562.9336 | 1376.1018 | 1754.3629 | 1391.5089 | 1207.1239 | 1426.7345 | 1816.6239 | 1260.9071 | 1296.1548 | 1471.5089 | 1874.0354 | 1216.8407 | 1399.531 | 1016.5708 | 1293.323 | 908.6947 | 949.2124 | 1511.438 | 1786.3894 | 1429.0974 | 1729.6371 | 1239.4513 | 1743.7876 | 1433.7035 | 1981.7743 | 1756.7478 | 1698.6593 | 1680.5664 | 3007.77 | 1495.2699 | 1353.3186 | 1246.3584 | 2541.8407 | 1077.7345 | 1027.6371 | 1044.2788 | 1488.0664 | 905.5752 | 912.3363 | 915.9602 | 938.3318 | 1185.4248 | 975.9292 | 990.0133 | 1599.5752 | 1558.1062 | 1650.2743 | 1922.531 | 4037.6593 | 3021.0841 | 2893.0796 | 2873.8097 | 5355.9779 | 844.2522 | 713.1991 | 704.2743 | 844.2522 | 900.6947 | 1077.1903 | 930.9514 | 1059.9602 | 856.4425 | 979.6328 | 900.6018 | 919.8274 | 765.8805 | 799.7611 | 834.7655 | 915.5752 | 752.2832 | 731.1726 | 746.4646 | 796.5354 | 1517.6371 | 971.5177 | 1011.708 | 1069.0796 | 1061.7035 | 988.562 | 1011.8274 | 835.5841 | 1966.031 | 2170.6504 | 2382.9425 | 2094.1416
245 | 4.6894 | 758.9754 | 1569.9813 | 1513.0206 | 2044.9774 | 1730.0698 | 1955.9833 | 1359.1553 | 1582.9096 | 808.0354 | 1077.0098 | 889.058 | 778.0727 | 1718.1307 | 1252.1032 | 1601.0875 | 1765.0305 | 1610.0187 | 1275.0216 | 1183.0354 | 1172.0983 | 1278.9853 | 2124.9823 | 1984.1278 | 1683.8506 | 2057.9646 | 1944.8663 | 1568.0118 | 2772.1288 | 1270.9656 | 1243.0521 | 1226.9204 | 1538.0246 | 1053.06 | 1542.9224 | 1547.0442 | 1294.9862 | 2834.8939 | 2745.002 | 2331.0737 | 3058.0423 | 1446.0845 | 1763.9587 | 1491.0403 | 1636.9243 | 840.9332 | 816.0029 | 872.0206 | 1678.886 | 1262.114 | 1741.0138 | 1351.0403 | 1185.0226 | 1421.0059 | 1871.945 | 1244.0167 | 1391.9047 | 1402.0688 | 1903.9695 | 1219.9971 | 1389.0098 | 943.0717 | 1254.0393 | 845.0629 | 926.0226 | 1567.944 | 1724.0619 | 1429.001 | 1656.0717 | 1210.0285 | 1827.9174 | 1346.0855 | 1946.0354 | 1667.0885 | 1651.0482 | 1565.115 | 2812.1946 | 1532.9636 | 1429.9243 | 1259.9862 | 2458.0826 | 986.0904 | 1064.9627 | 994.0491 | 1550.9381 | 891.0147 | 919.9931 | 983.9332 | 967.9705 | 1200.9843 | 1023.9518 | 1023.9666 | 1720.8811 | 1564.9931 | 1685.9636 | 1934.9872 | 3978.06 | 3004.0177 | 2836.056 | 2939.9341 | 5396.9607 | 833.0108 | 752.9607 | 659.0432 | 833.0108 | 838.0619 | 989.0875 | 838.0924 | 1027.0324 | 801.055 | 848.1307 | 898.0029 | 866.0531 | 802.9636 | 743.056 | 711.1229 | 856.059 | 753.999 | 739.9912 | 718.0275 | 744.0521 | 1477.0393 | 960.0108 | 1041.9705 | 1055.0138 | 1079.9813 | 986.0029 | 1057.9548 | 869.9656 | 1865.1002 | 2250.9204 | 2177.2044 | 2142.9509
246 | 4.7004 | 757.9095 | 1530.1922 | 1474.7222 | 2016.0339 | 1739.0792 | 1907.2827 | 1393.2488 | 1442.5202 | 859.5703 | 1108.9823 | 914.5929 | 887.706 | 1790.1406 | 1307.8078 | 1679.9144 | 1678.4071 | 1557.4248 | 1238.9661 | 1219.0452 | 1259.2035 | 1236.9047 | 1973.9273 | 1963.1357 | 1643.4459 | 2017.0452 | 1817.2262 | 1459.5492 | 2893.1067 | 1248.735 | 1290.6608 | 1208.8191 | 1452.9661 | 1092.1018 | 1467.7965 | 1511.6559 | 1311.2537 | 2814.5024 | 2690.9887 | 2315.3506 | 3074.2552 | 1444.7916 | 1737.6898 | 1558.1131 | 1559.029 | 853.6721 | 859.2149 | 859.1696 | 1533.147 | 1313.0452 | 1714.9661 | 1331.4749 | 1202.7965 | 1445.6333 | 1851.9435 | 1197.6106 | 1321.7851 | 1432.0565 | 1901.2149 | 1219.8304 | 1426.6834 | 1050.1632 | 1276.9548 | 871.0792 | 1017.0953 | 1424.9386 | 1761.2827 | 1368.9322 | 1698.3732 | 1199.5654 | 1711.6155 | 1374.4523 | 2040.2084 | 1775.2084 | 1712.1858 | 1543.3279 | 2900.2262 | 1541.643 | 1349.3279 | 1265.084 | 2452.1534 | 1032.2537 | 1038.3279 | 1066.4911 | 1568.8643 | 866.4297 | 887.9095 | 939.3215 | 880.3166 | 1114.6559 | 994.916 | 1051.1406 | 1714.5251 | 1567.2665 | 1694.4523 | 2062.6543 | 4074.6559 | 3012.3053 | 2961.3102 | 2753.2552 | 5405.1308 | 819.3392 | 781.6155 | 757.5137 | 819.3392 | 859.1809 | 939.147 | 862.2035 | 994.8594 | 810.588 | 873.2714 | 876.9095 | 958.231 | 753.1018 | 766.8191 | 760.3328 | 898.5703 | 733.2488 | 779.8756 | 743.1971 | 778.5703 | 1517.4911 | 953.8029 | 1024.9208 | 1055.0678 | 1110.3392 | 961.9435 | 987.4976 | 907.0113 | 1954.4523 | 2242.4976 | 2237.3441 | 2170.1519
247 | 4.7114 | 749.6932 | 1546.7168 | 1363.9912 | 2017.6076 | 1745.1976 | 1930.7964 | 1415.1888 | 1488.2832 | 822.472 | 1020.9912 | 878.2448 | 859.6637 | 1715.6608 | 1291.0708 | 1584.5752 | 1715.6991 | 1680.8908 | 1236.5428 | 1223.4484 | 1276.7876 | 1313.9941 | 2054.292 | 1976.8171 | 1596.298 | 2021.4484 | 1836.7404 | 1592.1917 | 2816.7463 | 1312.2684 | 1260.3156 | 1192.0088 | 1451.4868 | 1100.0088 | 1449.4808 | 1569.118 | 1246.354 | 2769.702 | 2690.236 | 2344.3952 | 3184.3776 | 1511.2861 | 1796.8436 | 1567.3156 | 1649.3776 | 824.3628 | 878.2124 | 873.3864 | 1548.8319 | 1317.0944 | 1710.4424 | 1370.3333 | 1186.534 | 1497.6725 | 1846.0088 | 1252.062 | 1304.2744 | 1437.2832 | 1919.1976 | 1204.0324 | 1398.0796 | 978.3924 | 1273.4248 | 878.944 | 938.1652 | 1507.6136 | 1785.5752 | 1365.1947 | 1731.2832 | 1246.8525 | 1676.7964 | 1413.056 | 1972.9587 | 1706.1652 | 1639.4424 | 1571.1504 | 2919.2684 | 1424.3452 | 1378.0472 | 1187.0767 | 2554.9115 | 968.2508 | 1065.7964 | 1023.3216 | 1555.7256 | 904.0708 | 880.6608 | 881.4396 | 908.826 | 1169.7109 | 1074.056 | 976 | 1673.8968 | 1674.7905 | 1732.938 | 1946.0856 | 4131.5516 | 3036.7581 | 2901.2596 | 2863.472 | 5504.0236 | 847.702 | 747.0088 | 714.41 | 847.702 | 873.3481 | 952.8024 | 879.174 | 1068.6784 | 861.0088 | 897.236 | 868.41 | 890.528 | 761.3156 | 751.4248 | 702.826 | 862.3216 | 755.354 | 767.4188 | 674.0384 | 741.4248 | 1473.6608 | 1023.292 | 1018.4012 | 1062.062 | 1140.4956 | 959.5015 | 1030.1976 | 906.3952 | 1993.1268 | 2284.9144 | 2180.0708 | 2098.1711
248 | 4.7224 | 736.6052 | 1536.1485 | 1407.1175 | 1962.2611 | 1713.3279 | 1880.7129 | 1419.088 | 1501.3997 | 840.9592 | 1062.4307 | 848.651 | 764.651 | 1742.8205 | 1293.3181 | 1566.1745 | 1784.2414 | 1635.6101 | 1258.6052 | 1240.0978 | 1268.1436 | 1233.8712 | 2022.7748 | 2048.7699 | 1649.9641 | 2039.5693 | 1828.3688 | 1478.4405 | 2887.8156 | 1281.1386 | 1233.0767 | 1153.7277 | 1513.1077 | 1060.2921 | 1428.5025 | 1573.5693 | 1259.9951 | 2714.6362 | 2700.1795 | 2281.1942 | 3200.6819 | 1399.0359 | 1747.6003 | 1537.7438 | 1663.5644 | 798.8614 | 886.8564 | 846.7797 | 1660.6523 | 1320.0669 | 1646.2871 | 1262.6559 | 1246.4827 | 1362.6559 | 1802.703 | 1291.3701 | 1356.7847 | 1449.646 | 1883.6671 | 1162.9233 | 1363.7685 | 1036.5742 | 1290.6052 | 914.6981 | 944.4258 | 1532.2822 | 1766.8874 | 1453.8466 | 1742.354 | 1160.5841 | 1630.0866 | 1375.2562 | 2050.1956 | 1713.2512 | 1577.7636 | 1537.0049 | 2889.5742 | 1480.4307 | 1379.677 | 1269.5904 | 2634.1795 | 1018.7438 | 1017.5792 | 1075.703 | 1504.5433 | 906.2463 | 904.9493 | 941.8255 | 943.7537 | 1076.9381 | 1033.3849 | 973.4877 | 1709.3132 | 1544.3329 | 1690.9381 | 1992.2203 | 4114.3638 | 2949.1323 | 2910.026 | 2881.6721 | 5502.8466 | 794.969 | 706.6101 | 691.974 | 794.969 | 806.1485 | 984.7079 | 845.1436 | 1015.5074 | 821.4715 | 905.8874 | 842.2055 | 872.5792 | 734.0767 | 768.3181 | 735.9951 | 914.0929 | 769.2104 | 702.0718 | 718.5433 | 760.1844 | 1500.8564 | 995.2921 | 1032.0211 | 1106.6819 | 1158.3082 | 1059.0521 | 998.974 | 841.1126 | 1959.6918 | 2235.3343 | 2184.0408 | 2187.031
249 | 4.7334 | 728.9872 | 1567.7227 | 1461.8318 | 2023.1927 | 1748.5664 | 1873.8673 | 1319.5409 | 1537.7473 | 814.9754 | 1070.7718 | 893.4818 | 810.1564 | 1736.3618 | 1277.5054 | 1543.47 | 1738.6863 | 1649.0609 | 1248.3855 | 1189 | 1271.5182 | 1307.1937 | 1969.4563 | 1964.6018 | 1649.0964 | 2024.0973 | 1892.0482 | 1546.8909 | 2860.0472 | 1258.7709 | 1262.0964 | 1222.9046 | 1517.4936 | 1091.8682 | 1441.9518 | 1558.1455 | 1235.0236 | 2681.6136 | 2700.76 | 2370.2763 | 3219.8191 | 1400.9636 | 1713.0836 | 1506.6018 | 1624.1327 | 820.8918 | 880.9764 | 815.5418 | 1519.12 | 1296.2045 | 1654.0964 | 1334 | 1152.4818 | 1434.5782 | 1746.1554 | 1169.4091 | 1348.2055 | 1466.9518 | 1820.7109 | 1136.2163 | 1457.291 | 1049.2537 | 1278.3137 | 827.9754 | 931.3609 | 1512.47 | 1740.2763 | 1396.4818 | 1724.2773 | 1228.6027 | 1713.3855 | 1407.8918 | 1952.0954 | 1719.0846 | 1650.7955 | 1563.8555 | 2906.7345 | 1492.9636 | 1374.18 | 1179.4936 | 2633.411 | 987.6982 | 1057.9764 | 1013.3736 | 1543.7355 | 888.8427 | 827.7218 | 962.4582 | 961.35 | 1129.5782 | 988.9882 | 909.71 | 1661.2537 | 1604.6873 | 1739.6755 | 2021.0009 | 4154.8063 | 3085.1209 | 2854.9764 | 2844.7473 | 5449.0954 | 821.5782 | 720.6873 | 741.8682 | 821.5782 | 838.4936 | 952.6018 | 848.0364 | 1056 | 858.7837 | 876.5782 | 797.4209 | 913.9882 | 759.9646 | 750.1927 | 707.8436 | 837.4454 | 746.5064 | 703.3255 | 757.6873 | 789.5909 | 1499.12 | 1028.6991 | 953.5418 | 1120.6637 | 1087.1564 | 954.4454 | 1047.5673 | 868.88 | 2031.7109 | 2114.8181 | 2210.6027 | 2159.41
250 | 4.7444 | 785.7625 | 1584.177 | 1414.4041 | 2029.177 | 1736.0088 | 1887.6135 | 1421.6224 | 1510.3392 | 851.5811 | 1012.3216 | 900.3997 | 831.3672 | 1699.1017 | 1323.5531 | 1530.5206 | 1772.9676 | 1592.9956 | 1256.7861 | 1190.2699 | 1260.1254 | 1234.1165 | 2094.5206 | 2018.0561 | 1653.3584 | 1951.354 | 1889.674 | 1603.4557 | 2938.6696 | 1273.6976 | 1264.0605 | 1145.6283 | 1546.7581 | 1029.1534 | 1441.6047 | 1489.68 | 1271.4321 | 2749.9159 | 2623.9543 | 2428.2876 | 3230.497 | 1424.6564 | 1771.3997 | 1553.7669 | 1603.0693 | 799.3024 | 843.5398 | 843.2183 | 1558.9439 | 1318.3672 | 1660.1121 | 1332.6696 | 1162.1209 | 1438.32 | 1845.6416 | 1214.9439 | 1296.8702 | 1457.171 | 1817.4189 | 1182.6003 | 1310.8054 | 993.6696 | 1244.9395 | 868.0885 | 973.3436 | 1502.3628 | 1795.9896 | 1404.6696 | 1708.972 | 1206.1402 | 1722.7537 | 1387.7537 | 2032.3953 | 1705.7168 | 1622.4646 | 1561.5443 | 2944.9912 | 1510.3068 | 1434.9439 | 1210.2094 | 2522.1534 | 1036.1017 | 1017.8791 | 1003.4794 | 1500.9956 | 943.3672 | 917.9159 | 930.3068 | 908.6976 | 1136.4041 | 970.9351 | 980.5206 | 1611.7168 | 1562.6888 | 1672.1254 | 1941.7581 | 4225.7065 | 3047.3348 | 2812.1578 | 2826.4101 | 5535.5635 | 834.0885 | 674.7581 | 679.6976 | 834.0885 | 861.7713 | 1000.4321 | 826.2787 | 1052.9764 | 814.9395 | 888.7257 | 863.9115 | 892.1858 | 701.4513 | 756.3584 | 687.4189 | 862.3112 | 719.298 | 691.6328 | 711.0324 | 743.9956 | 1536.4041 | 1000.9439 | 983.4557 | 1038.5074 | 1105.7669 | 979.5531 | 964.5959 | 828.2094 | 2024.4882 | 2210.5443 | 2182.8186 | 2130.298
251 | 4.7554 | 750.1996 | 1617.6362 | 1347.7296 | 2061.472 | 1674.9459 | 1838.7837 | 1310.2016 | 1474.0207 | 815.3825 | 1090.1248 | 860.7276 | 808.1101 | 1627.002 | 1266.5074 | 1495.2016 | 1785.5988 | 1624.3638 | 1260.1956 | 1211.9646 | 1228.0167 | 1263.5074 | 2057.6736 | 2044.2517 | 1688.3264 | 2015.5074 | 1837.1996 | 1592.0167 | 2918.6549 | 1268.0541 | 1257.5781 | 1194.4326 | 1544.9646 | 1020.3658 | 1449.1809 | 1606.8899 | 1215.9272 | 2702 | 2731.2724 | 2327.2744 | 3150.53 | 1494.0895 | 1734.9292 | 1489.3097 | 1551.241 | 800.0708 | 880.0167 | 803.5654 | 1530.6382 | 1295.3284 | 1721.9626 | 1310.1809 | 1167.8565 | 1323.4386 | 1810.8191 | 1183 | 1348.0501 | 1327.6402 | 1854.7817 | 1216.4513 | 1399.5447 | 971.6903 | 1245.9646 | 876.4867 | 902.528 | 1507.6362 | 1661.239 | 1383.9272 | 1705.8064 | 1279.5968 | 1741.5594 | 1414.4179 | 2023.3638 | 1772.3058 | 1555.4926 | 1571.4553 | 3012.4573 | 1486.8378 | 1403.4366 | 1227.8358 | 2511.1101 | 962.6195 | 1016.4366 | 1039.4533 | 1531.3451 | 912.8338 | 867.1622 | 904.7276 | 904.5821 | 1071.1121 | 1000.1996 | 942.0728 | 1684.1996 | 1619.6362 | 1646.7109 | 1937.2724 | 4167.8771 | 3027.3225 | 2767.5634 | 2923.0167 | 5619.3078 | 848.8899 | 672.9646 | 673.3992 | 848.8899 | 767.4012 | 940.5614 | 784.0728 | 1003.4533 | 812.2537 | 895.0521 | 845.6903 | 861.8171 | 721.1622 | 788.707 | 724.3451 | 806.6008 | 754.7089 | 718.1642 | 697.6736 | 774.2724 | 1463.8358 | 970.4553 | 976.528 | 1092.5634 | 1042.474 | 928.474 | 1033.8171 | 848.6008 | 2000.6362 | 2215.7984 | 2182.0374 | 2168.0374
252 | 4.7664 | 738.9793 | 1609.4277 | 1406.2139 | 2045.147 | 1687.4253 | 1891.0826 | 1348.6185 | 1532.3013 | 832.2345 | 1028.2552 | 871.8712 | 833.2989 | 1678.1701 | 1223.7861 | 1532.7679 | 1756.1908 | 1630.531 | 1158.7169 | 1230.8761 | 1198.2503 | 1215.853 | 2074.4071 | 1973.8068 | 1662.6622 | 1983.2479 | 1827 | 1557.7424 | 2919.9405 | 1255.7448 | 1167.1677 | 1104.209 | 1564.6416 | 1079.4872 | 1424.8943 | 1574.5565 | 1216.0207 | 2705.4022 | 2722.3402 | 2355.8165 | 3218.7886 | 1453.853 | 1777.0899 | 1502.5978 | 1668.0644 | 747.1081 | 838.2576 | 846.2783 | 1576.4897 | 1322.1081 | 1695.4046 | 1288.0207 | 1210.5359 | 1376.6829 | 1828.2163 | 1184.4459 | 1243.1859 | 1427.4666 | 1852.5954 | 1138.0182 | 1383.1264 | 954.4046 | 1271.1701 | 775.74 | 924.7217 | 1502.9149 | 1721.5359 | 1380.4484 | 1862.2807 | 1202.8481 | 1639.0182 | 1403.1725 | 2025.7035 | 1691.3378 | 1599.8299 | 1584.3864 | 3079.6185 | 1523.5954 | 1401.6853 | 1215.9769 | 2548.7169 | 998.1264 | 1024.2965 | 1007.0851 | 1522.8299 | 845.469 | 839.4204 | 912.6391 | 913.9563 | 1144.8323 | 989.3196 | 948.2734 | 1672.384 | 1621.3791 | 1696.3815 | 1928.5954 | 4279.7242 | 2911.4848 | 2773.8226 | 2880.322 | 5584.0644 | 818.7679 | 696.5541 | 659.1239 | 818.7679 | 800.2965 | 890.294 | 784.4046 | 973.6367 | 786.5978 | 832.7655 | 826.7035 | 830.938 | 682.0231 | 728.6416 | 711.8323 | 822.2369 | 747.7679 | 732.1495 | 711.0049 | 765.5103 | 1454.2734 | 993.4228 | 993.8736 | 1084.2783 | 1072.9793 | 960.6804 | 1004.7242 | 865.5978 | 1990.8967 | 2170.7473 | 2204.2576 | 2191.7035
253 | 4.7774 | 727.9027 | 1568.7611 | 1390.1151 | 2071.5044 | 1690.9292 | 1932.2743 | 1323.0973 | 1470.2212 | 825.7522 | 1032.5575 | 885.8672 | 812.177 | 1685.1593 | 1249.6195 | 1502.7699 | 1762.5752 | 1632.4955 | 1264.5575 | 1169.2832 | 1281.1593 | 1189.7876 | 2024.1239 | 2013.2478 | 1617.9469 | 2111.4956 | 1821.354 | 1587.7787 | 2857.8231 | 1254.3628 | 1213.1681 | 1175.0885 | 1508.0708 | 1107.3628 | 1416.6991 | 1520.6726 | 1231.9646 | 2738.0088 | 2719.9646 | 2222.0797 | 3193.3451 | 1419.708 | 1690.1504 | 1471.7434 | 1655.2743 | 721.9381 | 797.5133 | 815.1504 | 1569.9292 | 1289.0531 | 1685.0088 | 1292.8672 | 1142.3628 | 1345.8673 | 1765.0265 | 1201.7522 | 1339.6637 | 1451.2743 | 1848.0973 | 1197.2567 | 1403.177 | 950.531 | 1263.4601 | 859 | 960.3275 | 1500.7345 | 1663.3894 | 1359.0354 | 1778.1238 | 1253.2478 | 1689.4956 | 1361.4159 | 2004.9912 | 1739.2655 | 1596.4425 | 1533.4425 | 3037.6459 | 1527.9557 | 1325.6283 | 1242.938 | 2653.8761 | 1017.0088 | 1041.7964 | 1008.8761 | 1517.9823 | 827.6372 | 923.2832 | 892.7433 | 929.5486 | 1103.6018 | 985.2124 | 1005.0973 | 1654.708 | 1695.4514 | 1718.4956 | 1931.7876 | 4276.0266 | 2999.0089 | 2899.5398 | 2818.9999 | 5583.5399 | 790.2301 | 680.1239 | 720.2301 | 790.2301 | 828.0177 | 955.3717 | 777.2212 | 1004.531 | 751.5575 | 840.0265 | 809.885 | 804.0354 | 652.7257 | 674.3098 | 671.7699 | 773.1416 | 711.5398 | 720.8584 | 630.1593 | 769.5398 | 1514.8938 | 1041.5841 | 975.8673 | 1055.7788 | 1064.0443 | 964.1062 | 992.9469 | 830.2655 | 1937.5575 | 2134.7346 | 2162.0531 | 2174.3009
254 | 4.7884 | 725.794 | 1547.1229 | 1409.7905 | 2005.2891 | 1674.4218 | 1899.764 | 1326.0831 | 1463.6578 | 838.1627 | 1043.8205 | 829.2193 | 782.6013 | 1729.9302 | 1308.6976 | 1545.5747 | 1707.7144 | 1498.8771 | 1186.2095 | 1209.5482 | 1233.1397 | 1258.0531 | 2007.3387 | 2085.2954 | 1638.5216 | 2065.6243 | 1770.9735 | 1501.7709 | 2967.2521 | 1265.3555 | 1227.0698 | 1202.4749 | 1510.8736 | 1027.7577 | 1450.6711 | 1557.8603 | 1267.146 | 2699.3555 | 2664.1096 | 2289.2158 | 3182.0566 | 1425.1131 | 1674.9735 | 1522.4784 | 1623.8603 | 769.8736 | 799.9071 | 794.0531 | 1562.9602 | 1271.6578 | 1652.9336 | 1242.03 | 1158.1795 | 1389.1362 | 1749.0929 | 1212.8471 | 1314.382 | 1425.4518 | 1748.8638 | 1235.8038 | 1383.03 | 972.6844 | 1174.8904 | 854.3953 | 1012.0929 | 1485.382 | 1765.382 | 1399.2123 | 1677.648 | 1235.3555 | 1657.9867 | 1364.6613 | 2036.2989 | 1742.558 | 1581.5615 | 1613.5447 | 2896.854 | 1516.7842 | 1345.5049 | 1195.6711 | 2565.0433 | 971.8869 | 980.8107 | 1011.0796 | 1495.2626 | 820.8107 | 865.3855 | 850.2493 | 880.6445 | 1128.3387 | 1015.4385 | 998.3254 | 1655.1795 | 1726.8136 | 1683.1494 | 1905.8373 | 4363.8471 | 3072.442 | 2875.146 | 2727.9637 | 5678.9071 | 843.0664 | 667.7242 | 671.9434 | 843.0664 | 854.9036 | 923.3254 | 769.6711 | 1028.1096 | 772.9071 | 835.2758 | 880.9169 | 858.6844 | 681.9469 | 699.8471 | 709.4218 | 796.9469 | 701.1795 | 701.8771 | 681.9469 | 751.7242 | 1545.5182 | 1043.4916 | 1018.4784 | 1060.8073 | 1085.3024 | 929.6445 | 1010.3422 | 837.6145 | 1959.0167 | 2211.4051 | 2148.4951 | 2240.0265
255 | 4.7994 | 711.4887 | 1544.6333 | 1341.828 | 2024.1966 | 1619.6588 | 1839.7807 | 1342.2684 | 1471.8043 | 812.4867 | 1013.7325 | 825.1475 | 759.1455 | 1693.2949 | 1285.2448 | 1526.3186 | 1742.2202 | 1517.827 | 1253.5624 | 1207.172 | 1267.0275 | 1227.707 | 1951.1436 | 1972.293 | 1657.293 | 1954.5614 | 1794.4376 | 1572.0246 | 2825.8516 | 1246.3913 | 1245.294 | 1120.7798 | 1464.9489 | 1101.2694 | 1437.0265 | 1510.1947 | 1181.4386 | 2701.8506 | 2655.8554 | 2245.6304 | 3214.1465 | 1485.879 | 1693.6824 | 1495.172 | 1583.8761 | 726.9971 | 850.5133 | 769.4376 | 1584.9272 | 1284.3157 | 1683.2458 | 1283.0747 | 1178.5133 | 1371.8053 | 1703.8516 | 1175.2675 | 1291.828 | 1422.0973 | 1766.8043 | 1153.2439 | 1425.4159 | 964.0728 | 1176.5605 | 819.5605 | 952.3422 | 1457.9508 | 1727.8308 | 1281.7296 | 1756.7316 | 1224.5605 | 1667.707 | 1427.9282 | 1999.1446 | 1682.6569 | 1585.8034 | 1542.88 | 2989.6333 | 1564.759 | 1314.8525 | 1192.8781 | 2601.4887 | 910.2675 | 1036.0472 | 957.6352 | 1539.8289 | 870.5605 | 893.8063 | 895.3913 | 891.3176 | 1063.4867 | 1017.1711 | 934.4149 | 1683.0728 | 1585.0217 | 1654.7571 | 1941.8771 | 4326.5113 | 3132.7089 | 2777.1239 | 2815.293 | 5722.0757 | 805.1947 | 633 | 636.1957 | 805.1947 | 841.4887 | 864.292 | 773.0964 | 1022.5378 | 819.7335 | 867.8053 | 863.7561 | 847.7561 | 709.1229 | 661.292 | 657.4631 | 828.7561 | 718.8308 | 707.5369 | 704.3677 | 719.4386 | 1498.0737 | 1018.0482 | 999.2193 | 1049.706 | 1099.7325 | 943.9754 | 1001.4641 | 804.3166 | 2011.6106 | 2107.2183 | 2179.0255 | 2234.8516
256 | 4.8104 | 695.7773 | 1575.7183 | 1395.295 | 1972.5634 | 1626.9116 | 1828.3156 | 1330.4233 | 1490.3894 | 839.5177 | 979.2419 | 784.6106 | 788.7316 | 1615.8452 | 1251.4027 | 1472.208 | 1721.4425 | 1605.9528 | 1191.9469 | 1159.7448 | 1156.4292 | 1240.6711 | 2037.028 | 1968.59 | 1641.1077 | 1947.4691 | 1847.1947 | 1561.6445 | 2896.0811 | 1203.5841 | 1202.503 | 1148.9395 | 1558.382 | 1052.2153 | 1368.0664 | 1525.0605 | 1198.5973 | 2811.3023 | 2597.9189 | 2373.9808 | 3218.8717 | 1439.0678 | 1699.3097 | 1444.9262 | 1650.1681 | 817.82 | 812.6106 | 812.2625 | 1569.1947 | 1334.4498 | 1619.9794 | 1240.6298 | 1139 | 1363.2478 | 1775.9602 | 1202.3761 | 1330.6652 | 1430.1342 | 1779.0811 | 1156.7448 | 1379.6106 | 977.2552 | 1194.4631 | 836.792 | 921.9734 | 1477.503 | 1677.7242 | 1355.2419 | 1762.6711 | 1243.8053 | 1680.5369 | 1368.1416 | 2060.4366 | 1736.0073 | 1636.497 | 1477.2478 | 2999.2433 | 1461.5767 | 1351.6312 | 1194.4159 | 2577.9927 | 943.6844 | 1085.2758 | 931.7508 | 1551.5236 | 887.3894 | 837.5708 | 852.0472 | 871.7316 | 1107.9661 | 996.3495 | 926.2213 | 1693.1681 | 1685.3289 | 1617 | 1973.1077 | 4359 | 3082.2611 | 2725.1933 | 2805.6844 | 5650.3408 | 818.5841 | 640.9189 | 618.9528 | 818.5841 | 811.4159 | 898.4691 | 810.1548 | 973.5841 | 761.0531 | 847.0339 | 853.6784 | 842.7787 | 675.6106 | 681.6445 | 670.8186 | 828.0738 | 661.3422 | 694.3156 | 638.6239 | 735.5236 | 1477.6711 | 1034.3628 | 1010.1475 | 1099.7448 | 1062.0206 | 952.7448 | 981.6372 | 816.1681 | 1978.2552 | 2142.0678 | 2132.1077 | 2300.6519
257 | 4.8214 | 795.2094 | 1556.0187 | 1410.2536 | 2010.058 | 1716.0747 | 1811.1465 | 1287.414 | 1481.4651 | 761.0511 | 951.6303 | 796.7581 | 788.4651 | 1663.3697 | 1251.8024 | 1477.9489 | 1765.9302 | 1619.4326 | 1149.8162 | 1174.2232 | 1191.5929 | 1246.9046 | 1941.0325 | 2023.5605 | 1617.0256 | 2009.5791 | 1834.821 | 1529.8908 | 2854.3766 | 1194.0069 | 1221.0049 | 1151.7768 | 1599.4071 | 1014.3068 | 1426.7001 | 1515.0442 | 1188.528 | 2794.0394 | 2647.8279 | 2298.7906 | 3285.7768 | 1392.7277 | 1617.886 | 1435.9489 | 1613.1396 | 714.179 | 821.9744 | 750.472 | 1557.8466 | 1313.7512 | 1656.7955 | 1326.942 | 1143.5418 | 1404.4326 | 1742.1278 | 1174.6047 | 1273.3766 | 1434.8093 | 1738.5487 | 1167.707 | 1390.8791 | 966.7325 | 1194.6627 | 827.2419 | 903.472 | 1497.763 | 1731.6558 | 1336.5674 | 1783.2626 | 1250.3186 | 1677.9931 | 1321.6952 | 2030.5605 | 1683.7139 | 1601.7325 | 1521.2163 | 2864.4602 | 1504 | 1317.8791 | 1206.0442 | 2639.6627 | 964.9558 | 1028.414 | 1009.1652 | 1542.1465 | 880.6627 | 810.0256 | 831.3884 | 869.5605 | 1138.3884 | 956.6696 | 956.3766 | 1662.4395 | 1685.9813 | 1618.6116 | 1946.5349 | 4365.0069 | 3124.2723 | 2851.7837 | 2814.2094 | 5790.7837 | 803.7325 | 683.7955 | 645.472 | 803.7325 | 822.8977 | 962.1907 | 770.9813 | 959.1721 | 818.1278 | 819.8348 | 801.3186 | 823.1721 | 685.2675 | 662.7837 | 685.7768 | 828.5536 | 677.472 | 684.6185 | 662.7325 | 723.8024 | 1481.8535 | 1004.5929 | 1024.7325 | 1101.0373 | 1027.6952 | 958.8722 | 1005.8397 | 792.3255 | 1979.4533 | 2098.3649 | 2103.0442 | 2230.2301
258 | 4.8324 | 741.7768 | 1528.2704 | 1345.8457 | 2034.4503 | 1599.3648 | 1812.3176 | 1278.9184 | 1489.3648 | 778.7168 | 941.8457 | 804.1416 | 798.9056 | 1677.442 | 1232.4892 | 1468.5408 | 1776.648 | 1593.8284 | 1205.412 | 1190.7768 | 1224.442 | 1256.2532 | 1988.6696 | 2006.7079 | 1617.7596 | 1984.6696 | 1790.4592 | 1472.1116 | 2909.9528 | 1236.7812 | 1136.1333 | 1137.764 | 1564.0172 | 1077.7424 | 1393.4936 | 1494.0644 | 1158.3393 | 2859.9955 | 2644.4936 | 2343.8068 | 3272.8756 | 1477.2615 | 1664.5108 | 1439.5625 | 1592.8157 | 751.0772 | 807.1072 | 778.5963 | 1520.5708 | 1279.9484 | 1599.6307 | 1281.4292 | 1171.9056 | 1381.5281 | 1766.9184 | 1149.3176 | 1331.3348 | 1456.8412 | 1796.3476 | 1168.382 | 1398.0644 | 970.6824 | 1207.266 | 827.7985 | 938.266 | 1414.558 | 1711.0128 | 1355.6607 | 1862.9139 | 1265.7812 | 1640.9356 | 1382.2532 | 2032.2876 | 1727.7985 | 1601.236 | 1510.7168 | 2956.7551 | 1499.7124 | 1327.9228 | 1186.176 | 2650.9955 | 977.9955 | 1020.7124 | 976.5536 | 1538.8712 | 895.1716 | 829.4981 | 828.6696 | 861.7596 | 1116.6135 | 1009.03 | 1007.1888 | 1679.2921 | 1704.9955 | 1634.0344 | 1939.2704 | 4433.0306 | 3037.0389 | 2854.9783 | 2745.3731 | 5810.0172 | 804.6652 | 638.6996 | 676.1371 | 804.6652 | 801.8584 | 933.0128 | 802.382 | 977.1161 | 826.0855 | 851.0172 | 807.0944 | 828.8884 | 685.6052 | 676.3004 | 665.8884 | 788.4592 | 722.1116 | 718.8367 | 667.794 | 712.9056 | 1477.9828 | 1032.6779 | 1014.5491 | 1044.588 | 1091.1116 | 932.176 | 942.94 | 831.3476 | 2048.824 | 2208.7424 | 2084.2876 | 2274.3131
259 | 4.8434 | 735.8171 | 1535.6844 | 1389.2743 | 1929.2006 | 1604.8968 | 1815.882 | 1264.3186 | 1494.8968 | 799.9321 | 984.761 | 780.1032 | 807.6313 | 1657.239 | 1200.5664 | 1484.0531 | 1737.354 | 1601.0855 | 1194.0914 | 1169.761 | 1199.6195 | 1246.3156 | 2013.9056 | 1936.3038 | 1604.5428 | 2015.5516 | 1780.0797 | 1476.6313 | 2893.9557 | 1271.8407 | 1210.6578 | 1151.7079 | 1567.8525 | 1041.528 | 1404.643 | 1495.5899 | 1219.761 | 2801.5251 | 2655.472 | 2306.7227 | 3290.997 | 1372.5251 | 1701.0531 | 1510.0236 | 1653.3392 | 735.9499 | 760.1711 | 704.6991 | 1515.7375 | 1239.1859 | 1565.2272 | 1281.3009 | 1179.7758 | 1409.0737 | 1751.2921 | 1153.7375 | 1341.9852 | 1434.6135 | 1774.2625 | 1186.1328 | 1420.6343 | 960.1032 | 1181.3599 | 880.8938 | 903.6343 | 1432.1947 | 1680.1593 | 1304.6756 | 1789.5015 | 1293.6548 | 1618.5368 | 1372.4867 | 2055.6844 | 1770.115 | 1579.2507 | 1546.4749 | 2887.121 | 1488.121 | 1351.7758 | 1200.9469 | 2592.0118 | 932.0148 | 993.0383 | 963.4366 | 1505.879 | 860.7109 | 890.3923 | 862.4602 | 837.764 | 1060.7522 | 988.1711 | 982.5516 | 1731.9026 | 1663.2921 | 1672.6726 | 1938.1298 | 4592.2595 | 3119.6047 | 2777.6165 | 2681.5989 | 5798.6253 | 785.4602 | 653.6844 | 617.5811 | 785.4602 | 815.2891 | 894.2891 | 813.118 | 1029.1563 | 727.1092 | 851.4306 | 798.3687 | 818.8938 | 705.2094 | 672.9499 | 662.9086 | 778.764 | 712.0885 | 665 | 670.9734 | 709.8259 | 1466.4484 | 979.7109 | 946.0413 | 1047.8672 | 1099.5664 | 947.4602 | 975.0767 | 820.7257 | 2041.0679 | 2155.9321 | 2110.5929 | 2226.2507
260 | 4.8544 | 758.3151 | 1538.4499 | 1350.4843 | 1992.8653 | 1640.2979 | 1828.851 | 1325.2837 | 1531.5816 | 759.1175 | 950.0861 | 743.235 | 794.9484 | 1696.3151 | 1213.5501 | 1482.6303 | 1718.619 | 1470.3525 | 1175.7822 | 1132.5187 | 1221.6332 | 1243 | 1967.384 | 1963.1003 | 1658.6789 | 1991.2493 | 1815.3555 | 1471.8338 | 2807.9543 | 1240.3496 | 1191.2119 | 1097.3668 | 1499.4641 | 1074.3323 | 1377.2522 | 1455.8683 | 1183.4356 | 2770.7822 | 2618.8166 | 2283.616 | 3246.3383 | 1362.3211 | 1708.4326 | 1438.4297 | 1584.2979 | 777.2635 | 760.4499 | 747.6849 | 1532.1175 | 1278.4012 | 1628.9139 | 1252.8024 | 1165.3323 | 1310.6534 | 1785.2777 | 1161.4985 | 1305.4326 | 1349.3383 | 1756.7822 | 1206.9656 | 1466.3638 | 922.3181 | 1253.7762 | 857.8825 | 951.381 | 1374.8166 | 1704.4985 | 1389.5472 | 1818.4326 | 1231.5674 | 1577.1205 | 1374.2178 | 2050.5644 | 1701.4843 | 1575.0113 | 1575.5674 | 2936.4154 | 1538.3323 | 1351.2694 | 1183.1145 | 2563.5501 | 965.0831 | 979.6362 | 1007.0314 | 1468.1377 | 863.8653 | 836.616 | 853.1662 | 858.0172 | 1110.5674 | 987.7163 | 956.5988 | 1627.1519 | 1707.7306 | 1705.1834 | 1896.5157 | 4524.619 | 3092.0457 | 2751.6706 | 2808.6789 | 5700.8428 | 774.6991 | 651.6819 | 660.3323 | 774.6991 | 811.4012 | 886 | 792.4469 | 985.7306 | 803.9484 | 789.7537 | 799.6819 | 807.4356 | 676.1317 | 712.2463 | 675.5329 | 795.3496 | 647.6534 | 667.2006 | 670.9513 | 659.387 | 1492.5846 | 984.3323 | 974.9139 | 1030.1519 | 1104.2979 | 941.4671 | 936.1691 | 853.2807 | 2075.8309 | 2105.0314 | 2116.0801 | 2247.5015
261 | 4.8654 | 733.9784 | 1521.1288 | 1376.3707 | 1982.9145 | 1614.414 | 1829.3677 | 1318.2812 | 1501.826 | 776.7178 | 993.0433 | 770.2173 | 754.4985 | 1678.827 | 1217.174 | 1428.4779 | 1752.8908 | 1520.6097 | 1162.8702 | 1166.9784 | 1220.9784 | 1238.6952 | 1969.586 | 1964.8692 | 1598.1318 | 1982.9135 | 1905.2173 | 1486.4572 | 2889.5664 | 1236.3265 | 1118.2616 | 1105.6529 | 1467.3707 | 1063.6745 | 1421.6313 | 1477.8908 | 1207.1071 | 2755.7178 | 2623.2832 | 2267.5211 | 3310.585 | 1424.6293 | 1678.6066 | 1389.5448 | 1575.829 | 730.6745 | 759.5654 | 760.8682 | 1542.8692 | 1278.2369 | 1584.7827 | 1281.6293 | 1138.825 | 1367.7827 | 1692.0216 | 1136.2173 | 1291.6519 | 1421.6076 | 1757.3717 | 1186.4562 | 1415.8692 | 949.1308 | 1150.8929 | 836.3687 | 919.6529 | 1401.59 | 1674.3255 | 1336.7384 | 1803.8692 | 1239.4779 | 1642.5438 | 1403.5664 | 2030.8063 | 1721.5005 | 1511.829 | 1573.3028 | 2904.6509 | 1520.2389 | 1392.3038 | 1144.6735 | 2577.9361 | 966.6745 | 1016.0423 | 983.0443 | 1536.6735 | 840.8043 | 826.7827 | 843.4346 | 866.8043 | 1118.0865 | 981.6087 | 931.1298 | 1664.6745 | 1672.3707 | 1715.001 | 1888.827 | 4530.1268 | 3037.4366 | 2809.173 | 2733.0649 | 5859.3265 | 764.3048 | 632.1092 | 637.3471 | 764.3048 | 819.6509 | 885.6087 | 753.1318 | 950.3707 | 773.2822 | 852.5654 | 770.3255 | 843.4346 | 659.588 | 649.0433 | 675.4572 | 777.825 | 690.8898 | 677.587 | 676.172 | 700.6066 | 1508.4995 | 971.1308 | 937.0443 | 1071.001 | 1079.001 | 929.8023 | 972.1081 | 822.3275 | 2026.3903 | 2105.3078 | 2062.8043 | 2261.0423
262 | 4.8764 | 727.1858 | 1473.9337 | 1409.8495 | 2040.9912 | 1647.5442 | 1753.7212 | 1277.854 | 1485.146 | 784.0973 | 980.4248 | 767.6239 | 704.823 | 1698.3141 | 1231.5664 | 1426.2699 | 1744.1106 | 1545.1858 | 1184.5044 | 1170.1681 | 1225.6239 | 1220.5841 | 1933.9602 | 1940.146 | 1623.8761 | 2015.6372 | 1893.8894 | 1496.7787 | 2940.4867 | 1244.177 | 1147.5354 | 1127.5929 | 1511.2478 | 1079.0973 | 1441.146 | 1452.6814 | 1145.2168 | 2786.5974 | 2628.7787 | 2231.6902 | 3228.6017 | 1394.8938 | 1616.7257 | 1437.6239 | 1661.4115 | 751.9203 | 767.0398 | 722.646 | 1522.9292 | 1213.0133 | 1590.4956 | 1250.2301 | 1108.2301 | 1360.1858 | 1686.7522 | 1129.6726 | 1287.4071 | 1379.2566 | 1826.6017 | 1162.4823 | 1406.5354 | 953.0575 | 1205.9203 | 802.9203 | 934.1062 | 1495.4513 | 1648.7124 | 1330.5487 | 1779.7699 | 1227.0797 | 1649.0487 | 1430.9867 | 2080.7079 | 1747.4867 | 1609.6814 | 1530.1018 | 2894.1859 | 1506.969 | 1367.5929 | 1145.3496 | 2630.7434 | 986.8805 | 978.9027 | 998.0663 | 1514.6814 | 838.9602 | 836.8628 | 835.7522 | 861.2168 | 1113.4381 | 979.792 | 913.031 | 1690.9115 | 1715.8319 | 1732.2433 | 1907.4823 | 4424.947 | 3097.7212 | 2771.177 | 2724.9513 | 5864.8894 | 769.3141 | 649.8319 | 611.0398 | 769.3141 | 785.062 | 882.7522 | 799.6726 | 986.761 | 765.4867 | 849.8495 | 744.2965 | 836.3761 | 690.6814 | 647.6549 | 686.1947 | 742.8628 | 646.3584 | 667.7301 | 612.1991 | 641.8452 | 1488.7522 | 985.0398 | 966.4159 | 1098.2256 | 1093.3318 | 868.5177 | 947.6327 | 886.1018 | 1993.2876 | 2213.2035 | 2065.1195 | 2228.6859
263 | 4.8874 | 689.9646 | 1520.3923 | 1375.4405 | 2044.3618 | 1626.1416 | 1751.77 | 1334.0128 | 1451.8938 | 745.6303 | 968.9469 | 765.8151 | 724.5772 | 1656.8279 | 1258.529 | 1456.0403 | 1762.2251 | 1500.2556 | 1165.9695 | 1159.5113 | 1225.3569 | 1221.2026 | 1929.7797 | 1890.1544 | 1543.8938 | 2051.2477 | 1853.9695 | 1441.6077 | 2963.0531 | 1235.2154 | 1157.9164 | 1107.2203 | 1508.291 | 1051.4228 | 1387.8505 | 1426.9469 | 1143.7395 | 2808.7767 | 2562.3746 | 2206.3295 | 3206.8407 | 1432.3264 | 1648.2959 | 1426.1239 | 1587.6254 | 738.2733 | 759.3264 | 761.4936 | 1495.7797 | 1270.0354 | 1608.3697 | 1284.6303 | 1154.0836 | 1327.59 | 1695.8761 | 1126.6913 | 1305.7974 | 1373.8102 | 1782.5949 | 1123.7749 | 1425.2861 | 910.6303 | 1190.5113 | 789.9341 | 896.6736 | 1432.8938 | 1651.6736 | 1385.3972 | 1736.119 | 1233.0177 | 1619.9036 | 1366.0354 | 2009.4002 | 1760.3618 | 1580.2026 | 1549.6047 | 3021.2074 | 1495.4533 | 1349.4228 | 1205.0659 | 2651.6077 | 974.7797 | 988.5162 | 943.0177 | 1499.7395 | 829.7139 | 861.6785 | 836.7267 | 846.9646 | 1162.9469 | 973.0531 | 923.6657 | 1708.6077 | 1718.119 | 1657.4405 | 1852.4759 | 4562.1269 | 3085.4179 | 2757.59 | 2717.3136 | 5842.3697 | 731.5723 | 645.7315 | 620.0659 | 731.5723 | 816.0226 | 876.8987 | 780.6126 | 974.6736 | 803.0305 | 809.9341 | 746.5546 | 828.7139 | 668.6913 | 678.2861 | 626.6608 | 784.8584 | 674.9518 | 628.709 | 643.4533 | 697.1898 | 1490.8279 | 990.1013 | 979.0128 | 1059.8987 | 1016.2508 | 920.2203 | 943.3343 | 905.3844 | 2008.1672 | 2147.8761 | 2097.0433 | 2268.8505
264 | 4.8984 | 717.9528 | 1499.3255 | 1392.8132 | 1989.4651 | 1576.7439 | 1831.8377 | 1348.1175 | 1497.2557 | 732.5108 | 978.1401 | 746.8142 | 732.5108 | 1675.1863 | 1230.3486 | 1405.9995 | 1711.884 | 1485.3722 | 1130.7443 | 1129.9769 | 1218.0929 | 1222.9538 | 1929.4651 | 1871.8373 | 1582.5118 | 1983.2311 | 1817.5816 | 1471.5113 | 2943.3953 | 1273.1632 | 1167.4642 | 1113.116 | 1492.6274 | 1067.9066 | 1371.5575 | 1442.884 | 1198.8142 | 2719.091 | 2574.3708 | 2284.1401 | 3286.7434 | 1399.1392 | 1575 | 1389.0925 | 1604 | 714.0467 | 727.767 | 743.1627 | 1522.208 | 1268.5575 | 1643.116 | 1255.9302 | 1150.5821 | 1371.7203 | 1718.9547 | 1157.7443 | 1295.442 | 1397.791 | 1763.7443 | 1159.4656 | 1428.4193 | 894.7203 | 1167.7208 | 778.5113 | 935.2094 | 1467.0934 | 1670.908 | 1321.116 | 1763.4637 | 1233.209 | 1573.0231 | 1350.3722 | 2077.0462 | 1708.0231 | 1587.7675 | 1465.9524 | 2966.0241 | 1545.767 | 1350.3259 | 1229.5113 | 2671.2792 | 969.3491 | 923.4415 | 945.9995 | 1546.2099 | 808.6745 | 832.35 | 867.6972 | 860.7675 | 1185.6278 | 940.6052 | 890.9297 | 1724.3259 | 1732.2085 | 1667.6042 | 1857.6736 | 4621.2104 | 3124.1868 | 2785.442 | 2672.3472 | 5907.1146 | 765.2792 | 633.7203 | 635.442 | 765.2792 | 765.558 | 882.9995 | 755.6745 | 998.3264 | 814.0708 | 787.1165 | 782.3958 | 812.7906 | 694.8608 | 681.8844 | 637.8368 | 830.3028 | 645.1165 | 649.2561 | 661.908 | 662.5118 | 1514.0698 | 976.0939 | 1011.4887 | 1066.9297 | 1055.7439 | 921.9302 | 977.6981 | 805.6042 | 2049.3014 | 2169.3269 | 2142.8849 | 2224.4193
265 | 4.9094 | 761.7965 | 1498.6755 | 1416.1091 | 1983.1652 | 1606.3451 | 1759.4985 | 1272.121 | 1485.7197 | 784.7965 | 934.5929 | 735.9381 | 783.8171 | 1651.5929 | 1228.4129 | 1444.1416 | 1694.2655 | 1490.528 | 1127.8348 | 1107.0236 | 1224.7552 | 1224.3009 | 1920.7168 | 1882.6903 | 1585.4455 | 2072.472 | 1798.1032 | 1488.5929 | 2969.0561 | 1230.2891 | 1224.9616 | 1133.6726 | 1529.0413 | 1086.8997 | 1406.3894 | 1425.5103 | 1180.5929 | 2827.6401 | 2616.8288 | 2228.1062 | 3354.4868 | 1433.59 | 1609.5222 | 1430.469 | 1599.8377 | 701.2448 | 752.6549 | 737.2065 | 1598.0442 | 1294.0855 | 1624.0088 | 1248.2271 | 1117.9351 | 1407.941 | 1642.059 | 1140.879 | 1274.7581 | 1387.2213 | 1729.4955 | 1133.1416 | 1383.8555 | 929.9616 | 1181.9587 | 801.469 | 921.3481 | 1429.4926 | 1613.9145 | 1332.6136 | 1825.0707 | 1254.5516 | 1608.2183 | 1360.9145 | 2090.3894 | 1709.4307 | 1578.6755 | 1538.879 | 2915.8761 | 1580.6932 | 1327.2035 | 1235.3304 | 2667.8142 | 945.9203 | 961.9558 | 975.8171 | 1521.8732 | 814.7109 | 764.2213 | 885.6165 | 861.2242 | 1174.5958 | 903.4897 | 926.2655 | 1702.1829 | 1769.1681 | 1703.3451 | 1916.9794 | 4587.6225 | 3063.4307 | 2791.6903 | 2755.9468 | 5994.8436 | 762.059 | 671.41 | 625.7758 | 762.059 | 784.7729 | 912.5723 | 748.7345 | 941.2242 | 743.2861 | 775.1623 | 766.9764 | 819.3068 | 658.2065 | 630.0413 | 668.2242 | 802.59 | 626.0619 | 626.0206 | 594.3864 | 653.6932 | 1508.7965 | 926.7935 | 974.5752 | 1103.9794 | 1070.8997 | 918.879 | 942.8377 | 833.7552 | 2131.6077 | 2085.8348 | 2062.2035 | 2183.0384
266 | 4.9204 | 757.057 | 1475.5713 | 1332.914 | 1961.057 | 1633.3712 | 1685.4838 | 1281.4862 | 1416.4857 | 788.5423 | 940.5418 | 753.2281 | 775.7999 | 1633.3717 | 1165.943 | 1454.8574 | 1699.4567 | 1535.3422 | 1145.3146 | 1084.9991 | 1220.1431 | 1271.4867 | 1894.6568 | 1898.9444 | 1637.971 | 2024.058 | 1780.4287 | 1489.3992 | 3014.9724 | 1187.4857 | 1189.3717 | 1133.1721 | 1515.8004 | 1062.6858 | 1378.5433 | 1432.4853 | 1170.5998 | 2766.3983 | 2497.3123 | 2184.2281 | 3382.972 | 1436.6868 | 1690.2871 | 1423.1721 | 1581.8574 | 699.4287 | 728.9715 | 699.2861 | 1539.8294 | 1249.8855 | 1535.3702 | 1224.5138 | 1159.6288 | 1376.4002 | 1648.9154 | 1104.0009 | 1248.0566 | 1419.7719 | 1658.7999 | 1149 | 1396.5147 | 918.5418 | 1172.5722 | 808.3142 | 903.9144 | 1399.914 | 1666.0856 | 1312.9144 | 1726.5989 | 1242.1145 | 1677.7719 | 1382.4872 | 2055.8579 | 1690.4567 | 1554.5428 | 1517.943 | 2958.4273 | 1587.7144 | 1329.3717 | 1178.0005 | 2700.1426 | 925.6283 | 985.7153 | 964.7144 | 1559.2581 | 837.9724 | 813.7424 | 821.972 | 855.4862 | 1140.1991 | 912.4567 | 909.6008 | 1704.5433 | 1724.7134 | 1700.8013 | 1913.5718 | 4710.6003 | 3052.6849 | 2810.2586 | 2661.1996 | 5879.2015 | 771.4867 | 635.9149 | 625.5998 | 771.4867 | 804.115 | 879.6298 | 744.8289 | 946.7999 | 725.9149 | 784.1431 | 786.1725 | 799.4292 | 625.6858 | 634.5423 | 655.5437 | 815.2006 | 630.3707 | 626.3427 | 614.3427 | 640.144 | 1502.7714 | 950.0575 | 938.5708 | 1093.3721 | 1041.5433 | 905.914 | 939.7709 | 829.9144 | 2152.4876 | 2112.057 | 2052.0295 | 2225.8279
267 | 4.9314 | 763.5388 | 1469.529 | 1354.3107 | 1982.6264 | 1641.4199 | 1779.9174 | 1260.1967 | 1419.3107 | 796.6381 | 973.291 | 772.1504 | 776.9587 | 1644.5173 | 1163.5025 | 1444.3107 | 1712.1091 | 1546.7522 | 1121.9587 | 1115.9852 | 1222.471 | 1215.8446 | 1897.2596 | 1851.9106 | 1645.9125 | 1979.0777 | 1779.7355 | 1514.0265 | 2956.4975 | 1173.0678 | 1177.8859 | 1094.4199 | 1489.3422 | 1051.4199 | 1369.3933 | 1443.7984 | 1181.5388 | 2840.3323 | 2604.767 | 2230.6313 | 3338.3736 | 1381.7719 | 1620.0413 | 1403.0099 | 1557.0314 | 699.0049 | 729.233 | 696.6677 | 1506.5074 | 1280.1652 | 1603.233 | 1261.9686 | 1141.6116 | 1354.1455 | 1593.3884 | 1080.5586 | 1276.4975 | 1397.8859 | 1695.5222 | 1148.1917 | 1375.233 | 956.6794 | 1120.4562 | 808.2645 | 903.0826 | 1417.5388 | 1666.9174 | 1310.4661 | 1775.2596 | 1236.8909 | 1635.41 | 1324.2842 | 2041.942 | 1715.7719 | 1566.2743 | 1499.8761 | 3031.6942 | 1567.2959 | 1336.2065 | 1155.7719 | 2717.4248 | 944.119 | 929.7306 | 961.8083 | 1485.0993 | 786.7719 | 831.4562 | 788.9587 | 822.0728 | 1164.9075 | 945.0462 | 876.6215 | 1694.8545 | 1764.648 | 1643.8958 | 1904.4297 | 4693.3835 | 3062.9754 | 2696.6283 | 2698.9735 | 5848.2675 | 735.7817 | 614.7404 | 624.4149 | 735.7817 | 768.0413 | 815 | 732.119 | 947.1504 | 705.5487 | 764.9174 | 735.9686 | 790.5173 | 620.0777 | 656.1091 | 613.9272 | 772.9538 | 662.2645 | 639.233 | 618.3422 | 594.7817 | 1512.5438 | 949.6677 | 977.9272 | 1057.5025 | 1025.6578 | 914.648 | 964.2281 | 826.119 | 2052.3884 | 2118 | 2029.6726 | 2232.7885
268 | 4.9424 | 752.6047 | 1452.1564 | 1364.028 | 2001.0545 | 1617.7315 | 1743.9144 | 1285.1696 | 1434.944 | 759.1003 | 922.5221 | 755.056 | 757.8304 | 1680.2935 | 1196.9455 | 1445.0133 | 1674.4779 | 1508.0443 | 1111.8451 | 1066.3245 | 1236.7448 | 1201.8304 | 1853.8023 | 1931.1534 | 1601.9012 | 1998.0575 | 1796.8732 | 1468.2817 | 2923.0545 | 1180.2567 | 1185.8171 | 1080.8732 | 1456.8156 | 1052.8039 | 1393.7035 | 1419.2685 | 1180.4646 | 2792.9159 | 2617.6918 | 2259.3215 | 3306.705 | 1407.4941 | 1623.9572 | 1433.3938 | 1546.3393 | 715.7463 | 747.5782 | 739.0133 | 1538.944 | 1299.7728 | 1599.8864 | 1282.8023 | 1125.8732 | 1324.8171 | 1603.4926 | 1158.5634 | 1267.8451 | 1405.8171 | 1775.3658 | 1146.2817 | 1376.9572 | 919.3953 | 1103.944 | 790.2817 | 908.0973 | 1434.2124 | 1619.9277 | 1337.0148 | 1751.8039 | 1272.8319 | 1593.4513 | 1408.1992 | 2100.9144 | 1730.7891 | 1556.702 | 1460.8023 | 3004.3495 | 1504.5339 | 1385.0841 | 1181.4941 | 2718.3658 | 952.6475 | 932.944 | 943.9985 | 1437.3673 | 796.1549 | 809.8732 | 797.7198 | 813.0428 | 1118.2404 | 973.028 | 898.3097 | 1730.944 | 1748.4794 | 1687.2817 | 1939.2965 | 4713.5089 | 2984.0575 | 2612.0856 | 2709.7861 | 5995.0708 | 765.1253 | 646.0841 | 591.3097 | 765.1253 | 780.4086 | 824.2965 | 753.3245 | 936.2537 | 726.9705 | 764.1284 | 762.1549 | 832.2094 | 612.8584 | 624.6755 | 628.5914 | 749.4794 | 630.4779 | 643.4926 | 601.8732 | 652.0148 | 1558.9174 | 990.6047 | 983.5767 | 1069.2537 | 1043.2537 | 900.4513 | 940.4779 | 838.3097 | 2069.8171 | 2123.0708 | 2083.1253 | 2186.326
269 | 4.9534 | 712.4739 | 1487.7699 | 1376.9341 | 1954.7001 | 1579.0039 | 1727.1102 | 1315.9961 | 1445.294 | 807.4159 | 947.2979 | 740.002 | 737.294 | 1618.7581 | 1230.3638 | 1407.2979 | 1657.9341 | 1559.0659 | 1123.1141 | 1107.0502 | 1170.2979 | 1190.9961 | 1866.8122 | 1863.8879 | 1587.2379 | 2031.0619 | 1771.9498 | 1476.6441 | 2873.1721 | 1237.4316 | 1205.3481 | 1080.646 | 1448.8683 | 1075.3579 | 1367.354 | 1431.2419 | 1151.9498 | 2786.2438 | 2632.1337 | 2176.882 | 3348.7001 | 1421.1278 | 1612.1741 | 1423.4081 | 1596.2321 | 708.824 | 761.1219 | 701.0039 | 1522.5398 | 1248.2901 | 1552.2419 | 1255.8279 | 1123 | 1324.6499 | 1609.3461 | 1148.8279 | 1297.0482 | 1392.4198 | 1742.1838 | 1127.0078 | 1380.7562 | 945.12 | 1117.234 | 800.4081 | 869.998 | 1461.5339 | 1576.5802 | 1361.176 | 1773.7699 | 1290.3579 | 1620.586 | 1442.0717 | 2063.236 | 1749.7621 | 1492.8781 | 1457.0541 | 2915.412 | 1477.6401 | 1359.7099 | 1188.6597 | 2691.0639 | 937.53 | 953.8781 | 920.6362 | 1478.41 | 799.88 | 800.8259 | 835.0639 | 817.4759 | 1150.824 | 963.884 | 906.0541 | 1765.9902 | 1766.176 | 1675.0639 | 1938.8918 | 4701.0294 | 3039.1121 | 2687.9843 | 2628.2321 | 5987.7138 | 724.0561 | 629.236 | 578.7679 | 724.0561 | 783.706 | 841.8259 | 751.7138 | 931.4159 | 699.2282 | 795.0678 | 748.2399 | 785.1701 | 602.5821 | 614.5261 | 613.5339 | 763.06 | 599.822 | 633.764 | 600.176 | 664.4159 | 1585.8977 | 968.4081 | 962.586 | 1060.0059 | 1038.7099 | 916.414 | 919.8181 | 826.3559 | 2116.3963 | 2145.4081 | 2041.7621 | 2243.0639
270 | 4.9644 | 703.3329 | 1500.1063 | 1391.821 | 1967.6406 | 1558.0629 | 1786.1706 | 1326.8294 | 1479.0797 | 767.9105 | 921.2434 | 730.4503 | 737.6937 | 1665.7385 | 1175.9454 | 1390.7385 | 1718.4587 | 1519.6755 | 1140.4056 | 1160.3063 | 1150.6755 | 1203.0546 | 1924.6671 | 1839.4951 | 1586.0084 | 1992.7035 | 1735.8461 | 1489.4867 | 2927.5664 | 1150.9175 | 1213.2532 | 1074.7483 | 1514.7846 | 1033.6671 | 1358.8559 | 1412.4405 | 1110.9454 | 2737.3692 | 2571.3678 | 2203.5035 | 3354.2895 | 1383.1972 | 1608.3161 | 1425.7846 | 1592.1091 | 709.0629 | 737.4503 | 673.9371 | 1493.9175 | 1269.3245 | 1532.5217 | 1229 | 1115.649 | 1337.9356 | 1609.9567 | 1125.3692 | 1346.9916 | 1373.3231 | 1687.8112 | 1104.4224 | 1417.6853 | 942.0266 | 1121.9371 | 800.3343 | 878.6308 | 1442.8825 | 1605.4783 | 1346.5413 | 1766.5035 | 1253.2615 | 1611.6406 | 1355.0364 | 2044.9105 | 1768.7566 | 1551.1608 | 1468.235 | 2906.0364 | 1539.2965 | 1353.6839 | 1122.0266 | 2644.379 | 931.8559 | 970.8112 | 971.6853 | 1528.5119 | 812.1622 | 797.0266 | 811.9636 | 799.3958 | 1145.2434 | 959.207 | 921.5231 | 1790.47 | 1797.7916 | 1658.3958 | 1915.9175 | 4563.5748 | 3121.3944 | 2754.5958 | 2648.6126 | 5967.7916 | 745.2881 | 615.8112 | 560.5315 | 745.2881 | 776.8741 | 841.0895 | 703.6126 | 899.8741 | 742.3874 | 749.8014 | 701.0098 | 816.9552 | 632.649 | 634.6937 | 603.1524 | 758.4503 | 601.0364 | 632.1804 | 604.8378 | 646.0448 | 1492.6126 | 990.0811 | 972.0182 | 1013.2797 | 1031.1524 | 891.0546 | 954.8378 | 816.2699 | 2152.1916 | 2155.442 | 2006.5427 | 2200.6671
271 | 4.9754 | 736.9646 | 1559.2212 | 1347.7522 | 1938.3452 | 1565.3186 | 1782.5486 | 1277.2478 | 1532.9734 | 753.2036 | 932.2036 | 740.1416 | 744.3274 | 1648.1858 | 1187.7079 | 1419.6991 | 1721.0442 | 1544.2478 | 1128.4071 | 1144.115 | 1171.1062 | 1196.708 | 1874.469 | 1879.1593 | 1610.4425 | 1959.1947 | 1754.8849 | 1460.7522 | 2964.1947 | 1218.2124 | 1154.4425 | 1053.3717 | 1473.1416 | 1029.9823 | 1351.1327 | 1475.9735 | 1140.5487 | 2733.2124 | 2590.3716 | 2235.4247 | 3303.2478 | 1419.8407 | 1548.4778 | 1384.4602 | 1535.7699 | 705.1681 | 748.0974 | 668.2743 | 1573.3186 | 1267.3717 | 1584.4159 | 1218.4867 | 1127.354 | 1397.1681 | 1567.3364 | 1125.9911 | 1315.8673 | 1432.761 | 1661.9026 | 1143.2035 | 1427.9204 | 954.5752 | 1104.1681 | 749.6372 | 862.0796 | 1463 | 1580.6638 | 1318.469 | 1813.3982 | 1232.2655 | 1562.9115 | 1340.1858 | 2027.0177 | 1751.1062 | 1581.9468 | 1448.0708 | 2932.6018 | 1583.9558 | 1366.3716 | 1163.885 | 2623.9735 | 936.239 | 951.5929 | 938.2743 | 1590.2831 | 806.0885 | 804.7964 | 829.3628 | 831.3628 | 1146.0089 | 961.1239 | 906.9292 | 1727.2036 | 1839.4247 | 1685.9026 | 1983.531 | 4716.7081 | 3192.4513 | 2673.9823 | 2666.2478 | 6054.6638 | 754.0442 | 601.6903 | 566.2301 | 754.0442 | 745.1593 | 871.7257 | 707.8673 | 905.4337 | 730.4867 | 768.4956 | 677.1151 | 802.2743 | 634.7965 | 623.1681 | 614.9203 | 756.354 | 607.2124 | 635.5664 | 613.7788 | 646.7079 | 1499.7345 | 965.0619 | 960.1327 | 1009.0708 | 1063.9469 | 875.7876 | 963.1416 | 823.531 | 2040.9823 | 2077.6194 | 1935.239 | 2179.7788
272 | 4.9864 | 752.3456 | 1463.7281 | 1381.7281 | 1968.7797 | 1589.4922 | 1686.5152 | 1262.1318 | 1473.4415 | 778.4194 | 946.1696 | 750.6544 | 741.471 | 1606.7862 | 1154.6691 | 1440.7281 | 1667.8083 | 1579.3309 | 1111.4341 | 1104.3604 | 1183.4341 | 1231.4046 | 1876.2636 | 1926.2719 | 1581.0664 | 1972.3972 | 1740.8009 | 1428.647 | 2909.3309 | 1199.2866 | 1158.8157 | 1058.3972 | 1501.5217 | 1064.0959 | 1332.0959 | 1469.9493 | 1159.5659 | 2697.412 | 2483.4184 | 2179.0442 | 3307.8903 | 1378.0221 | 1533.8304 | 1394.9926 | 1565.4111 | 673.3972 | 768.2645 | 668.6839 | 1559.4415 | 1268.0959 | 1585.412 | 1193.6839 | 1172.118 | 1365.0664 | 1640.2581 | 1113.9779 | 1294.5373 | 1404.0664 | 1683.4258 | 1103 | 1505.471 | 932.6323 | 1093.5733 | 755.6323 | 856.235 | 1457.0442 | 1609.7945 | 1326.4562 | 1800.9558 | 1215.0369 | 1543.177 | 1355.2719 | 2049.7207 | 1717.5364 | 1505.8968 | 1502.5806 | 3008.8092 | 1595.2645 | 1320.5364 | 1155.9935 | 2674.1982 | 962.9484 | 949.0442 | 925.118 | 1519.0885 | 810.6028 | 771.9705 | 835.941 | 838.6028 | 1167.9484 | 937.6839 | 937.0295 | 1751.0959 | 1809.1834 | 1668.7134 | 1974.2203 | 4796.3097 | 3124.8894 | 2671.7936 | 2671.2212 | 6098.353 | 737.4415 | 618.647 | 578.2793 | 737.4415 | 728.176 | 836.9051 | 727.5659 | 946.3382 | 708.9926 | 760.6691 | 710.3899 | 805.9926 | 665.4046 | 605.9558 | 561.6323 | 735.0369 | 652.0516 | 652.8018 | 609.4415 | 624.588 | 1491.1548 | 934.529 | 952.0147 | 1038.4341 | 1031.1475 | 893.059 | 950.1622 | 810.9115 | 2085.6839 | 1990.8599 | 2012.8092 | 2207.8673
273 | 4.9974 | 747.6273 | 1421.9488 | 1352.9892 | 1942.8043 | 1586.2113 | 1648.6677 | 1291.6057 | 1443.6568 | 738.8712 | 898.2842 | 740.9597 | 695.4208 | 1621.0747 | 1161.5497 | 1396.2035 | 1707.3028 | 1565.2438 | 1088.8043 | 1107.295 | 1161.1475 | 1217.8338 | 1957.174 | 1924.469 | 1571.5792 | 1970.7857 | 1779.9007 | 1435.7158 | 2896.9597 | 1198.6165 | 1181.8898 | 1031.3913 | 1526.9302 | 1033.882 | 1322.1288 | 1429.3757 | 1166.41 | 2703.4533 | 2564.7267 | 2154.9892 | 3267.9332 | 1375.0885 | 1565.2438 | 1412.4425 | 1630.1957 | 680.3648 | 758.174 | 666.059 | 1537.8928 | 1241.6568 | 1572.5792 | 1196.8928 | 1145.8338 | 1363.1288 | 1590.9332 | 1105.9302 | 1231.1072 | 1394.5792 | 1728.7532 | 1106.7748 | 1436.0855 | 936.1662 | 1111.7935 | 766.0295 | 895.6972 | 1431.2733 | 1582.2065 | 1289.9115 | 1787.9488 | 1229.3648 | 1509.413 | 1393.2763 | 2056.9627 | 1758.4907 | 1537.0885 | 1497.174 | 2969.6568 | 1600.6165 | 1335.7532 | 1091.0217 | 2729.1878 | 937.879 | 944.5497 | 881.3323 | 1496.0295 | 822.5093 | 781.3727 | 846.1662 | 840.2143 | 1172.3913 | 936.0885 | 917.6755 | 1725.3432 | 1833.528 | 1672.1288 | 1960.7315 | 4740.1397 | 3128.9193 | 2729.0295 | 2626.4425 | 6077.9007 | 723.0993 | 603.7532 | 575.8417 | 723.0993 | 779.236 | 793.6195 | 727.2035 | 947.059 | 706.5388 | 749.0187 | 680.587 | 790.1553 | 640.8525 | 612.1662 | 591.59 | 735.9813 | 632.5093 | 621.4907 | 581.3727 | 634.3323 | 1478.1692 | 957.9302 | 946.2842 | 1013.7453 | 1002.6677 | 892.0993 | 926.59 | 838.0885 | 2125.6117 | 2014.587 | 1987.0403 | 2229.587
274 | 5.0084 | 749.2655 | 1403.503 | 1337.4646 | 1941.3363 | 1526.6785 | 1698.7257 | 1253.6637 | 1440.8673 | 714.3982 | 887.5354 | 706.7581 | 696.062 | 1564.7257 | 1179.91 | 1389.565 | 1662.3835 | 1516.028 | 1108.6637 | 1113.8481 | 1167.8009 | 1230.7773 | 1870.7197 | 1873.0044 | 1572.3599 | 1993.1799 | 1774.4602 | 1404.5206 | 2915.3658 | 1178.3319 | 1126.09 | 1023.6637 | 1517.062 | 1023.5118 | 1364.7301 | 1460.966 | 1104.6062 | 2747.8245 | 2579.2699 | 2149.062 | 3325.4174 | 1393.0856 | 1523.4926 | 1441.2891 | 1599.6725 | 719.5354 | 705.1327 | 678.8864 | 1548.379 | 1242.0664 | 1551.6769 | 1213.7773 | 1125.7197 | 1385.469 | 1599.7197 | 1090.7301 | 1251.5458 | 1408.1563 | 1681.1136 | 1127.8629 | 1387.0856 | 922.8673 | 1085.9144 | 774.7537 | 904.3599 | 1399.3127 | 1627.2803 | 1279.0236 | 1758.8393 | 1237.966 | 1523.4926 | 1378.5486 | 2111.7965 | 1770.0664 | 1528.0708 | 1442.7109 | 2948.3835 | 1598.8157 | 1293.4454 | 1145.1755 | 2726.3215 | 886.4602 | 954.379 | 863.2655 | 1493.379 | 812.3555 | 766.5826 | 798.7433 | 817.1092 | 1164.3083 | 932.4026 | 914.8393 | 1740.2183 | 1836.1239 | 1715.0856 | 1887.0192 | 4756.6769 | 3131.6077 | 2693.6769 | 2612.9896 | 6042.6018 | 717.8761 | 571.0428 | 547.621 | 717.8761 | 774.7389 | 840.3319 | 676.1327 | 954.91 | 713.2847 | 741.8629 | 678.593 | 777.3275 | 608.2507 | 617.7065 | 621.7581 | 761.621 | 615.2463 | 613.1607 | 596.7965 | 622.3083 | 1527.6018 | 936.6873 | 927.3599 | 1022.1991 | 1017.8908 | 922.4218 | 944.3171 | 810.9424 | 2198.9144 | 2017.9248 | 1992.0944 | 2222.2611
275 | 5.0194 | 749.2773 | 1438.6126 | 1350.6323 | 1896.9931 | 1580.29 | 1707.8121 | 1276.6775 | 1427.8063 | 731.1485 | 908.5684 | 695.5162 | 738.768 | 1613.8968 | 1172.29 | 1429.8701 | 1672.0718 | 1464.3933 | 1108.1416 | 1101.4002 | 1162 | 1225.9095 | 1844.1554 | 1831.1288 | 1570.1102 | 1968.3353 | 1784.6381 | 1369.5812 | 2937.3667 | 1210.4327 | 1131.5035 | 1036.7483 | 1528.8771 | 1007.9675 | 1374.5998 | 1378.3933 | 1087.1682 | 2720.8643 | 2517.1799 | 2181.1032 | 3363.9872 | 1420.1485 | 1514.4454 | 1423.0836 | 1557.3225 | 680.9931 | 695.3933 | 683.232 | 1523.4513 | 1260.9164 | 1512.6844 | 1196.4061 | 1070.1032 | 1386.1809 | 1569.8457 | 1122.6647 | 1313.058 | 1382.7738 | 1649.5546 | 1145.645 | 1411.2065 | 922.6775 | 1090.8457 | 786.2645 | 875.6323 | 1409.2262 | 1654.5801 | 1290.9617 | 1763.2123 | 1170.471 | 1524.0069 | 1304.4779 | 2140.1485 | 1776.413 | 1515.3422 | 1449.5162 | 2928.6519 | 1579.8505 | 1276.3353 | 1161.8387 | 2692.9361 | 920.9095 | 937.9095 | 872.471 | 1497.5035 | 815.2065 | 780.6008 | 770.8781 | 807.1996 | 1166.0836 | 918.6775 | 915.9026 | 1739.1543 | 1885.4965 | 1731.8515 | 1929.4327 | 4759.9755 | 3200.1543 | 2641.8132 | 2532.4258 | 6089.4721 | 679.1682 | 571.703 | 566.297 | 679.1682 | 756.0197 | 828.3028 | 695.4454 | 946.1868 | 716.3608 | 758.5418 | 711.761 | 807.174 | 582.4454 | 634.761 | 606.8387 | 743.8898 | 609.9872 | 585.2517 | 610.0708 | 654.239 | 1504.5998 | 935.232 | 910.0325 | 1019.174 | 1017.5035 | 925 | 923.0197 | 780.7877 | 2168.9095 | 2084.297 | 1977.3294 | 2259
276 | 5.0304 | 738.2802 | 1427.9022 | 1338.9788 | 1858.1804 | 1583.503 | 1658.6623 | 1283.38 | 1433.6623 | 729.3589 | 946.3377 | 696.6391 | 749.879 | 1617.7006 | 1137.499 | 1352.0998 | 1705.495 | 1462.6623 | 1086.1613 | 1115.6794 | 1191.2631 | 1231.7621 | 1902.4607 | 1819.2611 | 1595.2379 | 1965.2016 | 1732.0403 | 1431.4415 | 2851.0212 | 1197.3146 | 1149.1996 | 1025.2399 | 1487.2591 | 984.0595 | 1348.5393 | 1330.6775 | 1124.3397 | 2729.2631 | 2478.9233 | 2172.3589 | 3355.0978 | 1391.7561 | 1547.8407 | 1405.7812 | 1574.9425 | 647.881 | 682.6411 | 672.881 | 1507.5413 | 1283.5393 | 1534.6371 | 1193.3034 | 1092.1421 | 1419.7006 | 1591.2974 | 1105.6563 | 1303.0595 | 1369.2822 | 1623 | 1123.119 | 1411.1996 | 904.2974 | 1128.6391 | 774.7389 | 855.2379 | 1422.3377 | 1584.998 | 1298.497 | 1738.9828 | 1192.382 | 1586.7621 | 1337.4587 | 2113.2763 | 1766.5182 | 1530.879 | 1443.7984 | 2967.3609 | 1486.3014 | 1283.4627 | 1146.378 | 2749.1996 | 915.3609 | 926.6603 | 892.4818 | 1531.9214 | 802.6583 | 813.0152 | 818.1381 | 780.9214 | 1122.9385 | 936.4012 | 885.9828 | 1734.4243 | 1858.6794 | 1712.3589 | 1961.1593 | 4843.2359 | 3152.0998 | 2611.2339 | 2564.8639 | 6178.5988 | 720.1401 | 564.2611 | 573.7389 | 720.1401 | 779.3992 | 808.1018 | 728.4607 | 930.4012 | 674.7601 | 737.5393 | 699.2611 | 785.1593 | 611.2802 | 617.7006 | 608.0998 | 694.4395 | 633.7812 | 612.7218 | 611.2419 | 661.0383 | 1516.1633 | 944.2631 | 934.7006 | 1017.3014 | 1043.9405 | 923.8599 | 948.2994 | 822.119 | 2157.2802 | 2058.6755 | 1993.0978 | 2266.2208
277 | 5.0414 | 716.8908 | 1440.8141 | 1337.8702 | 1912.8644 | 1594.6902 | 1689.6844 | 1277.3304 | 1455.6607 | 739.4632 | 897.2684 | 697.9086 | 721.2684 | 1612.0914 | 1122 | 1339.5841 | 1638.2006 | 1484.2684 | 1088.3451 | 1080.7138 | 1217.4218 | 1238.1062 | 1896.9144 | 1846.2006 | 1570.062 | 2020.9941 | 1752.3009 | 1447.1534 | 2886.6549 | 1162.0414 | 1121.8141 | 1031.8466 | 1475.354 | 985.8083 | 1310.5368 | 1317.2862 | 1122.354 | 2813.4868 | 2527.8614 | 2194.2331 | 3346.1092 | 1346.3009 | 1571.5546 | 1391.2595 | 1608.115 | 652.6224 | 708.2389 | 676.0531 | 1501.7138 | 1244.3598 | 1507.2626 | 1251.1003 | 1104.2065 | 1399.5752 | 1598.0797 | 1078.2095 | 1278.5221 | 1386.3451 | 1617.5074 | 1121.5162 | 1413.6313 | 866.7463 | 1108.7227 | 760.354 | 849.2862 | 1394.062 | 1588.5251 | 1262.9853 | 1789.6696 | 1249.8938 | 1569.9586 | 1329.1623 | 2068.7316 | 1735.646 | 1518.354 | 1460.5871 | 3005.7168 | 1498.9764 | 1321.8614 | 1128.9469 | 2718.6755 | 932.1386 | 919.2301 | 897.0678 | 1527.4985 | 790.5162 | 752.1239 | 785.469 | 800.823 | 1132.9558 | 925.9675 | 917.8377 | 1820.8319 | 1862.9469 | 1696.9617 | 1947.646 | 4820.8024 | 3174.8938 | 2573.888 | 2620.6844 | 6137.2448 | 728.2773 | 566.8761 | 547.9764 | 728.2773 | 764.646 | 799.1209 | 743.3156 | 918.7905 | 702.2862 | 739.9469 | 711.6844 | 773.6077 | 614.6077 | 607.7758 | 604.2153 | 685.823 | 642.7994 | 622.7374 | 626.8761 | 622.4454 | 1556.0147 | 972.3835 | 945.5693 | 1031.1534 | 1059.062 | 902.7758 | 957.0472 | 791.4838 | 2170.0236 | 2000.5546 | 1930.3598 | 2280.354
278 | 5.0524 | 690.5433 | 1450.8899 | 1369.6647 | 1979.7168 | 1552.5433 | 1683.4046 | 1242.3181 | 1427.7168 | 759.7458 | 892.9941 | 697.5319 | 703.6416 | 1585.763 | 1137.3756 | 1354.1328 | 1629.7227 | 1476.7571 | 1084.4681 | 1070.3466 | 1186.4277 | 1242.1038 | 1872.4046 | 1830.0349 | 1559.1504 | 2028.9424 | 1765.8845 | 1448.3525 | 2896.6013 | 1210.8903 | 1098.3063 | 1046.4277 | 1487.0924 | 1007.0693 | 1287.4219 | 1342.2197 | 1119.526 | 2880.1155 | 2495.8903 | 2228.6996 | 3335.6361 | 1380.9248 | 1536.7807 | 1339.8092 | 1583.6244 | 611.1273 | 713.3122 | 640.6244 | 1444.0059 | 1228.7282 | 1520.2429 | 1246.4567 | 1072.8092 | 1366.5433 | 1670.5954 | 1141.1849 | 1266.7571 | 1362.6416 | 1617.9017 | 1138.7399 | 1386.1332 | 885.971 | 1082.4105 | 772.0924 | 851.5609 | 1422.8034 | 1607.6071 | 1265.4336 | 1807.6878 | 1282.3122 | 1504.1155 | 1340.3122 | 2078.6706 | 1761.1328 | 1529.2832 | 1474.8269 | 3015.2718 | 1501.5895 | 1342.4912 | 1153.4508 | 2694.7571 | 949.5433 | 886.8382 | 889.942 | 1506.208 | 808.1445 | 776.4277 | 783.6416 | 790.0983 | 1181.1676 | 901.2311 | 896.9769 | 1838.266 | 1875.7168 | 1731.1559 | 1910.0118 | 4863.3643 | 3159.1622 | 2676.1618 | 2583.2487 | 6075.6996 | 709.2718 | 542.1676 | 519.029 | 709.2718 | 757.3584 | 786.913 | 714.7979 | 879.6878 | 731.266 | 757.5723 | 730.0866 | 772.1908 | 625.734 | 580.8324 | 601.3815 | 713.5374 | 612.029 | 593.5723 | 584.7689 | 605.4277 | 1556.3987 | 952.1155 | 951.6416 | 1014.5492 | 1063.9307 | 878.2601 | 952.6302 | 812.2945 | 2203.8265 | 2003.0059 | 1924.0345 | 2279.1445
279 | 5.0634 | 731.5133 | 1480.647 | 1358.822 | 1950.6558 | 1559.7433 | 1652.7315 | 1256.0296 | 1409.0787 | 727.8338 | 880.8947 | 674.2567 | 721.0905 | 1562.6558 | 1163.7522 | 1380.7404 | 1582.7433 | 1489.4956 | 1114.8309 | 1094.5713 | 1157.8977 | 1279.9007 | 1840.8977 | 1825.942 | 1535.7552 | 1982.0905 | 1745.1632 | 1455.4926 | 2863.0787 | 1206.0207 | 1135.3471 | 1061.2567 | 1508.1721 | 1038.5133 | 1321.4199 | 1347.3501 | 1078.3146 | 2815.7865 | 2510.1987 | 2184.1632 | 3301.7729 | 1422.0148 | 1489.9361 | 1356.8427 | 1569.3383 | 604.767 | 691.9941 | 623.4199 | 1460.6913 | 1272.2507 | 1514.175 | 1237.1721 | 1113.6067 | 1375.411 | 1629.9764 | 1135.8161 | 1287.8338 | 1369.6677 | 1643.0905 | 1132.6677 | 1334.7552 | 902.5801 | 1063.8368 | 768.1573 | 815.0118 | 1414.3962 | 1552.8309 | 1267.3471 | 1813.5801 | 1265.997 | 1507.3264 | 1334.4199 | 2051.1603 | 1783.5713 | 1515.819 | 1453.531 | 2967.5801 | 1538.5074 | 1351.233 | 1143.7404 | 2710.8309 | 952.9911 | 917.6972 | 909.819 | 1515.0727 | 803.0846 | 768.7433 | 796.5044 | 804.5163 | 1144.0639 | 917.2419 | 901.9213 | 1825.1691 | 1872.9213 | 1734.058 | 1900.8575 | 4788.3903 | 3122.8604 | 2699.0846 | 2527.1603 | 6013.6529 | 677.8397 | 541.3353 | 528.2685 | 677.8397 | 761.589 | 827.9852 | 685.5133 | 901.0059 | 710.6647 | 782.3501 | 741.1514 | 776.003 | 610.4808 | 594.1721 | 586.0757 | 745.7433 | 593.7522 | 604.592 | 570.8427 | 608.1662 | 1580.7493 | 947.4051 | 957.4169 | 1020.0905 | 1040.825 | 870.9911 | 905.4169 | 847.233 | 2251.7611 | 1978 | 1992.9243 | 2221.5535
280 | 5.0744 | 756.9911 | 1434.8584 | 1336.7212 | 1944.7434 | 1553.8717 | 1641.1903 | 1290.8053 | 1422.8938 | 725.9956 | 838.4425 | 690.4292 | 713.5354 | 1555.4558 | 1161.7035 | 1351.1327 | 1574.7257 | 1503.0354 | 1105.5664 | 1062.5664 | 1158.4956 | 1230.8407 | 1788.7035 | 1868.4026 | 1528.6858 | 1987.4115 | 1777.7744 | 1422.1195 | 2849.4248 | 1229.677 | 1160.2743 | 1061.9779 | 1521 | 1069.1416 | 1339.3009 | 1394.7257 | 1050.3098 | 2718.3053 | 2544.7964 | 2166.1283 | 3349.9911 | 1383.6061 | 1551.4513 | 1360.9734 | 1587.0089 | 638.1327 | 674.9956 | 646.8805 | 1473.5044 | 1278.292 | 1543.0133 | 1228.1106 | 1116.2079 | 1361.5708 | 1557.6902 | 1067.1017 | 1295.4381 | 1356.1195 | 1677.5974 | 1132.4248 | 1343.1372 | 899.7212 | 1078.4381 | 763.1682 | 847.7212 | 1404.1992 | 1541.4203 | 1295.708 | 1771.2345 | 1261.146 | 1487.9956 | 1321.3982 | 2038.5752 | 1761.8673 | 1492.7301 | 1507.2566 | 2962.146 | 1590.7655 | 1301.9867 | 1134.3053 | 2745.3452 | 928.1372 | 946.6504 | 875.5708 | 1525.1947 | 792.2655 | 754.2876 | 800.1327 | 816.3893 | 1125.4691 | 920.7478 | 897.9779 | 1794.0885 | 1918.3363 | 1698.8938 | 1924.9513 | 4756.0443 | 3190.4336 | 2717.4513 | 2519.2965 | 6038.9425 | 708.4513 | 565.3142 | 552.2566 | 708.4513 | 760.4071 | 785.2743 | 730.7345 | 893.969 | 714.5841 | 785.5177 | 682.2522 | 762.8274 | 572.2965 | 584.6814 | 583.4513 | 731.2876 | 592.1327 | 612.5531 | 590.854 | 615.1549 | 1585.4381 | 933.3053 | 964.4381 | 1039.146 | 1030.8717 | 874.8938 | 884.969 | 807 | 2246.6681 | 1926.9247 | 1980.0973 | 2203.3629
281 | 5.0854 | 747.234 | 1421.8722 | 1327.5477 | 1959.5211 | 1600.9951 | 1702.5428 | 1228.9892 | 1423.7709 | 716.7021 | 869.9843 | 705.1278 | 708.8142 | 1572.0531 | 1160.3884 | 1349.9469 | 1611.0482 | 1520.8722 | 1114.6863 | 1054.9095 | 1192.5103 | 1240.2822 | 1813.4366 | 1807.4729 | 1538.3196 | 1983.0374 | 1775.4886 | 1421.7503 | 2844.6755 | 1196 | 1130.5742 | 1064.9577 | 1512.1701 | 1052.9892 | 1339.7551 | 1375.8613 | 1065.3776 | 2699.3559 | 2485.585 | 2145.4946 | 3357.4523 | 1353.9902 | 1534.4729 | 1364.3727 | 1571.1593 | 640.9204 | 689.1013 | 672.8564 | 1468.1869 | 1258.8348 | 1549.7021 | 1237.8516 | 1089.708 | 1372.0374 | 1561.6067 | 1021.1701 | 1299.1436 | 1351.7768 | 1680.2497 | 1113.9892 | 1361.8407 | 894.0796 | 1083.0265 | 767.4572 | 854.883 | 1468.6863 | 1529.1436 | 1279.4041 | 1737.8034 | 1269.2979 | 1495.4789 | 1320.9312 | 2064.9577 | 1749.8082 | 1553.357 | 1462.0157 | 2932.3294 | 1618.0157 | 1327.0856 | 1128.1593 | 2747.7866 | 915.4946 | 920.3462 | 890.4523 | 1535.3559 | 804.4632 | 747.1701 | 800.2714 | 788.5054 | 1153.6116 | 942.6116 | 853.2763 | 1760.8938 | 1894.6853 | 1736.8564 | 1913.2291 | 4804.8299 | 3215.3088 | 2673.9832 | 2555.9312 | 6040.9458 | 708.2497 | 579.0796 | 534.3884 | 708.2497 | 733.9361 | 782.0639 | 744.1436 | 900.8781 | 740.5693 | 745.8299 | 685.9951 | 739.0905 | 587.7394 | 560.2979 | 598.2606 | 717.9892 | 594.9204 | 609.6863 | 585.7925 | 596.4198 | 1571.9252 | 942.1701 | 968.1436 | 1022.5742 | 1053.2714 | 904.9095 | 872.4523 | 795.5211 | 2215.5959 | 1923.4583 | 1968.1387 | 2249.1593
282 | 5.0964 | 774.9499 | 1441.6643 | 1309.3614 | 1936.3781 | 1627.6637 | 1674.2684 | 1214.5379 | 1387.0841 | 716.8068 | 905.7984 | 695.319 | 721.7645 | 1574.2522 | 1146.0418 | 1353.1932 | 1601.7979 | 1460.3441 | 1127.4621 | 1072.3279 | 1159.4538 | 1284.6136 | 1869.2522 | 1804.0173 | 1549.4198 | 2016.5128 | 1703.5462 | 1436.6553 | 2828.0251 | 1194.6386 | 1092.8402 | 1069.9493 | 1489.2016 | 1022.6553 | 1315.2856 | 1319.4872 | 1083.7734 | 2688.8491 | 2487.2606 | 2134.1849 | 3363.8402 | 1405.966 | 1506.059 | 1401.7227 | 1559.6052 | 626.9577 | 681.1425 | 674.0084 | 1548.4371 | 1235.1682 | 1509.8736 | 1273.8402 | 1120.0418 | 1393.7144 | 1610.3363 | 1055.8324 | 1274.3948 | 1352.5796 | 1629.9493 | 1100.8992 | 1395.6303 | 878.5462 | 1053.0418 | 745.3614 | 846.4705 | 1460.588 | 1534.3447 | 1234.235 | 1765.7984 | 1252.4031 | 1506.8235 | 1380.8658 | 2088.5546 | 1741.1514 | 1567.9827 | 1434.8235 | 2895.0841 | 1591.2773 | 1343.1425 | 1120.2355 | 2664.8486 | 904.6386 | 923.6971 | 902.7394 | 1527.118 | 828.0251 | 750.0674 | 795.2099 | 799.3614 | 1139.3697 | 907.0418 | 851.0423 | 1789.6219 | 1865.9922 | 1719.4031 | 1929.9493 | 4837.3279 | 3233.2856 | 2562.1425 | 2517.3942 | 5970.0507 | 695.6136 | 554.0167 | 553.6219 | 695.6136 | 731.353 | 791.4538 | 694.4365 | 908.6971 | 743.5713 | 725.235 | 696.3274 | 750.2606 | 581.8402 | 573.353 | 577.8402 | 674.9493 | 606.8235 | 597.9577 | 582.5462 | 569.4371 | 1546.2689 | 910.5796 | 953.8318 | 1020.6893 | 1052.7478 | 853.8062 | 892.4538 | 769.2016 | 2244.4287 | 1995.0167 | 1957.3525 | 2224.8992
283 | 5.1074 | 783.5014 | 1460.8997 | 1328.7109 | 1886.8495 | 1553.469 | 1605.0088 | 1235.3805 | 1401.3215 | 740.8466 | 883.1799 | 688.5664 | 711.3127 | 1570.3569 | 1144.8968 | 1356.1859 | 1575.236 | 1440.2272 | 1098.6961 | 1073.764 | 1175.8466 | 1260.4513 | 1870.0177 | 1845.8141 | 1521.5841 | 1982.2064 | 1699.7375 | 1450.9646 | 2844.4897 | 1171.5604 | 1104.0325 | 1056.3805 | 1464.7463 | 1024.3805 | 1305.0678 | 1354.2597 | 1115.5339 | 2688.6548 | 2518.8378 | 2146.4484 | 3348 | 1358.4749 | 1495.4424 | 1411.5251 | 1558.6961 | 631.6166 | 656.5841 | 658.2124 | 1514.3923 | 1223.5604 | 1489.1092 | 1210.4601 | 1095.1268 | 1399.7375 | 1643.9823 | 1105.8732 | 1260.3894 | 1329.3982 | 1587.0177 | 1119.1859 | 1393.4749 | 886.3894 | 1041.6431 | 733.9498 | 841.9912 | 1421.1445 | 1586.6844 | 1212.3127 | 1754.3658 | 1245.9056 | 1543.0826 | 1362.1297 | 2091.3215 | 1781.6431 | 1471.1887 | 1417.4838 | 2922.8378 | 1627.1505 | 1312.5251 | 1180.0915 | 2678.8879 | 901.6018 | 904.7021 | 882.3481 | 1534.3156 | 815.1268 | 766.0678 | 781.059 | 794.941 | 1157.3717 | 880.2625 | 901.8466 | 1773.8997 | 1927.2124 | 1674.2212 | 1931.295 | 4835.5014 | 3206.7551 | 2588.8526 | 2486.8614 | 6044.6756 | 717.1268 | 529.4159 | 550.0177 | 717.1268 | 743.6608 | 785.941 | 650.3894 | 852.4159 | 704.0265 | 703.3127 | 663.9912 | 764.5929 | 582.7788 | 589.3894 | 562.9322 | 712.649 | 601.1357 | 552.7463 | 572.6785 | 601.5752 | 1541.5428 | 919.0915 | 953.7876 | 1037.7463 | 1026.8141 | 850.1328 | 913.0413 | 741.9498 | 2269.059 | 2001.6431 | 1905.5929 | 2240.8555
284 | 5.1184 | 725.6897 | 1428.1298 | 1356.1755 | 1854.9159 | 1539.7936 | 1607.5241 | 1239.6057 | 1410.5649 | 726.471 | 867.3103 | 703.2138 | 708.9518 | 1610.7095 | 1183.2621 | 1354.649 | 1593.911 | 1471.2163 | 1075.3486 | 1064.6564 | 1160.0359 | 1218.7837 | 1870.7429 | 1837.7837 | 1513.9951 | 1921.2621 | 1713.2213 | 1435.1632 | 2869.1706 | 1151.7837 | 1114.9085 | 1036.2089 | 1479.3103 | 1031.4759 | 1334.6156 | 1378.4252 | 1081.5142 | 2656.6996 | 2529.6922 | 2138.6032 | 3357.089 | 1377.7886 | 1465.1323 | 1356.8121 | 1530.5649 | 667.4326 | 661.4326 | 652.7404 | 1455.7379 | 1209.0457 | 1469.8603 | 1183.7454 | 1115.8294 | 1378.7787 | 1589.2015 | 1084.2138 | 1257.0841 | 1334.3486 | 1604.9641 | 1117.649 | 1368.3004 | 859.644 | 1040.0383 | 707.0408 | 872.4833 | 1414.8269 | 1547.7639 | 1245.8294 | 1729.3053 | 1271 | 1582.3028 | 1347.0964 | 2065.1656 | 1838.8776 | 1486.5365 | 1412.0457 | 2915.5142 | 1645.2114 | 1312.3461 | 1175.9036 | 2699.4302 | 901.4351 | 925.8245 | 861.3461 | 1511.0482 | 835.351 | 759.2596 | 800.4376 | 792.7837 | 1153.3412 | 917.9641 | 890.3412 | 1751.6539 | 1954.7478 | 1732.0173 | 1942.56 | 4784.8653 | 3198.9592 | 2604.5958 | 2507.9901 | 6082.1706 | 713.911 | 518.3486 | 552.178 | 713.911 | 723.6465 | 791.4376 | 661.4351 | 832.7379 | 702.0049 | 727.2621 | 689.2213 | 748.9951 | 601.5216 | 587.9975 | 575.4808 | 746.4252 | 589.4326 | 546.7404 | 550.9975 | 620.6514 | 1543.5241 | 970.8727 | 949.2138 | 1010.2138 | 1010.1731 | 902.3894 | 900.6007 | 760.4858 | 2244.9061 | 1991.9061 | 1892.3869 | 2168.7207
285 | 5.1294 | 703.8692 | 1432.8505 | 1355.1681 | 1854.7709 | 1538.5368 | 1626.3599 | 1237.2665 | 1381.0324 | 686.0934 | 895.0187 | 700.7198 | 705.1357 | 1606.3637 | 1159.0698 | 1346.8505 | 1562.9016 | 1447.3549 | 1080.472 | 1052.1258 | 1137.7198 | 1246.5703 | 1897.8319 | 1839.0747 | 1537.594 | 1904.9203 | 1739.8505 | 1387.4159 | 2875.5703 | 1170.7385 | 1090.5467 | 1010.7011 | 1485.9203 | 1028.3412 | 1330.7099 | 1333.472 | 1068.5516 | 2728.0099 | 2438.4198 | 2076.6725 | 3402.3275 | 1412.3973 | 1491.1731 | 1354.2577 | 1558.4396 | 631.9764 | 650.5093 | 680.0984 | 1453.9253 | 1254.8368 | 1440.7011 | 1208.1121 | 1101.5791 | 1370.7945 | 1530.1544 | 1087.6077 | 1218.9577 | 1382.65 | 1634.1495 | 1117.2104 | 1388.7434 | 840.6077 | 1048.3648 | 707.7385 | 872.9577 | 1407.2242 | 1484.9676 | 1269.8505 | 1752.6451 | 1230.7659 | 1524.9527 | 1355.7099 | 2059.3462 | 1855.4297 | 1531.5044 | 1450.4769 | 2834.35 | 1615.5093 | 1322.1544 | 1135.8879 | 2697.117 | 915.2665 | 908.6775 | 877.0423 | 1490.3313 | 826.9577 | 765.7571 | 794.8643 | 796.0374 | 1147.0049 | 947.6401 | 883.5143 | 1754.1121 | 1988.6588 | 1744.8544 | 1927.3038 | 4831.711 | 3227.2104 | 2550.8319 | 2465.8879 | 6013.4995 | 707.4346 | 553.4022 | 545.9951 | 707.4346 | 692.5654 | 808.4346 | 689.9863 | 872.6313 | 687.2379 | 727.6027 | 730.5703 | 744.6264 | 605.9066 | 576 | 581.1308 | 747.594 | 580.9626 | 552.5093 | 583.1593 | 619.944 | 1501.5181 | 936.4434 | 905.9951 | 986.1308 | 1026.117 | 885.6401 | 885.5516 | 784.2429 | 2263.9912 | 1928.0698 | 1919.0846 | 2160.9401
286 | 5.1404 | 714.5207 | 1433.4557 | 1332.4203 | 1831.9528 | 1503.1239 | 1628.9705 | 1233.9056 | 1360.5502 | 706.1475 | 871.3672 | 687.8348 | 722.4852 | 1532.503 | 1114.9233 | 1319.2905 | 1572.6859 | 1419.4498 | 1075.9646 | 1038.621 | 1132.379 | 1259.438 | 1882.8584 | 1840.5325 | 1524.1312 | 1905.6387 | 1730.8997 | 1430.2242 | 2854.2374 | 1171.4203 | 1086.9941 | 1042.5915 | 1468.5325 | 1009.3377 | 1320.6682 | 1328.9646 | 1069.3613 | 2734.2197 | 2407.2714 | 2073.1947 | 3405.8053 | 1401.4026 | 1479.8289 | 1383.7758 | 1597.0472 | 618.0708 | 658.0472 | 663.7758 | 1478.118 | 1277.9233 | 1493.2124 | 1193.9528 | 1106.7154 | 1370.4321 | 1558.5561 | 1084.8584 | 1217.1357 | 1379.7109 | 1619.9646 | 1124.9528 | 1403.8466 | 874.0708 | 1104.1239 | 754.6918 | 851.5207 | 1393.9941 | 1510.9823 | 1266.4321 | 1760.4321 | 1221.1888 | 1503.68 | 1302.9174 | 2066.8348 | 1835.5207 | 1519.5915 | 1453.3495 | 2797.5561 | 1619.0236 | 1337.9823 | 1151.5561 | 2707.438 | 908.8879 | 908.0531 | 909.0413 | 1483.2242 | 785.9056 | 762.4026 | 768.4616 | 797.5207 | 1182.5502 | 929.941 | 891.3908 | 1762.0826 | 1986.5384 | 1690.6151 | 1914.8643 | 4906.0236 | 3205.2787 | 2553.4616 | 2442.8289 | 5932.8466 | 678.3082 | 577.9528 | 512.9646 | 678.3082 | 700.5679 | 813.5089 | 692.8525 | 892.3377 | 691.7626 | 711.941 | 723.32 | 735.8289 | 594.9056 | 570.4675 | 583.5561 | 742.1784 | 573.4439 | 562.562 | 607.3731 | 619.4793 | 1468.6328 | 914.6741 | 903.1416 | 985.0354 | 1035.4321 | 894.0944 | 915.0295 | 785.0826 | 2331.5738 | 1969.4249 | 1952.4026 | 2174.581
287 | 5.1514 | 736.5487 | 1413.119 | 1321.061 | 1866.2498 | 1530.6981 | 1633.7896 | 1206.1494 | 1354.058 | 711.9361 | 862.3963 | 669.7896 | 709.6952 | 1562.884 | 1072.6312 | 1286.5762 | 1587.5123 | 1441.9145 | 1092.6706 | 1046.119 | 1128.4877 | 1232.5123 | 1887.8535 | 1828.997 | 1498.0669 | 1929.5152 | 1668.1435 | 1456.1494 | 2888.3167 | 1173.9725 | 1070.0275 | 1033.7837 | 1437.4179 | 1022.1888 | 1321.177 | 1344.6401 | 1059.3963 | 2725.6431 | 2461.7591 | 2132.7926 | 3411.1279 | 1377.2714 | 1461.8201 | 1352.1829 | 1592.6342 | 615.058 | 672.1829 | 636.8201 | 1443.3844 | 1261.3933 | 1535.1829 | 1195.7896 | 1111.5703 | 1403.826 | 1586.0334 | 1073.881 | 1185.3844 | 1351.0944 | 1611.4238 | 1119.5762 | 1342.4483 | 894.2439 | 1083.2654 | 778.7561 | 887.9755 | 1390.4238 | 1534.7926 | 1300.3412 | 1755.6981 | 1271.5487 | 1509.2684 | 1317.3383 | 2053.942 | 1848.2743 | 1528.8171 | 1416.6647 | 2848.2193 | 1624.0305 | 1342.2134 | 1133.177 | 2686.1799 | 909.5792 | 943.7926 | 910.8476 | 1489.5703 | 767.4238 | 760.4268 | 792.9145 | 781.4208 | 1154.9331 | 894.2379 | 882.7896 | 1746.6312 | 1971.8751 | 1745.0364 | 1917.2468 | 4880.116 | 3230.8595 | 2578.4297 | 2393.9056 | 5980.6795 | 654.7591 | 554.5428 | 517.3048 | 654.7591 | 710.9695 | 793.8751 | 700.4602 | 875.8201 | 717.9666 | 692.2104 | 694.0305 | 744.6126 | 574.8781 | 591.2773 | 569.8141 | 694.1829 | 572.2439 | 573 | 578.2714 | 598.997 | 1520.884 | 937.8171 | 932.6677 | 980.6952 | 1014.7257 | 876.5093 | 913.8171 | 742.3234 | 2310.5093 | 2038.939 | 1914.3599 | 2113.2743
288 | 5.1624 | 743.4494 | 1422.6288 | 1308.7566 | 1847.9714 | 1568.6519 | 1633.5044 | 1204.4956 | 1381.1475 | 690.8348 | 851.7886 | 695.3535 | 689.4494 | 1591 | 1030.2291 | 1283.7247 | 1558.2842 | 1469.4174 | 1092.3712 | 1030.3073 | 1133.4494 | 1193.7886 | 1807.3834 | 1797.7016 | 1511.9538 | 1945.826 | 1667.2026 | 1425.4813 | 2841.7226 | 1184.7797 | 1110.3909 | 1017.3304 | 1439.3997 | 1039.4262 | 1298.1652 | 1335.4582 | 1055.119 | 2767.8029 | 2438.8756 | 2158.5275 | 3352.0442 | 1324.6322 | 1485.1883 | 1365.3855 | 1546.1829 | 607.3304 | 695.4086 | 652.8029 | 1415.2114 | 1246.3943 | 1498.2699 | 1222.9361 | 1083.5826 | 1432.8117 | 1554.467 | 1084.9449 | 1208.9096 | 1335.7104 | 1593.174 | 1115.1421 | 1379.1618 | 909.3855 | 1039.2522 | 749.3712 | 870.467 | 1394.8579 | 1553.1421 | 1307.6553 | 1760.4174 | 1258.4031 | 1498.6608 | 1397.2808 | 2074.8579 | 1871.2434 | 1506.6465 | 1399.6696 | 2858.5132 | 1606.5364 | 1309.4351 | 1110.1652 | 2693.4086 | 963.8722 | 949.4813 | 888.5914 | 1456.3073 | 790.8491 | 810.3997 | 823.5826 | 775.6608 | 1142.1013 | 907.9823 | 886.7247 | 1744.8491 | 1960.9681 | 1757.9538 | 2028.0429 | 4768.1223 | 3237.2468 | 2548.9164 | 2431.446 | 6007.2699 | 646.6465 | 546.8579 | 545.4636 | 646.6465 | 715.8029 | 789.826 | 691.9626 | 874.3943 | 712.6057 | 719.9273 | 708.1883 | 756.7886 | 585.2434 | 589.6234 | 548 | 686.2842 | 561.0088 | 573 | 575.2203 | 580.8899 | 1583.8171 | 947.2984 | 930.6145 | 966.7797 | 1029.2114 | 857.1332 | 909.5826 | 741.2257 | 2306.9592 | 1988.467 | 1902.9681 | 2182.1387
289 | 5.1734 | 742.6372 | 1433.3628 | 1320.2921 | 1793.3009 | 1553.8141 | 1642.8142 | 1209.7611 | 1368.3362 | 679.9026 | 873.6461 | 700.2832 | 675.6814 | 1574.2655 | 1048.8496 | 1305.9912 | 1569.7522 | 1446.9469 | 1080.5575 | 1050.7522 | 1126.6991 | 1175.6991 | 1789.8142 | 1807.2124 | 1482.6637 | 1975.6549 | 1687.1858 | 1377.2301 | 2787.9824 | 1156.1681 | 1117.3539 | 1064.0266 | 1448.2478 | 988.6106 | 1303.1593 | 1317.0974 | 1031.3628 | 2759.5044 | 2407.7611 | 2147.1239 | 3312.885 | 1304.5133 | 1520.2124 | 1339.0708 | 1523.6372 | 629.1947 | 692.1062 | 628.8495 | 1458.6726 | 1269.4691 | 1472.4159 | 1192.5752 | 1073.6637 | 1419.9204 | 1522.3363 | 1095.3363 | 1244.6902 | 1335.2301 | 1571.9823 | 1132.2124 | 1405.292 | 892.7876 | 1035 | 758.0708 | 828.0796 | 1384.8053 | 1532.4248 | 1311.6284 | 1754.1416 | 1241.7965 | 1515.3982 | 1376.4424 | 2054.5486 | 1849.9911 | 1487.6814 | 1433.4071 | 2887.5222 | 1610.5133 | 1336.6195 | 1138.9115 | 2694.4248 | 958.7522 | 941.177 | 889.1947 | 1510.2213 | 774.5929 | 801.4336 | 795.8495 | 764.3274 | 1155.0619 | 917.1681 | 885.2389 | 1784.1947 | 1964.1416 | 1744.8584 | 2012.6105 | 4809.6019 | 3218.0708 | 2521.3717 | 2472.2832 | 5982.4956 | 656.2921 | 540.5841 | 527.9911 | 656.2921 | 748.531 | 790.5044 | 650.5221 | 896.9292 | 720.3009 | 742.7434 | 690.3097 | 719.2655 | 586.1239 | 553.5841 | 569.0531 | 703.6903 | 570.9912 | 590.8142 | 571.9823 | 599.5133 | 1557.2389 | 943.1858 | 912.6637 | 1002.4071 | 1041.3628 | 886.354 | 874.292 | 778.823 | 2302.3097 | 1945 | 1896.9646 | 2177.6991
290 | 5.1844 | 756.6254 | 1439.2508 | 1376.3353 | 1786.031 | 1518.1042 | 1688.9184 | 1231.8452 | 1318.8225 | 701.1578 | 913.8648 | 681.5098 | 683.5944 | 1530.1858 | 1122.9297 | 1276.5491 | 1607.6677 | 1454.9494 | 1105.2198 | 1089.7719 | 1135.6254 | 1189.5944 | 1884.6057 | 1859.0226 | 1484.0423 | 1972.4366 | 1692.9804 | 1374.1888 | 2859.9494 | 1169.2311 | 1075.4789 | 1053.7266 | 1453.9494 | 974.8029 | 1321.5634 | 1306.8844 | 1038.9071 | 2672.9124 | 2441.9917 | 2159.6367 | 3305.6224 | 1360.0959 | 1517.3633 | 1320.9381 | 1570.2001 | 621.4675 | 670.4789 | 611.7296 | 1450.9464 | 1269.29 | 1460.3633 | 1214.2847 | 1107.6171 | 1400.7802 | 1565.1269 | 1123.7719 | 1193.707 | 1369.8338 | 1573.1775 | 1180.71 | 1409.8874 | 887.9071 | 1046.0423 | 753.8422 | 821.2704 | 1384.2508 | 1538.2198 | 1304.4562 | 1735.645 | 1288.5438 | 1528.2817 | 1335.6677 | 1993.9041 | 1837.5634 | 1445.352 | 1459.2704 | 2873.2674 | 1630.7606 | 1355.7689 | 1124.5604 | 2650.5295 | 912.0929 | 942.9267 | 877.0506 | 1488.8814 | 779.648 | 789.7409 | 757.1971 | 769.293 | 1173.7409 | 896 | 924.9917 | 1783.645 | 1993.679 | 1729.676 | 1946.2508 | 4885.2591 | 3281.2001 | 2530.8535 | 2494.9184 | 5999.5211 | 674.7915 | 545.4592 | 497.8958 | 674.7915 | 728.9351 | 771.8225 | 669.293 | 902.9154 | 712.4789 | 715.2281 | 671.0733 | 723.7719 | 588.1465 | 556.8029 | 580.9267 | 743.9071 | 601.7296 | 560.1745 | 562.6873 | 619.7606 | 1520.6986 | 910.8844 | 916.2508 | 962.5688 | 1003.6337 | 887.0929 | 860.1156 | 781.4902 | 2346.0143 | 1964.8761 | 1877.2704 | 2164.4396
291 | 5.1954 | 768.7729 | 1427.4238 | 1310.7443 | 1793.2881 | 1531.4169 | 1659.2527 | 1201.2035 | 1318.0797 | 709.8899 | 894.9882 | 665.6627 | 663.0865 | 1562.4406 | 1124.878 | 1257.1406 | 1572.3323 | 1462.4543 | 1094.0373 | 1131.5271 | 1128.0187 | 1176.9882 | 1860.1543 | 1795.8239 | 1485.6627 | 1968.0305 | 1727.4966 | 1500.2371 | 2805.3696 | 1198.6932 | 1109.3796 | 1015.8712 | 1465.4051 | 986 | 1296.056 | 1312.1898 | 1061.1593 | 2672.9558 | 2470.3068 | 2138.4916 | 3235.2645 | 1394.3864 | 1513.1101 | 1310.939 | 1560.234 | 611.3373 | 646.8102 | 628.1593 | 1397.9577 | 1221.7306 | 1434.6627 | 1252.7119 | 1120.9882 | 1416.4779 | 1574.5034 | 1076.9154 | 1229.8762 | 1410.2999 | 1599.883 | 1175.2153 | 1424.3255 | 924.8338 | 1057.822 | 727.6136 | 839.8338 | 1392.7424 | 1569.3678 | 1306.1711 | 1705.939 | 1271.5899 | 1505.0865 | 1331.939 | 2036.9371 | 1847.0492 | 1467.6185 | 1418.5713 | 2826.2576 | 1648.8034 | 1307.6814 | 1113.061 | 2592.8289 | 942.3304 | 928.4051 | 872.9322 | 1449.4966 | 799.7119 | 774.5713 | 782.122 | 781.6627 | 1177.761 | 911.239 | 913.234 | 1769.178 | 2008.0187 | 1709.4848 | 1933.295 | 4811.2527 | 3279.7001 | 2510.2763 | 2480.4916 | 5981.4238 | 705.1711 | 555.7424 | 502.0797 | 705.1711 | 662.9577 | 784.6254 | 660.2153 | 866.3746 | 713.0797 | 729.9135 | 661.8899 | 753.1101 | 588.9695 | 555.0187 | 561.8899 | 736.2458 | 612.5152 | 545.0118 | 542.5034 | 629.9017 | 1496.5713 | 891.3559 | 940.5457 | 903.5644 | 997.061 | 855.8102 | 849.7306 | 766.8594 | 2369.2645 | 1947.1357 | 1885.1101 | 2173.6814
292 | 5.2064 | 754.969 | 1382.3186 | 1306.6269 | 1766.8215 | 1530.6195 | 1616.4233 | 1213.326 | 1353.7183 | 695.5398 | 858.969 | 677.8776 | 656.7242 | 1597.9262 | 1102.1903 | 1296.3378 | 1544.3805 | 1407.0177 | 1087.264 | 1080.4896 | 1092.0486 | 1209.5959 | 1831.4971 | 1799.0871 | 1474.233 | 1929.441 | 1712.7419 | 1499.379 | 2793.7124 | 1187.3127 | 1128.6195 | 1043.8348 | 1462.0369 | 1057.5104 | 1304.6387 | 1367.9455 | 1016.5206 | 2733.8407 | 2439.0118 | 2118.9572 | 3279.2036 | 1384.3127 | 1467.4837 | 1317.1844 | 1548.1844 | 608.9262 | 624.5029 | 637.7301 | 1411.4853 | 1271.6269 | 1438.2271 | 1251.4971 | 1109.767 | 1422.6991 | 1540.8156 | 1071.6756 | 1269.5398 | 1425.6195 | 1639.0679 | 1151.8466 | 1399.1593 | 917.472 | 1090.5649 | 777.4794 | 844.5826 | 1395.0369 | 1565.2758 | 1311.3923 | 1706.4174 | 1251.7242 | 1477.9631 | 1329.5339 | 2055.3613 | 1856.1106 | 1525.7242 | 1421.9882 | 2830.3068 | 1630.4351 | 1288.5398 | 1150.0679 | 2594.528 | 959.4292 | 893.3186 | 869.2758 | 1517.2095 | 807.7242 | 756.0738 | 815.1903 | 807.7183 | 1210.4056 | 896.472 | 871.1962 | 1799.295 | 2032.0251 | 1733.1047 | 1893.0855 | 4703.2566 | 3244.7729 | 2491.6195 | 2479.9882 | 5855.0043 | 683.8643 | 549.9631 | 514.6505 | 683.8643 | 709.357 | 807.6505 | 670.8717 | 879.1047 | 689.8953 | 717.8215 | 685.0251 | 757.0796 | 599.3009 | 579.0251 | 553.3068 | 706.5826 | 598.6991 | 545.8525 | 547.4543 | 619.2758 | 1524.2095 | 919.8717 | 953.8835 | 938.0251 | 1017.3437 | 833.5029 | 875.4056 | 780.7183 | 2341.1165 | 1904.8525 | 1881.5826 | 2119.938
293 | 5.2174 | 742.767 | 1406.1741 | 1344.6313 | 1733.9321 | 1532.4238 | 1617.767 | 1231.5634 | 1374.0551 | 703.0128 | 854.4238 | 676.8643 | 667.233 | 1564.3176 | 1084.7964 | 1303.7542 | 1568.4917 | 1392.2586 | 1102.3009 | 1081.5979 | 1097.1485 | 1233.6313 | 1785.1691 | 1894.4071 | 1483.5467 | 1933.1062 | 1691.411 | 1454.4917 | 2786.7709 | 1176.822 | 1131.0128 | 1040.8515 | 1482.6696 | 1064.6479 | 1320.0423 | 1374.4405 | 997.4366 | 2738.1101 | 2388.7286 | 2084.5929 | 3303.4956 | 1406.2164 | 1465.4622 | 1323.7542 | 1554.7542 | 568.2497 | 624.3009 | 633.1101 | 1429.5762 | 1299.6185 | 1442.0551 | 1201.6352 | 1129.9027 | 1412.356 | 1516.5762 | 1104.767 | 1266.411 | 1419.767 | 1602.4021 | 1107.4149 | 1361.3176 | 933.3393 | 1087.0845 | 764.1691 | 807.4956 | 1384.4533 | 1566.5467 | 1286.3982 | 1720.5339 | 1308.1741 | 1494.4366 | 1327.9872 | 2045.2586 | 1880.3137 | 1473.2113 | 1407.4828 | 2774.8682 | 1675.7336 | 1341.3648 | 1189.3815 | 2689.3904 | 969.2036 | 903.0384 | 881.1485 | 1530.4956 | 807.0423 | 759.0551 | 795.4405 | 809.2075 | 1267.7375 | 885.8348 | 882.2586 | 1781.7709 | 2056.411 | 1769.4917 | 1904.6824 | 4736.2803 | 3252.7247 | 2547.0217 | 2506.1229 | 5926.9588 | 670.6146 | 553.4789 | 526.8899 | 670.6146 | 759.233 | 811.644 | 709.1357 | 864.2497 | 668.233 | 717.9155 | 672.3048 | 734.3304 | 608.4661 | 578.6735 | 555 | 703.6568 | 580.6991 | 540.4789 | 583.9705 | 611.7119 | 1559.8771 | 966.9705 | 945.9872 | 983.6146 | 1002.5634 | 875.119 | 890.9194 | 809.8899 | 2379.6402 | 1908.3137 | 1833.8938 | 2145.5979
294 | 5.2284 | 745.2026 | 1446.0192 | 1312.5501 | 1721.7847 | 1486.0875 | 1633.4307 | 1223.4115 | 1357.5693 | 703.3859 | 909.7036 | 697.0639 | 612.4862 | 1543.7974 | 1057.9744 | 1259.5182 | 1534.3028 | 1394.177 | 1077.7399 | 1112.9872 | 1132.6396 | 1226.2026 | 1792.2921 | 1888.5246 | 1474.3604 | 1959.6013 | 1668.1514 | 1444.9552 | 2788.6652 | 1159.7655 | 1102.5246 | 992.1322 | 1499.209 | 1030.5949 | 1300.7783 | 1370.6269 | 972.5182 | 2765.8678 | 2400.2665 | 2047.1834 | 3286.2026 | 1412.968 | 1475.5629 | 1327.6205 | 1550.8038 | 595.71 | 659.8678 | 629.6013 | 1414.5821 | 1307.4499 | 1450.2217 | 1203.887 | 1122.3604 | 1427.8294 | 1509.3987 | 1053.0875 | 1229.322 | 1402.9616 | 1568.6077 | 1114.2793 | 1378.0767 | 933.9488 | 1094.855 | 712.3668 | 799.2217 | 1370.791 | 1540.5246 | 1318.6844 | 1744.646 | 1316.7527 | 1512.8102 | 1314.9872 | 2006.8913 | 1975.179 | 1483.3176 | 1388.6141 | 2742.8166 | 1728.4115 | 1361.3604 | 1174.9361 | 2655.4159 | 961.7527 | 902.354 | 893.791 | 1477.7271 | 804.6141 | 751.5885 | 806.0575 | 806.6269 | 1217.6568 | 919.0575 | 861.9297 | 1752.3987 | 2085.4627 | 1800.8422 | 1930.6141 | 4781.5757 | 3270.8166 | 2584.3859 | 2441.0555 | 5914.5481 | 695.2217 | 542.968 | 539.4179 | 695.2217 | 755.5949 | 797.9744 | 710.5821 | 900.7292 | 697.0639 | 712.5501 | 652.6141 | 735.8294 | 586.7591 | 576.032 | 553.7974 | 713.8102 | 578.8038 | 568.4499 | 572.3412 | 628.8294 | 1485.4478 | 952.9361 | 911.9425 | 973.9488 | 1015.4563 | 822.6184 | 845.9297 | 767.1003 | 2377.9361 | 1904.5693 | 1829.0448 | 2060.9404
295 | 5.2394 | 743.5457 | 1413.9587 | 1341.7906 | 1799.8319 | 1497.6313 | 1633.1829 | 1217.1829 | 1397.6992 | 716.7935 | 881.8702 | 690.0029 | 636.5163 | 1536.2507 | 1078.2242 | 1268.3599 | 1485.5929 | 1416.4514 | 1073.1121 | 1118.8171 | 1118.6165 | 1231.9086 | 1838.18 | 1826.3451 | 1439.2979 | 1908.4601 | 1659.1357 | 1443.885 | 2823.7257 | 1160.6578 | 1090.0443 | 1019.472 | 1452.3008 | 1030.2271 | 1240.6194 | 1367.3422 | 973.9971 | 2733.6194 | 2416.0914 | 2073.9056 | 3252.6401 | 1409.1357 | 1451.115 | 1307.6165 | 1532.3658 | 601.7758 | 690.2035 | 637.3628 | 1417.59 | 1261.643 | 1447.7965 | 1221.7965 | 1113.0678 | 1453.5664 | 1474.6401 | 1028.4307 | 1198.8643 | 1405.1121 | 1597.9971 | 1173.3599 | 1362.118 | 905.115 | 1085.6165 | 725.3157 | 834.8378 | 1400.5192 | 1522.5221 | 1300.823 | 1708.6873 | 1343.7228 | 1535.0207 | 1311.6136 | 2007.3599 | 1953.0766 | 1512.6372 | 1432.882 | 2704.4365 | 1697.6401 | 1314.0265 | 1137.2065 | 2692.4898 | 961.1121 | 907.0207 | 885.4779 | 1458.0678 | 822.4985 | 749.2271 | 801.9115 | 801.5015 | 1239.2655 | 914.9115 | 851.6578 | 1748.3186 | 2081.6843 | 1735.4631 | 1920.8879 | 4693.0088 | 3268.6372 | 2508.2123 | 2418.7021 | 5810.2301 | 696.4779 | 547.1121 | 524.1386 | 696.4779 | 747.2507 | 775.8879 | 677.3894 | 879.643 | 693.0708 | 703.7493 | 679.0885 | 752.3628 | 585.0443 | 578.0236 | 576.9292 | 737.8614 | 610.0649 | 574.1829 | 562.59 | 615.2979 | 1496.2891 | 937.295 | 909.9528 | 983.7699 | 1034.8407 | 815.6992 | 866.9499 | 758.0885 | 2357.3864 | 1924.4514 | 1881.0413 | 2086.7612
296 | 5.2504 | 722.5983 | 1358.707 | 1366.2448 | 1821.3328 | 1541.3948 | 1605.0949 | 1212.2517 | 1349.5157 | 704.2173 | 855.2655 | 657.9793 | 679.6327 | 1553.9051 | 1131.0482 | 1286.2448 | 1465.8638 | 1428.7414 | 1082.6259 | 1128.7552 | 1103.8776 | 1238.7552 | 1864.0207 | 1805.7552 | 1442.5241 | 1907.9189 | 1704.9327 | 1446.9931 | 2781.6794 | 1149.2242 | 1092.619 | 1053.1293 | 1455.4223 | 999.707 | 1278.3481 | 1357.4966 | 1011.0275 | 2682.4759 | 2388.5914 | 2123.4155 | 3215.9793 | 1383.5845 | 1432.3604 | 1311.6534 | 1524.381 | 604.6534 | 642.176 | 628.2311 | 1431.7621 | 1306.2395 | 1469.5379 | 1237.9051 | 1091.2173 | 1442.0949 | 1501.8172 | 1070.0482 | 1248.1982 | 1452.1775 | 1623.1362 | 1146.1828 | 1309.4621 | 890.1155 | 1035.2035 | 728.1086 | 864.2655 | 1464.1844 | 1487.8225 | 1314.9327 | 1690.7689 | 1352.4759 | 1540.1155 | 1330.15 | 1996.4552 | 1904 | 1531.1569 | 1432.9656 | 2680.1293 | 1706.6672 | 1360.8447 | 1120.3673 | 2678.7743 | 957.4828 | 945.912 | 904.531 | 1461.8776 | 818.4828 | 743.7414 | 807.4017 | 799.8845 | 1275.9931 | 912.8914 | 848.3604 | 1762.2724 | 2105.8034 | 1755.4706 | 1933.9051 | 4621.9725 | 3264 | 2459.2448 | 2407.3397 | 5891.6343 | 673.5983 | 536.5983 | 514.7552 | 673.5983 | 752.3879 | 791.4086 | 643.7276 | 862.9051 | 692.3948 | 734.912 | 689.2448 | 775.15 | 573.2242 | 541.0811 | 565.7138 | 730.9725 | 602.7138 | 571.1293 | 553.6052 | 604.1293 | 1603.7223 | 960.0275 | 884.9518 | 951.1828 | 1050.3948 | 861.3948 | 914.0275 | 780.7621 | 2407.6947 | 1896.6863 | 1862.4415 | 2159.0069
297 | 5.2614 | 705.8948 | 1384.0325 | 1351.8751 | 1760.2438 | 1539.7896 | 1586.8948 | 1214.2763 | 1286.0462 | 687.4474 | 840.941 | 679.6313 | 704.6972 | 1537.0069 | 1130.3028 | 1263.5791 | 1457.1711 | 1444.1446 | 1075.3422 | 1088.2438 | 1106.2763 | 1253.1249 | 1847.4739 | 1802.8948 | 1430.6647 | 1927.6382 | 1682.7768 | 1451.0197 | 2675.6981 | 1132.8948 | 1100.2104 | 1005.8623 | 1444.0924 | 1020.5654 | 1338.8683 | 1317.0727 | 1028.7434 | 2664.3422 | 2371.4474 | 2103.349 | 3157.8623 | 1366.6382 | 1413.237 | 1364.2036 | 1552.5261 | 596.5791 | 632.3353 | 595.1121 | 1428.704 | 1302.3952 | 1537.0522 | 1251 | 1068.704 | 1408.5791 | 1535.4671 | 1098.0197 | 1220.0137 | 1457.7699 | 1613.3225 | 1169.9666 | 1337.6706 | 883.4474 | 1023.296 | 735.0393 | 831.4346 | 1427.4611 | 1489.2498 | 1329.1514 | 1713.6775 | 1337.5329 | 1515.5791 | 1372.2694 | 1994.2301 | 1908.4671 | 1553.296 | 1374.6057 | 2674.3422 | 1666.9282 | 1386.6844 | 1126.5723 | 2630.5329 | 957.0197 | 956.8948 | 874.6057 | 1495.5458 | 848.651 | 785.9469 | 788.3687 | 791.237 | 1227.0531 | 942.3353 | 880.9273 | 1699.2507 | 2139.2104 | 1791.2566 | 1909.9872 | 4603.6185 | 3216.7768 | 2494.651 | 2434.7562 | 5844.7049 | 615.4149 | 533.296 | 511.2566 | 615.4149 | 745.237 | 805.3618 | 636.5526 | 890.0393 | 723.9735 | 719.0924 | 677.4277 | 739.4149 | 570.9341 | 571.7759 | 564.6775 | 731.1249 | 610.6116 | 578.1052 | 550.7434 | 656.414 | 1597.5005 | 951.5791 | 917.6048 | 960.9273 | 1019.4346 | 870 | 913.8751 | 759.1976 | 2423.3225 | 1859.4277 | 1840.8289 | 2114.1386
298 | 5.2724 | 717.3097 | 1421.3584 | 1316.3805 | 1700.8761 | 1517.7389 | 1533.9159 | 1190.9336 | 1308.8717 | 669.0885 | 864.7744 | 686.5929 | 661.7123 | 1493.9823 | 1123.2035 | 1241.5929 | 1462.8053 | 1412.7212 | 1104.1726 | 1064 | 1082.9336 | 1232.031 | 1830.3982 | 1827.3186 | 1464.1815 | 1972.8806 | 1688.5222 | 1457.6018 | 2715.6063 | 1172.9292 | 1085.1858 | 1020.9292 | 1463.0753 | 1055.3584 | 1273.5486 | 1350.2877 | 1013.1372 | 2730.2478 | 2417.4823 | 2073.5442 | 3204.1505 | 1370.9027 | 1452.1815 | 1297.8893 | 1566.3496 | 602.562 | 668.4159 | 602.4646 | 1405.7876 | 1305.823 | 1511.3628 | 1227.5841 | 1086.1151 | 1419.7655 | 1562.0664 | 1087.6903 | 1234.1328 | 1425.8849 | 1556.2168 | 1139.5088 | 1317.469 | 902.1637 | 1052.7168 | 743 | 798.5929 | 1414.7655 | 1495.1947 | 1297.6327 | 1671.9159 | 1324.5929 | 1526.1151 | 1356.8274 | 2004.6504 | 1912.3009 | 1514.4203 | 1372.7655 | 2702.5222 | 1693.1903 | 1320.1106 | 1142.4071 | 2534.9866 | 969.4558 | 968.9602 | 887.7257 | 1481.4778 | 819.2654 | 767.7699 | 776.2035 | 791.1549 | 1250.4337 | 908.1681 | 923.1637 | 1724.9956 | 2076.053 | 1774.3407 | 1921.5222 | 4657.8894 | 3275.208 | 2517.3009 | 2445.9911 | 5744.6725 | 668.354 | 530.1947 | 519.0575 | 668.354 | 734.0973 | 753.615 | 690.6859 | 891.4956 | 739.6018 | 731.5708 | 663.5442 | 702.5929 | 559.7389 | 555.1194 | 542.4292 | 721.0885 | 609.0398 | 596.6106 | 565.0575 | 616.053 | 1529.6725 | 937.3982 | 936.9425 | 973.8938 | 1040.5797 | 914.8806 | 899.8451 | 718.6327 | 2374.0221 | 1870.2611 | 1880.7257 | 2081.146
299 | 5.2834 | 747.533 | 1424.6863 | 1338.7758 | 1708.5094 | 1505.0236 | 1516.941 | 1207.1603 | 1255.7502 | 657.3491 | 843.5732 | 664.4316 | 640.6155 | 1500.5448 | 1118.0472 | 1243.2901 | 1519.356 | 1445.6697 | 1077.5968 | 1056.0472 | 1093.8584 | 1221.6391 | 1830.3255 | 1825.4316 | 1500.9292 | 1955.2478 | 1699.3727 | 1417.2478 | 2667.2172 | 1155.5732 | 1075.9882 | 1020.1416 | 1407.4483 | 1032.8397 | 1256.1957 | 1376.6863 | 1024.2075 | 2710.3421 | 2439.0236 | 2098.1603 | 3307.6225 | 1369.0236 | 1442.5378 | 1258.6155 | 1502.3775 | 574.8869 | 640.5732 | 610.0236 | 1415.2193 | 1300.1416 | 1439.585 | 1203.7335 | 1085.4081 | 1453.8702 | 1547.8279 | 1103.8584 | 1207.644 | 1374.236 | 1594.9479 | 1135.4976 | 1312.2193 | 908.3609 | 1041.4788 | 739.6863 | 809.5684 | 1446.882 | 1544.6815 | 1294.6037 | 1691.3914 | 1334.2429 | 1508.8397 | 1366.8348 | 2024.2193 | 1884.5024 | 1452.5614 | 1451.9361 | 2738.5566 | 1713.7335 | 1320.1485 | 1106.9105 | 2567.179 | 981.9528 | 908.7384 | 904.6863 | 1416.9459 | 808.5684 | 745.9882 | 786.9528 | 793.6745 | 1251.8161 | 892.2783 | 890.9223 | 1764.3137 | 2167.5704 | 1755.0708 | 2027.1436 | 4644.5968 | 3301.7807 | 2571.6815 | 2443.6391 | 5770.7169 | 698.4081 | 539.9174 | 524.3373 | 698.4081 | 725.3727 | 765.764 | 689.8397 | 871.4316 | 717.8043 | 739.3845 | 666.9528 | 748.6933 | 552.9882 | 575.4032 | 533.2901 | 721.2783 | 571.177 | 555.9577 | 557.0826 | 594.9056 | 1526.5566 | 965.1603 | 903.8279 | 941.8279 | 1027.5496 | 869.4129 | 865.5142 | 719.5801 | 2410.3206 | 1861.1062 | 1851.9577 | 2095.2311
300 | 5.2944 | 719.8485 | 1437.1755 | 1323.5737 | 1728.0752 | 1455.7482 | 1549.001 | 1190.9488 | 1251.0512 | 659.3751 | 847.9759 | 668.1505 | 671.326 | 1508.2999 | 1108.5998 | 1253.0752 | 1508.8736 | 1404.4724 | 1046.2247 | 1031.0742 | 1112.0752 | 1246.2758 | 1810.7492 | 1798.0993 | 1444.8977 | 2003.5276 | 1739.2016 | 1413.5507 | 2620 | 1181.5767 | 1112.7763 | 1038.3761 | 1414.2267 | 996.3992 | 1276.7753 | 1328.4233 | 1009.349 | 2644.5486 | 2388.3982 | 2094.7743 | 3301.5236 | 1369.7001 | 1442.5507 | 1279.2006 | 1555.0532 | 597.8516 | 644.9759 | 636.3511 | 1448.6259 | 1306.2257 | 1455.1264 | 1220.9509 | 1051.324 | 1430.5737 | 1533.325 | 1042.4223 | 1178.9749 | 1373.5757 | 1632.0502 | 1129.7492 | 1338.8756 | 886.4243 | 991.8736 | 738 | 817.325 | 1424.5988 | 1533.5236 | 1284.4744 | 1710.2748 | 1306.8987 | 1500.7503 | 1389.1003 | 2009.6991 | 1899.7011 | 1439.7503 | 1415.7974 | 2683.1976 | 1674.9238 | 1332.6249 | 1117.7262 | 2602.9248 | 983.9749 | 894.2508 | 875.999 | 1425.4012 | 793.3741 | 780.0762 | 792.3501 | 810.5507 | 1242.3751 | 896.325 | 861.5747 | 1743.0491 | 2254.6008 | 1736.8495 | 1958.5958 | 4508.5958 | 3277.1494 | 2472.0953 | 2412.2487 | 5759.8987 | 678.4995 | 542.9498 | 528.0502 | 678.4995 | 715.2497 | 765.0742 | 654.0742 | 887.3009 | 729.6259 | 714.7492 | 655.4744 | 736.5236 | 567.5005 | 581.7994 | 562.651 | 671.998 | 568.1505 | 543.1505 | 559.4504 | 569.2487 | 1604.1775 | 973.2497 | 864.349 | 893.5737 | 1002.9248 | 834 | 862.1505 | 751.9759 | 2439.325 | 1859.4253 | 1856.0261 | 2010.1715
301 | 5.3054 | 729.8673 | 1430.941 | 1334.4926 | 1757.118 | 1465.3039 | 1557.1888 | 1165.0029 | 1283.8702 | 655.5015 | 836.6312 | 690.4956 | 686.7493 | 1502.1888 | 1070.9469 | 1258.7493 | 1486.8761 | 1411.5517 | 1118.2921 | 1014.8142 | 1119.1239 | 1245.3156 | 1850.4868 | 1800.6844 | 1398.5044 | 1982.702 | 1649.3421 | 1437.8083 | 2687.357 | 1155.2006 | 1086.0118 | 1019.3215 | 1373.4542 | 1018.3039 | 1311.9292 | 1310.1239 | 963.3215 | 2665.4897 | 2325.1357 | 2039.2625 | 3227.3274 | 1387.4956 | 1440.0029 | 1261.5693 | 1507.3362 | 582.9469 | 624.6961 | 601.8879 | 1361.7138 | 1294.3805 | 1488.8083 | 1236.8112 | 1047.6224 | 1425.6844 | 1495.1976 | 1064.0472 | 1171.1268 | 1412.5546 | 1584.5132 | 1139.9322 | 1298.8879 | 913.4278 | 993.9322 | 717.3805 | 814.9381 | 1437.8673 | 1515.6873 | 1259.1298 | 1641.3303 | 1294.8112 | 1505.3746 | 1368.9439 | 1969.1327 | 1947.6785 | 1466.3688 | 1398.4956 | 2631.3805 | 1711.9233 | 1290.4483 | 1150.1829 | 2541.5781 | 945.8849 | 899.8761 | 905.6136 | 1411.4454 | 801.2449 | 785.0029 | 787.5015 | 781.1976 | 1301.0472 | 884.3156 | 888.2419 | 1729.9381 | 2195.5811 | 1811.416 | 1901.9381 | 4472.5575 | 3289.9322 | 2467.3599 | 2325.8937 | 5751.9351 | 678.1858 | 538.2507 | 517.6283 | 678.1858 | 720.9351 | 775.6195 | 634.3776 | 853.6371 | 703.1976 | 683.0059 | 645.2507 | 710.4395 | 581.5605 | 566.4395 | 537.1976 | 681.7404 | 573.3127 | 540.0649 | 570.1858 | 548.1888 | 1662.4307 | 974.8112 | 889.1151 | 913.4278 | 1041.9263 | 829.8761 | 883.8083 | 741.3186 | 2440.3746 | 1866.4366 | 1867.2507 | 1973.9351
302 | 5.3164 | 724.1106 | 1393.5177 | 1293.1302 | 1724.9247 | 1510.7827 | 1566.8933 | 1145.3082 | 1309.4902 | 687.9804 | 818.3043 | 678.411 | 680.3043 | 1479.0118 | 1078.7866 | 1258.6008 | 1512.9843 | 1471.2768 | 1058.553 | 984.7153 | 1107.0079 | 1223.909 | 1819.8299 | 1785.0118 | 1358.1185 | 1974.091 | 1666.8619 | 1407.9208 | 2664.1302 | 1170.7788 | 1116.3717 | 1016.5929 | 1340.8024 | 1081.2729 | 1299.4149 | 1291.7114 | 952.8973 | 2743.1662 | 2385.356 | 2119.443 | 3217.3914 | 1333.4346 | 1437.0988 | 1276.5851 | 1480.9921 | 618.7709 | 639.3835 | 575.1027 | 1367.3717 | 1267.5098 | 1428.8377 | 1233.4031 | 1062.4941 | 1469.4784 | 1509.4823 | 1103.9921 | 1144.2134 | 1432.9961 | 1579.4902 | 1133.6086 | 1291.6918 | 948.8894 | 986.2094 | 699.6047 | 868.5694 | 1446.8024 | 1504.1067 | 1286.68 | 1645.2808 | 1357.1662 | 1480.8141 | 1414.6682 | 1979.7866 | 1939.2134 | 1447.6165 | 1362.6244 | 2584.2212 | 1729.7035 | 1321.6721 | 1126.5177 | 2505.9051 | 898.9169 | 884.7075 | 881.6244 | 1441.3756 | 821.1937 | 726.1302 | 765.411 | 761.9012 | 1276.6283 | 873.4031 | 865.6204 | 1706.6126 | 2205.9764 | 1779.4385 | 1914.2925 | 4528.5733 | 3243.0314 | 2489.502 | 2334.5772 | 5660.1538 | 617.3357 | 490.8259 | 523.1937 | 617.3357 | 699.8141 | 778.7035 | 648.8894 | 837.8973 | 648.2212 | 707.2808 | 631.4071 | 745.4784 | 586.3992 | 560.9012 | 575.9686 | 694.502 | 587.6918 | 538.0988 | 547.1145 | 573.9843 | 1570.4581 | 979.0988 | 933.2847 | 923.0039 | 1023.2212 | 895.1623 | 877.7075 | 702.3161 | 2477.3796 | 1882.6918 | 1823.324 | 2010.1819
303 | 5.3274 | 726.6785 | 1421.2901 | 1284.238 | 1687.4906 | 1484.8456 | 1567.7286 | 1181.002 | 1324.119 | 690.8977 | 805.0334 | 657.8977 | 689.8142 | 1492.7807 | 1092.5595 | 1239.4906 | 1487.2694 | 1404.996 | 1072.8162 | 1017.1377 | 1062.5575 | 1231.9499 | 1822.6283 | 1744.6765 | 1396.6805 | 2015.883 | 1661.117 | 1414.0688 | 2644.5595 | 1211.9164 | 1103.117 | 1022.7119 | 1404.4946 | 1054.3028 | 1241.7266 | 1293.6785 | 911.5742 | 2736.2527 | 2406.1858 | 2091.0983 | 3248.9164 | 1367.5114 | 1423.0501 | 1258.8122 | 1491.2714 | 557.7079 | 618.5408 | 622.121 | 1342.0147 | 1314.8162 | 1485.4278 | 1212.4906 | 1032.8289 | 1392.0294 | 1481.7099 | 1127.7975 | 1153.2213 | 1371.2173 | 1502.996 | 1127.7119 | 1266.6765 | 881.5388 | 976.1524 | 688.881 | 789.4866 | 1448.5595 | 1496.8643 | 1267.8289 | 1655.0167 | 1342.4218 | 1430.8456 | 1379.2005 | 1938.7433 | 1921.1691 | 1471.3068 | 1394.5448 | 2585.0856 | 1734.8309 | 1347.2714 | 1137.2045 | 2519.0856 | 934.4091 | 883.9833 | 830.9646 | 1440.7787 | 793.9646 | 748.5615 | 776.9332 | 763.8476 | 1264.9666 | 862.4572 | 867.0856 | 1689.1691 | 2228.9833 | 1743.5929 | 1924.9833 | 4515.5408 | 3336.4632 | 2458.8122 | 2378.3402 | 5522.1818 | 682.5468 | 521.5448 | 495.2527 | 682.5468 | 700.3904 | 753.2193 | 654.8643 | 857.7975 | 655.4926 | 709.8977 | 611.0501 | 739.2193 | 560.6598 | 540.0668 | 555.1504 | 680.1858 | 555.5575 | 522.6263 | 540.9833 | 586.7119 | 1521.0167 | 948.6765 | 909.4051 | 917.8643 | 1018.3904 | 852.8102 | 849.9312 | 722.0354 | 2468.0836 | 1777.9439 | 1824.2213 | 2058.883
304 | 5.3384 | 731.7242 | 1374.9955 | 1284.2065 | 1655.3776 | 1468.2758 | 1501.5471 | 1178.4469 | 1313.9646 | 659.9292 | 771.3082 | 646.4823 | 634.8923 | 1499.1032 | 1120.795 | 1172.6165 | 1493.7257 | 1431.8319 | 1024.7536 | 1026.3097 | 1077.416 | 1253.9322 | 1850.3805 | 1676.2374 | 1401.3436 | 2003.205 | 1621.6534 | 1404.1711 | 2677.1401 | 1194.4808 | 1076.6888 | 969.41 | 1367.9955 | 1054.5531 | 1229.5177 | 1282.792 | 898.3451 | 2749.795 | 2393.7581 | 2044.6195 | 3206.8584 | 1351.4454 | 1418.4484 | 1279.5885 | 1516.8982 | 602.2124 | 591.2404 | 584.0649 | 1292.584 | 1302.9631 | 1446.1327 | 1142.7197 | 1026.9661 | 1368.4838 | 1513.5207 | 1022.0265 | 1184.1741 | 1372.5192 | 1521.8658 | 1143.9322 | 1324.2464 | 892.2788 | 1011.2095 | 678.0339 | 762.4838 | 1461.587 | 1482.9646 | 1273.5531 | 1642.1372 | 1334.6903 | 1443.9676 | 1315.8938 | 1925.2419 | 1931.7596 | 1449.0663 | 1456.2788 | 2534.0649 | 1746.6903 | 1294.6858 | 1144.1032 | 2557.1401 | 956.1726 | 905.5531 | 857.1062 | 1402.584 | 806.3466 | 757.5855 | 798.0354 | 761.3791 | 1290.7257 | 858.2758 | 839.2389 | 1630.9616 | 2268.6578 | 1738.8274 | 1903.8259 | 4452.7551 | 3308.6843 | 2479.5885 | 2419.416 | 5498.208 | 652.6504 | 520.9985 | 450.1357 | 652.6504 | 691.2404 | 725.1711 | 596.0649 | 842.5501 | 613.5826 | 710.0693 | 591.9646 | 751.0015 | 559.4145 | 522.5855 | 518.1711 | 677.1726 | 562.4499 | 550.0369 | 544.4484 | 552.2389 | 1531.3112 | 946.1387 | 900.7935 | 861.9616 | 995.4808 | 855.8643 | 841.7242 | 761.795 | 2568.801 | 1767.7611 | 1809.5501 | 2079.5177
305 | 5.3494 | 737.8918 | 1372.3579 | 1306.3039 | 1670.3039 | 1506.297 | 1571.2695 | 1159.2163 | 1264.8112 | 663.9931 | 762.9459 | 683.7699 | 614.2094 | 1455.2841 | 1092.703 | 1184.6146 | 1465.9125 | 1435.115 | 1041.4523 | 985.4395 | 1084.5811 | 1231.2773 | 1791.4532 | 1750.006 | 1447.2901 | 1939.2369 | 1637.7768 | 1333.6627 | 2632.291 | 1177.3717 | 1028.5477 | 949 | 1346.2094 | 1047.2704 | 1190.0206 | 1229.8653 | 911.7837 | 2662.0481 | 2308.5142 | 2004.8584 | 3148.703 | 1336.7365 | 1364.5005 | 1258.1219 | 1468.5545 | 598.1219 | 608.8309 | 540.6421 | 1304.5674 | 1295.8918 | 1372.3382 | 1165.3983 | 1001.2773 | 1417.4523 | 1425.2103 | 1020.2429 | 1145.1288 | 1373.4258 | 1513.2773 | 1123.4867 | 1280.2979 | 874.8584 | 966.0816 | 689.7296 | 741.4867 | 1459.3717 | 1475.0541 | 1188.2035 | 1634.3176 | 1363.7768 | 1449.8446 | 1249.4464 | 1885.2841 | 1930.3717 | 1443.8378 | 1459.4798 | 2483.2773 | 1708.0206 | 1287.2566 | 1119.4326 | 2460.0009 | 962.1554 | 857.764 | 842.5408 | 1409.412 | 771.0747 | 750.3176 | 761.7571 | 760 | 1241.5545 | 883.7768 | 810.5339 | 1635.7768 | 2182.1022 | 1852.8909 | 1928.8781 | 4375.0816 | 3299.1485 | 2472.115 | 2361.7708 | 5454.6558 | 619.6352 | 499.0069 | 471.8781 | 619.6352 | 699.2566 | 756.297 | 602.7227 | 837.6824 | 643.3442 | 681.0678 | 592.8918 | 725.8584 | 543.115 | 510.1622 | 494.7974 | 655.1691 | 525.3382 | 534.2773 | 516.2773 | 519.115 | 1540.4189 | 947.7906 | 929.5133 | 863.8309 | 1017.4051 | 866.527 | 842 | 747.6961 | 2568.2841 | 1802.3579 | 1810.5742 | 2023.818
306 | 5.3604 | 704.8077 | 1358.2856 | 1285.7979 | 1655.7881 | 1469.8323 | 1553.3348 | 1140.7733 | 1275.9558 | 699.6996 | 764 | 644.0885 | 638.9607 | 1503.6456 | 1060.2807 | 1176.0344 | 1406.5762 | 1379.3299 | 1048.0147 | 964.5123 | 1083 | 1206.0246 | 1757.5172 | 1744.3151 | 1430.059 | 1941.7144 | 1645.5025 | 1322.2316 | 2509.4183 | 1160.0246 | 1066.9115 | 979.6996 | 1309.5614 | 999.32 | 1199.7095 | 1165.8274 | 923.4877 | 2702.6259 | 2335.9066 | 2024.4484 | 3100.5565 | 1362.4582 | 1401.1579 | 1258.7291 | 1385.852 | 588.9951 | 614.7537 | 534.4975 | 1238.8716 | 1292.0098 | 1387.4484 | 1164.0295 | 995.4926 | 1406.0442 | 1367.5319 | 994.5663 | 1113.0246 | 1369.2512 | 1461.0688 | 1051.1032 | 1254.2512 | 892.9509 | 939.0098 | 695 | 782.9165 | 1391.1081 | 1409.6057 | 1217.8968 | 1587.3249 | 1308.6057 | 1409.8127 | 1312.6161 | 1887.2217 | 1923.5074 | 1449.9951 | 1361.4036 | 2414.5958 | 1725.9459 | 1326.4435 | 1092.2807 | 2403.7783 | 907.8421 | 894.1579 | 845.4779 | 1377.0639 | 740.7733 | 725.7881 | 726.7832 | 753.2611 | 1202.2807 | 863.0491 | 827.4631 | 1605.3151 | 2270.2961 | 1741.9994 | 1950.734 | 4356.2463 | 3376.3845 | 2457.5123 | 2360.9607 | 5324.6843 | 608.7635 | 457.3102 | 457.0442 | 608.7635 | 652.3348 | 692.1278 | 611.5025 | 798.5663 | 642.2807 | 648.2856 | 591.2561 | 686.2955 | 517.2807 | 506.2512 | 467.2856 | 671.7095 | 517.9853 | 509.7733 | 523.9705 | 538.9558 | 1571.9509 | 945.5025 | 884.3397 | 900.4533 | 1002.0442 | 897.4484 | 860.7193 | 724.7733 | 2530.5369 | 1844.4435 | 1768.3249 | 1952.8323
307 | 5.3714 | 766.062 | 1408.0797 | 1271.4336 | 1641.9115 | 1488.0089 | 1479.6814 | 1094.9026 | 1251.5132 | 636.1769 | 732.0354 | 626 | 632.5398 | 1463.3539 | 1094.3363 | 1171.2832 | 1464.8231 | 1352.3451 | 1054.1328 | 987.8761 | 1086.8053 | 1202.5221 | 1754.7611 | 1685.3362 | 1358.6371 | 2028.9116 | 1590.9646 | 1366.3363 | 2483.3717 | 1170.9823 | 982.2566 | 1003.6991 | 1338.0974 | 1000.2213 | 1226.2655 | 1185.7699 | 909.2566 | 2788.885 | 2355 | 1948.9999 | 3087.4779 | 1374.8053 | 1385.7522 | 1245.4956 | 1424.3629 | 539.7699 | 611.7168 | 562.3983 | 1307.6107 | 1274.7787 | 1356.1416 | 1183.1151 | 961.9911 | 1361.9911 | 1393.7257 | 984.0443 | 1114.8496 | 1353.0177 | 1503.3186 | 1054.354 | 1292.8142 | 839.8318 | 928.6283 | 676.7345 | 730.7433 | 1332.469 | 1400.177 | 1229.8672 | 1515.6814 | 1336.4691 | 1387.8672 | 1296.6194 | 1853.4247 | 1922 | 1468.5044 | 1306.4071 | 2430.0089 | 1702.7522 | 1337.2389 | 1117.2036 | 2359.9469 | 903.4602 | 851.3539 | 824.1239 | 1358.6726 | 773.2921 | 721.0443 | 691.8407 | 751 | 1194.4779 | 821.0354 | 843.3717 | 1639.1859 | 2235.8937 | 1730.5753 | 1922.7964 | 4321.2301 | 3302.584 | 2397.1592 | 2256.3627 | 5366.9116 | 650.1416 | 441.1947 | 441.9115 | 650.1416 | 657.0708 | 720.7965 | 569.9026 | 814.6814 | 610.8849 | 654.6991 | 568.6902 | 701.354 | 532.3098 | 490.7787 | 434.8849 | 647.2743 | 524.0443 | 483.6902 | 491.9469 | 522.1239 | 1648.9735 | 916.8407 | 869.8053 | 937.3983 | 1011.2655 | 850.1592 | 851.0177 | 731.4159 | 2561.8142 | 1838.4956 | 1801.708 | 1976.3363
308 | 5.3824 | 753.4257 | 1372.6386 | 1266.9066 | 1636.9066 | 1488.1731 | 1456.5865 | 1142.3216 | 1234.8132 | 649.3476 | 704.2129 | 615.1731 | 643.4671 | 1461.0138 | 1048.2251 | 1113.9985 | 1457.0659 | 1403.1347 | 976.5712 | 952.6922 | 1052.4257 | 1251.7213 | 1737.9862 | 1714.3078 | 1391.0413 | 1963.5176 | 1714.7503 | 1321.7718 | 2455.0659 | 1193.0138 | 1045.8958 | 911.3308 | 1307.6922 | 1015.8269 | 1167.0383 | 1214.7871 | 900.1332 | 2649.4625 | 2179.4487 | 1978.4548 | 3045.2389 | 1333.4656 | 1392.7871 | 1214.4793 | 1415.1593 | 544.8805 | 624.9204 | 519.9587 | 1223.317 | 1235.1991 | 1349.9602 | 1135.3185 | 873.6646 | 1357.1868 | 1355.2787 | 992.7335 | 1092.2527 | 1363.4671 | 1455.2649 | 1083.6539 | 1246.2251 | 830.8269 | 919.8132 | 643.1593 | 741.4808 | 1312.4931 | 1386.9862 | 1160.4916 | 1507.2802 | 1291.6784 | 1376.4931 | 1272.9464 | 1850.2802 | 1849.3048 | 1395.1179 | 1361.6416 | 2386.8653 | 1651.0123 | 1333.1332 | 1098.3859 | 2416.0551 | 902.5865 | 821.9464 | 815.2266 | 1337.6662 | 699.1578 | 706.5329 | 671.3997 | 709.2389 | 1260.5084 | 834.974 | 791.0919 | 1599.8653 | 2282.3752 | 1746.8668 | 1833.3446 | 4172.3431 | 3411.977 | 2385.9602 | 2265.8545 | 5284.5039 | 596.7182 | 458.5605 | 436.1332 | 596.7182 | 687.2006 | 710.1593 | 580.2006 | 790.7458 | 626.2006 | 623.4257 | 592.9341 | 640.1715 | 478.6784 | 488.3201 | 433.1868 | 583.6386 | 504.1195 | 473.1332 | 459.8928 | 520.1868 | 1584.1578 | 926.5605 | 857.853 | 976.1608 | 991.4793 | 869.8943 | 859.147 | 735.7734 | 2576.3201 | 1820.6263 | 1793.026 | 1942.5988
309 | 5.3934 | 759.4985 | 1351.5005 | 1267.5713 | 1668.9971 | 1457.5024 | 1386.6489 | 1068.8652 | 1205.5024 | 652.1436 | 684.0728 | 564.0757 | 633.0728 | 1487.2124 | 965.0059 | 1152.7808 | 1432.2871 | 1372.647 | 955.3569 | 938.4287 | 1024.7158 | 1212.5762 | 1734.5713 | 1694.7886 | 1337.6489 | 1968.4258 | 1667.2222 | 1309.7139 | 2478.9971 | 1212.9272 | 1020.5762 | 897.9272 | 1326.4258 | 1031.5703 | 1116.5742 | 1196.4307 | 866.0029 | 2701.2056 | 2173.5674 | 1998.9282 | 3019.644 | 1379.9233 | 1288.0098 | 1160.647 | 1353.5762 | 548.6431 | 535.5083 | 482.2163 | 1208.8555 | 1225.7856 | 1347.2861 | 1100.1445 | 924.0649 | 1337.7876 | 1407.7798 | 921.9351 | 1110.1406 | 1274.5083 | 1454.8555 | 1019.292 | 1205.4307 | 826.1436 | 916.4287 | 628.7148 | 714.0747 | 1312.3569 | 1347.4316 | 1198.3521 | 1530.4268 | 1303.8545 | 1402.2832 | 1265.2861 | 1777.5781 | 1906.5645 | 1340.5742 | 1241.7265 | 2399.2832 | 1715.2075 | 1279.3618 | 1012.2212 | 2365.292 | 860.1465 | 797.7876 | 760.0049 | 1312.3589 | 705.4258 | 659.5752 | 683.7129 | 682.8584 | 1262.7158 | 785.4336 | 780.4995 | 1518.0059 | 2248.7905 | 1729.1445 | 1788.002 | 4282.0581 | 3272.3017 | 2434.353 | 2194.1504 | 5234.1445 | 617.8535 | 459.2861 | 417.7158 | 617.8535 | 638.2192 | 662.7178 | 552.4316 | 744.0747 | 590.5752 | 597.2871 | 569.7886 | 671.0669 | 466.4995 | 470.9302 | 418.5015 | 595.4976 | 510.5703 | 434.2891 | 431.2163 | 497.645 | 1631.2793 | 941.4277 | 934.1358 | 891.5083 | 1038.2095 | 829.9331 | 797.0059 | 752.4985 | 2525.1475 | 1842.1406 | 1786 | 1970.6392
310 | 5.4044 | 748.0413 | 1311.6991 | 1260.0206 | 1643.6888 | 1452.3938 | 1399.1195 | 1068.736 | 1201.1917 | 677.9277 | 694.3628 | 595.6844 | 620.2227 | 1468.264 | 962.1608 | 1089.0044 | 1415.233 | 1351.2227 | 982.3319 | 973.1092 | 1068.6741 | 1241.8865 | 1701.4867 | 1626.9528 | 1344.5339 | 1979.3938 | 1615.8702 | 1319.7773 | 2396.2374 | 1185.0826 | 987.6475 | 876.4661 | 1236.056 | 1053.3525 | 1145.503 | 1144.5177 | 831.4661 | 2667.3569 | 2268.1874 | 1971.2743 | 3000.0413 | 1425.5236 | 1298.6947 | 1148.7979 | 1350.969 | 575.9277 | 524.3628 | 477.5959 | 1185.2743 | 1194.8805 | 1303.7094 | 1090.2124 | 959.1608 | 1285.7197 | 1347.0044 | 974.6121 | 1075.5073 | 1311.239 | 1428.0826 | 1020.7463 | 1102.4498 | 797.6681 | 870.5177 | 627.7979 | 695.2227 | 1319.3835 | 1326.8289 | 1200.4351 | 1532.0103 | 1298.4351 | 1343.7713 | 1285.5443 | 1776.9484 | 1850.6002 | 1445.3069 | 1279.006 | 2323.2168 | 1690.5177 | 1257.8186 | 1018.9277 | 2306.9012 | 844.2124 | 752.9012 | 744.2021 | 1283.8599 | 658.9424 | 676.7257 | 700.767 | 681.3938 | 1205.3362 | 812.2906 | 779.6062 | 1490.4248 | 2349.7007 | 1731.9794 | 1790.7773 | 4171.3922 | 3233.3732 | 2371.1961 | 2251.5915 | 5204.264 | 569.9528 | 410.1239 | 394.6475 | 569.9528 | 596.6681 | 632.8495 | 540.6062 | 748.3628 | 608.5236 | 608.1608 | 529.0826 | 645.9012 | 463.2124 | 449.2434 | 406.0206 | 582.2537 | 467.7197 | 431.1814 | 418.6165 | 494.9897 | 1616.6888 | 925.6475 | 897.7507 | 886.7463 | 1022.6784 | 844.7257 | 877.1461 | 721.8908 | 2561.2699 | 1818.6784 | 1773.233 | 1950.8805
311 | 5.4154 | 741.5693 | 1314.1534 | 1245.0364 | 1616.3648 | 1429.5034 | 1419.9636 | 1109.0295 | 1236.0295 | 657.4523 | 663.9705 | 572.9705 | 635.0148 | 1504.1318 | 972.4818 | 1111.942 | 1459.8034 | 1438.9352 | 948.4887 | 935.8171 | 1102.8761 | 1261.9125 | 1667.0727 | 1588.5034 | 1325.883 | 1991.3432 | 1560.3864 | 1310.6568 | 2310.6125 | 1185.1023 | 996.7739 | 846.5034 | 1262.0443 | 1060.4307 | 1124.2114 | 1145.9636 | 797.4523 | 2714.3875 | 2188.5319 | 1937.2625 | 3003.292 | 1395.7296 | 1320.3432 | 1147.3795 | 1382.3579 | 542.4887 | 528 | 463.4159 | 1169.8977 | 1233.9931 | 1311.6352 | 1112.4966 | 917.6273 | 1327.475 | 1333.4818 | 958.9705 | 1040.6932 | 1274.3353 | 1412.708 | 1027.6205 | 1140.3875 | 784.5182 | 860.6205 | 566.4228 | 667.883 | 1340.4454 | 1317.5182 | 1122.4591 | 1466.1819 | 1273.1239 | 1335.9125 | 1238.3353 | 1688.8239 | 1843.9636 | 1374.3421 | 1295.5693 | 2347.3727 | 1683.0511 | 1255.1898 | 986.2989 | 2287.2773 | 859.2045 | 785.942 | 756.9636 | 1276.3795 | 707.3875 | 621.6125 | 651.3353 | 645.5398 | 1196.0511 | 823 | 758.7443 | 1481.5182 | 2210.4728 | 1742.1023 | 1823.7886 | 4119.2261 | 3287.8545 | 2388.4602 | 2222.5762 | 5141.2841 | 572.825 | 408.5329 | 389.1898 | 572.825 | 589.1898 | 615.8466 | 554.5841 | 705.0069 | 548.7512 | 614.4307 | 493.4523 | 627.8977 | 470.1023 | 433.6568 | 398.3284 | 581.8614 | 432.5034 | 455.6863 | 393.5034 | 447.3864 | 1539.1307 | 925.8614 | 916.2193 | 901.7227 | 965.5762 | 843.0875 | 842.8534 | 769.7159 | 2570.7591 | 1867.7159 | 1745.6932 | 1972.9784
312 | 5.4264 | 754.1603 | 1293.9695 | 1250.2252 | 1612.8225 | 1457.7232 | 1372.8264 | 1044.6184 | 1179.0211 | 678.1431 | 646.1298 | 544.4371 | 638.1775 | 1489.3245 | 911.4886 | 1089.9218 | 1369.0078 | 1422.1642 | 925.8397 | 886.3417 | 1006.0078 | 1225.3417 | 1728.4465 | 1605.208 | 1369.3511 | 1921.6184 | 1560.5152 | 1267.6966 | 2273.792 | 1192.7576 | 911.1681 | 877.1908 | 1239.4543 | 1048.6622 | 1061.0688 | 1096.3589 | 773.7271 | 2761.0782 | 2214.3511 | 1926.8874 | 2930.1509 | 1450.3339 | 1324.645 | 1136.3073 | 1298.523 | 537.9351 | 520.5973 | 478.0954 | 1168 | 1194.0039 | 1303.6622 | 1044.7959 | 854.3589 | 1281.6012 | 1319.3722 | 923.8569 | 981.5363 | 1284.208 | 1373.1642 | 972.0688 | 1110.584 | 773.1298 | 825.6317 | 622.7365 | 588.7959 | 1300.584 | 1293.792 | 1121.4504 | 1406.584 | 1237.5668 | 1324.0172 | 1256.4332 | 1774.7498 | 1866.7404 | 1361.6927 | 1236.6012 | 2242.8475 | 1688.9351 | 1227.8569 | 989.3378 | 2220.8436 | 864.645 | 837.9484 | 751.7748 | 1234.0516 | 694.8569 | 636.6106 | 626.6622 | 624.6622 | 1240.5934 | 748.1509 | 721.9218 | 1518.6583 | 2209.6583 | 1725.9046 | 1807.147 | 4202.3644 | 3254.7615 | 2360.8092 | 2204.4199 | 5117.3073 | 544.9218 | 382.2119 | 352.8092 | 544.9218 | 560.2119 | 603.1298 | 534.147 | 691.5324 | 540.4027 | 573.8741 | 514.9656 | 566.7787 | 449.792 | 415.3722 | 397.8225 | 545.1642 | 453.3206 | 416.584 | 404.4504 | 473.8358 | 1601.1392 | 955.7881 | 912.4848 | 839.8436 | 1022.269 | 847.5801 | 841.3378 | 703.2158 | 2601.2557 | 1816.8436 | 1744.1126 | 1914.1986
313 | 5.4374 | 780.3746 | 1304.0266 | 1153.4926 | 1521.171 | 1471.6785 | 1323.5988 | 1040.6873 | 1191.7139 | 623.8938 | 668.2095 | 547.5133 | 601.2684 | 1512.3835 | 964.7847 | 1007.0324 | 1366.2006 | 1383.1121 | 885.4336 | 870.4867 | 1036.593 | 1231.1918 | 1641.1533 | 1567.4247 | 1319.059 | 1971.1151 | 1547.1386 | 1285.7139 | 2325.767 | 1220.7139 | 988.8378 | 898.3569 | 1260.3835 | 1022.6254 | 1117.4543 | 1034.9115 | 739.1121 | 2702.8849 | 2130.6666 | 1800.7787 | 2925.6873 | 1422.2596 | 1313.3127 | 1105.6165 | 1276.4956 | 528.9823 | 511.4867 | 456.1209 | 1093.702 | 1176.3215 | 1234.2153 | 1051.5398 | 858.8614 | 1329.1062 | 1313.4956 | 936.3658 | 956.6431 | 1319.0531 | 1432.6195 | 1017.6018 | 1121.0354 | 765.9911 | 807.1475 | 568.7109 | 635.6107 | 1315.2095 | 1271.469 | 1124.1652 | 1424.5487 | 1277.7493 | 1313.4867 | 1223.764 | 1709.5103 | 1806.7197 | 1336.4513 | 1233.1829 | 2221.3392 | 1638.2419 | 1219.4956 | 945.2507 | 2213.6785 | 851.6431 | 760.1799 | 703.2507 | 1160.7197 | 680.6519 | 616.2861 | 619.826 | 606.1386 | 1193.233 | 745.3569 | 738.3746 | 1513.6431 | 2203.8082 | 1645.1976 | 1771.1121 | 4023.1503 | 3288.4101 | 2364.6873 | 2192.1475 | 4972.2477 | 535.4956 | 347.6165 | 360.8614 | 535.4956 | 543.1475 | 613.5221 | 501.4513 | 704.3569 | 547.0089 | 571.6785 | 487.6076 | 549.826 | 424.9646 | 409.4956 | 413.0266 | 537.8348 | 405.5722 | 396.1475 | 392.9734 | 406.8672 | 1589.6165 | 973.6873 | 883.2861 | 915.3245 | 1008.6165 | 804.7551 | 876.5575 | 686 | 2584.6254 | 1745.3982 | 1765.8702 | 1863.2684
314 | 5.4484 | 725.703 | 1247.703 | 1163.6485 | 1477.587 | 1480.6239 | 1299.9051 | 1031.1406 | 1232.5782 | 670.4499 | 608.6204 | 532.9051 | 659.2267 | 1473.1037 | 919.5501 | 1057.2267 | 1313.2443 | 1422.7434 | 937.297 | 822.4095 | 1040.2232 | 1271.2724 | 1646.413 | 1579.3304 | 1352.7434 | 1887.4306 | 1555.8594 | 1260.5044 | 2417.4745 | 1163.3146 | 886.8647 | 853.5624 | 1225.3392 | 1069.6731 | 1055.6081 | 1061.4253 | 723.6819 | 2652.7276 | 2169.6028 | 1820.8963 | 2983.0615 | 1411.4587 | 1314.3884 | 1082.2109 | 1339.5325 | 537.1652 | 517.6239 | 457.7768 | 1056.1283 | 1190.2478 | 1207.4464 | 1057.6942 | 866.2355 | 1327.4464 | 1297.7522 | 927.2935 | 1004.826 | 1268.2443 | 1358.7487 | 971.3972 | 1071.6327 | 824.297 | 816.8594 | 543.1406 | 609.8805 | 1289.5044 | 1303.5782 | 1101.1283 | 1389.3392 | 1286.1529 | 1289.1283 | 1185.5044 | 1660.6573 | 1887.181 | 1367.4253 | 1268.0369 | 2228.7768 | 1633.0826 | 1188.5044 | 978.355 | 2237.0246 | 850.6942 | 769.2601 | 707.5536 | 1263.2056 | 666.2935 | 616.0826 | 625.7768 | 581.8225 | 1190.471 | 761.5536 | 707.4218 | 1502.529 | 2273.8506 | 1686.7557 | 1792.1072 | 3986.5413 | 3213.1371 | 2289.0668 | 2135.7856 | 4941.4587 | 522.2935 | 340.3058 | 358.0703 | 522.2935 | 523.2109 | 588.8928 | 502.7768 | 689.2109 | 515.8102 | 499.3023 | 489.471 | 565.0949 | 407.4464 | 391.2109 | 380.4218 | 560.0246 | 416.1775 | 398.2355 | 372.2109 | 419.9543 | 1608.5659 | 992.942 | 903.413 | 899.9631 | 945.3972 | 844.2847 | 839.7979 | 687.6942 | 2558.8225 | 1801.768 | 1729.3392 | 1873.7891
315 | 5.4594 | 740.7817 | 1185.4365 | 1083.9203 | 1506.5162 | 1520.6726 | 1357.1574 | 1035.8751 | 1259.6254 | 687.0157 | 625.3599 | 526.5624 | 604.8269 | 1530.3137 | 948.532 | 1032.9056 | 1331.3599 | 1331.3107 | 925.6558 | 878.4543 | 1039 | 1326.9853 | 1650.1406 | 1583.8594 | 1312.8741 | 1925.8604 | 1522.765 | 1200.8584 | 2249.8091 | 1194.9695 | 929.4543 | 860.469 | 1164.7178 | 1050.6401 | 1046.297 | 1063.703 | 735.7503 | 2646.8594 | 2085.467 | 1932.1269 | 2918.233 | 1406.703 | 1250.5457 | 1045.4838 | 1260.764 | 555.3284 | 468.2959 | 446.9685 | 1086.3756 | 1174.9528 | 1233.3599 | 1050.2655 | 810.2802 | 1284.6086 | 1277.8122 | 890.6244 | 977.9056 | 1233.9371 | 1312.9213 | 979.0472 | 1093.7975 | 743.0452 | 814.703 | 549.5939 | 615.8909 | 1284 | 1308.2812 | 1153.7198 | 1388.0157 | 1300.6096 | 1253.2025 | 1191.1721 | 1655 | 1908.297 | 1344.7807 | 1273.1406 | 2139.764 | 1670.9538 | 1204.4848 | 928.4208 | 2233.2655 | 840.6873 | 739.6244 | 737.5005 | 1239.3117 | 662.2812 | 601.531 | 651.0629 | 559.9528 | 1179.9685 | 751.1091 | 650.2959 | 1496.703 | 2244.8898 | 1665.2025 | 1827.0787 | 3932.8584 | 3255.5792 | 2257.8122 | 2108.8122 | 4822.4041 | 504.531 | 349.4533 | 341.531 | 504.531 | 536.3284 | 586.5782 | 502.2812 | 673.1091 | 483.3589 | 532.4071 | 459.2025 | 542.2183 | 423.9066 | 363.9371 | 387.7503 | 508.999 | 397.6558 | 360.3274 | 390.4848 | 408.6716 | 1594.9528 | 962.4838 | 909.7188 | 900.0157 | 946.1721 | 853.8594 | 855.2036 | 668.234 | 2599.6883 | 1803.6873 | 1739.1878 | 1859.8122
316 | 5.4704 | 745 | 1238.0177 | 1173.7302 | 1512 | 1486.0309 | 1309.469 | 1001.261 | 1181.9336 | 660.1062 | 626.5133 | 532.1018 | 607.0752 | 1441.6283 | 960.2301 | 1023.5133 | 1369.9956 | 1360.3275 | 926.3584 | 855.0044 | 1041.615 | 1280.084 | 1666.562 | 1560.4646 | 1305 | 1912.3363 | 1517 | 1171.0796 | 2250.5089 | 1215.0752 | 880.7256 | 797.6239 | 1182.1505 | 1113.9912 | 1028.6947 | 1078.6903 | 724.9248 | 2712.3761 | 2059.5398 | 1825.2212 | 2801.3982 | 1378.9779 | 1178.1106 | 1054.8186 | 1238.1549 | 537.0796 | 447.7965 | 398.8009 | 1032.7256 | 1150.208 | 1250.2035 | 1028.9513 | 809.7168 | 1284.1018 | 1232.2876 | 894.5885 | 938.8761 | 1238.7168 | 1294.0531 | 1000.3053 | 1140.8407 | 734.9735 | 822.7168 | 532.6947 | 638.0487 | 1248.261 | 1295.7965 | 1040.0928 | 1409.0487 | 1248.9557 | 1250.615 | 1172.0796 | 1604.4424 | 1903.7699 | 1300.7743 | 1203.2654 | 2131.1018 | 1627.3141 | 1210.615 | 918.1283 | 2069.8672 | 812.8495 | 720.9248 | 694.9292 | 1224.1549 | 648.9248 | 617.3053 | 616.646 | 591.8673 | 1160.5664 | 723.7212 | 705.6328 | 1402.73 | 2281.5841 | 1673.9469 | 1864.2522 | 3916.1549 | 3270.2301 | 2219.261 | 2150.4558 | 4792.5398 | 469.7478 | 323.1062 | 328.5398 | 469.7478 | 533.7699 | 589.615 | 433.1371 | 670.1283 | 468.5398 | 577.4823 | 474.9204 | 559.792 | 393.8761 | 393.9956 | 337.6991 | 505.2301 | 400.9735 | 373.177 | 379.1814 | 416.4602 | 1549.2876 | 980.5354 | 881.2345 | 876.5929 | 974.1505 | 840.9248 | 876.4336 | 681.562 | 2631.4071 | 1762.7743 | 1724.5664 | 1864.8451
317 | 5.4814 | 821.0217 | 1315.9499 | 1100.6027 | 1495.2045 | 1462.3205 | 1345.1987 | 989.8122 | 1221.2704 | 676.3314 | 610.9725 | 546.2596 | 599.2763 | 1448.2154 | 942.4366 | 948.7463 | 1377.6519 | 1370.5359 | 864.2379 | 827.9007 | 1012.8289 | 1297.6352 | 1608.8898 | 1602.9666 | 1203.3431 | 1971.7621 | 1443.6302 | 1186.5634 | 2288.591 | 1194.0167 | 870.2321 | 772.9725 | 1132.9616 | 1127.5359 | 1068.4307 | 1107.5192 | 718.5801 | 2805.9774 | 2045.6244 | 1783.1327 | 2860.2537 | 1423.6185 | 1262.7011 | 1063.1878 | 1240.5359 | 553.4474 | 422.1327 | 376.9725 | 1063.7788 | 1096.6135 | 1264.3756 | 1019.8122 | 802.1603 | 1288.5359 | 1256.0551 | 826.1661 | 933.116 | 1234.6962 | 1356.53 | 974.7129 | 1063.0226 | 728.9282 | 792.177 | 541.4916 | 621.5526 | 1231.5084 | 1295.7679 | 1102.4415 | 1274.9842 | 1209.177 | 1254.5359 | 1199.9391 | 1696.889 | 1853.4975 | 1305.6077 | 1204.4916 | 2103.7129 | 1607.6244 | 1217.1878 | 945.4032 | 2101.6903 | 794.8564 | 741.9833 | 717.1711 | 1193.8289 | 634.6244 | 603.2045 | 595.0885 | 562.5251 | 1128.8289 | 744.7512 | 684.9449 | 1411.0993 | 2310.9833 | 1725.5025 | 1826.5693 | 3888.4808 | 3312.5467 | 2298.8614 | 2173.7955 | 4787.4641 | 495.0551 | 297.7846 | 334.0718 | 495.0551 | 514.4366 | 569.6686 | 486.6461 | 631.9892 | 492.6352 | 514.9341 | 465.6244 | 532.9449 | 372.2045 | 364.5251 | 352.9833 | 505.116 | 390.5084 | 392.7955 | 353.1327 | 388.8289 | 1634.9332 | 986.6519 | 855.7846 | 865.9282 | 1010.707 | 883.1987 | 852.4808 | 706.9833 | 2647.3756 | 1751.6962 | 1656.586 | 1830.6411
318 | 5.4924 | 809.4897 | 1271.2242 | 1111.4066 | 1521.6805 | 1431.3195 | 1365.3402 | 1013.1991 | 1174.2242 | 653.0083 | 634.0954 | 534.556 | 631.1618 | 1444.7262 | 911.3195 | 1021.4563 | 1355.5934 | 1442.7011 | 882.8879 | 810.6598 | 979.4233 | 1249.0167 | 1583.0747 | 1632.3029 | 1316.3732 | 2003.3029 | 1469.8506 | 1239.1908 | 2229.3652 | 1123.7802 | 870 | 758.4523 | 1226.3816 | 1088.5644 | 1115.2281 | 1002.552 | 704.556 | 2689.7305 | 2100.4646 | 1756.6971 | 2835.0457 | 1375.3279 | 1250.8009 | 1048.7635 | 1242.7925 | 506.7055 | 444.9917 | 391.1328 | 1007.1578 | 1141.9833 | 1153.967 | 1036.9253 | 759.7719 | 1349.0497 | 1247.4523 | 825.0664 | 941.9626 | 1259.0954 | 1313.9129 | 939.6308 | 1065.444 | 684.9793 | 797.8589 | 533.1411 | 609.1411 | 1243.444 | 1259.2532 | 1106.7262 | 1233.6971 | 1280.2862 | 1258.5851 | 1174.4233 | 1639.1951 | 1838.0374 | 1273.8382 | 1197.0374 | 2039.9503 | 1614.9626 | 1158.8466 | 963.3402 | 2117.0664 | 767.9046 | 722.5934 | 720.1037 | 1190.8963 | 625.8299 | 591.1411 | 587.6224 | 570.5477 | 1140.2365 | 713.942 | 723.1244 | 1370.9793 | 2308.6224 | 1654.9213 | 1782.4606 | 3697.6808 | 3315.3112 | 2268.7719 | 2158.0747 | 4756.527 | 487.3486 | 325.473 | 313.4897 | 487.3486 | 548.7468 | 515.0167 | 474.3859 | 609.0747 | 470.0083 | 483.5934 | 455.9336 | 481.498 | 407.6431 | 365.3776 | 353.3112 | 436.8839 | 386.3112 | 393.2075 | 354.4813 | 385 | 1560.0624 | 981.6224 | 918.4273 | 821.9793 | 1058.0207 | 875.556 | 802.3943 | 675.0457 | 2646.3112 | 1798.502 | 1677.5767 | 1817.0374
319 | 5.5034 | 708.876 | 1219.5722 | 1135.8053 | 1437.7787 | 1503.4101 | 1332.4749 | 987.8348 | 1137.1091 | 659.0856 | 656.0797 | 536.6342 | 594.115 | 1458.5369 | 917.9941 | 1001.9262 | 1333.9203 | 1406.4808 | 879.6401 | 813.5428 | 985.9941 | 1240.2743 | 1561.0118 | 1551.4129 | 1322.823 | 1961.4808 | 1468.5487 | 1217.7433 | 2188.3835 | 1208.6726 | 838.2006 | 800.6106 | 1115.3451 | 1066.7375 | 1053.6755 | 1045.4219 | 704.8171 | 2752.2272 | 2106.0914 | 1773.0797 | 2770.1268 | 1407.1593 | 1219.8348 | 1066.0797 | 1194.8466 | 570.9587 | 414.3835 | 392.0914 | 1042.7021 | 1207.0561 | 1192.7876 | 1037.1829 | 815.8732 | 1316.0236 | 1232.3717 | 883.239 | 895.7551 | 1268.3599 | 1316.177 | 965.9823 | 1068.8171 | 740.8732 | 785.3717 | 525.6401 | 575.292 | 1255.9027 | 1247.7316 | 1113.2684 | 1348.2036 | 1289.9086 | 1249.9144 | 1190.9882 | 1626.4572 | 1822.4631 | 1261.823 | 1176.9203 | 2001.2006 | 1621.4513 | 1121.1091 | 929.5663 | 2033.5044 | 773.177 | 746.3481 | 692.7433 | 1205.5369 | 628.6342 | 545.4808 | 586.0914 | 572 | 1165.6224 | 820.8437 | 700.7433 | 1325.115 | 2233.4985 | 1635.0973 | 1878.8496 | 3661.4631 | 3326.8112 | 2235.8348 | 2125.1091 | 4621.2595 | 481.4572 | 334.4513 | 278.292 | 481.4572 | 526.6519 | 575.3245 | 419.3038 | 679.6844 | 497.8909 | 512.7994 | 452.2743 | 512.3422 | 334.7728 | 344.1947 | 372.0797 | 456.2567 | 364.1947 | 370.2861 | 350.4572 | 420.4337 | 1641.7641 | 961.9203 | 939.6283 | 860.6106 | 1020.2979 | 892.1711 | 852.4219 | 691.8968 | 2630.5546 | 1757.6637 | 1681 | 1855.9764
320 | 5.5144 | 805.8182 | 1273.0133 | 1107.6121 | 1527.5305 | 1441.0157 | 1352.0212 | 934.3535 | 1152.4169 | 645.2665 | 638.6622 | 525.029 | 607.4961 | 1413.9577 | 861.9075 | 975.058 | 1305.2955 | 1407.5251 | 875.3166 | 829.6544 | 1049.6175 | 1346.8182 | 1699.8895 | 1587.2006 | 1254.7782 | 2042.6387 | 1397.0948 | 1222.3668 | 2209.8628 | 1212.4749 | 896.6967 | 776.4538 | 1161.9342 | 1025.4036 | 975.174 | 1079.5462 | 643.3033 | 2651.3401 | 1990.8942 | 1806.3088 | 2731.7704 | 1425.7335 | 1273.1716 | 1046.8205 | 1255.3801 | 548.5329 | 461.6465 | 416.8628 | 985.3033 | 1122.7571 | 1217.3378 | 1020.4248 | 729.9154 | 1296.3456 | 1195.087 | 909.2586 | 884.5541 | 1236.7704 | 1277.4036 | 949.6622 | 1066.2375 | 717.5329 | 803.3378 | 537.8918 | 591.3378 | 1228.4538 | 1270.942 | 1105.7124 | 1335.1372 | 1228.3033 | 1237.9498 | 1149.7202 | 1620.4749 | 1840.3378 | 1303.3589 | 1207.2296 | 2024.7045 | 1730.6599 | 1124.4459 | 937.0502 | 2088.5384 | 855.9553 | 720.4538 | 713.0212 | 1184.8997 | 592.1662 | 587.0423 | 603.4961 | 558.1873 | 1156.9498 | 718.7359 | 699.8417 | 1358.7547 | 2293.2218 | 1701.2218 | 1759.1607 | 3694.9922 | 3339.0502 | 2245.8918 | 2102.6622 | 4678.9052 | 473.6332 | 323.9498 | 284.2085 | 473.6332 | 502.8205 | 570.9498 | 431.4961 | 658.4538 | 494.5541 | 520.6042 | 459.3668 | 500.3456 | 347.2586 | 337.3958 | 305.8574 | 477.4169 | 339.8997 | 347.7414 | 349.0792 | 367.8284 | 1590.2242 | 944.3456 | 910.6121 | 898.1505 | 1042.7045 | 875.5831 | 833.1372 | 686.6332 | 2563.6199 | 1761.2876 | 1718.7547 | 1908.8048
321 | 5.5254 | 732.8839 | 1266.8024 | 1139.5261 | 1459.4827 | 1464.8604 | 1313.8751 | 989.7208 | 1162.3982 | 657.9971 | 609.0059 | 522.1337 | 587.5378 | 1470.6539 | 875.6627 | 925.41 | 1348.7237 | 1390.2704 | 863.8024 | 854.3284 | 951.4218 | 1282.2822 | 1600.0235 | 1527.5467 | 1267.6627 | 1981.881 | 1459.119 | 1245.3953 | 2115.8869 | 1226.9302 | 863.7414 | 794.529 | 1156.7355 | 1043.4622 | 1053.9155 | 990.5525 | 645.4651 | 2722.2498 | 2088.3108 | 1748.3461 | 2730.8663 | 1524.9794 | 1217.3461 | 1057.1308 | 1259.1337 | 548.7994 | 423.0757 | 367.6775 | 995.8633 | 1161.6569 | 1182.6076 | 974.2094 | 732.2645 | 1271.6716 | 1260.119 | 842.881 | 905.4622 | 1238.6657 | 1302.9273 | 939.6018 | 1034.2733 | 730.8633 | 800.3343 | 558.5959 | 527.6804 | 1311.8486 | 1273 | 1049.0118 | 1230.3549 | 1250.061 | 1180.0786 | 1234.648 | 1610.6686 | 1758.9508 | 1225.8171 | 1221.1976 | 2005.5378 | 1642.9537 | 1196.8516 | 969.7267 | 1973.6253 | 848.0698 | 710.5349 | 660.878 | 1094.352 | 586.2006 | 567.6716 | 578.8722 | 577.529 | 1106.5437 | 741.6598 | 711.1976 | 1387.1947 | 2297.1337 | 1675.2733 | 1789.058 | 3673.7384 | 3293.3431 | 2218.0727 | 2009.5525 | 4430.2535 | 491.6627 | 305.2704 | 295.2645 | 491.6627 | 506.5988 | 542.939 | 441.3982 | 623.3402 | 464.1396 | 507.0029 | 432.939 | 470.0727 | 364.8633 | 323.9361 | 330.7935 | 477.1337 | 381.8575 | 328.2704 | 334.0698 | 377.9302 | 1643.7876 | 957.9302 | 888.4041 | 866.4739 | 1011.4071 | 894.529 | 870.1918 | 638.41 | 2736.2292 | 1828.2527 | 1751.8604 | 1885.0059
322 | 5.5364 | 776.1623 | 1212.1106 | 1145.7817 | 1489.9263 | 1478.3451 | 1305.3274 | 975.0914 | 1161.1091 | 666.5634 | 625.9086 | 470.9469 | 594.5089 | 1512.8171 | 940.3437 | 942.7994 | 1283.9292 | 1339.8377 | 883.7994 | 793.6017 | 1002.6166 | 1268.4911 | 1612.7994 | 1509.764 | 1278.4543 | 1943.91 | 1405.3834 | 1196.8923 | 2147.7626 | 1190.1829 | 854.382 | 785.6003 | 1115.3466 | 1051.618 | 964.5117 | 1005.7449 | 662.0723 | 2693.0192 | 2088.4366 | 1835.7065 | 2630.7846 | 1490.4012 | 1195.0369 | 975.7477 | 1187.1474 | 552.7817 | 423.7817 | 388.5811 | 1006.4543 | 1164.0546 | 1210.2537 | 972.8909 | 770.8171 | 1370.2154 | 1197.8746 | 855.0177 | 905.1091 | 1256.0177 | 1261.5103 | 921.0369 | 1041.4543 | 733.8909 | 807.5634 | 527.8554 | 583.5074 | 1268.8377 | 1227.6194 | 1052.5634 | 1254.1991 | 1221.7463 | 1250.6888 | 1209.8009 | 1638.3628 | 1926.9588 | 1335.3423 | 1208.764 | 1952.9469 | 1715.416 | 1179.3097 | 920.0014 | 2021.7257 | 823.3643 | 712.8363 | 701.4351 | 1105.0177 | 646.5074 | 569.7817 | 587.3997 | 613.0354 | 1175.798 | 741.1637 | 675.1283 | 1416.4174 | 2372.6343 | 1658.8186 | 1737.1652 | 3614.3289 | 3224.7654 | 2231.1269 | 1968.0192 | 4497.0885 | 458.9646 | 322.9086 | 304.5089 | 458.9646 | 501.3274 | 542.8909 | 469.4174 | 607.764 | 487.5266 | 505.0546 | 419.6549 | 491.6357 | 331.9646 | 322.0546 | 310.3097 | 501.5811 | 387.8363 | 343.0723 | 332.0546 | 402.6357 | 1644.2183 | 941.0369 | 931.4351 | 830.91 | 1007.1091 | 848.7286 | 846.528 | 682.2714 | 2724.4189 | 1798.9646 | 1635.8215 | 1803.584
323 | 5.5474 | 767.5074 | 1228.1544 | 1096.1986 | 1420.171 | 1427.2832 | 1281.057 | 1040.0826 | 1134.1839 | 734.9981 | 591.5644 | 458.4228 | 634.2665 | 1506.3805 | 897.0718 | 972.7316 | 1359.4907 | 1325.5074 | 797.8475 | 817.7739 | 1044.3088 | 1303.4356 | 1702.1102 | 1506.1268 | 1210.0442 | 1965.943 | 1455.6323 | 1232.3088 | 2127.9725 | 1202.3658 | 906.6746 | 807.0275 | 1084.2684 | 1117.1249 | 1112.2351 | 1064.4632 | 632.353 | 2754.2095 | 2053.5221 | 1700.2153 | 2716.9411 | 1472.6765 | 1189.2114 | 1038.0404 | 1198.3235 | 561.6195 | 436.4504 | 360.3107 | 984.9725 | 1186.0275 | 1143.0442 | 1025.6746 | 729.9026 | 1326.1986 | 1221.7739 | 846.4228 | 901.1691 | 1178.467 | 1302.0974 | 921.9154 | 965.3825 | 741.6618 | 779.2684 | 546.1121 | 567.8456 | 1232.4798 | 1251.8161 | 1085.5625 | 1266.5349 | 1288.9558 | 1249.2537 | 1179.2684 | 1594.9872 | 1845.0589 | 1296.9872 | 1188.8456 | 1996.9282 | 1683.6067 | 1168.4228 | 934.2389 | 2112.1525 | 749.2133 | 721.6195 | 644.6214 | 1096.4228 | 628.9302 | 556.5919 | 547.7758 | 527.8475 | 1191.4926 | 728.5496 | 655.7611 | 1350.0019 | 2304.2133 | 1744.1947 | 1887.2351 | 3613.8732 | 3284.2095 | 2186.9872 | 2058.8988 | 4319.0756 | 433.057 | 330.704 | 293.5074 | 433.057 | 492.3805 | 543 | 416.41 | 608.9154 | 454.5221 | 498.296 | 430.4926 | 512.1544 | 351.0698 | 354.5625 | 319.5349 | 457.9872 | 374.5919 | 407.2095 | 326.2537 | 359.9449 | 1691.886 | 935.2114 | 916.7611 | 822.296 | 1019.4504 | 879.5202 | 859.3658 | 676.3805 | 2640.6361 | 1781.6765 | 1766.9116 | 1790.3805
324 | 5.5584 | 754.3899 | 1235.7901 | 1142.5005 | 1416.03 | 1408.5098 | 1312.9804 | 884.8883 | 1127.1799 | 674.9493 | 599.7001 | 503.5905 | 623.3899 | 1465.2596 | 888.2099 | 974 | 1332.9297 | 1398.7208 | 885.181 | 816.09 | 1074.1303 | 1373.8707 | 1596.3888 | 1528.3102 | 1261.3206 | 2019.3806 | 1447.3299 | 1226.2399 | 2007.6887 | 1260.2306 | 929.3702 | 782.7797 | 1056.8097 | 1051.1293 | 1020.0589 | 1022.3795 | 658.1603 | 2763.7901 | 2010.2896 | 1776.4509 | 2740.4002 | 1376.939 | 1166.6898 | 1019.6598 | 1178.6298 | 549.3899 | 467.0703 | 365.7901 | 1033.4705 | 1155.9597 | 1157.4602 | 974.6494 | 750.3102 | 1300.7197 | 1168.6794 | 894.5005 | 880.6298 | 1279.7611 | 1328.2503 | 969.5305 | 951.3299 | 744.91 | 765.3899 | 495.5894 | 561.1799 | 1205.7797 | 1209.3495 | 1089.91 | 1238.8697 | 1232.8294 | 1165.5791 | 1149.8697 | 1540.6194 | 1919.5708 | 1296.94 | 1193.8201 | 1940.7994 | 1691.7001 | 1182.5502 | 933.06 | 1946.2482 | 781.8904 | 702.5998 | 677.8904 | 1113.4602 | 581.4395 | 583.1603 | 578.0103 | 536.6101 | 1157.0796 | 685.3195 | 717.0807 | 1352.8201 | 2369.8707 | 1689.7994 | 1770.8087 | 3491.7787 | 3320.9504 | 2147.1696 | 2037.7797 | 4331.3702 | 425.2099 | 301.8997 | 290.09 | 425.2099 | 505.5801 | 521.6598 | 385.8697 | 622.5801 | 485.0103 | 510.6101 | 445.5502 | 517.8501 | 342.2999 | 331.7497 | 332.6101 | 483.1603 | 359.4498 | 386.7197 | 331.8201 | 338.5998 | 1646.4695 | 910.7497 | 918.91 | 866.6205 | 1012.2399 | 886.8201 | 856.12 | 683.7601 | 2625.3599 | 1699.5191 | 1652.7187 | 1785.1499
325 | 5.5694 | 777.5752 | 1268.4159 | 1160.6991 | 1484.7611 | 1375.5841 | 1243.2743 | 995.9116 | 1116.1947 | 647.4602 | 655.9912 | 477.4956 | 621.0354 | 1464 | 910.5929 | 997.5752 | 1278.9557 | 1322.4159 | 815.3097 | 801.2655 | 1012.1327 | 1316.0796 | 1566.4779 | 1527.0354 | 1237.4602 | 1994.4779 | 1457.8053 | 1184.7433 | 2102.2301 | 1246.2832 | 845.5221 | 811.469 | 1059.9292 | 1114.8142 | 1061.2036 | 983.6725 | 630.5133 | 2776.7699 | 1969.7079 | 1767.2124 | 2697.7787 | 1461.4248 | 1159.1239 | 1002.3009 | 1175.0531 | 532.3009 | 400.2212 | 376.8053 | 960.3451 | 1230.6372 | 1177.646 | 1003.4513 | 777.5221 | 1284.2832 | 1192.5398 | 817.4159 | 935.0089 | 1283 | 1300.5133 | 941.531 | 1008.9558 | 773.4867 | 771.8761 | 538.2036 | 558.0531 | 1264.9204 | 1186.3894 | 1129.2921 | 1218.354 | 1293.8673 | 1239.6195 | 1177.4867 | 1555.6991 | 1886.6371 | 1249.8495 | 1164.531 | 1991.062 | 1752.9027 | 1141.7433 | 941.8407 | 1939.0354 | 761.3894 | 739.3275 | 722.2213 | 1092.3894 | 619.2921 | 553.5487 | 578.0177 | 542.8938 | 1234.5841 | 730.1682 | 671.8495 | 1282.2743 | 2364.1416 | 1666.3894 | 1799.4159 | 3407.4513 | 3182.5132 | 2137.1593 | 1949.5752 | 4327.0885 | 460.3629 | 304.9292 | 281.1593 | 460.3629 | 506 | 520.0177 | 424.2921 | 588.6194 | 507.6106 | 477.6017 | 440.1062 | 533.7168 | 376.3806 | 354.5752 | 339.8761 | 508.5575 | 383.5575 | 303.4867 | 345.7522 | 351.7522 | 1664.646 | 984.6549 | 938.646 | 889.6106 | 991.3717 | 916.469 | 904.1328 | 692.8407 | 2783.1505 | 1715.6637 | 1629.354 | 1863.5841
326 | 5.5804 | 818.7783 | 1228.2217 | 1125.1947 | 1430.3028 | 1503.3024 | 1337.4809 | 975.1244 | 1113.0162 | 606.2217 | 557.5407 | 488.9351 | 609.0649 | 1455.0487 | 912.9892 | 910.4759 | 1271.0379 | 1417.4755 | 882.6269 | 840.7837 | 991.1081 | 1240.4056 | 1566.9946 | 1478.265 | 1235.0108 | 2001.9567 | 1518.6701 | 1193.9459 | 2095.0487 | 1206.2163 | 888.7567 | 791.1136 | 1117.6864 | 1093.1244 | 1058.0216 | 1034.7188 | 633.9784 | 2723.292 | 2028.6756 | 1788.881 | 2657.2163 | 1471.9513 | 1125.1839 | 969.1784 | 1170.027 | 530.0108 | 418.8919 | 352.1352 | 1011.7134 | 1111.6543 | 1229.7188 | 1002.0108 | 793.9135 | 1239.2433 | 1170.1244 | 871.6972 | 864.3893 | 1173.5948 | 1294.0324 | 946.9676 | 1004.0324 | 718.3028 | 715.3082 | 567.8432 | 568.9405 | 1225.2217 | 1226.7783 | 1029.5462 | 1231.9243 | 1260.1893 | 1215.1406 | 1159.1027 | 1549.0379 | 1948.6593 | 1303.7026 | 1168.973 | 2010.8973 | 1712.2271 | 1180.7837 | 902.2163 | 2037.4647 | 814.708 | 695.2433 | 685.2055 | 1121.8378 | 642.8756 | 555.9838 | 540.2055 | 558.9135 | 1165.3839 | 719.0649 | 654.0919 | 1385.4322 | 2335.1568 | 1667.9892 | 1751.265 | 3460.7026 | 3320.2375 | 2148.9351 | 1979.827 | 4293.1839 | 440.1136 | 298.0379 | 280.0054 | 440.1136 | 509.9784 | 544.8648 | 380.2433 | 586.0108 | 476.173 | 500.8702 | 433.0379 | 485.265 | 356.1136 | 334.1136 | 283.3082 | 417.4975 | 407.8702 | 299.0162 | 328.0973 | 355.9784 | 1628.2001 | 929.3082 | 903.1947 | 859.1676 | 1014.8702 | 879.2055 | 811.5083 | 665.1514 | 2708.4218 | 1727.9351 | 1659.8324 | 1800.3515
327 | 5.5914 | 810.8623 | 1262.4149 | 1128.9518 | 1447.0069 | 1502.5526 | 1231.4337 | 886.0964 | 1166.6145 | 668.9862 | 576.0964 | 539.4287 | 601.2065 | 1455.2616 | 825.4337 | 948.4956 | 1312.6352 | 1413.0964 | 850.3992 | 771.1446 | 1012.117 | 1310.9312 | 1532.296 | 1521.6352 | 1182.709 | 1973.2753 | 1386.7503 | 1227.6283 | 2021.6489 | 1195.9793 | 897.7315 | 790.7591 | 1090.1033 | 1088.6834 | 1039.0069 | 951.4887 | 632.9518 | 2758.53 | 1978.2822 | 1795.2566 | 2582.587 | 1433.6627 | 1142.1583 | 994.2203 | 1189.6214 | 470.3924 | 394.1514 | 388.0207 | 970.1859 | 1252.3392 | 1148.3373 | 999.6008 | 778.6696 | 1368.7247 | 1234.8348 | 874.8348 | 922.9656 | 1228.5369 | 1250.9243 | 943.7315 | 992.7522 | 697.0482 | 798.5113 | 525.8141 | 539.9312 | 1215.1652 | 1243.7591 | 1028.2891 | 1244.3511 | 1273.587 | 1264.3048 | 1195.9449 | 1518.0069 | 1897.9518 | 1281.3442 | 1166.1652 | 1848.7797 | 1683.4061 | 1213.4011 | 983.6214 | 1925.9105 | 804.2409 | 731.7384 | 607.5369 | 1120.7591 | 633.2822 | 564.1239 | 589.8692 | 511.1308 | 1154.0344 | 762.2841 | 681.5113 | 1282.8692 | 2445 | 1680.3029 | 1787.3185 | 3465.3874 | 3339.5507 | 2204.8486 | 2082.5526 | 4220.2703 | 460.8417 | 280.1377 | 279.0138 | 460.8417 | 510.8073 | 498.8967 | 368.179 | 604.8348 | 470.7315 | 440 | 401.2547 | 454.062 | 376.8967 | 341.8967 | 320.9174 | 435.0413 | 362.0688 | 324.9587 | 319.0757 | 355.0275 | 1645.8967 | 977.3855 | 923.0344 | 925.4768 | 1073.1033 | 874.7797 | 904.3529 | 664.179 | 2787.0551 | 1778.3717 | 1572.9775 | 1765.5801
328 | 5.6024 | 791.4602 | 1179.7257 | 1122.0575 | 1446.1401 | 1438.0192 | 1293.5015 | 900.8629 | 1110.8466 | 666.7507 | 589.6549 | 456.7699 | 630.7316 | 1491.9454 | 887.8274 | 1019.2138 | 1260.6711 | 1425.3318 | 907.3097 | 791.9808 | 1028.885 | 1302.0354 | 1574.3289 | 1469.5944 | 1284.2714 | 2012.7699 | 1493.1017 | 1173.885 | 1970.8658 | 1194.534 | 859.115 | 756.1917 | 1104.5398 | 1043.1917 | 1040.326 | 1020.351 | 625.8083 | 2833.9646 | 2018.7124 | 1688.7095 | 2524.1313 | 1386.1888 | 1164.9617 | 1024.907 | 1135.7478 | 527.115 | 415.7507 | 390.1755 | 995.9262 | 1158.2464 | 1196.5206 | 944.1092 | 731.7095 | 1328.5015 | 1210.6932 | 850.2905 | 917.1755 | 1306.4027 | 1241.5148 | 905.9204 | 956.4056 | 704.2301 | 728.4794 | 498.9233 | 530.3643 | 1239.5561 | 1209.1726 | 1069.7316 | 1292.6223 | 1215.3997 | 1163.9425 | 1188.0383 | 1517.6386 | 1890.8466 | 1329.5427 | 1188.4277 | 1817.5752 | 1743.0163 | 1127.8245 | 930.0738 | 1913.7862 | 838.3864 | 694.8437 | 683.5236 | 1086.2109 | 672.8112 | 582.7478 | 571.2876 | 531.0354 | 1159.5752 | 659.6136 | 612.0163 | 1263.2138 | 2444.1947 | 1723.9617 | 1690.4381 | 3374.143 | 3415.9321 | 2181.4852 | 2017.233 | 4115.5752 | 438.2109 | 299.3672 | 280.9233 | 438.2109 | 483.4602 | 484.0383 | 394.2684 | 581.1917 | 433.093 | 440.5561 | 437.9233 | 463.1534 | 362.3451 | 325.9454 | 309.0767 | 442.1121 | 372.2876 | 319.0767 | 330.3451 | 358.9041 | 1631.7478 | 1032.4631 | 927.0604 | 849.6357 | 1086.7345 | 843.8437 | 811.2109 | 689.1371 | 2796.0738 | 1829.7949 | 1712.7536 | 1706.5148
329 | 5.6134 | 814.4965 | 1267.6539 | 1122.7974 | 1353.2341 | 1439.3776 | 1267.4651 | 944.3707 | 1103.6608 | 654.4405 | 573.5418 | 445.944 | 617.1573 | 1437.7797 | 880.1328 | 977.5211 | 1292.3569 | 1341.353 | 871.7797 | 791.7866 | 1020.6401 | 1355.3776 | 1539.9056 | 1498.8043 | 1247.0629 | 1999.4896 | 1395.056 | 1167.9371 | 1964.0315 | 1269.9793 | 863.8987 | 766.3147 | 1080.2448 | 1054.1328 | 1054.6401 | 936 | 614.9931 | 2781.3461 | 2004.0629 | 1724.9056 | 2578.7905 | 1450.7759 | 1161.9302 | 970.5733 | 1172.2065 | 533.1573 | 403.0944 | 347.9685 | 939.9302 | 1220.2694 | 1170.584 | 1049.6715 | 766.4267 | 1303.6608 | 1198.1151 | 814.8496 | 875.944 | 1324.7974 | 1317.2065 | 951.8672 | 927.1013 | 717.0069 | 753.2203 | 496.0698 | 549.2517 | 1269.0069 | 1216.9931 | 1056.7237 | 1171.7237 | 1288.6048 | 1161.944 | 1188.3009 | 1448.472 | 1882.4965 | 1255.1573 | 1107.7797 | 1844.9862 | 1791.9793 | 1170 | 982.9862 | 1950.1612 | 806.8427 | 735.7345 | 609.3845 | 1095.3323 | 612.7552 | 617.8534 | 585.4336 | 583.1504 | 1188.9371 | 691.5664 | 663.9685 | 1222.4336 | 2406.5242 | 1721.8427 | 1763.9685 | 3229.1642 | 3262.713 | 2104.7237 | 1978 | 4141.5702 | 448.4965 | 268.416 | 276.8427 | 448.4965 | 508.0384 | 486.5035 | 408.9754 | 590.9371 | 488.1819 | 469 | 434.2517 | 472.353 | 379.4336 | 272.472 | 312.7483 | 497.8043 | 385.4582 | 323.5349 | 348.472 | 353.119 | 1668.3009 | 1003.8358 | 879.7237 | 831.2832 | 1022.5664 | 887.1573 | 824.0767 | 646.4789 | 2848.3569 | 1719.0138 | 1648.7483 | 1779.406
330 | 5.6244 | 796.2871 | 1256.825 | 1056.444 | 1420.5998 | 1448.5934 | 1252.1377 | 926.5315 | 1123.825 | 668.6126 | 617.2369 | 476.3127 | 619.4621 | 1430.9125 | 913.3372 | 931.0811 | 1213.2754 | 1386.1505 | 864.2999 | 817.2689 | 948.6126 | 1361.0491 | 1541.2882 | 1463.4877 | 1246.4621 | 1952.9125 | 1368.6254 | 1167.9691 | 1964.4749 | 1179.0566 | 831.619 | 774.2935 | 1057.7001 | 1091.1878 | 982.7876 | 939.3693 | 584.8751 | 2790.556 | 2005.0811 | 1721.825 | 2483.4568 | 1561.2679 | 1130.7503 | 1017.5998 | 1118.6628 | 538.0438 | 406.4376 | 346.0811 | 908.0064 | 1165.9498 | 1128.3127 | 979.0939 | 718.0438 | 1320.9808 | 1292.1622 | 840.3874 | 900.7118 | 1255.5998 | 1262.8751 | 917.5251 | 962.7065 | 746.7748 | 758.9061 | 529.3063 | 556.6062 | 1297.8122 | 1187.1377 | 1076.8997 | 1192.381 | 1338.4376 | 1190.5187 | 1135.5624 | 1462.9125 | 1868.9253 | 1261.9691 | 1102.2689 | 1783.3063 | 1698.0374 | 1169.825 | 923.2754 | 1833.2689 | 801.6937 | 663.8559 | 650.4685 | 1043.9125 | 635.7812 | 522.3127 | 568.2625 | 553.9808 | 1187.3938 | 710.9253 | 661.9936 | 1211.1697 | 2516.8815 | 1714.2433 | 1765.4504 | 3264.4312 | 3374.9434 | 2130.1067 | 1979.7065 | 4032.8933 | 432.5187 | 311.1686 | 272.5688 | 432.5187 | 538.9936 | 501.7375 | 438.825 | 588.6937 | 463.7001 | 468.5187 | 439.9872 | 512.9498 | 362.2188 | 288.4877 | 305.7439 | 460.2561 | 338.9626 | 339.5624 | 359.587 | 327.2689 | 1617.794 | 966.2935 | 901.825 | 840.2188 | 1040.3938 | 894.2316 | 885.9872 | 688.937 | 2768.9691 | 1777.3682 | 1640.0374 | 1640.1441
331 | 5.6354 | 738.9056 | 1250.1416 | 1089.4012 | 1389.5516 | 1398.6342 | 1274.1416 | 960.5605 | 1118.879 | 681.8319 | 539.5959 | 502.0295 | 566.9705 | 1430.8496 | 881.2036 | 913.0738 | 1260.4218 | 1451.8053 | 850.121 | 842.0855 | 963.0089 | 1233.0738 | 1629.8997 | 1494.2625 | 1190.4956 | 1947.8613 | 1426.2242 | 1207.1887 | 1952.3835 | 1223.1593 | 864.4778 | 740.0266 | 1069.8613 | 1134.7404 | 1000.5605 | 1012.7758 | 603.3835 | 2751.5456 | 1988.0826 | 1717.4956 | 2554.9646 | 1437.0089 | 1163.2831 | 987.7847 | 1154.7758 | 536.3097 | 414.4867 | 327.177 | 928.9911 | 1141.2065 | 1157 | 1004.8496 | 720.2331 | 1260.1564 | 1252.5044 | 828.7375 | 807.6165 | 1222.702 | 1280.2065 | 930.3451 | 998.0383 | 719.6077 | 736.2331 | 537.121 | 546.2625 | 1251.0089 | 1209.5339 | 1034.531 | 1158.4012 | 1345.3097 | 1182.6342 | 1128.0738 | 1460.4956 | 1910.944 | 1297.994 | 1133.0266 | 1790.9528 | 1675.3746 | 1166.9528 | 968.292 | 1860.7198 | 794.4395 | 623.5015 | 616.0472 | 1044.354 | 630.2153 | 551.2802 | 575.4012 | 489.6814 | 1196.3923 | 753.1681 | 644.0915 | 1339.3864 | 2553.7964 | 1656.9705 | 1818.3274 | 3250.6549 | 3328.5988 | 2193.1238 | 2016.8879 | 3964.0472 | 424.0177 | 292.8879 | 284.2714 | 424.0177 | 516.6637 | 496.4484 | 435.3363 | 580.9351 | 478.7198 | 459.5133 | 396.1681 | 491.0089 | 364.5899 | 321.1681 | 319.4218 | 444.8968 | 360.2242 | 328.2625 | 288.3629 | 352.9262 | 1655.2065 | 933.9322 | 900.5222 | 841.7492 | 1050.5693 | 937.9557 | 840.9351 | 626.5399 | 2810.5428 | 1693.7964 | 1619.056 | 1687.9528
332 | 5.6464 | 772.3668 | 1203.7768 | 1114.6401 | 1377.353 | 1443.41 | 1227.3668 | 964.5329 | 1098.4282 | 679.205 | 614.2664 | 452.8884 | 614.7448 | 1458.0363 | 945.1661 | 948.9567 | 1223.4239 | 1394.7881 | 864.763 | 792.4533 | 974.7379 | 1266.4965 | 1544.8884 | 1464.2483 | 1186.3737 | 1902.1072 | 1427.7448 | 1146.8953 | 1942.3417 | 1205.2414 | 820.2915 | 789.5649 | 1019.7336 | 1066.622 | 1008.5649 | 1004.763 | 594.0934 | 2635.9775 | 2037.1298 | 1713.1687 | 2485.7699 | 1448.2595 | 1074.5173 | 999.9749 | 1148.4031 | 495.1687 | 390.3668 | 345.2915 | 913.8451 | 1117.4239 | 1153.5147 | 1032.0363 | 738.1981 | 1347.2734 | 1260.4533 | 859.2984 | 818.3417 | 1192.8157 | 1249.4533 | 949.0363 | 964.3166 | 713.5467 | 754.0614 | 552.9248 | 517.9567 | 1267.795 | 1171.622 | 1091.6081 | 1183.1618 | 1296.5147 | 1223.173 | 1159.8564 | 1452.6834 | 1910.0683 | 1186.2872 | 1178.8062 | 1805.0182 | 1652.442 | 1182.795 | 889.0683 | 1854.9066 | 785.6834 | 729.7881 | 604.647 | 1081.6765 | 619.1185 | 558.6652 | 595.173 | 517.385 | 1206.5536 | 754.7016 | 710.4031 | 1226.4965 | 2483.2734 | 1705.1549 | 1803.5147 | 3227.59 | 3304.1116 | 2111.3417 | 2010.4213 | 3947.7266 | 454.378 | 290.1799 | 272.3417 | 454.378 | 508.2664 | 503.0433 | 437.7881 | 565.0502 | 474.7517 | 485.4784 | 400.8451 | 504.9248 | 325.6401 | 324 | 278.7336 | 456.378 | 366.7336 | 300.2301 | 342.8132 | 320.1185 | 1625.41 | 1015.8382 | 940.6081 | 787.9567 | 1070.2345 | 855.3235 | 824.7768 | 700.3348 | 2779.0294 | 1633.1825 | 1617.8132 | 1706.2552
333 | 5.6574 | 788.9036 | 1231.5477 | 1132.9685 | 1313.1937 | 1445.0491 | 1241.5654 | 893.2252 | 1074.6617 | 678.2104 | 575.7906 | 464.2743 | 583.4356 | 1406.9027 | 853.8545 | 957.7266 | 1310.7581 | 1351.5644 | 811.1928 | 787.2576 | 906.1446 | 1256.9657 | 1558.2094 | 1461.5792 | 1237.7757 | 1998.4356 | 1399.7414 | 1196.5329 | 1942.1622 | 1154.4838 | 801.2734 | 768.8712 | 1058.4513 | 1111.8879 | 990.2901 | 947.4838 | 606.5811 | 2771.8387 | 1995.7424 | 1767.2586 | 2458.1956 | 1377.2567 | 1176.1937 | 982.9361 | 1116.8702 | 547.5654 | 407.7099 | 321.5162 | 933.114 | 1203.7099 | 1109.2252 | 976.0649 | 698.7099 | 1346.0816 | 1245.7433 | 874.6617 | 823.3225 | 1309.1298 | 1233.1307 | 897.822 | 950.8072 | 718.9843 | 715.6932 | 512.3707 | 530.9194 | 1262.5005 | 1220.5969 | 1062.709 | 1196.9036 | 1256.9824 | 1174.6283 | 1121.7572 | 1461.5969 | 1904.7915 | 1267.4847 | 1114.2891 | 1775.0973 | 1714.1288 | 1184.1928 | 893.7896 | 1841.0482 | 795.7257 | 687.1288 | 663.646 | 1052.177 | 645.823 | 591.6135 | 543.8869 | 507.7099 | 1249.0334 | 733.2104 | 673.468 | 1209.0639 | 2488.3707 | 1680.8869 | 1754.7099 | 3209.0992 | 3298.531 | 2115.5162 | 1948.8869 | 3933.9213 | 405 | 279.2419 | 274.3392 | 405 | 473.0964 | 490.645 | 394.8063 | 590.5005 | 482.178 | 490.6293 | 424.2261 | 464.1288 | 347.7581 | 322.2261 | 310.5172 | 408.3707 | 401.178 | 320.0639 | 312.7581 | 347.5487 | 1633.3383 | 1001.4513 | 909.6126 | 803.822 | 983.8054 | 884.9843 | 848.7581 | 665.8545 | 2687.3058 | 1773.6932 | 1584.3717 | 1730.1465
334 | 5.6684 | 768.9027 | 1263.6371 | 1094.8983 | 1303.5354 | 1400.0929 | 1192.3009 | 930 | 1049.5797 | 637.8805 | 563.3009 | 458.6018 | 576.7212 | 1435.1681 | 876.3009 | 911.2301 | 1300.7522 | 1350.8451 | 863.4602 | 808.4469 | 943.5797 | 1382.6504 | 1576.739 | 1543.0752 | 1147.2522 | 1985.8672 | 1432.5797 | 1152.0752 | 1887 | 1165.0177 | 849.1637 | 754.0487 | 1073.9779 | 1071.9691 | 1007.323 | 952.0708 | 582.0664 | 2767.3982 | 1961.9159 | 1737.4867 | 2341.5045 | 1453.0309 | 1164.1195 | 946.5133 | 1161.4646 | 496.6283 | 392.5354 | 318.2301 | 873.9912 | 1186.3982 | 1141.9115 | 956.6018 | 688.3673 | 1295.854 | 1164.3938 | 846.3274 | 827.2788 | 1260.7256 | 1140.27 | 929.0708 | 902.4425 | 700.5531 | 731.5619 | 528.3717 | 529.1637 | 1238.646 | 1214.2345 | 1095.9867 | 1173.5575 | 1339.3584 | 1210.8893 | 1171.4469 | 1453.8407 | 1848.6815 | 1222.7832 | 1182.6548 | 1742.0929 | 1728.8584 | 1226.146 | 932.1681 | 1853.6504 | 804.9292 | 701.6726 | 617.4381 | 1070.1238 | 622.9513 | 552.4425 | 555.4867 | 496.0664 | 1186.2036 | 693.7168 | 661.2036 | 1241.3982 | 2507.6239 | 1688.8628 | 1741.6726 | 3062 | 3368.6327 | 2116.1328 | 1956.1195 | 3816.1195 | 404.6283 | 282.1858 | 245.7655 | 404.6283 | 496.9071 | 499.6726 | 403.9071 | 563.5797 | 453.7168 | 475.4646 | 403.0708 | 472.1681 | 345.3717 | 304.5089 | 252.9248 | 422.6991 | 368.0708 | 351.5619 | 311.3009 | 332.4867 | 1660.4646 | 1012.7035 | 926.792 | 838.6947 | 1044.3318 | 871.1063 | 843.1195 | 680.4026 | 2731.1328 | 1784.2654 | 1552.5709 | 1652.7434
335 | 5.6794 | 805.1062 | 1187.6814 | 1161.4788 | 1299 | 1407.3127 | 1209.8398 | 933.1564 | 1061.3176 | 666.8417 | 569.3716 | 484.6313 | 571.1593 | 1381.3716 | 879.7413 | 934.8938 | 1229.3687 | 1282.7886 | 904.5801 | 768.945 | 957.527 | 1376.2693 | 1620.3697 | 1496.1573 | 1173.7394 | 1937.5733 | 1445.2645 | 1112.7326 | 1887.8417 | 1217.8928 | 839.6322 | 739.8398 | 1066.3147 | 1134.7394 | 1028.6322 | 946.0511 | 610.001 | 2741.5792 | 1894.0473 | 1750.9479 | 2406.2625 | 1386.999 | 1136.6313 | 932.4208 | 1158.4247 | 489.7384 | 386.6322 | 341.473 | 948.3176 | 1164.9981 | 1102.9469 | 981.0531 | 728.8958 | 1347.3186 | 1180.7355 | 826.528 | 824.0029 | 1222.3668 | 1221.3156 | 918.7375 | 894.0541 | 741.4228 | 700.6303 | 530.6322 | 520.7896 | 1284 | 1191.2133 | 1016.2065 | 1174.055 | 1253.2114 | 1165.5241 | 1134.7857 | 1413.9518 | 1983.7935 | 1284.2655 | 1148.3166 | 1734.7935 | 1753.1023 | 1169.4208 | 888.472 | 1850.2153 | 818.4189 | 717.3687 | 652.473 | 1008.7867 | 645.2616 | 546.7896 | 570.3156 | 522.2124 | 1228.3726 | 726.2124 | 703.3726 | 1215.0531 | 2548.5762 | 1674.3127 | 1755.8958 | 3060.001 | 3333.8909 | 2174.1062 | 1930.3147 | 3774.9537 | 401 | 284 | 259.1072 | 401 | 499.6824 | 515.2635 | 398.8967 | 577.8427 | 488.3166 | 483.2616 | 393.5792 | 432.4701 | 346.3697 | 329.8417 | 301.7365 | 411.0521 | 360.7886 | 321.0511 | 323.2635 | 344.4749 | 1659.9498 | 965.7345 | 916.7886 | 863.001 | 994.6814 | 948.3215 | 820.5762 | 619.2616 | 2787.1072 | 1696.4749 | 1638.7404 | 1656.6332
336 | 5.6904 | 792.0551 | 1230.9474 | 1073.175 | 1292.7723 | 1456.5349 | 1240.0722 | 959.8825 | 1033.8776 | 668.5949 | 517.7001 | 489.295 | 552.2876 | 1338.0428 | 806.7577 | 939.7197 | 1214.8348 | 1293.1824 | 877.8898 | 806.4174 | 943.2375 | 1279.0453 | 1593.7699 | 1501.0649 | 1126.7674 | 2031.6524 | 1418.3574 | 1186.6475 | 1885.2472 | 1246.0649 | 827.295 | 774.6548 | 1074.0147 | 1096.6977 | 1009.3623 | 970.59 | 592.06 | 2741.8776 | 1993.0526 | 1736.4774 | 2405.3672 | 1399.5349 | 1137.6524 | 945.4002 | 1089.1824 | 467.2876 | 375 | 353.4125 | 910.6548 | 1204.4051 | 1116.0526 | 979.9327 | 712.2925 | 1302.7601 | 1206.4676 | 802.6401 | 780.5202 | 1253.7026 | 1224.94 | 906.4725 | 877.3427 | 712.88 | 722.8825 | 518.8825 | 520.88 | 1279.4174 | 1143.5275 | 1082.4749 | 1130.7075 | 1246.2252 | 1204.1175 | 1201.235 | 1349.8078 | 1905.5398 | 1252.1077 | 1135.5275 | 1670.6401 | 1802.8849 | 1169.2399 | 919.88 | 1772.5226 | 851.4174 | 717.1701 | 665.3525 | 1048.2472 | 666.6548 | 550.8752 | 574.41 | 500.3403 | 1158.4076 | 699.1701 | 637.2852 | 1207.5875 | 2592.2399 | 1730.7026 | 1739.41 | 3038.2999 | 3395.3599 | 2155.5324 | 1955.4051 | 3658.8299 | 401.1175 | 279.2999 | 231.2876 | 401.1175 | 523.0098 | 503.4774 | 369.7577 | 570.4676 | 473.1775 | 506.1824 | 389.6475 | 493.5324 | 325.2925 | 333.2399 | 304.0049 | 416.9449 | 377.0575 | 344.4149 | 316.8825 | 322.53 | 1619.8201 | 998.7148 | 928.2399 | 842.9449 | 1032.8947 | 853.8176 | 864.2399 | 645.355 | 2752.9425 | 1674.53 | 1575.4725 | 1621.4198
337 | 5.7014 | 804.7463 | 1200.9174 | 1089.9174 | 1251.8407 | 1418.587 | 1158 | 960.5575 | 1055.8613 | 624.7551 | 548.7168 | 482.1829 | 588.7551 | 1433.7935 | 854.9322 | 862.6579 | 1145.6106 | 1263.8791 | 827.4572 | 778.8791 | 934.1534 | 1353.6932 | 1559.4041 | 1463.8702 | 1111.5575 | 1990.4336 | 1389.1829 | 1176.7315 | 1815.764 | 1202.6401 | 823.1681 | 730.6342 | 989.1918 | 1144.4041 | 959.2685 | 964.8024 | 584.7404 | 2761.0059 | 2023.2124 | 1695.7788 | 2324.7493 | 1367.5575 | 1104.7404 | 1003.2802 | 1054.4277 | 504.2743 | 374.7404 | 347.7021 | 869.8791 | 1248.3746 | 1152.9616 | 1019.0206 | 724.9853 | 1326.1976 | 1222.5339 | 836.823 | 833.3658 | 1277.764 | 1232.7788 | 901.7493 | 938.0531 | 724.7935 | 717.3274 | 514.7551 | 532.5339 | 1244.8702 | 1152.0147 | 1060.0236 | 1121.3894 | 1312.0236 | 1193.5781 | 1192.8761 | 1438.4719 | 1850.3068 | 1304.5634 | 1149.3363 | 1705.3421 | 1790.5044 | 1153.0709 | 930.6254 | 1808.4189 | 819.4661 | 750.6784 | 659.3038 | 977.5959 | 616.4041 | 587.4719 | 584.6106 | 576.5723 | 1182.0531 | 732.0295 | 684.9911 | 1206.938 | 2568.2832 | 1740.4867 | 1754.1534 | 3004.0384 | 3348.9764 | 2133.6254 | 1993.9233 | 3622.0767 | 403.0384 | 251.3983 | 266.587 | 403.0384 | 462.2212 | 469.3983 | 409.236 | 585.236 | 461.6784 | 462.6017 | 388.5575 | 473.9616 | 332.2743 | 303.0531 | 281.5281 | 400.2301 | 383.7021 | 325 | 315.351 | 317.7021 | 1639.7021 | 951.5281 | 902.7257 | 810.9144 | 966.6048 | 881.9528 | 830.6784 | 626.3805 | 2759.6667 | 1669.5723 | 1567.2449 | 1572
338 | 5.7124 | 747.0826 | 1254.7104 | 1138.7374 | 1280.0594 | 1386.8687 | 1166.0989 | 960.2015 | 1008.1421 | 601.295 | 527.1043 | 471.9892 | 568.705 | 1415.4479 | 879.9011 | 910.732 | 1144.5629 | 1306.8741 | 796.0054 | 820.1691 | 955.6062 | 1295.3004 | 1539.7266 | 1465.4154 | 1130.5251 | 1940.1043 | 1386.946 | 1145.9027 | 1848.2463 | 1158.0108 | 831.8525 | 753.295 | 1024 | 1125.7212 | 1008.5683 | 992.3112 | 571.349 | 2691.3004 | 1970.4317 | 1693.3274 | 2359.0594 | 1391.9243 | 1102.8579 | 916.8633 | 1067.1259 | 483.7266 | 373.5683 | 341.1313 | 911.1691 | 1266.1691 | 1129.6602 | 936.8309 | 735.6224 | 1300.9622 | 1197.617 | 781.7158 | 804.7374 | 1296.4641 | 1242.6888 | 954.8147 | 922.8687 | 708.5791 | 689.5521 | 496.5521 | 509.6062 | 1249.2572 | 1141.8039 | 1074.2123 | 1133.3776 | 1335.59 | 1135.8973 | 1132.1043 | 1385.0162 | 1904.1529 | 1237.4857 | 1169.973 | 1647.4641 | 1809.5954 | 1227.322 | 906.295 | 1757.7861 | 833.1583 | 714.7158 | 603.2734 | 1003.383 | 600.2788 | 540.268 | 571.6278 | 560.8903 | 1173.6888 | 696.2626 | 642.5467 | 1179.8201 | 2588.3328 | 1674.5467 | 1781.4317 | 3007.59 | 3315.7483 | 2111 | 1956.698 | 3636.4263 | 405.3112 | 294.6062 | 244.2463 | 405.3112 | 461.1205 | 509.4803 | 386.9838 | 565.1151 | 431.6834 | 416.9784 | 400.2788 | 464.3004 | 317.5521 | 288.4371 | 328.732 | 443.2788 | 368.7483 | 320.4533 | 309.295 | 308.1475 | 1631.7104 | 1000.0108 | 916.4533 | 750.2842 | 1029.8633 | 827.4317 | 805.6564 | 586.973 | 2841.376 | 1655.027 | 1578.6278 | 1580.5251
339 | 5.7234 | 805.6637 | 1250.5093 | 1100.1868 | 1211.53 | 1371.9931 | 1217.3156 | 901.3294 | 997.5762 | 597.6382 | 562.1268 | 479.8318 | 577.4562 | 1390.6037 | 835.1682 | 893.7788 | 1149.6775 | 1293.4032 | 796.6382 | 789.1475 | 908.9724 | 1287.7994 | 1527.6981 | 1485.2488 | 1165.3274 | 1970.0324 | 1436.5024 | 1240.878 | 1874.8642 | 1151.8663 | 837.3088 | 740.9931 | 1022.9194 | 1121.5507 | 1016.0138 | 961.6844 | 510.1082 | 2666.2006 | 1975.6451 | 1651.3225 | 2314.4583 | 1471.06 | 1112.1888 | 915.5575 | 1082.9125 | 468.6106 | 375.1475 | 351.9931 | 878.4493 | 1228.3549 | 1046.4376 | 966.7719 | 677.3019 | 1341.9194 | 1147.9263 | 781.6981 | 772.6706 | 1272.4376 | 1265.7512 | 980.6961 | 906.1406 | 703.2419 | 712.3894 | 524.175 | 477.4838 | 1278.2212 | 1192.3549 | 1005.6451 | 1176.4149 | 1303.5762 | 1244.4425 | 1176.8525 | 1371.4769 | 1896.705 | 1186.6589 | 1193.9793 | 1613.0944 | 1778.2143 | 1178.9724 | 903.1013 | 1674.1406 | 830.5114 | 710.8387 | 618.235 | 1045.3549 | 611.3963 | 557.6844 | 518.8801 | 519.06 | 1206.3225 | 718.6037 | 659.5024 | 1214.5162 | 2558.5389 | 1719.5989 | 1771.0531 | 2981.5968 | 3285.2557 | 2105.2881 | 1833.4445 | 3600.0972 | 383.175 | 261.5575 | 280.5899 | 383.175 | 488.47 | 459.2557 | 406.6981 | 586.5831 | 466.9263 | 437.6037 | 401.3618 | 452.4769 | 336.2212 | 286.5644 | 302.6706 | 442.3549 | 339.0276 | 299.175 | 296.8387 | 304.8525 | 1639.5507 | 978.1199 | 885.6037 | 751.5369 | 1016.2074 | 831.5644 | 857.0324 | 612.2143 | 2753.3501 | 1634.8801 | 1511.9862 | 1624.5899
340 | 5.7344 | 771.6667 | 1209 | 1033.3333 | 1193.8333 | 1349.6666 | 1222 | 935.6667 | 938 | 624.1666 | 532.6667 | 458.3333 | 563 | 1377.3334 | 856.6667 | 895.8333 | 1124.5 | 1315.5 | 836 | 774.6667 | 871.3334 | 1336.6667 | 1543.6667 | 1501.6666 | 1073.3333 | 1909.6667 | 1383 | 1159.3333 | 1794 | 1171.8333 | 841.3333 | 730.5 | 1018.3333 | 1132.8333 | 1025.3333 | 949.6667 | 563 | 2636.3334 | 2015.5 | 1679.5 | 2411.9999 | 1438.5 | 1162.5 | 935 | 1068.8333 | 464.1667 | 366.1667 | 335.5 | 858.3333 | 1183.3333 | 1131.8333 | 963.6666 | 662.3333 | 1338.8333 | 1156.1667 | 786.6667 | 749.6667 | 1348.3333 | 1233.1667 | 877.3334 | 888.6667 | 725.6667 | 722.3333 | 547.8333 | 510.3333 | 1259.1667 | 1134.3333 | 1047.6666 | 1095.1666 | 1267.6668 | 1212.6667 | 1182.8333 | 1381.6666 | 1918.6667 | 1261.6667 | 1161.1667 | 1644.6667 | 1754.8333 | 1138.8333 | 938.5 | 1655.8333 | 892.6666 | 701.3333 | 621.1667 | 995.1667 | 630.6667 | 547.6667 | 563.5 | 494.6667 | 1225.6667 | 702.5 | 608.8334 | 1169 | 2674 | 1806.6666 | 1736.8333 | 2949.6667 | 3329.5 | 2066.1666 | 1928.6667 | 3399.1667 | 411 | 284.5 | 240.5 | 411 | 496.1666 | 504.1667 | 414.6667 | 535.5 | 470.6667 | 426.5 | 367.5 | 470.3333 | 323.5 | 330.1667 | 271.5 | 397 | 379.8333 | 330.3333 | 284.6667 | 309.5 | 1649.6666 | 927.3333 | 867.3333 | 759 | 980.8333 | 863.6666 | 789 | 584.3333 | 2803.4999 | 1677.8333 | 1483.3333 | 1586
341 | 5.7454 | 818.7256 | 1240.8004 | 1054.295 | 1179.9794 | 1301.411 | 1178.2419 | 932.0315 | 963.3796 | 584.4848 | 592.7679 | 463.5683 | 530.6735 | 1398.4523 | 841.231 | 847.8958 | 1120.9371 | 1278.8948 | 846.3471 | 760.1583 | 904.4317 | 1338.5683 | 1565.7256 | 1443.7581 | 1067.1367 | 1962.6421 | 1411.59 | 1117.0846 | 1831.9773 | 1123.5271 | 849.0629 | 736.5792 | 1015.8731 | 1081.4415 | 962.8948 | 954.7896 | 572.4523 | 2722.2733 | 1976.3894 | 1700.5683 | 2289.4512 | 1403.8417 | 1123.3677 | 888.9479 | 1078.3471 | 474.0944 | 374.6844 | 315.5163 | 854.821 | 1198.5792 | 1138.8633 | 929.3265 | 645.4533 | 1348.3471 | 1168.2419 | 775.0217 | 804.4208 | 1303.9262 | 1205.0521 | 940.4208 | 904.4946 | 745.7787 | 728.6312 | 490.9154 | 492.3471 | 1241.0738 | 1083.8633 | 1016.1681 | 1067.9794 | 1394.9154 | 1201.1681 | 1165.0423 | 1320.0217 | 1916.2625 | 1292.1475 | 1163.2527 | 1663.179 | 1767.2104 | 1169.0954 | 918.4946 | 1673.0423 | 812.4425 | 711.3894 | 608.9794 | 982.0738 | 627.5575 | 559.9262 | 529.2429 | 508.3579 | 1201.6735 | 711.4208 | 644.6106 | 1126.1583 | 2640.4425 | 1759.821 | 1769.4219 | 2938.5369 | 3350.7473 | 2042.5694 | 1909.4621 | 3466.4425 | 385.474 | 252.9056 | 246.4002 | 385.474 | 473.0846 | 519.2625 | 407.821 | 539.4317 | 467.5792 | 465.2527 | 383.2213 | 440.0423 | 345.8427 | 330.7994 | 285.7787 | 403.7787 | 356.3048 | 318.7896 | 287.474 | 286.116 | 1604.2527 | 944.7158 | 868.821 | 799.3362 | 1000.2006 | 817.8948 | 795.7679 | 580.8319 | 2709.7365 | 1620.9262 | 1453.4002 | 1591.705
342 | 5.7564 | 772.355 | 1250.1976 | 1075.7212 | 1201.8638 | 1353.9612 | 1170.7537 | 935.24 | 999.7163 | 628.765 | 546.9675 | 451.4538 | 509.4813 | 1372.6175 | 819.1687 | 917.3775 | 1142.2512 | 1289.3112 | 830.1637 | 782.29 | 900.8638 | 1327.6013 | 1519.9287 | 1454.6888 | 1050.175 | 1945.1263 | 1458.8412 | 1155.71 | 1715.65 | 1168.5688 | 826.41 | 751.47 | 970.94 | 1051.1687 | 971.2075 | 960.705 | 549.295 | 2688.7325 | 1984.8737 | 1693.6175 | 2192.3112 | 1388.4037 | 1079.7537 | 936.5025 | 1053.94 | 440.88 | 384.705 | 343.115 | 842.9075 | 1214.235 | 1148.175 | 944.0162 | 696.8525 | 1329.295 | 1138.9513 | 828.0762 | 779.9675 | 1301.8687 | 1169.355 | 927.825 | 885.1588 | 730.6175 | 695.9675 | 469.5025 | 470.0438 | 1261.1087 | 1099.6775 | 994.5737 | 1101.1475 | 1364.3225 | 1176.1313 | 1189.9012 | 1344.5801 | 1886.1249 | 1285.7438 | 1162.3063 | 1646.7876 | 1758.8088 | 1212.35 | 897.055 | 1689.1037 | 837.41 | 705.53 | 640.4263 | 976.4813 | 612.3437 | 554.8088 | 582.35 | 499.6725 | 1206.8737 | 687.5413 | 656.3987 | 1143.7 | 2667.1313 | 1757.47 | 1806.3288 | 2923.7876 | 3351.12 | 2106.1588 | 1837.0325 | 3487.3938 | 405.235 | 270.82 | 254.35 | 405.235 | 494.115 | 475.5462 | 404.5138 | 552.5025 | 477.88 | 470.4263 | 388.055 | 452.8525 | 376.8475 | 289.9287 | 274.2512 | 392.825 | 328.65 | 323.7488 | 306.6613 | 301.3663 | 1613.2512 | 947.0438 | 864.7488 | 751.3663 | 1033.0388 | 822.9562 | 768.9012 | 602.4263 | 2695.3437 | 1614.8525 | 1474.5462 | 1564.35
343 | 5.7674 | 796.6283 | 1180.7965 | 1082.1858 | 1177.3186 | 1389.1858 | 1204.6725 | 976.4159 | 956.885 | 625.6902 | 520.531 | 437.7699 | 471.5133 | 1363.0619 | 880.938 | 866.8496 | 1143.177 | 1305.2124 | 849.9646 | 738.0177 | 864.3098 | 1287.6283 | 1552.3275 | 1435.531 | 1026.6725 | 1879.0177 | 1469.6106 | 1191.6371 | 1722.3539 | 1136.5576 | 811.0708 | 738.2832 | 963.0885 | 1115.9911 | 989.6106 | 951 | 559.4071 | 2679.1682 | 2042.8407 | 1691.5929 | 2211.8761 | 1447.7522 | 1123.8496 | 961.7611 | 1036.115 | 413.9823 | 383.5487 | 314.6106 | 803.9381 | 1204.8938 | 1125.8938 | 959.1681 | 697.6991 | 1334.5221 | 1090.4425 | 848.6371 | 740.5044 | 1316.8496 | 1193.0177 | 942.5752 | 852.6106 | 729 | 675.0266 | 493.1327 | 467.8319 | 1311.3628 | 1103.469 | 1002.0885 | 1102.1327 | 1349.9469 | 1156.8761 | 1229.4867 | 1266.6195 | 1959.2655 | 1223.5133 | 1128.9115 | 1566.6725 | 1751.6902 | 1185.6991 | 864.7699 | 1677.4424 | 810.2655 | 707.7257 | 638.8142 | 938.3098 | 660.3805 | 554.8142 | 549.3894 | 453.354 | 1271.354 | 668.823 | 670.9292 | 1099.354 | 2654.7965 | 1749.9823 | 1700.3717 | 2845.5044 | 3366.2124 | 2070.5575 | 1871.2035 | 3428.407 | 400.5752 | 233.0177 | 230.9558 | 400.5752 | 462.354 | 486.7876 | 393.7965 | 572.2655 | 453.4248 | 467.5929 | 359.2301 | 443.3186 | 323.7168 | 313.5752 | 282.7079 | 415.1062 | 359.3717 | 306.5398 | 299.3717 | 292.9204 | 1633.7168 | 935.8761 | 849.5752 | 745.1593 | 984.9912 | 814.584 | 804.4159 | 591.8584 | 2770.6549 | 1613.2566 | 1491.4867 | 1540.1416
344 | 5.7784 | 802.1367 | 1258.1293 | 1080.295 | 1231.9493 | 1318.0148 | 1166.5034 | 929.82 | 970.3884 | 602.5108 | 550.6475 | 467.7984 | 507.1441 | 1332.4317 | 827.18 | 879.5108 | 1098.3958 | 1258.4675 | 843.295 | 750.8491 | 896.5541 | 1208.9641 | 1571.5034 | 1458.9567 | 1008.7916 | 1900.2665 | 1378.3309 | 1151.9784 | 1654.1589 | 1202.7551 | 833.0433 | 719.7699 | 986.885 | 1085.4032 | 987.0359 | 951.8633 | 553.4459 | 2741.9277 | 1977.8627 | 1688.0575 | 2165.0068 | 1410.3457 | 1118.8126 | 957.9641 | 1021.5325 | 420.5757 | 415.4892 | 322.8275 | 826.7984 | 1186.18 | 1114.7341 | 947.6834 | 708.9926 | 1326.2808 | 1093.7409 | 814.3741 | 716.8417 | 1322.1725 | 1193.115 | 931.6042 | 862.1224 | 717.9926 | 714.9926 | 481.6691 | 473.4892 | 1264.0433 | 1083.4966 | 1018.115 | 1069.086 | 1339.5108 | 1127.9208 | 1194.7192 | 1323.3383 | 1932.0216 | 1261.9499 | 1111.59 | 1538.6475 | 1735.8491 | 1204.7625 | 904.5108 | 1605.2016 | 782.59 | 693.3741 | 632.9351 | 964.9425 | 607.115 | 554.9784 | 546.3525 | 442.7842 | 1259.9208 | 711.4459 | 648.7409 | 1174.4675 | 2655.7267 | 1742.5467 | 1707.3816 | 2800.0427 | 3368.6771 | 2082.5831 | 1875.0216 | 3340.2158 | 391.115 | 247.7916 | 253.2659 | 391.115 | 457.18 | 451.5182 | 396.2375 | 549.9858 | 465.0433 | 456.7192 | 406.2517 | 426.0792 | 340.5034 | 303.6834 | 267.8275 | 430.3667 | 356.4533 | 304.4391 | 294.9425 | 320.4317 | 1662.1583 | 906.2733 | 845.7267 | 772.3667 | 1020.8776 | 785.4248 | 763.8633 | 565.8275 | 2820.4459 | 1611.5393 | 1475.0649 | 1565.9567
345 | 5.7894 | 796.2625 | 1222.9843 | 1068.8033 | 1164.2969 | 1380.2064 | 1125.1563 | 911.9125 | 958.6687 | 597.5064 | 547.2969 | 446.8968 | 533.9626 | 1338.1062 | 833.6814 | 885.9125 | 1102.2595 | 1249.7562 | 850.2812 | 777.0688 | 912.5939 | 1238.6343 | 1536.3156 | 1430.6126 | 1050.7375 | 1944.294 | 1377.0905 | 1147.3874 | 1751.1533 | 1150.0875 | 845.1907 | 749.825 | 987.997 | 1135.1907 | 967.175 | 948.8407 | 525.0875 | 2693.4091 | 1897.7778 | 1648.3785 | 2130.4465 | 1449.8534 | 1067.0718 | 960.4877 | 1027.9626 | 462.4218 | 401.7876 | 331.2251 | 794.2625 | 1211.8437 | 1140.0127 | 937.9312 | 660.7876 | 1350.8751 | 1073.7375 | 817.5251 | 714.2468 | 1291.647 | 1183.1062 | 912.6844 | 889.1406 | 684.8437 | 682.528 | 487.9282 | 477.353 | 1289.65 | 1131.2311 | 1002.175 | 1008.5625 | 1345 | 1127.8781 | 1178.5654 | 1327.4159 | 1936.8407 | 1304.2094 | 1119.4031 | 1548.2999 | 1769.3687 | 1187.5251 | 901.7876 | 1644.1249 | 807.7405 | 696.2969 | 603.4218 | 966.2969 | 589.8063 | 539.4375 | 545.1406 | 447.4749 | 1247.3313 | 684.6844 | 626 | 1149.3313 | 2703.9745 | 1725.2094 | 1754.4031 | 2729.6529 | 3499.5191 | 2141.1533 | 1878.2438 | 3361.9873 | 382.2468 | 277.8751 | 220.4592 | 382.2468 | 476.4562 | 487.647 | 410.7375 | 574.4375 | 480.8407 | 424.3687 | 384.6657 | 427.2625 | 306 | 284.5782 | 278.7345 | 407.8968 | 356.9095 | 333.2251 | 307.9312 | 326.1062 | 1668.6313 | 903.8781 | 866.828 | 730.0501 | 992.0344 | 854.8594 | 775.2124 | 588.5969 | 2782.9626 | 1550.7876 | 1507.5438 | 1516.6253
346 | 5.8004 | 792.5295 | 1231.2477 | 1033.2331 | 1137.7876 | 1321.6829 | 1116.1844 | 907.3068 | 960.8333 | 575.7036 | 518.143 | 465.941 | 493.028 | 1348.1534 | 811.2331 | 870.4882 | 1064.6269 | 1282.0103 | 860.6445 | 750.5088 | 869.4321 | 1268.143 | 1531.3555 | 1415.4912 | 1045.3348 | 1864.6519 | 1403.0944 | 1170.6829 | 1684 | 1143.8746 | 805.2684 | 740.6135 | 952.6549 | 1109.4528 | 980.7139 | 906.6032 | 524.6445 | 2749.3141 | 1983.5575 | 1575.7449 | 2213.8229 | 1377.9543 | 1082.8259 | 917.857 | 986.3068 | 462.7006 | 375.6829 | 305.3098 | 793.4145 | 1234.7596 | 1085.5885 | 956.5575 | 644.5399 | 1321.2021 | 1083.2404 | 801.1637 | 732.7212 | 1277.5502 | 1197.9617 | 912.3555 | 890.115 | 720.9823 | 697.1814 | 466.4041 | 490.2477 | 1282.7669 | 1170.9513 | 1011.8673 | 1029.2891 | 1353.6549 | 1133.7876 | 1227.1814 | 1267.1077 | 1928.5015 | 1284.6342 | 1111.3555 | 1574.6829 | 1780.972 | 1176.4528 | 875.6829 | 1676.6829 | 854.2684 | 677.4808 | 617.1637 | 946.2788 | 587.8953 | 537.7036 | 544.4321 | 471.7699 | 1231.944 | 704.0694 | 622.6342 | 1104.1327 | 2834.4601 | 1697.4601 | 1736.2581 | 2758.3524 | 3384.5855 | 2095.8776 | 1841.4218 | 3371.9349 | 403.1254 | 249.885 | 253.0177 | 403.1254 | 487.857 | 461.2891 | 409.9026 | 548.1814 | 463.1254 | 429.4808 | 381.059 | 414.8746 | 315.1357 | 275.885 | 240.8746 | 432.4705 | 319.3658 | 292.6445 | 321.5088 | 333.7492 | 1670.7316 | 914.3555 | 898.969 | 695.1637 | 999.0694 | 839.7006 | 788.8156 | 589.913 | 2745.3938 | 1540.0694 | 1485.3244 | 1435.059
347 | 5.8114 | 809.4946 | 1182.6136 | 1087.4297 | 1119.8014 | 1326.4946 | 1147.1012 | 883.7581 | 885.3609 | 604.0904 | 495.592 | 447.9312 | 481.6136 | 1345.0974 | 870.7365 | 828.4258 | 1105.9135 | 1286.3431 | 856.3648 | 746.6785 | 857.7798 | 1251.6569 | 1533.6136 | 1424.3756 | 1001.7798 | 1846.8624 | 1350.9459 | 1171.1986 | 1690.0649 | 1150.4258 | 783.4946 | 692.7581 | 976.2094 | 1133.2527 | 1014.9498 | 866.0394 | 521.881 | 2681.1731 | 1990.3717 | 1670.2349 | 2121.6352 | 1417.2458 | 1071.1515 | 935.3539 | 976.6569 | 435.0865 | 369.3756 | 348.1229 | 810.6499 | 1236.527 | 1119.2999 | 952.2527 | 660.5309 | 1335.3609 | 1099.6677 | 809.2635 | 733.2311 | 1318.3431 | 1195.6932 | 899.9567 | 888.3284 | 748.8594 | 659.4621 | 476.527 | 434.2852 | 1219.7257 | 1133.0147 | 1056.4405 | 1051.5595 | 1376.1986 | 1112.5162 | 1174.2999 | 1303.3431 | 1970.7798 | 1274.2527 | 1085.3107 | 1572.1662 | 1795.4189 | 1196.7148 | 899.7001 | 1671.6391 | 825.1662 | 680.5162 | 624 | 942.7365 | 613.2202 | 573.6716 | 528.9892 | 477.1121 | 1279.7906 | 733.9931 | 625.3933 | 1052.9892 | 2808.1082 | 1678.4366 | 1714.4474 | 2697.7867 | 3350.4189 | 2138.2133 | 1804.9675 | 3193.1515 | 405.9243 | 255.2094 | 211.4946 | 405.9243 | 507.1229 | 473.1986 | 379.0108 | 511.7257 | 464.1553 | 440.3501 | 403.1121 | 411.5703 | 344 | 289.5487 | 235.5487 | 401.6067 | 339.0974 | 301.5054 | 301.5054 | 300.8594 | 1672.7944 | 909.2852 | 844.7798 | 699.9784 | 1054.5162 | 820.1731 | 750.3176 | 570.4297 | 2741.3284 | 1589.1986 | 1425.3068 | 1471.2635
348 | 5.8224 | 794.7502 | 1207.1101 | 1069.325 | 1135.855 | 1339.3098 | 1120.8398 | 886.795 | 904.3402 | 570.5998 | 517.8201 | 482.6302 | 509.8201 | 1391.3447 | 866.47 | 861.1003 | 1080.675 | 1263.1397 | 840.7296 | 793.35 | 863.615 | 1283.8702 | 1570.5649 | 1389.205 | 1012.915 | 1940.9051 | 1340.5049 | 1150.3751 | 1680.1755 | 1192.1101 | 784.915 | 709.3098 | 953.0752 | 1121.2802 | 955.8496 | 942.8201 | 497.4799 | 2711.5103 | 1980.0107 | 1607.8648 | 2133.0304 | 1360.5148 | 1105.8451 | 898.1799 | 1017.88 | 459.795 | 350.1048 | 311.94 | 770.0152 | 1220.6804 | 1089.7046 | 922.2605 | 602.385 | 1366.53 | 1133.0152 | 812.47 | 716.8702 | 1279.2399 | 1186.3599 | 906.0197 | 894.8952 | 715.8299 | 653.2399 | 492.4799 | 450.8899 | 1210.085 | 1092.6446 | 1049.0099 | 1055.8102 | 1357.7601 | 1114.615 | 1121.12 | 1244.1003 | 1976.88 | 1280.2999 | 1069.3044 | 1550.59 | 1777.7699 | 1193.0949 | 935.3903 | 1598.6007 | 808.8899 | 688.9749 | 621.615 | 924.4252 | 598.6401 | 557.8299 | 538.7198 | 446.9499 | 1280.9302 | 668.2999 | 671.8899 | 1059.2748 | 2848.5998 | 1749.8747 | 1744.8854 | 2731.4207 | 3372.2543 | 2065.0949 | 1809.5148 | 3201.3455 | 378.9051 | 242.41 | 218.2149 | 378.9051 | 473.06 | 459.795 | 379.35 | 514.2748 | 438.795 | 474.8899 | 362.615 | 443.4548 | 346.915 | 304.06 | 254.0349 | 363.325 | 351.6902 | 302.47 | 294.2551 | 270.4799 | 1634.8147 | 933.8398 | 867.0448 | 700.0949 | 1036.7404 | 856.3402 | 763.1701 | 562.9248 | 2774.3948 | 1575.53 | 1455.3098 | 1463.8702
349 | 5.8334 | 781.5694 | 1199.3688 | 1075.2183 | 1138.5044 | 1364.5811 | 1150.9085 | 899.2596 | 887.8053 | 595.8407 | 516.2095 | 438.059 | 498.5044 | 1436.9587 | 859.7316 | 824.2095 | 1072.5634 | 1315.3953 | 872.7167 | 773.7994 | 863.9322 | 1271.3009 | 1571.3038 | 1398.3688 | 1029.8732 | 1928.1593 | 1369.0826 | 1130.3363 | 1596.8997 | 1173 | 804.3688 | 740.6814 | 916.118 | 1078.7905 | 1012.8495 | 933.4454 | 516.2861 | 2683.2419 | 1871.3098 | 1646.0413 | 2064.0502 | 1398.5546 | 1094.5959 | 900.3451 | 1008.3953 | 454.236 | 383.9587 | 315.1003 | 752.0236 | 1185.3924 | 1137.9852 | 878.3157 | 611.2183 | 1347.4808 | 1095.8909 | 813.2183 | 692.1003 | 1315.1239 | 1205.6637 | 945.4631 | 859.6549 | 726.177 | 683.3009 | 511.5634 | 473.6047 | 1226.9558 | 1140.7581 | 1055.4218 | 1010.354 | 1380.3836 | 1119.3688 | 1123.6726 | 1203.6047 | 1953.2536 | 1282.5221 | 1146.4306 | 1522.2802 | 1799.8732 | 1191.4218 | 878.5044 | 1518.9883 | 846.5782 | 695.0088 | 614.1681 | 912.8378 | 587.6549 | 569.2861 | 567.0678 | 478.1505 | 1271.5782 | 690.4867 | 676.0265 | 1086.3864 | 2873.2861 | 1781.4306 | 1673.5634 | 2597.236 | 3473.6489 | 2072.5722 | 1848.6637 | 3133.2331 | 367.823 | 232.4543 | 241.5546 | 367.823 | 482.1003 | 465.0502 | 373.1091 | 529.1858 | 434.3451 | 478.4719 | 360.7139 | 463.0265 | 339.7492 | 300.0678 | 265.5044 | 373.6549 | 327.5281 | 297.6726 | 281.3688 | 296.2183 | 1680.6224 | 967.5133 | 889.6401 | 730.5871 | 981.3776 | 829.2684 | 746.1681 | 586.295 | 2825.9646 | 1598.9057 | 1480.3038 | 1449.6372
350 | 5.8444 | 840.3938 | 1206.7025 | 1087.6603 | 1120.1273 | 1314.8377 | 1134.1623 | 922.8146 | 918.8648 | 589.0541 | 528.9189 | 439.0811 | 498.9499 | 1431.9268 | 843.973 | 835.1814 | 1062.7375 | 1281.0963 | 800.1583 | 738.2355 | 867.8107 | 1314.0772 | 1529.5251 | 1436.6873 | 1038.8107 | 1941.3166 | 1393.6253 | 1125.7104 | 1581 | 1170.6834 | 845.8727 | 719.5792 | 908.3166 | 1055.6563 | 988.1313 | 925.6293 | 498.498 | 2719.2045 | 1917.8727 | 1655.7065 | 2084.1623 | 1408.9769 | 1102.2315 | 898.6372 | 1011.4518 | 414.3087 | 371.1082 | 319.1273 | 795.2355 | 1266.9228 | 1063.5211 | 979.9189 | 634.3747 | 1305.5831 | 1099.9499 | 814.6563 | 719.004 | 1360.6253 | 1209.7104 | 927.2665 | 848.911 | 705.498 | 705.8067 | 493.3397 | 479.8146 | 1285.6293 | 1087.7104 | 1005.5521 | 1014.4598 | 1427.4321 | 1141.471 | 1140.7724 | 1217.6332 | 1924.3747 | 1256.9499 | 1115.1161 | 1450.3206 | 1839.7955 | 1148.502 | 880.1082 | 1611.9189 | 872.4129 | 675.6293 | 613.4479 | 960.2005 | 609.3437 | 545.7065 | 545.2355 | 478.2546 | 1322.027 | 701.3127 | 628.3127 | 1103.8997 | 2847.1003 | 1725.7916 | 1675.3206 | 2592.6372 | 3394.5442 | 2052.3437 | 1861.1082 | 3176.7025 | 383 | 222.6332 | 236.6293 | 383 | 494.2355 | 476.4749 | 377.7375 | 507.3707 | 436.4017 | 436.5521 | 375.6062 | 427.7645 | 304.3437 | 272.1544 | 272.031 | 385.498 | 352.0732 | 299.1352 | 307.8146 | 306.6332 | 1656.471 | 950.8958 | 866.4017 | 740.5172 | 1000.027 | 858.3476 | 752.3978 | 562.4479 | 2856.8608 | 1626.1735 | 1442 | 1503.7144
351 | 5.8554 | 827.3589 | 1172.5271 | 1096.1514 | 1112.1308 | 1281.6794 | 1160.0757 | 907.1514 | 872.8692 | 595.0757 | 501.9252 | 435.8879 | 511.7738 | 1441.4169 | 848.3579 | 815.6981 | 1080.886 | 1227.8486 | 796.3589 | 730.2075 | 835.1131 | 1292.9439 | 1516.472 | 1444.114 | 1021.8103 | 1938.8495 | 1354.6224 | 1135.8673 | 1601.2252 | 1180.6972 | 869.5664 | 726.4336 | 910.9813 | 1052.5467 | 976.6421 | 911.5093 | 494.0561 | 2695.8103 | 1939.7551 | 1600.5299 | 2089.2468 | 1436.4336 | 1071.7355 | 932.1701 | 1024.1897 | 377.7729 | 373.7935 | 299.9617 | 787.5093 | 1256.3598 | 1038.5654 | 969.2262 | 685.7542 | 1352.7159 | 1116.3963 | 790.4159 | 737.2458 | 1309.8495 | 1203.8682 | 938.529 | 792.0748 | 679.9252 | 661.8103 | 475.4907 | 468.6794 | 1263.057 | 1099.0748 | 977.2468 | 1080.2841 | 1339.6028 | 1117.8682 | 1204 | 1232.3206 | 1991.7532 | 1252.8692 | 1164.2468 | 1466.0953 | 1878.585 | 1162.3019 | 886.4159 | 1571.643 | 814.8869 | 668.1505 | 623.943 | 880.8309 | 621.8112 | 520.1131 | 545.3579 | 435.6972 | 1294.7 | 698.414 | 617.5654 | 1121.3972 | 2808.1701 | 1724.9066 | 1697.4346 | 2612.586 | 3346.0561 | 2056.3589 | 1860.4729 | 3101.171 | 381.1888 | 241.8486 | 222.8112 | 381.1888 | 483.7925 | 483.4907 | 371.4346 | 513.9439 | 474.0383 | 426.6608 | 383.8112 | 437.9056 | 318.6224 | 254.6411 | 285.2841 | 377.1318 | 325.0748 | 311.6234 | 289.7365 | 313.7738 | 1637.3963 | 934.4542 | 894.6804 | 692.4897 | 1000.7738 | 883.0196 | 765.7551 | 567.2075 | 2802.2832 | 1554.2075 | 1468.5645 | 1515.8122
352 | 5.8664 | 809.1726 | 1213.4026 | 1077.3717 | 1128.0132 | 1294.1062 | 1172.677 | 870.1504 | 850.7036 | 600.4514 | 490.6505 | 405.3098 | 521.4779 | 1368.6018 | 847.3805 | 831.3363 | 1096.6637 | 1264.8496 | 798.593 | 743.6814 | 834.3451 | 1287.708 | 1504.4292 | 1404.4602 | 1046.3495 | 1919.6858 | 1390.9602 | 1144.3495 | 1618.1548 | 1187.1548 | 831.8672 | 723.6062 | 899.9159 | 1088.5354 | 962.5398 | 914.3717 | 504.3186 | 2737.9425 | 1915.7743 | 1582.8673 | 2023.5177 | 1416.9557 | 1085.3142 | 898.0663 | 986.7699 | 402.1726 | 345.6239 | 314.2035 | 770.2654 | 1225.9912 | 1082.1638 | 952.9247 | 682.1238 | 1390.031 | 1109.6593 | 771.8806 | 710.4026 | 1310.4779 | 1168.7522 | 909.4867 | 821.885 | 671.1638 | 688.5486 | 474.8274 | 457.8584 | 1253.5974 | 1138.9381 | 940.7346 | 1046.1106 | 1378 | 1104.4292 | 1192.3761 | 1231.8318 | 2030.4646 | 1214.3673 | 1084.6946 | 1421.031 | 1837.1858 | 1155.7743 | 872.9071 | 1511.4292 | 831.9912 | 714.0178 | 635.9159 | 862.5398 | 621.2876 | 539.4514 | 565.115 | 463.5177 | 1216.2301 | 715.2389 | 621.5796 | 1078.9159 | 2836.5842 | 1717.9602 | 1685.5664 | 2556.177 | 3377.0531 | 2037.5442 | 1787.2389 | 3017.4999 | 384.854 | 263.1416 | 221.031 | 384.854 | 472.8318 | 470.261 | 346.6858 | 501.7965 | 456.969 | 418.5974 | 382.031 | 421.4646 | 316.0044 | 279.739 | 242.9425 | 387.4292 | 344.2035 | 291.1682 | 268.1991 | 312.4823 | 1643.5398 | 906.7036 | 857.2257 | 734.1682 | 1001.0531 | 853.6283 | 750.5708 | 559.0044 | 2813.2788 | 1569.2522 | 1494.4999 | 1474.8318
353 | 5.8774 | 812.1406 | 1199.6529 | 1078.9381 | 1068.4051 | 1345.1819 | 1165.8091 | 844.9961 | 914.5497 | 634.9381 | 515.2478 | 425.4297 | 540.4955 | 1367.4877 | 813.0167 | 869.2852 | 1061.9794 | 1275.9665 | 843.8555 | 751.3304 | 824.2852 | 1305.6775 | 1499.3677 | 1385.3884 | 1025.3432 | 1934.3432 | 1424.2517 | 1132.1574 | 1567.6903 | 1133.4258 | 792.0167 | 699.1406 | 907.649 | 1117.5742 | 976.7148 | 904.8387 | 484.2645 | 2738.3677 | 1885.6362 | 1669.6155 | 1989.3304 | 1378.3393 | 1100.6903 | 856.1613 | 988.471 | 406.7729 | 368.0619 | 294.0206 | 739.0374 | 1262.9007 | 1102.7807 | 913.9548 | 622.8181 | 1360.0245 | 1087.5536 | 799.0206 | 709.058 | 1338.6362 | 1123.5497 | 918.0374 | 836.4464 | 716.9174 | 675.2813 | 489.8142 | 458.8348 | 1255.9213 | 1136.2723 | 985.6735 | 1037.5988 | 1372.4503 | 1118.6284 | 1165.3677 | 1207.1445 | 1979.88 | 1237.3265 | 1040.8968 | 1404.7355 | 1780.5458 | 1154.6903 | 891.76 | 1497.5536 | 864.6568 | 735.1317 | 613.6155 | 891.0787 | 610.0206 | 566.7561 | 555.1239 | 453.1613 | 1221.3845 | 668.7935 | 608.3393 | 1065.7561 | 2899.1691 | 1775.409 | 1726.7109 | 2509.351 | 3371.9046 | 1996.702 | 1765.2774 | 2936.7355 | 404.2852 | 259.88 | 218.12 | 404.2852 | 469.3638 | 463.6903 | 356.12 | 508.5949 | 467.3884 | 426.471 | 370.6323 | 389.1613 | 292.6529 | 296.4297 | 240.7561 | 391.1819 | 328.2645 | 303.1239 | 279.529 | 330.0206 | 1668.4877 | 958.7974 | 869.0826 | 728.822 | 1011.5742 | 833.7807 | 758.2439 | 518.6775 | 2837.6155 | 1627.6657 | 1416.6735 | 1463.8555
354 | 5.8884 | 815.4179 | 1212.1785 | 1079.9189 | 1096.3712 | 1353.2252 | 1080.7606 | 856.998 | 914.8215 | 641.6775 | 497.0974 | 423.4523 | 495.3874 | 1372.1278 | 833.1603 | 856.3854 | 1051.1278 | 1227.2109 | 842.0791 | 744.0669 | 835.4361 | 1282.1623 | 1485.499 | 1390.501 | 1037.353 | 1968.7099 | 1362.5964 | 1172.1765 | 1577.3367 | 1160.0487 | 796.5781 | 725.7911 | 938.9513 | 1110.2596 | 968.2738 | 886.9838 | 487.2738 | 2713.6593 | 1854.9838 | 1622.6917 | 1979.357 | 1445.7099 | 1119.8215 | 901.1156 | 978.3874 | 433.4036 | 356.9189 | 296.0649 | 741.5781 | 1266.4219 | 1086.0994 | 947.6126 | 603.5964 | 1330.0811 | 1068.7099 | 786.1603 | 722.7891 | 1314.7586 | 1139.4036 | 961.5659 | 858.3387 | 723.9838 | 713.3387 | 497.8377 | 466.1927 | 1241.501 | 1089.5335 | 972.8053 | 983.4036 | 1359.9838 | 1137.0467 | 1174.1947 | 1187.9189 | 1964.7586 | 1247.4848 | 1053.0325 | 1396.3063 | 1805.0142 | 1183.9838 | 862.1947 | 1509.1967 | 845.5335 | 664.7911 | 578.8864 | 908.2089 | 615.4523 | 539.7079 | 537.2414 | 459.1603 | 1251.3712 | 704.9209 | 654.7079 | 1075.2921 | 2842.8864 | 1754.4503 | 1762.4523 | 2453.3205 | 3374.9411 | 1974.0689 | 1837.355 | 2902.5619 | 384.6106 | 241.0325 | 225.7891 | 384.6106 | 499.6775 | 479.0954 | 374.9675 | 498.5821 | 474.5335 | 407.9189 | 364.5152 | 407.6937 | 305.1785 | 291.4036 | 234.0325 | 400.5801 | 331.2738 | 297.7748 | 291.645 | 313.4341 | 1694.8073 | 963.2758 | 853.2089 | 677.6126 | 999.5172 | 805.5821 | 771.4036 | 511.0831 | 2797.1278 | 1640.1562 | 1418.0325 | 1437.351
355 | 5.8994 | 804.3068 | 1219.6076 | 1064.0324 | 1086.9587 | 1317.1828 | 1129.8407 | 848.6461 | 888.4366 | 636.6814 | 506.5428 | 424.3097 | 504.4572 | 1342.7876 | 817.6165 | 834.9882 | 1024.5959 | 1277.5723 | 813.8083 | 768.4277 | 840.6283 | 1304.8083 | 1463.5959 | 1412.4041 | 1012.3393 | 1968.7345 | 1394.3688 | 1154.7552 | 1545.4248 | 1186.711 | 756.2331 | 712.2153 | 906.3215 | 1128.3097 | 955.3835 | 882.8938 | 470.1711 | 2681.8289 | 1892.3157 | 1584.649 | 1992.1091 | 1427.1297 | 1089.9262 | 929.6725 | 949.1946 | 436.3835 | 355.7758 | 310.6696 | 701.2331 | 1271.6636 | 1084.9409 | 950.8289 | 607.6372 | 1360.3599 | 1089.7965 | 770.6165 | 705.7758 | 1313.3628 | 1166.2537 | 968.2359 | 831.9705 | 714.2773 | 690.8318 | 485.0207 | 450.4041 | 1263.7552 | 1122.7758 | 976.7316 | 991.649 | 1353.0855 | 1104.1062 | 1179.7463 | 1174.8407 | 2000.2213 | 1254.4159 | 1057.7021 | 1429.1564 | 1770.9145 | 1173.5752 | 891.9676 | 1556.0973 | 872.1062 | 685.9676 | 566.0531 | 888.7876 | 621.5752 | 519.4455 | 527.0531 | 439.4041 | 1245.82 | 732.4897 | 640.7641 | 1103.2773 | 2863.0502 | 1753.6254 | 1742.5988 | 2445.0502 | 3420.8465 | 2082.9558 | 1842.8083 | 2876.7316 | 353.1475 | 252.7227 | 205.6165 | 353.1475 | 495.0324 | 449.2979 | 356.9587 | 510.3952 | 476.1828 | 404.6696 | 363.2006 | 410.2979 | 322.7876 | 295.4366 | 241.8614 | 386.8407 | 323.2979 | 317.3186 | 280.5752 | 279.0324 | 1698.3097 | 1009.6372 | 863.6254 | 695.5723 | 1005.5014 | 835.2979 | 770.1711 | 557.7876 | 2783.9351 | 1553.8496 | 1424.4572 | 1391.6992
356 | 5.9104 | 834.9135 | 1181.3736 | 1063.5364 | 1050.4636 | 1308.4258 | 1142.5295 | 910.4533 | 901.7994 | 615.1902 | 509.2665 | 419.7095 | 501.0069 | 1340.4499 | 809.8338 | 859.7266 | 1018.7301 | 1293.9341 | 794.294 | 726.9966 | 828.8166 | 1326.7129 | 1498.4224 | 1413.5467 | 995.9444 | 1961.9066 | 1455.9203 | 1167.4464 | 1572.6264 | 1192.6642 | 782.8132 | 672.0797 | 867.1731 | 1116.443 | 941.3668 | 893.533 | 465.443 | 2730.1559 | 1944.6538 | 1584.3633 | 1929.9135 | 1389.9032 | 1073.91 | 868.647 | 891.7198 | 419.0968 | 368.3668 | 300.1106 | 730.7198 | 1182.0034 | 1027.8722 | 964.6401 | 635.7232 | 1379.8338 | 1109.1902 | 791.1731 | 714.0069 | 1300.5364 | 1159.3077 | 913.8304 | 784.9066 | 690.5536 | 644.0034 | 487.4636 | 454.09 | 1268.09 | 1107.0275 | 1010.9238 | 980.4636 | 1357.443 | 1119.6264 | 1169.2527 | 1196.6126 | 2029.0137 | 1222.4636 | 1064.4499 | 1452.7644 | 1781.4533 | 1172.533 | 898.6401 | 1503.3805 | 882.5398 | 680.2871 | 571.6332 | 880.6367 | 615.8098 | 544.0034 | 544.6229 | 448.9032 | 1250.9794 | 735.4464 | 681.3633 | 1062.4773 | 2895.4808 | 1803.5433 | 1675.7301 | 2493.8304 | 3342.0488 | 2076.2177 | 1792.0481 | 2906.9272 | 368.1037 | 267 | 212.3668 | 368.1037 | 469.467 | 453.8132 | 352.4361 | 486.1731 | 455.0728 | 422.0865 | 377.5604 | 411.5433 | 305.9169 | 286.6367 | 258.4464 | 385.3599 | 328.5398 | 317.3736 | 276.6264 | 271.9966 | 1683.5364 | 1016.2871 | 892.8338 | 719.2006 | 990.2596 | 838.3599 | 759.9931 | 561.9897 | 2796.5536 | 1569.3771 | 1419.9169 | 1431.8372
357 | 5.9214 | 811.351 | 1181.6588 | 1087.6342 | 1058.6687 | 1385.7808 | 1150.9606 | 893.0924 | 921.2192 | 598 | 532.8928 | 439.4729 | 477.9852 | 1392.4445 | 804.2931 | 868.8879 | 993.8436 | 1272.6736 | 770.1564 | 747.5221 | 823.7365 | 1334.9655 | 1575.2685 | 1413.2537 | 950.6687 | 2011.3226 | 1405.2045 | 1150.5702 | 1585.4877 | 1154.7808 | 797.5074 | 682.9803 | 889.5074 | 1140.3756 | 965.8004 | 901.3461 | 489.3756 | 2758.9606 | 1891.7217 | 1617.6785 | 1907.1022 | 1417.4002 | 1073.8781 | 853.7611 | 923.5271 | 397.234 | 348.8535 | 268.6391 | 729.8485 | 1194.9076 | 1054.9213 | 950.6145 | 626.3461 | 1342.7414 | 1089.3707 | 793.2242 | 718.4051 | 1309.9852 | 1094.0098 | 907.4051 | 804.2734 | 693.1564 | 646.766 | 475.7611 | 435.7168 | 1280.5172 | 1075.7906 | 981.3756 | 972.8928 | 1394.1465 | 1133.2389 | 1211.8781 | 1223.9655 | 1991.234 | 1219.4002 | 1076.2537 | 1375.1219 | 1783.2488 | 1180.3461 | 908.2783 | 1465.2438 | 867.3313 | 657.0098 | 585.8879 | 898.0295 | 630.3707 | 540.0049 | 558.4779 | 457.9951 | 1283.9852 | 747.4926 | 677.117 | 987.8289 | 2848.3855 | 1797.1072 | 1689.9076 | 2433.6922 | 3414.4779 | 1976.7069 | 1746.5861 | 2840.702 | 352.8928 | 268.5025 | 216.1416 | 352.8928 | 449.2586 | 465.5025 | 394.5074 | 492.3559 | 485.2389 | 432.5025 | 365.4051 | 417.8781 | 292.2537 | 286.7512 | 251.7119 | 379.4779 | 331.7365 | 300.0049 | 288.3609 | 285.7611 | 1699.3707 | 976.4828 | 848.9803 | 685.5074 | 996.3412 | 851.6342 | 766.8633 | 572.4828 | 2795.7758 | 1536.3658 | 1420.9027 | 1408.7758
358 | 5.9324 | 797.8614 | 1205.1386 | 1067.907 | 1083.6047 | 1422.0428 | 1087.3377 | 853.4056 | 910.1711 | 611.9646 | 527.7256 | 424.413 | 470.3702 | 1421.494 | 849.5516 | 859.8893 | 986.826 | 1300.2772 | 792.6902 | 754.4837 | 803.4056 | 1314.7404 | 1565.6828 | 1414.2847 | 967.4587 | 2029.1415 | 1362.8791 | 1117.4986 | 1542.3554 | 1173.587 | 805.3879 | 666.4912 | 904.7581 | 1159.6195 | 989.4587 | 852.9926 | 486.1209 | 2757.4027 | 1899.4632 | 1628.5545 | 1892.9218 | 1440.6726 | 1073.8967 | 862.6726 | 973.8865 | 408.5413 | 335.5339 | 286.1033 | 720.4558 | 1238.0856 | 1016.1181 | 929.5265 | 588.854 | 1364.6446 | 1089.6549 | 779.0251 | 728.0531 | 1287.2419 | 1117.1888 | 917.0531 | 813.2596 | 723.0605 | 669.9823 | 490.1033 | 390.7507 | 1290.2419 | 1105.7507 | 983.5516 | 997.2065 | 1388.4587 | 1115.0177 | 1199.4837 | 1167.2772 | 1945.907 | 1225.2168 | 1095.1283 | 1344.4587 | 1811.093 | 1181.2567 | 912.9498 | 1436.5014 | 849.3776 | 676.6977 | 600.1637 | 914.0354 | 631.4307 | 541.1209 | 534.8363 | 455.3274 | 1291.4986 | 734.1814 | 669.7758 | 956.1283 | 2925.7199 | 1803.0502 | 1685.3731 | 2409.9116 | 3347.3908 | 1964.8083 | 1822.3628 | 2816.826 | 366.733 | 268.2847 | 227.5693 | 366.733 | 469.2493 | 467.6121 | 389.9749 | 474.9395 | 473.6121 | 429.9572 | 381.6475 | 416.3451 | 310.3525 | 281.0177 | 221.5339 | 358.9395 | 323.4307 | 315.4735 | 281.3702 | 303.5944 | 1707.8009 | 968.3098 | 842.1888 | 712.0074 | 968.2065 | 843.9321 | 749.5693 | 532.8893 | 2848.7655 | 1502.9498 | 1404.7787 | 1418.32
359 | 5.9434 | 824.2045 | 1184.3982 | 1006.7925 | 1075.0109 | 1411.9911 | 1075.6126 | 837.3972 | 920.3893 | 620.3933 | 508.6087 | 437.5968 | 486.002 | 1381.6126 | 815.9822 | 849.8122 | 1030.002 | 1283.2085 | 792.4071 | 743.2065 | 806.6067 | 1311.587 | 1527.4032 | 1387.1888 | 920.3982 | 1967.413 | 1397.8142 | 1143.7955 | 1513.2095 | 1170.4081 | 802.7984 | 676.4022 | 906.5949 | 1146.9793 | 967.2105 | 849.0128 | 469.1947 | 2822.4022 | 1935.3893 | 1540.7905 | 1903.7984 | 1424.9941 | 1057.3943 | 855.7984 | 979.1848 | 427.9882 | 355.8034 | 292.2006 | 745.999 | 1256.001 | 1062.5909 | 928.6077 | 579.2016 | 1429.8083 | 1067.7856 | 775.589 | 694.5998 | 1319.4219 | 1126.3943 | 907.6116 | 821.6156 | 719.5998 | 641.3805 | 482.9941 | 400.6136 | 1275.5988 | 1091.2016 | 1006.0099 | 986.9901 | 1365.4101 | 1101.999 | 1184.6047 | 1157.412 | 1913.6067 | 1193.1996 | 1112.5968 | 1315.8073 | 1870.6067 | 1183.7777 | 873.7994 | 1449.8241 | 885.7974 | 700.8043 | 591.1947 | 891.5998 | 631.9961 | 532.999 | 505.7866 | 456.5988 | 1282.1779 | 718.997 | 649.3893 | 992.8063 | 3027.7984 | 1810.9852 | 1658.416 | 2423.7797 | 3344.2331 | 2025.8053 | 1820.4081 | 2810.3775 | 368.0039 | 241.9891 | 218.5998 | 368.0039 | 489.001 | 466.5998 | 366.998 | 484.003 | 469.3982 | 423.6008 | 355.3982 | 403.7974 | 336.2016 | 267.5988 | 239.8024 | 367.2026 | 333.2006 | 313.9911 | 281.3943 | 315.2006 | 1720.7964 | 1006.2114 | 858.5978 | 714.7866 | 973.1976 | 823.8034 | 745.4022 | 520.0109 | 2868.1888 | 1475.0049 | 1375.8201 | 1411.2095
360 | 5.9544 | 839.2375 | 1192.1996 | 1031.8486 | 1077.4759 | 1378.8127 | 1094.2006 | 837.2498 | 888 | 612.7871 | 520.1504 | 425.9376 | 500.2498 | 1419.5875 | 778.4995 | 859.3633 | 1007.9508 | 1316.8373 | 816.8874 | 761.35 | 825.35 | 1296.3368 | 1529.5752 | 1362.4872 | 944.2867 | 1987.2006 | 1447.1873 | 1134.5374 | 1504.0511 | 1185.5128 | 778.613 | 680.5251 | 894.2999 | 1108.7493 | 991.5752 | 900.3747 | 450.1126 | 2805.0757 | 1945.8988 | 1523.9622 | 1858.9886 | 1456.9111 | 1039.5875 | 847.7001 | 917.3254 | 404.7871 | 356.1003 | 289.4626 | 732.6878 | 1269.7247 | 1030.0501 | 939.6253 | 600.912 | 1432.3131 | 1047.0993 | 754.3747 | 675.0255 | 1352.176 | 1125.087 | 942.175 | 834.7635 | 704.1504 | 609.2242 | 475.3122 | 434.1627 | 1268.2876 | 1078.6755 | 1015.3756 | 982.1618 | 1390.2252 | 1102.2999 | 1191.1627 | 1205.9622 | 1927.4002 | 1203.9622 | 1119.912 | 1349.5497 | 1869.5506 | 1143.6367 | 888.087 | 1476.3888 | 866.8628 | 683.8883 | 584.0747 | 887.7001 | 625.3623 | 531.65 | 489.7493 | 459.5998 | 1269.036 | 705.4626 | 627.3122 | 1026.0246 | 3009.8004 | 1808.8988 | 1685.1381 | 2344.263 | 3425.613 | 2028.3879 | 1796.3264 | 2728.8004 | 387.0123 | 220.5497 | 215.9376 | 387.0123 | 502.3122 | 468.0624 | 354.4002 | 495.8874 | 450.8004 | 427.2375 | 359.8997 | 391.4626 | 333.9877 | 266.0624 | 251.1249 | 368.8127 | 326.7502 | 291.5374 | 269.7748 | 306.2753 | 1708.7625 | 994.8013 | 862.7247 | 681.8373 | 1010.1363 | 843.4872 | 760.6623 | 543.9253 | 2856.2744 | 1527.5364 | 1418.8505 | 1430.1003
361 | 5.9654 | 827.4071 | 1195.7079 | 1070.2654 | 1052.6903 | 1358.531 | 1078.7434 | 833.7079 | 867.1858 | 639.0708 | 509.1947 | 428.0974 | 512.8496 | 1378.2212 | 813.0443 | 826.3717 | 995.1859 | 1308.4601 | 826.0974 | 732.6637 | 823.8496 | 1309.761 | 1524.5221 | 1355.1858 | 974.9557 | 1989.1593 | 1410.1238 | 1111.2654 | 1489.0797 | 1166.9026 | 785.4071 | 681.6726 | 889.2301 | 1143.5752 | 987.7965 | 889.115 | 465.5399 | 2742.9646 | 1961.1238 | 1550.3451 | 1790.9557 | 1498.6902 | 1046.646 | 840.4513 | 917.0885 | 416.6283 | 359.416 | 291.9204 | 705.2566 | 1259.0885 | 1031.9204 | 926.2478 | 612.6814 | 1407.2213 | 1074.9026 | 786.0177 | 656.0708 | 1328.2124 | 1139.9292 | 952.8938 | 797.0443 | 678.7522 | 650.9735 | 492.9735 | 429.3717 | 1282.1151 | 1052.3009 | 1001.2921 | 1011.2035 | 1346.3185 | 1093.4071 | 1169.8053 | 1199.2124 | 1890.8141 | 1241.8142 | 1112.5663 | 1354.531 | 1855.9912 | 1200.416 | 895.7079 | 1427.8407 | 836.5575 | 653.3186 | 599.5487 | 880.0265 | 617.7079 | 537.3982 | 501.6371 | 465.6018 | 1295.4512 | 687.531 | 623.7345 | 1037.0796 | 2999.9204 | 1801.6105 | 1686.2832 | 2336.416 | 3407.3451 | 1985.6371 | 1761.4071 | 2698.531 | 403.7965 | 235.7257 | 214.2743 | 403.7965 | 501.2832 | 473.1239 | 354.7699 | 487.6814 | 452.8142 | 439.1947 | 371.0265 | 396.8938 | 317.6283 | 274.0974 | 257.4248 | 359.5752 | 317.5487 | 295.4513 | 271.6283 | 283.7787 | 1675.5663 | 961.708 | 862.2832 | 712.5399 | 1036.8937 | 822.5929 | 753.4601 | 511.8849 | 2841.2743 | 1545.7433 | 1413.3894 | 1402.8318
362 | 5.9764 | 816.8673 | 1136.1007 | 1037.4302 | 1057.8948 | 1334.2586 | 1109.1121 | 837.4759 | 832.4439 | 641.2586 | 513.3844 | 423.6362 | 516.6224 | 1350.9725 | 801.3387 | 843.3043 | 1035.4966 | 1299.4828 | 812.8948 | 699.1854 | 825.4371 | 1333.0114 | 1535.0595 | 1337.9268 | 971.8009 | 2045.2975 | 1364.1716 | 1080.1785 | 1522.7689 | 1155.1922 | 821.9474 | 673.1396 | 871.6293 | 1153.4897 | 980.8284 | 829.4577 | 479.762 | 2769.2494 | 1881.5172 | 1581.794 | 1843.7459 | 1469.0526 | 1050.0069 | 876.3363 | 978.4691 | 405.8146 | 363.7689 | 289.1396 | 705.4828 | 1239.2975 | 1059.2975 | 926.3181 | 590.3181 | 1420.9336 | 1060.1327 | 783.5835 | 670.9336 | 1359.2517 | 1128.5629 | 956.325 | 826.8604 | 676.4966 | 644.0732 | 488.2586 | 440.6224 | 1321.1121 | 1050.4371 | 1021.1259 | 975.7758 | 1322.4759 | 1105.1579 | 1159.9336 | 1154.762 | 1913.4416 | 1257.325 | 1058.1991 | 1336.3776 | 1901.2906 | 1205.643 | 887.675 | 1447.6293 | 848.9061 | 685.2975 | 597.3181 | 838.4119 | 602.2448 | 528.0137 | 469.1259 | 448.762 | 1278.1372 | 661.9474 | 626.9725 | 1003.5835 | 3015.4966 | 1720.8009 | 1725.1922 | 2332.4646 | 3381.9749 | 1988.9656 | 1776.5309 | 2648.0344 | 379.7826 | 245.8009 | 222.5561 | 379.7826 | 466.0137 | 469.4439 | 356.1991 | 490.2311 | 474.0069 | 444.6957 | 369.1854 | 398.1396 | 316.4302 | 270.0732 | 252.3181 | 362.4966 | 307.4508 | 284.5698 | 275.675 | 290.6087 | 1644.3638 | 1019.9863 | 846.6819 | 713.2128 | 1000.238 | 803.5354 | 749.2906 | 501.7345 | 2802.4759 | 1489.2586 | 1450.6156 | 1365.3318
363 | 5.9874 | 830.1455 | 1118.8034 | 1064.5251 | 1031.5506 | 1323 | 1091.3992 | 873.943 | 851.8604 | 626.4051 | 483.9243 | 390.5762 | 502.4749 | 1387.4494 | 792.6136 | 872.8732 | 1030.5634 | 1326.0443 | 788.3481 | 713.6834 | 814.766 | 1344.2792 | 1556.0757 | 1360.9872 | 973.9685 | 2060.5064 | 1392.6391 | 1104.5821 | 1484.5192 | 1170.12 | 803.1583 | 650.8545 | 865.5506 | 1146.3294 | 966.7404 | 823.6706 | 455.0128 | 2806.6145 | 1927.303 | 1612.2911 | 1912.1268 | 1449.1966 | 1059.8732 | 899.5889 | 991.1396 | 406.5949 | 347.5379 | 254.2723 | 712.7217 | 1247.3166 | 1067.3166 | 900.1268 | 572.2153 | 1397.2409 | 1045.9557 | 768.1328 | 682.2911 | 1375.8732 | 1132.0443 | 935.057 | 832.5192 | 676.5064 | 639.4366 | 473.4051 | 434.5634 | 1294.8613 | 1074.8162 | 1013.7089 | 959.5251 | 1362.9872 | 1105.3549 | 1152.4179 | 1144.8417 | 1999.4366 | 1263.4681 | 1037.6077 | 1328.9557 | 1855.8553 | 1159.5319 | 901.3038 | 1450.5379 | 890 | 684.3294 | 614.2655 | 812.9174 | 600.9115 | 501.9115 | 493.6136 | 433.8987 | 1324.2281 | 674.6136 | 627.9498 | 981.3923 | 3024.4936 | 1772.8476 | 1692.4877 | 2303.7404 | 3341.5379 | 2009.2851 | 1750.3796 | 2555.817 | 344.6391 | 250.2153 | 221.5634 | 344.6391 | 465.9745 | 456.057 | 349.9872 | 498.823 | 469.0443 | 429.9557 | 370.2026 | 392.0315 | 321.057 | 270.3796 | 241.8545 | 374.1898 | 294.0443 | 272.3038 | 272.2281 | 298.0698 | 1694.9489 | 1020.4749 | 858.4936 | 689.886 | 1015.5506 | 817.4306 | 781.5949 | 558.8476 | 2866.8034 | 1501.3668 | 1457.1583 | 1379.9932
364 | 5.9984 | 839.5398 | 1213.0134 | 1076.7669 | 1022.2286 | 1307.3038 | 1080.854 | 872.6136 | 861.5324 | 646.0059 | 470.3864 | 396.1608 | 468.997 | 1376.4587 | 839.6962 | 844.7566 | 1017.1534 | 1285.2964 | 802.9263 | 735.8481 | 802.385 | 1318.9337 | 1536.3746 | 1349.379 | 951.9985 | 2007.9114 | 1410.3097 | 1114.4602 | 1462.4602 | 1195.5383 | 799.6932 | 667.3127 | 861.7684 | 1139.9278 | 965.3053 | 824.5324 | 454.618 | 2795.7803 | 1936.6739 | 1562.2934 | 1862.6858 | 1419.618 | 1054.3775 | 866.9233 | 926.9926 | 410.0767 | 331.6858 | 258.0826 | 697.2316 | 1236.6932 | 1036.8422 | 924.5502 | 587.1608 | 1391.5458 | 1036.8451 | 784.385 | 688.615 | 1351.4498 | 1121.7655 | 922.5369 | 806.3053 | 697.7758 | 633.4572 | 479.1564 | 420.6917 | 1295.09 | 1099.3835 | 994.4587 | 967.1504 | 1385.0797 | 1104.6269 | 1160.0826 | 1162.618 | 1984.2242 | 1257.2994 | 1050.3894 | 1301.379 | 1814.3968 | 1170.3172 | 886.4587 | 1419.9129 | 845.6814 | 659.4617 | 625.9955 | 870.0841 | 597.4631 | 510.0045 | 510.2271 | 432.2301 | 1294.4602 | 701.8451 | 602.0811 | 989.385 | 3069.2405 | 1805.0737 | 1696.4706 | 2277.3761 | 3350.615 | 2020.5487 | 1720.4498 | 2519.236 | 353.0767 | 232.3835 | 208.1534 | 353.0767 | 481.9985 | 445.8451 | 367.4631 | 475.3068 | 447.6858 | 435.1564 | 364.7684 | 411.2316 | 322.8481 | 269.5339 | 230.1519 | 357.1475 | 291.6136 | 282.8496 | 260.9263 | 277.7684 | 1730.2286 | 1008.6947 | 875.0797 | 715.8496 | 1036.236 | 809.4661 | 761.5324 | 594.3053 | 2879.7625 | 1535.0782 | 1440.3053 | 1410.9955
365 | 6.0094 | 845.5261 | 1289.2704 | 1049.412 | 994.4641 | 1300.3746 | 1107.0993 | 838.881 | 859.0669 | 659.3077 | 515.0197 | 423.7866 | 491.9676 | 1392.3845 | 820.5585 | 818.65 | 1039.1711 | 1262.3324 | 793.3549 | 728.412 | 795.943 | 1385.1141 | 1527.3894 | 1355.2282 | 938.8387 | 1940.6775 | 1390.1416 | 1127.9056 | 1442.2556 | 1173.6725 | 817.7444 | 666.3549 | 861.7915 | 1171.0049 | 948.3127 | 812.588 | 471.2655 | 2803.9232 | 1835.7817 | 1522.0148 | 1837.2134 | 1452.8535 | 1039.9577 | 877.0049 | 896.7394 | 402.8387 | 313.3225 | 255.6155 | 675.407 | 1226.3077 | 1027.4789 | 930.2359 | 582.5585 | 1360.1218 | 1038.1613 | 786.9479 | 672.3077 | 1320.6451 | 1092.1514 | 932.062 | 801.1613 | 710.5162 | 620.1092 | 498.1092 | 417.5261 | 1316.7148 | 1048.0275 | 986.6873 | 951 | 1450.6077 | 1123.4542 | 1193.1613 | 1162.2035 | 1964.3746 | 1217.0521 | 1051.1986 | 1300.119 | 1809.4394 | 1224.4789 | 886.2704 | 1386.6923 | 843.7641 | 674.2183 | 591.7296 | 866.5585 | 609.2655 | 533.3697 | 501.531 | 434.2134 | 1268.5683 | 714.062 | 624.8437 | 974.8859 | 3065.2359 | 1813.2232 | 1714.3549 | 2246.0049 | 3349.9479 | 2007.3873 | 1728.9824 | 2533.0993 | 325.4592 | 232.6873 | 196.5211 | 325.4592 | 460.7817 | 434.8387 | 370.7345 | 501.8063 | 426.0049 | 447 | 376.1662 | 417.8437 | 319.6775 | 262.2183 | 232.3746 | 330.3176 | 306.6972 | 286.2035 | 275.0521 | 296.7493 | 1740.8535 | 995.8289 | 901.2704 | 729.1563 | 1037.7768 | 868.9155 | 723.0472 | 553.6155 | 2860.7965 | 1538.1563 | 1436.1092 | 1413.2232
366 | 6.0204 | 855.6972 | 1245.1656 | 998.9085 | 1002.2291 | 1287.7158 | 1086.734 | 835.3117 | 870.4956 | 621.6057 | 512.2113 | 415.0826 | 517.4405 | 1410.413 | 791.6421 | 832.1745 | 1037.3122 | 1315.9356 | 778.1559 | 722.6421 | 805.4769 | 1365.0737 | 1536.4498 | 1344.3215 | 946.587 | 1979.0271 | 1377.9081 | 1132.9818 | 1443.532 | 1162.1559 | 832.0275 | 659.9081 | 853.8441 | 1169.8166 | 952.6972 | 816.8992 | 460.9267 | 2802.1648 | 1867.651 | 1560.266 | 1807.5049 | 1497.8806 | 1083.853 | 870.9543 | 928.7429 | 399.8899 | 348.8806 | 249.2567 | 676.6603 | 1257.1648 | 1019.0368 | 912.4401 | 585.2474 | 1361.3939 | 1063.9632 | 800.1559 | 669.1283 | 1330.4956 | 1075.1101 | 937.3854 | 773.4773 | 709.9356 | 623.1377 | 504.0551 | 432.0733 | 1291.4494 | 1028.119 | 981.8166 | 935.4405 | 1480.3122 | 1081.5688 | 1182.4956 | 1146.0551 | 1995.9632 | 1204.3304 | 1054.7154 | 1312.9267 | 1800.7793 | 1234.0275 | 886.8992 | 1385.0093 | 893.6514 | 683.1927 | 590.1283 | 840.0733 | 636.3668 | 520.9267 | 512.4587 | 434.1652 | 1268.3943 | 724.2478 | 606.1194 | 959.4312 | 2997.8441 | 1839 | 1695.7522 | 2190.8446 | 3346.6239 | 1956.5044 | 1755.7526 | 2561.3575 | 331.4951 | 234.6239 | 199.6146 | 331.4951 | 445.4862 | 445.5044 | 372.8073 | 514.1194 | 439.4036 | 450.4036 | 366.2571 | 443.853 | 308.4312 | 276.5413 | 245.4862 | 348.587 | 325.7065 | 292.9081 | 276.9174 | 313.734 | 1763.5138 | 996.2291 | 908.7065 | 719.6789 | 1020.7522 | 870.6425 | 728.532 | 522.9449 | 2878.8899 | 1548.1283 | 1436.7065 | 1436.5688
367 | 6.0314 | 871.4867 | 1189.9794 | 1005.8791 | 1000.5899 | 1300.8821 | 1075.5074 | 852.0265 | 844.0236 | 599.4985 | 490.9292 | 393.056 | 496.0147 | 1397.4926 | 826.4366 | 792.5428 | 1013.4808 | 1291.6342 | 773.5516 | 715.5604 | 816.5251 | 1338.4602 | 1487.0413 | 1330.3953 | 938.0708 | 2060.9204 | 1371.584 | 1145.4101 | 1455.0177 | 1178.4897 | 832.9941 | 664.5516 | 872.9056 | 1185.9174 | 992.416 | 824.9322 | 446.9941 | 2851.0029 | 1888.587 | 1588.4425 | 1816.3629 | 1492.584 | 1064.6283 | 871.4484 | 927.587 | 411.9764 | 365.5192 | 282.9823 | 674.1002 | 1285.0649 | 1027.416 | 926.5752 | 597.6017 | 1404.1061 | 1066.0531 | 806.5192 | 677.0177 | 1319.4779 | 1126.3364 | 916.0206 | 757.4396 | 709.0649 | 629.9646 | 486.0472 | 436.0354 | 1330.944 | 1069.9735 | 979.9705 | 942.9292 | 1428.5634 | 1087.9735 | 1174.469 | 1105.6135 | 1992.5693 | 1251.8231 | 1067.5487 | 1318.4366 | 1847.5516 | 1184.1445 | 871.0029 | 1401.8939 | 863.0737 | 657.056 | 594.0295 | 860.9853 | 618.1238 | 516.4661 | 523.9705 | 426.0059 | 1299.9735 | 711.5103 | 599.4513 | 942.5575 | 3055.2921 | 1811.5811 | 1683.0649 | 2179.3305 | 3326.0472 | 2035.3216 | 1733.4278 | 2596.5221 | 360.5251 | 232.4926 | 210.0118 | 360.5251 | 448.9912 | 464.4985 | 372.0236 | 475.0472 | 433.0295 | 428.0767 | 345.9764 | 427.1209 | 305.0177 | 276.9941 | 237.0265 | 349.0442 | 307.0324 | 294.5604 | 259.0619 | 292.0383 | 1750.0383 | 993.0944 | 881.059 | 694.056 | 1012.5516 | 846.944 | 751.9823 | 540.9528 | 2847.1061 | 1584.4337 | 1404.0737 | 1420.5398
368 | 6.0424 | 864.762 | 1211.3029 | 1025.0349 | 962.8486 | 1346.1082 | 1064.3161 | 824.0998 | 850.8137 | 626.0517 | 510.9135 | 398.0084 | 488.4459 | 1387.2296 | 851.0649 | 797.411 | 1031.7488 | 1262.3461 | 784.0949 | 698.5757 | 814.6406 | 1396.9519 | 1515.3978 | 1346.0781 | 937.4976 | 2030.2777 | 1401.7439 | 1117.2561 | 1452.5757 | 1197.3245 | 834.4892 | 659.7704 | 885.0781 | 1193.0565 | 1007.2079 | 806.6238 | 430.6106 | 2796.8751 | 1861.0433 | 1557.4291 | 1820.0914 | 1472.1731 | 1024.5757 | 885.9351 | 915.8786 | 422.0433 | 356.9567 | 266.0132 | 677.8004 | 1282.9219 | 1027.9051 | 930.1166 | 600.8004 | 1373.619 | 1011.732 | 783.6539 | 636.2645 | 1341.8137 | 1138.5806 | 904.4026 | 814.7788 | 694.1514 | 622.0781 | 478.1731 | 430.1298 | 1286.1479 | 1048.351 | 958.3294 | 945.5457 | 1445.8221 | 1099.5541 | 1148.2212 | 1091.5192 | 1984.3245 | 1286.97 | 1053.0433 | 1301.1779 | 1876.4843 | 1180.4627 | 878.6839 | 1403.2645 | 881.9302 | 688.06 | 581.9567 | 822.5806 | 642.4062 | 510.6322 | 507.9483 | 415.3161 | 1313.5974 | 727.411 | 626.2163 | 956.7139 | 3123.9567 | 1813.1166 | 1683.9868 | 2190.0049 | 3282.4159 | 2023.9748 | 1756.9784 | 2560.9051 | 358.6406 | 226.8269 | 188.1214 | 358.6406 | 476.5192 | 461.4243 | 359.4026 | 460.0216 | 438.3245 | 405.0649 | 361.1514 | 397.7488 | 310.238 | 279 | 228.5108 | 344.2163 | 309.2812 | 276.5541 | 240.5541 | 257.5457 | 1738.5325 | 996.7572 | 877.3461 | 699.5192 | 1046.0817 | 872.1298 | 727.8618 | 526.8618 | 2788.0132 | 1570.2212 | 1382.0649 | 1397.2945
369 | 6.0534 | 852.9538 | 1217.1534 | 1016.5546 | 994.71 | 1343.1534 | 1046.5841 | 826.4012 | 846.1681 | 630.5988 | 474.0442 | 406.8761 | 489.1386 | 1355.645 | 816.8593 | 840.063 | 992.2585 | 1300.4936 | 772.8131 | 741.2331 | 831.9853 | 1401.29 | 1529.8446 | 1307.645 | 936.4601 | 2052.9431 | 1413.5358 | 1111.2184 | 1431.937 | 1192.645 | 797.9056 | 659.9685 | 878.0315 | 1167.2458 | 963.6597 | 812.5104 | 414.5693 | 2751.2311 | 1836.3215 | 1569.6805 | 1740.8131 | 1539.6343 | 1057.294 | 876.1996 | 919.9223 | 408.8299 | 360.2311 | 267.5399 | 690.9833 | 1306.1849 | 1015.647 | 952.2458 | 594.6283 | 1393.647 | 1026.6952 | 789.4474 | 617.1239 | 1369.6008 | 1094.9076 | 893.1996 | 824.8908 | 697.3382 | 604.0462 | 478.6765 | 429.1996 | 1300.8506 | 1054.0482 | 958.3403 | 903.0295 | 1507.5861 | 1070.0904 | 1152.2331 | 1128.6008 | 1975.9833 | 1264.0462 | 1026.7522 | 1285.4474 | 1909.4916 | 1216.6765 | 906.355 | 1360.061 | 897.3215 | 696.7669 | 560.645 | 798.7857 | 673.0442 | 484.061 | 490.3235 | 434.7247 | 1312.2458 | 729.7837 | 636.5231 | 984.292 | 3107.8761 | 1813.7984 | 1699.7227 | 2140.292 | 3275.8781 | 2025.6196 | 1773.7857 | 2525.1072 | 358.1996 | 232.0777 | 197.6785 | 358.1996 | 490.061 | 460.6155 | 349.7689 | 466.7542 | 435.2143 | 406.4307 | 363.8151 | 406.9076 | 316.8151 | 273.7689 | 237.3697 | 354 | 311.0147 | 276.9076 | 237.7689 | 266.294 | 1690.8279 | 1033.6155 | 857.4287 | 716.7227 | 1062.5673 | 865.9685 | 739.2793 | 515.7857 | 2848.651 | 1507.7984 | 1361.4601 | 1397.4159
370 | 6.0644 | 863.2434 | 1198.2212 | 1027.4646 | 1006.9778 | 1366.5177 | 1049.2434 | 840.7168 | 835.4248 | 619.4956 | 456.8496 | 401.4956 | 480.7566 | 1369.5885 | 797.9912 | 813.2787 | 963.1328 | 1266.3893 | 793.6815 | 736.8451 | 813.6239 | 1387.416 | 1523.1328 | 1319.4469 | 924.1018 | 2050.6504 | 1362.8274 | 1087.3981 | 1451.7611 | 1164.8274 | 776.2434 | 668.2744 | 866.5132 | 1164.1416 | 969.7301 | 851.8584 | 407.969 | 2799.9027 | 1851.6195 | 1574.1371 | 1737.0531 | 1565.6239 | 1062.0486 | 861.4336 | 883.907 | 400.354 | 367.677 | 286.323 | 672.1416 | 1275.7345 | 1024.1814 | 921.8363 | 553.323 | 1373.2699 | 1025.9557 | 770.4513 | 634.2832 | 1320.4291 | 1061.9469 | 935.7213 | 796.1106 | 727.1947 | 599.1106 | 455.938 | 442.8098 | 1355.0088 | 1076.8982 | 969.938 | 930.1151 | 1491.4513 | 1068.3053 | 1168.7256 | 1143.1416 | 1940.0088 | 1239.8363 | 1025.6283 | 1305 | 1873.3009 | 1201.4336 | 902.584 | 1336.5442 | 873.5354 | 682.9514 | 557.9912 | 815 | 643 | 501.7699 | 517.3452 | 444.4734 | 1304.323 | 735.8806 | 625.2212 | 981.0708 | 3106.7788 | 1784.7168 | 1684.8982 | 2147.7788 | 3326.6593 | 2017.1857 | 1759.4823 | 2469.3495 | 354.677 | 242.3938 | 211.8274 | 354.677 | 479.3938 | 473.1726 | 342.323 | 468.2522 | 447.9115 | 417.8496 | 351.969 | 402.8584 | 313.5354 | 270.0708 | 231.615 | 343.8274 | 325.6283 | 281.4248 | 247.9912 | 269.4424 | 1677.5885 | 1001.1991 | 857.1239 | 708.8584 | 1037.2832 | 868.9204 | 761.0088 | 542.1726 | 2857.6105 | 1509.2345 | 1353.385 | 1391.5442
371 | 6.0754 | 860.9508 | 1177.5939 | 1031.9292 | 984.8092 | 1416.5507 | 1050.7847 | 816.1878 | 812.5201 | 609.0216 | 460.3323 | 403.0708 | 487.9784 | 1370.1632 | 820.8584 | 787.4307 | 984.5723 | 1292.3905 | 812.9724 | 715.6922 | 805.4307 | 1387.9262 | 1503.4955 | 1315.9724 | 935.2861 | 1999.0492 | 1357.8338 | 1057.6706 | 1420.0864 | 1148.3323 | 781.6185 | 665.2832 | 844.8092 | 1162.1662 | 1000.0954 | 859.3107 | 433.3137 | 2830.88 | 1869.5693 | 1543.1446 | 1717.994 | 1522.9262 | 1018.354 | 845.237 | 856.4277 | 412.12 | 351.9262 | 269.7601 | 681.6677 | 1229.0216 | 1031.5015 | 908.2153 | 566.3353 | 1405.889 | 990.3815 | 763.2891 | 634.1662 | 1322.8644 | 1067.3137 | 941.3048 | 775.117 | 705.0648 | 609.3353 | 464.4553 | 442.4739 | 1380.886 | 1086.5261 | 987.766 | 932.114 | 1467.8584 | 1128.6952 | 1146.5231 | 1117.0678 | 1933.0246 | 1228.4061 | 1032.7601 | 1303.9046 | 1839.4985 | 1187.9754 | 870.8554 | 1325.6431 | 901.9322 | 711.8584 | 596.7414 | 836.9076 | 626.5693 | 534.3107 | 533.9046 | 457.8368 | 1302.5477 | 735.9262 | 576.6617 | 956.8063 | 3151.5045 | 1824.2675 | 1691.7876 | 2141.5693 | 3297.4926 | 1963.3107 | 1741.763 | 2448.8122 | 361.3815 | 244.2861 | 199.1662 | 361.3815 | 479.0954 | 480.3569 | 340.5477 | 454.0462 | 468.6185 | 430.6431 | 326.9262 | 403.1908 | 308.5231 | 275.9292 | 228.763 | 332.8092 | 346.4523 | 274.4031 | 258.8092 | 269.8122 | 1762.5074 | 1001.1229 | 848.8525 | 710.2861 | 1028.4031 | 837.1848 | 746.3569 | 533.4739 | 2798.9046 | 1527.6185 | 1362.8092 | 1366.4031
372 | 6.0864 | 868.3594 | 1187.4794 | 1023.6804 | 983.56 | 1376.642 | 1085.9986 | 795.6401 | 808.5993 | 607.3196 | 472.4794 | 394.5202 | 496.9602 | 1393.3585 | 811.6406 | 767.4011 | 998.3997 | 1314.9213 | 814.7591 | 704.6008 | 777.5614 | 1387.5595 | 1493.4794 | 1327.2788 | 928.4803 | 2086.2774 | 1358.7601 | 1056.1209 | 1392.559 | 1153.1996 | 823.3585 | 674.1991 | 847.4798 | 1137.7207 | 965.1612 | 846.2807 | 435.3608 | 2829.4798 | 1822.9611 | 1567.4395 | 1716.9587 | 1528.7193 | 1010.8796 | 839.12 | 848.7601 | 423.5998 | 338.12 | 268.0797 | 676.8004 | 1256.4386 | 1012.441 | 922.2798 | 597.4794 | 1397.082 | 1009.6795 | 785.56 | 647.7994 | 1387.4794 | 1113.0394 | 914.3599 | 735.1209 | 717.8781 | 610.2404 | 480.2001 | 404.4808 | 1362.6017 | 1081.9606 | 982.6411 | 894.4002 | 1429.5216 | 1105.8018 | 1148.7198 | 1091.2001 | 1962.0792 | 1207.881 | 1055.1588 | 1271.6411 | 1833.9199 | 1179.0801 | 887.4391 | 1344.3594 | 913.2404 | 731.1996 | 590.5609 | 839.8805 | 607.9602 | 517.9208 | 512.2807 | 454.2409 | 1337.7188 | 720.44 | 560.6795 | 924.8004 | 3182.2399 | 1816.8815 | 1673.201 | 2142.5595 | 3351.3566 | 1959.8004 | 1725.4808 | 2440.2409 | 349.3205 | 247 | 200.4798 | 349.3205 | 482.7999 | 453.8009 | 343.2399 | 429.3205 | 462.2001 | 425.8402 | 339.9991 | 401.6401 | 301.2001 | 278.88 | 239.3599 | 352.8393 | 339.2802 | 285.2793 | 257.56 | 266.2807 | 1756.2418 | 1023.2001 | 864.3585 | 695.6406 | 1019.12 | 832.3589 | 746.12 | 528.5197 | 2798 | 1493.761 | 1364.3599 | 1376.7193
373 | 6.0974 | 837.9351 | 1209.4395 | 1041.4691 | 1013.7581 | 1312.2596 | 1086.3805 | 841.0709 | 829.1504 | 623.295 | 491.5782 | 404.1623 | 499.8673 | 1435.8673 | 827.7581 | 735.8496 | 997.5605 | 1282.7109 | 791.5073 | 688.1445 | 768.4632 | 1409.3009 | 1503.4218 | 1359.5782 | 908.9823 | 2080.0442 | 1385.6136 | 1086.5045 | 1369.2182 | 1128.2359 | 821.23 | 680.2655 | 864.1623 | 1142.7463 | 958.7463 | 829.4277 | 433.7404 | 2799.8082 | 1827.3304 | 1532.2182 | 1716.8141 | 1518.6755 | 1031.3009 | 835.9941 | 846.4277 | 413.6932 | 336.7109 | 276.5723 | 680.0177 | 1282.705 | 995.1504 | 875.351 | 627.4513 | 1407.2272 | 1024.8614 | 760.82 | 648.9823 | 1400.2832 | 1123.9941 | 918.1445 | 749.4809 | 716.9351 | 605.5782 | 463.6932 | 400.8791 | 1293.2419 | 1060.1327 | 966.5841 | 912.3186 | 1427.3363 | 1073.1564 | 1110.5073 | 1100.4454 | 2029.0709 | 1204.4573 | 1063.6873 | 1269.3186 | 1877.059 | 1197.7463 | 855.7787 | 1347.6991 | 893.1268 | 702.8023 | 565.5664 | 844.0236 | 599.9941 | 502.4336 | 502.8673 | 443.7286 | 1334.0973 | 728.0118 | 577.5782 | 888.5309 | 3180.5664 | 1787.0059 | 1678.1859 | 2096.2005 | 3342.2978 | 1989.9145 | 1669.23 | 2371.9291 | 349.0177 | 245.8555 | 194.2714 | 349.0177 | 485.5723 | 442.5841 | 348.4336 | 441.6077 | 456.5664 | 413.5605 | 381.4573 | 397.8555 | 310.4454 | 270.5605 | 217.9646 | 352.2596 | 324.2714 | 282.5427 | 256.8555 | 279.1859 | 1763.2154 | 993.5309 | 883.4041 | 664.8318 | 1042.3186 | 831.82 | 730.9764 | 509.6755 | 2798 | 1487.1682 | 1361.2773 | 1368.3923
374 | 6.1084 | 857.6107 | 1156.5437 | 1017.5732 | 1000.9262 | 1328.8127 | 1086.5517 | 808.0216 | 834.7537 | 622.7389 | 479.8014 | 412.2463 | 488.0477 | 1400.6651 | 826.0329 | 769.412 | 961.0182 | 1311.9818 | 774.1986 | 687.2463 | 837.4416 | 1422.8456 | 1497.6617 | 1326.7571 | 953.7798 | 1982.6504 | 1429.798 | 1106.8785 | 1387.6107 | 1153.857 | 811.706 | 666.4926 | 854.1249 | 1156.6617 | 967.4007 | 846.5369 | 449.9228 | 2831.1181 | 1873.3825 | 1498 | 1662.849 | 1556.8717 | 1006.849 | 850.5369 | 855.9376 | 438.6583 | 337.3383 | 271.1544 | 668.5403 | 1297.2906 | 1028.2282 | 914.3315 | 609.1873 | 1449.4154 | 1020.5698 | 773.5664 | 653.445 | 1369.7718 | 1088.849 | 913.7389 | 781.8456 | 710.3053 | 630.0443 | 473.4597 | 413.1691 | 1291.4597 | 1046.2316 | 954.294 | 941.8604 | 1474.0443 | 1077.0772 | 1138.21 | 1103.1544 | 2035.1873 | 1228.9376 | 1028.3712 | 1291.2611 | 1918.6913 | 1209.9081 | 840.5517 | 1349.353 | 896.1067 | 656.294 | 585.798 | 841.3859 | 599.3383 | 522.5369 | 523.1986 | 425.2792 | 1295.8014 | 724.0624 | 600.8751 | 910.0885 | 3174.4926 | 1829.7503 | 1716.9523 | 2137.1771 | 3269.5698 | 2023.6765 | 1730.4382 | 2324.3087 | 376.706 | 235.0624 | 188.7537 | 376.706 | 481.9081 | 455.4302 | 325.8638 | 475.7832 | 444.0624 | 411.5074 | 362.0953 | 410.445 | 304.3859 | 271.4302 | 233.9047 | 327.8933 | 322.8456 | 265.0772 | 246.647 | 283.4631 | 1816.3678 | 1007.4892 | 851.1249 | 681.2282 | 1011.6504 | 876.1328 | 727.7684 | 542.6107 | 2820.798 | 1509.353 | 1334.7866 | 1350.5846
375 | 6.1194 | 853.2202 | 1190.8005 | 1036.7326 | 964.0629 | 1341.0472 | 1077.1416 | 818.6854 | 800.1887 | 624.3717 | 479.7434 | 384.7384 | 501.0679 | 1367.2252 | 799.4503 | 796.5969 | 965.8584 | 1271.9272 | 802.3088 | 686 | 805.9586 | 1418.644 | 1502.5654 | 1306.9685 | 953.8013 | 1952.5811 | 1364.231 | 1111.3088 | 1385.6076 | 1197.7591 | 810.6755 | 665.1937 | 817.3195 | 1196.1829 | 974.5339 | 819.6076 | 427.529 | 2899.0521 | 1855.3825 | 1495.0157 | 1682.9735 | 1568.7541 | 1004.293 | 856.4346 | 858.2252 | 452.6598 | 321.6912 | 281.9214 | 634.3038 | 1306.822 | 1047.2094 | 912.1051 | 578.4503 | 1438.8535 | 1016 | 781.6755 | 675.5339 | 1330.4975 | 1100.6175 | 911.7906 | 769.2881 | 737.9528 | 605.0108 | 503.5182 | 402.6598 | 1350.1672 | 1020.9007 | 993.5232 | 928.3195 | 1489.5969 | 1109.0207 | 1202.4868 | 1060.8013 | 2011.0157 | 1193.6234 | 1048.2144 | 1265.1573 | 1896.5604 | 1192.6755 | 893.2144 | 1341.0786 | 892.6598 | 661.5025 | 564.995 | 820.2566 | 614.7276 | 534.4031 | 508.3353 | 420.7276 | 1285.5969 | 728.9371 | 609.0315 | 925.6755 | 3210.1987 | 1844.2724 | 1715.8692 | 2118.8121 | 3324.6854 | 2009.0944 | 1718.7964 | 2362.5546 | 354.7856 | 223.2252 | 213.2773 | 354.7856 | 485.5654 | 465.3717 | 327.0993 | 485.8063 | 442.9685 | 414 | 358.293 | 390.1573 | 279.0786 | 258.6912 | 258.5969 | 339.6962 | 341.3245 | 273.1465 | 234.6283 | 268.0157 | 1765.9272 | 1035.5811 | 860.2773 | 708.5654 | 965.4661 | 888.7699 | 753.6962 | 532.2517 | 2860.2773 | 1527.9371 | 1316.0157 | 1345.6283
376 | 6.1304 | 887.1328 | 1209.1062 | 1038.1062 | 956.6091 | 1347.792 | 1062.0457 | 850.472 | 805.8481 | 656.6298 | 483.3451 | 379.792 | 485.1062 | 1399.3304 | 813.1018 | 758.0147 | 1016.8938 | 1287.2597 | 799.3805 | 683.5634 | 786.8024 | 1419.7006 | 1564.2597 | 1334.7567 | 903.6342 | 1981.1932 | 1393.6712 | 1089.497 | 1405.9042 | 1188.9897 | 808.3097 | 653.208 | 800.7817 | 1193.5427 | 964.1165 | 785.0811 | 387.8525 | 2946.0207 | 1798.497 | 1510.6652 | 1709.5634 | 1518.5427 | 1064.295 | 859.3097 | 839.1622 | 423.6799 | 333.1475 | 285.2994 | 621.6091 | 1292.4262 | 1036.2537 | 944.8791 | 582.3555 | 1410.6799 | 1072.0414 | 742.1519 | 644.4056 | 1368.5591 | 1087.4513 | 899.5988 | 727.5427 | 740.3451 | 615.767 | 478.6696 | 402.3097 | 1325.1769 | 1009.2183 | 997.3244 | 893.5074 | 1437.6135 | 1096.6696 | 1132.3288 | 1065.4573 | 1948.6946 | 1221.0414 | 1062.8628 | 1271.1932 | 1863.208 | 1211.0207 | 881.6696 | 1327.3451 | 903.2744 | 692.8835 | 566.1932 | 827.6195 | 638.0103 | 526.0811 | 495.264 | 440.3555 | 1295.7463 | 738.8274 | 616.5738 | 923.3097 | 3199.4513 | 1832.5634 | 1680.1976 | 2053.208 | 3307.1312 | 1990.2994 | 1707.5944 | 2372.0708 | 354.1475 | 226.7006 | 208.2891 | 354.1475 | 454.0604 | 447.6799 | 339.3909 | 486.8274 | 454.3097 | 448.1121 | 345.8068 | 360.8628 | 292.1475 | 275.0207 | 235.8525 | 357.264 | 322.9793 | 273.2994 | 248.0561 | 262.3451 | 1730.7006 | 1054.4469 | 888.1829 | 685.5885 | 989.8938 | 890.7567 | 760.9085 | 504.8732 | 2937.5237 | 1544.528 | 1353.5944 | 1323.1165
377 | 6.1414 | 858.0206 | 1155.0639 | 1003.9498 | 989.9361 | 1323.6853 | 1052.1927 | 854.943 | 824.3786 | 654.4926 | 478.5143 | 408.5074 | 476.4287 | 1464.8791 | 812.9498 | 737.9715 | 1013.764 | 1263.7207 | 785.0285 | 708.7218 | 780.3854 | 1401.0069 | 1551.9065 | 1307.7423 | 935.2292 | 2020.9931 | 1392.6922 | 1076.7001 | 1429.1642 | 1166.9498 | 800.5713 | 645.6214 | 819.886 | 1189.5359 | 948.8928 | 837.1721 | 410.2862 | 2877.6204 | 1833.5507 | 1490.9282 | 1656.4208 | 1495.2714 | 1050.4995 | 822.3638 | 850.7503 | 404.4071 | 346.1357 | 260.4926 | 656.8004 | 1307.9931 | 1046.7788 | 959.9282 | 558.5496 | 1424.3432 | 1081.8997 | 732.943 | 633.9145 | 1387.9498 | 1060.2714 | 953.1436 | 754.6077 | 724.9498 | 602.7138 | 457.2429 | 399.5428 | 1286.8928 | 1036.7218 | 979.4857 | 912.0718 | 1433.8289 | 1087.0502 | 1156.9224 | 1077.5428 | 1928.6431 | 1224.0639 | 1053.8928 | 1268.114 | 1817.0924 | 1186.7423 | 866.4572 | 1305.114 | 952.2862 | 678.8997 | 567.4641 | 764.292 | 635.8141 | 535.9145 | 498 | 425.2497 | 1314.4287 | 735.6499 | 601.0069 | 932.9715 | 3144.3067 | 1804.8997 | 1709.6077 | 2082.9076 | 3270.6214 | 1994.7001 | 1641.6489 | 2239.9271 | 354.707 | 216.0855 | 211.7788 | 354.707 | 457.1003 | 435.8643 | 317.8712 | 464.3854 | 455.2714 | 445.1425 | 346.4002 | 376.1288 | 308.8643 | 276.8427 | 221 | 338.8712 | 300.6499 | 285.1573 | 272.0502 | 257.5143 | 1824.4868 | 1063.7571 | 875.4926 | 690.9931 | 1049.529 | 827.699 | 753.2714 | 512.9715 | 2928.528 | 1489.0992 | 1339.2281 | 1333.3717
378 | 6.1524 | 848.3201 | 1178.0935 | 999.8024 | 966.2728 | 1288.1568 | 1055.7734 | 816.6391 | 832.2385 | 656.5757 | 474.4651 | 382.3426 | 487.1686 | 1411.1504 | 829.7154 | 768.884 | 967.3136 | 1236.5059 | 802.2793 | 691.4125 | 797.0526 | 1420.5875 | 1543.5467 | 1308.6165 | 912.0344 | 2025.4941 | 1394.4887 | 1064.8894 | 1388.8775 | 1210.3319 | 776.5231 | 671.3491 | 832.6337 | 1200.2675 | 959.6745 | 818.7669 | 420.9592 | 2831.971 | 1869.2675 | 1511.2911 | 1663.6574 | 1487.9302 | 1042.3781 | 843.2621 | 848.1568 | 397.1976 | 348.1686 | 258.1396 | 664.6917 | 1285.2438 | 1020.8775 | 959.3781 | 554.5757 | 1470.7852 | 1046.3545 | 744.0698 | 631.5929 | 1414.855 | 1085.2503 | 954.2438 | 753.8485 | 729.6745 | 605.9828 | 466.8722 | 401.3373 | 1302.7444 | 1037.7905 | 960.087 | 898.0462 | 1439.9592 | 1048.7379 | 1165.0752 | 1055.2556 | 1946.3373 | 1229.3491 | 1043.1278 | 1223.0462 | 1868.7798 | 1181.2793 | 852.6219 | 1327.855 | 934.4415 | 655.3953 | 551.029 | 746.6165 | 627.0988 | 573.5177 | 470.7497 | 408.8314 | 1304.2556 | 753.913 | 599.6745 | 927.7615 | 3152.2503 | 1763.6509 | 1694.2728 | 2035.0225 | 3290.0118 | 1979.087 | 1669.0183 | 2218.0645 | 350.4361 | 212.7035 | 217.5639 | 350.4361 | 471.7325 | 443.3373 | 336.0526 | 466.4769 | 471.3781 | 424.9302 | 356.3663 | 392.1686 | 312.5349 | 280.942 | 219.7325 | 341.8432 | 298.6337 | 288.6627 | 248.9474 | 250.9302 | 1819.4296 | 1045.2556 | 842.087 | 689.7905 | 1062.2266 | 821.1923 | 750.3663 | 526.1396 | 2791.2664 | 1519.1751 | 1365.0645 | 1314.0462
379 | 6.1634 | 868.9381 | 1204.292 | 1029.7788 | 954.5664 | 1322.2832 | 1079.4425 | 827.177 | 816.3274 | 665.646 | 522.0974 | 416.2036 | 474.7876 | 1374.6903 | 829.1416 | 759.8053 | 1016.8319 | 1263.3186 | 798.4956 | 672.646 | 817.5841 | 1432.2478 | 1539.2035 | 1315.6018 | 967.9204 | 2078.7877 | 1445.3629 | 1063.3982 | 1380.2743 | 1133.407 | 774.1062 | 640.7787 | 831.708 | 1208.7522 | 986.3806 | 778.9557 | 414.646 | 2905.77 | 1891.3186 | 1518.4956 | 1644.0088 | 1472.0796 | 1077.5487 | 868.0619 | 857.7345 | 398.8761 | 351.9381 | 264.646 | 656.354 | 1276.9204 | 987.0796 | 948.6814 | 563.646 | 1432.4424 | 1031.7699 | 764.4425 | 639.8584 | 1414.9734 | 1074.5132 | 911.6814 | 698.2389 | 735.708 | 613.2478 | 469.4159 | 411.4602 | 1333.6726 | 1003.8053 | 965.4425 | 931.9735 | 1470.4691 | 1047.1947 | 1145.8761 | 1043.0619 | 1963.5664 | 1229.1416 | 1069.6549 | 1262.1417 | 1907.4779 | 1160.6991 | 853.6283 | 1312.469 | 929.7345 | 700.0974 | 576.7168 | 776.2124 | 638.9204 | 548.0088 | 495.6992 | 394.0796 | 1341.1593 | 744.0354 | 625.0885 | 908.7876 | 3236.4779 | 1840.7346 | 1649.6194 | 2020.4868 | 3275.1593 | 1937.9292 | 1692.4513 | 2206.9469 | 318.115 | 232.7965 | 215.646 | 318.115 | 454.2035 | 456.6903 | 356.5841 | 466.3097 | 474.8938 | 427.8142 | 348.2478 | 404.3363 | 299.1416 | 278.0177 | 219 | 352.292 | 309.3363 | 270.1416 | 236.8142 | 253.8142 | 1794.6903 | 1039.5221 | 869.4071 | 667.4336 | 1040.4956 | 852.7522 | 753.2301 | 523.6018 | 2758.0708 | 1516.6548 | 1386.8938 | 1276.9115
380 | 6.1744 | 877.2414 | 1224.7493 | 1028.5423 | 937.6426 | 1332.8338 | 1116.649 | 863.4577 | 786.5673 | 682.4577 | 491.7271 | 403.5516 | 512.4234 | 1441.8811 | 816.3918 | 751.2665 | 1038.8589 | 1315.8152 | 824.2321 | 649.9592 | 803.2006 | 1413.5172 | 1551.0659 | 1276.7679 | 955.7522 | 2118.2414 | 1376.4198 | 1085.7744 | 1334.3352 | 1142.1476 | 783.9248 | 666.0315 | 848.116 | 1250.5236 | 956.1848 | 800.9907 | 415 | 2861.053 | 1908.2665 | 1527.7744 | 1631.533 | 1527.5394 | 1081.1254 | 836.7178 | 837.0595 | 404.9499 | 331.9341 | 265.6583 | 654.0251 | 1302.4327 | 993.8245 | 892.8933 | 573.2163 | 1403.9749 | 995.1096 | 741.0845 | 669.0659 | 1377.6676 | 1038.2758 | 948.7149 | 739.5394 | 754.091 | 597.8087 | 467.3417 | 423.558 | 1325.909 | 1018.3073 | 964.467 | 972.8496 | 1452.1597 | 1102.4735 | 1094.6769 | 1093.3481 | 1959.8087 | 1214.4169 | 1061.3259 | 1257.1347 | 1866.1597 | 1208.8309 | 838.2758 | 1280.4921 | 903.1347 | 744.7744 | 568.6426 | 789.2665 | 653.2414 | 545.4076 | 531.1662 | 412.6741 | 1333.4262 | 733.9749 | 616.5673 | 940.4986 | 3256.2256 | 1856.1096 | 1618.5924 | 2001.4262 | 3297.8653 | 1977.9563 | 1676.3668 | 2167.884 | 322.0408 | 238.0501 | 214.6834 | 322.0408 | 473.3073 | 462.6583 | 336.2758 | 459.6834 | 476.2665 | 431 | 340.7085 | 384.3259 | 285.0752 | 283.8496 | 211.7586 | 359.5831 | 332.7744 | 294.2572 | 238.6834 | 293.207 | 1822.3825 | 1044.6332 | 837.6769 | 684.3574 | 1018.8338 | 866.2163 | 710.235 | 534.7994 | 2770.0501 | 1518.3574 | 1319.144 | 1269.6332
381 | 6.1854 | 865.5762 | 1200.7994 | 1026.0354 | 928.353 | 1352.824 | 1092.7758 | 871.3294 | 781.6942 | 654.117 | 506.9302 | 406.4828 | 488.3874 | 1467.9646 | 805.9528 | 758.6942 | 974.6696 | 1339.0118 | 848.0354 | 629.2822 | 815.118 | 1405.3294 | 1502.9872 | 1305.2714 | 948.471 | 2079.0816 | 1378.942 | 1072.1996 | 1366.6834 | 1182.3648 | 807.118 | 662.4936 | 820.117 | 1253.235 | 1010.7779 | 801.588 | 425.7296 | 2920.6019 | 1832.5397 | 1539.0118 | 1598.7758 | 1555.9764 | 1057.2232 | 852.5182 | 829.0472 | 394.9292 | 331.059 | 272.0354 | 669.0944 | 1297.9174 | 1017.0944 | 916.2714 | 579.3412 | 1398.9646 | 1007.177 | 740.765 | 647.1406 | 1416.9774 | 1053.1534 | 932.4228 | 706.9518 | 724.4346 | 602.059 | 467.6706 | 429.3412 | 1351.5418 | 1016.9056 | 1004.6008 | 923.3402 | 1489.6716 | 1077.058 | 1130.707 | 1085.7994 | 2041.1781 | 1200.941 | 1053.3648 | 1264.1534 | 1866.118 | 1233.6116 | 859.8594 | 1278.6942 | 868.5762 | 691.6342 | 596.236 | 805.412 | 633.529 | 517.1052 | 489.0462 | 408.5644 | 1318.6824 | 758.471 | 637.8476 | 938.1996 | 3299.9774 | 1860.8004 | 1659.6362 | 2002.1062 | 3309.6234 | 1937.9281 | 1633.4228 | 2214.6598 | 327.9646 | 225.2704 | 189.8584 | 327.9646 | 465.1996 | 463.6706 | 355.8476 | 477.7768 | 456.8702 | 414.9056 | 349.7296 | 385.7532 | 305.471 | 249.0934 | 216.7178 | 363 | 323.8938 | 314.353 | 237.3294 | 263.7168 | 1818.5762 | 1048.6824 | 849.5654 | 684.9292 | 1069.6716 | 871 | 713.1534 | 536.294 | 2833.0482 | 1487.411 | 1307.4592 | 1294.471
382 | 6.1964 | 883.9499 | 1170.3422 | 1018.0737 | 980.534 | 1290.199 | 1079.4631 | 816.3687 | 792.1947 | 640.0973 | 503.733 | 392.3672 | 493.6814 | 1459.5369 | 784.2448 | 770.5605 | 979.0959 | 1302.441 | 807.9778 | 693.972 | 810.6593 | 1412.5118 | 1513.6077 | 1310.9277 | 949.0737 | 2031.3672 | 1409.5118 | 1081.4867 | 1344.1002 | 1185.6829 | 799.2448 | 623.4159 | 829.9986 | 1214.2699 | 987.9055 | 778.6106 | 416.6593 | 2945.733 | 1782.3908 | 1524.2935 | 1647.6063 | 1605.8998 | 1048.6829 | 790.8333 | 827.5369 | 371.2448 | 344.8776 | 272.9513 | 615.5398 | 1288.2684 | 999.733 | 936.9513 | 560.879 | 1455.4116 | 1048.6328 | 727.1475 | 618.0251 | 1407.4896 | 1089.9012 | 931.8039 | 735.7773 | 720.0236 | 628.1696 | 476.8776 | 440.2434 | 1377.8776 | 1069.7773 | 996.9528 | 902.8289 | 1484.5383 | 1037.6106 | 1177.7065 | 1051.9277 | 1994.5412 | 1200.0973 | 1042.3422 | 1249.0014 | 1866.4395 | 1252.2183 | 873.2684 | 1309.6814 | 927.9234 | 688.9986 | 566.1489 | 789.4646 | 630.1947 | 536.9248 | 499.1446 | 408.8776 | 1333.6578 | 754.9764 | 621.587 | 905.1475 | 3290.4896 | 1895.9012 | 1668.6106 | 2047.1933 | 3244.5398 | 1930.6328 | 1657.388 | 2224.2448 | 320.9027 | 213.1711 | 197.8039 | 320.9027 | 491.559 | 433.2699 | 317.7345 | 460.3672 | 456.9263 | 391.2935 | 345.4395 | 383.4395 | 296.5133 | 247.4381 | 229.8776 | 326.8067 | 323.1947 | 281.8554 | 240.4145 | 252.9749 | 1782.3186 | 1039.0737 | 835.5383 | 714.5104 | 1054.2949 | 851.1961 | 727.3658 | 517.2935 | 2805.5884 | 1515.1682 | 1348.0472 | 1331.9499
383 | 6.2074 | 910.5988 | 1168.8663 | 1001.1819 | 976.8879 | 1262.647 | 1060.4493 | 788.2193 | 809.2085 | 658.6844 | 489.9145 | 375.8289 | 504.2193 | 1435.4493 | 765.5723 | 770.8289 | 936.995 | 1313.5025 | 824.1495 | 718.353 | 820.6844 | 1431.9892 | 1538.9518 | 1323.8555 | 876.5673 | 2081.5664 | 1440.8073 | 1068.7542 | 1335.294 | 1197.8181 | 827.8447 | 628.6844 | 819.3638 | 1241.711 | 984.3314 | 761.9626 | 426.6844 | 2941.647 | 1849.9942 | 1475.9842 | 1632.4228 | 1592.4601 | 1057.3422 | 816.1761 | 875.0531 | 396.3688 | 337.1819 | 252.5349 | 653.7375 | 1277.9626 | 968.5349 | 923.4867 | 588.8447 | 1449.6312 | 1037.3638 | 734.3796 | 623.8181 | 1439.6628 | 1065.5457 | 927.0963 | 758.4385 | 711.0108 | 624.7915 | 455.2782 | 433.1819 | 1372.2674 | 1001.3696 | 959.3638 | 910.8663 | 1463.4385 | 1073.7965 | 1155.8023 | 1005.0698 | 1943.2674 | 1188.0963 | 1032.5241 | 1224.8771 | 1839.059 | 1266.4759 | 835.8505 | 1301.4493 | 917.289 | 664.4601 | 555.5133 | 777.6096 | 643.7326 | 527.8879 | 500.4012 | 400.4867 | 1324.0108 | 772.3314 | 630.6362 | 899.8663 | 3313.6254 | 1885.4493 | 1711.0531 | 2032.7168 | 3339.0473 | 1926.3156 | 1677.3048 | 2159.6902 | 321.7807 | 191.9253 | 196.5723 | 321.7807 | 493.706 | 428.7326 | 311.7699 | 484.8447 | 466.8663 | 402.7699 | 343.0855 | 379 | 308.5507 | 260.1711 | 231.9145 | 303.7434 | 308.9253 | 263.9145 | 235.0482 | 251.353 | 1805.8447 | 1032.6096 | 897.1653 | 679.5565 | 1081.9676 | 868.4169 | 730.7807 | 521.1229 | 2733.5083 | 1522.0108 | 1304.2998 | 1307.1553
384 | 6.2184 | 920.5374 | 1203.5438 | 1014.3943 | 934.5786 | 1304.9111 | 1084.9587 | 774.9729 | 799.4356 | 623.5516 | 465.6534 | 395.9317 | 495.9253 | 1430.9523 | 768.7822 | 781.0271 | 957.2783 | 1301.9459 | 844.2448 | 670.4292 | 823.7552 | 1448.1972 | 1515.116 | 1309.6534 | 901.4755 | 2104.0477 | 1413.3815 | 1072.442 | 1361.8093 | 1209.3673 | 844 | 648.0271 | 783.5309 | 1216.8441 | 1007.0271 | 801.1559 | 430.4626 | 2911.7487 | 1832.4768 | 1457.2925 | 1627.5644 | 1543.4562 | 1056.7552 | 829.558 | 881.4356 | 384.116 | 336.9523 | 260.9794 | 654.7011 | 1288.1495 | 1009.5232 | 914.4626 | 570.3338 | 1444.9046 | 976.7694 | 720.0683 | 638.1972 | 1405.8171 | 1030.1907 | 923.8299 | 760.9523 | 693.3879 | 611.9253 | 460.9794 | 420.9253 | 1365.1701 | 986.5915 | 963.8569 | 910.8776 | 1473.7345 | 1093 | 1164.9317 | 1001.442 | 1970.1288 | 1216.6662 | 1009.7758 | 1221.4149 | 1887.0128 | 1222.0142 | 878.7203 | 1248.1366 | 944.2783 | 680.5438 | 572.3196 | 784.0747 | 643.7552 | 511.0477 | 485.5103 | 419.0541 | 1361.5708 | 764.6057 | 621.1907 | 892.0954 | 3259.5452 | 1825.0619 | 1684.8918 | 1979.3815 | 3277.6818 | 1955.7886 | 1629.5993 | 2060.7153 | 316.6327 | 218.6662 | 212.5167 | 316.6327 | 460.8235 | 442.9046 | 316.4626 | 478.3608 | 473.9523 | 429.3943 | 346.8299 | 376.8776 | 312.3402 | 262 | 235.2448 | 343.4484 | 322.2242 | 259.4626 | 239.7822 | 255.7822 | 1838.9794 | 1017.1431 | 882.8918 | 727.0399 | 1027.3003 | 837.5516 | 768.7886 | 524.5851 | 2810.2164 | 1473.9665 | 1315.7886 | 1291.7075
385 | 6.2294 | 903.2861 | 1166.6165 | 996.8082 | 962.3068 | 1288.6106 | 1083.1652 | 793.7522 | 834.0266 | 660.7021 | 448.8024 | 410.7581 | 500.9174 | 1447.3953 | 805.1092 | 755.7699 | 944.3304 | 1335.1859 | 807.1327 | 677.9115 | 850.7906 | 1439.764 | 1514.0767 | 1302.1593 | 948.472 | 2113.5162 | 1422.9115 | 1072.2419 | 1346.2478 | 1165.9351 | 810.8908 | 628.528 | 786.6755 | 1252.9824 | 983.9292 | 794.528 | 401.6489 | 2899.4395 | 1799.764 | 1539.7729 | 1649.1151 | 1556.6313 | 1072.9941 | 824 | 797.2654 | 406.8289 | 320.2861 | 245.6873 | 633.6431 | 1331.4278 | 995.4513 | 919.4779 | 571.8348 | 1435.3245 | 984.7906 | 716.7581 | 632.6431 | 1462.8939 | 1058.708 | 922.8407 | 758.6814 | 701.236 | 596.0442 | 496.7906 | 425.1976 | 1364 | 1025.3511 | 958.4838 | 932.3127 | 1479.7198 | 1125.3894 | 1150.9675 | 1025.7139 | 1932.6165 | 1213.7256 | 1001 | 1204.7256 | 1877.8908 | 1213.7965 | 867.0118 | 1273.9056 | 938.528 | 696.1593 | 566.2035 | 752.4513 | 629.0442 | 503.9616 | 502.1534 | 414.6047 | 1328.8967 | 749.0826 | 595.7256 | 929.0266 | 3265.2685 | 1811.9174 | 1662.8407 | 1942.8466 | 3250.1534 | 1955.1652 | 1637.3511 | 2119.2508 | 320.4779 | 190.5339 | 203.7256 | 320.4779 | 460.7965 | 405.9734 | 311.4012 | 475.4779 | 459.4454 | 410.3687 | 350.1593 | 388.9558 | 319.3569 | 268.4779 | 220.4454 | 357.4012 | 323.8024 | 272.3953 | 237.9616 | 288.5104 | 1816.4896 | 1026.8348 | 851.4838 | 735.1268 | 1030.5487 | 835.1151 | 766.0059 | 524.7198 | 2898.944 | 1492.2685 | 1338.1976 | 1246.6548
386 | 6.2404 | 872.5742 | 1231.9185 | 987 | 967.0192 | 1328.708 | 1055.7704 | 815.4449 | 849.2679 | 654.9139 | 473.8181 | 399.823 | 470.3254 | 1500.5836 | 771.8805 | 730.0909 | 941.5167 | 1335.1627 | 814.6937 | 643.3446 | 842.2345 | 1441.5885 | 1464.5599 | 1309.5885 | 908.9523 | 2128.2488 | 1424.2153 | 1070 | 1315.7704 | 1158.9808 | 822.1578 | 645.354 | 776.6268 | 1225.6843 | 1026.4401 | 819.4066 | 391 | 2903.6603 | 1857.2247 | 1495.1341 | 1665.2488 | 1536.2537 | 1088.7847 | 807.8948 | 793.354 | 419.2679 | 323.2488 | 243.5885 | 599.9857 | 1342.3397 | 950.5216 | 922.7321 | 572.1436 | 1473.1913 | 1004.0526 | 741.6937 | 634.1244 | 1486.6794 | 1036.3972 | 953.4784 | 743.0909 | 754.3157 | 560.2537 | 499.9474 | 425.0718 | 1348.6268 | 1034.5359 | 936.4307 | 895.6125 | 1486.5885 | 1075.0432 | 1122.2345 | 1007.378 | 1945.9425 | 1274.3491 | 996.6077 | 1235.3348 | 1902.3348 | 1249.4784 | 837.6268 | 1309.3014 | 943.6411 | 691.8756 | 599.3348 | 758.7655 | 638.6411 | 518.1052 | 518.0526 | 431.6937 | 1326.5693 | 724.0383 | 583.1436 | 901.8087 | 3293.6411 | 1899.9185 | 1708.8516 | 1924.7512 | 3258 | 1914.5934 | 1660.4449 | 2148.8756 | 341.3014 | 188.6411 | 188.2153 | 341.3014 | 462.8038 | 438.3157 | 309.2679 | 468.4833 | 433.9666 | 402.1962 | 340.7512 | 388.1436 | 303.2345 | 293.6937 | 220.5885 | 341.3589 | 332.7129 | 265.5551 | 204.5216 | 275.1101 | 1803.5359 | 1076.1913 | 902.6362 | 716.0192 | 1046.9282 | 819.0383 | 769.2488 | 525.7321 | 2801.6031 | 1500.8756 | 1309.5216 | 1278.7798
387 | 6.2514 | 910.1495 | 1229.5044 | 1033.8938 | 1013.6156 | 1323.4366 | 1097.8712 | 766.407 | 828.9026 | 604.8053 | 486.233 | 394.2557 | 478.8417 | 1471.1062 | 798.1721 | 714.5791 | 933.8348 | 1317.0905 | 821.5113 | 664.9617 | 809.6922 | 1499.0817 | 1476.2626 | 1324.6539 | 899.9322 | 2103.7148 | 1373.3618 | 1066.2783 | 1333.3078 | 1179.3756 | 848.12 | 634.9026 | 769.5113 | 1278.7129 | 1003.5722 | 814.88 | 401.4209 | 2879.6922 | 1799.6105 | 1426.0609 | 1631.0383 | 1600.9234 | 1068.1809 | 781.9026 | 845.4278 | 410.8122 | 329.233 | 277.2626 | 618.7739 | 1352.9095 | 1017.9008 | 905.1357 | 541.7148 | 1458.4818 | 998.0678 | 744.0452 | 615.1583 | 1437.1062 | 1080.3147 | 937.4592 | 726.8348 | 798.7965 | 598.8712 | 480.6244 | 416.5565 | 1328.1131 | 1014.647 | 944.8869 | 893.1426 | 1514.3078 | 1063.1652 | 1124.6765 | 1005.4435 | 1964.6991 | 1270.7148 | 1003.9322 | 1234.1131 | 1878.0383 | 1288.5408 | 868.473 | 1261.6627 | 972.0748 | 671.3913 | 568.3392 | 780.1426 | 627.6244 | 537.3982 | 465.174 | 395.3392 | 1321.8348 | 727.1652 | 641.2921 | 868.88 | 3262.527 | 1865.4975 | 1697.7148 | 2009.5773 | 3231.2035 | 1945.9165 | 1647.88 | 2161.1426 | 343.5339 | 205.1652 | 189.4661 | 343.5339 | 447.1131 | 424.7374 | 332.0748 | 459.0452 | 447.8417 | 409.6991 | 319.88 | 395.6765 | 294.5113 | 294.5565 | 229.6991 | 321.8574 | 333.2783 | 279.6087 | 245.1043 | 258.5565 | 1843.1947 | 1151.5477 | 861.9636 | 705.3009 | 1033.1131 | 880.2243 | 769.2783 | 539.3982 | 2710.6175 | 1500.4887 | 1239.1966 | 1279.88
388 | 6.2624 | 930.5398 | 1174.8716 | 1018.2212 | 979.5486 | 1355.1018 | 1052.1991 | 824.4204 | 820.4867 | 645.4779 | 459.0044 | 387.9469 | 512.4823 | 1495.1018 | 818.2965 | 730.6726 | 961.7788 | 1320.5664 | 785.438 | 685.5664 | 809.323 | 1496.0575 | 1501.3761 | 1302.5177 | 923.4292 | 2127.2921 | 1354.7301 | 1071.8097 | 1365.4558 | 1161.5442 | 853.2434 | 657.5089 | 793.2124 | 1251.5486 | 986.7301 | 774.1947 | 422.4027 | 2982.9824 | 1796.2124 | 1423.3496 | 1563.5221 | 1594.4646 | 1084.2124 | 818.1284 | 810.8185 | 421.6195 | 325.4602 | 257.7345 | 659.2655 | 1297.9822 | 976.9026 | 945.9115 | 556.969 | 1445.7301 | 979.1106 | 732.9203 | 626.9159 | 1447.4823 | 1076.5442 | 921.1903 | 736.6195 | 739.146 | 574.3849 | 488.6195 | 391.3009 | 1367.6416 | 1031.4558 | 962.1062 | 876.0575 | 1491.2212 | 1079.1062 | 1137.0797 | 1064.7478 | 1997.2655 | 1216.3584 | 976.7345 | 1193.438 | 1860.4602 | 1254.8716 | 898.4027 | 1256.6195 | 914.1725 | 687.6991 | 566.6195 | 750.9513 | 644.6991 | 553.3496 | 485.3452 | 373.1903 | 1294.5442 | 760.5089 | 595.4159 | 875.8628 | 3194.7654 | 1884.6151 | 1618.3893 | 1971.146 | 3194.0044 | 1995.0487 | 1652.5929 | 2069.907 | 326.1106 | 212.7832 | 181.1637 | 326.1106 | 446.5398 | 423.8363 | 342.2699 | 463.8097 | 436.8407 | 399.8938 | 338.2124 | 376.3009 | 287.9469 | 278.3805 | 210.0575 | 343.4823 | 363.7788 | 289.7832 | 247.8672 | 280.969 | 1806.3053 | 1099.3628 | 873.5354 | 733.2655 | 985.6283 | 846.4955 | 748.3274 | 537.1903 | 2740.9115 | 1454.0885 | 1258.3452 | 1296.6991
389 | 6.2734 | 912.0078 | 1196.3707 | 1013.3825 | 956.8486 | 1387.9922 | 1032.7689 | 830.5457 | 813.0796 | 679.9164 | 472.2989 | 369.8525 | 511.7729 | 1474.1671 | 827.6893 | 748.5339 | 998.1436 | 1336.0718 | 809.3707 | 652.6293 | 808.9243 | 1525.1397 | 1539.8329 | 1275.5457 | 883.0953 | 2057.2625 | 1350.9243 | 1042.4739 | 1360.6971 | 1228.5861 | 845.3107 | 640.0875 | 824.0679 | 1291.9007 | 1001.3786 | 789.1436 | 411.0836 | 2972.8643 | 1847.06 | 1486.2075 | 1563.6854 | 1583.2311 | 1088.9243 | 785.0953 | 821.9125 | 400.6253 | 336.3029 | 254.1514 | 635.1671 | 1295.1475 | 981.1436 | 864.8839 | 570.3825 | 1484.2154 | 1000.9125 | 736.1514 | 649.9164 | 1486.7572 | 1046.7011 | 909.0039 | 785.5979 | 703.9282 | 618.6697 | 479.9282 | 407.8368 | 1379.6932 | 999.0157 | 975.9961 | 917.4425 | 1473.3107 | 1097.6096 | 1146.1514 | 1023.7925 | 2001.6175 | 1251.2115 | 945.47 | 1205.8368 | 1855.9243 | 1305.59 | 880.9321 | 1234.8564 | 922.9882 | 683.4661 | 583.3029 | 746.1514 | 644.3107 | 547.7729 | 491.8486 | 407.9086 | 1310.8368 | 760.7729 | 582.4582 | 891.5339 | 3272.3472 | 1867.2428 | 1615.9922 | 1927.47 | 3268.8133 | 2008.3068 | 1623.705 | 2097.4386 | 336.3786 | 191.7011 | 178 | 336.3786 | 459.5339 | 427.7689 | 335.3107 | 460.6175 | 440.9961 | 418.2989 | 301.4021 | 372.8446 | 302.9164 | 246.3186 | 217.6096 | 322.7807 | 333.2468 | 284.8486 | 236.0796 | 278.235 | 1770.0078 | 1142.1279 | 837.0953 | 779.1397 | 1072.7297 | 819.5418 | 728.9282 | 537.3068 | 2741.3904 | 1448.9961 | 1307.1397 | 1289.3904
390 | 6.2844 | 915.3746 | 1171.0639 | 977.5014 | 961.2498 | 1354.939 | 1015.8132 | 787.5014 | 857.8732 | 686.5624 | 446.97 | 411.092 | 483.2197 | 1486.6554 | 830.7812 | 759.031 | 976.3136 | 1327.5005 | 852.03 | 638.0939 | 793.1568 | 1514.1258 | 1510.939 | 1298.1239 | 947.122 | 2121.2779 | 1447.6525 | 1065.8112 | 1299.1897 | 1208.0329 | 821.1258 | 636.6873 | 808.3446 | 1313.6873 | 1022.4056 | 761.0639 | 406.7812 | 2903.5334 | 1846.7193 | 1500.0939 | 1550.2507 | 1627.5295 | 1055.1888 | 795.2178 | 812.0319 | 401.5934 | 331.4066 | 251.3127 | 642.9681 | 1330.4676 | 995.2498 | 861.7803 | 572 | 1423.3465 | 1026.1868 | 709.8761 | 598.7522 | 1446.0019 | 1069.6864 | 888.8132 | 795.0939 | 753.123 | 575.0648 | 493.9681 | 425.9371 | 1384.4685 | 1015.9051 | 924.3146 | 939.031 | 1528.0291 | 1095.7502 | 1113.6263 | 1020.8432 | 1899.2226 | 1241.001 | 1019.4657 | 1207.5315 | 1842.5005 | 1217.1916 | 879.4685 | 1255.9051 | 937.4685 | 683.1249 | 562.001 | 745.6563 | 644.3437 | 516.0949 | 488.4376 | 388.5324 | 1311.7502 | 745.5005 | 606.0929 | 872.3446 | 3343.9671 | 1910.3417 | 1615.7502 | 1962.4047 | 3207.2846 | 2030.0929 | 1619.4685 | 2038.4404 | 371.4676 | 213.1239 | 190.4995 | 371.4676 | 498.1549 | 441.2807 | 311.1258 | 448.0629 | 458.8432 | 437.4995 | 355.9661 | 370.0939 | 304.0939 | 238 | 231.3746 | 347.5925 | 347.5615 | 272.8441 | 219.938 | 250.7822 | 1789.7803 | 1123.7202 | 877.6854 | 643.1306 | 1091.2817 | 855.2798 | 743.7493 | 560.6554 | 2728.4066 | 1498.092 | 1328.9371 | 1274.0629
391 | 6.2954 | 917.2065 | 1210.9912 | 1003.5015 | 912.2153 | 1328.9233 | 1050.6756 | 810.3275 | 890.8378 | 672.7168 | 491.1652 | 388.292 | 470.4454 | 1476.3363 | 857.1859 | 735.6076 | 944.1947 | 1324 | 811.0088 | 671.121 | 844.5458 | 1477.0531 | 1507.1416 | 1221.8878 | 948.3687 | 2175.3599 | 1402.7905 | 1049.6076 | 1290.9351 | 1198.3805 | 794.3687 | 657.0443 | 769.466 | 1288.8141 | 1015.0973 | 786.708 | 422.8702 | 2958.6225 | 1846.9676 | 1548.9912 | 1552.1416 | 1588.4218 | 1066.6313 | 798.0324 | 795.8908 | 396.8584 | 324.2389 | 246.826 | 616.2596 | 1353.2832 | 970.0206 | 804.7256 | 553.7492 | 1433.9469 | 1014.3363 | 729.7729 | 621.882 | 1433.587 | 1043.879 | 911.7729 | 762.2596 | 759.2714 | 588.1534 | 457.7374 | 426.6195 | 1401.0767 | 971.6283 | 945.5015 | 910.8466 | 1473.3775 | 1045.5958 | 1143.6756 | 1003.5752 | 1921.7847 | 1230.826 | 942.8111 | 1147.2802 | 1918.3511 | 1191.7611 | 884.174 | 1241.7817 | 938.2065 | 672.0973 | 574.0443 | 790.2301 | 652.1416 | 526.2508 | 490.3805 | 414.121 | 1315.5546 | 729.3038 | 654.0561 | 829.4985 | 3401.2301 | 1946.4248 | 1691.7641 | 1952.5752 | 3140.9439 | 2017.7492 | 1645.5988 | 2106.6667 | 350.6401 | 221.7935 | 177.3363 | 350.6401 | 492.1298 | 441.826 | 293.8908 | 474.3599 | 468.5546 | 421.1622 | 336.118 | 400.7404 | 301.413 | 256.2508 | 221.3038 | 344.3038 | 323.2596 | 283.4897 | 256.4691 | 273.3599 | 1873.1771 | 1139.9794 | 890.6195 | 675.7965 | 1034.8348 | 893.7729 | 746.6195 | 521.7699 | 2672.8348 | 1507.826 | 1292.3244 | 1314.6431
392 | 6.3064 | 932.3102 | 1219.971 | 1013.8058 | 969.9106 | 1259.73 | 1057.7768 | 794.8259 | 844.4287 | 659.3304 | 472.7679 | 412.0379 | 493.1741 | 1416.3997 | 828.5447 | 743.5044 | 1000.0468 | 1324 | 784.3014 | 666.3014 | 826.7679 | 1500.4263 | 1543.6495 | 1278.9686 | 962.3393 | 2174.1362 | 1412.2031 | 993.0403 | 1293 | 1194.1652 | 776.1072 | 630.5737 | 764.2232 | 1236.8751 | 973.3215 | 763.7679 | 415.7188 | 2907.701 | 1890.7075 | 1473.3618 | 1498.3997 | 1624.1539 | 1099.3973 | 787.3304 | 850.2119 | 427.2321 | 330.2522 | 237.942 | 632.1741 | 1364.2522 | 988.7857 | 882.2788 | 553.029 | 1505.6584 | 1026.116 | 717.6607 | 648.116 | 1453.3393 | 1069.6495 | 941.5625 | 782.2031 | 790.0379 | 600.8348 | 491.3191 | 412.3014 | 1382.4375 | 998.2901 | 981.5915 | 924.7567 | 1522.6874 | 1148.2297 | 1174.9509 | 1040.7075 | 1932.5826 | 1225.1652 | 978.241 | 1250.4529 | 1912.4085 | 1272.7744 | 857.6027 | 1289.9887 | 943.6406 | 731.8526 | 578.1942 | 818.9218 | 659.6406 | 516.4956 | 501.4754 | 381.0983 | 1317 | 776.7655 | 602.1474 | 865.1249 | 3407.3594 | 1923.7969 | 1674.9331 | 1941.3594 | 3180.7655 | 1992.0491 | 1650.3884 | 2066.3416 | 338.971 | 214.7478 | 182.6696 | 338.971 | 488.8058 | 450.6696 | 311.1451 | 414.3126 | 478.8638 | 427.2812 | 339.6986 | 394.4956 | 276.0201 | 282.7567 | 214.7768 | 349.058 | 319.029 | 287.8058 | 240.4085 | 271.3304 | 1867.2143 | 1122.8259 | 839.2345 | 767.5222 | 1106.415 | 856.6809 | 751.6406 | 577.6874 | 2741.1918 | 1476.3795 | 1341.6294 | 1307.4666
393 | 6.3174 | 923.7286 | 1201.002 | 1043.4514 | 936.2782 | 1371.6225 | 1041.4562 | 783.2733 | 847.5438 | 690.5418 | 469.0905 | 397.0934 | 489.1829 | 1411.1809 | 747.1897 | 735.5467 | 985.5486 | 1290.4582 | 814.5418 | 688.3609 | 851.7237 | 1511.2724 | 1558.5447 | 1323.3609 | 920.3687 | 2122.9144 | 1369.9144 | 1012.0875 | 1336.3589 | 1167.6391 | 827.8122 | 624.6362 | 821.4484 | 1259.5418 | 1016.358 | 759.2724 | 404.001 | 2977.7179 | 1860.0953 | 1485.5418 | 1578.2626 | 1625.3648 | 1110.0905 | 813.6333 | 853.3648 | 369.5525 | 344.2714 | 204.9125 | 606.9125 | 1401.178 | 966.3667 | 881.2753 | 545.001 | 1507.0924 | 1014.4562 | 723.6352 | 654.4543 | 1508.7218 | 1050.1848 | 991.177 | 753.822 | 762.0039 | 537.3706 | 456.1868 | 428.6342 | 1341.0039 | 966.2772 | 911.9173 | 954.5428 | 1533.2733 | 1090.9183 | 1147.2763 | 1051.8181 | 1989.2665 | 1241.998 | 1024.9056 | 1216.8249 | 1928.9066 | 1247.823 | 877.9971 | 1314.4533 | 977.7237 | 726.5477 | 560.82 | 751.8259 | 641.3658 | 517.9086 | 481.0934 | 384.0895 | 1307.1829 | 804.5438 | 578.8191 | 931.6303 | 3429.7247 | 1936.6342 | 1641.4572 | 1965.3609 | 3196.2724 | 1958.3667 | 1600.0963 | 2024.821 | 322.4562 | 223.6352 | 189.9086 | 322.4562 | 477.5467 | 426.821 | 333.18 | 441.3589 | 476.0914 | 414.4562 | 333.1829 | 348.4592 | 309.2685 | 243.823 | 198.4562 | 345.2733 | 335.5438 | 281.4553 | 251.998 | 263.2733 | 1917.4484 | 1069.5506 | 867.0866 | 736.7325 | 1144.18 | 837.8191 | 749.7276 | 533.4611 | 2711.9144 | 1452.6381 | 1310.1868 | 1306.2724
394 | 6.3284 | 905.2227 | 1175.41 | 1034.2227 | 1004.7169 | 1380.0708 | 1072.0458 | 800.2685 | 809.4808 | 673.0885 | 485.7773 | 423.5546 | 470.3923 | 1486.9042 | 774.18 | 703.9366 | 954.0885 | 1348.6004 | 802.0708 | 697.3215 | 803.3643 | 1516.1519 | 1589.0634 | 1319.0354 | 925.7773 | 2152.6888 | 1342.3923 | 1072.975 | 1304.4808 | 1204.3496 | 815.9189 | 627.4912 | 788.3289 | 1308.5192 | 1020.5265 | 754.1873 | 397.0177 | 3039.3496 | 1926.565 | 1458.9542 | 1608.9646 | 1599.7492 | 1064.4985 | 796.7492 | 820.2758 | 402.3319 | 341.1873 | 237.8584 | 624.9115 | 1350.8731 | 947.5442 | 905.8938 | 537.1873 | 1486.5619 | 1024.2861 | 711.8835 | 653.3392 | 1488.4454 | 1049.8127 | 961.9719 | 752.6431 | 732.41 | 560.3673 | 436.8658 | 409.7492 | 1403.5827 | 1036.7169 | 914.9292 | 940.0708 | 1527.8481 | 1078.6431 | 1125.0531 | 1073.59 | 1951.3466 | 1246.8304 | 1003.7669 | 1189.7139 | 1926.5265 | 1255.4381 | 854.9366 | 1317.8304 | 970.8835 | 696.2581 | 580.2508 | 771.7065 | 637 | 549.7242 | 474.3392 | 405.2685 | 1344.8584 | 737.4173 | 589.4558 | 875.9085 | 3427.5265 | 1953.4558 | 1546.3185 | 1986.7773 | 3174.5796 | 1923.7669 | 1594.8127 | 2058.8584 | 346.4027 | 216.0354 | 179.3746 | 346.4027 | 481.6431 | 426.8127 | 321.2227 | 480.8938 | 504.0634 | 412.6608 | 355.9115 | 397.9573 | 295.5796 | 266.385 | 188.3569 | 341.5088 | 327.3746 | 292.4558 | 238.5619 | 280.2685 | 1931.6608 | 1083.0811 | 839.4631 | 755.0634 | 1092.3643 | 891.6357 | 724.0885 | 520 | 2757.4661 | 1463.9469 | 1275.9189 | 1298.6962
395 | 6.3394 | 918.8535 | 1241.7847 | 1025.9439 | 1020.1573 | 1371.7866 | 1073.1013 | 809.0561 | 825.4376 | 727.8299 | 467.9331 | 407.2478 | 522.6166 | 1515.4828 | 790.5841 | 699.6854 | 984.3924 | 1398.235 | 780.3038 | 679.4611 | 840.8752 | 1543.1229 | 1533.4847 | 1306.0452 | 957.651 | 2149.2025 | 1408.942 | 1075.5732 | 1338.9764 | 1217.2134 | 827.8535 | 623.7866 | 789.1121 | 1272.3382 | 1028.2694 | 718.4503 | 438.1338 | 2990.6981 | 1833.2949 | 1471.3815 | 1621.1121 | 1664.0994 | 1067.6401 | 759.9783 | 838.4376 | 401.0452 | 340.8427 | 235.0452 | 599.6637 | 1354.1682 | 945 | 892.6185 | 579.8191 | 1506.5949 | 1010.1465 | 749.4484 | 693.4484 | 1467.6745 | 1069.5389 | 938.8318 | 788.5497 | 783.6166 | 609.2911 | 499.0433 | 410.8987 | 1391.0344 | 1009.1809 | 935.0108 | 914.0904 | 1551.4376 | 1065.6745 | 1135.4828 | 1090.6401 | 1935.9439 | 1285.7631 | 978.6185 | 1229.1338 | 1932.5841 | 1275.0108 | 832.1465 | 1350.0217 | 961.2586 | 666.5624 | 539.4955 | 755.4051 | 656.3815 | 562.7414 | 500.1229 | 432.5949 | 1337.8318 | 793.7847 | 597.8987 | 875.4828 | 3381.3382 | 1913.0236 | 1600.4701 | 1897.3057 | 3212.9764 | 1980.358 | 1568.1917 | 2059.4159 | 345.2586 | 224.1121 | 203.1229 | 345.2586 | 490.5841 | 433.8987 | 328.9548 | 467.6185 | 469.5516 | 421.4268 | 328.1357 | 388.0904 | 277.5172 | 250.2478 | 201.3255 | 314.8771 | 335.9548 | 292.472 | 239.2134 | 260.4051 | 1958.1229 | 1136.8752 | 870.7631 | 729.821 | 1048.1357 | 893.7306 | 712.2586 | 570.5605 | 2748.7758 | 1443.4051 | 1240.6637 | 1354.3019
396 | 6.3504 | 929.6947 | 1179.763 | 966.8623 | 985.8013 | 1366.7252 | 1104.4887 | 839.9238 | 861.6338 | 731.3053 | 462.2901 | 356.8318 | 504.7861 | 1468.412 | 761.2212 | 713.6795 | 998.6947 | 1419.5344 | 849.6721 | 668.3053 | 876.504 | 1517.2212 | 1527.9848 | 1274.0762 | 980.0993 | 2111.9464 | 1406.6106 | 1057.7556 | 1368.3742 | 1247.0689 | 842.1146 | 573.412 | 774.7556 | 1281.0993 | 1028.2901 | 753.8933 | 445.145 | 3050.1072 | 1866.8629 | 1442.5113 | 1505.0147 | 1700.9391 | 1099.0689 | 812.1377 | 868.649 | 371.6411 | 343.5649 | 256.9391 | 627.3437 | 1404.8781 | 1005.7026 | 898.4046 | 592.2748 | 1524.6795 | 988.1907 | 746.7404 | 644.5722 | 1424.8166 | 1036.2364 | 912.0609 | 800.9848 | 799.9848 | 533.2133 | 463.9769 | 413.7099 | 1370.6106 | 1027.504 | 864.4729 | 945.0536 | 1550.8702 | 1076.6795 | 1155.0993 | 1110.0993 | 1935 | 1251.9616 | 960.4656 | 1201.9464 | 1898.9464 | 1286.5497 | 860.6338 | 1329.496 | 974.5344 | 703.8933 | 572.8933 | 760.4046 | 708.7331 | 539.206 | 460.5418 | 400.2364 | 1337.5649 | 746.2974 | 581.0457 | 841.2364 | 3368.7252 | 1949.4582 | 1586.6411 | 2005.68 | 3172.2669 | 1874.8696 | 1558.7252 | 2001.5722 | 351.6947 | 202.0609 | 176.3662 | 351.6947 | 479.1755 | 444.4046 | 313.9007 | 446.0457 | 474.8245 | 403.3358 | 329.6947 | 400.2444 | 301.504 | 259.6795 | 227.9391 | 311.7099 | 298.6716 | 240.702 | 228.8854 | 226.0762 | 1936.496 | 1148.5649 | 875.4351 | 700.0609 | 1087.3132 | 891.145 | 761.443 | 571.4503 | 2802.2827 | 1442.4199 | 1248.6795 | 1296.6027
397 | 6.3614 | 913.6549 | 1134.9115 | 963.0708 | 1000.8142 | 1317.4336 | 1068.3716 | 810.3097 | 814.9646 | 711.7876 | 500.1593 | 386.2921 | 491.3274 | 1465.2035 | 776.8142 | 710.7965 | 1019.0797 | 1370.8318 | 841.1858 | 675.6726 | 868.8584 | 1566.6372 | 1534.2035 | 1296.7522 | 924.0265 | 2180.5841 | 1468.177 | 1048.9292 | 1406.6903 | 1239.8584 | 910.7788 | 660.3983 | 825.77 | 1253.646 | 984.6371 | 749.7257 | 395.5663 | 3026.4424 | 1855.1858 | 1451.7434 | 1573.4603 | 1658.1681 | 1147.3629 | 804.6549 | 813.1592 | 379.1416 | 343.1327 | 228.9115 | 634.7345 | 1344.4867 | 983.0442 | 886.9911 | 559.177 | 1532.2035 | 952.0442 | 704.2389 | 652.3451 | 1505.593 | 1056.885 | 947.0266 | 799.3982 | 785.3894 | 577.1062 | 485.4868 | 396.6549 | 1420.0354 | 1069.2921 | 917.9116 | 1007.3717 | 1608.9735 | 1105.0177 | 1113.7699 | 1026.2743 | 1910.7168 | 1261.4779 | 951.9292 | 1232.4248 | 1912.0797 | 1288 | 845.1858 | 1320.6637 | 956.1858 | 717.0708 | 560.0531 | 757.6637 | 670.1681 | 515.0531 | 472.9469 | 394.8673 | 1331.0619 | 749.0089 | 597.0797 | 895.708 | 3398.354 | 1902.3628 | 1683.4691 | 1961.0884 | 3152.7257 | 1919.177 | 1572.7434 | 1978.1239 | 371.2124 | 198.8673 | 175.3363 | 371.2124 | 494.3451 | 472.885 | 339.6195 | 448.2035 | 440.5752 | 435.5576 | 324.9292 | 377.8495 | 321.6106 | 269.8053 | 213.7876 | 347.5576 | 310.2124 | 273.6284 | 222.6637 | 256.5576 | 1954.5487 | 1110.8407 | 900.1505 | 752.3717 | 1096.7345 | 921.354 | 743.115 | 583.0089 | 2769.5044 | 1470.7522 | 1280.4868 | 1347.4425
398 | 6.3724 | 899.5659 | 1173.9774 | 980.7114 | 960.0226 | 1375.9661 | 1042.6499 | 806.7591 | 823.7114 | 685.2522 | 483.1317 | 365.6135 | 489.1205 | 1524.9297 | 838.0502 | 708.2409 | 1021.1205 | 1360.3614 | 828.325 | 652.3726 | 894.266 | 1561.8068 | 1556.9887 | 1279.8908 | 940.5069 | 2105.8043 | 1495.591 | 1054.1568 | 1335.4793 | 1226.5659 | 858.5521 | 626.6248 | 812.0113 | 1313.207 | 1005.266 | 792.857 | 414.3865 | 3057.0615 | 1846.7227 | 1437.2886 | 1592.2773 | 1629.8908 | 1072.202 | 779.1317 | 856.7729 | 403.8683 | 348.2773 | 244.2296 | 636.7591 | 1384.1343 | 975.3614 | 849.8181 | 549.6023 | 1518.9272 | 1004.1706 | 734.941 | 640.9272 | 1480.2999 | 1070.675 | 911.6612 | 821.8683 | 776.8432 | 583.3614 | 471.5295 | 441.4956 | 1409.5295 | 999.3588 | 947.3501 | 970.2635 | 1553.793 | 1094.0477 | 1080.6135 | 1083.3639 | 2009.0276 | 1220.9021 | 954.5182 | 1218.6499 | 1907.0841 | 1253.6976 | 907.9661 | 1321.7591 | 944.2045 | 770.7729 | 537.6499 | 706.8657 | 689.3865 | 512 | 473.3614 | 387.0841 | 1317.6863 | 740.4454 | 592.9636 | 821.4429 | 3512.064 | 1906.3137 | 1584.6587 | 1983.5433 | 3202.8934 | 1970.3388 | 1612.8206 | 2004.1455 | 315.9498 | 221.8683 | 176 | 315.9498 | 454.7817 | 451.4931 | 335.2045 | 461.3137 | 481.6161 | 419.8908 | 330.1568 | 374 | 315.2045 | 230.5408 | 216.2773 | 317.8181 | 358.7365 | 262.409 | 218.4818 | 246.1681 | 1965.0364 | 1099.7227 | 852.1066 | 766.2773 | 1089.9636 | 872.3475 | 738.1205 | 528.7089 | 2688.2384 | 1467.9636 | 1277.0841 | 1288.3952
399 | 6.3834 | 967.5635 | 1145.2183 | 966.0551 | 994.1328 | 1352.0019 | 1068.5389 | 761.5162 | 864.3491 | 653.4611 | 452.353 | 394.9981 | 514.8633 | 1529.4326 | 751.465 | 656.2733 | 1050.4307 | 1383.1878 | 813.6224 | 663.2694 | 855.1917 | 1547.5142 | 1555.5408 | 1258.2714 | 970.7552 | 2242.9372 | 1409.7079 | 1031.8122 | 1339.2694 | 1192.8938 | 869.6205 | 636.9449 | 825.0531 | 1289.8938 | 1041.9981 | 749.9488 | 416.2163 | 3069.1347 | 1804.9754 | 1431.4326 | 1590.3245 | 1653.7552 | 1086.8633 | 783.1347 | 823.8672 | 454.2675 | 342.7571 | 226.4877 | 601.3264 | 1390.1082 | 1061.5084 | 866.4041 | 566.8368 | 1442.0855 | 1009.3245 | 724.8387 | 674.6736 | 1556.1574 | 980.1405 | 928.296 | 765.2468 | 792.9449 | 565.1632 | 478.8102 | 417.6775 | 1423.9449 | 1024.6736 | 939.2979 | 980.9449 | 1553.9184 | 1089.3245 | 1175.1023 | 1115.296 | 1953.4365 | 1341.6412 | 979.9715 | 1210.649 | 1871.2183 | 1272.1878 | 854.5712 | 1278.2999 | 994.7267 | 756.5959 | 555.5123 | 786.5084 | 656.4346 | 485.2448 | 466.7571 | 418.1062 | 1375.7532 | 777.3491 | 592 | 834.9715 | 3387.873 | 1940.9981 | 1677.8044 | 1969.2714 | 3242.1062 | 2004.5389 | 1534.1671 | 1994.6224 | 362.4022 | 235.7021 | 180.4592 | 362.4022 | 488.2409 | 436.4061 | 345.5939 | 472.8102 | 494.2429 | 422.351 | 353.296 | 401.647 | 315.7837 | 275.8348 | 215.2163 | 347.7817 | 348.0551 | 298.3491 | 224.2429 | 263.6205 | 2022.1859 | 1098.1082 | 859.2694 | 810.7001 | 1064.9203 | 882.8368 | 751.3776 | 563.8083 | 2698.5123 | 1501.7817 | 1254.5959 | 1270.9734
400 | 6.3944 | 947.0678 | 1242.2626 | 1004.6859 | 979.1091 | 1374.2198 | 1088.2743 | 784.9322 | 813.848 | 673.5074 | 491.4941 | 386.3422 | 514.3835 | 1491.9306 | 862.0576 | 698.823 | 1038.6298 | 1470.9883 | 869.8644 | 626.1224 | 878.9322 | 1517.972 | 1634.5635 | 1235.205 | 1009.2611 | 2195.9026 | 1409.8496 | 1054.3157 | 1334.6711 | 1204.3702 | 812.3274 | 614.5885 | 881.2478 | 1322.1652 | 1053.233 | 782.8776 | 409.6711 | 3105.2611 | 1878.6594 | 1449.9867 | 1552.0265 | 1641.6298 | 1121.6446 | 779.4793 | 828.0413 | 413.8893 | 357.3702 | 244.795 | 646.7272 | 1369.205 | 989.7241 | 841.8761 | 575.3289 | 1490.8644 | 988.205 | 688.643 | 679.9041 | 1549.7257 | 981.6578 | 888.5059 | 767.0413 | 795.9041 | 578.3702 | 484.5207 | 430.2743 | 1442.2743 | 957.5737 | 933.4793 | 1063.4676 | 1603.823 | 1068.205 | 1120.9985 | 1042.9572 | 1940.4793 | 1321.7389 | 958.5885 | 1246.1652 | 1824.5059 | 1264.1504 | 840.6711 | 1335.385 | 1015.4661 | 773.8909 | 572.4661 | 789.767 | 641.1504 | 529.0148 | 476.8496 | 449.1239 | 1350.4513 | 802.795 | 574.8215 | 900.385 | 3462.3172 | 2006.4809 | 1598.7787 | 1940.7802 | 3199.8893 | 1963.6976 | 1469.7522 | 1952.3141 | 349.1091 | 204.4513 | 167.438 | 349.1091 | 480.3422 | 452.1785 | 311.7389 | 496.6033 | 512.1785 | 421.1917 | 308.0811 | 367.9306 | 340.4115 | 244.9306 | 214.0959 | 355.6165 | 317.0678 | 315.6578 | 245.795 | 235.2596 | 1886.1474 | 1174.8511 | 860.0959 | 730.1076 | 1118.0561 | 883.1917 | 760.233 | 569 | 2801.3585 | 1512.3289 | 1273.6991 | 1295.3157
401 | 6.4054 | 956.8299 | 1184.2684 | 965.0118 | 1011.824 | 1394.412 | 1088.1672 | 802.6627 | 840.9912 | 656.7551 | 467.5909 | 363.9223 | 520.415 | 1542.0689 | 830.0953 | 664.5939 | 1031.5015 | 1426.8475 | 852.173 | 671.4868 | 839.8446 | 1507.6686 | 1535.7787 | 1302.648 | 919.525 | 2167.0059 | 1399.0865 | 1099.1554 | 1312.9223 | 1268.3167 | 881.1466 | 645.9076 | 793.525 | 1312.2537 | 1068.6627 | 753.0924 | 410.8328 | 3076.0088 | 1876.0029 | 1469.412 | 1637.8093 | 1721.5615 | 1031.525 | 788.1642 | 846.412 | 408.0836 | 369.9971 | 235.0865 | 665.7463 | 1392.6598 | 926.9311 | 839 | 565.0029 | 1525.4091 | 1010.7434 | 753.7316 | 646.0924 | 1557.1642 | 1069.1436 | 914.2419 | 783.5792 | 771.2566 | 586.415 | 480.4179 | 400.8417 | 1405.5103 | 955.4985 | 949.4956 | 958.3637 | 1556.764 | 1086.1613 | 1148.824 | 1065.2419 | 1950.0806 | 1302.4208 | 986.2419 | 1153.7757 | 1888.7316 | 1266.6657 | 872.9912 | 1308.0924 | 994.0895 | 720.0983 | 540.0924 | 802.7463 | 705.9824 | 535.8328 | 463.3373 | 443.7522 | 1361.6627 | 777.5074 | 614.239 | 872.176 | 3438.0924 | 2005.6686 | 1640.3196 | 1899.5103 | 3149.1789 | 1984.6598 | 1572.1377 | 1997.4868 | 336.9194 | 230.3255 | 201.7404 | 336.9194 | 489.9971 | 473.2449 | 329.0777 | 472.4238 | 473.6775 | 443.9105 | 306.6657 | 419.9017 | 294.4297 | 246.4985 | 222.2478 | 333.0895 | 335.9941 | 287.6745 | 219.5909 | 225.585 | 1961.7257 | 1155.5074 | 878.3284 | 759.4897 | 1044.2713 | 896.7463 | 765.5821 | 584.5792 | 2773.5103 | 1489.173 | 1199.937 | 1304.415
402 | 6.4164 | 924.5663 | 1209.5762 | 1019.5089 | 1040.0752 | 1385.7842 | 1103.7881 | 800.2851 | 877.4337 | 723.7247 | 471.501 | 392.6475 | 537.7168 | 1468.0594 | 790.7089 | 726.0099 | 1062.5762 | 1361.6337 | 828.6396 | 675.0713 | 849.002 | 1575.7247 | 1615.1564 | 1214.2713 | 963.0079 | 2287.5901 | 1451.8653 | 1015.701 | 1356.5069 | 1239.6376 | 828.5624 | 614.6376 | 827.7208 | 1371.3663 | 1077.4297 | 773.2178 | 372.9228 | 3026.5644 | 1912.1485 | 1491.4317 | 1583.7763 | 1688.1366 | 1135.3742 | 772.2832 | 870.2891 | 421.9307 | 344.0673 | 221.9267 | 680.9307 | 1400.5723 | 957.2911 | 901.2237 | 582.5743 | 1513.1406 | 1024.1446 | 741.4257 | 661.5743 | 1500.4198 | 1063.0693 | 948.5762 | 806.3604 | 798.7188 | 560.0673 | 476.2851 | 402.6436 | 1467.0099 | 970.8594 | 924.0673 | 971.2178 | 1591.9346 | 1121.4337 | 1128.7822 | 1093.0752 | 1932.4257 | 1329.7901 | 965.7822 | 1266.6614 | 1894.0713 | 1288.3604 | 889.9307 | 1278.996 | 1008.7168 | 708.499 | 543.501 | 823.503 | 629.3446 | 547.1446 | 537.2257 | 431.8554 | 1374.1446 | 814.0059 | 616.1426 | 863.4277 | 3450.7881 | 1924.202 | 1636.6416 | 1905.2871 | 3140.3564 | 2011.1465 | 1583.8574 | 1984.3545 | 337.8574 | 223.7129 | 171.5663 | 337.8574 | 487.2851 | 433.2079 | 311.497 | 445.8535 | 460.7129 | 433.9267 | 323.7168 | 402.7109 | 289.0713 | 260.002 | 214.6416 | 331.9287 | 362.1465 | 298.002 | 215.1426 | 268.6495 | 1994.1465 | 1214.295 | 840.9941 | 769.501 | 1153.0891 | 882.2119 | 738.1386 | 578.5703 | 2773.7149 | 1507.4317 | 1280.299 | 1309.6436
403 | 6.4274 | 984.1062 | 1213.882 | 969.4218 | 1013.7699 | 1390.646 | 1065.4778 | 793.413 | 894.1151 | 672.5398 | 467.295 | 358.3009 | 488.1858 | 1446.0029 | 821.8761 | 707.4749 | 1020.7728 | 1467.0974 | 851.4661 | 673.118 | 905.5192 | 1601.7021 | 1611.649 | 1209.823 | 1046.0443 | 2243.3628 | 1485.1711 | 998.649 | 1357.177 | 1245.469 | 850.3481 | 616.705 | 837.587 | 1312.9528 | 1068.59 | 715.7168 | 418.9322 | 3118.0414 | 1883.9469 | 1385.7256 | 1571.472 | 1698.1741 | 1074.3657 | 782.292 | 871.059 | 399.4749 | 382.4632 | 225.705 | 665.0619 | 1450.8732 | 1004.2272 | 924.8201 | 555.7699 | 1514.823 | 1041.9381 | 708.9469 | 648.8849 | 1577.8673 | 1077.9971 | 950.059 | 806.118 | 816.0561 | 495.8938 | 496.7021 | 436.8761 | 1468.236 | 971.059 | 909.767 | 1005.9351 | 1644.8732 | 1116.472 | 1134.528 | 1106.292 | 1964.8761 | 1275.5398 | 998.8171 | 1296.7021 | 1869.5339 | 1259.8879 | 900.41 | 1359.8083 | 1008.118 | 723.9971 | 555.292 | 796.7699 | 627.705 | 556.469 | 452.6636 | 461.1121 | 1284.6636 | 817 | 602.8259 | 905.3452 | 3451.059 | 2018.6874 | 1658.5841 | 1922.9381 | 3203.0472 | 1975.3599 | 1535.0678 | 1935.0088 | 319.1799 | 240.8791 | 192.5251 | 319.1799 | 517.1121 | 439.41 | 330.7021 | 440.236 | 449.649 | 430.177 | 339.1151 | 370.8879 | 305.9381 | 227.1239 | 218.705 | 316.0029 | 348.9439 | 289.59 | 244.1711 | 261.649 | 2053.4012 | 1217.118 | 877.5222 | 825.5192 | 1136.5929 | 851.8289 | 733.177 | 574.236 | 2739.1829 | 1495.8259 | 1205.1327 | 1359.8732
404 | 6.4384 | 953.6814 | 1249.2719 | 959.3269 | 995.794 | 1439.618 | 1025.8215 | 797.7665 | 870.2143 | 703.3186 | 496.5521 | 392.2252 | 529.8049 | 1520.3103 | 819.2335 | 713.6264 | 1051.3653 | 1466.3736 | 809.1484 | 662.5138 | 840.3628 | 1622.0659 | 1570.9616 | 1303.4229 | 1057.6731 | 2210.4479 | 1513.6923 | 1060.9175 | 1334.1209 | 1208.8215 | 875.8324 | 640.8324 | 868.5054 | 1289.9341 | 1098.5054 | 769.1977 | 423.9066 | 3255.5547 | 1861.0275 | 1464.7965 | 1626.2911 | 1778.1235 | 1151.0301 | 775.3736 | 863.3736 | 408.4862 | 328.7556 | 245.0659 | 635.4012 | 1419.6814 | 960.2886 | 888.8215 | 600.7114 | 1581.7306 | 1040.1401 | 721.2994 | 691.8516 | 1612.5521 | 1055.1676 | 907.1017 | 825.0659 | 836.0659 | 544.4312 | 462.7281 | 438.0467 | 1468 | 954.794 | 900.4203 | 1041.3653 | 1607.9616 | 1148.412 | 1120.7006 | 1151.8049 | 2008.945 | 1334.9175 | 917.1101 | 1271.3077 | 1959.5163 | 1308.5246 | 877.1676 | 1365.9533 | 987.0275 | 779.3378 | 539.794 | 784.5138 | 695.6839 | 538.8874 | 485.1318 | 424.8682 | 1331.4312 | 819.8599 | 599.1401 | 904.1868 | 3473.8791 | 2037.3928 | 1671.4395 | 1905.8874 | 3213.6731 | 1994.9258 | 1557.739 | 1963.4587 | 322.7665 | 206.7281 | 197.8132 | 322.7665 | 468.4754 | 458.1126 | 336.7665 | 478.1318 | 469.9725 | 428.0934 | 329.5138 | 453.8432 | 323.206 | 245.9725 | 237.1126 | 350.2719 | 316.5413 | 295.6731 | 233.6072 | 239.0742 | 2055.0934 | 1239.8791 | 957.2169 | 791.8215 | 1147.3928 | 905.2911 | 737.7665 | 576.8599 | 2712.2143 | 1444.4754 | 1251.4779 | 1279.1101
405 | 6.4494 | 980.002 | 1175.6843 | 992.7955 | 984.3786 | 1435.2409 | 1053.9332 | 746.824 | 871.8968 | 735.8987 | 485.4474 | 345.7207 | 487.5831 | 1569.3825 | 828.6559 | 718.8279 | 1059.7591 | 1395.5122 | 897.7651 | 678.415 | 900.7287 | 1592.1013 | 1588.3117 | 1225.9252 | 1052.2065 | 2148.1681 | 1437.7531 | 1030.2045 | 1461.4229 | 1263.004 | 919.4858 | 612.0668 | 815.9272 | 1413.5606 | 1070.0668 | 735.3077 | 393.1013 | 3164.4758 | 1945.9371 | 1400.4434 | 1673.4169 | 1741.4454 | 1127.9636 | 801.0708 | 833.0668 | 406.1032 | 326.9656 | 237.3097 | 678.4169 | 1477.8663 | 987.9332 | 897.6214 | 592.3786 | 1610.1052 | 1056.415 | 726.8279 | 707.5182 | 1568.6175 | 1024.0668 | 888.585 | 834.6903 | 794.5142 | 557.6214 | 469.6903 | 395.5142 | 1450.6195 | 972.3461 | 903.8623 | 1066.1741 | 1746.9755 | 1095.9272 | 1152.8299 | 1071.9252 | 2014.8623 | 1317.7227 | 1029.8359 | 1249.7227 | 1971.7247 | 1335.1396 | 876.9656 | 1437.4533 | 1024.6234 | 737.5831 | 540.9312 | 741.5142 | 699 | 560.2085 | 481.2065 | 427.8279 | 1377.4514 | 797.7915 | 625.0708 | 853.7895 | 3521.3481 | 2060.2085 | 1643.0324 | 1988.0404 | 3172.4798 | 1895.5791 | 1587.002 | 1903.2025 | 301.7571 | 192.4474 | 153.5831 | 301.7571 | 475.6559 | 439.6883 | 303.2045 | 506.0708 | 482.587 | 430.8968 | 315.4818 | 387.5122 | 313.3786 | 236.3786 | 195.5142 | 319.1701 | 341.0708 | 267.998 | 228.1721 | 237.0344 | 1943.9576 | 1281.5546 | 839.3362 | 844.0728 | 1135.4474 | 921.5182 | 723.5162 | 608.8643 | 2732.2429 | 1471.9332 | 1270.415 | 1323.2793
406 | 6.4604 | 941.8849 | 1232.6505 | 952.9292 | 1045.6062 | 1416.4203 | 1094.1151 | 766.5133 | 929.6947 | 768.2921 | 510.4248 | 350.8451 | 552.4956 | 1559.2655 | 849.5354 | 693.5752 | 1106.9381 | 1367.5752 | 866.7743 | 664.3318 | 901.0442 | 1606.646 | 1560.6416 | 1230.823 | 1109.6947 | 2263.3452 | 1549.4116 | 1082.7832 | 1399.5044 | 1230.7743 | 839.8362 | 632.5133 | 871.6947 | 1376.9292 | 1045.531 | 785.8275 | 432.093 | 3245.0974 | 1927.4867 | 1415.6018 | 1535.1636 | 1711.6416 | 1188.6062 | 742.3495 | 823.1991 | 408.9336 | 338.7345 | 246.7788 | 669.2434 | 1459.4646 | 944.9955 | 886.2655 | 605.6903 | 1542.5486 | 968.9911 | 760.2478 | 691.3761 | 1625.6726 | 1092.4293 | 923.2036 | 851.6239 | 767.5752 | 563.8673 | 438.7079 | 416.4912 | 1504.7611 | 1014.0708 | 901.0664 | 994.6371 | 1685.5044 | 1074.4425 | 1113.907 | 1100.2921 | 2003.2655 | 1377.6284 | 1017.3761 | 1281.2699 | 1963.1991 | 1313.5088 | 881.8894 | 1407.7301 | 1015.2434 | 722.3097 | 536.1106 | 810.4071 | 698.0221 | 539.4867 | 497.6239 | 428 | 1309.5707 | 846.8717 | 589.8186 | 940.0089 | 3611.9868 | 1981.792 | 1603.8628 | 2000.7788 | 3077.1238 | 1971.208 | 1587.0221 | 1897.0885 | 329.3584 | 206.6682 | 165.6903 | 329.3584 | 514.1372 | 433.1327 | 316.6682 | 460.0619 | 444.8628 | 429.0442 | 285.6637 | 402.6018 | 311.0442 | 283.916 | 229.2036 | 302.354 | 328.3097 | 274.823 | 245.6018 | 226.2434 | 2019.208 | 1236.0619 | 840.8673 | 823.5088 | 1163.3584 | 889.7301 | 737.6682 | 604.1327 | 2746.6903 | 1520.916 | 1234.8186 | 1256.5486
407 | 6.4714 | 996.4494 | 1248.8525 | 949.0295 | 1046.0098 | 1527.889 | 1047.4719 | 799.6755 | 881.4916 | 741.2753 | 455.5506 | 382.6854 | 571.823 | 1588.705 | 848.0197 | 682.1082 | 1119.882 | 1487.8005 | 887.7837 | 659.0492 | 907.9312 | 1656.5084 | 1640.2036 | 1281.4986 | 989.2093 | 2174.9046 | 1549.0295 | 985.9734 | 1419.7837 | 1330.007 | 912.2626 | 591.4129 | 880.9213 | 1377.9803 | 1110.3511 | 747.3933 | 422.1082 | 3173.7276 | 1914.1278 | 1414.0197 | 1554.7739 | 1717.9312 | 1156.3343 | 779.6165 | 828.941 | 402.0688 | 348.9017 | 236.1082 | 642.2655 | 1475.8329 | 894.4916 | 867.1868 | 566.3933 | 1498.4228 | 1053.1446 | 778.823 | 721.6952 | 1654.7247 | 1054.3933 | 920.0393 | 824.2753 | 825.4199 | 522.4129 | 453.8427 | 439.7739 | 1425.7964 | 984.3048 | 901.9902 | 1050.4297 | 1605.7767 | 1146.2823 | 1143.6952 | 1128.7247 | 1967.354 | 1418.6067 | 901.1504 | 1306.7542 | 1981.8132 | 1327.8525 | 862.1966 | 1390.1671 | 1003.118 | 742.7935 | 503.3245 | 828.8329 | 711.8624 | 510.2851 | 445.5211 | 400.2753 | 1328.7935 | 800.4719 | 631.5772 | 940.0197 | 3496.17 | 2021.5871 | 1626.764 | 2027.7345 | 3181.9381 | 1928.4424 | 1509.7669 | 1934.6264 | 351.7837 | 200.0688 | 178.8722 | 351.7837 | 552.6264 | 430.0295 | 308.0885 | 488.705 | 480.6362 | 460.6854 | 314.705 | 389.1376 | 305.059 | 236.4818 | 196.3343 | 339.6264 | 377.5084 | 285.8918 | 263.823 | 242.8329 | 2136.8496 | 1261.7345 | 836.0492 | 878.4494 | 1087.7571 | 868.2065 | 769.6854 | 614.8918 | 2658.8751 | 1417.0422 | 1319.1544 | 1371.8398
408 | 6.4824 | 1011.968 | 1316.823 | 974.8377 | 1054.0664 | 1430.2779 | 1086.9115 | 780.9385 | 874.1082 | 728.1647 | 506.9484 | 367.9656 | 508.0418 | 1621.9385 | 837.9926 | 737.9361 | 1077.9533 | 1477.1033 | 906.0172 | 667.059 | 958.9484 | 1694.0246 | 1674.059 | 1221.0664 | 1045.2066 | 2224.177 | 1497.091 | 1058.9287 | 1441.1918 | 1255.1229 | 910.1057 | 612.8844 | 910.9852 | 1390.8967 | 1094.9779 | 710.0885 | 445.032 | 3153.6565 | 1913.0123 | 1396.968 | 1643.9164 | 1745.0221 | 1112.1721 | 756.1426 | 829.0369 | 418.0787 | 352.9459 | 240.9754 | 685.9656 | 1447.0295 | 935.0984 | 878.9066 | 601.1057 | 1646.7344 | 972.0885 | 760.909 | 722.0393 | 1560.2656 | 1136.9016 | 905.9287 | 941.7762 | 805.1475 | 535.0467 | 448.9656 | 442.0885 | 1445.0738 | 957.9951 | 957.7934 | 1076.0295 | 1658.968 | 1033.0689 | 1060.1303 | 1134.9238 | 2039.9139 | 1300.0664 | 987.9508 | 1275.9705 | 2027.7861 | 1372 | 888.9213 | 1435.9016 | 1050.9287 | 721.1451 | 547.9607 | 809.8746 | 712.0049 | 524.0172 | 472.9803 | 427.9779 | 1353.1377 | 767.3147 | 611.909 | 982.8131 | 3637.9631 | 2051.1008 | 1610.9705 | 2050.0664 | 3094.182 | 2023.1647 | 1608.9213 | 1914.9139 | 345.0172 | 202.0689 | 158.0246 | 345.0172 | 495.0541 | 464.0934 | 318.9508 | 497.9459 | 485.0984 | 440.9951 | 312.0098 | 426.0393 | 310.1082 | 265.0221 | 188.059 | 326.9951 | 323.0566 | 299.9877 | 215.0148 | 267.0221 | 2030.0492 | 1320.8697 | 848.1623 | 849.0984 | 1165.9877 | 939.973 | 742.973 | 549.1844 | 2717.1721 | 1448.118 | 1245.1303 | 1306.0787
409 | 6.4934 | 998.469 | 1244.705 | 909.8998 | 1080.4247 | 1540.758 | 1050.6607 | 755.3657 | 917.5427 | 794.2035 | 485.8377 | 354.2803 | 525.944 | 1597.8998 | 834.0412 | 711.4837 | 1059.7965 | 1517.5987 | 912.1445 | 690.5132 | 938.118 | 1703.9115 | 1698.3098 | 1247.1445 | 1128.351 | 2296.2655 | 1533.4837 | 1029.9557 | 1518.351 | 1304.3362 | 952.4837 | 566.7228 | 904.9705 | 1350.0325 | 1086.8703 | 745.7345 | 458.0738 | 3416.7992 | 1919.1358 | 1384.6933 | 1610.0738 | 1753.4985 | 1181.9852 | 813.115 | 843.9557 | 449.6165 | 331.708 | 230.8377 | 672 | 1459.9588 | 973.9822 | 841.3835 | 642.4955 | 1539.8703 | 1008.6638 | 723.0412 | 737.7787 | 1667.7787 | 1096.115 | 877.2655 | 851.118 | 864.2772 | 554.1918 | 434.7787 | 478.2065 | 1476.0473 | 956.059 | 875.4308 | 1088.3983 | 1647.1948 | 1061.8113 | 1113.1475 | 1104.2213 | 2003.82 | 1327.177 | 968 | 1264.3393 | 1942.4308 | 1372.1033 | 857.3688 | 1396.6785 | 1022.3098 | 779.1887 | 532.7228 | 759.4573 | 714.2213 | 531.1475 | 465.5605 | 419.1623 | 1408.9262 | 894.8082 | 575.8113 | 907.2065 | 3622.9705 | 2091.8967 | 1600.1653 | 2075.9527 | 3167.6017 | 2089.41 | 1577.9145 | 1881.0768 | 351.823 | 229.6017 | 168.0885 | 351.823 | 516.5575 | 501.4837 | 299.6048 | 476.472 | 525 | 438.8672 | 315.9262 | 441.7935 | 353.8525 | 274.1328 | 212.1475 | 325.767 | 346.6343 | 294.7787 | 220.8377 | 275.587 | 2050.649 | 1267.8967 | 912.8642 | 889.2508 | 1161.3245 | 928.7787 | 732.0885 | 623.41 | 2784.4335 | 1496.0738 | 1296.4955 | 1338.2065
410 | 6.5044 | 963.0811 | 1225.8924 | 970.027 | 1043.3521 | 1394.3531 | 1029.7847 | 715.0822 | 889.3796 | 742.0816 | 475.1352 | 372.8648 | 587.6209 | 1657.4862 | 772.2714 | 676.2158 | 1113.5679 | 1425.0005 | 857.2714 | 657.7025 | 946.1622 | 1700.2714 | 1718.8918 | 1189.4862 | 1086.1357 | 2313.7025 | 1498.8918 | 1027.2704 | 1474.5133 | 1260.1352 | 917.5403 | 614.973 | 902.0265 | 1419.5138 | 1144.2699 | 729.2709 | 461.5127 | 3209.5439 | 1993.6479 | 1429.4046 | 1615.1352 | 1719.0536 | 1182.3791 | 753.6214 | 841.9194 | 422.9454 | 378.7296 | 220.3245 | 673.0005 | 1522.5398 | 907.7577 | 866.5944 | 543.4602 | 1596.5398 | 1049.9984 | 661.0551 | 721.1342 | 1654.2709 | 1038.1898 | 896.2709 | 860.6495 | 817.7036 | 566.2699 | 421.7036 | 490.6479 | 1545.1076 | 960.7031 | 969.8102 | 1114.1888 | 1724.377 | 1116.4327 | 1122.5403 | 1119.7031 | 1927.7582 | 1338.6214 | 966.9454 | 1288.325 | 2038.2434 | 1378.5403 | 879.999 | 1440.9995 | 1042.5944 | 725.4327 | 580.3781 | 790.8924 | 728.8107 | 539.9724 | 501.6209 | 430.027 | 1402.2153 | 880.7832 | 627.4581 | 921.3786 | 3624.4342 | 2085.649 | 1674.8903 | 2007.8929 | 3138.4581 | 2049.5673 | 1636.5933 | 1952.973 | 340.0541 | 203.2163 | 173.8378 | 340.0541 | 487.5679 | 466.4051 | 339.3781 | 507.7296 | 524.5944 | 430.8112 | 310.0536 | 428.0541 | 344.9735 | 282.351 | 221.351 | 376.2429 | 387.9724 | 281.1628 | 210.7842 | 248.2704 | 2095.0546 | 1261.9194 | 837.3786 | 905.351 | 1183.5949 | 914.1352 | 737.2969 | 584.3245 | 2614.6765 | 1498.7556 | 1197.0816 | 1352.1622
411 | 6.5154 | 967.1013 | 1257.646 | 971.9833 | 1092.9833 | 1514.6067 | 1092.1121 | 788.2694 | 973.0727 | 781.118 | 480.6293 | 366.1907 | 539.1013 | 1640.2193 | 852.938 | 648.7473 | 1133.056 | 1460.9607 | 936.7974 | 647.6686 | 952.5113 | 1779.3647 | 1714.2134 | 1172.1013 | 1126.764 | 2308.3098 | 1496.6126 | 1036.528 | 1458.1406 | 1266.9273 | 901.59 | 611.8367 | 867.118 | 1438.118 | 1121.2646 | 771.6961 | 409.298 | 3443.0727 | 1944.0954 | 1371.882 | 1618.6627 | 1692.8271 | 1230.4267 | 738.82 | 875.354 | 386.416 | 366.1681 | 232.236 | 708.2527 | 1472.7139 | 970.8033 | 852.7473 | 596.6293 | 1547.5399 | 946.0119 | 735.354 | 658.1299 | 1699.4494 | 1080.4887 | 938.8141 | 918.6234 | 879.236 | 540.8653 | 483.0393 | 444.0846 | 1514.2586 | 986.6293 | 927.236 | 1085.9833 | 1633.3039 | 1131.0954 | 1106.7866 | 1148.3432 | 2025.7414 | 1327.8712 | 929.118 | 1338.2813 | 2042.6735 | 1363.1013 | 807.6853 | 1408.6579 | 1030.3993 | 739.5447 | 555.5221 | 823 | 722.0787 | 503.4327 | 453.18 | 431.1967 | 1339.0226 | 837.0393 | 538.7926 | 935.7473 | 3741.9035 | 2104.3981 | 1566.1859 | 2077.298 | 3049.5959 | 2037.6186 | 1549.3707 | 1951.292 | 342.236 | 212.298 | 169.1013 | 342.236 | 508.118 | 446.8093 | 316.4494 | 499.1406 | 509.0954 | 460.5113 | 278.3993 | 429.8033 | 377.9154 | 260.8486 | 198.354 | 351.1239 | 350.528 | 323.354 | 239.7866 | 259.6126 | 2135.9499 | 1294.3707 | 850.0167 | 882.1573 | 1204.0167 | 919.5506 | 714.6353 | 597.4553 | 2674.5673 | 1420.5959 | 1237.2193 | 1360.7926
412 | 6.5264 | 993.2773 | 1250.8584 | 995.6903 | 1118.6003 | 1503.1932 | 1049.6136 | 748.0649 | 925.7419 | 784.0516 | 497.3422 | 323.2906 | 567.4971 | 1669.7935 | 826.587 | 666.6903 | 1113.6652 | 1461.9101 | 881.9159 | 662.1932 | 967.1682 | 1720.9883 | 1695.6711 | 1201.736 | 1122.0325 | 2432.7344 | 1536.9292 | 1024.9484 | 1484.8127 | 1314.5162 | 918.6844 | 559.6711 | 870.826 | 1438.9867 | 1197.2831 | 694.3555 | 441.7227 | 3394.0383 | 1920.5029 | 1368.4838 | 1588.4263 | 1878.3731 | 1190.7094 | 713.6652 | 884.4646 | 420.6903 | 300.2006 | 237.5354 | 689.587 | 1538.2257 | 963.1091 | 870.6386 | 612.8776 | 1633.7935 | 1043.174 | 742.6062 | 760.1357 | 1689.0458 | 1066.2448 | 864.9234 | 886.7036 | 882.8318 | 563.2906 | 483.4321 | 492.9351 | 1542.4188 | 1001.8451 | 932.0192 | 1113.3555 | 1715.8584 | 1111.5295 | 1125.4321 | 1224.9926 | 1995.7227 | 1396.0899 | 932.826 | 1389.199 | 1938.6003 | 1392.1682 | 801.9101 | 1497.2257 | 1089.3097 | 707.4588 | 497.7877 | 824.3422 | 724.7743 | 516.736 | 482.4646 | 436.1549 | 1365.6519 | 836.1932 | 606.4838 | 953.587 | 3597.8526 | 1998.1432 | 1640.2315 | 2111.4263 | 3110.2123 | 2119.9601 | 1536.2198 | 1930.851 | 347.2257 | 243.3289 | 195.5354 | 347.2257 | 510.4838 | 489.9675 | 303.2581 | 525.8127 | 487.0516 | 471.8643 | 336.4838 | 426.7036 | 325.761 | 305.5354 | 206.7419 | 400.9867 | 339.0516 | 331.2773 | 257.8127 | 297.761 | 2208.3614 | 1280.1873 | 828.5295 | 888.8909 | 1179.7419 | 930.2123 | 777.3997 | 630.7802 | 2666.4263 | 1484.0516 | 1266.2257 | 1422.7536
413 | 6.5374 | 978.2104 | 1281.4208 | 988.7856 | 1146.7217 | 1476.5193 | 1113.763 | 786.1917 | 924.707 | 785.8309 | 519.1878 | 349.8761 | 593.2743 | 1667.9174 | 822.1318 | 661.3196 | 1178.5073 | 1493.7404 | 863.7257 | 629.3196 | 1009.4287 | 1827.9321 | 1701.7364 | 1252.1917 | 1138.6991 | 2344.5645 | 1558.7936 | 1025.7896 | 1469.1358 | 1325.3422 | 965.1878 | 576.2823 | 890.2823 | 1406.2823 | 1133.5379 | 754.6578 | 455.5113 | 3361.004 | 1895.8761 | 1361.4927 | 1629.8043 | 1779.0266 | 1170.0679 | 778.2517 | 895.237 | 415.8948 | 357.1465 | 228.8083 | 685.707 | 1528.9214 | 909.8722 | 865.5339 | 627.0826 | 1628.53 | 1026.8495 | 718.6618 | 778.5713 | 1746.1505 | 1036.5566 | 933.7404 | 913.526 | 847.3462 | 586.6804 | 471.4661 | 455.9587 | 1496.4061 | 998.1052 | 915.2144 | 1174.2969 | 1741.819 | 1153.526 | 1113.7217 | 1137.4101 | 2005.4208 | 1363.6844 | 949.1505 | 1301.5752 | 1971.2596 | 1431.0413 | 839.5566 | 1482.1052 | 1092.8269 | 768.5713 | 550.0226 | 848.233 | 740.9813 | 566.9361 | 492.4061 | 437.8495 | 1395.7443 | 805.8122 | 593.8682 | 941.0786 | 3688.0865 | 2112.0865 | 1574.0679 | 2157.5979 | 3066.6058 | 2009.2557 | 1572.2703 | 1872.1691 | 331.4021 | 230.233 | 187.4474 | 331.4021 | 503.3009 | 474.3235 | 307.8722 | 505.1505 | 488.767 | 450.9361 | 327.4474 | 455.6991 | 305.5566 | 295.7856 | 203.5979 | 366.6205 | 340.767 | 318.3835 | 236.1917 | 277.0453 | 2215.3196 | 1306.2823 | 874.1691 | 942.5113 | 1173.8495 | 878.8083 | 747.2969 | 595.4327 | 2713.4927 | 1485.3835 | 1249.1878 | 1313.6245
414 | 6.5484 | 959.1633 | 1233.8771 | 973.7434 | 1129.5241 | 1564.9891 | 1087.0202 | 784.1229 | 979.9907 | 797.619 | 480.7434 | 390.4759 | 567.6952 | 1702.4012 | 866.3997 | 666.9145 | 1145.4292 | 1456.2877 | 889.1524 | 634.4572 | 1046.9813 | 1783.6858 | 1621.9642 | 1259.4199 | 1107.972 | 2465.5801 | 1623.8942 | 1047.8087 | 1559.3048 | 1343.9331 | 926.591 | 637.3997 | 948.3515 | 1466.3328 | 1202.4386 | 734.3142 | 461.0187 | 3408.2006 | 1938.0762 | 1394.4184 | 1588.944 | 1840.972 | 1214.6952 | 740.1337 | 923.0467 | 430.7527 | 332.4386 | 227.7527 | 738.0949 | 1596.0467 | 909.0669 | 884.0373 | 603.2675 | 1611.8771 | 1016.4012 | 742.8476 | 744.1431 | 1760.5614 | 1071.4852 | 888.7434 | 852.4199 | 907.8756 | 579.79 | 452.5054 | 468.7807 | 1518.2955 | 985.3048 | 927.2084 | 1166.8196 | 1626.4199 | 1090.8196 | 1095.9145 | 1205.353 | 1961.7636 | 1403.5334 | 963.8663 | 1320.8974 | 2022.1244 | 1417.9145 | 873.4184 | 1466.79 | 1059.0187 | 738.0295 | 561.9518 | 839.7434 | 765.9611 | 566.2286 | 508.8087 | 420.5428 | 1390.7045 | 851.1151 | 549.7527 | 877.3826 | 3705.4479 | 2126.1711 | 1620.0669 | 2177.4184 | 3168.6268 | 2013.7621 | 1503.8383 | 1868.4588 | 309.9051 | 217.7045 | 193.6952 | 309.9051 | 535.4759 | 520.5614 | 307.2193 | 523.3717 | 500.1524 | 452.21 | 333.9238 | 421.7714 | 341.7045 | 280.1337 | 226.0093 | 399.7418 | 351.3904 | 320.7232 | 236.79 | 305.1711 | 2226.6299 | 1368.2379 | 862.2286 | 951.5241 | 1154.2566 | 877.9813 | 741.6485 | 674.4666 | 2748.3235 | 1491.1524 | 1212.0389 | 1394.8569
415 | 6.5594 | 1042.1062 | 1293.0885 | 1001.6017 | 1151.2744 | 1478.2301 | 1157.9646 | 724.7345 | 991.3805 | 795.5664 | 509.4867 | 371.7345 | 567.0089 | 1775.0442 | 848.1505 | 674.5486 | 1190.0974 | 1577.9911 | 897.4602 | 643.3629 | 1071.0885 | 1799.2567 | 1774.6107 | 1311.1416 | 1148.2478 | 2403.6813 | 1569.2567 | 998.2389 | 1558.9292 | 1328.9646 | 961.3629 | 618.9735 | 895.3629 | 1438.0708 | 1233.5929 | 726 | 438.1505 | 3451.7788 | 1948.6461 | 1352.5929 | 1663.5664 | 1879.9204 | 1209.1416 | 759.9822 | 876.7169 | 452 | 363.3274 | 246.9646 | 716.5929 | 1541.5753 | 923.5929 | 835.115 | 651.7256 | 1678.1682 | 1093.823 | 736.4867 | 759.3274 | 1737.5487 | 1043.6549 | 921.0265 | 904.4956 | 874.4248 | 555.3009 | 491.8141 | 454.7965 | 1532.938 | 994.8408 | 857.354 | 1202.6814 | 1683.8053 | 1126.8584 | 1107.708 | 1245.3805 | 2062.5132 | 1410.531 | 940.5664 | 1443.8849 | 2062.6549 | 1429.9735 | 831.6814 | 1444.5133 | 1035.0885 | 796.0796 | 524.646 | 872.823 | 717.885 | 564.1327 | 460.8319 | 419.3363 | 1374.5575 | 904.0089 | 575.5133 | 978.2035 | 3725.6549 | 2104.177 | 1633.2655 | 2134.885 | 3058.7168 | 2020.2566 | 1510.1947 | 1967.1947 | 335.0354 | 202.531 | 189.469 | 335.0354 | 516.646 | 487.4601 | 318.7256 | 539.469 | 502.0885 | 478.9115 | 332.469 | 421.0354 | 325.0265 | 304.4956 | 213.5575 | 339.0974 | 345.7699 | 285.6018 | 212.5664 | 281.0531 | 2305.4956 | 1360.2389 | 865.4425 | 970.7965 | 1132.1505 | 904.1239 | 791.5398 | 666.4071 | 2738.5311 | 1489.1062 | 1256.6637 | 1380.646
416 | 6.5704 | 1013.389 | 1296.6205 | 954.8693 | 1190.3417 | 1469.978 | 1130.9528 | 724.0157 | 985.8063 | 817.9685 | 466.0236 | 366.5811 | 597.0929 | 1726.5197 | 859.4961 | 643.1323 | 1232.0787 | 1552.5433 | 960.3717 | 675.4803 | 1066.0535 | 1875.3102 | 1844.6346 | 1293.6126 | 1199.2567 | 2329.0866 | 1638.6583 | 1017.3496 | 1516.2945 | 1314.3496 | 996.4031 | 623.9606 | 927.7827 | 1482.7213 | 1244.2551 | 726.7055 | 451.3102 | 3500.4961 | 2044.4724 | 1376.0551 | 1691.7134 | 1914.3339 | 1189.4961 | 701.0236 | 1013.6661 | 447.9685 | 370.9528 | 229.326 | 733.4031 | 1605.0472 | 950.9858 | 833.5591 | 622.1323 | 1747.7669 | 1090.589 | 690.9071 | 773 | 1809.737 | 1069.2866 | 907.2709 | 891.822 | 916.1701 | 578.1937 | 455.3717 | 464.0079 | 1529.1559 | 1043.9134 | 863.5197 | 1223.0236 | 1725.1165 | 1147.411 | 1118.8299 | 1241.5197 | 2007.1087 | 1421.7449 | 963.674 | 1361.5905 | 2091.411 | 1441.3102 | 849.9071 | 1499.1638 | 1041.5433 | 771.5433 | 519.3874 | 869.085 | 733.5512 | 516.4189 | 486.4724 | 464.8063 | 1365.6661 | 933.3874 | 619.4803 | 974.8378 | 3742.6189 | 2112.7528 | 1637.7134 | 2145.1087 | 3102.0394 | 1987.0614 | 1473.2173 | 1918.3244 | 352.3717 | 210.3181 | 182.2866 | 352.3717 | 515.4189 | 456.5433 | 286.1087 | 535.6126 | 505.1165 | 480.2173 | 325.3874 | 439.5811 | 316.7134 | 287.1858 | 208.6976 | 370.8992 | 362.4882 | 332.4961 | 238.2945 | 261.5039 | 2292.3181 | 1370.6819 | 868.589 | 979.2016 | 1240.0457 | 926.8063 | 723.9921 | 649.8992 | 2841.0535 | 1452.9307 | 1210.7906 | 1381.8378
417 | 6.5814 | 1110.8161 | 1312.1976 | 1030.7207 | 1231.6676 | 1568.0463 | 1076.6608 | 744.94 | 999.1062 | 780.7561 | 499.3324 | 348.9577 | 584.703 | 1746.9646 | 861.6716 | 679.7738 | 1313.8584 | 1600.9823 | 932.7493 | 652.6608 | 1029.0571 | 1884.7384 | 1776.2193 | 1308.4523 | 1264.1131 | 2433.2016 | 1597.6716 | 1068.7984 | 1519.7247 | 1363.5368 | 1002.94 | 576.5123 | 913.767 | 1504.9646 | 1195.2616 | 730.8515 | 462.2085 | 3494.4169 | 2007.1199 | 1432.7807 | 1701.9115 | 1951.3038 | 1194.2724 | 730.6008 | 973.6431 | 413.8093 | 320.6185 | 257.1907 | 736.2085 | 1656.8515 | 1038.0314 | 896.0245 | 657.4169 | 1705.5899 | 1084.7847 | 696.8515 | 775.2616 | 1843.293 | 1051.173 | 874.4523 | 916.3215 | 885.8447 | 567.8338 | 425.1485 | 475.4877 | 1596.0531 | 946.8093 | 885.0954 | 1253.3924 | 1741.5123 | 1150.3883 | 1150.9646 | 1261.7384 | 2032.6048 | 1465.5368 | 937.2793 | 1460.9292 | 2096.8761 | 1462.3855 | 854.9469 | 1574.7139 | 1094.9577 | 812.9714 | 509.297 | 852.5654 | 786.0068 | 540.1485 | 456.9223 | 472.6785 | 1332.9932 | 915.7207 | 595.4169 | 1013.703 | 3662.5517 | 2168.5899 | 1639.7698 | 2157.827 | 3145.4169 | 1947.6362 | 1494.8093 | 1837.9932 | 323.1662 | 228.7384 | 167.9046 | 323.1662 | 533.9469 | 505.8869 | 296.3392 | 546.0423 | 524.6785 | 498.8692 | 312.3569 | 421.8447 | 331.6608 | 272.5831 | 205.5477 | 373.7316 | 349.2262 | 326.53 | 230.9823 | 267.2439 | 2308.3638 | 1355.9932 | 855.5477 | 975.0068 | 1182.5231 | 936.1485 | 720.2016 | 650.8093 | 2778.6145 | 1486.1907 | 1198.1907 | 1421.7207
418 | 6.5924 | 1070.7286 | 1265.919 | 994.3436 | 1180.6209 | 1621.1975 | 1082.0236 | 757.5943 | 1047 | 774.8865 | 543.233 | 375.7374 | 559.2744 | 1741.7818 | 907.3554 | 682.8894 | 1253.8865 | 1586.7315 | 966.4897 | 647.9941 | 1151.2389 | 1861.8599 | 1814.3436 | 1311.8746 | 1263.8717 | 2468.9852 | 1646.615 | 1028.6268 | 1590.4837 | 1306.7552 | 1025.4985 | 562.6416 | 964.2655 | 1496.8983 | 1221.5251 | 714.7522 | 458.5074 | 3480.6209 | 1977.4234 | 1383.7493 | 1668.2685 | 1935.1165 | 1263.9911 | 755.8894 | 965.1313 | 426.4956 | 351.5074 | 242.3732 | 733.1342 | 1638.8717 | 959.6475 | 858.7552 | 647.8835 | 1669.4956 | 1130.2271 | 681.2537 | 792.8746 | 1780.6416 | 1029.1254 | 876.3702 | 917.5192 | 925.152 | 545.7581 | 442.8776 | 490.8628 | 1625.2655 | 964.385 | 873.6121 | 1280.2507 | 1731.4027 | 1089.919 | 1138.7611 | 1251.5222 | 2151.4897 | 1404.4926 | 958.1106 | 1431.8835 | 2042.2537 | 1527.0988 | 831.7522 | 1590.7345 | 1122.2389 | 753.1342 | 543.5015 | 862.767 | 740.2537 | 563.6446 | 472.3761 | 472.4837 | 1387.5222 | 894.5133 | 591.5251 | 983.7611 | 3756.4867 | 2146.6623 | 1550.2921 | 2174.9852 | 3138.6416 | 1988.764 | 1520.534 | 1876.3495 | 346.3702 | 211.1254 | 174.6239 | 346.3702 | 510.6268 | 505 | 300.8805 | 515.8776 | 530.2507 | 488.3702 | 323 | 454.382 | 331.0059 | 287.2567 | 203.5044 | 396.9823 | 350.7493 | 314.2478 | 224.6298 | 274.8687 | 2287.4204 | 1408.1401 | 864.0355 | 939.534 | 1217.2478 | 957.6387 | 758.618 | 660.6121 | 2714.6535 | 1468.3643 | 1188.5133 | 1391.6121
419 | 6.6034 | 1025.5064 | 1364.2183 | 933.4907 | 1178.9843 | 1506.0856 | 1135.3844 | 700.005 | 1050.4415 | 810.3973 | 503.5506 | 350.2675 | 614.3844 | 1807.9144 | 863.0649 | 715.4857 | 1278.3844 | 1550.296 | 942.5506 | 635.6519 | 1127.4779 | 1833.296 | 1741.2389 | 1299.0235 | 1253.9351 | 2445.5635 | 1620.7974 | 1026.9429 | 1559.0207 | 1330.1505 | 1013.1897 | 601.587 | 1010.9116 | 1538.7846 | 1280.8623 | 719.174 | 478.6156 | 3491.3738 | 2077.408 | 1383.9272 | 1705.2546 | 1913.7611 | 1250.709 | 788.8987 | 983.4415 | 418.2026 | 367.7611 | 236.4857 | 754.0364 | 1624.6676 | 1010.3844 | 879.3973 | 664.9715 | 1670.7011 | 1069.3923 | 695.7817 | 794.0649 | 1821.2389 | 1032.3402 | 878.4907 | 968.6441 | 976.1976 | 563.6234 | 456.1583 | 465.8545 | 1653.1897 | 989.3766 | 848.7817 | 1284.0649 | 1811.179 | 1195.9272 | 1163.8987 | 1300.0078 | 2122.7325 | 1398.7011 | 925.7896 | 1472.2361 | 2050.0364 | 1472.4337 | 834.6598 | 1570.9322 | 1104.122 | 770.5949 | 541.6312 | 897.4936 | 749.826 | 599.5378 | 477.6156 | 440.47 | 1434.7689 | 918.291 | 643.291 | 1006.4208 | 3738.2154 | 2220.5456 | 1644.0806 | 2159.9608 | 3176.3481 | 2025.3039 | 1586.9586 | 1836.5114 | 339.5428 | 209.2468 | 171.8623 | 339.5428 | 517.0649 | 513.6726 | 313.174 | 526.6805 | 531.0364 | 479.6156 | 318.7325 | 468.6598 | 345.826 | 301.7611 | 214.413 | 353.7611 | 348.8623 | 310.9272 | 229.9429 | 257.2468 | 2389.6312 | 1444.5506 | 933.2703 | 1026.883 | 1205.9429 | 988.3481 | 745.5792 | 632.6156 | 2777.7325 | 1453.7375 | 1214.0806 | 1368.8467
420 | 6.6144 | 1062.999 | 1332.1986 | 1012.6017 | 1232.149 | 1559.2512 | 1125.8978 | 784.1032 | 1066.1519 | 853.853 | 511.6485 | 376.2502 | 599.4995 | 1759.9971 | 873.9007 | 701.2015 | 1265.2989 | 1603.4976 | 947.1996 | 646 | 1141.8004 | 1890.6962 | 1762.4956 | 1226.8491 | 1255.6953 | 2521.4489 | 1607.15 | 1001.5959 | 1604.6495 | 1393.702 | 973.445 | 602.5988 | 1061.4537 | 1468.296 | 1273.702 | 714.5998 | 491.4012 | 3592.0087 | 2006.6933 | 1393.9007 | 1686.0925 | 1919.3457 | 1299.7498 | 778.1013 | 993.9036 | 433.7994 | 361.0497 | 226.2502 | 757.447 | 1599.8452 | 993.1003 | 921.0535 | 662.7946 | 1740.2512 | 1032.1986 | 736.0487 | 804.9007 | 1836.9474 | 1049.3992 | 951.6037 | 1001.1509 | 900.4985 | 558.5988 | 492.7527 | 494.6485 | 1605.4976 | 990.5024 | 889.4985 | 1304.4976 | 1885.4547 | 1194.5044 | 1150.1022 | 1273.5005 | 2094.3505 | 1472.8997 | 935.6504 | 1586.8539 | 2057.9455 | 1511.2025 | 832.8958 | 1664.2541 | 1150.8014 | 741.5005 | 512.4985 | 863.15 | 753.5005 | 549.595 | 490.7011 | 483.999 | 1402.5998 | 879.297 | 593.3505 | 991.3476 | 3835.8452 | 2146.0448 | 1604.7507 | 2261.4576 | 3166.4499 | 2043.9007 | 1496.001 | 1959.7517 | 366.2006 | 195.5988 | 173.8491 | 366.2006 | 540.0468 | 555.2541 | 302.4518 | 552.7011 | 531.4479 | 491.8014 | 297.3983 | 457.2989 | 349.6504 | 291.3009 | 215.5005 | 351.8481 | 350.5492 | 322.7001 | 196.0487 | 249.297 | 2358.2493 | 1443.0516 | 847.8978 | 1059.5034 | 1182.5453 | 977.2502 | 761.1022 | 654.5482 | 2749.6504 | 1511.703 | 1189.1461 | 1421.2989
421 | 6.6254 | 1095.8998 | 1348.0501 | 982.9735 | 1255.5163 | 1549.6815 | 1168.8938 | 729.2979 | 1046.9529 | 811.4101 | 528.4277 | 365.997 | 610.4336 | 1816.1918 | 866.0561 | 671.7286 | 1289.6785 | 1653.7876 | 956.4336 | 644.7021 | 1129.646 | 1952.295 | 1827.9675 | 1246.4071 | 1334.6194 | 2539.5398 | 1616.5989 | 1056.938 | 1618.0561 | 1378.2803 | 1056.9439 | 621.8643 | 1003.6254 | 1527.4542 | 1252.6018 | 715.4041 | 475.8673 | 3466.7435 | 2115.0767 | 1384.9204 | 1811.9705 | 2012.1977 | 1308.3806 | 764.4366 | 935.5428 | 449.6785 | 364.2419 | 221.0265 | 806.295 | 1689.8407 | 983.1327 | 852.997 | 751.3481 | 1724.5163 | 1051.6194 | 751.0501 | 798.8407 | 1878.7551 | 1061.2153 | 886.6755 | 980.1327 | 932.003 | 578.8378 | 447.9469 | 518.7286 | 1644.2684 | 961.3039 | 915.4867 | 1351.3806 | 1806.5693 | 1128.2213 | 1104.3481 | 1264.0265 | 2084.6755 | 1485.6254 | 917.1297 | 1524.6519 | 2141.0561 | 1481.1151 | 899.3481 | 1589.8908 | 1130.4071 | 746.4661 | 537.1888 | 870.003 | 743.3776 | 629.7256 | 469.7286 | 492.2389 | 1416.8703 | 940.5989 | 579.9439 | 1029.4307 | 3911.5634 | 2224.6696 | 1590.0531 | 2157.8526 | 3161.6991 | 2043.1947 | 1491.1682 | 1935.8407 | 349.6725 | 216.1622 | 188.7021 | 349.6725 | 593.7551 | 488.8407 | 278.1387 | 535.9469 | 563.6194 | 470.2714 | 320.8613 | 473.4041 | 342.649 | 278.5959 | 212.1918 | 386.1357 | 369.3806 | 323.1092 | 221.9204 | 292.2654 | 2392.416 | 1414.8378 | 889.2714 | 1011.357 | 1260.6194 | 976.8938 | 725.8938 | 689.0826 | 2743.9469 | 1468.8968 | 1243.8849 | 1438.0531
422 | 6.6364 | 1161.4193 | 1310.3161 | 993.4105 | 1279.4621 | 1618.9012 | 1179.2261 | 740.7124 | 1124.5851 | 851.7969 | 490.6938 | 337.9671 | 616.7168 | 1844.0472 | 887 | 658.2547 | 1323.5379 | 1687.8913 | 965.3348 | 640.618 | 1118.4062 | 1948.0714 | 1768.382 | 1255.9714 | 1307.9857 | 2539.6795 | 1670.2503 | 986.0428 | 1642.6652 | 1460.4852 | 1035.1273 | 611.9857 | 1034.7969 | 1496.213 | 1340.9155 | 703.9199 | 499.6037 | 3607.7311 | 2068.4764 | 1393.5423 | 1793.3304 | 2074.6454 | 1336.4105 | 799.7969 | 977.1515 | 474.1131 | 355.5851 | 219.4907 | 791.0758 | 1718.6037 | 959.0472 | 819.382 | 718.3963 | 1751.7783 | 1034.7596 | 706.1603 | 821.146 | 1870.7267 | 1053.5708 | 892.1559 | 956.9199 | 964.0472 | 577.6509 | 461.8397 | 507.6981 | 1644.0242 | 1014.292 | 904.5609 | 1375.7453 | 1840.359 | 1171.2261 | 1079.2503 | 1268.0758 | 2075.4951 | 1529.1888 | 874.7882 | 1551.146 | 2159.7311 | 1558.7124 | 868.6652 | 1599.595 | 1145.9056 | 820.3633 | 537.3633 | 911.9956 | 746.6938 | 588.0472 | 467.9484 | 452.0615 | 1466.0187 | 971.0373 | 569.472 | 1020.4248 | 3847.405 | 2182.8167 | 1578.2547 | 2320.0286 | 3117.5423 | 2100.8255 | 1547.3677 | 1963.0329 | 325.5522 | 213.6839 | 181.8255 | 325.5522 | 577 | 515.5093 | 311.5093 | 549.1416 | 541.8727 | 484.2689 | 291.854 | 463.6652 | 340.9385 | 287.6367 | 236.2075 | 394.618 | 386.4149 | 335.2217 | 233.3348 | 247.7596 | 2507.3633 | 1434.595 | 895.764 | 1066.8584 | 1240.618 | 996.8255 | 737.9714 | 708.3161 | 2759.4105 | 1526.2832 | 1169.6037 | 1419.9714
423 | 6.6474 | 1097.7592 | 1343.0011 | 1000.1307 | 1269.0227 | 1649.23 | 1136.7365 | 766.7365 | 1150.2527 | 819.1652 | 534.0432 | 356.1101 | 590.9233 | 1858.5281 | 887.9341 | 665.8801 | 1361.0669 | 1663.8574 | 944.4503 | 649.6372 | 1190.0216 | 2035.7138 | 1763.2862 | 1214.0767 | 1308.0561 | 2579.176 | 1718.1101 | 1034.7797 | 1679.6166 | 1390.9913 | 1031.3736 | 594.8682 | 1009.825 | 1602.6372 | 1382.1966 | 712.8132 | 503.2754 | 3628.1997 | 2062.0767 | 1373.2754 | 1825.1976 | 1980.8261 | 1352.2851 | 755.5821 | 1051.9773 | 450.6264 | 381.0659 | 233.7581 | 828.9557 | 1709.5713 | 955.4061 | 817.8358 | 721.2635 | 1755.7149 | 1062.1868 | 685.9449 | 803.9902 | 1922.649 | 1059.1318 | 933.6598 | 979.4514 | 967.6922 | 570.5821 | 467.2959 | 519.2203 | 1726.1534 | 951.77 | 851.4622 | 1393.1544 | 1902.5378 | 1128.9233 | 1122.2527 | 1319.5939 | 2074.9362 | 1553.4838 | 954.3855 | 1521.473 | 2163.5723 | 1578.7581 | 876.4719 | 1701.1101 | 1141.4179 | 830.5378 | 561.7365 | 946.2516 | 796.7689 | 594.6814 | 507.7138 | 478.4287 | 1426.8348 | 912.473 | 621.9892 | 1058.2743 | 3774.474 | 2276.8456 | 1576.1652 | 2326.053 | 3086.2527 | 2097.0108 | 1508.1534 | 1958.9017 | 370.1101 | 200.6155 | 176.5162 | 370.1101 | 570.2743 | 503.1101 | 299.297 | 568.4838 | 544.1318 | 497.0551 | 328.8132 | 480.8132 | 321.1101 | 269.8024 | 220.5713 | 401.7689 | 363.5497 | 342.3078 | 220.6706 | 282.8466 | 2528.3736 | 1542.649 | 858.1091 | 1058.1652 | 1254.1209 | 986.2851 | 715.6933 | 719.8899 | 2797.5173 | 1556.3736 | 1170.8466 | 1400.3088
424 | 6.6584 | 1166.1814 | 1416.0044 | 929.7478 | 1331.5355 | 1621.4957 | 1170.9204 | 781.6062 | 1148.3938 | 845.2035 | 493.7478 | 372.6062 | 623.3319 | 1902.5885 | 893.1947 | 722.8539 | 1418.4557 | 1691.3717 | 941.7434 | 635.7965 | 1171.5664 | 2013.1593 | 1794.0044 | 1204.5664 | 1394.7832 | 2600.7655 | 1734.6062 | 1001.3894 | 1765.7965 | 1551.0044 | 1038.5752 | 605.9867 | 1067.6106 | 1588 | 1311.7301 | 714.9867 | 552.3805 | 3749.8628 | 2077.4559 | 1408.4424 | 1824.7478 | 2097.2787 | 1334.093 | 726.6239 | 1009.177 | 450.1947 | 377.9956 | 227.7832 | 804.4115 | 1711.1461 | 924.7699 | 858.6283 | 702.0133 | 1796.0088 | 1059.8141 | 679.9867 | 875.6681 | 1967.0088 | 1062.1505 | 967 | 1047.7567 | 968.7699 | 549.9867 | 424.7434 | 545.0442 | 1697.9735 | 992.0486 | 890.8274 | 1433.1991 | 1856.6018 | 1152.5708 | 1137.1195 | 1348.6371 | 2212.8628 | 1575.5885 | 1004.6548 | 1569.3319 | 2212.0663 | 1576.5664 | 839.5797 | 1738.3142 | 1158.8274 | 786.5929 | 573.8186 | 872.1859 | 766.0354 | 591.646 | 476.7965 | 480.438 | 1422.5001 | 958.739 | 634.0044 | 1037.9602 | 3882.1947 | 2252.1328 | 1595.2345 | 2223.04 | 3105.6992 | 2094.1549 | 1482.9602 | 1993.8363 | 388.7965 | 212.1991 | 162.5752 | 388.7965 | 555.9337 | 522.792 | 317.2035 | 588.3982 | 542.5708 | 501.4203 | 326.208 | 476.615 | 344.177 | 286.5797 | 209.0044 | 376.8097 | 378.0044 | 354.3717 | 241.5973 | 308.0398 | 2532.5885 | 1569.8849 | 819.3231 | 1085.9956 | 1265.7699 | 946.3894 | 766.2787 | 706.1814 | 2847.2212 | 1535.699 | 1211.1726 | 1484.5797
425 | 6.6694 | 1186.8053 | 1391.4935 | 975.3087 | 1370.1278 | 1714.5202 | 1213.6144 | 781.708 | 1158.4395 | 849.6106 | 531.4867 | 370.1711 | 680.9833 | 1929.2783 | 913.0099 | 686.2616 | 1372.0167 | 1720.1711 | 982.6411 | 643.0167 | 1212.7818 | 2058.0335 | 1804.1643 | 1231.4061 | 1421.416 | 2647.2213 | 1730.057 | 1005.7148 | 1765.1947 | 1534.1042 | 1065.8624 | 613.1376 | 1057.3658 | 1600.0502 | 1382.3992 | 731.4395 | 553.7217 | 3731.2646 | 2182.8426 | 1355.6243 | 1895.0433 | 2050.7346 | 1418.3156 | 710.2852 | 1023.0236 | 455 | 379.6745 | 248.0738 | 812.8152 | 1767.1878 | 946.7148 | 842.7247 | 686.1947 | 1808.0837 | 1069.8319 | 702.9932 | 815.057 | 1961.9597 | 1100.0236 | 977.9932 | 1082.5605 | 997.4798 | 558.6175 | 467.1209 | 505.4228 | 1724.5369 | 954.3422 | 864.7316 | 1426.1779 | 1887.3453 | 1183.4395 | 1181.0373 | 1332.7582 | 2155.1544 | 1592.5536 | 933.0137 | 1619.7954 | 2158.5536 | 1614.1878 | 853.9833 | 1801.6243 | 1134 | 796.8053 | 558.2684 | 898.7414 | 757.6776 | 568.6038 | 481.4798 | 448.1141 | 1498.8357 | 1003.4297 | 622.6008 | 1059.4129 | 3913.2145 | 2303.2753 | 1571.1878 | 2363.9833 | 3185.1612 | 2148.7414 | 1501.6646 | 1978.5468 | 381.4297 | 211.3087 | 183.5202 | 381.4297 | 606.0639 | 531.057 | 309.349 | 586.1947 | 569 | 482.3658 | 317.3087 | 466.8053 | 357.6008 | 309.8624 | 200.5605 | 369.4798 | 378.2281 | 375.5605 | 244 | 275.3825 | 2564.1407 | 1547.9658 | 890.057 | 1100.7818 | 1290.4631 | 968.6844 | 701.9764 | 726.1711 | 2833.2852 | 1459.3186 | 1218.2418 | 1502.5772
426 | 6.6804 | 1179.9724 | 1323.9479 | 970.7896 | 1308.5271 | 1717.7242 | 1131.7385 | 789.3544 | 1175.4474 | 873.9734 | 497.1185 | 379.763 | 652.3544 | 1980.6047 | 964.9734 | 723.7232 | 1396.619 | 1735.8029 | 958.7641 | 657.7763 | 1233.0909 | 2098.7375 | 1871.7641 | 1219.3422 | 1466.6844 | 2703.7763 | 1744.9612 | 983.5393 | 1760.8427 | 1491.1451 | 1083.7763 | 595.4474 | 1064.619 | 1649.0266 | 1421.856 | 729.4331 | 489.1971 | 3880.7109 | 2106.2247 | 1314.9745 | 1959.9724 | 2145.0133 | 1380.8427 | 730.2237 | 1009.0797 | 448.5128 | 368.9867 | 253.1307 | 859.7354 | 1794 | 938.856 | 883.474 | 682.5138 | 1876.9602 | 1149.4076 | 740.6445 | 838.6854 | 1975.0133 | 1083.619 | 1001.3279 | 1074.0531 | 992.4862 | 551.1451 | 442.2768 | 504.5393 | 1721.2492 | 979.5659 | 858.6578 | 1423.3964 | 2001.9602 | 1210.5138 | 1090.8958 | 1408.8427 | 2155.3555 | 1584.7875 | 872.2513 | 1592.093 | 2195.9745 | 1642.1185 | 836.3156 | 1758.0664 | 1147.8693 | 825.7641 | 566.3555 | 973.2646 | 832.8018 | 631.2891 | 486.3289 | 452.4208 | 1451.5924 | 966.9081 | 599.9214 | 1005.6711 | 4021.9612 | 2265.6578 | 1596.4341 | 2378.5281 | 3162.0276 | 2205.8161 | 1453.1063 | 2052.2768 | 343.5659 | 208.3555 | 176.6578 | 343.5659 | 568.3289 | 534.7763 | 299.8294 | 577.5924 | 561.8417 | 480.6711 | 313.237 | 480.7763 | 333.8029 | 324.6445 | 188.2503 | 382.382 | 388.6312 | 356.763 | 240.4208 | 292.474 | 2633.1573 | 1713.8529 | 884.6047 | 1149.0531 | 1285.474 | 1021.5924 | 721.5128 | 743.3688 | 2858.8161 | 1514.1451 | 1162.9612 | 1506.8161
427 | 6.6914 | 1129.8909 | 1353.5841 | 985.4956 | 1364.2714 | 1748 | 1207.233 | 731.6165 | 1186.2036 | 870.1475 | 506.1918 | 377.4277 | 609.4838 | 1952.2478 | 941.7964 | 693.5841 | 1434.9558 | 1742.3923 | 1012.5959 | 649.6843 | 1202.2302 | 2152.2832 | 1918.0442 | 1223.056 | 1489.6874 | 2708.8997 | 1796.6874 | 983.7316 | 1819.5429 | 1514.6637 | 1101.1712 | 603.6077 | 1099.1799 | 1675.2714 | 1466.3038 | 670.8997 | 487.0915 | 3902.8201 | 2159.1563 | 1360.472 | 1924.9942 | 2166.1593 | 1420.4277 | 724.8643 | 1070.5871 | 437.0915 | 362.472 | 213.7316 | 808.0266 | 1792.5841 | 964.8967 | 890.9764 | 724.8997 | 1863.0797 | 1119.1445 | 721.4484 | 907.944 | 1996.3953 | 1111.1003 | 931.0324 | 1115.9233 | 959.7758 | 564.9882 | 473.3599 | 499.3038 | 1714.5222 | 985 | 842.4365 | 1506.5044 | 1979.5841 | 1221.9852 | 1146.8643 | 1437.8082 | 2160.6283 | 1517.1299 | 953.7876 | 1639.1121 | 2245.7198 | 1654.0236 | 837.1475 | 1794.4159 | 1181.8407 | 859.7728 | 568.3245 | 1037.0088 | 787.944 | 607.528 | 478.8082 | 445.1918 | 1483.2714 | 981.2478 | 608.7404 | 1017.9115 | 4044.8967 | 2293.3305 | 1584.6283 | 2472.5044 | 3235 | 2216.3245 | 1495.3126 | 2074.8643 | 355.1357 | 220.236 | 172.708 | 355.1357 | 577.0915 | 538.7198 | 325.1121 | 597.236 | 556.3511 | 490.3157 | 317.9321 | 482.3599 | 346.764 | 317.2478 | 205.5959 | 410.8319 | 368.1121 | 345.2242 | 223.7522 | 309.1799 | 2631.0945 | 1595.2272 | 871.1151 | 1155.997 | 1317.7552 | 1028.7286 | 709.6195 | 760.1121 | 2856.581 | 1542.6195 | 1196.5162 | 1516.1445
428 | 6.7024 | 1191.3377 | 1356.4484 | 1000.586 | 1409.7104 | 1745.2689 | 1155.1451 | 719.1038 | 1227.4209 | 906.4759 | 531.9381 | 391.0275 | 630.7036 | 2002.4972 | 927.6416 | 690.9862 | 1458.4484 | 1721.4966 | 1027.7036 | 630.469 | 1321.7792 | 2204.5034 | 1884.3721 | 1206.5929 | 1507.1858 | 2724.7724 | 1815.6754 | 1018.4553 | 1881.8687 | 1536.4759 | 1162.9794 | 619.7861 | 1125.8418 | 1718.4759 | 1492.1451 | 702.414 | 534.1726 | 3931.2764 | 2145.1932 | 1358.7724 | 2015.9719 | 2212.4966 | 1445.945 | 716.6692 | 1119.848 | 477.9656 | 376.6623 | 233.3102 | 882.1376 | 1791.7242 | 922.7036 | 843.3996 | 745.9862 | 1886.2483 | 1080.2758 | 686.5722 | 919.6898 | 2047.9656 | 1093.738 | 906.2208 | 1161.2483 | 977.4622 | 558.207 | 477.7724 | 520.4278 | 1808.9375 | 973.0826 | 818.1588 | 1495.1382 | 1968.2964 | 1166.3246 | 1114.0894 | 1433.1657 | 2132.3033 | 1641.8756 | 955.4134 | 1631.5103 | 2262.738 | 1664.6066 | 860.5654 | 1805.7036 | 1152.3865 | 809.9587 | 540.4759 | 983.9862 | 795.2208 | 606.4966 | 471.4346 | 464.9794 | 1511.8274 | 992.0206 | 589.2001 | 1073.09 | 4032.4972 | 2410.9106 | 1552.3308 | 2461.883 | 3227.055 | 2194.1864 | 1448.2689 | 2051.5241 | 378.7242 | 227.9518 | 173.262 | 378.7242 | 619.7036 | 556.7311 | 315.2758 | 604.207 | 614.0344 | 518.469 | 315.4484 | 487.7655 | 352.7036 | 330.5034 | 207.5447 | 409.8205 | 363.738 | 328.531 | 216.7036 | 304.3102 | 2753.1244 | 1656.7104 | 904.6961 | 1106.3727 | 1376.2208 | 1005.883 | 740.7724 | 753.7517 | 2785.2414 | 1568.5585 | 1188.207 | 1495.152
429 | 6.7134 | 1156.3874 | 1399.064 | 951.0137 | 1441.592 | 1747.6834 | 1197.0914 | 777.2674 | 1254.1829 | 929.2468 | 554.8309 | 374.3097 | 663.2468 | 2093.7177 | 965.3589 | 693.352 | 1498.9794 | 1766.5154 | 1042.2674 | 651.5566 | 1323.3314 | 2209.5143 | 1860.4445 | 1187.3668 | 1458.5714 | 2737.4091 | 1765.8388 | 1037.4789 | 1837.5566 | 1552.472 | 1175.1269 | 594.88 | 1187.9646 | 1736.0354 | 1504.1258 | 736.0914 | 559.1406 | 4040.2046 | 2184.0708 | 1344.3097 | 1948.9657 | 2230.1554 | 1478.5988 | 754.3451 | 1080.9086 | 487.7886 | 413.4366 | 216.3806 | 847.0148 | 1798.3097 | 941.6971 | 805.36 | 754.8663 | 1900.8103 | 1104.6982 | 665.4228 | 954.3383 | 2065.3451 | 1085.7109 | 919.2606 | 1140.3726 | 999.472 | 572.8309 | 469.824 | 542.3589 | 1761.1348 | 944.5566 | 865.5566 | 1546.6057 | 1951.3874 | 1231.2743 | 1089.8526 | 1472.6057 | 2147.5645 | 1592.9931 | 900.416 | 1635.0777 | 2278.6834 | 1628.4228 | 847.1131 | 1828.6057 | 1204.3451 | 826.6971 | 550.4789 | 970.4572 | 804.352 | 602.528 | 490.3589 | 488.0708 | 1486.8594 | 998.1131 | 600.8377 | 1136.4857 | 4138.5703 | 2430.7177 | 1542.0148 | 2526.5497 | 3211.0777 | 2209.7532 | 1459.2823 | 2060.0217 | 380.0983 | 240.4012 | 169.0423 | 380.0983 | 626.4503 | 564.4789 | 307.9086 | 612.8377 | 599.7817 | 528.0914 | 333.0491 | 494.2891 | 379.2537 | 325.0914 | 193.9931 | 378.2183 | 371.8663 | 311.6126 | 232.8309 | 280.6057 | 2809.0423 | 1672.176 | 841.6697 | 1182.5851 | 1389.7817 | 1046.5772 | 731 | 759.2537 | 2816.36 | 1550.296 | 1191.1052 | 1529.6617
430 | 6.7244 | 1164.6357 | 1435.5353 | 924.3673 | 1480.5442 | 1785.3673 | 1184.9086 | 771.0959 | 1249.736 | 942.4439 | 540.8171 | 354.7242 | 710.8245 | 2050.7331 | 930.9056 | 671.5398 | 1546.7257 | 1876.084 | 1017.8141 | 657.4469 | 1356.6327 | 2264.5635 | 1920.5339 | 1257.823 | 1541.177 | 2807.003 | 1840.09 | 1038.7286 | 1902.3851 | 1579.7331 | 1200.0989 | 565.9985 | 1156.8127 | 1742.3569 | 1432.5576 | 718.997 | 552.8275 | 4102.8334 | 2224.8186 | 1344.0974 | 2018.174 | 2295.1003 | 1524.2832 | 759.3628 | 1105.0974 | 466.4513 | 401.9056 | 237.5428 | 912.5428 | 1793.4587 | 953.826 | 883.8186 | 770.0944 | 1976.2907 | 1193.7212 | 705.9145 | 974.4528 | 2070.6357 | 1070.6476 | 916.1799 | 1107.3777 | 1022.0944 | 556.9071 | 477.1829 | 519.6387 | 1868.3688 | 948.5368 | 892.1844 | 1549.6342 | 1978.4632 | 1194.1799 | 1130.09 | 1489.823 | 2257 | 1628.8275 | 951.0929 | 1662.6387 | 2300.2684 | 1659.3643 | 879.0855 | 1845.823 | 1213.4558 | 824.91 | 533.174 | 936.5561 | 795.91 | 578.8141 | 465.7286 | 512.9926 | 1519.3673 | 1027.9026 | 571.7271 | 1122.6387 | 4109.6283 | 2389.3702 | 1602.0855 | 2577.5517 | 3239.1844 | 2187.5517 | 1501.9941 | 2107.4513 | 359.0944 | 221.4572 | 170.9086 | 359.0944 | 584.4528 | 563 | 309.9955 | 598.1888 | 602.2729 | 526.2773 | 317.3628 | 531.4543 | 387.9955 | 316.7286 | 211.8186 | 393.1903 | 374.5428 | 308.8216 | 217.7257 | 288.5457 | 2808.9985 | 1675.1859 | 887.7182 | 1224.6283 | 1380.8127 | 1015.6357 | 728.2714 | 773.9985 | 2887.4513 | 1583.5517 | 1164.5546 | 1514.1859
431 | 6.7354 | 1176.9686 | 1384.1367 | 941.4336 | 1485.4061 | 1802.1485 | 1212.5114 | 774.9026 | 1296.7148 | 902.3943 | 529.0118 | 342.5625 | 725.8947 | 2070.0236 | 920.8397 | 644.1367 | 1536.8594 | 1864.9489 | 996.8555 | 626.5507 | 1340.4218 | 2352.0196 | 1899.6324 | 1260.0471 | 1519.5625 | 2836.117 | 1830.1524 | 1027.465 | 1996.0432 | 1601.0196 | 1221.6135 | 560.2812 | 1144.6795 | 1725.9607 | 1516.2498 | 695.586 | 590.1681 | 4165.6096 | 2250.9529 | 1372.4415 | 2010.5035 | 2330.3166 | 1563.0353 | 747.8712 | 1157.1092 | 463.2577 | 372.1642 | 220.4336 | 875.4729 | 1820.5625 | 987.7306 | 884.5742 | 799.8358 | 2075.2812 | 1150.1642 | 738.1367 | 955.3048 | 2080.117 | 1127.293 | 911.1328 | 1176.4415 | 1038.7188 | 555.9804 | 482.4297 | 527.7227 | 1885.7345 | 925.6677 | 907.9961 | 1550.8319 | 2028.4218 | 1196.2616 | 1125 | 1508.0157 | 2257.8555 | 1662.176 | 955.4415 | 1711.7848 | 2313.7848 | 1685.5271 | 856.5585 | 1856.6017 | 1221.4258 | 843.6874 | 503.5467 | 999.7031 | 812.9764 | 566.9804 | 472.7148 | 480.1288 | 1549.8358 | 1010.1092 | 574.1367 | 1157.2419 | 4120.4643 | 2411.883 | 1594.6716 | 2607.293 | 3248.4375 | 2256.0738 | 1448.3284 | 2092.5664 | 376.0039 | 230.2969 | 170.2812 | 376.0039 | 584.4101 | 571.2694 | 286 | 615.1838 | 603.8555 | 560.4101 | 313.8555 | 523.7266 | 377.4022 | 319.1524 | 208.0118 | 436.1485 | 369.4101 | 328.9961 | 216.1288 | 294.5625 | 2813.2616 | 1681.6017 | 848.9804 | 1196.6834 | 1369.8201 | 1009.4336 | 724.707 | 781.3983 | 2874.8476 | 1618.7109 | 1206.8712 | 1556.2459
432 | 6.7464 | 1202.7227 | 1426.5035 | 941.8412 | 1529.7645 | 1836.0473 | 1261.9326 | 733.3594 | 1304.6259 | 932.9125 | 525.0054 | 367.0875 | 678 | 2088.7809 | 947.1642 | 661.5182 | 1559.5802 | 1928.5128 | 1025.4725 | 662.8157 | 1349.8923 | 2346.6313 | 1987.938 | 1200.8923 | 1569.3703 | 2881.7901 | 1815.6204 | 971.1843 | 1952.8978 | 1566.1843 | 1172.4617 | 564.8102 | 1198.8157 | 1788.1952 | 1560.5128 | 692.1697 | 577.3339 | 4103.615 | 2307.4307 | 1357.7847 | 2074.6917 | 2282.3741 | 1549.2316 | 727.0256 | 1190.7792 | 490.0767 | 360.1697 | 216.0821 | 846.0674 | 1876.3192 | 971.1642 | 875.0511 | 817.1843 | 2054.23 | 1109.0201 | 744.5128 | 928.0256 | 2125.1952 | 1106.6971 | 927.5693 | 1162.677 | 1029.1332 | 574.1588 | 470.8869 | 520.5693 | 1883.2007 | 986.4872 | 909.7281 | 1605.0109 | 2048.0054 | 1225.5128 | 1119.3486 | 1498.1131 | 2253.7537 | 1625.9489 | 930.6716 | 1780.7993 | 2392.9125 | 1724.1533 | 872.1077 | 1837.9288 | 1244.0364 | 869.969 | 532.323 | 1019.0566 | 839.7281 | 603.0054 | 472.2974 | 517.6313 | 1560.6406 | 1053.5949 | 596.5748 | 1204.3741 | 4198.9984 | 2404.5802 | 1648.1077 | 2627.7444 | 3256.3339 | 2320.2045 | 1412.4415 | 2101.677 | 380.8157 | 216.8923 | 169.7537 | 380.8157 | 605.0256 | 574.7482 | 287.7847 | 565.9744 | 601.5383 | 580.1332 | 318.0821 | 513.354 | 398.8048 | 319.4927 | 197.4617 | 417.3997 | 394.4872 | 329.2409 | 233.7026 | 306.2974 | 2864.8211 | 1673.6368 | 885.3029 | 1241.0566 | 1408.0767 | 1029.1751 | 723.8869 | 795.0147 | 2883.7281 | 1600.3942 | 1192.8668 | 1593.3176
433 | 6.7574 | 1178.6018 | 1486.6283 | 956.7611 | 1554.9203 | 1923.0885 | 1215.3452 | 761.469 | 1329.0266 | 918.7257 | 542.6371 | 388.0885 | 676.4513 | 2191.9203 | 968.1505 | 671.6106 | 1594.292 | 1904.2477 | 1091.4071 | 660.292 | 1360.3629 | 2412.8231 | 1986.8585 | 1222.2036 | 1656.5929 | 2887.8406 | 1857.3895 | 968.1328 | 1986.5752 | 1581.4072 | 1204.9912 | 564.7876 | 1226.646 | 1810.1858 | 1559.1681 | 727.0708 | 622.3805 | 4115.9824 | 2286.3009 | 1360.4513 | 2081.1594 | 2299.478 | 1613.3982 | 721.2389 | 1160.3274 | 469 | 407.4602 | 219.1416 | 933.5399 | 1948.7875 | 992.4602 | 860.3805 | 794 | 2017.3275 | 1099.8496 | 730.469 | 922.2389 | 2142.8495 | 1081.2389 | 914.6283 | 1169.9292 | 1047.5664 | 563.885 | 451.0266 | 552.5399 | 1933.0442 | 990.3097 | 890.2301 | 1639.3451 | 2082.9823 | 1239.0354 | 1092.0619 | 1530.6726 | 2226.6106 | 1661.6815 | 904.1858 | 1764.7169 | 2363.5487 | 1701.8231 | 869.0708 | 1927.4337 | 1267.3185 | 867.1062 | 526.3097 | 1030.3805 | 824.8761 | 607.3185 | 477.0266 | 545.4159 | 1532.531 | 1050.354 | 618.9203 | 1172.8495 | 4102.5753 | 2427.8318 | 1637.9469 | 2684.0619 | 3294.2566 | 2264.6195 | 1468.6991 | 2141.7611 | 394.3982 | 222.0974 | 161.1947 | 394.3982 | 614.416 | 555.9912 | 296.0266 | 578.5752 | 580.7787 | 575.0266 | 317.4248 | 521.6637 | 371.8584 | 337.2301 | 214.8142 | 409.8231 | 400.4779 | 324.1947 | 231.1416 | 306.6903 | 2929.4868 | 1729.584 | 886.9114 | 1265.3894 | 1388.8584 | 1080.469 | 705.5752 | 755.8231 | 2914.7169 | 1552.2655 | 1202.9292 | 1557.9026
434 | 6.7684 | 1206.1185 | 1466.3722 | 963.3388 | 1555.0502 | 1897.7797 | 1263.9828 | 732.8476 | 1358.6097 | 978.8639 | 534.9489 | 365.678 | 674.9322 | 2192.0502 | 983.6612 | 643.3559 | 1563.0841 | 1876.6775 | 1106.4066 | 608.1524 | 1383.5084 | 2452.4916 | 2078.0167 | 1244.0683 | 1651.8982 | 2869.7449 | 1956.2374 | 995.2886 | 1999.9661 | 1651.2207 | 1253.0339 | 580.9828 | 1221.8648 | 1802.7286 | 1581.7114 | 708.2546 | 629.6947 | 4223.0167 | 2302.7617 | 1371.1692 | 2239.0841 | 2357.2379 | 1614.2198 | 718.2035 | 1168.6943 | 469.9661 | 411.7463 | 223.3727 | 962.1701 | 1898.2537 | 1001.2546 | 853.2035 | 782.085 | 2078.8982 | 1144.3053 | 700.8815 | 925.9661 | 2159.8634 | 1082.0678 | 871.1863 | 1187.4573 | 1089.3564 | 560.8987 | 451.3053 | 600.8137 | 1947.7621 | 1003.237 | 872.0507 | 1637.033 | 2094.6273 | 1250.5938 | 1095.7788 | 1551.6101 | 2235.067 | 1689.7123 | 902.2031 | 1864.8471 | 2392.1181 | 1755.0167 | 873.4405 | 2002.6784 | 1216.1018 | 893.3392 | 511.2207 | 1013.5427 | 831.5418 | 574.1352 | 475.3731 | 499.8137 | 1574.6775 | 1024.2207 | 596.8304 | 1180.2026 | 4184.9493 | 2388.4736 | 1623.8643 | 2662.0167 | 3285.0507 | 2265.3559 | 1481.2207 | 2195.4242 | 370.4238 | 227.0678 | 176.2881 | 370.4238 | 645.5256 | 609.5084 | 303.7119 | 604.2031 | 588.2705 | 566.2886 | 301.4577 | 507.8137 | 368.9154 | 329.6775 | 224.5256 | 449.4916 | 406.9661 | 345.0846 | 229.5762 | 305.3559 | 2955.9493 | 1740.6599 | 848.3894 | 1266.5938 | 1420.3722 | 1051.5256 | 720.2542 | 803.2203 | 2961.0687 | 1577.8471 | 1175.3736 | 1572.5758
435 | 6.7794 | 1249.0886 | 1494.5801 | 971.6332 | 1614.3402 | 1909.7218 | 1273.2989 | 733.7011 | 1394.645 | 1006.2723 | 555.6598 | 362.6598 | 710.7602 | 2235.293 | 986.3993 | 653.0325 | 1601.6184 | 1912.3225 | 1100.6037 | 638.0561 | 1408.7129 | 2458.3786 | 2090.3137 | 1200 | 1671.0384 | 2933.2989 | 1946.0118 | 970.3697 | 2046.7985 | 1628.3638 | 1262.7011 | 599.6598 | 1197.0207 | 1792.9232 | 1623.6804 | 679.352 | 640.9793 | 4236.9852 | 2366.5771 | 1390.9705 | 2321.0797 | 2309.352 | 1656.3196 | 700.9852 | 1224.6598 | 480.6923 | 383.7159 | 233.3107 | 920.3579 | 1969.4111 | 992.7454 | 835.3166 | 757.6686 | 2070.6244 | 1146.7277 | 702.0413 | 945.0502 | 2254.0177 | 1084.6863 | 876.3225 | 1224.6804 | 1089.4554 | 537.6982 | 430.3255 | 574.6155 | 1997.9882 | 1014.6244 | 895.0177 | 1701.293 | 2090.3727 | 1188.9823 | 1165.9911 | 1539.6214 | 2316.3491 | 1661.3756 | 944.0413 | 1910.7159 | 2463.0531 | 1798.7395 | 875.9675 | 1995.0974 | 1213.3432 | 883.0295 | 504.7454 | 1008.7483 | 853.9262 | 599.9823 | 457.0118 | 509.3875 | 1580.5683 | 1014.0679 | 602.6539 | 1219.9173 | 4198.4229 | 2451.9173 | 1611.941 | 2702.062 | 3267.231 | 2253.9705 | 1460.3697 | 2171.062 | 379.71 | 225.6745 | 183.0118 | 379.71 | 650.3816 | 606.9646 | 310.0266 | 634.6745 | 631.6066 | 547.7218 | 286.9793 | 513.3756 | 363.6952 | 350.9469 | 205.9793 | 431.9764 | 409.6686 | 369.6952 | 226.9764 | 292.9675 | 3006.8664 | 1834.9823 | 876.9764 | 1234.0679 | 1461.9528 | 1029.3018 | 713.2989 | 814.6804 | 2878.003 | 1619.0354 | 1137.7395 | 1615.9941
436 | 6.7904 | 1269.8332 | 1483.4854 | 961.4779 | 1647.4219 | 1961.5576 | 1237.5221 | 750.761 | 1388.3186 | 974.3584 | 549.6136 | 363.5457 | 762.2153 | 2242.6608 | 1012.9764 | 674.6534 | 1631.3895 | 1964.27 | 1059.733 | 646.8849 | 1434.8009 | 2507.0236 | 2055.8569 | 1197.9203 | 1697.6932 | 2950.8998 | 1990.3939 | 990.8407 | 2127.3348 | 1662.5177 | 1291.1991 | 591.5339 | 1220.705 | 1786.4057 | 1623.295 | 677.7891 | 616.6017 | 4302.9073 | 2337.1948 | 1361.295 | 2312.6105 | 2349.7287 | 1671.9558 | 687.1873 | 1227.6254 | 493.1475 | 427.0516 | 223.5457 | 952.944 | 2019.5737 | 1029.3628 | 835.4381 | 780.8363 | 2063.4897 | 1182.8009 | 716.1357 | 1001.1829 | 2266.6932 | 1083.4424 | 889.1033 | 1258.2626 | 1146.0044 | 561.7729 | 425.3466 | 546.239 | 1965.3908 | 981.1475 | 890.3628 | 1682.3186 | 2138.4897 | 1222.4057 | 1139.2035 | 1542.621 | 2339.4779 | 1729.86 | 960.5619 | 1909.6858 | 2447.3553 | 1833.0958 | 839.7891 | 2033.2713 | 1243.2228 | 903.3466 | 556.9602 | 1033.1518 | 816.1313 | 597.705 | 477.8245 | 540.8009 | 1517.1195 | 1073.5177 | 610.9838 | 1173.705 | 4299.2066 | 2450.4175 | 1581.3584 | 2760.6372 | 3270.4412 | 2288.4175 | 1470.4424 | 2201.2153 | 400.6814 | 240.0118 | 176.0958 | 400.6814 | 669.1357 | 580.2271 | 320.3746 | 637.2271 | 613.8762 | 577.7212 | 299.6889 | 527.4424 | 378.8407 | 313.267 | 199.2788 | 427.7847 | 421.7449 | 384.1475 | 223.1313 | 286.2508 | 3079.2639 | 1800.8171 | 864.4661 | 1306.3422 | 1429.6534 | 1027.4101 | 706.2906 | 835.0915 | 2894.9042 | 1648.8923 | 1181.8009 | 1616.1593
437 | 6.8014 | 1229.7366 | 1572.6411 | 954.4935 | 1671.5702 | 2001.5615 | 1224.0492 | 740.7945 | 1416.2026 | 1008.1505 | 562.4327 | 369.6932 | 722.4139 | 2314.4602 | 981.1534 | 681.5369 | 1721.0058 | 2060.5673 | 1079.1505 | 614.9696 | 1453.712 | 2534.1809 | 2052.6078 | 1171.3893 | 1724.123 | 3035.178 | 2039.2605 | 974.0275 | 2137.9175 | 1684.411 | 1312.8437 | 564.5427 | 1240.7945 | 1862.3097 | 1612.8958 | 662.8437 | 611.7149 | 4496.2459 | 2390.6165 | 1341.5644 | 2252.7828 | 2435.7366 | 1720.7945 | 700.8162 | 1258.5094 | 487.2301 | 440.2576 | 227.1809 | 1006.6874 | 2021.8191 | 976.8249 | 855.9233 | 830.9725 | 2112.6136 | 1165.822 | 674.9262 | 1043.3589 | 2314.298 | 1098.709 | 908.3372 | 1267.7511 | 1096.2822 | 576.6903 | 442.5094 | 581.6874 | 1968.5281 | 970.9233 | 863.3068 | 1691.8987 | 2168.233 | 1286.1071 | 1109.3314 | 1614.6136 | 2389.1997 | 1819.7149 | 931.1042 | 1911.0375 | 2370.4573 | 1824.0217 | 815.1534 | 2018.6844 | 1286.6411 | 899.3343 | 552.8466 | 1004.7887 | 861.2518 | 613.1288 | 496.2576 | 533.1534 | 1561.0433 | 1123.7641 | 622.9016 | 1202.767 | 4391.5094 | 2558.7424 | 1599 | 2785.8987 | 3384.7482 | 2382.0275 | 1458.4327 | 2219.5556 | 376.7453 | 245.8741 | 160.9204 | 376.7453 | 639.0521 | 592.945 | 307.0767 | 632 | 643.7207 | 574.0492 | 329.9016 | 493.5398 | 384.6382 | 325.4327 | 208.3343 | 459.3314 | 434.1042 | 369.6165 | 242.411 | 316.0767 | 3035.2922 | 1770.8437 | 865.2576 | 1327.288 | 1425.0521 | 1050.7207 | 733.0521 | 848.822 | 2945.9696 | 1672.5615 | 1173.4356 | 1610.0029
438 | 6.8124 | 1298.1662 | 1570.144 | 927.4681 | 1644.9833 | 2047.7562 | 1237.1745 | 745.1136 | 1409.7895 | 1034.0305 | 589.2576 | 372.7645 | 709.9307 | 2329.7119 | 992.0526 | 677.554 | 1706.6343 | 2029.6593 | 1134.7257 | 630.759 | 1502.0305 | 2535.3629 | 2087.5624 | 1187.7922 | 1774.5679 | 3078.4376 | 2004.7424 | 984.3545 | 2167.4405 | 1683.1967 | 1348.4764 | 540.2493 | 1251.4238 | 1871.7257 | 1657.8088 | 696.9917 | 641.7119 | 4546.6538 | 2382.856 | 1320.0693 | 2346.349 | 2489.6898 | 1750.3545 | 737.6288 | 1275.3795 | 497.4543 | 447.0748 | 219.0831 | 1042.3186 | 2010.9862 | 957.9307 | 858.8255 | 842.9169 | 2136.5374 | 1181.9529 | 663.9612 | 1064.4155 | 2385.1219 | 1178.3407 | 877.9695 | 1230.4681 | 1102.795 | 571.7286 | 463.4626 | 608.0388 | 2093.0055 | 976.0526 | 883.2493 | 1749.6703 | 2163.1052 | 1268.9002 | 1128.2271 | 1638.1662 | 2415.2133 | 1849.7119 | 914.7729 | 1988.5402 | 2432.9833 | 1855.7729 | 837.5595 | 2081.9169 | 1302.3324 | 906.9612 | 541.205 | 1063.1967 | 901 | 597.2355 | 464.0998 | 538.4848 | 1615.2659 | 1154.2659 | 593.6814 | 1214.9474 | 4418.4017 | 2555.6371 | 1609.0221 | 2793.1883 | 3363.9307 | 2319.23 | 1477.0914 | 2288.6593 | 368.3102 | 240.4764 | 183.2576 | 368.3102 | 635.3407 | 623.144 | 308.2576 | 636.0831 | 622.9695 | 609.0748 | 289.9169 | 503.5291 | 393.7812 | 353.7424 | 206.3102 | 458.1829 | 408.4931 | 365.5679 | 247.8781 | 306.8643 | 3135.7064 | 1845.0803 | 863.1662 | 1315.3462 | 1446.2105 | 1035.9086 | 715.2355 | 851.59 | 2960.2743 | 1691.2881 | 1165.2355 | 1607.6426
439 | 6.8234 | 1313.1504 | 1586.1741 | 971.6991 | 1708.5576 | 2074.1947 | 1248.1741 | 747.767 | 1389.6107 | 1036.6696 | 581.4808 | 370.292 | 740.8496 | 2367.062 | 1018.1888 | 652.6018 | 1761.6313 | 2042.0236 | 1157.7581 | 607.1741 | 1475.1415 | 2573.7227 | 2055.5576 | 1248.7197 | 1790.5192 | 3151.0915 | 2004.0826 | 1009.4424 | 2161.6933 | 1706.9558 | 1354.5929 | 563.6844 | 1276.6549 | 1920.0679 | 1737.9942 | 697.525 | 664.4897 | 4511.062 | 2446.2066 | 1279.9469 | 2325.0147 | 2487.4366 | 1748.2152 | 716.292 | 1266.2714 | 494.2625 | 469.4808 | 222.1652 | 1045.5723 | 2004.5959 | 993.0679 | 869.6844 | 836.767 | 2164.0384 | 1201.0884 | 693.3009 | 1058.7344 | 2349.2773 | 1167.1268 | 882.233 | 1275.8496 | 1142.4602 | 559.3098 | 461.6549 | 588.2862 | 2142.4956 | 1004.1062 | 901.3156 | 1828.4188 | 2185.8496 | 1272.2536 | 1146.9558 | 1667.3393 | 2375.0147 | 1840.6607 | 940.9116 | 1939.5576 | 2494.6402 | 1859.6696 | 867.8496 | 2102.9942 | 1315.0442 | 898.7197 | 516.879 | 1065.4808 | 868.7875 | 595.1062 | 443.4897 | 529.1121 | 1579.3451 | 1139.8201 | 606.5339 | 1214.121 | 4391.7581 | 2527.2478 | 1639.0384 | 2843.1741 | 3346.5309 | 2317.0798 | 1466.6018 | 2236.2152 | 392.4219 | 252.3451 | 169.7286 | 392.4219 | 636.0147 | 620 | 309.6844 | 630.7286 | 617.2625 | 614.2241 | 298.7227 | 547.2862 | 366.0826 | 354.6165 | 216.6165 | 421.4808 | 404.7227 | 378.8348 | 263.3982 | 295.6696 | 3167.6668 | 1888.2152 | 893.4897 | 1352.4719 | 1499.7522 | 1072.475 | 697.3835 | 883.7522 | 2974.8231 | 1645.5014 | 1165.4071 | 1683.9647
440 | 6.8344 | 1304.3141 | 1599.1902 | 952.5649 | 1703.9847 | 2053.9789 | 1295.6225 | 744.2507 | 1452.2055 | 1045.438 | 568.6254 | 346.4985 | 730.293 | 2452.6195 | 1026.7522 | 667.5197 | 1779.9125 | 2106.2238 | 1135.4563 | 626.9184 | 1454.6436 | 2662.7194 | 2148.7405 | 1195.562 | 1799.0211 | 3196.4774 | 2019.2055 | 991.4563 | 2221.2055 | 1716.7522 | 1366.9519 | 566.8127 | 1320.7464 | 1957.8491 | 1778.2507 | 696.6407 | 683.0845 | 4623.5321 | 2516.4832 | 1277.6859 | 2348.266 | 2511.2901 | 1735.1363 | 678.2507 | 1288.7704 | 480.562 | 454.6465 | 237.7675 | 1066.6254 | 2076.7886 | 990.0845 | 876.3746 | 844.3324 | 2216.266 | 1198.9096 | 713.2267 | 1036.3477 | 2360.0153 | 1130.5015 | 888.9155 | 1282.3112 | 1158.879 | 581.7945 | 442.027 | 602.3746 | 2095.5044 | 990.8185 | 899.3746 | 1782.2478 | 2145.6101 | 1252.0364 | 1179.3112 | 1664.6917 | 2355.9184 | 1814.0211 | 960.9001 | 1983.6647 | 2514.6888 | 1851.4198 | 858.8761 | 2119.1086 | 1292.5831 | 884.7434 | 527.8914 | 1077.9548 | 851.828 | 620.2085 | 480.2901 | 521.8732 | 1614.8039 | 1128.6705 | 612.9577 | 1226.855 | 4388.4534 | 2504.5437 | 1622.0058 | 2847.4832 | 3327.5532 | 2434.6676 | 1452.6283 | 2201.7646 | 394.6071 | 235.1662 | 163.852 | 394.6071 | 653.3295 | 620 | 303.3141 | 632.3717 | 616.6225 | 579.562 | 329.145 | 540.0029 | 370.9155 | 343.7099 | 209.6676 | 436.3295 | 428.4169 | 388.9577 | 233.8367 | 306.8127 | 3295.9184 | 1812.6042 | 909.3141 | 1312.9155 | 1501.8761 | 1083.2536 | 725.3083 | 893.3958 | 3036.5678 | 1624.3477 | 1183.1268 | 1751.3324
441 | 6.8454 | 1314.9322 | 1594.4435 | 948.5871 | 1679.6048 | 2071.6903 | 1309.9823 | 726.3903 | 1502.8516 | 1033.0629 | 579.2822 | 349.5742 | 723.121 | 2522.0984 | 1040.9096 | 658.5968 | 1835.6903 | 2132.6597 | 1150.2871 | 620.6097 | 1490.4661 | 2752.2468 | 2205.8662 | 1219.7709 | 1822.9322 | 3207.0629 | 2070.6677 | 1008.3274 | 2240.8387 | 1759.4741 | 1422.8387 | 562.5516 | 1338.6371 | 2017.3855 | 1785.6903 | 710.8662 | 655.0855 | 4703.4936 | 2507.0128 | 1303.7935 | 2407.0177 | 2490.6823 | 1818.5742 | 688.5467 | 1284.4307 | 502.7306 | 454.529 | 259.1662 | 1038.1081 | 2112.3048 | 967.7984 | 851.4355 | 851.5516 | 2247.2694 | 1248.1887 | 734.3274 | 1093.6549 | 2443.9725 | 1127.4258 | 903.0403 | 1274.1887 | 1140.6274 | 552.7758 | 415.2774 | 590.9016 | 2089.9499 | 965.6854 | 881.3726 | 1810.529 | 2161.6757 | 1307.2645 | 1157.7226 | 1658.9096 | 2352.8742 | 1859.5596 | 905.7581 | 1975.6597 | 2538.3177 | 1825.7129 | 843.5113 | 2136.6176 | 1304.7935 | 929.1887 | 564.2468 | 1112.1839 | 912.7532 | 634.0984 | 487.8387 | 524.9016 | 1674.7 | 1098.9774 | 594.5162 | 1212.879 | 4421.2694 | 2537.0483 | 1567.6726 | 2846.9902 | 3393.0678 | 2425.8161 | 1418.7758 | 2302.358 | 375.6097 | 240 | 173.1662 | 375.6097 | 683.0226 | 622.0403 | 293.9371 | 691.1161 | 648.121 | 575.2065 | 324.35 | 513.8564 | 388.7129 | 338.2016 | 195.9194 | 460.7178 | 434.9774 | 399.0806 | 218.3629 | 321.3225 | 3289.6097 | 1858.1112 | 881.5742 | 1332.3451 | 1515.8919 | 1078.0128 | 747.0049 | 866.2517 | 2997.8742 | 1682.879 | 1173.9774 | 1716.5211
442 | 6.8564 | 1377.1195 | 1653.062 | 958.084 | 1773.6019 | 2096.7832 | 1307.8717 | 718.9867 | 1534.3186 | 1025.9425 | 596.6372 | 335.1018 | 736.6239 | 2555.7699 | 1073.6637 | 654.6416 | 1820.8495 | 2130.8319 | 1184.0355 | 639.3628 | 1536.4824 | 2723.3451 | 2170.916 | 1219.6769 | 1868.7257 | 3188.5929 | 2113.4071 | 1000.8982 | 2224.3761 | 1831.9823 | 1414.7832 | 556.4779 | 1327.9735 | 2017.969 | 1830.5399 | 700.7169 | 645.031 | 4724.2965 | 2571.7567 | 1335.2345 | 2438.1328 | 2481.2965 | 1842.7744 | 719.885 | 1301.8452 | 537.0885 | 453.4911 | 238.2168 | 1022.5974 | 2127.9867 | 962.2478 | 814.6239 | 850.9425 | 2223.8407 | 1258.3541 | 694.5309 | 1075.0133 | 2458.2566 | 1132.6504 | 915.2478 | 1255.3496 | 1178.7257 | 558.9867 | 464.5796 | 579.5089 | 2146.2611 | 998.3451 | 857.9867 | 1822.5221 | 2266.6372 | 1309.4779 | 1103.3761 | 1699.6504 | 2444.04 | 1918.2301 | 903.5531 | 1994.8497 | 2591 | 1854.5974 | 847.0885 | 2194.8937 | 1337.4956 | 903.6239 | 542.4911 | 1085.7787 | 908.8672 | 633.3008 | 468.854 | 513.5089 | 1654.3938 | 1118.3053 | 569.0885 | 1248.5576 | 4446.1815 | 2594.2433 | 1578.885 | 2959.2258 | 3399.3983 | 2431.6239 | 1428.7699 | 2352.5354 | 391 | 250.0885 | 164.8275 | 391 | 689.3363 | 627.1018 | 297.6947 | 720.1725 | 645.6504 | 598.8319 | 306 | 521.7832 | 403.7257 | 351.7389 | 190 | 445.4646 | 439.1725 | 386.084 | 244.5796 | 342.0575 | 3320.9735 | 1924.5796 | 876.7699 | 1351.8275 | 1527.1991 | 1106.6062 | 702.4469 | 845.3363 | 3069.2833 | 1687.8717 | 1178.5929 | 1696.5576
443 | 6.8674 | 1415.9223 | 1673.8574 | 934.5959 | 1851.0856 | 2105.8938 | 1325.7021 | 738.2203 | 1519.4533 | 1029.1347 | 590.7021 | 338.4326 | 753.3835 | 2563.7797 | 1065.1711 | 657.9842 | 1774.8938 | 2237.946 | 1214.3107 | 634.0491 | 1535.3598 | 2705.9223 | 2197.8653 | 1216.179 | 1897.4691 | 3248.8192 | 2148.3629 | 995.4041 | 2291.5498 | 1796.0777 | 1435.5753 | 559.4897 | 1385.9096 | 2026.824 | 1884.1996 | 735.0412 | 664.7306 | 4778.3756 | 2595.9715 | 1332.6086 | 2471.9794 | 2573.0158 | 1890.3471 | 727.5674 | 1378.0443 | 535.4247 | 439.8653 | 228.6735 | 1039.3186 | 2134.6735 | 973.2694 | 833.5467 | 880.9588 | 2258.3137 | 1295.8289 | 660.8447 | 1068.7591 | 2402.6942 | 1114.4041 | 918.0491 | 1301.2852 | 1164.6371 | 560.9144 | 479.1426 | 590.5388 | 2180.4484 | 988.71 | 889.767 | 1893.0601 | 2256.3756 | 1308.1632 | 1098.7306 | 1694.8161 | 2540.7021 | 1902 | 910.7227 | 2034.8986 | 2558.5516 | 1915.4484 | 854.0777 | 2135.5595 | 1356.943 | 892.2615 | 505.5024 | 1067.4405 | 899.2488 | 598.5388 | 458.5674 | 534.4897 | 1658.3962 | 1147.3344 | 594.2488 | 1261.9763 | 4467.2742 | 2568.2979 | 1589.1632 | 3009.3392 | 3447.7227 | 2451.8447 | 1413.8495 | 2342.4041 | 397.9223 | 254.0491 | 172.6735 | 397.9223 | 683.1189 | 655.9588 | 295.4897 | 699.3471 | 636.057 | 673.4976 | 316.8161 | 543.4405 | 401.3186 | 357.355 | 204.71 | 446.0777 | 436.9509 | 373.8447 | 233.1839 | 333.6735 | 3339.1062 | 1936.9794 | 885.2124 | 1386.4976 | 1490.8368 | 1112.122 | 702.1996 | 870.7021 | 3118.8004 | 1707.8653 | 1189.4976 | 1733.3392
444 | 6.8784 | 1422.3304 | 1620.6297 | 956.6873 | 1838.9734 | 2120.3437 | 1311.9867 | 741.881 | 1547.4937 | 1029.3348 | 578.3215 | 340.7797 | 757.6563 | 2539.207 | 1097.0089 | 620.0044 | 1809.3658 | 2300.6563 | 1178 | 632.5684 | 1522.4848 | 2732.7975 | 2169.1892 | 1221.4006 | 1900.7532 | 3309.6297 | 2177.2203 | 994.2247 | 2359.524 | 1837.2779 | 1444.8766 | 563.1057 | 1391.1759 | 2038.8633 | 1877.7399 | 710.3437 | 678.1234 | 4847.5291 | 2583.9911 | 1319.357 | 2516.5861 | 2604.3215 | 1904.8063 | 712.762 | 1376.8367 | 524.3437 | 440.5551 | 227.0044 | 1009.3171 | 2137.0089 | 971.8899 | 832.2025 | 911.4405 | 2309.119 | 1298.2513 | 665.6519 | 1082.119 | 2436.9519 | 1106.9956 | 916.1234 | 1354.793 | 1178.9386 | 554.3392 | 482.0924 | 600.238 | 2182.5107 | 994.207 | 900.9823 | 1948.5196 | 2250.3348 | 1319.4538 | 1120.1323 | 1739.7006 | 2506.0747 | 1861.0658 | 884.881 | 2017.9298 | 2534.6873 | 1938.4228 | 847.6696 | 2170.9519 | 1300.9468 | 873.8943 | 498.6918 | 1099.6785 | 896.6475 | 598.4494 | 453.9956 | 522.088 | 1626.4361 | 1182.4538 | 608.1101 | 1265.2956 | 4478.0798 | 2586.0177 | 1602.2335 | 2970.7709 | 3427.2203 | 2490.4671 | 1405.0355 | 2348.3437 | 409.2247 | 243.6696 | 165.2203 | 409.2247 | 662.6918 | 680.6563 | 285.3127 | 694.6741 | 631.0968 | 659.3962 | 317.207 | 567.2247 | 410.4715 | 357.2114 | 215.5462 | 440.5595 | 443.3259 | 396.4494 | 218.1101 | 312.8721 | 3357.1367 | 1950.4405 | 857.2025 | 1408.881 | 1528.0443 | 1132.9823 | 738.4538 | 878.7886 | 3082.5462 | 1709 | 1171.3918 | 1703.6696
445 | 6.8894 | 1444.6048 | 1620.1505 | 981.3717 | 1813.1151 | 2209.2715 | 1306.1446 | 733.4572 | 1577.3952 | 1052.0325 | 573.2301 | 358.5458 | 751.944 | 2556.2921 | 1104.2566 | 635.9735 | 1836.8554 | 2301.3452 | 1184.4012 | 633.2832 | 1577.6018 | 2818.0975 | 2184.124 | 1204.9204 | 1875.2861 | 3332.9264 | 2201.062 | 1021.2331 | 2340.9174 | 1898.1711 | 1460.7198 | 559.7434 | 1386.0945 | 2062.124 | 1879.3511 | 738.4602 | 671.4247 | 4983.6933 | 2596.4042 | 1372.062 | 2550.5428 | 2626.2066 | 1932.8849 | 684.0265 | 1322.1151 | 520.9115 | 435.4277 | 232 | 1034.4367 | 2160.7168 | 982.1446 | 826.003 | 885.0531 | 2314.3422 | 1278.876 | 667.2891 | 1073.6253 | 2577.4986 | 1113.8879 | 901.1946 | 1344.1091 | 1204.3657 | 573.3452 | 453.2271 | 610.6844 | 2249.5076 | 1008.0915 | 882.8289 | 1929.9469 | 2289.9499 | 1351.9469 | 1157.7168 | 1756.0531 | 2531.2981 | 1861.6668 | 918.9794 | 2066.2861 | 2576.7463 | 1940.062 | 851.4572 | 2208.1651 | 1327.4161 | 891.8879 | 527.9145 | 1104.7699 | 887.8024 | 617.2596 | 464.0885 | 517.9469 | 1642.6048 | 1182.4838 | 601.5133 | 1294.9616 | 4513.5899 | 2599.1416 | 1613.1711 | 2991.8054 | 3485.3541 | 2494.5103 | 1418.4808 | 2389.2626 | 406.056 | 264.3747 | 170.7729 | 406.056 | 690.0855 | 672.2006 | 273.8879 | 687.9115 | 642.4337 | 592.6548 | 299.0855 | 550.7964 | 430.0265 | 360.5458 | 201.7994 | 444.5428 | 441.8584 | 394.6844 | 220.6578 | 297.2006 | 3386.9145 | 2004.5252 | 840.4867 | 1395.4277 | 1561.0826 | 1145.9204 | 714.2507 | 899.2301 | 3082.5162 | 1699.8554 | 1149.9499 | 1704.7139
446 | 6.9004 | 1469.8353 | 1667.1219 | 983.4695 | 1827.8599 | 2293.0182 | 1339.5369 | 746.2075 | 1594.8909 | 1085.4513 | 580.1829 | 374.6524 | 746.7925 | 2607.7989 | 1116.0246 | 648.5305 | 1853.7016 | 2270.9262 | 1199.0428 | 632.4941 | 1553.3166 | 2866.7069 | 2276.0492 | 1246.3476 | 1907.5797 | 3379.4813 | 2228.2438 | 1025.585 | 2362.8781 | 1865.9508 | 1499.2866 | 544.2192 | 1428.3839 | 2119.3048 | 1979.8909 | 729.3593 | 660.1465 | 5033.8781 | 2662.9091 | 1368.7251 | 2641.0867 | 2648.2802 | 1957.1102 | 683.2075 | 1350.4759 | 528.8235 | 431.4086 | 245.1465 | 1050.3166 | 2130.0481 | 979.3839 | 831.3839 | 857.5914 | 2346.6225 | 1221.2257 | 682.8963 | 1088.6834 | 2608.8224 | 1147.2503 | 869.3476 | 1328.8417 | 1178.3294 | 568.7497 | 460.9027 | 603.4086 | 2367.4021 | 1010.2556 | 903.3112 | 1947.0791 | 2272.554 | 1362.9754 | 1161.323 | 1743.1465 | 2570.5604 | 1917.8353 | 954.5059 | 2070.8781 | 2568.4813 | 1964.4267 | 863.738 | 2168.4695 | 1416.3839 | 936.5187 | 540.2685 | 1091.7134 | 915.677 | 638.1647 | 471.6706 | 549.6342 | 1674.4086 | 1162.9572 | 608.6161 | 1275.6653 | 4530.2268 | 2572.6888 | 1585.1765 | 3057.4578 | 3544.323 | 2504.6834 | 1400.1465 | 2427.2866 | 391.9572 | 280.7743 | 168.7315 | 391.9572 | 703.5551 | 666.6096 | 276.262 | 685.0246 | 670.7743 | 619.7872 | 291.8963 | 534.5733 | 412.3048 | 393.0856 | 199.8781 | 441.6524 | 465.4941 | 391.6342 | 225 | 335.2749 | 3426.9701 | 2025.6395 | 856.3904 | 1413.4759 | 1583.7262 | 1160.585 | 698.8235 | 914.1647 | 3052.2738 | 1709.1893 | 1170.585 | 1733.4578
447 | 6.9114 | 1510.7356 | 1694.3904 | 965.2094 | 1878.9902 | 2252.6823 | 1352.9547 | 740.2458 | 1600.819 | 1107.0453 | 582.7817 | 375.8545 | 752.2724 | 2643.1269 | 1135.8909 | 656.1908 | 1868.1366 | 2265.2901 | 1201.2183 | 639.4002 | 1524.5634 | 2922.2804 | 2292.6548 | 1254.7817 | 1903.4188 | 3370.0541 | 2220.6273 | 1001.8545 | 2432.9352 | 1892.7081 | 1491.9184 | 534.5634 | 1416.8545 | 2117.3727 | 1975.701 | 688.8092 | 688.9725 | 5072.1358 | 2662.228 | 1316.9911 | 2710.7551 | 2659.7267 | 2036.2538 | 685.9184 | 1386.1544 | 550.7728 | 451.7906 | 240.2458 | 1034.5723 | 2129.9991 | 977.0089 | 839.1269 | 846.3638 | 2361.9823 | 1250.5723 | 694.3088 | 1093.8279 | 2578.4632 | 1166.1091 | 897.5359 | 1346.1091 | 1190 | 578.7995 | 461.2094 | 605 | 2342.0364 | 985.9361 | 897.5004 | 1925.4366 | 2265.1171 | 1331.5457 | 1140.2547 | 1734.3913 | 2522.1455 | 1947.1721 | 900.8642 | 2080.2272 | 2604.8261 | 1983.0177 | 849.055 | 2206.0994 | 1380.2822 | 961.3638 | 533.7276 | 1054.7551 | 945.4179 | 641.5909 | 474.4543 | 537.2094 | 1662.0275 | 1164.6362 | 623.5182 | 1289.3461 | 4593.0914 | 2611.2626 | 1544.9184 | 3009.6469 | 3530.4818 | 2540.6637 | 1418.3727 | 2441.118 | 396.0453 | 262.0275 | 171.4179 | 396.0453 | 717.5909 | 678.1633 | 278.8997 | 699.1091 | 691.5359 | 654.9547 | 299.4543 | 555.9725 | 407 | 391.173 | 202 | 442.3727 | 468.0639 | 386.9184 | 215.8456 | 342.6823 | 3477.7453 | 1985.4543 | 874.2272 | 1435.1819 | 1574.6646 | 1116.1366 | 708.7276 | 890.6096 | 3076.7533 | 1738.2635 | 1148.2999 | 1792.6087
448 | 6.9224 | 1530.3657 | 1764.5399 | 941.5413 | 1936.823 | 2241.2345 | 1376.5635 | 738.2699 | 1555.9174 | 1118.5177 | 594.4351 | 363.6947 | 810.5517 | 2638.3539 | 1105.8952 | 633.8716 | 1937.4559 | 2343.6461 | 1202.9056 | 662.5871 | 1610.6211 | 2932.9674 | 2291.2935 | 1232.3421 | 1903.5163 | 3446.9042 | 2222.9292 | 1011.9292 | 2473.8834 | 1923.907 | 1493.3407 | 547.4941 | 1393.3303 | 2155.093 | 1992.0562 | 686.9174 | 704 | 5140.152 | 2637.211 | 1341.5399 | 2712.5517 | 2702.9174 | 2077.6003 | 660.6593 | 1389.5767 | 527.1888 | 474.4705 | 217.5177 | 1077.7994 | 2162.789 | 971.7301 | 838.3067 | 860.211 | 2366.1166 | 1232.5309 | 692.7537 | 1071.8112 | 2635.6697 | 1158.059 | 895.3539 | 1366.2228 | 1220.6343 | 579.4011 | 428.6475 | 607.9646 | 2379.5399 | 1007.4351 | 853.6829 | 1900.6475 | 2290.0132 | 1356.0222 | 1130.5295 | 1720.8584 | 2554.5635 | 2013.1638 | 856.0708 | 2107.2699 | 2643.9424 | 2033.2345 | 819 | 2171.8848 | 1360.4705 | 959.3997 | 538.6933 | 1089.1624 | 979.2935 | 642.4587 | 466.7772 | 517.9882 | 1676.152 | 1194.764 | 610.6475 | 1265.823 | 4669.714 | 2669.0708 | 1541.3997 | 2994.6093 | 3503.3303 | 2543.9056 | 1439.2935 | 2447.059 | 414.9292 | 243.5413 | 185.0354 | 414.9292 | 708.5767 | 718.7994 | 295.0118 | 724.1638 | 708.1298 | 674.6107 | 314.5059 | 578.9056 | 404.5295 | 357.6711 | 191.1298 | 458.8466 | 432.7301 | 384.882 | 206.9882 | 317.3997 | 3549.2331 | 2012.8584 | 860.2596 | 1442.3525 | 1561.916 | 1126.1638 | 679.1062 | 878.764 | 3176.2463 | 1722.2713 | 1179.1992 | 1811.5767
449 | 6.9334 | 1530.9027 | 1790.3274 | 967.8771 | 1940.8594 | 2282.2301 | 1415.0128 | 740.7827 | 1589.3324 | 1083.5781 | 604.5575 | 341.4297 | 817.0973 | 2652.0561 | 1152.01 | 624.2813 | 1950.0284 | 2346.6548 | 1187.1996 | 680.2763 | 1620.6754 | 2874.8132 | 2305.9616 | 1199.4553 | 1949.0511 | 3563.3246 | 2232.0384 | 1027.1151 | 2469.358 | 1925.4602 | 1552.0689 | 545.468 | 1402.473 | 2213.8004 | 2058.821 | 684.3785 | 710.0767 | 5214.5625 | 2639.7187 | 1316.1818 | 2774.294 | 2738.358 | 2070.5192 | 664.7237 | 1385.532 | 536.826 | 488.647 | 223.0895 | 1057.0206 | 2171.9283 | 966.6854 | 832.2557 | 890.0639 | 2414.6982 | 1242.1967 | 675.0511 | 1092.0511 | 2675.647 | 1130.7187 | 897.4091 | 1375.3146 | 1237.3146 | 568.294 | 433.7237 | 597.3274 | 2420.5192 | 1022.1151 | 870.9155 | 1919.3963 | 2260.179 | 1356.995 | 1146.7365 | 1745.1534 | 2584.4041 | 1980.2202 | 867.7237 | 2151.3068 | 2617.1406 | 2075.7493 | 834.6982 | 2183.0995 | 1375.1534 | 971.5447 | 531.5781 | 1134.6676 | 955.4758 | 654.6087 | 460.7237 | 528.1151 | 1696.8849 | 1207.9105 | 603.5703 | 1267.8338 | 4688.8494 | 2704.5242 | 1539.3658 | 3024.4119 | 3531.2095 | 2597.5753 | 1441.3018 | 2498.2095 | 429.6087 | 238.9872 | 175.9233 | 429.6087 | 721.7493 | 712.1996 | 307.1662 | 733.7571 | 704.5959 | 676.6037 | 312.4808 | 584.9744 | 417.6982 | 346.5447 | 180 | 448.706 | 433.7621 | 410.1918 | 236.358 | 321.4425 | 3648.9283 | 2028.0384 | 835.468 | 1418.0334 | 1620.6087 | 1152.9744 | 716.5881 | 868.1101 | 3180.7827 | 1695.4553 | 1211.5831 | 1836.9027
450 | 6.9444 | 1558.8879 | 1764.5516 | 993.3692 | 1949.1913 | 2353.9431 | 1397.3274 | 693.3977 | 1620.1868 | 1076.9342 | 635.9715 | 334.9626 | 736.7251 | 2757.7607 | 1221.7011 | 665.1913 | 1898.3407 | 2348.5276 | 1212.2375 | 651.0703 | 1525.879 | 2933.7518 | 2286.9208 | 1180.9582 | 1947.8132 | 3586.5427 | 2237.0747 | 1014.2153 | 2527.6023 | 1936.6219 | 1568.5098 | 560.1165 | 1452.4858 | 2220.9208 | 2027.3648 | 672.9253 | 701.9956 | 5283.4155 | 2662.0881 | 1248.3319 | 2799.0747 | 2725.024 | 2133.7233 | 705.8594 | 1438.831 | 565.6308 | 514.0044 | 234.1495 | 1012.2198 | 2238.7144 | 1006.4155 | 862.8923 | 881.5142 | 2450.2286 | 1298.5934 | 710.6397 | 1092.8879 | 2635.6504 | 1137.5276 | 925.3692 | 1377.5979 | 1232.3363 | 586.3319 | 457.7429 | 577.7758 | 2439.6352 | 1040.3363 | 883.8087 | 1946.8879 | 2317.0507 | 1348.0836 | 1158.7758 | 1753.3692 | 2607.6023 | 2007.6575 | 912.4902 | 2139.7295 | 2646.1539 | 2044.6877 | 824.0658 | 2243.6219 | 1403.5979 | 981.2242 | 507.8176 | 1131.3034 | 943.0089 | 657.8132 | 472.8132 | 533.8879 | 1691.6308 | 1202.556 | 610.556 | 1264.3363 | 4606.7429 | 2722.6966 | 1570.4947 | 3024.6441 | 3556.9537 | 2618.6593 | 1417.7011 | 2530.6966 | 425.5516 | 248.8923 | 165.3319 | 425.5516 | 692.2438 | 723.2331 | 315.3319 | 713.444 | 712.5187 | 628.8835 | 299.589 | 570.5516 | 419.5142 | 372.8221 | 190.8923 | 449.3407 | 451.9253 | 410.4769 | 243.7714 | 333.0792 | 3660.734 | 2010.7714 | 850.1165 | 1446.3167 | 1613.4395 | 1141.6637 | 740.2153 | 858.2286 | 3168.6681 | 1723.6397 | 1196.2153 | 1833.7669
451 | 6.9554 | 1566.0885 | 1788.5753 | 965.0442 | 1953.1947 | 2415.1505 | 1384.2478 | 683.0531 | 1615.4425 | 1120.646 | 610.0265 | 347.8053 | 733.4779 | 2778.8318 | 1194.8849 | 669.1947 | 1954.5133 | 2381.9027 | 1250.0973 | 646.3717 | 1536.2921 | 3016.9735 | 2263.5044 | 1220.6284 | 1960.5222 | 3539.9734 | 2268.2036 | 1007.8053 | 2550.1593 | 1918.2478 | 1579.6372 | 552.6371 | 1452.7876 | 2229.3629 | 2011.531 | 677.3717 | 685.2832 | 5353.7257 | 2699.2832 | 1302.4072 | 2893.3895 | 2703.9912 | 2130.9999 | 704.6371 | 1438.7433 | 564.2212 | 493.0176 | 232.6903 | 1048.0089 | 2293.3629 | 1014.0708 | 848.5752 | 881.2124 | 2448.292 | 1309.4956 | 699.115 | 1097.9646 | 2634.1948 | 1123.646 | 935.2744 | 1403.6195 | 1237.2832 | 577.2212 | 456.5664 | 574.1239 | 2497.416 | 1029.8849 | 873.9646 | 1974.7965 | 2362.8407 | 1370.6903 | 1120.6106 | 1768.5841 | 2581.3097 | 2038.5574 | 946.4956 | 2110.7168 | 2630.6106 | 2027.2036 | 843.5399 | 2217.8142 | 1418.469 | 980.0973 | 531.3983 | 1120.6725 | 923.8672 | 660.4336 | 455.2566 | 552.7699 | 1699.2478 | 1208.2478 | 599.7876 | 1282.5575 | 4581.6725 | 2709.9646 | 1578.7345 | 3074.9646 | 3598.4957 | 2653.8142 | 1416.3717 | 2526.9912 | 407.6283 | 253.1593 | 163.1239 | 407.6283 | 678.3275 | 759.2478 | 318.4336 | 716.2478 | 708.7522 | 638.4868 | 301.7434 | 583.9558 | 401.1593 | 373.292 | 195.6903 | 475.2478 | 458.4956 | 397 | 229.7788 | 330.4778 | 3665.4248 | 2017.4868 | 871.8407 | 1462.4601 | 1640.416 | 1135.1239 | 729.0266 | 866.9115 | 3186.2744 | 1741.8495 | 1197.7699 | 1806.2478
452 | 6.9664 | 1628.24 | 1840.6441 | 962.5576 | 2030.6388 | 2443.471 | 1414.62 | 707.6922 | 1618.4086 | 1128.2212 | 609.0045 | 359.4567 | 776.4327 | 2757.4086 | 1193.7547 | 634.7261 | 1938.2792 | 2407.0143 | 1238.0241 | 635.1588 | 1608.4853 | 3061.4567 | 2353.7878 | 1203.8947 | 1948.2935 | 3495.6012 | 2341.8073 | 1008.0482 | 2492.8465 | 1907.8751 | 1573.9714 | 528.1971 | 1412.6637 | 2258.6922 | 2087.5139 | 678.1106 | 715.1392 | 5345.0196 | 2734.0286 | 1344.2212 | 2976.6298 | 2798.86 | 2104.0767 | 721.4853 | 1412.7547 | 552.3943 | 470.0624 | 243.2114 | 1097.9518 | 2333.5576 | 993.4808 | 810.1588 | 908.8412 | 2427.8269 | 1322.4327 | 677.6682 | 1103.9135 | 2701.0767 | 1132.0045 | 918.7502 | 1372.2551 | 1267.6824 | 564.851 | 446.1971 | 601.5335 | 2557.9037 | 995.6155 | 879.3702 | 1966.1445 | 2338.1204 | 1342.0963 | 1106.1249 | 1763.8127 | 2561.7931 | 1985.851 | 901.8465 | 2112.3702 | 2589.5049 | 2082.8796 | 879.2837 | 2170.9955 | 1418.5433 | 947.9233 | 534.678 | 1082.1347 | 922.0286 | 666.4567 | 446.1731 | 561.1347 | 1703.5433 | 1205.4808 | 586.2837 | 1291.5433 | 4588.769 | 2699.0722 | 1543.3078 | 3033.6494 | 3637.2694 | 2702.3702 | 1442.6441 | 2534.9376 | 403.6298 | 256.149 | 166.0865 | 403.6298 | 697.851 | 722.7984 | 305.4424 | 697.1829 | 677.2935 | 641.4424 | 317.3461 | 559.8224 | 409.0384 | 372.9278 | 193.1731 | 486.0624 | 465.4567 | 392.1106 | 229.6539 | 336.7306 | 3704.3845 | 2026.9616 | 859.0722 | 1446.8171 | 1649.8367 | 1155.471 | 719.6155 | 836.8127 | 3176.2694 | 1727.6539 | 1216.4567 | 1781.2069
453 | 6.9774 | 1660.7778 | 1794.6666 | 953.4444 | 2027.6666 | 2447.3333 | 1433.1111 | 715.5556 | 1615.2222 | 1120 | 648.2222 | 361.7778 | 774.8889 | 2772.5556 | 1185.3333 | 624.3333 | 1911.2222 | 2483.0001 | 1206.3333 | 632.1111 | 1639.4444 | 3052.1111 | 2410.4444 | 1185.0001 | 1965.0001 | 3623.3336 | 2360.3333 | 975.7777 | 2562.8891 | 1874.8889 | 1519.2222 | 527.7778 | 1425.7778 | 2264.8889 | 2094.8888 | 704.5556 | 729.1111 | 5340.2223 | 2788.2223 | 1347.6667 | 2981.8889 | 2834.3333 | 2122.1111 | 732.4444 | 1423.7778 | 559.3334 | 491.5556 | 237.1111 | 1108 | 2331.6666 | 991.3333 | 783.7778 | 902.4444 | 2506.0001 | 1354.2223 | 681.6667 | 1084.1111 | 2694.6666 | 1157.3333 | 903.3333 | 1393.5556 | 1275.4444 | 586.3334 | 455.7778 | 613 | 2598.5556 | 1004.8889 | 897.4445 | 1954.8889 | 2333 | 1377.0001 | 1119.6667 | 1797.1112 | 2606.5556 | 1998.4445 | 912.4445 | 2129.3334 | 2611.7778 | 2086.3333 | 863.1111 | 2217.889 | 1393.4444 | 958.2223 | 521 | 1074.1111 | 941.7778 | 649.8889 | 470.7778 | 533.6666 | 1689 | 1232.2223 | 595.1111 | 1292 | 4671.889 | 2702.8889 | 1534.3333 | 2989.6667 | 3541.2221 | 2760.4445 | 1416.6666 | 2569.5556 | 413.3333 | 248 | 168.1111 | 413.3333 | 717.6667 | 692.8889 | 305.6667 | 704.6667 | 655.6667 | 657.2223 | 314.3333 | 546 | 433.3334 | 395.7778 | 186.1111 | 498.6667 | 447.2222 | 422.4445 | 215.4444 | 340.5555 | 3762.3334 | 2049.3334 | 866.2223 | 1468.7778 | 1659.2223 | 1143.2222 | 696.1111 | 810.1111 | 3183.5556 | 1724.5556 | 1176.5555 | 1728.6666
454 | 6.9884 | 1695.2921 | 1733.7787 | 919.9395 | 1991.6888 | 2502.7346 | 1458.3924 | 729.0605 | 1601.5929 | 1139.3068 | 639.7787 | 356.618 | 747.854 | 2876.9411 | 1196.9602 | 625.8894 | 1948.2567 | 2500.3362 | 1225.3172 | 645.5428 | 1578.8082 | 3024.5575 | 2357.0648 | 1218.1106 | 1975.9896 | 3808.4883 | 2371.3924 | 976.739 | 2559.1695 | 1903.4278 | 1565.7346 | 515.8038 | 1401.4866 | 2257.1858 | 2119.2774 | 700.6076 | 727.6534 | 5436.4381 | 2833.8245 | 1317.8185 | 2985.5678 | 2890.77 | 2224.6196 | 748.3673 | 1444.146 | 527.4114 | 517.0605 | 218.0251 | 1086.9896 | 2298.8289 | 1032.0458 | 787.0605 | 882.0251 | 2547.854 | 1319.7831 | 673.8643 | 1045.7183 | 2703.7995 | 1154.2463 | 905.4572 | 1371.5619 | 1217.5471 | 589.1253 | 466.5678 | 603.2965 | 2618.4071 | 1026.2714 | 855.0796 | 2004.0708 | 2336.5074 | 1414.1858 | 1157.9101 | 1811.3112 | 2631.5177 | 2094.6343 | 931.0147 | 2186.0207 | 2652.1962 | 2041.1755 | 845.1858 | 2192.4409 | 1366.1858 | 985.1357 | 510.7537 | 1101.6888 | 946.764 | 673.9101 | 494.1106 | 542.2817 | 1662.236 | 1261.4071 | 593.7787 | 1298.8142 | 4672.9793 | 2702.1003 | 1545.4926 | 3047.1918 | 3551.6845 | 2758.0294 | 1445.5989 | 2620.739 | 422.2714 | 242.8142 | 157.6431 | 422.2714 | 741.8392 | 713.0354 | 307.6431 | 727.7035 | 701.8599 | 663.1253 | 322.6283 | 573.382 | 429.8038 | 382.6932 | 178.4322 | 476.7433 | 437.5428 | 438.0752 | 212.6785 | 310.261 | 3808.3569 | 2007.8583 | 878.3215 | 1489.739 | 1630.4866 | 1116.4572 | 692.9749 | 835.3673 | 3163.6828 | 1749.4174 | 1182.4779 | 1765.624
455 | 6.9994 | 1711.0157 | 1758.4878 | 919.8624 | 1986.4739 | 2545.0383 | 1478.2596 | 734.3589 | 1606.5418 | 1139.4965 | 635.4808 | 346.939 | 769.2891 | 2917.6116 | 1218.4199 | 642.3048 | 1938.2527 | 2486.2822 | 1253.6794 | 654.2212 | 1575.9302 | 3010.7404 | 2407.9981 | 1222 | 1990.1672 | 3779.9026 | 2352.9704 | 952.7109 | 2506.1446 | 1954.824 | 1626.6028 | 537.4878 | 1414.533 | 2286.8083 | 2150.5349 | 666.8171 | 754.7247 | 5450.4513 | 2857.9007 | 1281.0383 | 2971.4965 | 2889.7649 | 2266.1307 | 724.2684 | 1455.1986 | 543.1515 | 516.5575 | 217.3206 | 1073.9007 | 2329.0905 | 1029.0924 | 830.0679 | 859.8554 | 2554.7857 | 1278.1603 | 681.7021 | 1081.9529 | 2711.7561 | 1169.4425 | 911.2822 | 1348.8466 | 1189.4808 | 588.5035 | 436.8328 | 589.237 | 2713.8222 | 1017.5575 | 843.4878 | 1988.23 | 2322.939 | 1415.9617 | 1144.23 | 1723.7039 | 2608.9931 | 2115.9704 | 935.2055 | 2146.643 | 2651.3972 | 2050.1672 | 826.6568 | 2151.0069 | 1388.8466 | 970.9164 | 517.0226 | 1128.9617 | 967.2822 | 685.176 | 487.5575 | 548.9774 | 1651 | 1265.1603 | 609.7857 | 1302.2596 | 4712.357 | 2768.3187 | 1578.7857 | 3047.7492 | 3611.6411 | 2786.1358 | 1456.291 | 2666.9076 | 397.3136 | 254.9617 | 166.4042 | 397.3136 | 747.4808 | 743.4042 | 287.3972 | 680.2684 | 742.9617 | 672.3658 | 300.5122 | 544.1917 | 421.3206 | 378.4425 | 186.7021 | 449.0383 | 450.8624 | 393.6116 | 222.2212 | 304.7021 | 3752.6046 | 1986.7473 | 868.7788 | 1445.406 | 1611.6254 | 1121.1219 | 705.8624 | 877.4042 | 3136.4808 | 1787.7247 | 1148.345 | 1814.4808
456 | 7.0104 | 1719.9199 | 1800.3653 | 925.8151 | 2005.9155 | 2512.7482 | 1490.0334 | 715.0044 | 1605.4833 | 1146.1805 | 623.1893 | 345.1849 | 778.6347 | 2909.7729 | 1206.1514 | 640.412 | 1908.8908 | 2518.4745 | 1228.8952 | 663.9243 | 1579.7482 | 3069.4657 | 2378.2826 | 1190.0088 | 1992.4498 | 3673.0889 | 2348.882 | 974.1717 | 2512.8908 | 1920.7905 | 1627.2984 | 546.7817 | 1462.8864 | 2310.2227 | 2133.8151 | 676.9577 | 770.1849 | 5376.6056 | 2861.7773 | 1290.2183 | 3035.6461 | 2902.6505 | 2275.3231 | 705.588 | 1442.706 | 568.5211 | 521.441 | 221.3697 | 1064.7438 | 2324.2694 | 1000.6681 | 877.6972 | 891.0246 | 2531.7148 | 1325.8019 | 716.4366 | 1096.118 | 2681.6347 | 1149.1558 | 888.3785 | 1329.3072 | 1159.4164 | 585.9666 | 440.4745 | 593.6259 | 2725.3116 | 998.7438 | 852.1893 | 1955.0379 | 2306.9666 | 1402.6347 | 1157.8398 | 1653.0801 | 2596.294 | 2082.2606 | 948.4454 | 2116.9199 | 2715.2764 | 2081.4789 | 826.4789 | 2197.4745 | 1398.3363 | 970.0713 | 545.1426 | 1139.3319 | 1007.7685 | 693.1426 | 484.7394 | 543.9621 | 1703.1338 | 1289.1048 | 613.3741 | 1273.7482 | 4696.463 | 2798.3407 | 1607.2562 | 3024.4789 | 3515.4718 | 2786.0132 | 1452.1048 | 2687.0379 | 372.6681 | 248.1514 | 158.265 | 372.6681 | 742.2984 | 758.3697 | 275.4454 | 679.3609 | 698.8283 | 685.853 | 301.2896 | 538.5502 | 400.4877 | 385.4076 | 206.6259 | 444 | 431.3407 | 391.6972 | 226.5924 | 302.1136 | 3749.8398 | 2055.8019 | 861.4454 | 1395.2984 | 1665.9155 | 1115.1514 | 712.4076 | 883.4833 | 3191.1338 | 1759.9375 | 1144.4701 | 1810.4833
457 | 7.0214 | 1771.2832 | 1786.997 | 949.7935 | 2009.6253 | 2529.0738 | 1455.4601 | 705.2094 | 1615.5133 | 1141.6725 | 646.2596 | 352.0472 | 797.4455 | 2983.0768 | 1196.1858 | 618.115 | 1926.3747 | 2535.3953 | 1245.0502 | 636.5545 | 1556.767 | 3047.9232 | 2369.5871 | 1166.1858 | 1977.7906 | 3726.1034 | 2416.1918 | 992.882 | 2554.5635 | 1927.0974 | 1677.6343 | 544.2566 | 1440.5781 | 2324.7227 | 2168.0738 | 691.8142 | 752.2537 | 5432.8939 | 2872.6755 | 1296.3717 | 3095.4661 | 2920.6253 | 2287.0914 | 706.115 | 1458.1652 | 586.7935 | 519.7197 | 237.5133 | 1083.6077 | 2332.2596 | 978.2537 | 883.115 | 898.4159 | 2545.121 | 1352.5575 | 704.5545 | 1074.767 | 2724.0296 | 1161.4455 | 926.6579 | 1342.5163 | 1169.6549 | 592.7227 | 444.6961 | 562.4837 | 2790.4072 | 997.0472 | 894.6107 | 1887.0855 | 2313.7227 | 1391.8142 | 1186.1858 | 1656.7463 | 2643.7257 | 2122.5399 | 918.7876 | 2159.8171 | 2660.3833 | 2066.2301 | 869.888 | 2227.6991 | 1404.3039 | 1004.8171 | 531.1829 | 1144.8142 | 974.1563 | 648.3421 | 476.5103 | 530.2773 | 1718.4395 | 1260.2507 | 600.9764 | 1279.1888 | 4640.9292 | 2813.2596 | 1594.4395 | 2986.2507 | 3511.7051 | 2719.1356 | 1423.8554 | 2656.5781 | 402.2596 | 258.1416 | 157.9086 | 402.2596 | 743.0472 | 775.7227 | 295.9794 | 692.7197 | 678.9086 | 670.9056 | 313.767 | 565.9086 | 397.6991 | 392.8614 | 200.882 | 451.8614 | 429.3275 | 414.6519 | 222.1622 | 302.0472 | 3772.1386 | 2077.115 | 858.3953 | 1388.7906 | 1636.9704 | 1127.5605 | 716.8378 | 811.2713 | 3273.5635 | 1784.2154 | 1156.6725 | 1775.764
458 | 7.0324 | 1793.766 | 1800.8506 | 938.2345 | 1988.2129 | 2587.7867 | 1494.8713 | 726.9784 | 1595.3834 | 1076.086 | 634.6814 | 370.1913 | 782.2133 | 2965.6179 | 1179.5747 | 638.7655 | 1988.5099 | 2583.1269 | 1264.766 | 646.8506 | 1541.6598 | 3001.8299 | 2442.1058 | 1174.2552 | 1936.2772 | 3733.788 | 2437.8938 | 976.2129 | 2512.8096 | 2004.446 | 1705.3618 | 568.7444 | 1448.5315 | 2296.7453 | 2173.7237 | 680.0428 | 727.0428 | 5510.1274 | 2887.6382 | 1279.1917 | 3107.7871 | 2875.1494 | 2332.297 | 691.5959 | 1491.1274 | 603 | 501.2768 | 227.8726 | 1075.0855 | 2289.8308 | 977.7232 | 848.8515 | 876.2129 | 2487.4694 | 1330.4897 | 659.3195 | 1062.1278 | 2766.7871 | 1175.8299 | 948.3407 | 1340.6175 | 1222.4676 | 580.1706 | 441.5531 | 553.2763 | 2885.1485 | 1065.1692 | 900.6175 | 1840.9149 | 2343.1274 | 1426.9356 | 1130.703 | 1664.6809 | 2625.9582 | 2133.1067 | 914.5743 | 2168.1278 | 2625.9356 | 1981.0437 | 865.6175 | 2136.0441 | 1416.7021 | 1025.468 | 513.617 | 1080.5968 | 961.6805 | 611.851 | 484.4253 | 541.5954 | 1722.8933 | 1294.5522 | 586.6598 | 1232.0648 | 4621.4257 | 2812.1706 | 1540.8942 | 2998.9568 | 3553.383 | 2800.658 | 1436.3398 | 2580.6179 | 421.532 | 253.9577 | 164.532 | 421.532 | 726.6386 | 766.8726 | 309.766 | 695.8722 | 674.4257 | 680.8933 | 304.2129 | 550.9365 | 418.9572 | 407.2552 | 186.6809 | 456.383 | 432.5959 | 404.1917 | 214.0639 | 305.383 | 3821.914 | 2053.3407 | 877.1274 | 1370.7237 | 1610.2552 | 1124.3407 | 723.0851 | 802.7867 | 3170.1085 | 1797.4046 | 1157.1062 | 1722.0644
459 | 7.0434 | 1752.0953 | 1851.4651 | 947.4307 | 1984.6293 | 2557.7679 | 1523.7846 | 742.5172 | 1604.06 | 1058.3619 | 601.4139 | 372.707 | 733.5693 | 2989.1545 | 1216.5684 | 637.7502 | 1987.1642 | 2602.9312 | 1313.0856 | 630.1297 | 1515.7502 | 2994.2242 | 2414.1386 | 1206.4307 | 1881.4228 | 3641.5869 | 2388.2074 | 970.1121 | 2484.9912 | 1961.4837 | 1657.2074 | 564.6381 | 1422.2074 | 2306.9823 | 2178.1033 | 705.353 | 734.7326 | 5565.4219 | 2957.1889 | 1247.3451 | 3078.0177 | 2865.4395 | 2336.4228 | 652.3884 | 1457.0609 | 580.9744 | 508.474 | 214.1121 | 1064.0344 | 2306.0079 | 1013.3186 | 823.1898 | 875.7758 | 2487.8614 | 1357.3875 | 669.8014 | 1075.5084 | 2791.1377 | 1199.06 | 927.9312 | 1331.1809 | 1197.9488 | 589.9135 | 456.3274 | 541.3451 | 2866.759 | 1068.5004 | 873.5605 | 1823.8967 | 2334.233 | 1413.3884 | 1099.5172 | 1677.7326 | 2652.12 | 2092.4228 | 898.4572 | 2126.759 | 2680.353 | 1974.1289 | 827.8623 | 2094.767 | 1422.7758 | 1032.6637 | 514 | 1045.6637 | 975.1121 | 632.5084 | 482.0432 | 548.0777 | 1699.3884 | 1277.3628 | 561.3795 | 1219.4307 | 4664.6028 | 2777.233 | 1563.1977 | 2913.6134 | 3561.8102 | 2936.1112 | 1409.7679 | 2628.6973 | 408.6726 | 259.7326 | 168.0344 | 408.6726 | 761.1977 | 762.5516 | 322.2154 | 686.2674 | 673.5516 | 683.8191 | 304.4051 | 568.7237 | 432.6637 | 397.8967 | 197.2154 | 469.2154 | 441.2154 | 428.9823 | 226.7326 | 317.5861 | 3818.8711 | 2046.7758 | 880.1898 | 1366.9223 | 1570.0609 | 1134.767 | 697.3107 | 798.5693 | 3151.1889 | 1830.5684 | 1141.9744 | 1663.8711
460 | 7.0544 | 1781.2523 | 1923.8939 | 928.9955 | 1995.9071 | 2568.7036 | 1460.1813 | 724.1106 | 1591.7699 | 1103.2877 | 637.6284 | 365.792 | 741.739 | 3030.4159 | 1237.5841 | 610.9602 | 1949.1106 | 2522.5751 | 1283.615 | 639.1549 | 1473.5619 | 3030.4956 | 2404.8142 | 1191.2035 | 1869.1328 | 3679.7833 | 2363.6416 | 1014.5532 | 2449.4468 | 1894.5442 | 1608.9026 | 585.1549 | 1403.4159 | 2369.9868 | 2187.6416 | 694.8451 | 709.9955 | 5635.9116 | 3029.0045 | 1253.2478 | 3058 | 2815.8229 | 2251.6327 | 649.323 | 1435.7743 | 586.6062 | 535.6062 | 228.3983 | 1037.4867 | 2320.3938 | 988.2964 | 848.7965 | 842.0707 | 2473.4468 | 1346.6371 | 679.2257 | 1067.7522 | 2797.8672 | 1226.5487 | 909.8097 | 1257.084 | 1237.4426 | 573.8451 | 461.7168 | 536.3407 | 2896.6859 | 1013.6371 | 876.9823 | 1772.0132 | 2329.6991 | 1414.8142 | 1106.7743 | 1644.0132 | 2638.4468 | 2153.2744 | 916.7213 | 2097.6593 | 2602.4777 | 2058.2346 | 847.4381 | 2142.2877 | 1426.2832 | 1007.7699 | 526.8319 | 1070.1726 | 977.208 | 610.6371 | 473.7168 | 525.3938 | 1710.4381 | 1301.1018 | 561.2655 | 1245.9071 | 4664.8451 | 2816.3275 | 1528.7654 | 2906.385 | 3615.1195 | 2922.1813 | 1392.4735 | 2658.1106 | 396.2257 | 263.2257 | 175.4912 | 396.2257 | 778.0929 | 781.7567 | 310.1106 | 672.5088 | 701.0974 | 673.2257 | 293.5265 | 541.1637 | 416.7522 | 366.5442 | 217.8319 | 459.677 | 457.9823 | 436.1681 | 223.1681 | 310.8849 | 3769.4115 | 2039.3761 | 888.4735 | 1326.0884 | 1605.8762 | 1127.7522 | 697.6814 | 792.6239 | 3236.4779 | 1865.6991 | 1125.792 | 1701.6682
461 | 7.0654 | 1829.9627 | 1928.4985 | 962.3108 | 2032.0787 | 2622.7326 | 1477.3875 | 721.5005 | 1543.9596 | 1119.7306 | 621.4591 | 353.5378 | 768.7316 | 3075.8486 | 1244.4622 | 620.3088 | 1934.7306 | 2486.2311 | 1283.6942 | 618.7286 | 1501.7719 | 3019.3058 | 2414.3845 | 1178.5772 | 1842.5752 | 3730.1544 | 2413.6952 | 1038.6922 | 2424 | 1953.2734 | 1689.5055 | 560.1514 | 1314.8022 | 2372.4218 | 2220.0404 | 645.3058 | 706.0394 | 5676.0394 | 2989.1504 | 1290.1554 | 3116.1957 | 2862.1584 | 2322.5449 | 661.6155 | 1453.3855 | 597 | 517.2291 | 235.6922 | 1044.5399 | 2321.7699 | 976.116 | 892.1554 | 797.8063 | 2455.8466 | 1307.0757 | 658.3451 | 988.996 | 2721.111 | 1179.3814 | 920.6549 | 1205.1534 | 1227.5742 | 565.5005 | 459.6922 | 539.6922 | 2982.5419 | 1019.1937 | 887.885 | 1753.2311 | 2340.5388 | 1384.4975 | 1136.1554 | 1669.6952 | 2624.116 | 2146.6509 | 919.3068 | 2040.6892 | 2626.4662 | 2068.1524 | 834.9213 | 2122.113 | 1364.8809 | 970.4218 | 516.3451 | 1091.1544 | 898.5722 | 632.5409 | 463.8456 | 533.9627 | 1686.1514 | 1276.112 | 593.1554 | 1250.6922 | 4670.2321 | 2732.5318 | 1554.1574 | 2912.1907 | 3665.5399 | 2860.9213 | 1340.4965 | 2644.3845 | 377.3068 | 273.0777 | 176.0383 | 377.3068 | 768.0767 | 814.1937 | 321.2321 | 685.6932 | 699.6145 | 652.999 | 289.2694 | 571.2341 | 422.3471 | 355.3078 | 209.9606 | 440.8456 | 445.3451 | 408.0757 | 225.1544 | 294.4995 | 3731.7679 | 2067.3481 | 884.884 | 1312.7699 | 1691.3491 | 1126.1544 | 721.3855 | 795.3845 | 3163.6095 | 1868.0383 | 1145.5782 | 1668.3037
462 | 7.0764 | 1949.919 | 1957.6353 | 961.3544 | 2014.025 | 2637.6676 | 1530.9794 | 728.6647 | 1561.6411 | 1133.322 | 625.9809 | 378.6441 | 730.7 | 3090.6721 | 1273.3162 | 640.6588 | 1917.678 | 2490.3324 | 1272.3544 | 643.6338 | 1527.3324 | 3021.3191 | 2399.0103 | 1156.3515 | 1839.6573 | 3744.3265 | 2394.7 | 995.3647 | 2391.3574 | 1961.6883 | 1764.3147 | 538.6662 | 1338.6132 | 2301.7103 | 2241.6632 | 652.6485 | 696.6794 | 5566.0869 | 2959.9985 | 1296.0059 | 3132.3441 | 2806.4045 | 2374.6706 | 707.6338 | 1421.0294 | 595.6676 | 534.6426 | 213.6824 | 1029.6868 | 2324.0029 | 989.3279 | 854.0383 | 804.3191 | 2483.9823 | 1352.6249 | 652.9971 | 1018.2852 | 2769.269 | 1176.9809 | 904.0191 | 1239.6382 | 1190.3427 | 581.325 | 490.9764 | 531.0059 | 2863.7766 | 1003.3544 | 873.6794 | 1783.3132 | 2312.0236 | 1431.2823 | 1107.3647 | 1653.3647 | 2688.6235 | 2145.3147 | 913.6647 | 2037.6485 | 2606.0456 | 2028.353 | 832.3235 | 2103.6647 | 1363.3103 | 957.0015 | 524.9867 | 1056.0294 | 913.2911 | 660.6618 | 453.0044 | 503.3633 | 1650.6765 | 1280.6411 | 595.675 | 1233.3456 | 4579.4089 | 2781.5867 | 1542.0324 | 2869.3544 | 3688.6588 | 2846.9985 | 1386.9426 | 2650.3279 | 369.6662 | 254.0191 | 175 | 369.6662 | 728.6942 | 822.3368 | 315.6794 | 662.6897 | 682.6735 | 663.3191 | 281.6735 | 544.3751 | 392.3618 | 359.9971 | 209.9941 | 445.9926 | 419.3456 | 395.6676 | 218.0088 | 309.3191 | 3798.6088 | 2007.7236 | 831.0353 | 1360.9691 | 1618.0736 | 1121.0074 | 710.3471 | 807.9897 | 3202.2661 | 1813.0398 | 1127.0221 | 1657.6485
463 | 7.0874 | 1899.8376 | 1961.3599 | 955.8201 | 1983.4307 | 2636.3215 | 1526.6814 | 732.0354 | 1543.7197 | 1101.6489 | 649.8555 | 374.3245 | 709.3569 | 3115.4956 | 1310.7817 | 638.5369 | 1901.8584 | 2577.8438 | 1250.5369 | 663.2861 | 1477.115 | 3007.2537 | 2406.9263 | 1182.6726 | 1835.823 | 3738.1445 | 2378.1062 | 961.7876 | 2410.9587 | 1954.4631 | 1701.0117 | 536.9646 | 1350.5752 | 2318.1358 | 2193.7934 | 682.6401 | 683.9292 | 5585.0237 | 2965.7493 | 1299.4631 | 3200.3098 | 2812.2774 | 2344.8613 | 707.6106 | 1397.6077 | 580.7522 | 541.5015 | 228.1033 | 1037.0325 | 2295.5398 | 978.7522 | 870.7434 | 809.9292 | 2466.1475 | 1357.2537 | 660.4277 | 1027.9321 | 2751.2595 | 1189.3215 | 934.4219 | 1247.4661 | 1151.4336 | 586.3215 | 500.2153 | 558.2095 | 2924.0149 | 1022.2449 | 844.646 | 1760.3628 | 2233.581 | 1421.2212 | 1103.6401 | 1625.2153 | 2618.2654 | 2072.5132 | 907.5369 | 1999.115 | 2521.4011 | 2033.3186 | 852.5693 | 2001.873 | 1362.7168 | 943.1091 | 509.5752 | 1046.8555 | 927.0737 | 642.2891 | 456.7847 | 479.6077 | 1679.9587 | 1302.0708 | 590 | 1216.8584 | 4515.7876 | 2820.3953 | 1573.5989 | 2851.6077 | 3608.5132 | 2896.1712 | 1400.2183 | 2696.7434 | 391.0325 | 251.8555 | 172.9646 | 391.0325 | 734.4248 | 798.9675 | 293.4307 | 640.8938 | 698.354 | 673.6785 | 306.1741 | 512.6077 | 409.6372 | 381.6755 | 211.2861 | 438.1091 | 426.6047 | 414.6755 | 222.177 | 302.0383 | 3681.2741 | 1984.6047 | 857.2066 | 1297.8701 | 1631.0974 | 1118.0354 | 682.0029 | 801.4307 | 3176.0825 | 1813.1387 | 1128.9617 | 1688.3186
464 | 7.0984 | 1920.4573 | 1954.3815 | 918.5545 | 1957.6588 | 2627.0201 | 1477.2453 | 733.6908 | 1557.3934 | 1096.8152 | 648.7832 | 377.4336 | 713.4538 | 3059.9242 | 1286.3898 | 610.1327 | 1891.0924 | 2572.7192 | 1263.5782 | 657.1647 | 1477.8673 | 2992.5462 | 2401.5664 | 1218.2689 | 1823.4017 | 3739.9076 | 2455.6019 | 977.4135 | 2423.8555 | 1918.6268 | 1693.3211 | 580.8993 | 1354.8874 | 2373.1564 | 2163.0924 | 683.4017 | 670.2571 | 5603.3495 | 2976.2891 | 1283.6588 | 3180.0486 | 2811.7512 | 2405.9111 | 714.9597 | 1382.8756 | 571.237 | 567.3934 | 220.6588 | 1033.1848 | 2299.5782 | 989.2689 | 828.141 | 788.6588 | 2460.9799 | 1345.3092 | 711.3531 | 1035.6505 | 2659.141 | 1239.4253 | 937.0403 | 1191.5747 | 1129.0924 | 594.2891 | 493.5462 | 577.8353 | 2909.2536 | 1049.0521 | 886.1161 | 1703.6268 | 2208.1445 | 1406.2168 | 1120.9799 | 1626.1445 | 2633.1682 | 2087.9514 | 919.1967 | 1934.2453 | 2552.6422 | 2037.8555 | 836.8957 | 1964.7429 | 1375.7227 | 983.2808 | 538.064 | 1019.5344 | 942.1043 | 634.763 | 469.6706 | 491.1244 | 1676.968 | 1284.6588 | 595.5261 | 1188.1327 | 4437.4704 | 2810.9277 | 1567.2251 | 2808.5545 | 3544.5142 | 2912.0924 | 1425.9917 | 2615.4586 | 381.6588 | 240.4218 | 165.7832 | 381.6588 | 687.2856 | 740.6469 | 293.2168 | 658.032 | 719.7429 | 662.9479 | 285.4621 | 521.3614 | 419.4017 | 406.1967 | 198.2571 | 465.8471 | 416.7311 | 391.5344 | 227 | 300.2168 | 3656.0438 | 2007.8874 | 861.6588 | 1248.3294 | 1621.6268 | 1119.6908 | 665.4017 | 778.4218 | 3110.2251 | 1839.8152 | 1163.2488 | 1679.0403
465 | 7.1094 | 1960.4366 | 1921.2536 | 936.5585 | 1895.5683 | 2580.114 | 1432.9872 | 726.2665 | 1509.1317 | 1100.8909 | 640.3756 | 369.6421 | 715 | 3070.4799 | 1268.5939 | 575.0963 | 1854.2536 | 2559.0305 | 1298.4189 | 656.4061 | 1415.1799 | 3009.9823 | 2383.3451 | 1209.8299 | 1785.5506 | 3706.4415 | 2403.9311 | 982.7817 | 2393.878 | 1912.9518 | 1682.0963 | 573.5811 | 1349.3451 | 2310.6646 | 2198.6677 | 683.5152 | 677.6549 | 5703.9726 | 3014.4494 | 1291.8732 | 3178.9037 | 2830.2311 | 2402.4415 | 717.3756 | 1371.3756 | 572.8122 | 565.5329 | 230.2793 | 1020.6726 | 2254.9744 | 957.6293 | 803.5152 | 768.0482 | 2412.9744 | 1386.4799 | 692.9262 | 1017.9872 | 2695.6805 | 1184.6646 | 910.0177 | 1171.1091 | 1090.8475 | 589.5634 | 451.2232 | 542.7384 | 2874 | 1045.5634 | 901.9695 | 1625.5378 | 2150.9439 | 1397.0482 | 1124.297 | 1626.5939 | 2613.1927 | 2091.2055 | 895.7689 | 1931.4671 | 2596.2793 | 2012.7994 | 863.5585 | 1963.6726 | 1318.9134 | 1004.703 | 537.7207 | 1015.5457 | 937.236 | 635.297 | 490.1701 | 511.061 | 1645.5024 | 1284.4366 | 592.3756 | 1129.8957 | 4496.3964 | 2776.9567 | 1589.5585 | 2742.1927 | 3526.9695 | 2953.2921 | 1379.0225 | 2596.7159 | 397.6067 | 223.1573 | 171.4366 | 397.6067 | 717.947 | 759.0738 | 308.6549 | 657.8604 | 670.1622 | 695.9646 | 285.4671 | 524.5939 | 425.8427 | 414.4061 | 184.5634 | 421.9439 | 424.4671 | 370.6726 | 222.0787 | 314.9518 | 3616.1317 | 2076.1653 | 895.886 | 1234.1573 | 1599.7817 | 1136.1701 | 671.1396 | 761.8604 | 3125.5281 | 1825.6116 | 1151.7994 | 1635.1445
466 | 7.1204 | 1956.1313 | 1979.2493 | 954.8614 | 1864.4159 | 2647.1255 | 1494.2493 | 712.9852 | 1427.2035 | 1089.2699 | 640.8614 | 366.146 | 697.8318 | 3052.9483 | 1283.0221 | 605.059 | 1827.1239 | 2500.4808 | 1274.233 | 661.292 | 1341.0796 | 2990.1017 | 2381.5767 | 1198.708 | 1744.5324 | 3741.0738 | 2413.7818 | 967.4085 | 2379.854 | 1825.0058 | 1669.1386 | 534.1017 | 1304.6563 | 2344.5266 | 2192.8245 | 695.0148 | 682.2847 | 5594.7712 | 3044.0221 | 1260.3716 | 3261.9647 | 2855.1608 | 2415.6136 | 703.5546 | 1368.2847 | 571.1386 | 527.525 | 240.8614 | 967.9262 | 2241.5841 | 1001.0886 | 820.0221 | 787.3216 | 2367.3938 | 1350.3495 | 693.8835 | 971.233 | 2622.8524 | 1197.0664 | 901.2847 | 1126.2182 | 1113.1977 | 560.9631 | 449.0221 | 538.4454 | 3002.7612 | 1004.8023 | 862.5177 | 1610.7301 | 2161.0517 | 1368.3938 | 1104.6858 | 1579.5029 | 2600.1534 | 2087.0074 | 915.1903 | 1938.7153 | 2505.9984 | 1963.6563 | 891.8761 | 1945.9779 | 1346.7965 | 982.8245 | 517.1239 | 992.8171 | 894.5177 | 638.5767 | 467.6711 | 483.379 | 1627.4159 | 1276.5546 | 597.8687 | 1124.3068 | 4391.2166 | 2701.4808 | 1565.6563 | 2628.2905 | 3436.7285 | 2959.8392 | 1375.6062 | 2655.2124 | 361.5029 | 224.7227 | 182.1534 | 361.5029 | 733.4159 | 771.708 | 278.233 | 621.8097 | 670.6062 | 694.1165 | 309.1682 | 531.8687 | 416.4085 | 373.5103 | 196.7375 | 410.5915 | 426.708 | 409.4897 | 220.7153 | 301.4012 | 3576.4085 | 1967.5161 | 903.9852 | 1202.1017 | 1587.2699 | 1149.4381 | 714.3437 | 742.2625 | 3107.9557 | 1862.7818 | 1141.2847 | 1595.6784
467 | 7.1314 | 2045.4092 | 2026.1908 | 912.3421 | 1851.5408 | 2614.0579 | 1539.7355 | 705.3618 | 1431.9224 | 1040.6145 | 636.9066 | 330.6184 | 637.1553 | 3004.0737 | 1311.5566 | 606.4474 | 1806.7197 | 2473.8171 | 1273.0974 | 650.6303 | 1346.5566 | 2938.9803 | 2415.0158 | 1152.6145 | 1669.2408 | 3768.2763 | 2386.1592 | 982.3737 | 2301.8711 | 1850.5763 | 1708.7473 | 536.0974 | 1259.8948 | 2254.579 | 2213.8329 | 718.646 | 682 | 5560.65 | 3017.2566 | 1256.6421 | 3260.4395 | 2784.3264 | 2469.2026 | 685.9026 | 1323.6145 | 586.7355 | 516.4553 | 247.821 | 937.8132 | 2215.1671 | 990.8013 | 828.9105 | 748.2487 | 2352.179 | 1326.0895 | 647.8829 | 965.0039 | 2605.5566 | 1150.2408 | 860.2527 | 1148.0197 | 1146.9184 | 564.0974 | 436.0816 | 517.5329 | 3021.2566 | 1025.6539 | 880.7434 | 1631.0079 | 2115.3342 | 1388.8329 | 1113.0079 | 1497.6961 | 2529.5093 | 2072.9921 | 904.7158 | 1909.8948 | 2442.2605 | 1966.1908 | 879.5369 | 1934.179 | 1314.4277 | 981.2763 | 499.6303 | 975.4513 | 894.5526 | 648.0039 | 464.7316 | 470 | 1586.1632 | 1249.4434 | 571.8948 | 1125.6342 | 4355.1013 | 2807.246 | 1614.0315 | 2723.4289 | 3337.9685 | 3040.8605 | 1411.1947 | 2654.354 | 356.0974 | 251.7395 | 155.1671 | 356.0974 | 724.9066 | 748.8987 | 289.4671 | 630.8289 | 686.5487 | 644.0698 | 305.0855 | 500.8013 | 427.0079 | 389.0158 | 201.5447 | 386.1671 | 417.7237 | 407.8092 | 215.179 | 294 | 3571.7276 | 2044.9736 | 905.821 | 1217.9224 | 1544.4356 | 1134.5369 | 741.6421 | 756.3737 | 3126.5605 | 1880.2724 | 1134.4513 | 1612.9224
468 | 7.1424 | 2078.5202 | 1965.6671 | 915.238 | 1805.0845 | 2614.1574 | 1491.9857 | 732.857 | 1390.9454 | 1078.7542 | 632.0403 | 364.355 | 662.8752 | 2961.4031 | 1316.3004 | 623.9376 | 1789.1612 | 2570.6697 | 1260.2419 | 640.8205 | 1324.6632 | 2895.883 | 2459.5566 | 1161.8973 | 1603.5644 | 3641.5136 | 2411.9155 | 975.9415 | 2297.4174 | 1805.1872 | 1692.9233 | 580.9558 | 1275.6372 | 2226.5384 | 2163.5865 | 684.5644 | 642.0442 | 5518.4655 | 3031.6372 | 1252.8609 | 3209.7841 | 2673.1287 | 2460.5826 | 692.8388 | 1292.2016 | 594.4798 | 549.5566 | 227.8023 | 929.5605 | 2235.8167 | 1026.0949 | 896.593 | 758.117 | 2337.4213 | 1371.6151 | 666.476 | 919.9051 | 2599.2016 | 1159.2965 | 893.0949 | 1121.5644 | 1147.3407 | 567.5202 | 435.6593 | 513.1795 | 3045.9961 | 1039 | 870.8023 | 1660.6775 | 2085.6008 | 1408.8791 | 1131.5787 | 1486.5787 | 2514.5787 | 2075.1391 | 879.7217 | 1841.2861 | 2424.1209 | 2009.2562 | 824.1651 | 1816.5721 | 1311.0182 | 938.1248 | 501.6593 | 968.5605 | 894.0806 | 650.2601 | 511.6554 | 487.7581 | 1586.7984 | 1291.3147 | 605.3953 | 1172.0949 | 4271.2496 | 2768.6892 | 1546.9896 | 2654.6125 | 3485.6619 | 3070.0403 | 1350.7074 | 2638.8609 | 360.2601 | 238.0624 | 181.736 | 360.2601 | 718.5605 | 766.1574 | 307.8791 | 610.143 | 652.0039 | 675.355 | 279.5423 | 513.8973 | 411.5423 | 378.8023 | 196.5605 | 386.8388 | 409.8205 | 418.7581 | 211.5202 | 282.9012 | 3541.6632 | 1935.6567 | 874.7035 | 1177.6853 | 1543.3186 | 1125.0403 | 720.8426 | 757.3407 | 3063.2679 | 1768.2718 | 1141.619 | 1544.5682
469 | 7.1534 | 2073.4867 | 2066.8674 | 967.885 | 1766.6814 | 2562.8229 | 1507.6018 | 724.9469 | 1412.8673 | 1086.7168 | 658.8319 | 381 | 663.2035 | 2858.4778 | 1339.823 | 653.062 | 1742.1238 | 2570.3982 | 1206.6106 | 622.4513 | 1327.5398 | 2790.4689 | 2346.8848 | 1221.3983 | 1635.6372 | 3648.6549 | 2403.6814 | 971 | 2225.5132 | 1764.4513 | 1727.885 | 616.062 | 1248.8938 | 2195.8938 | 2152.531 | 685.0443 | 637.0266 | 5432.3008 | 3110.2036 | 1261.2832 | 3171.9292 | 2718.4691 | 2445.2301 | 665.407 | 1247.8849 | 566.4159 | 529.407 | 214.7345 | 924.9911 | 2155.9822 | 1006.1416 | 850.7964 | 743.177 | 2331.4956 | 1404.5487 | 672.4779 | 957.9116 | 2563.9115 | 1210.8673 | 884.4247 | 1098.7257 | 1084.8229 | 553.4602 | 466.5841 | 507.4779 | 2996.5663 | 1084.1328 | 842.6902 | 1557.6724 | 2109.3275 | 1386.4247 | 1087.8584 | 1451.8849 | 2534.7965 | 2053.177 | 883.7788 | 1830.7788 | 2446.5753 | 1989.6548 | 862.6638 | 1774.4956 | 1268.8584 | 922 | 522.8053 | 943.6814 | 881.469 | 648.4956 | 489.3893 | 460.1504 | 1567.9292 | 1286.4336 | 580.3805 | 1116.0353 | 4190.1061 | 2759.8142 | 1581.947 | 2548.3097 | 3515.1593 | 3047.1858 | 1360.6018 | 2566.7964 | 351.7257 | 230.7522 | 170.177 | 351.7257 | 716.2478 | 750.9292 | 317.7699 | 590.2212 | 641.2566 | 664.1681 | 288.5487 | 478.6194 | 418.2921 | 401.8407 | 188.2301 | 398.5221 | 405.7434 | 412.2124 | 215.5133 | 279 | 3464.8052 | 1926.6372 | 895.3452 | 1123.8938 | 1555.5221 | 1099.177 | 711.2478 | 721.1504 | 2930.1857 | 1782.4071 | 1103.6283 | 1531.2921
470 | 7.1644 | 2151.5084 | 2114.8805 | 953.9488 | 1763.587 | 2517.0069 | 1507.5324 | 711.355 | 1403.8294 | 1030.2492 | 635.8908 | 368.7679 | 645.0034 | 2947.2627 | 1328.1229 | 665.3515 | 1731.8225 | 2447.3824 | 1209.3481 | 630.7611 | 1290.7168 | 2765.8805 | 2287.6519 | 1229.5905 | 1608.1878 | 3677.4676 | 2457.9248 | 988.5836 | 2125.6076 | 1718.4813 | 1642.0854 | 525.1435 | 1184.2492 | 2225.2832 | 2122.3618 | 676.2389 | 646.8805 | 5390.8874 | 3023.6759 | 1308.5767 | 3106.6622 | 2697.8942 | 2389.4847 | 690.9317 | 1203.6553 | 544.0034 | 534.2901 | 213.7645 | 955.3447 | 2086.2458 | 1033.7542 | 831.058 | 730.413 | 2267.5461 | 1377.8908 | 660.0614 | 888.8465 | 2557.058 | 1220.942 | 894.3481 | 1086.8259 | 1048.1809 | 600.2218 | 498.1706 | 476.7133 | 3132.7234 | 1089.0614 | 894.8668 | 1477.4233 | 2011.7339 | 1425.3993 | 1109.4607 | 1390.8362 | 2530.8259 | 1993.7782 | 893.8805 | 1769.0171 | 2380.8431 | 1974.413 | 868.0034 | 1803.8157 | 1225.4778 | 922 | 490.7748 | 944.4095 | 907.0512 | 665.5836 | 481.587 | 465.8191 | 1554.6485 | 1233.8942 | 615.1638 | 1084.355 | 4141.1263 | 2792.2286 | 1568.8328 | 2496.4198 | 3399.4984 | 2981.1332 | 1355.9454 | 2573.0512 | 363.5256 | 243.2321 | 164.2935 | 363.5256 | 700.7099 | 763.6416 | 289.4198 | 578.5905 | 642.7645 | 645.8259 | 294 | 495.3447 | 413.0614 | 394.4164 | 211.9931 | 382.6519 | 402.7065 | 393.4744 | 230.7611 | 272.884 | 3452.1741 | 1919.8294 | 873.1263 | 1080.6553 | 1512.8942 | 1062.7133 | 733.9351 | 702.355 | 2961.5119 | 1723.5496 | 1137.3993 | 1503.8908
471 | 7.1754 | 2211.7326 | 2036.2133 | 955.0983 | 1751.7945 | 2564.9293 | 1476.0354 | 690.3569 | 1282.8111 | 1065.2685 | 667.1702 | 351.7945 | 661.9735 | 2949.0442 | 1302.5801 | 665.4464 | 1595.5072 | 2429.3127 | 1190.1425 | 640.4376 | 1175.463 | 2758.6785 | 2338.7061 | 1199.0354 | 1528.1956 | 3640.6066 | 2351.9359 | 982.3481 | 2161.2597 | 1756.4917 | 1597.0177 | 541.6077 | 1169.5536 | 2241.1072 | 2129.8663 | 703.0718 | 626.4729 | 5460.3491 | 3059.1348 | 1316.5624 | 3077.3481 | 2690.2144 | 2464.6619 | 696.5624 | 1188.5624 | 573.4022 | 507.1514 | 226.4287 | 890.4276 | 2133.8132 | 978.4188 | 845.2055 | 715.0177 | 2265.8663 | 1347.8033 | 677.1967 | 877.5359 | 2559.5536 | 1205.2409 | 889.1249 | 1080.116 | 1032.1249 | 571.7227 | 464.6066 | 468 | 3118.0796 | 1051.0442 | 850.8564 | 1500.5094 | 2011.6254 | 1405.821 | 1176.9293 | 1369.0088 | 2492.2674 | 2040.1436 | 869.589 | 1722.589 | 2357.2232 | 1938.0442 | 818.1691 | 1743.0884 | 1266.7149 | 952.295 | 489.0983 | 931.4641 | 877.937 | 629.0531 | 467.4641 | 478.768 | 1511.0531 | 1184.2674 | 588.8298 | 1042.3834 | 4128.6696 | 2753.3923 | 1526.1514 | 2494.768 | 3340.589 | 2981.6431 | 1370.4199 | 2556.3658 | 358.6785 | 263.3127 | 168.1072 | 358.6785 | 710.7591 | 761.4553 | 298.6431 | 596.9735 | 674.8486 | 679.5094 | 288.5624 | 464.6066 | 399.9017 | 382.7856 | 218.4464 | 373.116 | 372.4818 | 399.8751 | 221.0177 | 285.7591 | 3406.0619 | 1874.937 | 863.7768 | 1083.4287 | 1528.5182 | 1099.0541 | 741 | 703.884 | 2963.3569 | 1727.9646 | 1141.9017 | 1507.9823
472 | 7.1864 | 2203.851 | 2060.3452 | 941.4292 | 1663.5678 | 2575.4764 | 1475.1018 | 687.3673 | 1260.2581 | 1007.4041 | 633.233 | 373.2508 | 634.0693 | 2856.935 | 1205.466 | 683.9381 | 1570.9926 | 2381.9203 | 1237.447 | 671.9853 | 1250.9485 | 2737.0619 | 2326.5929 | 1205.2036 | 1482.1711 | 3475.1282 | 2387.0178 | 971.1091 | 2095.935 | 1762.32 | 1593 | 543.1637 | 1086.3568 | 2128.3716 | 2152.3599 | 656.764 | 625.5236 | 5376.2624 | 3113.7198 | 1281.0693 | 3083.469 | 2600.466 | 2429.4513 | 709.2036 | 1183.0546 | 605.8835 | 505.8909 | 230 | 864.2655 | 2147.0546 | 996.0871 | 820.5929 | 681.8038 | 2294.6726 | 1360.149 | 634.0766 | 833.8657 | 2380.8774 | 1171.0147 | 908.8835 | 1074.2655 | 1049.5162 | 558.2109 | 447.4764 | 487.7272 | 3115.7817 | 988.1312 | 866.1416 | 1456.553 | 1953.9277 | 1410.9926 | 1158.2802 | 1366.2109 | 2415.7167 | 2095.8761 | 954.3231 | 1747.3525 | 2310.4439 | 1905.9056 | 835.7744 | 1694.5929 | 1254.1711 | 941.2728 | 553.5487 | 932.5236 | 882.7817 | 617.7891 | 492.1962 | 437.6002 | 1470.5929 | 1247.385 | 584.1018 | 1040.469 | 4028.5751 | 2777.8761 | 1586.4396 | 2408.6223 | 3312.4838 | 3034.3452 | 1402.8289 | 2473.2477 | 348.8982 | 242.7492 | 165.0546 | 348.8982 | 655.8185 | 740.851 | 291.3746 | 571.4365 | 644.5457 | 644.233 | 325.6652 | 421.4365 | 388.32 | 396.7817 | 214.8437 | 386.9926 | 383.7272 | 410.1018 | 208.32 | 279.7419 | 3296.5205 | 1830.0693 | 881.3599 | 1091.7345 | 1485.7094 | 1102.531 | 717.3274 | 703.4218 | 2991.9307 | 1779.3998 | 1121.6401 | 1504.8982
473 | 7.1974 | 2254.295 | 2142.7198 | 951.4248 | 1628.9793 | 2610.8634 | 1450.5545 | 731.2744 | 1253.3835 | 1010.4386 | 626.6096 | 359.1642 | 654.651 | 2887.0895 | 1269.9529 | 606.4581 | 1604.651 | 2462.5625 | 1203.9174 | 656.7601 | 1194.054 | 2710.3628 | 2309.3835 | 1195.3766 | 1458.9725 | 3522.9529 | 2304.2251 | 973.0069 | 2110.062 | 1638.1906 | 1559.3421 | 535.1917 | 1132.1171 | 2115.6303 | 2151.3904 | 670.0482 | 595.7463 | 5257.4443 | 2935.4777 | 1238.3421 | 3057.1504 | 2576 | 2392.9656 | 704.9862 | 1189.2124 | 592.7669 | 503.9931 | 246.8289 | 893.4523 | 2054.643 | 996.3835 | 853.8703 | 702.8496 | 2252.9174 | 1343.3628 | 679.5006 | 917.1653 | 2465.2273 | 1148.5752 | 947.0551 | 1070.5959 | 1002.9105 | 600.473 | 461.2262 | 492.1986 | 3169.6854 | 1044.5212 | 919.0758 | 1397.9242 | 1931.1917 | 1426.2193 | 1132.1711 | 1440.5281 | 2310.6499 | 2085.1642 | 926.1091 | 1658.232 | 2273.1573 | 1879.9656 | 829.5752 | 1687 | 1234.9793 | 907.1504 | 527.527 | 876.3008 | 844.527 | 597.1642 | 472.7532 | 486.1033 | 1472.6165 | 1230.938 | 599.6234 | 1001.3284 | 4007.6096 | 2785.5959 | 1587.3697 | 2335.3146 | 3399.1584 | 3020.5545 | 1410.8014 | 2481.0482 | 353.411 | 245.6165 | 146.3697 | 353.411 | 680.0689 | 767.2537 | 283.1917 | 581.4317 | 623.5821 | 644.822 | 288.7187 | 465.0895 | 402.0275 | 394.589 | 209.1917 | 388.5959 | 420.2537 | 391.9656 | 236.4523 | 265.7807 | 3284.6234 | 1847.4455 | 851.5408 | 1017.6705 | 1458.3766 | 1100.8014 | 731.8358 | 709.411 | 2867.17 | 1819.8496 | 1126.6165 | 1501.3972
474 | 7.2084 | 2278.5777 | 2111.1799 | 948.4907 | 1599.1489 | 2540.2605 | 1432.6086 | 737.9317 | 1222.5216 | 953.3475 | 628.8137 | 376.7827 | 609.1115 | 2782.0181 | 1305.0187 | 647.8388 | 1552.6022 | 2388.3598 | 1176.5403 | 628.2173 | 1183.5777 | 2678.9626 | 2293.9813 | 1183.0496 | 1458.2547 | 3394.4713 | 2320.4971 | 962.6086 | 2094.5899 | 1621.0187 | 1524.9626 | 572.2424 | 1116.4533 | 2131.3914 | 2056.4281 | 659.913 | 621.3604 | 5120.2727 | 2996.4042 | 1244.646 | 3041.8633 | 2491.3785 | 2396.9504 | 702.1863 | 1152.7576 | 553.0123 | 507.0683 | 234.7266 | 872.5216 | 2038.5093 | 973.6583 | 837.1489 | 720.5777 | 2268.6647 | 1339.6273 | 668.7763 | 802.6764 | 2405.2418 | 1140.118 | 920.0123 | 1034.1986 | 977.795 | 579.2669 | 465 | 480.6086 | 3124.2295 | 1022.9439 | 886.8756 | 1382.9317 | 1819.3411 | 1388.3166 | 1165.4288 | 1291.3843 | 2371.9942 | 2065.354 | 897.5403 | 1701.5345 | 2245.8446 | 1897.7827 | 852.0374 | 1616.2108 | 1254.7827 | 941.4971 | 502.354 | 866.0683 | 823.4223 | 587.9317 | 463.559 | 461.9439 | 1492.9007 | 1194.3353 | 587.5403 | 981.236 | 3962.6209 | 2804.528 | 1512.2108 | 2445.8639 | 3321.1051 | 3070.1431 | 1416.6957 | 2480.677 | 314.3166 | 244.7453 | 172.9194 | 314.3166 | 641.9936 | 737.5712 | 268.9813 | 555.0806 | 592.5216 | 647.1863 | 329.8885 | 437.3166 | 379.9626 | 384.0496 | 201.4907 | 370.913 | 389.7576 | 411.4101 | 236.677 | 267.882 | 3211.5152 | 1779.9562 | 869.0374 | 1065.7208 | 1415.1302 | 1074.9626 | 747.5777 | 684.9626 | 2886.0748 | 1684.608 | 1132.2547 | 1451.3662
475 | 7.2194 | 2304.1269 | 2024.0117 | 919.7433 | 1590.5221 | 2482.3539 | 1496.0768 | 724.6106 | 1211.8702 | 923.3068 | 616.6106 | 325.8347 | 573.8731 | 2736.8289 | 1360.8614 | 660.3481 | 1484.4867 | 2397.3865 | 1117.6607 | 659.3422 | 1175.2625 | 2634.1799 | 2283.5664 | 1230.5576 | 1428.4395 | 3371.5634 | 2309.351 | 978.1711 | 2030.3569 | 1624.826 | 1519.826 | 545.4837 | 1085.0472 | 2090.2241 | 2102.7287 | 679.9528 | 598.4395 | 5125.3865 | 2902.8878 | 1311.5989 | 2952.4483 | 2520.7316 | 2365.1357 | 716.0413 | 1124.177 | 552.4336 | 522.8673 | 226.0442 | 812.3126 | 1976.4011 | 962.3923 | 788.0501 | 679.2241 | 2191.5781 | 1326.7846 | 687.7788 | 795.7788 | 2362.5251 | 1173.6903 | 878.9616 | 950.8377 | 1001.4307 | 556.3068 | 446.0029 | 445.7876 | 3174.5989 | 1002.4366 | 846.4395 | 1318.4011 | 1865.0325 | 1287.3185 | 1149.2212 | 1316.5989 | 2341.4424 | 2032.2655 | 889.8702 | 1599.7137 | 2212.0914 | 1915.3894 | 844.7846 | 1642.9499 | 1195.5752 | 877.4896 | 480.8289 | 900.8643 | 807.7846 | 596.0856 | 448.1327 | 427.3953 | 1438.3569 | 1158.2655 | 604.649 | 995.5192 | 3816.5426 | 2749.5309 | 1582.7258 | 2317.0677 | 3308.7375 | 3057.3953 | 1351.9232 | 2455.0472 | 359.5133 | 253.9115 | 159.351 | 359.5133 | 619.4366 | 716.6106 | 292.4307 | 570.3009 | 605.823 | 629.6548 | 312.2655 | 454.4307 | 360.7846 | 379.5221 | 194.2183 | 371.1298 | 405.7788 | 368.2684 | 263.0826 | 274.7817 | 3165.0914 | 1748.9587 | 866.7404 | 1009.9675 | 1348.1828 | 1037.6135 | 756.6077 | 650.0914 | 2867.4395 | 1786.6314 | 1139.6077 | 1453.2154
476 | 7.2304 | 2320.7355 | 2050.7508 | 971.8393 | 1573.235 | 2494.9563 | 1452.999 | 749.6623 | 1198.4262 | 965.2655 | 600.588 | 365.1333 | 589.118 | 2736.353 | 1308.293 | 639.882 | 1542.5748 | 2330.2339 | 1134.5005 | 664.4852 | 1102.5869 | 2553.0722 | 2193.9837 | 1216.6907 | 1366.6755 | 3277.6012 | 2258.8962 | 942.6022 | 1898.9684 | 1573.8667 | 1536.765 | 557.2798 | 1052.3377 | 2052.4995 | 2121.1913 | 695.8973 | 565.176 | 5159.6623 | 2966.0163 | 1307.3967 | 2893.764 | 2469.8077 | 2264.9542 | 709.7792 | 1074.7345 | 584.177 | 533.5443 | 225 | 826.1475 | 2017.4863 | 984.471 | 773.9705 | 673.353 | 2107.7782 | 1333.2208 | 678.7497 | 829.5005 | 2312.7345 | 1186.0295 | 878.706 | 985.295 | 956.705 | 567.2503 | 430.2645 | 458.2798 | 3120.1017 | 976.529 | 824.0732 | 1223.6897 | 1872.4557 | 1345.1191 | 1140.647 | 1267.9695 | 2279.0285 | 2006.7202 | 866.3672 | 1513.8077 | 2176.6612 | 1867.705 | 853.7355 | 1613.4405 | 1155.176 | 898.883 | 518.9125 | 816.6307 | 836.0153 | 643.2655 | 433.2645 | 437.9268 | 1428.5148 | 1110.0875 | 610.6765 | 1008.2208 | 3830.7508 | 2699.2787 | 1542.3225 | 2248.7935 | 3165.1444 | 2965.6602 | 1343.0295 | 2480.177 | 359.2645 | 226.6612 | 148.294 | 359.2645 | 630.412 | 715.6765 | 297.1617 | 524.705 | 610.8383 | 605.882 | 306 | 432.3377 | 353.8087 | 347.9847 | 193.8383 | 382.8973 | 401.7792 | 362.1913 | 233.7935 | 259.235 | 3074.3367 | 1725.882 | 943.7956 | 969.8525 | 1354.4415 | 1005.0142 | 701.8372 | 644.8383 | 2830.823 | 1799.9115 | 1131.7792 | 1363.7924
477 | 7.2414 | 2271.117 | 2080.2635 | 954.0816 | 1507.0599 | 2463.2772 | 1468.3668 | 715.8751 | 1119.4512 | 953.587 | 615.0108 | 346.0816 | 559.8289 | 2703.8289 | 1303.6549 | 604.53 | 1458.1877 | 2215.6361 | 1117.0354 | 657.1957 | 1046.4729 | 2408.8671 | 2315.6383 | 1190.7365 | 1371.3098 | 3327.0433 | 2229.587 | 941.8043 | 1910.0217 | 1476.3942 | 1543.4022 | 569.5054 | 1028.2881 | 2008.7256 | 1999.522 | 695.4484 | 571.057 | 4910.688 | 2891.992 | 1298.0462 | 2841.7718 | 2394.2094 | 2231.6902 | 691.8397 | 1035.3805 | 493.0382 | 505.231 | 254.769 | 776.5653 | 1954.0029 | 996.6549 | 829.1358 | 682.5054 | 2067.7827 | 1333.2989 | 691.3098 | 775.4621 | 2275.9321 | 1163.1848 | 924.2282 | 957.5762 | 952.1033 | 602.3206 | 463.7228 | 429.6794 | 3002.3834 | 1022.182 | 853.3206 | 1199.4946 | 1856.5408 | 1325.9783 | 1139.1495 | 1224.53 | 2297.0679 | 1991.943 | 927.4917 | 1576.8477 | 2070.6361 | 1814.6224 | 809.2202 | 1512.4404 | 1117.6794 | 900.0462 | 518.4946 | 836.4238 | 810.53 | 616.2772 | 429.2989 | 402.7256 | 1430.7011 | 1122.413 | 605.0462 | 963.2202 | 3794.7718 | 2664.7827 | 1567.5733 | 2240.2989 | 3142.9538 | 2931.9892 | 1304.8751 | 2436.6686 | 360.4022 | 203.9892 | 156.356 | 360.4022 | 633.8505 | 680.2772 | 315.7119 | 476.7256 | 568.4729 | 597.7473 | 294.0924 | 406.587 | 370.8614 | 377.7228 | 205.057 | 342.4729 | 393.1957 | 347.6902 | 213.3913 | 257.7011 | 2927.8671 | 1710.943 | 919.8751 | 1021.1358 | 1327.3805 | 1059.5379 | 704.6087 | 682.4238 | 2796.0816 | 1784.2419 | 1143.6087 | 1376.6195
478 | 7.2524 | 2314.6328 | 1946.8096 | 930.0177 | 1450.2699 | 2289.6105 | 1473 | 708.1372 | 1129.2965 | 921.5265 | 580.0354 | 339.2743 | 548.823 | 2729.9248 | 1295.5088 | 661.124 | 1440.1372 | 2254.947 | 1096.6062 | 668.0797 | 1026.9203 | 2407.5708 | 2213.1636 | 1168.0177 | 1259.2432 | 3279.4645 | 2202.7035 | 1012.8895 | 1831.4424 | 1439.4292 | 1503.4469 | 573.5885 | 980.9955 | 2045.1416 | 1988.354 | 658.761 | 524.6814 | 4978.1683 | 2836.3097 | 1359.124 | 2821.9203 | 2348.4867 | 2241.9425 | 702.8053 | 1065.3761 | 507.062 | 504.3142 | 218.584 | 735.0752 | 1816.7521 | 1007.4912 | 845.0398 | 660.7035 | 2018.1327 | 1343.354 | 716.4337 | 778.8053 | 2266.6858 | 1217.6726 | 892.3097 | 928.5663 | 940.9203 | 565.7212 | 439.8009 | 432.6283 | 3076.3231 | 970.6017 | 954.7744 | 1219.5708 | 1810.8584 | 1346.7478 | 1124.3318 | 1189.6637 | 2206.2256 | 1952.8982 | 919.1548 | 1501.1282 | 1981.3849 | 1771.6239 | 882.9691 | 1434.7389 | 1118.9027 | 876.5663 | 508.3717 | 809.3495 | 815.354 | 572.8982 | 455.7478 | 410.6682 | 1370.6017 | 1140.6682 | 614.354 | 893.7389 | 3633.4158 | 2617.7212 | 1573.1372 | 2146.8141 | 3094.8185 | 3034.2656 | 1334.239 | 2338.2654 | 352.3717 | 242.416 | 178.708 | 352.3717 | 640.0398 | 698.1593 | 292.2522 | 470 | 554.0973 | 626.3363 | 296.3142 | 435.7876 | 356.7433 | 343.4469 | 193.1947 | 343.6283 | 350.584 | 370.0221 | 202.3717 | 251.0973 | 2922.2567 | 1690.8805 | 890.5663 | 966.2875 | 1367.7301 | 1040.6637 | 736.3363 | 637.23 | 2650.3583 | 1775.0973 | 1126.1548 | 1312.1105
479 | 7.2634 | 2330.001 | 2041.3915 | 885.8692 | 1404.4946 | 2299.7552 | 1411.7413 | 720.2517 | 1067.3657 | 874.9941 | 564.3736 | 341.6254 | 583.005 | 2737.6254 | 1294.8751 | 583.4876 | 1397.1189 | 2149.3588 | 1037.9921 | 652.4975 | 974.2428 | 2386.4965 | 2214.8781 | 1208.7562 | 1222.6224 | 3145.7324 | 2168.3707 | 970.6174 | 1777.7443 | 1468.3796 | 1431.3657 | 535.4946 | 872.4856 | 2015.2448 | 2025.005 | 648.6244 | 500.3726 | 4799.9723 | 2698.7314 | 1314.7423 | 2777.9941 | 2234.4847 | 2232.6234 | 720.7522 | 982.6125 | 501.4985 | 494.4985 | 239.1288 | 719.4985 | 1806.6264 | 987.997 | 812.7453 | 640.3726 | 1974.2448 | 1213.7314 | 659.6165 | 722.3667 | 2153.9842 | 1158.7403 | 936.7572 | 865.4916 | 909.2458 | 546.7483 | 433.2498 | 426.999 | 3047.2438 | 976.7522 | 898.2399 | 1162.6165 | 1715.6125 | 1351 | 1151.7542 | 1193.7512 | 2202.3766 | 1841.9851 | 928.2517 | 1472.1229 | 2029.5084 | 1796.6293 | 814.6135 | 1464.3806 | 1028.9871 | 942.1348 | 512.2507 | 784.7473 | 802.998 | 595.004 | 445.1229 | 411.2498 | 1314.6184 | 1113.2458 | 621.2507 | 890.1259 | 3591.4975 | 2599.6234 | 1588.7522 | 2109.2468 | 3085.2498 | 2958.2369 | 1375.8801 | 2333.6264 | 352.7502 | 220.996 | 164.4975 | 352.7502 | 625.2478 | 702.8751 | 290.6254 | 486.6274 | 622.1348 | 603.8712 | 278.6224 | 431.3736 | 350.4995 | 327.3736 | 202.3766 | 345.8751 | 398.2577 | 360.8731 | 209.7512 | 242.999 | 2826.111 | 1640.7433 | 860.7463 | 927.1209 | 1294.4886 | 1069.2547 | 735.7493 | 593.1199 | 2605.2468 | 1749.4965 | 1169.3816 | 1286.998
480 | 7.2744 | 2307.9272 | 2048.2384 | 950.9932 | 1425.6224 | 2352.9204 | 1387.9139 | 738.8609 | 1068.649 | 886.7483 | 578.9735 | 368.6224 | 510.7949 | 2564.9541 | 1256.8412 | 641.9932 | 1363.4901 | 2054.9075 | 1053.9602 | 667.7483 | 982.9735 | 2262.3116 | 2155.881 | 1218.5497 | 1250.1721 | 3136.8741 | 2155.1259 | 967.4371 | 1826.1323 | 1465.0133 | 1387.3908 | 552.1854 | 904.1456 | 1951.4306 | 1905.7619 | 662.1986 | 559.2316 | 4697.4572 | 2757.2051 | 1287.4769 | 2623.8017 | 2276.682 | 2185.7418 | 683.9537 | 1001.9469 | 502.6622 | 508.0861 | 267.8476 | 741.0728 | 1782.265 | 972.5762 | 824.8609 | 625.5762 | 1922.8544 | 1288.1786 | 680.2847 | 695.3643 | 2073.2187 | 1131.2517 | 920.0398 | 835.702 | 927.1854 | 515.7153 | 453.4105 | 436.649 | 3049.8741 | 1026.033 | 903.0861 | 1181.5099 | 1666.616 | 1303.967 | 1135.5895 | 1123.1194 | 2054.9143 | 1828.7748 | 907.8146 | 1481.5364 | 2003.3908 | 1737.1061 | 859.7945 | 1462.9007 | 1063.033 | 884.5565 | 467.7418 | 719.881 | 755.7418 | 606.0993 | 477.6092 | 383.4901 | 1308 | 1111.6622 | 679.682 | 882.1259 | 3581.0133 | 2611.3112 | 1581.2384 | 2020.8082 | 3171.9667 | 2935.2384 | 1357.0398 | 2302.1657 | 357.4371 | 230.3112 | 177.9735 | 357.4371 | 627.4371 | 703.8874 | 304.3112 | 484.5629 | 577.8677 | 589.351 | 337.2316 | 403.3776 | 342.0133 | 357.9469 | 176.4901 | 348.6622 | 383.8146 | 342.1391 | 206.5629 | 240.2252 | 2777.3908 | 1621.5762 | 864.9867 | 916.7881 | 1251.9405 | 995.0201 | 712.8146 | 623.4966 | 2587.8014 | 1751.3245 | 1171.5629 | 1383.503
481 | 7.2854 | 2322.0944 | 1953.4306 | 938.4071 | 1344.4276 | 2242.9379 | 1371.6047 | 745.4985 | 1008.82 | 875.5044 | 536.9144 | 348.6076 | 531.59 | 2478.2212 | 1257.3982 | 622.2094 | 1289.8229 | 2109.7788 | 1069.5959 | 657.4041 | 940.0147 | 2288.3865 | 2098.5162 | 1232.4956 | 1263.8968 | 3157.7935 | 2084.7226 | 1018.3835 | 1726.8347 | 1432.5103 | 1378.5015 | 511.8141 | 925.295 | 1916.1091 | 1878.3038 | 698.1888 | 544.5074 | 4661.0088 | 2794.8909 | 1268.8053 | 2586.1062 | 2167.9379 | 2185.3982 | 688.8968 | 938.5221 | 529.9912 | 492.9056 | 267.4012 | 704.4129 | 1787.9971 | 998.8909 | 789.2124 | 591.6106 | 1959.2862 | 1277.5074 | 642.6135 | 669.6076 | 2081.1947 | 1171.2862 | 902.6047 | 816.8053 | 897.6106 | 560.5841 | 447.0029 | 428.1032 | 3024.9085 | 1024.8024 | 896.0029 | 1100.4276 | 1606.2182 | 1284.5044 | 1120.4041 | 1149.0885 | 2026.1032 | 1865.8879 | 862.7138 | 1341.7462 | 1859.5456 | 1710.1062 | 838.2094 | 1379.2271 | 1058.2035 | 877.7994 | 500.6874 | 720.9971 | 733.8053 | 649.2862 | 476.6018 | 396.1947 | 1304.4012 | 1045.4218 | 623.1209 | 892.6962 | 3616.8879 | 2522.1297 | 1561.1062 | 1984.8082 | 3048.0441 | 2907.9085 | 1421.4779 | 2257.5132 | 317.5132 | 218.5044 | 183.5988 | 317.5132 | 583.0147 | 669.8112 | 291.6047 | 488.4985 | 563.8024 | 573.6047 | 300.9144 | 398.2006 | 351.7965 | 346.705 | 202.6903 | 328.3068 | 373.8024 | 367.8909 | 210.4985 | 236.4012 | 2757.705 | 1609.2035 | 872.2979 | 868.3156 | 1274.9912 | 1006.5929 | 751.3865 | 585.7138 | 2577.0029 | 1714.2124 | 1116.1179 | 1293.4335
482 | 7.2964 | 2281.1361 | 2055.1756 | 1004.3998 | 1310.3761 | 2195.3441 | 1371.824 | 691.2802 | 1001.088 | 876.736 | 593.1038 | 383.3919 | 543.208 | 2509.3039 | 1257.088 | 650.0079 | 1265.5841 | 2117.912 | 1020.8402 | 625.9041 | 953.2399 | 2208.1843 | 2046.4882 | 1119.0884 | 1161.0323 | 3063.3283 | 2090.68 | 1003.9361 | 1768.8078 | 1404.3761 | 1375.264 | 552.5998 | 849.4803 | 1817.2403 | 1844.1681 | 681.0241 | 529.232 | 4374.369 | 2755.1361 | 1330.8398 | 2573.056 | 2101.1922 | 2126.7202 | 653.5201 | 902.7281 | 551.2399 | 427.1602 | 239.6401 | 696.352 | 1716.0402 | 971.9041 | 864.3437 | 610.7999 | 1889.2163 | 1274.088 | 687.2478 | 710.7759 | 2062.9361 | 1193.3279 | 901 | 860.5998 | 890.352 | 554.144 | 449.648 | 387.7841 | 2946.3043 | 1006.6721 | 875.8481 | 1042.6642 | 1603.648 | 1303.0639 | 1168.2478 | 1153 | 2025.912 | 1825.3122 | 886.2719 | 1355.1839 | 1918.7837 | 1734.4479 | 821.32 | 1368.176 | 1010.4882 | 870.704 | 505 | 670.9282 | 776.6878 | 586.5122 | 476.912 | 377.0241 | 1273.9041 | 1095.4558 | 605.056 | 872.1121 | 3428.5686 | 2528.4159 | 1602.7759 | 1979.264 | 2935.4163 | 2990.7277 | 1397.9921 | 2309.5438 | 313 | 213.352 | 181.264 | 313 | 596.2399 | 690.6239 | 276.32 | 482.616 | 549.32 | 564.704 | 290.528 | 382.4961 | 349.352 | 345.912 | 180.4641 | 342.4159 | 377.56 | 351.8481 | 219.208 | 241.472 | 2626.4964 | 1571.5201 | 836.5201 | 873.032 | 1268.88 | 987.1121 | 730.4641 | 589.208 | 2604.2719 | 1702.704 | 1144.6559 | 1298.4159
483 | 7.3074 | 2327.8358 | 1965.2526 | 980.4985 | 1265.4828 | 2105.1169 | 1404.3501 | 692.47 | 982.5142 | 847.4228 | 556.4828 | 363.8928 | 563.41 | 2427.9655 | 1183.9812 | 641.9085 | 1217.7856 | 2067.1642 | 1003.9843 | 668.2901 | 967.94 | 2160.2556 | 2116.8673 | 1171.7759 | 1149.1514 | 2904.2653 | 2078.1357 | 944.6941 | 1570.6566 | 1304.0255 | 1351.8928 | 558.8486 | 834.6057 | 1849.593 | 1886.2901 | 727.9873 | 462.3756 | 4395.9873 | 2706.6342 | 1258.4355 | 2554.4385 | 2050.7099 | 2121 | 686.9715 | 844.5427 | 506.7856 | 466.2901 | 262.8801 | 698.7729 | 1703.4543 | 975.47 | 829.4828 | 578.8013 | 1842.2556 | 1324.8358 | 679.06 | 638.2841 | 2033.2714 | 1139.5427 | 890.8328 | 836.3471 | 864.1199 | 520.6499 | 480.5015 | 384.9243 | 2926.06 | 1029.9558 | 901.7286 | 991.7856 | 1569.8013 | 1266.1799 | 1119.3913 | 1118.8013 | 1987.1799 | 1817.3028 | 865.8928 | 1320.1042 | 1803.994 | 1682.467 | 796.8928 | 1269.1012 | 1045.7444 | 907.8958 | 495.7571 | 660.4543 | 729.2399 | 595.7129 | 468.6814 | 401.8044 | 1246.9685 | 1031.6784 | 617.8643 | 811.7699 | 3329.5869 | 2491.1799 | 1573.7256 | 1923.5427 | 2828.9498 | 2909.3441 | 1406.0915 | 2118.1267 | 307.4543 | 223.1672 | 173.6057 | 307.4543 | 534.2241 | 640.3156 | 290.7129 | 505.1072 | 511.9528 | 560.3028 | 279.8328 | 384.6972 | 317.5742 | 335.8328 | 181.6972 | 341.2271 | 309.6027 | 351.8486 | 226.47 | 266.0315 | 2658.3658 | 1529.1799 | 881.9873 | 834.2556 | 1164.4797 | 971.1357 | 726.1514 | 601.0915 | 2578.3471 | 1738.9715 | 1129.5142 | 1297.2271
484 | 7.3184 | 2310.4587 | 2021.624 | 948.966 | 1251.6976 | 2070.8392 | 1372.3466 | 703.3024 | 945.41 | 817.8392 | 544.5708 | 339.5221 | 523.7905 | 2328.2787 | 1238.8776 | 604.4734 | 1179.3466 | 1983.3908 | 995.5074 | 609.2492 | 943.7124 | 2122.3466 | 2032.1518 | 1177 | 1067.5176 | 2896.6829 | 2051.7124 | 1018.6726 | 1598.9558 | 1337.2095 | 1292.8687 | 575.8584 | 828.3805 | 1786.5029 | 1806.6445 | 707.6489 | 447.6342 | 4283.8642 | 2717.9853 | 1280.0974 | 2523.0295 | 1880.2889 | 2077.9174 | 748.0679 | 833.4439 | 499.2537 | 435.3466 | 236.9026 | 695.2537 | 1702.0634 | 978.8097 | 815.6976 | 569.4439 | 1752.8347 | 1228.786 | 677.0634 | 627.3171 | 1944.8347 | 1101.2832 | 927.4632 | 807.7758 | 857.3171 | 536.7316 | 482.0634 | 409.3511 | 2824.786 | 960.82 | 840.3126 | 1018.9071 | 1572.6195 | 1202.1224 | 1091.5855 | 1044.82 | 1991.4926 | 1779.5368 | 872.4292 | 1280.4734 | 1816.4779 | 1671.4439 | 852.1313 | 1239.4587 | 982.5029 | 809.8495 | 524.9705 | 675.9218 | 694.0929 | 591.3805 | 440.8392 | 375.9026 | 1223.4587 | 1022.2537 | 638.6682 | 799.5074 | 3238.7079 | 2562.9263 | 1589.7316 | 1845.9468 | 2884.6874 | 2829.8834 | 1370.4734 | 2144.1461 | 312.6195 | 202.4587 | 161.7611 | 312.6195 | 520.5708 | 605.0929 | 299.4926 | 445.1858 | 479.966 | 548.7611 | 293.0487 | 347.5368 | 306.5708 | 304.0929 | 204.4779 | 338.1903 | 335.8437 | 325.7758 | 234.4926 | 234.2832 | 2502.7816 | 1487.6003 | 814.82 | 795.41 | 1177.5413 | 917.5516 | 718.5074 | 534.5663 | 2471.1031 | 1716.7124 | 1113.9513 | 1220.2005
485 | 7.3294 | 2279.585 | 1904.5446 | 968.824 | 1204.5054 | 2113.5969 | 1396.5684 | 684.0737 | 910.7384 | 786.585 | 561.0797 | 345.591 | 491.585 | 2276.4542 | 1224.9715 | 632.3638 | 1156.1249 | 1975.1534 | 942.8121 | 650.5458 | 894.5565 | 1992.8514 | 2036.4376 | 1205.4661 | 1049.6647 | 2804.0107 | 1984.528 | 960.4257 | 1530.8347 | 1242.2664 | 1309.8751 | 576.0511 | 834.6421 | 1684.2664 | 1807.6421 | 667.0963 | 485.9037 | 4171.6243 | 2762.648 | 1248.7895 | 2430.8573 | 1972.4268 | 2013.3234 | 687.4768 | 838.6932 | 462.9429 | 474.7503 | 215.0737 | 605.8062 | 1643.1701 | 977.1023 | 875.7277 | 564.2556 | 1672.0393 | 1259.0059 | 686.4887 | 609.9203 | 1873.528 | 1138.8525 | 911.0226 | 740.528 | 838.0226 | 520.9203 | 475.3579 | 412.8977 | 2822.7444 | 964.5398 | 874.9037 | 1033.3353 | 1536.9429 | 1222.6706 | 1108.0285 | 1031.4602 | 1830.6921 | 1803.5684 | 873.9489 | 1253.3294 | 1768.6588 | 1592.2438 | 853.1534 | 1238 | 959.9715 | 828.6195 | 446.346 | 705.4661 | 734.6992 | 621.3638 | 457.9774 | 397.7218 | 1178.352 | 1008.7158 | 586.8633 | 816.0797 | 3142.8573 | 2483.5506 | 1536.9144 | 1817.2782 | 2759.9537 | 2885.7277 | 1377.4887 | 2095.761 | 307.3068 | 193.409 | 183.7729 | 307.3068 | 559.8525 | 589.7158 | 284.8181 | 444.7955 | 506.4661 | 559.3864 | 295.8977 | 375.3638 | 335.415 | 295.3579 | 194.6136 | 305.7384 | 318.0737 | 348.6706 | 237.8466 | 221.5624 | 2388.5732 | 1456.5339 | 893.5006 | 815.7729 | 1158.1249 | 998.4494 | 715.1534 | 612.5517 | 2438.3805 | 1613.4709 | 1116.7955 | 1229.2331
486 | 7.3404 | 2338.5532 | 1978.7375 | 896.9518 | 1213.5339 | 1983.3598 | 1370.1263 | 688.767 | 958.9804 | 779.233 | 553.3495 | 328.6991 | 514.029 | 2211.5245 | 1192.2817 | 672.4464 | 1107.9031 | 1810.6414 | 963.0679 | 648.1942 | 920.8348 | 2058.0871 | 2021.6214 | 1164.7089 | 1054.767 | 2744.2138 | 1981.9612 | 938.7379 | 1460.6799 | 1252.3786 | 1304.2719 | 559.6603 | 799.437 | 1674.1942 | 1719.5732 | 629.437 | 441.8643 | 4032.3986 | 2665.0392 | 1297.9415 | 2314.5053 | 1890.5343 | 1989.8156 | 637.8643 | 860.1455 | 447.5437 | 423.175 | 239.9514 | 629.8348 | 1574.641 | 1042.359 | 846.3205 | 557.2719 | 1634.3594 | 1258.1165 | 657.204 | 620.5339 | 1881.5339 | 1139.0777 | 889.8156 | 750.4563 | 812.0098 | 530.5728 | 442.3205 | 369.7477 | 2724.0004 | 1044.7764 | 846.2428 | 1004.204 | 1572.4853 | 1298.0094 | 1112.8446 | 1020.4272 | 1854.6795 | 1744.4468 | 837.4759 | 1189.5245 | 1764.0777 | 1576.4661 | 866.4563 | 1197.6312 | 940.7379 | 800.204 | 512.1648 | 675.2817 | 700.4759 | 549.9518 | 462.8446 | 411.4951 | 1155.8156 | 1043.563 | 580.1554 | 774.7089 | 3085.1361 | 2544.359 | 1552.2621 | 1727.5732 | 2780.8737 | 2824.6022 | 1434.7085 | 2086.2719 | 304.1165 | 202.6116 | 147.5147 | 304.1165 | 525.4759 | 655.3201 | 259.0098 | 434.4272 | 511.8446 | 535.971 | 279.6603 | 401.029 | 334.1165 | 344.0192 | 203.6116 | 283.8156 | 320.8446 | 339.4272 | 230.3107 | 213.3107 | 2407.9902 | 1495.3688 | 870.1263 | 809.3107 | 1212.7473 | 1010.6893 | 758.2523 | 560.2915 | 2469.6795 | 1670.4755 | 1185.2424 | 1157.9129
487 | 7.3514 | 2363.3894 | 1915.8052 | 941.6992 | 1158.5132 | 2003.3098 | 1326.1681 | 708.469 | 906.4867 | 784.8407 | 552.0265 | 328 | 546.1505 | 2128.2035 | 1188.0796 | 674.9735 | 1171.2213 | 1899.3983 | 982.4956 | 673.3098 | 872.3539 | 1924.7167 | 1955.7787 | 1222.3806 | 985.8849 | 2773.1505 | 2007.3098 | 979.8584 | 1405.4336 | 1202.3805 | 1254.3539 | 554.1327 | 799.9469 | 1662.3186 | 1676.0885 | 692.2478 | 445.8407 | 3961.7787 | 2557.8141 | 1336.0177 | 2275.9292 | 1865.5841 | 1970.5044 | 700.2478 | 821.0885 | 486.9115 | 447.2832 | 242.9469 | 657.2832 | 1493.1504 | 966.1504 | 886.8584 | 557.9735 | 1613.531 | 1212.2477 | 662.8142 | 601.531 | 1835.2743 | 1149.708 | 921.1239 | 749.0531 | 826.5752 | 553.3894 | 463.3894 | 380.6549 | 2644.0707 | 958.4424 | 896.593 | 998.1327 | 1462.053 | 1251.3539 | 1159.7257 | 967.4336 | 1824.8495 | 1677.7522 | 914.8496 | 1217.177 | 1711.4336 | 1565.292 | 812.4867 | 1165.823 | 927.3451 | 826.2567 | 517.9204 | 686.6549 | 654.2212 | 587.885 | 472.7345 | 361.3805 | 1178.3629 | 1001.1947 | 643.2744 | 731.1416 | 3130.6992 | 2472.0442 | 1555.9204 | 1731.7876 | 2803.4159 | 2826.8673 | 1422.3982 | 2020.7787 | 294.2655 | 203 | 155.7345 | 294.2655 | 484.0885 | 630.7433 | 285.2567 | 427.1858 | 504.2124 | 529.1593 | 309.177 | 384.4779 | 297.9823 | 333.3451 | 206.9204 | 318.0443 | 321 | 312.7168 | 218.3186 | 208.1327 | 2252.2742 | 1368.5043 | 884.5752 | 745.7256 | 1075.7963 | 979.8495 | 713.2743 | 580.3894 | 2249.0529 | 1721.6726 | 1076.053 | 1178.3629
488 | 7.3624 | 2325.556 | 1891.3279 | 936.0998 | 1156.0143 | 1961.613 | 1317.1141 | 703.0855 | 874.442 | 741.6273 | 543.1283 | 347.7149 | 495.7414 | 2074.7414 | 1170.2566 | 638.5275 | 1115.8269 | 1758.0815 | 908.0835 | 623.7271 | 861.1426 | 1857.9124 | 1878.0978 | 1191.4705 | 996.8147 | 2693.1691 | 1930.1263 | 1013.5295 | 1375.4134 | 1181.2851 | 1238.2139 | 532.3136 | 735.9267 | 1584.1263 | 1631.6273 | 666.3992 | 432.1996 | 3872.2688 | 2505.7128 | 1321.2281 | 2168.5397 | 1831.4847 | 1953.2424 | 693.1283 | 749.0265 | 432.7984 | 449.9715 | 214.4134 | 593.9267 | 1514.6578 | 991.6008 | 907.7149 | 541.2424 | 1502.5968 | 1248.4582 | 632.442 | 612.8289 | 1808.3707 | 1008.053 | 854.9695 | 737.1711 | 774.7556 | 517.5275 | 438.3707 | 334.6701 | 2558.2118 | 994.444 | 828.998 | 887.5968 | 1413.6558 | 1213.5275 | 1156.0713 | 1002.4725 | 1806.2566 | 1647.4134 | 912.0713 | 1197.2994 | 1697.1853 | 1550.2139 | 832.6864 | 1155.1426 | 894.4705 | 760.9552 | 505.1853 | 639.6843 | 660.8859 | 544.6416 | 476.943 | 370.8432 | 1066.6253 | 942.8269 | 557.2688 | 745.7719 | 3084.6843 | 2383.2546 | 1689.0753 | 1649.1976 | 2704.4399 | 2789.5418 | 1345.112 | 1913.5255 | 317.6578 | 189.1996 | 159.943 | 317.6578 | 512.5723 | 564.9409 | 274.1711 | 397.4277 | 491.1853 | 531.9572 | 289.2994 | 363.2994 | 318.6864 | 343.8432 | 215.8717 | 317.0285 | 309.1711 | 334.6721 | 244.6151 | 243.4868 | 2221.3849 | 1358.0998 | 877.1141 | 729.2139 | 1072.9857 | 891.2688 | 694.2566 | 550.442 | 2285.387 | 1705.2566 | 1128.2016 | 1216.4582
489 | 7.3734 | 2292.0649 | 1902.9764 | 956.9587 | 1104.1022 | 1879.1612 | 1179.2713 | 698.0098 | 798.1494 | 792.8978 | 566.9528 | 341.0138 | 483.0236 | 2094.9587 | 1114.1101 | 675.9253 | 1061.1061 | 1720.0708 | 961.8919 | 700.8467 | 772.1749 | 1858.9961 | 1923.9076 | 1166.0491 | 955.0826 | 2555.2693 | 1902.0531 | 960.1061 | 1393.9627 | 1184.9921 | 1183.1081 | 497.0688 | 727.0157 | 1561.0432 | 1603.055 | 665.002 | 427.0098 | 3715.3066 | 2415.1769 | 1258.1238 | 2141.0511 | 1662.3322 | 1851.2005 | 637.1101 | 783.9292 | 434.9941 | 399.1002 | 272.884 | 629.9273 | 1480.0688 | 950.0826 | 815.1828 | 565.9509 | 1539.9233 | 1193.1101 | 668.9273 | 592.0413 | 1715.1828 | 1074.8644 | 983.7445 | 681.1101 | 792.9627 | 506.0216 | 405.0649 | 377.9135 | 2506.1002 | 948.0924 | 844.9666 | 947.8781 | 1344.1356 | 1211.0039 | 1120.0708 | 970.0649 | 1659.289 | 1580.1317 | 888.0472 | 1138.116 | 1630.1317 | 1526.0472 | 802.0609 | 1141.0275 | 885.0177 | 824.8722 | 536.9371 | 677.9233 | 704.9135 | 544 | 416.1199 | 347.0472 | 1079.9705 | 959.9646 | 580.9509 | 731.0295 | 2919.3243 | 2356.0511 | 1563.2516 | 1573.1474 | 2591.2202 | 2858.8624 | 1393.9017 | 2034.7582 | 283.0688 | 219.9391 | 169.9803 | 283.0688 | 455.114 | 559.0098 | 279.9882 | 464.8663 | 477.0275 | 517.0295 | 328.9214 | 357.0118 | 320.9961 | 246.1926 | 216 | 291.0511 | 322.9725 | 339.9902 | 210.0688 | 215.057 | 2184.0727 | 1396.9233 | 908.9371 | 748.9607 | 1038.0688 | 941.8978 | 741.9056 | 537.0256 | 2305.9607 | 1604.1985 | 1110.0373 | 1165.1022
490 | 7.3844 | 2311.7713 | 1845.5663 | 939.0929 | 1066.4734 | 1770.062 | 1295.143 | 767.9778 | 813.5059 | 769.525 | 504.2581 | 356.062 | 466.1859 | 1998.081 | 1149.6902 | 680.587 | 1033.413 | 1635.0929 | 900.3098 | 697.3908 | 769.7109 | 1742.0413 | 1885.8864 | 1222.6593 | 923.2049 | 2646.2431 | 1854.4837 | 963.6489 | 1344.9793 | 1063.3304 | 1221.0413 | 540.1962 | 721.7728 | 1486.8554 | 1515.9071 | 696.7315 | 407.1033 | 3615.9897 | 2475.7403 | 1317.5973 | 2015.7434 | 1638.3201 | 1783.1033 | 734.7109 | 705.3304 | 478.7315 | 428.031 | 226.382 | 622.3908 | 1473.7832 | 978.5044 | 863.5147 | 531.9793 | 1486.8332 | 1174.9174 | 647.382 | 534.4233 | 1690.5354 | 1087.2685 | 869.2375 | 656.6608 | 740.4941 | 517.0723 | 468.8245 | 353.3511 | 2529.761 | 939.2581 | 842.9587 | 942.3805 | 1350.3805 | 1208.3805 | 1179.4011 | 963.7919 | 1663.8554 | 1601.618 | 886.1342 | 1135.8038 | 1540.618 | 1529.4424 | 793.8054 | 1094.5782 | 930.1843 | 764.9174 | 504.969 | 676.7419 | 675.031 | 568.6593 | 465.2272 | 383.9793 | 1031.1342 | 976.7419 | 653.2669 | 733.0207 | 3051.9659 | 2303.2168 | 1591.8245 | 1602.4853 | 2552.9587 | 2803.5989 | 1381.8554 | 1944.6489 | 280.1652 | 201.9277 | 178.8141 | 280.1652 | 454.2168 | 535.2375 | 275.0723 | 437.8038 | 453.2168 | 471.1342 | 295.1239 | 342.1962 | 310.8141 | 298.969 | 202.8761 | 258.8967 | 324.7625 | 280.062 | 245.6076 | 206.9793 | 2163.5854 | 1369.4631 | 861.9793 | 700.6386 | 1039.6696 | 992.8038 | 712.2685 | 567 | 2191.2272 | 1613.4233 | 1177.5354 | 1196.0294
491 | 7.3954 | 2193.2537 | 1804.3166 | 948.4749 | 1015.3441 | 1775.118 | 1212.706 | 670.7916 | 860.3264 | 724.1357 | 528.118 | 362.588 | 483.5929 | 1909.0049 | 1120.7011 | 640.8643 | 1071.643 | 1643.6608 | 929.4346 | 639.3619 | 743.2891 | 1743.1729 | 1875.0452 | 1190.4523 | 847.769 | 2383.2891 | 1803.3716 | 930.769 | 1341.2133 | 1094.9774 | 1222.2408 | 557.9144 | 699.7738 | 1471.7787 | 1507.5654 | 670.6381 | 416.0049 | 3612.2408 | 2357.5557 | 1280.1082 | 2084.0599 | 1666.8288 | 1791.8013 | 706.7965 | 735.8013 | 452.5477 | 429.9823 | 262.4572 | 565.2665 | 1450.9419 | 932.4248 | 816.4572 | 529.5024 | 1376.2439 | 1165.055 | 683.3667 | 574.118 | 1644.8692 | 1111.8466 | 892.4297 | 718.1228 | 691.7512 | 523.2989 | 451.412 | 386.5703 | 2410.7738 | 963.5929 | 839.6333 | 881.937 | 1291.8142 | 1148.9095 | 1122.4297 | 849.1082 | 1649.7512 | 1565.2714 | 899.2262 | 1116.1406 | 1503.7286 | 1474.4169 | 869.937 | 1148.4847 | 851.9774 | 756.8417 | 501.8869 | 650.8918 | 676.9823 | 536.3619 | 486.3215 | 380.7787 | 1042.4395 | 950.2812 | 583.0678 | 733.8918 | 2855.5299 | 2321.8514 | 1575.1809 | 1645.7885 | 2549.4749 | 2859.5299 | 1367.5703 | 1911.5428 | 295.5703 | 195.0452 | 160.6155 | 295.5703 | 474.2763 | 557.231 | 282.0226 | 419.0905 | 473.4346 | 483.6834 | 307.1583 | 360.1858 | 292.7512 | 295.4798 | 191.2488 | 249.882 | 302 | 286.3392 | 208.0678 | 204.5703 | 2026.2812 | 1319.1308 | 861.3344 | 667.4475 | 1008.8142 | 973.4346 | 738.9273 | 565.2133 | 2212.7738 | 1652.0776 | 1133.3392 | 1103.7463
492 | 7.4064 | 2162.6057 | 1818.5934 | 966.4867 | 987.3491 | 1736.7906 | 1199.3368 | 704.6858 | 785.9754 | 728.5339 | 491.885 | 386.9179 | 465.0924 | 1864.499 | 1149.4292 | 634.8953 | 1010.4168 | 1624.4271 | 903.8131 | 653.9179 | 795.2998 | 1622.5708 | 1875.3943 | 1209.9302 | 880.3019 | 2432.6817 | 1735.4086 | 964.9774 | 1263.1273 | 1089.8809 | 1103.0698 | 511.768 | 689.5811 | 1419.8378 | 1529.731 | 655.8378 | 373.384 | 3489.2649 | 2465.9405 | 1326.2895 | 1957 | 1573.1047 | 1737.8706 | 698 | 684.1047 | 433.1745 | 387.7105 | 239.3142 | 617.5565 | 1361.9302 | 995.3614 | 793.4189 | 509.1869 | 1427.2423 | 1084.4209 | 656 | 535.3717 | 1594.9179 | 1063.5708 | 910.0575 | 636.5462 | 681.9774 | 495.6181 | 426.7331 | 367.3019 | 2398.347 | 945.7906 | 866.9302 | 836.5585 | 1324.1376 | 1145.8029 | 1139.7084 | 897.4189 | 1637.2546 | 1574.4518 | 908.3019 | 1080.9897 | 1492.1745 | 1405.5462 | 823.3019 | 1081.0103 | 850.0575 | 750.3142 | 498.0472 | 603.384 | 633.3142 | 552.8029 | 457.9425 | 329.0472 | 974.3019 | 877.3265 | 586.0698 | 684.9528 | 2793.6057 | 2230.9897 | 1582.3717 | 1504.7659 | 2564.5893 | 2792.4743 | 1349.8727 | 1931.9281 | 277.3491 | 196.5113 | 143.8953 | 277.3491 | 444.4661 | 524.9425 | 281.7782 | 423.2793 | 448.7207 | 470.0698 | 313.4415 | 326.4312 | 282.1047 | 272.8604 | 200.9877 | 285.8357 | 301.9302 | 300.7207 | 212.7454 | 186.6981 | 1952.768 | 1367.1971 | 919.1396 | 727.6612 | 1041.7659 | 948.3717 | 777.6037 | 491.386 | 2200.6612 | 1569.9425 | 1145.9405 | 1134.499
493 | 7.4174 | 2207.8496 | 1835.7552 | 900.0679 | 995.4424 | 1722.2211 | 1178.3569 | 696.3304 | 731.7463 | 687.8584 | 542.4896 | 356.3864 | 440.6608 | 1819.1769 | 1103.9233 | 630.9144 | 935.4601 | 1499.0384 | 868.5929 | 625.0856 | 696.2095 | 1661.8937 | 1831.8024 | 1206.2537 | 859.4896 | 2257.056 | 1797.3745 | 990.5368 | 1235.1297 | 977.2272 | 1102.7345 | 532.0088 | 678.3304 | 1442.292 | 1462.7993 | 676.6873 | 384.472 | 3378.9202 | 2408.1976 | 1280.0384 | 1953.9321 | 1573.5457 | 1675.5545 | 700.4543 | 687.1416 | 436.0649 | 431.6017 | 247.5752 | 551.416 | 1362.4071 | 1032.056 | 803.3481 | 538.7433 | 1350.6224 | 1174.5161 | 657.1328 | 518.236 | 1564.8112 | 1106.5752 | 885.2272 | 624.5192 | 710.2272 | 565.4985 | 446.3864 | 348.5192 | 2324.528 | 943.5871 | 865.9912 | 850.6784 | 1243.6312 | 1166.584 | 1104.3776 | 904.8937 | 1584.9881 | 1502.7728 | 890.7935 | 1134.0649 | 1495.7256 | 1394.6991 | 806.1239 | 1026.9912 | 822.1121 | 756.7345 | 524.3569 | 612.5368 | 642 | 577.2655 | 483.3392 | 358.7552 | 954.1504 | 915.2272 | 586.7256 | 655.4808 | 2825.7284 | 2282.3657 | 1564.8584 | 1446.3039 | 2414.764 | 2695.0472 | 1373.528 | 1845.8672 | 287.3776 | 183.3304 | 142.5104 | 287.3776 | 483.2537 | 551.3776 | 249.5104 | 428.6873 | 441.0472 | 473.4632 | 299.2272 | 363.8289 | 284.764 | 269.1416 | 173.708 | 227.6519 | 301.0856 | 294.6048 | 259.1121 | 205.764 | 1967.8642 | 1343.3481 | 920.0265 | 662.4808 | 980.8496 | 930.1032 | 738.0088 | 582.5929 | 2130.2418 | 1598.7375 | 1091.4071 | 1088.3273
494 | 7.4284 | 2216.4335 | 1846.0029 | 959.6799 | 964.5354 | 1547.4986 | 1128.0935 | 704.8442 | 765.6317 | 686.8442 | 512.3569 | 344.8329 | 451.9179 | 1742 | 1066.2238 | 611.2011 | 865.2493 | 1519.4051 | 817.6204 | 648.9037 | 756.8357 | 1573.1105 | 1842.1105 | 1168.5609 | 825.8471 | 2230.8018 | 1699.2408 | 962.7847 | 1158.7847 | 1000.9887 | 1057.5949 | 513.9037 | 681.5496 | 1421.9916 | 1374.6091 | 632.915 | 392.5722 | 3229.5156 | 2395.3909 | 1301.9518 | 1887.4193 | 1520.1332 | 1607.0113 | 745.3598 | 688.5128 | 398.5467 | 364.0708 | 239.0595 | 577.323 | 1404.8131 | 1010.4419 | 775.592 | 476.0935 | 1343.238 | 1084.0595 | 680.1785 | 523.5354 | 1559.5722 | 1094.3116 | 911.9037 | 638.7478 | 731.6091 | 517.796 | 435.2493 | 357.9887 | 2309.3485 | 1013.6204 | 884.7507 | 823.5836 | 1217.4278 | 1140.7365 | 1110.9292 | 823.7252 | 1503.2861 | 1481.5694 | 923.966 | 1088.3967 | 1468.8697 | 1428.1671 | 848.9405 | 1047.5949 | 787.0935 | 715.8187 | 472.7847 | 582.4051 | 639.0255 | 615.4448 | 472.2011 | 372.3938 | 939.4391 | 937.085 | 561.119 | 641.0142 | 2579.3088 | 2213.6997 | 1560.8102 | 1516.9858 | 2405.5978 | 2692.1927 | 1364.4759 | 1820.7394 | 293.7507 | 191.2493 | 172.4051 | 293.7507 | 447.2861 | 558.2861 | 280.6544 | 383.3088 | 440.2153 | 500.5014 | 323.2861 | 319.4646 | 280.6544 | 273.9632 | 187.7026 | 259.8697 | 320.323 | 325.2153 | 222.4872 | 201.8329 | 1842.7507 | 1307.6799 | 863.4504 | 653.3088 | 995.7394 | 912.4164 | 719.1303 | 533.9405 | 2016.5581 | 1649.9547 | 1137.323 | 998.8781
495 | 7.4394 | 2082.2921 | 1763.411 | 923.0718 | 970.8466 | 1590.3737 | 1175.4208 | 732.5644 | 745.0767 | 710.2575 | 516.9233 | 358.7896 | 421.3688 | 1733.7452 | 1128.1907 | 642.1238 | 885.4307 | 1502.5348 | 784.9529 | 677.6362 | 691 | 1482.6559 | 1749.5025 | 1134.9529 | 774.0619 | 2118.8663 | 1632.9008 | 986.6932 | 1186.7847 | 983.6362 | 1062.6932 | 542.4926 | 625.2055 | 1324.9901 | 1322.5792 | 678.4258 | 368.9233 | 3125.8515 | 2443.6165 | 1308.4581 | 1809.9754 | 1455.1238 | 1618.2625 | 675.0866 | 662.9184 | 422.4258 | 379.6362 | 241.2921 | 530.9803 | 1304.703 | 965.6214 | 864.1436 | 521.6263 | 1346.2822 | 1086.0049 | 680.3441 | 530.0619 | 1532.2625 | 1038.5841 | 941.4258 | 691.6115 | 676.9951 | 558.0521 | 453.9951 | 340.0619 | 2213.3244 | 974.9233 | 863.9951 | 812.9803 | 1235.7649 | 1178.1858 | 1098.2351 | 847.6263 | 1483.171 | 1534.9134 | 875.5693 | 985.8663 | 1448.3441 | 1409.4011 | 850.3688 | 1051.3293 | 832.2674 | 786.114 | 501.646 | 574.8663 | 592.646 | 503.6087 | 502.1907 | 347.5025 | 973.3343 | 892.4356 | 567.5222 | 581.9479 | 2598.6362 | 2281.8986 | 1544.5123 | 1570.3814 | 2340.3068 | 2629.9184 | 1374.7277 | 1779.1858 | 274.5742 | 207.4827 | 163.5025 | 274.5742 | 435.6411 | 544.2006 | 291.5742 | 399.5496 | 414.5841 | 459.7797 | 311.0669 | 311.1436 | 289.2773 | 299.4729 | 185.0099 | 244.2252 | 279.1485 | 295.8515 | 243.7748 | 213.349 | 1823.5742 | 1265.6882 | 914.6981 | 672.1337 | 956.0521 | 922.9479 | 734.4877 | 533.7178 | 2054.2006 | 1587.1485 | 1098.3737 | 1071.6932
496 | 7.4504 | 2097.6549 | 1718.8496 | 927.531 | 939.3849 | 1622.1459 | 1151.7212 | 713.8319 | 758.8319 | 635.0752 | 502.2434 | 369.0841 | 449.7301 | 1625.0664 | 1069.4469 | 594.0177 | 890.0752 | 1401.8805 | 848.4115 | 659.2434 | 687.3008 | 1508.6769 | 1752.0486 | 1199.2522 | 755.0532 | 2136.7787 | 1738.668 | 938.3142 | 1178.8716 | 961.0398 | 1007.2522 | 526.1947 | 604.6903 | 1290.845 | 1342.0752 | 668.8141 | 357.1859 | 3104.022 | 2358.6149 | 1202.9602 | 1740.1238 | 1402.0575 | 1555.5708 | 715.4204 | 634.9911 | 415.2522 | 360.6549 | 256.7389 | 482.0266 | 1278.2788 | 909.9248 | 869.4469 | 480.2699 | 1329.5176 | 1096.1416 | 647.9336 | 500.9646 | 1464.1061 | 1070.9071 | 931.2256 | 605.9734 | 663.4956 | 507.3363 | 440.1593 | 318.6992 | 2119.0132 | 959.0664 | 851.0841 | 762.7655 | 1178.4778 | 1106.6992 | 1148.5442 | 805.177 | 1366.3938 | 1486.761 | 873.531 | 1001.5928 | 1400.8007 | 1338.0089 | 879.0664 | 980.6062 | 785.0929 | 710.8407 | 506.1327 | 597.0664 | 603.0177 | 597.3761 | 440.9248 | 355.0089 | 909.2168 | 909.5708 | 626.5885 | 632.7478 | 2566.2875 | 2201.3186 | 1575.0664 | 1445.8319 | 2385.407 | 2598.6283 | 1409.2434 | 1703.5796 | 280.823 | 165.0266 | 169.5796 | 280.823 | 428.5619 | 508 | 297.9911 | 344.4381 | 445.0575 | 445.0929 | 296.7301 | 315.1859 | 266.0752 | 231.1107 | 209.9911 | 272.9159 | 294.5885 | 281.5885 | 215.6726 | 192.5044 | 1823.0132 | 1199.4381 | 871.4115 | 647.3274 | 909.4558 | 967.3583 | 706.0354 | 541.7389 | 2019.7655 | 1595.2743 | 1144.8584 | 1015.4115
497 | 7.4614 | 2080.2055 | 1738.3255 | 958.6764 | 903.5654 | 1489.6273 | 1120.4327 | 712.2891 | 753.1455 | 621.82 | 492.5546 | 368.0727 | 425.1927 | 1599.0855 | 1104.2055 | 618.3127 | 877.9764 | 1363.3599 | 846.6254 | 650.9036 | 648.2527 | 1417.0364 | 1710.5418 | 1199.6873 | 817.7237 | 2074.9745 | 1637.2891 | 1009.231 | 1126.7581 | 905.0727 | 1003.9291 | 556.5673 | 630.7237 | 1208.9381 | 1330.7473 | 649.3363 | 377.7237 | 3018.6528 | 2191.2891 | 1260.8545 | 1688.1564 | 1371.1564 | 1529.3037 | 716.0491 | 624 | 418.0964 | 347.8309 | 242.2291 | 516.2527 | 1312.94 | 918.94 | 895.7256 | 504.6145 | 1244.2419 | 1066.9273 | 665.82 | 457.1927 | 1389.9636 | 1054.5673 | 902.06 | 579.0709 | 658.4818 | 511.3855 | 435.3363 | 355.9164 | 2031.1327 | 939.2527 | 855.18 | 781.4218 | 1144.8328 | 1140.2546 | 1095.8564 | 815.2055 | 1350.5272 | 1394.3018 | 904.4837 | 915.7818 | 1265.4799 | 1339.2674 | 856.5546 | 1003.1219 | 791.3146 | 720.3855 | 469.7818 | 574.2655 | 623.3618 | 552.3963 | 453.4091 | 363.6873 | 879.2744 | 894.8673 | 628.8564 | 629.0364 | 2414.2872 | 2183.6745 | 1557.6618 | 1445.5418 | 2278.1908 | 2539.8219 | 1402.1564 | 1696.9764 | 267.6145 | 197.6145 | 164.5182 | 267.6145 | 401.8073 | 504.6273 | 287.6745 | 368.7837 | 412.5182 | 456.1327 | 273.8309 | 334.3746 | 255.2291 | 264.2291 | 199.7709 | 267.9528 | 300.3255 | 287.7109 | 223.6745 | 198.4818 | 1731.0855 | 1229.9509 | 866.3491 | 639.4218 | 954.6273 | 898.18 | 747.3982 | 525.0128 | 2037.4582 | 1520.0727 | 1169.8328 | 1011.3127
498 | 7.4724 | 2044.1125 | 1723.7423 | 893.151 | 926.026 | 1450.3942 | 1158.5231 | 716.3038 | 717.2173 | 594.6077 | 482.8029 | 349.7827 | 427.2173 | 1557.9981 | 1071.0462 | 640.2173 | 854.3279 | 1420.3923 | 877.5048 | 644.4587 | 659.0221 | 1408.8692 | 1718.4808 | 1177.7827 | 779.3942 | 2114.2673 | 1630.3298 | 929.6077 | 1156.1327 | 885.6962 | 947.7606 | 515.5433 | 601.0865 | 1266.3942 | 1311.7202 | 693.9115 | 344.9356 | 2918.0702 | 2197.3683 | 1292.9798 | 1689.9817 | 1377.0019 | 1425.6317 | 682.2596 | 623.674 | 417.8048 | 362.8269 | 238.8029 | 529.1952 | 1307.8894 | 916.2798 | 791.4125 | 497.4144 | 1208.674 | 1087.6077 | 639.6942 | 458.2394 | 1395.0885 | 1019.3019 | 904.5029 | 631.7625 | 662.674 | 514.674 | 470.2413 | 330.6942 | 2018.2192 | 951.326 | 889.0202 | 774.8048 | 1095.475 | 1072.1106 | 1063.5192 | 785.4144 | 1452.7885 | 1367.5875 | 837.2596 | 969.1308 | 1338.2635 | 1258.7567 | 843.5433 | 950.7163 | 738.7606 | 724.4346 | 523.5654 | 560.2817 | 607.5654 | 594.525 | 475.3481 | 346.7808 | 976.199 | 897.7846 | 588.2615 | 616.0885 | 2484.8288 | 2193.3683 | 1581.5692 | 1453.1548 | 2338.8106 | 2443.5837 | 1412.8913 | 1671.4808 | 259.8712 | 192.1529 | 159.674 | 259.8712 | 397.7183 | 474.9558 | 294.1087 | 350.2817 | 406.4788 | 442.4587 | 293.0644 | 295.0202 | 249.5212 | 259.3904 | 203.3058 | 230.1087 | 291.2837 | 287.3904 | 230.8692 | 200.6096 | 1678.2635 | 1256.7865 | 874.7625 | 630.3058 | 916.9154 | 925.3923 | 727.7808 | 496.8913 | 2014.6337 | 1508.9538 | 1117.324 | 1024.7423
499 | 7.4834 | 1967.475 | 1649.7109 | 920.8142 | 831.1622 | 1391.8702 | 1126.3717 | 728.2094 | 714.1622 | 617.4513 | 512.9115 | 345.9439 | 430.3304 | 1593.7994 | 1016.6608 | 641.2743 | 815.5162 | 1383.2772 | 776.4513 | 607.3452 | 653.0561 | 1409.2271 | 1666.3304 | 1180.1121 | 719.177 | 1965.2683 | 1550.8791 | 956.8054 | 1103.7669 | 867.2595 | 955.9351 | 526.7729 | 611.2094 | 1197.348 | 1248.2507 | 717.4129 | 330.177 | 2815.5074 | 2219.6784 | 1265.0797 | 1613.9439 | 1323.6636 | 1404.2891 | 712.3865 | 615.3156 | 408.1209 | 345.6047 | 264.1946 | 535.4926 | 1335.1798 | 971.7669 | 820.2951 | 465.5811 | 1204.9115 | 1104.4041 | 668.9115 | 453.3304 | 1358.0413 | 1068.2183 | 849.531 | 596.8791 | 649.8406 | 512.2419 | 422.5162 | 355.9203 | 1960.4954 | 953.3953 | 914.0176 | 762.3392 | 1161.4926 | 1038.2891 | 1114.0973 | 760.7168 | 1334.7729 | 1354.3068 | 865.6932 | 954.9852 | 1283.4838 | 1346.8791 | 849.5723 | 965.0176 | 748.5074 | 722.6784 | 520.6047 | 571.5014 | 603.1534 | 525.8466 | 471.5723 | 378.7433 | 900.3717 | 855.7257 | 571.823 | 588.354 | 2415.5987 | 2226.6844 | 1497.3304 | 1383.2123 | 2198.3805 | 2502.6312 | 1408.6135 | 1620.5398 | 230.475 | 178.4513 | 159.0561 | 230.475 | 384.3216 | 490.0324 | 298.8702 | 364.646 | 394.8466 | 404.3776 | 308.7906 | 326.3068 | 262.4838 | 277.7906 | 174.9351 | 235.354 | 268.0324 | 302.525 | 238.1209 | 180.6372 | 1611.8731 | 1171.9115 | 841.9351 | 594.9203 | 872.4425 | 902.4277 | 757.2035 | 504.1033 | 1949.5074 | 1570.8378 | 1158.2183 | 960.2655
500 | 7.4944 | 1993.205 | 1605.9985 | 932.4012 | 797.5988 | 1408.6749 | 1091.6061 | 739.1318 | 677.0015 | 626.0015 | 496.2064 | 329.6676 | 430.0073 | 1545.7408 | 1030.2709 | 634.6676 | 813.7321 | 1270.3324 | 778.2064 | 670.1318 | 669.8653 | 1354.6706 | 1686.2592 | 1214.7965 | 735.2665 | 1878.1347 | 1548.0015 | 992.0717 | 1078.7291 | 817.795 | 964.937 | 567.4729 | 612.8053 | 1153.53 | 1220.3338 | 654.8668 | 343.4685 | 2808.3572 | 2171.8067 | 1294.8053 | 1589.2738 | 1232.0015 | 1435.0776 | 729.8053 | 572.2679 | 399.8067 | 346.1362 | 227.7291 | 509.7335 | 1238.8023 | 947.795 | 936.8067 | 454.1289 | 1194.202 | 1061.407 | 661.6662 | 463.1332 | 1312.6735 | 1044.2694 | 899.798 | 592.6691 | 581.9956 | 506.4056 | 418.0673 | 315.5359 | 1834.2606 | 946.3353 | 857.3324 | 749.6618 | 1126.0073 | 1086.9312 | 1100.0615 | 795.9327 | 1368.8111 | 1334.9385 | 879.5988 | 905.3324 | 1276.7408 | 1332.0102 | 860.0673 | 903.2694 | 764.5403 | 685.5988 | 520.2035 | 518.5988 | 592.2035 | 536.2635 | 482.8668 | 332.1289 | 872.8009 | 858.6018 | 558.3982 | 612.4056 | 2341.3324 | 2251.2123 | 1512.1962 | 1290.8609 | 2154.1918 | 2426.795 | 1384.2665 | 1635.1362 | 265.3982 | 190.5988 | 176.9312 | 265.3982 | 413.2738 | 484.8653 | 322.7365 | 338.4627 | 398.6018 | 454.5359 | 294.6706 | 317.0644 | 266.6662 | 269 | 191.1318 | 259.5388 | 265.6633 | 270.6662 | 242.5974 | 181.1977 | 1538.8609 | 1172.6603 | 860.3968 | 562.7965 | 906.2709 | 887.2694 | 693.4612 | 557.6706 | 1946.3543 | 1595.3411 | 1140.9312 | 984.5329
501 | 7.5054 | 1933.1376 | 1604.7384 | 929.1308 | 804.3618 | 1333.6903 | 1036.2724 | 746.3058 | 655.1229 | 603.5407 | 440.1642 | 327.6942 | 367.474 | 1474.3431 | 998.53 | 634.1495 | 817.1415 | 1275.7984 | 718.8171 | 681.235 | 676.9066 | 1308.5034 | 1733.3539 | 1228.8278 | 738.1642 | 1853.1936 | 1524.3765 | 957.4779 | 1115.4179 | 857.3431 | 924.8692 | 516.2203 | 565.2724 | 1179.3618 | 1212.2163 | 673.064 | 328.7463 | 2604.295 | 2119.7129 | 1256.8771 | 1525.7089 | 1220.1642 | 1333.3431 | 676.4513 | 558.6716 | 344.9479 | 327.1121 | 261.5074 | 502.1789 | 1222.3658 | 983.8505 | 871.9066 | 493.4179 | 1163.3392 | 1000.8919 | 670.3658 | 462.1081 | 1256.3658 | 1019.7089 | 914.7984 | 567.7984 | 620.5634 | 461.5113 | 413.4553 | 300.5674 | 1866.462 | 929.7276 | 866.1455 | 789.5074 | 1062.1642 | 1097.1229 | 1141.3431 | 791.231 | 1265.7797 | 1297.004 | 888.8358 | 918.0748 | 1212.0374 | 1244.0561 | 851.0895 | 879.7276 | 702.4179 | 685.5221 | 495.2577 | 534.53 | 559.1081 | 565.4553 | 484.3284 | 372.8732 | 859.9066 | 841.5447 | 574 | 560.9626 | 2351.4553 | 2151.8958 | 1544.6903 | 1337.4887 | 2213.466 | 2466.7797 | 1380.1789 | 1601.705 | 275.761 | 199.9813 | 194.5447 | 275.761 | 351.8919 | 495.5447 | 293.2537 | 367.8879 | 384.7463 | 431.4179 | 260.4366 | 337 | 272.7463 | 267.3992 | 203.3992 | 213.9105 | 293.6716 | 272.8171 | 260.761 | 200.1081 | 1576.0294 | 1215.466 | 893.3658 | 590.0708 | 874.0934 | 860.8171 | 739.3805 | 524.0187 | 1781.2469 | 1531.3845 | 1155.6342 | 980.7237
502 | 7.5164 | 1841.1047 | 1538.6756 | 971.7891 | 789.6284 | 1305.168 | 1032.6313 | 711.4189 | 610.7331 | 561.2567 | 447.1578 | 365.1062 | 383.4204 | 1439.4247 | 1030.2109 | 675.3716 | 772.2625 | 1243.5855 | 720.2036 | 666.5782 | 643.0545 | 1241.4735 | 1612.9498 | 1142.8349 | 740.2655 | 1786.5222 | 1467.208 | 1013.792 | 1105.056 | 814.6917 | 881.4218 | 540.5265 | 562.1047 | 1148.2153 | 1177.848 | 763.6946 | 331.7404 | 2634.2197 | 2138.1709 | 1299.8466 | 1524.7891 | 1213.112 | 1300.0029 | 649.0473 | 554.1534 | 375.5782 | 357.4764 | 242.1047 | 468.3156 | 1241.5295 | 932.0545 | 842.7891 | 486.6858 | 1093.7345 | 1010.4218 | 692.0545 | 448.4749 | 1269.6902 | 1021.472 | 905.1534 | 561.3098 | 616.7375 | 496.0487 | 430.8378 | 341.3702 | 1739.6844 | 934.9469 | 871.7345 | 757.3215 | 1059.2153 | 1044.6844 | 1124.8894 | 725.4749 | 1239.1062 | 1340.3702 | 880.264 | 935.056 | 1223.3687 | 1250.9527 | 832.4218 | 878.0029 | 699.7891 | 633.7847 | 525.056 | 567.3156 | 549.5782 | 567.5295 | 494.8436 | 374.0545 | 824.9498 | 836.7876 | 578.2611 | 549.3156 | 2337.432 | 2196.7433 | 1548.9942 | 1307.0545 | 2116.1666 | 2421.0102 | 1347.1047 | 1551.5796 | 252.7331 | 192.5796 | 183.0015 | 252.7331 | 362.2109 | 478.6356 | 285.2095 | 337.6814 | 390.1076 | 425.0015 | 269.3127 | 334.1593 | 275.7404 | 254.6858 | 196.8407 | 233.6829 | 278.4218 | 260 | 229.5265 | 186.6327 | 1498.4189 | 1136.1578 | 904.4807 | 570.6298 | 898.3186 | 856.5222 | 734.2582 | 523.6873 | 1897.1622 | 1553.2153 | 1124.0029 | 946.68
503 | 7.5274 | 1836.6233 | 1617.5201 | 968.6853 | 808.4994 | 1199.6136 | 1043.9951 | 723.0924 | 654.7217 | 623.3244 | 447.3196 | 363.6756 | 392.531 | 1410.2272 | 1010.1799 | 651.3973 | 776.1081 | 1180.8457 | 789.1956 | 669.7266 | 623.3609 | 1239.2369 | 1593.4848 | 1218.5978 | 705.644 | 1824.8554 | 1499.9332 | 995.6962 | 1062.3864 | 740.6549 | 869.5978 | 534.2576 | 564.7679 | 1086.644 | 1116.2321 | 643.8761 | 298.9538 | 2521.2418 | 2002.3147 | 1251.8505 | 1513.8603 | 1140.1908 | 1253.3706 | 705.7886 | 595.8554 | 379.4071 | 332.7631 | 246.1288 | 463.1956 | 1210.2321 | 903.8554 | 839.5152 | 469.5103 | 1106.9223 | 998.2576 | 670.1495 | 435.3402 | 1199.4071 | 1029.7935 | 944.4739 | 617.8141 | 612.8919 | 538.1495 | 471.0255 | 322.1701 | 1737.6598 | 936.938 | 888.4945 | 694.6598 | 1010.4022 | 1048.1033 | 1173.3038 | 720.6756 | 1227.8092 | 1312.1543 | 874.9125 | 899.3609 | 1208.0924 | 1187.7473 | 830.4641 | 847.0207 | 707.5723 | 688.0413 | 490.7217 | 562.0255 | 556.469 | 533.6804 | 474.938 | 348.2369 | 804.9745 | 844.9223 | 604.4022 | 558.6549 | 2210.5359 | 2115.7935 | 1609.7679 | 1289.5723 | 2011.1337 | 2313.4228 | 1365.7582 | 1536.8554 | 291.6185 | 178.8761 | 164.7886 | 291.6185 | 360.7217 | 431.1081 | 295.469 | 367.4897 | 361.4484 | 405.5978 | 301.1288 | 312.3657 | 235.6391 | 241.0826 | 209.6185 | 243.0668 | 273.2321 | 256.4277 | 225.6391 | 179.1033 | 1511.113 | 1147.887 | 838.3973 | 584.7266 | 860.6853 | 893.4945 | 778.6233 | 492.2114 | 1822.2722 | 1488.0717 | 1099.825 | 989.7679
504 | 7.5384 | 1782.1696 | 1539.8112 | 934.4592 | 732.1952 | 1224.0816 | 1080.8112 | 667.3712 | 657.056 | 583.5536 | 444.1824 | 394.5088 | 403.2832 | 1446.5776 | 943.5792 | 692.0432 | 759.0816 | 1208.1312 | 749.9808 | 643.2704 | 617.232 | 1211.6224 | 1628.7232 | 1201.088 | 654.7104 | 1778.144 | 1516.8864 | 1024.3952 | 1022.0496 | 777.4528 | 860.296 | 516.7296 | 549.5472 | 1040.2704 | 1124.1568 | 626.4032 | 322.528 | 2435.8352 | 2025.0496 | 1248.1568 | 1437.4464 | 1143.8112 | 1194.9504 | 702.4784 | 555.1632 | 361.7424 | 346.6976 | 241.3776 | 432.9184 | 1209.2192 | 876.1136 | 802.4528 | 472.1824 | 1040.8864 | 980.7296 | 671.2896 | 435.9056 | 1171.528 | 990.2144 | 900.3712 | 569.1696 | 615.8736 | 525.0624 | 437.912 | 315.1568 | 1716.4832 | 919.6288 | 859.9504 | 690.4464 | 1014.4144 | 1051.112 | 1119.4336 | 753.3328 | 1211.7856 | 1262.7984 | 889.44 | 900.1632 | 1144.3456 | 1177.944 | 861.6352 | 849.6288 | 716.4272 | 698.3648 | 504 | 520.1568 | 548.2704 | 540.648 | 464.9248 | 330.2896 | 838.088 | 786 | 611.6976 | 593.8112 | 2170.9744 | 2089.8944 | 1606.9504 | 1296.4208 | 1969.0192 | 2329.3456 | 1418.1504 | 1502 | 270.5152 | 153.9248 | 157.1008 | 270.5152 | 373.6352 | 419.9184 | 290.3712 | 352.5968 | 353.7424 | 381.5216 | 289.8112 | 287.1136 | 226.6288 | 261.3648 | 192.3456 | 222.8304 | 279.1504 | 244.1072 | 220.2768 | 202.5408 | 1456.5408 | 1194.2512 | 884.6976 | 566.4784 | 818.2512 | 854.9184 | 732.0128 | 503.1824 | 1688.2768 | 1422.9936 | 1136.176 | 969.9872
505 | 7.5494 | 1752.7345 | 1543.239 | 937.8849 | 753.1416 | 1203.2124 | 1058.9291 | 686.3717 | 608 | 585.2654 | 446.9469 | 347.2832 | 418.1947 | 1359.9115 | 974.2566 | 625.3894 | 737.2389 | 1137.1947 | 724.7522 | 647.9823 | 577.3275 | 1178.1769 | 1617.5486 | 1196.4159 | 633.9204 | 1726.1681 | 1488.3098 | 951.708 | 967.3098 | 781.4248 | 878.0088 | 533.4337 | 540.5132 | 1032.7168 | 1080.0442 | 681.6106 | 296.2478 | 2353.6107 | 1965.2478 | 1229.9381 | 1420.1593 | 1133.4159 | 1245.6017 | 730.354 | 533.6726 | 377.1593 | 309.4779 | 257.6902 | 448.7257 | 1142.5664 | 911.5044 | 804.6726 | 463.6548 | 1011.7257 | 979.3274 | 676.3096 | 434.8938 | 1152.646 | 1042.0354 | 908.2743 | 546.5486 | 576.5133 | 487.3363 | 441.6106 | 294.0178 | 1626.6637 | 905.8849 | 916.6372 | 680.5575 | 958.1947 | 955.9204 | 1093.8938 | 705.1327 | 1178.2655 | 1228.3805 | 869.7788 | 869.9115 | 1118.0265 | 1222.9115 | 844.8141 | 842.3098 | 677.8053 | 692.407 | 510.2301 | 491.0354 | 550.4514 | 555.6372 | 484.6637 | 348.1593 | 826.407 | 783.2743 | 586.7434 | 575.6283 | 2137.1416 | 2191.354 | 1657.0177 | 1243.938 | 1988.9911 | 2293.0973 | 1373.0265 | 1501.2212 | 240.4602 | 173.0796 | 174.1151 | 240.4602 | 377.2567 | 433.5841 | 304.5044 | 309.4602 | 372.8584 | 359.0974 | 286.6195 | 317.6372 | 221.2567 | 258.9115 | 171.6726 | 243.6991 | 237.3363 | 262.9115 | 219.1858 | 195.2212 | 1448.2478 | 1163.2212 | 846.6991 | 601.7522 | 790.3363 | 868.5841 | 748.6372 | 499.1327 | 1705.8762 | 1414.7788 | 1123.1327 | 948.1062
506 | 7.5604 | 1690.409 | 1577.3018 | 930.5885 | 749.5811 | 1174.1047 | 1012.5187 | 712.5137 | 609.2419 | 550.7955 | 445.0649 | 363.3441 | 418.793 | 1343 | 903.5511 | 645.2144 | 706.5885 | 1122.0624 | 733.2419 | 675.2045 | 652.1047 | 1124.3117 | 1594.586 | 1205.177 | 655.1072 | 1727.6883 | 1480.4738 | 981.9252 | 975.6185 | 784.818 | 825.1072 | 580.7207 | 492.586 | 1000.3043 | 1041.7581 | 670.8928 | 286.9302 | 2405.9502 | 1943.3043 | 1295.4464 | 1381.9252 | 1131.414 | 1213.5511 | 702.4464 | 548.6534 | 337.0624 | 329.6559 | 244.414 | 460 | 1162.8903 | 933.6583 | 821.7232 | 435.7631 | 1018.6534 | 980.9052 | 607.0723 | 449.9651 | 1174.7581 | 1051.4763 | 894.5162 | 528.1396 | 613.7905 | 503.6234 | 439.1721 | 348.2743 | 1604.995 | 900.4514 | 882.6434 | 684.9302 | 962.9327 | 971.1047 | 1108.3441 | 697.6259 | 1218.8579 | 1197.4813 | 882.3142 | 849.2743 | 1114.414 | 1202.6883 | 806.2169 | 851.4464 | 674.9302 | 663.6908 | 528.1347 | 509.5836 | 571.1372 | 548.5511 | 476.6559 | 325.2419 | 790.2394 | 776.1396 | 653.9302 | 537.4813 | 2135.4439 | 2155.1671 | 1617.7905 | 1243.2843 | 1970.1421 | 2241.9975 | 1367.5511 | 1500.2768 | 281.2369 | 166.5885 | 179.5486 | 281.2369 | 401.6908 | 429.9975 | 312.5836 | 356.4464 | 373.7282 | 401.0673 | 320.2095 | 311.3791 | 232.2369 | 243.6559 | 195.5536 | 251.7556 | 253.2095 | 250.1396 | 193.6908 | 187.4464 | 1443.4838 | 1157.3092 | 874.6259 | 593.0624 | 808.0723 | 866.0324 | 746.7257 | 491.4189 | 1762.0374 | 1419.6559 | 1105.4838 | 925.5836
507 | 7.5714 | 1646.8771 | 1491.2025 | 946.5182 | 702.6578 | 1179.4218 | 1022.0531 | 685.7011 | 594.9232 | 567.2989 | 405.6844 | 357.8771 | 409.4483 | 1328.528 | 897.3422 | 700.7472 | 727.1229 | 1062.9735 | 722.8701 | 647.0796 | 647.7738 | 1127.1327 | 1576.7738 | 1232.6509 | 681.6676 | 1693.5112 | 1401.5084 | 939.1829 | 953.6844 | 701.4818 | 853.8603 | 534.528 | 501.0866 | 949.9204 | 1034.3687 | 648.06 | 285.7975 | 2282.6382 | 1901.0335 | 1265.94 | 1325.8073 | 1116.116 | 1197.0363 | 690.2626 | 534.1927 | 286.0866 | 330.2989 | 246.6578 | 457.3687 | 1116.8408 | 966.9469 | 821.7709 | 475.6844 | 1014.9371 | 1045.6509 | 636.6844 | 440.7374 | 1168.0265 | 981.4749 | 890.7011 | 540.7109 | 588.6844 | 525.4916 | 440.1662 | 325.0098 | 1559 | 920.2291 | 798.7779 | 685.5517 | 976.6844 | 987.3855 | 1096.2989 | 742.2724 | 1169.0363 | 1184.8771 | 911.2025 | 843.7709 | 1118.8506 | 1197.6746 | 894.3953 | 832.4651 | 661.2989 | 659.9665 | 494.9302 | 486.0964 | 557.8603 | 545.6313 | 477.2989 | 326.7109 | 768.2724 | 792.2193 | 644.4651 | 524.6578 | 2091.0531 | 2107.9735 | 1612.6382 | 1307.8436 | 2004.6145 | 2219 | 1362.8771 | 1493.6509 | 253.5949 | 188.4385 | 155.1662 | 253.5949 | 398.1858 | 413.7975 | 285.3687 | 331.3254 | 394.9931 | 381.3953 | 335.6676 | 306.8073 | 193.8506 | 248.2458 | 208.8338 | 224 | 270.6411 | 265.1229 | 210.5782 | 180.0866 | 1425.1523 | 1132.0894 | 922.3422 | 533.3156 | 830.8869 | 851.6047 | 757.0531 | 539 | 1783.9567 | 1431.2626 | 1113.6844 | 913.9371
508 | 7.5824 | 1618.9704 | 1463.5265 | 937.4204 | 710.326 | 1147.8952 | 996.1902 | 718.0413 | 539.2729 | 544.2212 | 402.8525 | 350.8112 | 389.8835 | 1262.643 | 886.6637 | 643.9793 | 728.1682 | 1063.9793 | 672.705 | 635.2626 | 586.9897 | 1088.295 | 1524.326 | 1181.705 | 647.1371 | 1599.4011 | 1418.3997 | 981.0841 | 944.8318 | 721 | 827.8318 | 498.0517 | 542.0944 | 974.705 | 1020.6003 | 685.4425 | 305.3687 | 2311.0234 | 1930.6223 | 1249.8615 | 1322.3997 | 1048.3584 | 1141.7905 | 708.3053 | 535.9793 | 318.295 | 320.1888 | 247.379 | 438.5059 | 1124.0338 | 964.0958 | 846.9793 | 464.2846 | 1043.8642 | 987.9897 | 635.0103 | 420.8422 | 1152.2006 | 963.9145 | 911.4631 | 526.6844 | 584 | 516.705 | 463.1048 | 296.0103 | 1545.8008 | 891.5472 | 885.8422 | 697.0117 | 971.3053 | 1006.2419 | 1066.7374 | 719.0427 | 1132.0841 | 1182.9056 | 888.8525 | 863.4218 | 1107.4645 | 1210.1799 | 868.7374 | 826.5369 | 644.9366 | 627.5989 | 489.326 | 495.326 | 536 | 546.6637 | 469.0413 | 325.8835 | 755 | 774.938 | 627.6533 | 522.6003 | 2064.4955 | 2122.6416 | 1685.2522 | 1263.0841 | 1956.8126 | 2208.1165 | 1358.8215 | 1447.3363 | 293.4321 | 196.4838 | 170.2316 | 293.4321 | 352.4735 | 435.9159 | 274.6106 | 313.031 | 353.0517 | 370.0944 | 309.9366 | 299 | 202.6106 | 231.348 | 192.1475 | 232.3363 | 247.3053 | 267.7891 | 226.5899 | 209.0531 | 1334.5679 | 1086.6416 | 913.7478 | 538.2316 | 806.6313 | 888.8408 | 748.0944 | 527.6533 | 1757.8952 | 1458.3363 | 1118.9573 | 962.0841
509 | 7.5934 | 1509.6598 | 1448.8299 | 966.3913 | 724.586 | 1102.2231 | 963.5841 | 766.0472 | 576.3422 | 545.4858 | 418.7542 | 333.0945 | 415.0265 | 1213.7807 | 854.6106 | 631.3157 | 726.3933 | 1037.172 | 691.5841 | 679.1682 | 590.412 | 1067.2212 | 1541.7561 | 1190.586 | 637.0964 | 1541.4877 | 1404.4689 | 996.0472 | 929.0492 | 712.4651 | 808.3913 | 546.293 | 559 | 994.5595 | 1003.5614 | 658.4651 | 289.7561 | 2180.6106 | 1891.5311 | 1328.7581 | 1307.9812 | 1034.9489 | 1108.24 | 698.9282 | 512.0983 | 309.9017 | 344.4896 | 247.121 | 412.3648 | 1035.7995 | 917.5104 | 836.2665 | 448.3402 | 968.9547 | 973.1229 | 632.8072 | 420.414 | 1113.0019 | 1017.8072 | 913.7316 | 522.9017 | 584 | 530.9508 | 476.8318 | 309.6578 | 1501.7561 | 922.4347 | 872.7335 | 637.1701 | 972.9017 | 1009.1966 | 1060.8733 | 685.7561 | 1097.5446 | 1175.2665 | 891.8299 | 831.0964 | 1026.9774 | 1165.6333 | 847.5104 | 824.7561 | 647.5142 | 678.4405 | 506.2684 | 504.7089 | 549.6559 | 524.6087 | 491.1986 | 365.414 | 767.6805 | 715.4877 | 597.3157 | 515.0718 | 2043.4839 | 2131.6617 | 1677.4689 | 1251.9547 | 1873.1663 | 2154.4424 | 1350.6067 | 1454.2723 | 265 | 211.172 | 173.9263 | 265 | 364.4877 | 417.7827 | 276.0964 | 330.8072 | 387.6371 | 383.586 | 307.881 | 291.9282 | 206.2911 | 206.9735 | 194.7788 | 255.1229 | 240.6106 | 277.4896 | 209.8034 | 183.8053 | 1353.2458 | 1112.9754 | 908.8506 | 540.4631 | 810.7827 | 930.2005 | 764.7561 | 496.0964 | 1702.4689 | 1456.4935 | 1144.1966 | 958.7581
510 | 7.6044 | 1494.6106 | 1426.9611 | 944.1332 | 692.2276 | 1052.9651 | 975.291 | 765.6415 | 568.3299 | 569.0349 | 444.6455 | 368.1435 | 388.4282 | 1206.4139 | 842.0246 | 661.1721 | 684.1475 | 1018.5615 | 724.0389 | 702.3259 | 612.5329 | 1037.1824 | 1539.9754 | 1189.7336 | 642.1681 | 1536.3402 | 1323.1721 | 1018.1824 | 922.7439 | 685.4877 | 788.1824 | 541.2561 | 536.203 | 990.5923 | 987.0349 | 642.7582 | 283.6209 | 2170.8422 | 1965.6558 | 1308.2704 | 1222.3545 | 1066.3648 | 1135.1395 | 676.4877 | 493.3402 | 317.3894 | 309.8771 | 265.5615 | 425.1578 | 1104.582 | 945.4631 | 869.3545 | 461.3402 | 913.5123 | 962.3545 | 616.7724 | 431.5123 | 1090.4528 | 997.418 | 917.793 | 532.9508 | 581.6947 | 521.1189 | 444.4528 | 317.9262 | 1501.7684 | 964.08 | 861.9303 | 657.2459 | 959.3854 | 981.6701 | 1127.2418 | 680.9017 | 989.2356 | 1198.3648 | 890.709 | 838.2173 | 1015.9057 | 1182.5615 | 872.1332 | 800.6907 | 625.502 | 650.9017 | 503.7439 | 478.2316 | 576.8874 | 533.8033 | 448.5512 | 373.1824 | 799.3156 | 723.1475 | 608.9857 | 530.5963 | 2069.3465 | 2086.1229 | 1596.9405 | 1200.3545 | 1920.1538 | 2110.2459 | 1381.0697 | 1388.7193 | 259.8771 | 185.9017 | 180.6947 | 259.8771 | 363.9508 | 387.8668 | 301.6598 | 301.1967 | 347.1721 | 368.9017 | 276.8811 | 267.9508 | 230.7193 | 234.5963 | 219.9508 | 241.0246 | 230.0738 | 241.8525 | 226.8279 | 184.6844 | 1316.084 | 1103.1927 | 931.8731 | 542.1824 | 785.7336 | 850.916 | 754.0103 | 518.5861 | 1617.8422 | 1382.6741 | 1137.418 | 918.5472
511 | 7.6154 | 1510.9528 | 1379.9203 | 937.7847 | 653.6963 | 1062.0884 | 967.18 | 691.1947 | 568.3362 | 538.705 | 431.4631 | 331.1741 | 410 | 1171.3037 | 840.1681 | 659.9881 | 711.9322 | 1020.3628 | 757.3894 | 644.8259 | 562.4425 | 1039.2006 | 1525.7256 | 1214.5103 | 586.6991 | 1508.9469 | 1325.2094 | 1006.7787 | 923.3422 | 698.7641 | 770.8731 | 546.8319 | 469.4631 | 899.7256 | 960.4631 | 675.6106 | 281.1534 | 2139.3184 | 1895.0087 | 1335.0944 | 1204.6666 | 1037.3834 | 1054.0856 | 691.1063 | 476.6844 | 299.9734 | 291.5103 | 263.8731 | 418.8319 | 1069.3009 | 939.9262 | 869.3834 | 443.8791 | 930.8378 | 969.6312 | 680.2684 | 414.4778 | 1082.5191 | 1022.4984 | 909.8466 | 523.2419 | 564.5309 | 482.5163 | 469.5369 | 310.8791 | 1465.0766 | 898.0325 | 906.7522 | 677.8731 | 887.4159 | 972 | 1100.5634 | 667.6578 | 1048.8053 | 1180.6578 | 909.5575 | 798.0944 | 1065 | 1191.0738 | 853.705 | 761.4544 | 652.7787 | 634.705 | 509.7109 | 471.705 | 527.3628 | 521.8113 | 486.0738 | 369.0266 | 760.3569 | 744.4897 | 578.3422 | 509.0472 | 1979.8644 | 2087.3097 | 1623.7581 | 1213.8053 | 1838.3362 | 2119.0619 | 1347.5222 | 1375.0678 | 250.9056 | 174.8053 | 168.6312 | 250.9056 | 364.1741 | 400.1209 | 312.9469 | 323.3894 | 365.0472 | 352.705 | 316.0206 | 268.7109 | 224.8525 | 222.7109 | 211.0472 | 235.9469 | 233.0944 | 233.6638 | 214.59 | 198.9469 | 1297.0738 | 1070.5634 | 868.0797 | 547.4219 | 812.3894 | 884.1534 | 729.0266 | 516.115 | 1633.8319 | 1426.4837 | 1175.3834 | 887.2891
512 | 7.6264 | 1487.5584 | 1364.6859 | 951.9818 | 716.7891 | 1006.1745 | 988.4537 | 700.5099 | 540.8771 | 521.1047 | 423.5432 | 356.4219 | 406.6313 | 1096.7011 | 829.2458 | 608.1411 | 734.5083 | 1013.7026 | 749.0698 | 631.3687 | 588.4219 | 1052.9636 | 1498.9302 | 1210.6146 | 558.475 | 1483.0698 | 1277.088 | 995.7542 | 909.8756 | 712.0501 | 750.4917 | 555.2276 | 479.7906 | 879.387 | 946.9287 | 677.1927 | 310.824 | 2024.4036 | 1821.3338 | 1313.3854 | 1146.4234 | 987.5599 | 1098.4719 | 722.0865 | 479.2276 | 296.8604 | 284.2458 | 247.4219 | 421.3323 | 1104.2625 | 936.7193 | 830.2276 | 435.8058 | 956.8756 | 945.4552 | 672.0167 | 371.5797 | 1017.3005 | 960.5797 | 878.7724 | 512.4385 | 542.176 | 502.1578 | 470.1578 | 310.6662 | 1359.2807 | 930.5781 | 888.6844 | 656.088 | 879.1214 | 965.824 | 1115.2458 | 669.6146 | 1049.8771 | 1178.4036 | 886.3854 | 796.5979 | 1035.2428 | 1181.0865 | 850.7026 | 830.8589 | 642.5964 | 631.7026 | 524.3505 | 464.7724 | 519.2989 | 546.7011 | 489.8422 | 367.0865 | 722.8604 | 754.5615 | 588.7375 | 500.4385 | 2010.4188 | 2057.0182 | 1636.2458 | 1180.6282 | 1810.3156 | 2047.1214 | 1385.9287 | 1336.7709 | 277.1062 | 176.7193 | 159.8953 | 277.1062 | 366.4021 | 400.8953 | 291.5615 | 326.0182 | 349.9818 | 353.9135 | 300.6495 | 273.8058 | 229.9484 | 233.9818 | 197.9469 | 218.7724 | 222.3338 | 256.912 | 239.1229 | 184.5797 | 1308.7026 | 1091.9833 | 891.1745 | 582.8771 | 802.3854 | 900.9105 | 734.3854 | 482.4036 | 1635.7709 | 1407.1229 | 1149.9833 | 880.5099
513 | 7.6374 | 1407.8791 | 1394.1769 | 947.7719 | 686.6607 | 987.4464 | 933.5487 | 714.763 | 547.6696 | 505.414 | 410.2232 | 373.8977 | 394.4838 | 1094.7021 | 825.0373 | 629.293 | 705.1258 | 1007.5771 | 717.6096 | 641.1721 | 605.3117 | 1032.53 | 1505.4838 | 1217.3815 | 592.1209 | 1448.6794 | 1303.4602 | 992.6372 | 861.0187 | 656.4789 | 746.1769 | 549.7395 | 498.8653 | 889.8839 | 931.53 | 667.3628 | 289.7256 | 2023.3117 | 1850.237 | 1289.9164 | 1204.8977 | 947.1209 | 1058.0325 | 689 | 479.5998 | 311.7256 | 294.6883 | 260.7955 | 401.0325 | 1078.2094 | 929.5536 | 837.0463 | 419.8742 | 893.9537 | 905.47 | 658.3766 | 387.7768 | 1076.414 | 973.2605 | 895.3953 | 505.1396 | 548.9164 | 504.3953 | 460.6745 | 297.293 | 1345.349 | 923.0649 | 895.9302 | 663.707 | 847.4464 | 955.8604 | 1125.3953 | 681.9489 | 1031.1769 | 1175.5536 | 876.3953 | 830.7768 | 964.5674 | 1135.6932 | 880.9115 | 797.4002 | 628.1583 | 652.5349 | 518.8093 | 483.7395 | 539.4651 | 531.2232 | 476.47 | 333.1209 | 730.1071 | 755.414 | 600.1396 | 494.0187 | 1919.526 | 2069.7256 | 1632.6234 | 1124.5763 | 1804.5349 | 2021.8929 | 1347.3815 | 1308.8791 | 286.5909 | 173.0698 | 177.6372 | 286.5909 | 337.2045 | 406.3304 | 293 | 334.9164 | 347.8231 | 378.7955 | 306.1907 | 259.6323 | 238.3677 | 212.4838 | 189.9813 | 231 | 243.0325 | 240.7955 | 235.5536 | 196.0885 | 1308.7305 | 1078.9586 | 873.9115 | 567.6932 | 785.3628 | 876.9067 | 734.0649 | 495.9627 | 1638.6461 | 1394.1769 | 1168.0187 | 912.3442
514 | 7.6484 | 1383.438 | 1317.8362 | 977.031 | 621.6814 | 990.7876 | 935.4911 | 688.9912 | 560.0088 | 529.0355 | 416.7788 | 361.2522 | 390.5797 | 1140.1903 | 799.1681 | 641.9071 | 668.7168 | 933.8097 | 685.2346 | 644.6106 | 583.3141 | 1056.8716 | 1478.376 | 1195.9469 | 600.0753 | 1411.3717 | 1338.2301 | 991.0619 | 843.8097 | 677.7035 | 710.1372 | 561.115 | 481.1593 | 843.6725 | 960.146 | 663.1372 | 283.4115 | 1996.1238 | 1874.177 | 1254.5663 | 1171.1858 | 937.0619 | 1053.3229 | 693.2743 | 495.0398 | 303.885 | 297.4513 | 251.3806 | 404.5663 | 1070.739 | 915.5266 | 874.7566 | 463.3053 | 872.4249 | 870.7478 | 689.2434 | 395.8762 | 1064.6194 | 954.5575 | 919.1284 | 489.1681 | 538.9071 | 499.4292 | 444.5266 | 295.5575 | 1284.3539 | 952.115 | 893.2301 | 651.0929 | 866.3584 | 954.7345 | 1122.5663 | 680.0531 | 1001.5487 | 1148.0929 | 888.8319 | 798.5751 | 973.7257 | 1128.2877 | 884.7345 | 780.1284 | 597.7257 | 655.4115 | 538.9912 | 489.0088 | 551.2434 | 542.6637 | 457.0133 | 324.2832 | 686.4823 | 739.3451 | 590.5797 | 490.854 | 1868.2877 | 2053.6416 | 1621.2965 | 1192.4292 | 1815.6549 | 2034.0707 | 1359.8364 | 1323.823 | 252.1903 | 162.3363 | 178.8097 | 252.1903 | 343.5663 | 386.3584 | 291.4734 | 327.0443 | 371.3451 | 361.7478 | 301.2743 | 295.3982 | 192.5266 | 219.8762 | 198.6416 | 233.1372 | 251.146 | 233.823 | 225.1903 | 173.3363 | 1267.823 | 1019.7833 | 889.6416 | 532.8097 | 767.7035 | 929.0619 | 767.0841 | 511.7566 | 1692.407 | 1370.3496 | 1173.4027 | 911.8938
515 | 7.6594 | 1330.3048 | 1253.8673 | 962.9184 | 617.7946 | 966.8112 | 896.9273 | 715.2704 | 536.2704 | 558.051 | 405.4248 | 359.4464 | 391.7896 | 1141.5191 | 815.0089 | 669.3176 | 673.2015 | 932.4809 | 710.5791 | 643.1416 | 562.8112 | 1014.6135 | 1442.0344 | 1175.9783 | 615.3776 | 1449.8367 | 1325.2104 | 994.5969 | 843.1161 | 655.176 | 717.8584 | 538.5663 | 485.8023 | 832.3992 | 915.8623 | 687.0561 | 281.8023 | 1949.8367 | 1846.2704 | 1244.0089 | 1092.5447 | 906.2959 | 1004.6352 | 715.0689 | 454.9656 | 312.1416 | 285.94 | 263.2487 | 387.3393 | 1097.8712 | 937.7385 | 857.588 | 452.8839 | 903.8623 | 875.0944 | 700.3776 | 428.5064 | 1039.0944 | 949.9528 | 931.5319 | 488.176 | 552.3431 | 487.1849 | 451.176 | 292.5408 | 1236.4809 | 931.7896 | 886.8801 | 649.7257 | 864.4719 | 948.2666 | 1089.1416 | 685.7296 | 1008.6135 | 1110.5191 | 876.1671 | 765.8673 | 970.0344 | 1169.4719 | 855.7168 | 795.3903 | 582.9184 | 649.3559 | 554.1543 | 452.5663 | 556.6607 | 548.4592 | 489.1072 | 318.8457 | 672.6097 | 724.4209 | 604.4936 | 506.5536 | 1905.3431 | 2035.2015 | 1607.2998 | 1135.7207 | 1817.875 | 2002.7513 | 1393.4758 | 1331.4031 | 259.236 | 161.94 | 170.176 | 259.236 | 324.1161 | 386.06 | 284.1888 | 335.4975 | 343.7168 | 364.824 | 294.8023 | 273.764 | 207.1161 | 237.8967 | 198.94 | 228.1543 | 244.9745 | 253.4719 | 230.0128 | 162.1416 | 1265.5574 | 1066.8023 | 881.0472 | 527.0344 | 743.2704 | 887.8151 | 753.2487 | 506.6569 | 1638.1416 | 1415.3865 | 1200.2959 | 922.1377
516 | 7.6704 | 1329.1755 | 1283.2064 | 954.6391 | 686.794 | 966.0929 | 925.3712 | 700.8456 | 535.3712 | 498 | 387 | 368.0103 | 378.9277 | 1112.4021 | 841.1755 | 682.206 | 660.8865 | 947.7424 | 696.4017 | 629.0826 | 576.9381 | 988.3299 | 1458.2989 | 1149.9794 | 595.0413 | 1412.1651 | 1322.8972 | 1026.8146 | 857.1032 | 666.2886 | 723.3402 | 524.2267 | 453.969 | 843.0103 | 878.8452 | 721.0723 | 255.2473 | 1894.6808 | 1832.5054 | 1288.9794 | 1057.8142 | 879.3299 | 1003.3609 | 736.3815 | 441.0103 | 300.0619 | 292.4434 | 250.5359 | 421.9381 | 1091.969 | 970.7424 | 854.732 | 445.4946 | 857.6082 | 878.2576 | 678.7217 | 399.2271 | 1011.9075 | 943.4228 | 888.3918 | 500.9381 | 533.2267 | 522.4847 | 450.7424 | 309.8761 | 1235.5776 | 927.1751 | 894.9385 | 687.1858 | 851.2576 | 1000.0929 | 1106.0925 | 690.2576 | 987.6082 | 1101.8555 | 835.0516 | 793.2271 | 964.8559 | 1126.2374 | 871.6288 | 755.7527 | 568.3712 | 618.5256 | 531.6804 | 443.505 | 540.206 | 540.3609 | 475.2783 | 346.268 | 684.2064 | 744.2064 | 613.6495 | 488.0929 | 1849.6086 | 2018.9277 | 1539.4843 | 1114.742 | 1750.6701 | 2029.3196 | 1322.9174 | 1281.3712 | 256.5875 | 179 | 175.3505 | 256.5875 | 339.7527 | 366.031 | 273.031 | 295.2267 | 346.1032 | 345.1342 | 270.8865 | 285.3196 | 230.6701 | 207.9071 | 207.7527 | 212.6082 | 274.3402 | 233.6598 | 215 | 157.3196 | 1208.5566 | 1026.0619 | 888.5669 | 552.206 | 739.0723 | 838.9897 | 741.1957 | 520.495 | 1615.1755 | 1416.2684 | 1159.8043 | 953.5466
517 | 7.6814 | 1253.2536 | 1239.9203 | 948.0236 | 658.236 | 942.8672 | 908.4189 | 700.5606 | 531.4218 | 495.9469 | 404.4514 | 362.9203 | 388.5516 | 1082.5752 | 789.5486 | 686.8112 | 688.1328 | 928.2094 | 733.6283 | 626.6874 | 588.6578 | 975.997 | 1438.6518 | 1173.5045 | 572.2861 | 1394.7139 | 1302.2951 | 1011.3452 | 835.2891 | 658.9174 | 724.0531 | 554.9233 | 468.8968 | 841.6844 | 889.1533 | 701.8672 | 270.2891 | 1829.3953 | 1800.3687 | 1310.4514 | 1095.7109 | 869.3923 | 1017.1623 | 734.8702 | 432.8407 | 294.8437 | 277.5015 | 276.3422 | 424.761 | 1090.472 | 953.6047 | 845.4218 | 486.8702 | 860.4454 | 901.2124 | 672.4484 | 358.0797 | 980.3983 | 939.6874 | 866.7345 | 518.8171 | 552.6313 | 506.6874 | 433.2625 | 335.3687 | 1172.6017 | 946.7079 | 860.4779 | 633.2064 | 847.5221 | 982.6844 | 1139.6548 | 697.1298 | 983.2595 | 1125.8938 | 835.6077 | 758.239 | 933.4749 | 1093.003 | 875.5782 | 756.4514 | 566.8171 | 634.3657 | 525.5782 | 472.6047 | 555.761 | 545.9499 | 487.7109 | 355.5782 | 666.9263 | 697.4985 | 625.5546 | 488.1357 | 1825.77 | 2039.5015 | 1597.9263 | 1138.3097 | 1739.3923 | 1999.4925 | 1355.0826 | 1290.0826 | 258.997 | 179.3422 | 172.3422 | 258.997 | 325.3953 | 369.9233 | 268.0266 | 316.6844 | 332.1593 | 328.9233 | 288.5516 | 291.764 | 219.7345 | 222.236 | 201.2655 | 223.6578 | 265.1298 | 239.1062 | 219.1062 | 175.0561 | 1194.0707 | 1043.77 | 894.3126 | 573.236 | 749.9794 | 852.9764 | 783.3983 | 489.6844 | 1594.6874 | 1391.7169 | 1154.5576 | 893.1799
518 | 7.6924 | 1194.623 | 1252.4327 | 919.3284 | 644.265 | 935.1268 | 898.5074 | 742.1868 | 539.7724 | 498.3805 | 421.3397 | 353.8097 | 379.6455 | 1053.8358 | 757.4179 | 642.6829 | 690.2016 | 895.2502 | 701.7463 | 632.7949 | 593.3171 | 946.3918 | 1395.7984 | 1221.2798 | 562.7872 | 1416.9739 | 1311.0712 | 1022.1121 | 833.7724 | 621.9739 | 709.9218 | 569.7463 | 488.8358 | 827.4031 | 863.4066 | 694.1268 | 260.265 | 1810.6047 | 1760.1085 | 1316.7054 | 1079.1494 | 841.9105 | 1035.2128 | 757.7984 | 433.4066 | 304.8097 | 283.6195 | 281.9626 | 405.5821 | 1072.4553 | 956.7984 | 855.8992 | 491.7202 | 851.444 | 897.5039 | 676.3918 | 365.4179 | 979.1607 | 956.7837 | 876.7759 | 505.1868 | 550.3545 | 526.6195 | 429.7724 | 327.0113 | 1151.8619 | 939.179 | 909.8245 | 599.7089 | 826.179 | 980.4553 | 1111.5187 | 666.0486 | 941.6568 | 1111.9739 | 850.6829 | 784.5074 | 942.9366 | 1099.5673 | 873.2537 | 796.3805 | 585.444 | 614.4813 | 508.3658 | 470.4813 | 546.1529 | 577.9887 | 488.1642 | 341.2016 | 690.086 | 691.735 | 641.6455 | 514.8471 | 1858.1868 | 2007.3362 | 1633.1381 | 1133.7907 | 1703.7576 | 1966.7759 | 1374.3023 | 1309.6568 | 242.8619 | 174.3284 | 172.6455 | 242.8619 | 329.6455 | 386.1642 | 279.5708 | 316.5821 | 346.1268 | 355.0895 | 283.8992 | 291.9105 | 203.5447 | 218.5447 | 213.556 | 221.5821 | 254.6082 | 253.735 | 230.8992 | 196.556 | 1146.2685 | 1082.9218 | 879.3432 | 552.8844 | 801.8992 | 866.4292 | 816.5447 | 498.0895 | 1604.6942 | 1435.4292 | 1169.0599 | 858.7724
519 | 7.7034 | 1214.5614 | 1248.9566 | 922.2036 | 622.4641 | 925.4297 | 889.0599 | 708.003 | 527.9626 | 500.8309 | 393.9341 | 339.9311 | 388.1691 | 1043.9626 | 760 | 657.2066 | 672.3323 | 868.8712 | 691.0659 | 639.2784 | 582.8279 | 936.1662 | 1409.2665 | 1211.3923 | 577.5614 | 1371.524 | 1247.5015 | 1003.0689 | 841.768 | 606.5988 | 693.5045 | 545.1288 | 487.3982 | 827.9745 | 878.6647 | 698 | 263.0374 | 1763.9656 | 1695.9371 | 1284.6048 | 1055.1033 | 821.6647 | 991.0973 | 749.6273 | 450.8653 | 303.4012 | 267.232 | 274.4955 | 397.8653 | 1060.4985 | 950.0943 | 868.1347 | 455.3952 | 852.53 | 853.5644 | 651.4955 | 373.8683 | 944.2724 | 960.1721 | 916.6961 | 465.1347 | 536.6961 | 507.6647 | 439.9686 | 307.7335 | 1162.4985 | 940.0599 | 911.8937 | 612.3697 | 847.232 | 962.997 | 1142.5763 | 637 | 923.2006 | 1133.6422 | 850.5359 | 808.0689 | 940.6332 | 1088.7395 | 887.6018 | 809.8338 | 593.1317 | 605.4297 | 498.1691 | 483.8024 | 558.0659 | 584.7964 | 483 | 323.3323 | 663.7365 | 709.5015 | 633.6647 | 499.2665 | 1800.53 | 1976.2469 | 1626.5359 | 1195.1602 | 1698.7395 | 1947.6468 | 1390.7769 | 1285.6991 | 259 | 164.7335 | 175.6677 | 259 | 334.5015 | 376.9656 | 283.7964 | 309.232 | 354.0344 | 363.9656 | 284.3982 | 280.8338 | 207.7994 | 229.768 | 202.9341 | 223.4012 | 256.3982 | 250.9626 | 238 | 183.3668 | 1160.4267 | 1058.4356 | 904.4012 | 526.7994 | 783.3264 | 874.3068 | 817.1317 | 512.4671 | 1565.8309 | 1405.53 | 1124.768 | 890.9745
520 | 7.7144 | 1149.9173 | 1200.5413 | 955.6136 | 618.8879 | 874.3288 | 859.267 | 713.733 | 508.0649 | 506.9853 | 404.6859 | 344.1947 | 391.1121 | 1026.7183 | 755.4513 | 688.7655 | 668.6534 | 905.9602 | 691.6859 | 685.5707 | 556.3393 | 946.6461 | 1396.2492 | 1186.2596 | 563.711 | 1304.4189 | 1244.4911 | 1009.7979 | 851.5162 | 591.4189 | 731.8555 | 525.4764 | 453.2566 | 833.6503 | 862.9675 | 680.5634 | 282.9351 | 1758.8953 | 1739.711 | 1300.2168 | 1038.4734 | 817.7183 | 958.0796 | 707.2094 | 446.9277 | 305.3141 | 265.0649 | 250.5487 | 396.5811 | 1031.289 | 931.0576 | 883.0649 | 448.5015 | 820.3864 | 831.0398 | 649.1549 | 375.6607 | 982.5162 | 952.5456 | 887.8304 | 481.2021 | 513.8953 | 491.9675 | 447.8304 | 306.7257 | 1163.2345 | 925.8083 | 876.1696 | 628.9675 | 823.0796 | 935.6607 | 1157.755 | 669.5989 | 936.4764 | 1152.8524 | 875.7183 | 801.91 | 938.8628 | 1085.5456 | 869.8126 | 786.8451 | 597.4042 | 602.4691 | 528.0251 | 487.2492 | 545.0398 | 560.8628 | 489.0649 | 322.3068 | 675.2817 | 704.2168 | 612.2817 | 482.0796 | 1786.2021 | 2027.7477 | 1617.9822 | 1141.8702 | 1723.2168 | 1869.1402 | 1465.8555 | 1291.3717 | 258.6209 | 171.6859 | 178.2094 | 258.6209 | 334.5236 | 374.9277 | 270.4764 | 316.9204 | 349.6283 | 360.7906 | 297.2596 | 277.9351 | 205.5162 | 226.2492 | 198.5236 | 228.3466 | 236.6607 | 233.7183 | 232.6932 | 188.1696 | 1124.823 | 1040.7802 | 911.6209 | 526.0325 | 741.2743 | 900.5634 | 787.289 | 514.2419 | 1550 | 1374.1445 | 1170.9057 | 920.91
521 | 7.7254 | 1119.3501 | 1191.17 | 987.0885 | 625.9115 | 845.828 | 878.2606 | 711.1761 | 513.6952 | 538.3038 | 408.5211 | 370.0865 | 372.7394 | 1018.3018 | 765.2193 | 663.7817 | 658.5201 | 899.8239 | 690.0423 | 644.4789 | 560.6076 | 944.7364 | 1388.3491 | 1230.6962 | 588.9557 | 1293.174 | 1238.0885 | 1001.9135 | 865.7404 | 592.3058 | 726.4759 | 539.826 | 437.5241 | 766.0875 | 862.3048 | 657.4789 | 275.3481 | 1740.8672 | 1775.9135 | 1289.3058 | 1002.0463 | 828.4779 | 919.4769 | 692.7817 | 444.0423 | 288.5644 | 269.9125 | 266.6097 | 407.0433 | 998.2173 | 948.0423 | 873.8239 | 485.4346 | 828.0463 | 835.2193 | 666.5624 | 367.4789 | 992.827 | 924.0915 | 901.7414 | 502.7837 | 505.2596 | 508.1328 | 434.7384 | 300.6519 | 1149.8662 | 943.6489 | 866.172 | 620.0865 | 795.0433 | 919.6519 | 1121.9175 | 676.0382 | 937.9115 | 1097.0875 | 885.3038 | 765.4779 | 943.2616 | 1064.1358 | 837.2626 | 742.826 | 615.9125 | 631.3038 | 541.2163 | 465.6519 | 552.7414 | 577.002 | 500.5654 | 348.6107 | 639.8672 | 699.5674 | 600.3481 | 466.175 | 1794.8692 | 1991.5241 | 1591.0885 | 1117.6952 | 1702.1308 | 1893.6026 | 1465.172 | 1288.4759 | 250.173 | 184.1308 | 172 | 250.173 | 328.7837 | 366.9547 | 289.9135 | 330.8682 | 334.9135 | 357.6076 | 316.2586 | 268 | 207.2173 | 198.7817 | 183 | 230.2616 | 237.8712 | 234.3028 | 218.1298 | 186.3903 | 1158.6982 | 1007.5674 | 919.2183 | 524.3471 | 761.828 | 871.2173 | 745.2163 | 501.2596 | 1581.6992 | 1365.827 | 1237.6962 | 893.4789
522 | 7.7364 | 1113.8652 | 1162.0256 | 1017.4764 | 598.7527 | 865.9454 | 874 | 748.2655 | 531.2183 | 512.8362 | 378.2035 | 355.1199 | 360.4508 | 973.3127 | 771.0108 | 660.6873 | 629.5418 | 880.8983 | 680.2655 | 665.5128 | 531.0654 | 900.9454 | 1431.8329 | 1247.2837 | 563.0654 | 1265.1307 | 1276.1456 | 1014.8619 | 845.6543 | 609.3672 | 711.971 | 525.8908 | 465.3491 | 769.56 | 871.4218 | 681.7419 | 266.4872 | 1692.6617 | 1755.7635 | 1318.4764 | 1011.938 | 836.7237 | 897.5782 | 699.3747 | 427.8073 | 278.6146 | 269.2655 | 273.5418 | 397.9454 | 1016.7783 | 930.1927 | 859.7419 | 474.1199 | 841.5707 | 867.2473 | 651.2183 | 391.7419 | 987.2399 | 974.2763 | 936.4218 | 524.0546 | 506.3747 | 537.7709 | 432.2655 | 301.8437 | 1096.1381 | 912.531 | 859.7783 | 634.182 | 804.5128 | 927.3019 | 1170.2763 | 628.0256 | 915.1927 | 1113.4764 | 860.2399 | 747.0835 | 921.1125 | 1124.469 | 860.1927 | 741.4036 | 599.1199 | 611.8908 | 537.3019 | 464.8255 | 578.5418 | 579.9744 | 481.6071 | 354.2035 | 613.8619 | 723.9454 | 608.44 | 461.0108 | 1792.8255 | 2008.4508 | 1613 | 1091.6253 | 1700.5418 | 1831.6038 | 1456.7601 | 1256.6146 | 250.9165 | 185.3672 | 175.6328 | 250.9165 | 337.5418 | 342.4036 | 298.4218 | 319.6691 | 342.6146 | 332.909 | 295.3019 | 260.3309 | 204.3854 | 200.531 | 190.6691 | 221.029 | 241.2217 | 231.5674 | 214.6509 | 171.7891 | 1170.9926 | 1037.1927 | 937.2473 | 545.4838 | 767.0108 | 854.6617 | 768.3457 | 499.9528 | 1609.6071 | 1373.56 | 1228.4508 | 895.1381
523 | 7.7474 | 1078.9646 | 1205.6549 | 1014.6106 | 598.9027 | 824.9734 | 874 | 741.9469 | 547.0354 | 483.0885 | 382.6019 | 356.5576 | 351.8053 | 964.823 | 727.1062 | 643.8052 | 624.1593 | 877.0354 | 663.9645 | 682.7522 | 528.9646 | 872.8672 | 1452.5309 | 1250.6017 | 557.6372 | 1238.4691 | 1277.5398 | 1022.5044 | 793.9912 | 587.7787 | 715.0442 | 542.5222 | 448.2655 | 765.4956 | 835.0707 | 701.2566 | 255.9027 | 1652.115 | 1707.4336 | 1336.4071 | 981.062 | 848.6814 | 892.7522 | 710.5309 | 444.8053 | 273.5133 | 255.8761 | 269.823 | 393.1593 | 1045.5841 | 933.9912 | 904.6284 | 448.938 | 803.1681 | 876.6637 | 676.6017 | 399.1947 | 975.9735 | 950.2566 | 933.3451 | 518.0265 | 525.4336 | 518.4424 | 450.0708 | 301.9292 | 1073.5221 | 948.5576 | 904.8053 | 643.6194 | 808.8584 | 928.6106 | 1139.1858 | 650.8584 | 906.9292 | 1108.115 | 873.0974 | 735.0708 | 889.3894 | 1089.2743 | 880.6284 | 741.1681 | 588.0797 | 621.0355 | 543.6017 | 474.823 | 562.761 | 561.7169 | 474.5399 | 322.4159 | 638.1416 | 715.8319 | 597.9469 | 450.3806 | 1760.3981 | 1991.0708 | 1638.7877 | 1082.9646 | 1706.3894 | 1872.0708 | 1458.4601 | 1263.1593 | 252.5309 | 170.8495 | 179.3363 | 252.5309 | 330.9115 | 343 | 308.239 | 330.1681 | 344.5841 | 324.239 | 300.7699 | 266.0531 | 202.4159 | 220.7522 | 213.2301 | 204.2301 | 222.8142 | 250.2655 | 233.3983 | 179.5664 | 1148.2301 | 1028.9292 | 947.4956 | 553.7344 | 740.9912 | 865.2744 | 789.6459 | 515.0177 | 1570.5133 | 1374.4867 | 1197.761 | 892.4867
524 | 7.7584 | 1087.73 | 1191.0196 | 984.8638 | 601.8638 | 803.9804 | 890.7006 | 735.5571 | 533.5595 | 483.7104 | 419.8687 | 373.146 | 380.4233 | 987.5595 | 706.8589 | 627.4062 | 648.8417 | 858.4209 | 663.6835 | 692.562 | 556.1436 | 851.1387 | 1440.8466 | 1244.5644 | 581 | 1238.5914 | 1265.6982 | 1004.4405 | 805.7104 | 560.2896 | 686.4233 | 567.4478 | 447.562 | 769 | 796.1288 | 677.1583 | 264.8589 | 1642.8442 | 1740.2479 | 1349.9951 | 999.7129 | 822.0123 | 899.9926 | 694.9926 | 459.292 | 280.5571 | 251.119 | 273.7178 | 392.1485 | 1021.8761 | 937.5742 | 923.438 | 427.2896 | 798.5742 | 875.5742 | 668.8515 | 386.8491 | 973.0196 | 941.146 | 952.2626 | 488.8638 | 510.7153 | 505.2798 | 460.8564 | 298.4233 | 1076.1264 | 941.5938 | 916.2945 | 623.1362 | 798.4209 | 908.2847 | 1154.4037 | 659.2626 | 880.7276 | 1077.4331 | 904.8761 | 756.9902 | 917.4233 | 1086.2724 | 898.1534 | 754.1313 | 571.3018 | 645.7251 | 520.4356 | 474.4356 | 568.5595 | 586.2871 | 500.4356 | 328.9951 | 653.7178 | 685.3116 | 600.5497 | 481.5693 | 1687.1607 | 2006.5669 | 1674.1436 | 1121.2773 | 1689.0319 | 1846.8859 | 1431.2798 | 1248.8736 | 245.9853 | 168.7055 | 162.0123 | 245.9853 | 325.5791 | 332.7227 | 311.4331 | 314.4405 | 354.2773 | 328.2896 | 286.2847 | 287.0025 | 205.1411 | 201.8712 | 208.8736 | 215.7178 | 229.0074 | 229.0074 | 234.7251 | 189.146 | 1154.1509 | 1019.4282 | 945.0049 | 515.416 | 742.0049 | 900.438 | 776.827 | 507.8515 | 1545.8687 | 1369.1558 | 1183.0049 | 917.9877
525 | 7.7694 | 1072.1318 | 1173.0964 | 995.1446 | 613.4661 | 829.3923 | 872.4729 | 749.7492 | 543.7492 | 519.5143 | 393.5113 | 364.9164 | 375.3441 | 992.0226 | 698.2733 | 647.2605 | 671.3569 | 872.762 | 701.2251 | 689.058 | 558.5241 | 861.9646 | 1447.8682 | 1270.5369 | 574.8328 | 1239.0384 | 1237.817 | 981.0354 | 826.5369 | 546.5113 | 676.058 | 533.3569 | 479.299 | 760.6303 | 834.8584 | 696.0512 | 251.8682 | 1660.4759 | 1786.5821 | 1361.9646 | 986.5821 | 808.6077 | 900.9872 | 725.061 | 432.701 | 296.9518 | 284.2733 | 255.2251 | 391.3697 | 1014.7877 | 935.2026 | 905.7974 | 435.0964 | 775.939 | 858.6657 | 675.7974 | 393.8328 | 944.9646 | 938.6431 | 993.5241 | 493.8584 | 502.8328 | 508.3923 | 473.4533 | 298.1898 | 1082.7364 | 906.3215 | 909.0482 | 613.7364 | 806.1544 | 903.3923 | 1173.4149 | 667.0452 | 863.4051 | 1060.3087 | 888.5369 | 744.3667 | 906.1769 | 1099.1898 | 885.4051 | 769.0354 | 556.0128 | 642.0964 | 537.6333 | 463.881 | 587.5595 | 566.3795 | 481.5113 | 337.881 | 648 | 662.3825 | 629.3923 | 471.3441 | 1684.8712 | 2028.5369 | 1610.8878 | 1147.7749 | 1657.7041 | 1820.7041 | 1410.6047 | 1234.7749 | 265.5949 | 172.4277 | 148.1672 | 265.5949 | 308.5949 | 325.1672 | 284.6175 | 303.8102 | 359.1898 | 325.9646 | 293.2861 | 286.9646 | 208.881 | 180.762 | 184.2026 | 208.6785 | 226.2508 | 216.119 | 228.8938 | 182.238 | 1156.4179 | 1027.4887 | 954.6205 | 512.2123 | 734 | 892.9292 | 785.9134 | 513.9164 | 1517.3087 | 1362.5369 | 1183.3697 | 947.6913
526 | 7.7804 | 1061.4542 | 1180.5855 | 998.0014 | 629.7743 | 839.5649 | 860.8467 | 754.3009 | 508.0014 | 520.7757 | 392.1505 | 375.4897 | 356.6416 | 969.8304 | 683.7551 | 625.5472 | 644.4542 | 878.1696 | 684.0944 | 673.6033 | 539.4542 | 850.2463 | 1391.0781 | 1285.3776 | 564.2832 | 1255.643 | 1214.2803 | 976.9056 | 807.4173 | 546.9617 | 659.2449 | 540.4145 | 473.2846 | 744.1136 | 826.9262 | 718.4926 | 242.1696 | 1628.379 | 1707.6828 | 1369.264 | 946.171 | 812.9248 | 864.8126 | 722.1902 | 411.5472 | 280.3023 | 269.3023 | 258.9233 | 407.8864 | 1047.264 | 958.829 | 880.3598 | 462.9617 | 737.1519 | 840.736 | 657.6047 | 374.2271 | 936.8687 | 953.3201 | 970.8318 | 507.6608 | 503.1505 | 500.9056 | 458.115 | 302.8481 | 1047.0383 | 905.736 | 908.4911 | 618.3378 | 816.7537 | 913.5649 | 1118.379 | 648.4705 | 897.0738 | 1062.0369 | 903.3776 | 713.4336 | 884.6224 | 1068.5295 | 885.5664 | 757.208 | 560.4911 | 615.6253 | 554.0767 | 467.7168 | 553.9454 | 567.5458 | 485.5841 | 355.7537 | 609.9645 | 680.6608 | 615.1136 | 462.1505 | 1718.5472 | 2016.6622 | 1589.506 | 1149.9631 | 1687.1696 | 1849.264 | 1386.9425 | 1251.4528 | 266.3392 | 169.4145 | 164.1505 | 266.3392 | 307.528 | 362.885 | 273.5841 | 318.1696 | 349.3584 | 318.7743 | 308.1505 | 292.4528 | 189.6239 | 193.4145 | 196.5089 | 201.5664 | 238 | 218.6224 | 234.4911 | 166.4159 | 1157.4351 | 1031.5664 | 962.9631 | 491.5089 | 721.7743 | 891.2271 | 761.7154 | 493.6991 | 1522.2065 | 1378.2832 | 1200.3392 | 941.4351
527 | 7.7914 | 1000.9085 | 1154.9774 | 1014.1142 | 623.8368 | 837.35 | 900.1632 | 770.7443 | 494.8132 | 496.9991 | 395.0236 | 384.6283 | 372.2085 | 956.233 | 708.9519 | 627.3255 | 622.9764 | 850.3736 | 688.116 | 690.0462 | 520.3019 | 842.5575 | 1367.3009 | 1272.1632 | 545.6519 | 1204.5358 | 1279.2321 | 978.9302 | 788.9755 | 578.0689 | 677.3019 | 540.4425 | 447.5113 | 738.3953 | 794.3953 | 674.7443 | 258.5575 | 1596.8604 | 1652.2783 | 1343.6292 | 942.8594 | 792.1868 | 857.534 | 728.2075 | 423.0925 | 273.3255 | 276.2783 | 271.3726 | 404.3028 | 1009.0717 | 979.8142 | 888.3472 | 450.8151 | 734.2783 | 827.3028 | 654.8132 | 388.3708 | 922.8604 | 966.0698 | 949.6274 | 507.4415 | 536.2547 | 509.9066 | 435.4877 | 314.5585 | 1036.8604 | 947.184 | 916.0689 | 652.2094 | 814.2104 | 952.7434 | 1109.65 | 658.0472 | 883.5604 | 1068.7679 | 906.4415 | 731.6274 | 872.7217 | 1062.6736 | 907.8359 | 758.767 | 549.4651 | 601.533 | 541.6274 | 484.4877 | 530.7434 | 593.6745 | 492.6981 | 353.2104 | 586.3245 | 707.417 | 605.2094 | 465.4887 | 1696.1406 | 1981.3953 | 1652.4887 | 1161.3708 | 1691.9302 | 1843.6283 | 1394.3028 | 1245.9538 | 281.9066 | 175.4887 | 174.4651 | 281.9066 | 305.0934 | 388.3028 | 285.8142 | 313.1632 | 347.2792 | 323.534 | 310.5585 | 277.1868 | 191.0453 | 201.8142 | 208 | 205.6972 | 246.3717 | 228.5811 | 219.2792 | 161.6274 | 1136.3953 | 1030.116 | 924.6057 | 490.9075 | 703.7443 | 886.767 | 795.3953 | 480.9764 | 1526.7217 | 1382.4415 | 1183.8151 | 898.0708
528 | 7.8024 | 929.1509 | 1126.1558 | 1029.7187 | 613.9498 | 826.1859 | 880.4523 | 766.5226 | 498.6784 | 502.3367 | 388.4322 | 372.1357 | 377.6784 | 917.3116 | 744.1809 | 661.9848 | 619.5427 | 829.2763 | 677.7689 | 694.1357 | 520.0201 | 844.588 | 1405.7739 | 1265.8191 | 523.2261 | 1208.6884 | 1253.7639 | 978.5678 | 814.2714 | 581.8594 | 671.7689 | 540.7989 | 456.9095 | 745.2512 | 829.8944 | 680.5477 | 256.7689 | 1608.5477 | 1664.794 | 1327.2311 | 953.6784 | 769.7035 | 886.6834 | 747.1307 | 451.7989 | 268.3618 | 289.2714 | 242.2463 | 396.2965 | 979.6633 | 944.9247 | 943.3216 | 444.9597 | 763.5025 | 848.417 | 680.1608 | 389.4926 | 899.6986 | 936.8343 | 938.6533 | 508.8392 | 537.2212 | 523.2965 | 452.9095 | 316.5025 | 1007.4926 | 987.4975 | 901.7187 | 647.5678 | 778.9297 | 941.4473 | 1123.2011 | 634.3618 | 855.2562 | 1095.1908 | 927.8894 | 753.5929 | 852.5678 | 1053.4021 | 932.4071 | 749.4724 | 552.3869 | 634.1106 | 544.9749 | 478.0403 | 566.7085 | 573.8141 | 482.4774 | 346.5728 | 630.1406 | 728.382 | 609.6382 | 465.1608 | 1648.5629 | 2000.6633 | 1632.1106 | 1119.0505 | 1686.794 | 1855.005 | 1433.0802 | 1278.1456 | 295.2965 | 164.6583 | 169.2714 | 295.2965 | 301.7286 | 363.588 | 279.5678 | 315.412 | 343.4774 | 344.0905 | 314.8894 | 259.4774 | 218.1357 | 206.0703 | 198.9297 | 227.5477 | 250.2714 | 227.3618 | 223.1156 | 165.4523 | 1122.2463 | 1067.5074 | 919.0954 | 483.206 | 713.8442 | 877.4724 | 785.5427 | 475.6332 | 1537.5979 | 1385.7488 | 1200.8742 | 882.1156
529 | 7.8134 | 900.3039 | 1137.9322 | 1011.5458 | 578.2891 | 839.6136 | 858.2478 | 757.1858 | 489.8555 | 506.5045 | 374.82 | 376.3451 | 353.6755 | 900.2213 | 739.5103 | 667.6194 | 619.2419 | 835.7168 | 672.7316 | 676.3097 | 535.4219 | 832.469 | 1410.5103 | 1274.8968 | 531.7109 | 1217.9439 | 1222.3039 | 1020.5812 | 787.4336 | 560.0413 | 671.1387 | 556.5103 | 456.5516 | 757.3658 | 843.0059 | 691.4897 | 237.0413 | 1595.4955 | 1662.8968 | 1351.8555 | 955.1387 | 757.6549 | 905.7522 | 716.5929 | 427.8465 | 263.9174 | 284.9587 | 251.4425 | 397.0826 | 996.5516 | 960.9676 | 911.9911 | 458.2271 | 794.5871 | 849.2536 | 707.0974 | 376.5458 | 878.3658 | 906 | 942.8496 | 513.7316 | 503.6136 | 522.1239 | 447.1652 | 312.3304 | 973.9794 | 975.5103 | 910.9676 | 643.5516 | 770.9587 | 916.3245 | 1150.5812 | 660.767 | 874.2065 | 1128.6755 | 924.3716 | 745.3716 | 848.5516 | 1029.3245 | 895.0884 | 723.0826 | 570.6696 | 622.9381 | 545.7522 | 465.2626 | 585.4897 | 556.469 | 480.0619 | 368.9794 | 622.5781 | 699.1858 | 601.8407 | 473.4897 | 1644.0767 | 1997.9646 | 1649.3393 | 1061.9735 | 1652.0884 | 1878.9794 | 1464.5162 | 1296.5162 | 270.1297 | 171.2832 | 164.469 | 270.1297 | 318.7729 | 360.2832 | 276.0413 | 325.0826 | 334.6961 | 326.3304 | 313.3304 | 260.9794 | 219.4071 | 216 | 192.9174 | 220.8613 | 253.7935 | 233.6903 | 244.2419 | 166.2626 | 1096.1858 | 1056.4601 | 929.2271 | 504.6903 | 745.1593 | 868.7109 | 781.3245 | 473.2419 | 1518.0943 | 1378.0413 | 1266.649 | 889.531
530 | 7.8244 | 891.9292 | 1154.9567 | 1056.0866 | 581.0669 | 792.3367 | 833.4508 | 744.4862 | 484.4508 | 531.502 | 381.0866 | 352.878 | 368.0905 | 916.5098 | 703.3603 | 654.5689 | 629.9921 | 822.9921 | 671.4429 | 677.5649 | 564 | 852.5728 | 1417.0275 | 1277.9921 | 539.4744 | 1201.0197 | 1240.5334 | 1069.0118 | 772.0236 | 553.4823 | 671.9331 | 550.4784 | 469.561 | 742.4154 | 805.4508 | 680.9567 | 236.5059 | 1562.9685 | 1653.9449 | 1337.9174 | 965.4705 | 768.0708 | 892.3997 | 714.5098 | 403.0669 | 275.5295 | 276.4587 | 269.9646 | 361.376 | 992.4941 | 1004.4705 | 861.4941 | 438.4666 | 755.7678 | 807.4823 | 694.3721 | 394.9843 | 878.4744 | 920.5571 | 973.4941 | 504.4115 | 489.4666 | 530.5571 | 447.5334 | 297.0039 | 950.4036 | 978.0118 | 934.3918 | 660.5767 | 752.4193 | 922.9921 | 1187.4666 | 675.4154 | 860.4862 | 1131.0433 | 899.9764 | 729.0079 | 851.0197 | 1055.0669 | 889.5925 | 738.5767 | 572.9449 | 614.9567 | 538.9252 | 473.4862 | 555.38 | 550.9725 | 490.5531 | 364.4784 | 612.0708 | 721.624 | 537.3131 | 484.5413 | 1661.4862 | 2004.1141 | 1670.9567 | 1096.0315 | 1597.3997 | 1856.4075 | 1440.0039 | 1235.3603 | 267.1102 | 184.0079 | 174.0315 | 267.1102 | 301.3879 | 360.4587 | 277.5138 | 313.4705 | 332.5295 | 306.4862 | 300.0118 | 268.0236 | 211.9528 | 214.4941 | 199.0079 | 204.5059 | 241.4115 | 246.9921 | 238.4272 | 176.4941 | 1056.38 | 1015.0236 | 944.1023 | 507.4508 | 753.9685 | 909.6043 | 777.9528 | 489.5138 | 1486.5885 | 1360.439 | 1262.4154 | 894.5256
531 | 7.8354 | 907.9411 | 1078.1749 | 1070.2861 | 591.3146 | 731.4582 | 822.0285 | 749.2281 | 455.5437 | 501.1445 | 385.5152 | 341.5418 | 392.5428 | 922.0856 | 695.2557 | 633.4306 | 629.5428 | 814.8289 | 665.7424 | 682.6863 | 554.7433 | 862.7719 | 1425.5428 | 1276 | 526.8289 | 1202.4002 | 1191.9173 | 1034.9734 | 803.7129 | 572.6559 | 652.943 | 566.0846 | 442.3165 | 705.058 | 779.6293 | 699.827 | 239.5428 | 1514.8878 | 1728.9667 | 1296.9439 | 936.9154 | 779.3146 | 885.5133 | 741.1702 | 415.3717 | 270.1435 | 260.8574 | 247.115 | 356.2272 | 991 | 974.8869 | 909.8831 | 434.1141 | 703.6854 | 833.3413 | 662 | 382.7719 | 872 | 930.3717 | 980.7424 | 481.4857 | 497.9705 | 567.1131 | 453.4287 | 307.7709 | 950.6844 | 977.9144 | 913.1711 | 653.7728 | 754.6274 | 935.3992 | 1159.9724 | 647.8289 | 858.0285 | 1096.231 | 902.7424 | 742.8279 | 853.4287 | 1054.5152 | 894.4867 | 734.8584 | 553.3432 | 619.942 | 530.2852 | 452.5152 | 561.5124 | 579.4838 | 491.1435 | 354.3717 | 632.5713 | 737.5722 | 555.3108 | 481.6293 | 1718.6825 | 2014.4867 | 1653.3146 | 1111.1996 | 1606.9696 | 1794.9449 | 1425.058 | 1218.5133 | 267.7443 | 182.4002 | 187.6568 | 267.7443 | 288.4278 | 344.3432 | 286.1426 | 299.8289 | 320.4582 | 329.2272 | 285.5152 | 268.3432 | 187.1435 | 198.6008 | 191.2291 | 210.1141 | 220.5428 | 241.4002 | 216.9144 | 178.5998 | 1053.2557 | 1034.885 | 968.9715 | 514.0276 | 737.7719 | 896.9163 | 789.6559 | 522.3127 | 1500.2576 | 1343.4572 | 1254.3707 | 903.5713
532 | 7.8464 | 933.1549 | 1054.4381 | 1041.938 | 583.9469 | 727.2168 | 861.9647 | 749.1018 | 490.0222 | 498.3275 | 365.8407 | 343.677 | 380.8363 | 931.8451 | 740.4823 | 629.1195 | 622.0487 | 827.4292 | 670.3142 | 703.593 | 550.7389 | 844.469 | 1457.0133 | 1270.7345 | 535.2168 | 1168.4601 | 1222.7745 | 998.9469 | 817.469 | 597.6327 | 636.2566 | 582.3142 | 425.6947 | 695.792 | 795.9602 | 713.7832 | 250.4779 | 1487.531 | 1761.3981 | 1298.5354 | 901.2035 | 761.4159 | 840.3406 | 784.5354 | 404.6814 | 253.7876 | 249.6814 | 240.8451 | 382.0531 | 986.7876 | 967.6903 | 976.4823 | 447.4779 | 719.4779 | 882.5354 | 671.4779 | 382.8982 | 853.0442 | 898.6194 | 967.4115 | 474.1549 | 500.8363 | 556.407 | 461.531 | 311.7345 | 942.407 | 999.7478 | 913.208 | 652.1682 | 746.5132 | 907.4026 | 1154.1106 | 638.8407 | 841.6239 | 1060.8982 | 910.4735 | 735.5708 | 834.1504 | 1056.4292 | 889.1106 | 746.1726 | 562.7434 | 640.2655 | 532.1018 | 437.5796 | 568.6194 | 596.1504 | 491.6372 | 372.1151 | 622.8894 | 706.1504 | 637.5354 | 438.9867 | 1699.6503 | 2011.2168 | 1652.792 | 1129.5841 | 1619.4646 | 1762.6858 | 1420.531 | 1223.3628 | 258.323 | 173.208 | 186.1549 | 258.323 | 325.6416 | 344.792 | 306.7965 | 291.3673 | 311.4779 | 357.6858 | 292.1682 | 277.7434 | 193.9558 | 194.4779 | 190.9513 | 205.0487 | 233.5841 | 239.0531 | 213.4735 | 183.0531 | 1044.2477 | 1053.792 | 996.9602 | 508.8318 | 730.5265 | 905.8098 | 824.1106 | 534.2035 | 1480.4159 | 1364.1151 | 1239.0929 | 911.2655
533 | 7.8574 | 902.7522 | 1094.3884 | 1033.7768 | 581.3835 | 747.0049 | 860.8977 | 735.5339 | 501.5919 | 510.6843 | 383.4759 | 356.4809 | 344.7522 | 943.9272 | 709.5043 | 650.9951 | 623.6991 | 834.3009 | 659.4562 | 703.4414 | 565.7768 | 811.8299 | 1497.471 | 1263.3058 | 559.3157 | 1146.3884 | 1279.0629 | 1051.8841 | 802.6116 | 578.4464 | 675.8064 | 548.2861 | 456.6991 | 730.0974 | 817.6893 | 698.8446 | 252.5339 | 1491.0728 | 1690.2133 | 1332.6264 | 890.233 | 752.0777 | 803.3983 | 787.3736 | 398.4612 | 270.4759 | 252.6215 | 255.6215 | 373.8397 | 990.5438 | 984.3107 | 975.1406 | 464.6215 | 698.9026 | 873.5191 | 698.3206 | 362.5191 | 853.7817 | 896.0925 | 964.7041 | 482.5487 | 503.0099 | 504.8299 | 448.3687 | 310.233 | 960.3403 | 992.9026 | 955.418 | 660.6067 | 729.6991 | 900.7139 | 1170.3884 | 663.4809 | 866.4131 | 1076.0826 | 909.4612 | 734.0875 | 849.1751 | 1052.5241 | 890.3009 | 787.3157 | 582.4661 | 616.4414 | 525.5388 | 446.5438 | 561.2429 | 588.5438 | 481.9852 | 384.4562 | 603.3785 | 692.0777 | 637.6794 | 438.4032 | 1662.7916 | 1986.2812 | 1708.651 | 1122.7571 | 1628.4809 | 1783.7817 | 1447.2429 | 1272.889 | 283.7768 | 156.1455 | 175.6893 | 283.7768 | 335.2232 | 339.762 | 311.8397 | 298.1603 | 335.6264 | 321.6646 | 307.6116 | 280.762 | 217.4661 | 200.8446 | 195.767 | 205.6215 | 241.8446 | 256.1652 | 224.3157 | 185.0777 | 1009.767 | 1022.8977 | 988.5142 | 505.7817 | 721.8397 | 948.0777 | 845.238 | 527.5438 | 1486.1652 | 1391.5438 | 1213.4612 | 890.1356
534 | 7.8684 | 847.8298 | 1102.8579 | 1043 | 568.9774 | 747.4887 | 863.619 | 740.4715 | 473.5511 | 501.5511 | 379.5059 | 381.6534 | 359.2724 | 965.534 | 696.6589 | 648.3068 | 639.1647 | 824.8977 | 654.6534 | 669.6308 | 577.2045 | 818.4543 | 1473.1764 | 1241.1593 | 548.6081 | 1166.9828 | 1236.6479 | 1070.4434 | 818.943 | 521.4208 | 667.5402 | 536.3294 | 454.5285 | 752.2443 | 789.1991 | 690.3864 | 258.5738 | 1493.1647 | 1690.4941 | 1367.4941 | 898.5113 | 733.1593 | 808.9774 | 751.9547 | 425.5566 | 267.057 | 251.1819 | 260.7955 | 353.8751 | 996.4489 | 1002.3294 | 938.6081 | 467.5909 | 704.8634 | 805.3583 | 695.2615 | 352.7783 | 856.324 | 917.2045 | 1008.7611 | 493.3466 | 509.4887 | 498.7783 | 457.6589 | 307.4887 | 969.6878 | 946.5059 | 993.7955 | 620.79 | 761.1475 | 959.5168 | 1179.4091 | 645.114 | 875.0968 | 1067.5683 | 906.7955 | 733.2217 | 861.7728 | 1036.6534 | 884.2045 | 768.3411 | 562.0968 | 585.3864 | 546.943 | 471.7385 | 552.8525 | 610.4317 | 470.6136 | 359.8127 | 599.8579 | 689.1421 | 634.1249 | 442.8127 | 1665.1991 | 1958.9602 | 1743.1819 | 1127.8407 | 1665.2272 | 1754.3583 | 1504.9883 | 1280.3583 | 259.381 | 168.5964 | 169.7955 | 259.381 | 328.7557 | 339.8181 | 285.8127 | 296.6308 | 342.3864 | 302.8806 | 296.3864 | 270.8977 | 204.2615 | 196.2443 | 195.2045 | 217.4091 | 225.6706 | 254.0172 | 229.0398 | 172.2217 | 1037.8634 | 1060.3403 | 959 | 501.7104 | 739.9032 | 923.6479 | 841.8751 | 528.4887 | 1529.7611 | 1385.8751 | 1206.3864 | 892.2498
535 | 7.8794 | 819.4366 | 1107.1268 | 1035.1121 | 596.0591 | 739.6195 | 884.2389 | 733.7876 | 491.4661 | 503.1268 | 368.7611 | 386.944 | 373.351 | 940.6872 | 714.3658 | 644.8023 | 625.8436 | 793.5752 | 670.0856 | 665.3805 | 549.702 | 790.9822 | 1453.5782 | 1263.3128 | 539.8879 | 1135.4896 | 1206.4218 | 1027.7168 | 828.1121 | 493.3805 | 669.6372 | 547.4366 | 445.4366 | 708.8701 | 765.6195 | 679.3658 | 261.8023 | 1464.0678 | 1702.3658 | 1368.0973 | 917.0856 | 746.8614 | 803.3805 | 766.1151 | 419.2654 | 256.7611 | 245.6903 | 269.1416 | 363.1564 | 972.3362 | 1012.8732 | 938.3393 | 456.1121 | 726.056 | 841.6815 | 713.8054 | 367.3805 | 850.5074 | 943.7906 | 1008.7728 | 490.8732 | 515.1416 | 531.4101 | 490.0148 | 305.2537 | 960.2124 | 953.2272 | 975.6607 | 611.649 | 764.7876 | 922.3332 | 1167.7876 | 608.7463 | 862.4513 | 1025.702 | 910.0708 | 737.2124 | 846.4926 | 1021.0973 | 935.5812 | 764.7198 | 531.9852 | 591.2685 | 558.9292 | 490.3805 | 543.5782 | 590.8849 | 507.5517 | 352.3245 | 583.2802 | 691.6342 | 630.4779 | 467.4396 | 1629.2242 | 1974.8319 | 1728.6755 | 1143.4366 | 1620.2211 | 1801.95 | 1544.5074 | 1263.7198 | 264.1151 | 180.2389 | 175.3245 | 264.1151 | 316.9144 | 331.2242 | 298.0443 | 294.6342 | 342.6342 | 322.8879 | 297.1977 | 276.1977 | 193.5074 | 190.6195 | 201.5074 | 217.6195 | 239.9174 | 224.6607 | 226.1268 | 168.3245 | 1048.9144 | 1077.6872 | 968.5782 | 495.5782 | 755.9292 | 934.5517 | 844.3953 | 514.4218 | 1528.646 | 1392.9027 | 1215.0856 | 920.4513
536 | 7.8904 | 832.2414 | 1090.1529 | 1023.2429 | 612.4857 | 718.0014 | 896.2714 | 788.2071 | 515.3628 | 516.6657 | 365.94 | 377.3343 | 367.1814 | 906.3943 | 695.0314 | 675.1185 | 609.8486 | 741.79 | 681.5757 | 703.1185 | 544.8471 | 786.3914 | 1467.9086 | 1310.2714 | 552.3329 | 1148.7842 | 1206.4843 | 1058.6328 | 812.2129 | 522.4828 | 660.0943 | 569.4528 | 445.5757 | 699.4528 | 780.2714 | 704.3614 | 244.91 | 1445.6357 | 1706.6357 | 1330.82 | 907.2743 | 789.5743 | 817.5143 | 788.1214 | 415.9986 | 260.8486 | 248.1514 | 267.2129 | 382.03 | 1007.5414 | 1003.3643 | 957.3329 | 468.9986 | 730.4843 | 877.2743 | 704.3057 | 375.1814 | 886.24 | 944.5772 | 975.6072 | 494.03 | 548.9071 | 557.2114 | 486.6372 | 297.7886 | 958.1829 | 977.1814 | 949.7586 | 650.6343 | 755.8786 | 906.3914 | 1145.6372 | 623.12 | 856.91 | 1025.4528 | 925.5143 | 727.6986 | 835.5157 | 1058.6628 | 929.67 | 760.4257 | 535.9386 | 652.39 | 545.2129 | 486.0914 | 544.0914 | 568.3029 | 538.3329 | 352.2429 | 551.3043 | 686.2129 | 603.3043 | 492.2429 | 1616.4243 | 1994.9386 | 1761.5728 | 1143 | 1568.3643 | 1824.9729 | 1487.55 | 1237.5486 | 273.4557 | 176.7271 | 176.3943 | 273.4557 | 306.5757 | 314.3943 | 300.0014 | 306.4843 | 323.9714 | 329 | 311.2114 | 275.8186 | 201.6057 | 193.7571 | 195.7886 | 208.0914 | 253.0914 | 227.5443 | 226.4243 | 165.3643 | 1029.94 | 1045.6971 | 954.6986 | 543.8757 | 753.1514 | 959.5757 | 832.6986 | 527.15 | 1488.2743 | 1368.5786 | 1220.8186 | 905.1229
537 | 7.9014 | 831.4159 | 1129.9204 | 1048.9883 | 571.6597 | 742.2203 | 854.9035 | 774.6715 | 518.884 | 528.94 | 383.5202 | 374.4681 | 383.9361 | 898.884 | 703.4042 | 701 | 611.352 | 744.5202 | 686.116 | 727.236 | 551.5319 | 799.8279 | 1441.8357 | 1333.5841 | 569.3481 | 1141.6636 | 1270.2125 | 1089.472 | 782.0678 | 550.8201 | 621.12 | 607.7562 | 453.0561 | 720.116 | 768.8918 | 727.176 | 254.2282 | 1450.236 | 1674.7797 | 1355.3364 | 920.2242 | 761.0756 | 800.7158 | 770.7119 | 400.5998 | 237.6558 | 236.6519 | 267.6441 | 386.176 | 1019.9518 | 1008.1721 | 945.5359 | 445.9557 | 718.9479 | 839.0678 | 666.2399 | 361.9479 | 866.4876 | 911.0078 | 1004.3963 | 495.824 | 552.7119 | 535.7758 | 496.0521 | 308.0521 | 940.3559 | 1000.9922 | 989.3924 | 682.1721 | 746.2399 | 932.176 | 1182.6884 | 640.292 | 852.116 | 1062.5802 | 945.1721 | 727.8761 | 831.176 | 1057.7797 | 950.5084 | 754.8201 | 531.0639 | 635.1395 | 541.528 | 487.4681 | 561.3442 | 545.8918 | 515.9518 | 345.06 | 577.8083 | 677.236 | 636.2764 | 466.2438 | 1598.3599 | 1998.296 | 1757.6597 | 1137.12 | 1570.8201 | 1792.236 | 1439.472 | 1199.3559 | 250.2399 | 165.0639 | 166.5319 | 250.2399 | 312.88 | 318.0561 | 278.3559 | 310.236 | 316.4081 | 331.352 | 300.3599 | 280.2321 | 212.0561 | 198 | 180.7679 | 215.3481 | 223.3038 | 234.884 | 207.1838 | 184.2842 | 1056.9283 | 1015.1878 | 966.6362 | 550.4798 | 740.4759 | 985.8722 | 848.7523 | 551.4081 | 1458.2999 | 1336.296 | 1265.2164 | 922.5163
538 | 7.9124 | 829.4027 | 1114.7743 | 1062.1962 | 514.1828 | 741.3864 | 826.0044 | 744.8068 | 499.792 | 518.9941 | 382.5914 | 360.9911 | 393 | 852.1784 | 715.7994 | 720.2095 | 603.3953 | 783.211 | 672.792 | 706.7905 | 554.8038 | 787.1976 | 1413.997 | 1334.3967 | 540.1814 | 1100.5944 | 1265.9778 | 1088.8009 | 798.416 | 555.7994 | 624.2036 | 576.1769 | 442.3923 | 689.3834 | 774.8112 | 723.7979 | 261.5973 | 1448.3997 | 1601.7772 | 1370.5929 | 893.7772 | 734.0015 | 772.795 | 767.2036 | 400.4101 | 234.0059 | 228.6003 | 280.8038 | 393.6033 | 1029.8142 | 1015.3997 | 959.2124 | 432.6062 | 703.3982 | 867.4219 | 658 | 371.0104 | 868.416 | 904.2036 | 1016.9926 | 496.8009 | 529.9941 | 520.8024 | 506.1991 | 325.4027 | 903.1843 | 1008.3967 | 985.5811 | 642.5766 | 761.4115 | 900.584 | 1181.584 | 627.792 | 825.5855 | 1010.3671 | 949.9985 | 765.413 | 847.6077 | 1046.6077 | 954.9837 | 718.379 | 531.0044 | 608.0059 | 530.1932 | 478.1932 | 580.0015 | 566.2183 | 495.1991 | 339.9985 | 613.4041 | 706.6151 | 666.7994 | 458.8038 | 1612.413 | 2054.4293 | 1728.3967 | 1132.3997 | 1569.7964 | 1819.2139 | 1492.2272 | 1218.6121 | 254.6062 | 163.8038 | 172.8053 | 254.6062 | 312.7979 | 335.0059 | 281.4041 | 315.6033 | 333.205 | 320.9941 | 301.8068 | 286 | 195.3893 | 216.0089 | 181.6033 | 204.7905 | 215.0044 | 232.6003 | 209.6077 | 192.9956 | 1040.3805 | 1019.6121 | 974.5944 | 525.5959 | 732.8009 | 975.7861 | 859.5944 | 500.9719 | 1440.7964 | 1351.6092 | 1267.7846 | 950.8024
539 | 7.9234 | 831.7748 | 1091.6392 | 1054.5497 | 558.9972 | 717.4867 | 852.2154 | 758.2881 | 487.1622 | 498.7483 | 355.6853 | 341.0364 | 391.7748 | 859.7552 | 707.0364 | 709.1091 | 608.6392 | 799.8378 | 679.1524 | 712.9902 | 566.1259 | 783.3874 | 1465.1328 | 1325.7119 | 535.828 | 1097.9007 | 1299.709 | 1094.2881 | 808.3608 | 522.5329 | 667.593 | 549.9007 | 455.2783 | 693.9538 | 783.8741 | 707.5231 | 259.2252 | 1421.0462 | 1621.2321 | 1392.8545 | 862.3874 | 721.2979 | 779.4769 | 797.1161 | 420.1259 | 260.3776 | 252.2783 | 276.1986 | 379.6224 | 986.5161 | 1029.0895 | 934.9567 | 444.6755 | 716.9273 | 833.291 | 686.179 | 391.7385 | 839.7679 | 924.9273 | 963.5063 | 469.2084 | 522.6126 | 528.2881 | 475.5958 | 297.7581 | 898.5398 | 985.0098 | 945.2979 | 620.8014 | 752.4965 | 924.3314 | 1204.1063 | 651.9175 | 820.7021 | 1017.4573 | 940.0364 | 776.2615 | 829.2084 | 1050.2615 | 940.351 | 709.6028 | 540.6755 | 611.7119 | 555.3049 | 456.7483 | 575.2615 | 592.8378 | 499.5133 | 349.0266 | 610.4238 | 700.0462 | 634.1455 | 481.765 | 1635.5133 | 2125.2419 | 1718.4867 | 1130.1622 | 1577.8643 | 1840.4503 | 1451.2014 | 1209.8839 | 266.0629 | 167.7748 | 183.0629 | 266.0629 | 316.1259 | 349.7385 | 309.6657 | 299.172 | 332.649 | 343.0168 | 303.0364 | 279.2615 | 190.8014 | 215.1357 | 195.8014 | 216.5035 | 219.1622 | 214.0098 | 224.9007 | 172.9105 | 1048.9175 | 1060.5035 | 985.3776 | 523.6755 | 748.0895 | 958.8378 | 865.4769 | 488.7287 | 1454.9902 | 1373.8014 | 1221.2713 | 956.4503
540 | 7.9344 | 811.6288 | 1104.6239 | 1031.7541 | 591.3771 | 751.4927 | 885.621 | 753.5015 | 507.2459 | 462.8805 | 360.9971 | 354.7463 | 379.7522 | 887.999 | 665.1322 | 682.7522 | 632.8717 | 775.38 | 700.6229 | 721.2507 | 558.1273 | 804.8707 | 1465.0068 | 1342.3712 | 560.8727 | 1103.4995 | 1301.8824 | 1084.5024 | 818.4966 | 523.8707 | 675.1283 | 544.2517 | 446.2546 | 681.6307 | 794.3722 | 730.8688 | 246.8776 | 1418.9971 | 1646.1263 | 1383.6307 | 874.4976 | 715.1244 | 820.4937 | 793.6288 | 423.3751 | 247.0049 | 238.8805 | 286.2468 | 372.3741 | 1002.6141 | 1050.4976 | 975.862 | 463.8717 | 713.2527 | 821.1195 | 682.1293 | 382.8776 | 834.2449 | 928.1263 | 984.1156 | 494.7415 | 513.002 | 524.7512 | 466.998 | 314.8678 | 884.0049 | 1006.7434 | 992.239 | 612.6278 | 748.2478 | 941.1273 | 1185.1341 | 648.38 | 831.8746 | 1047.5005 | 910.0049 | 760.7522 | 840.3693 | 1065.3712 | 906.8824 | 741.9961 | 558.622 | 623.3722 | 555.7541 | 453.3741 | 577.2488 | 618.3702 | 512.3732 | 359.7493 | 587.5034 | 676.7512 | 614 | 481.7522 | 1589.0098 | 2116.8805 | 1741.2449 | 1106.5044 | 1595.998 | 1802.0078 | 1477.6102 | 1224.619 | 274.8737 | 163.8756 | 185.6249 | 274.8737 | 318.7502 | 338.378 | 308.378 | 309.1205 | 339.2478 | 350.7522 | 298.6249 | 285.6229 | 192.001 | 216.3732 | 198.8756 | 217.0029 | 226.1234 | 230.1195 | 226.1254 | 167.1244 | 1009.7619 | 1095.9961 | 988.2517 | 522.8756 | 765.7483 | 964.3741 | 884.6229 | 541.8678 | 1470.1244 | 1392.4976 | 1212.4985 | 945.5024
541 | 7.9454 | 785.9823 | 1074.23 | 1080.1682 | 592.0089 | 754.8849 | 867.4159 | 773.2124 | 492.6017 | 475.3098 | 364.2655 | 392.7611 | 353.2478 | 891 | 676.1328 | 684.2832 | 638.5398 | 795.0443 | 665.584 | 701.1593 | 562.469 | 810.9911 | 1463.1151 | 1322.1415 | 555.8938 | 1113.2832 | 1276.4513 | 1096.7522 | 837.9204 | 526.8053 | 628.0442 | 540.2743 | 438.3717 | 691.3983 | 781.9734 | 754.4602 | 228.1681 | 1468.1417 | 1668.1239 | 1395.3098 | 874.354 | 713.177 | 834.2655 | 750.1415 | 437.0177 | 261.9469 | 229.0089 | 262.8672 | 373.7257 | 1018.7964 | 1052.9026 | 1005.8053 | 457.4336 | 711.3717 | 863.4868 | 702.1328 | 388.3806 | 869.1151 | 922.7257 | 1004.7168 | 510.1681 | 513.3717 | 517.177 | 501.6726 | 345.2832 | 878.5575 | 977.938 | 984.584 | 642.2301 | 765.1947 | 935.2743 | 1176.1151 | 661.0354 | 820.8938 | 1026.8849 | 952.3452 | 748.9026 | 833.7876 | 1038.1327 | 927.3275 | 736.1593 | 557.8053 | 614.7964 | 557.0177 | 462.3717 | 553.6017 | 620.2566 | 512.9026 | 346.0708 | 580.823 | 697.8496 | 617.1858 | 490.2921 | 1570.469 | 2042.0176 | 1721.9557 | 1110.2036 | 1601.3628 | 1786.2832 | 1505.1593 | 1208.5929 | 253.5132 | 162.6372 | 182.177 | 253.5132 | 303.3451 | 332.1858 | 302.1858 | 328.7345 | 367.0266 | 332.531 | 323.2124 | 277.9823 | 188.3628 | 200.9734 | 193.177 | 211.823 | 230.3628 | 237.4425 | 223.7257 | 176.646 | 997.8496 | 1027.0795 | 980.4513 | 524.823 | 737.23 | 980.3806 | 874.4336 | 525.7699 | 1459.8938 | 1419.7522 | 1259.0532 | 936.0885
542 | 7.9564 | 778 | 1085.5246 | 1063.0432 | 592.1032 | 720.6194 | 862.9892 | 773.3613 | 473.1032 | 476.413 | 376.587 | 387.2173 | 365.3806 | 889.0516 | 680.0624 | 671.4646 | 621.0624 | 766.9399 | 666.6795 | 715.6795 | 567.0516 | 762.8367 | 1472.0516 | 1330.6279 | 554.8451 | 1088.7743 | 1241.2281 | 1102.8043 | 811.1765 | 543.0816 | 608.9484 | 557.2365 | 441.3505 | 690.114 | 777.1441 | 742.0624 | 225.8968 | 1412.4153 | 1684.2989 | 1384.7227 | 875.8451 | 723.3397 | 826.4538 | 746.9376 | 424.8667 | 268.6086 | 242.0924 | 261.5354 | 365.2065 | 991.4646 | 1011.0324 | 1035.0708 | 464.2365 | 691.6194 | 839.6818 | 736.587 | 370.0216 | 833.2389 | 928.4946 | 989.6086 | 491.0108 | 507.9076 | 518.2473 | 496.5678 | 341.557 | 909.9784 | 986.3698 | 965.5462 | 630.8775 | 756.0624 | 939.2473 | 1183.7527 | 645.1765 | 802.3097 | 1078.3482 | 968.7635 | 743.4022 | 820 | 1034.1849 | 944.8559 | 739.6387 | 547.4538 | 598.5054 | 577.9892 | 475.7419 | 567.2257 | 616.2473 | 521.0408 | 384.4106 | 559.6194 | 677.5786 | 603.413 | 492.5054 | 1529.5894 | 2059.0084 | 1757.2041 | 1086.2281 | 1577.6194 | 1767.6194 | 1529.4214 | 1216.9268 | 250.0408 | 166.2473 | 172.8559 | 250.0408 | 311.886 | 332.7011 | 321.5354 | 326.8043 | 351.0732 | 339.6387 | 322.7119 | 265.4538 | 189.9484 | 218.8752 | 205.9376 | 221.1441 | 253.3806 | 227.4538 | 215.8559 | 159.5678 | 1019.7935 | 1025.8128 | 991.9892 | 509.4646 | 745.2773 | 992.5462 | 859.1549 | 512.443 | 1495.2149 | 1447.2365 | 1283 | 928.5054
543 | 7.9674 | 771.3825 | 1056.3392 | 1056.8378 | 597.2793 | 684.8387 | 819.0452 | 787.176 | 486.2203 | 497.1171 | 371.765 | 375.3382 | 393.2203 | 916.4553 | 687.2203 | 675.8968 | 633.5143 | 739.0147 | 683.9705 | 727.0147 | 558.7207 | 800.5429 | 1482.2498 | 1317.8682 | 559.3235 | 1054.471 | 1262.0433 | 1108.6028 | 812.529 | 548.3825 | 641.1023 | 592.47 | 468.7936 | 694.235 | 749.2359 | 731.353 | 233.9558 | 1375.294 | 1723.3816 | 1441.4691 | 873.706 | 692.2802 | 787.6037 | 820.5134 | 417.9705 | 266.9705 | 243.0295 | 276.294 | 377.5438 | 1021.0433 | 1001.2498 | 1013.6332 | 480.2793 | 727.3078 | 838.1908 | 727.1327 | 365.9263 | 842.4111 | 907.5152 | 1020.3963 | 503.8525 | 529.47 | 518.6765 | 466.471 | 330.3825 | 875.0304 | 1015.9263 | 964.6912 | 632.8525 | 721.53 | 918.5005 | 1207.1465 | 644.5438 | 825.1171 | 1101.412 | 980.5143 | 712.2212 | 804.118 | 1028.7797 | 895.3392 | 748.9853 | 537.059 | 637.3521 | 553.2359 | 501.5143 | 554.5447 | 598.8092 | 522.3677 | 370.9567 | 529.1327 | 662.3088 | 635.3668 | 473.118 | 1564.8811 | 2088.647 | 1789.647 | 1156.6747 | 1580.2203 | 1833.0866 | 1544.0295 | 1265.0876 | 252.0295 | 170.647 | 168.3382 | 252.0295 | 309.0885 | 328.6912 | 307.8387 | 334.588 | 311.5152 | 331.118 | 322.941 | 276.8968 | 191 | 214.4857 | 212.6765 | 219.0442 | 260.0442 | 239.5585 | 251.7051 | 186.3816 | 1019.3677 | 1079.7788 | 1047.3078 | 524.4848 | 755.3677 | 970.5152 | 889.7493 | 504.4267 | 1469.6922 | 1465.9263 | 1333.2931 | 907.7945
544 | 7.9784 | 773.3923 | 1102.4264 | 1120.5753 | 579.4306 | 704.68 | 851.6417 | 800.0221 | 491.3038 | 495.823 | 398.0281 | 369.6077 | 388.8672 | 881.7949 | 731.3983 | 721.4204 | 666.2434 | 793.2715 | 642.8613 | 705.1047 | 553.6519 | 829.9779 | 1504.8511 | 1334.3319 | 566.0664 | 1107.0119 | 1242.9055 | 1103.5634 | 813.9115 | 539.6077 | 623.2979 | 610.0443 | 462.1047 | 717.177 | 739.7404 | 750.2213 | 253.177 | 1401.5915 | 1760.5251 | 1464.5192 | 862.2375 | 693.2213 | 806.0723 | 848.6077 | 385.6238 | 266.3038 | 236.9557 | 275.6298 | 396.1106 | 1006.6238 | 1028.5694 | 981.5413 | 490.0664 | 696.0545 | 839.9115 | 664.7507 | 354.823 | 808.7728 | 949.5974 | 1047.7625 | 499.8451 | 529.5192 | 511.2596 | 452.9557 | 293.2979 | 844.6077 | 1031.1106 | 972.4366 | 657.8289 | 743.7242 | 948.1166 | 1232.8068 | 675.2434 | 864.2655 | 1067.0162 | 897.3081 | 746.8791 | 799.3702 | 1023.6962 | 911.7685 | 739.8894 | 520.1932 | 625.2979 | 568.0059 | 483.6902 | 560.8953 | 601.1328 | 558.4204 | 346.2817 | 573.9234 | 720.6196 | 662.7404 | 473.0885 | 1569.7124 | 2065.7345 | 1795.7183 | 1173.6902 | 1644.6196 | 1763.8052 | 1578.3983 | 1275.2596 | 268.8732 | 177.3923 | 178.7847 | 268.8732 | 303.674 | 326.326 | 291.2817 | 308.9941 | 308.4366 | 322.326 | 331.0443 | 270.5192 | 181.5634 | 198.6077 | 193.8009 | 206.5634 | 223.9719 | 234.8672 | 238.9719 | 194.5413 | 1027.1106 | 1091.6298 | 1026.1651 | 530.2596 | 735.4749 | 956.6519 | 882.7345 | 514.8511 | 1525.8851 | 1481.1106 | 1336.0826 | 907.7625
545 | 7.9894 | 775.3137 | 1094.1337 | 1134.3323 | 584.7857 | 711.764 | 880 | 794.823 | 498.236 | 500.354 | 412.6273 | 373.8633 | 396.7267 | 868.9813 | 753.118 | 717.351 | 611.7983 | 757.5437 | 661.121 | 676.469 | 563.295 | 800.174 | 1510.941 | 1383.3167 | 554.587 | 1066.3077 | 1220.5683 | 1134.003 | 759.8977 | 507.4877 | 648.18 | 589.351 | 440.3323 | 669.407 | 739.5683 | 779.2173 | 261 | 1397.2143 | 1742.233 | 1409.9567 | 868.295 | 719.472 | 834.9813 | 830.2143 | 407.4347 | 267.059 | 244.6087 | 274 | 381.7827 | 1032.5527 | 1056.1583 | 1007.5713 | 472.41 | 723.2203 | 833.941 | 676.062 | 362.0993 | 851.947 | 965.705 | 1043.764 | 485.4503 | 535.354 | 530.649 | 461.9813 | 286.6087 | 852.9813 | 990.0157 | 956.41 | 636.8603 | 716.3883 | 941.233 | 1223.9007 | 643.4473 | 849.233 | 1031.0373 | 899.357 | 760.705 | 831.1987 | 1037.354 | 972.18 | 711.6647 | 520.177 | 606.941 | 543.1927 | 490.59 | 570.3137 | 576.8013 | 506.994 | 368.0217 | 558.1337 | 696.525 | 659.823 | 479.059 | 1576.2173 | 2051.2547 | 1771.233 | 1155.1957 | 1685.0403 | 1770.0247 | 1567.8603 | 1235.6243 | 271.1367 | 177.941 | 170.9597 | 271.1367 | 287.528 | 335.6087 | 305.472 | 315.9627 | 310.941 | 305.528 | 333 | 274.295 | 196.9037 | 185.705 | 191.8633 | 218.472 | 227.2173 | 248.9037 | 220.941 | 191.059 | 1064.2577 | 1126.3757 | 1039.6893 | 532.4907 | 742.7857 | 966.9813 | 870.3137 | 553.885 | 1500.4847 | 1467.469 | 1377.8477 | 918.8633
546 | 8.0004 | 766.6165 | 1083.3835 | 1086.3839 | 572.4385 | 686.3426 | 884.1917 | 797.589 | 518.0684 | 497.6165 | 396.6303 | 375.6028 | 394.6165 | 874 | 710.288 | 702.8083 | 632.301 | 708.0413 | 648.2605 | 681.2739 | 545.644 | 792.589 | 1475.0688 | 1419.9587 | 559.8766 | 1088.4927 | 1281.8762 | 1101.0963 | 780.3147 | 500.9862 | 662.1096 | 575.5069 | 452.4656 | 670.5477 | 765.945 | 762.1509 | 246.3289 | 1420.0408 | 1741.8766 | 1404.767 | 910.0271 | 720.0138 | 819.7399 | 839.767 | 431.9862 | 247.7399 | 240.6165 | 263.5206 | 374.3972 | 1041.5206 | 1068.8904 | 1008.7261 | 464.3972 | 697.699 | 869.3285 | 736.3147 | 373.8904 | 820.1238 | 970.0821 | 1056.767 | 500.8629 | 535.411 | 563.356 | 475.3835 | 299.3835 | 841.233 | 1010.2188 | 925.3426 | 684.0817 | 705.7807 | 948.5615 | 1206.5206 | 670.8078 | 796.4798 | 1041.2601 | 937.767 | 756 | 807.2743 | 1036.0138 | 923.9321 | 761.0817 | 561.4243 | 599.0138 | 569.0133 | 462.2743 | 591.6574 | 590.2464 | 499.4519 | 375.5069 | 552.9725 | 688.3697 | 667.4794 | 456.9454 | 1522.8224 | 2050 | 1782.7532 | 1139.0275 | 1622.5348 | 1838.2188 | 1566.178 | 1274.89 | 270.7945 | 163.0275 | 172.6849 | 270.7945 | 305.1505 | 334.411 | 312.3014 | 312.7261 | 314.1917 | 313.3697 | 320.4248 | 272.0138 | 196.2193 | 197.767 | 203.3835 | 217.6165 | 231.1096 | 251.7124 | 235.3697 | 168.2468 | 1040.5757 | 1093.3977 | 1037.8358 | 528.0138 | 741.6165 | 963.6165 | 927.2872 | 547.0413 | 1501.0408 | 1393.3293 | 1346.11 | 910.1234
547 | 8.0114 | 733.1416 | 1102.6401 | 1057.3363 | 579.6401 | 725.3423 | 884.5782 | 772.2743 | 536.3746 | 486.1799 | 389.2891 | 361.4926 | 391.7109 | 890.3511 | 683.8908 | 726.593 | 620.587 | 723.1711 | 678.3658 | 689.1327 | 575.8113 | 787.1799 | 1469.9528 | 1417.6254 | 550.0708 | 1077.1651 | 1251.4158 | 1131.1859 | 812.5074 | 511.8201 | 630.1416 | 586.0856 | 429.4306 | 693.3746 | 782.2655 | 768.7729 | 256.3511 | 1359.7757 | 1709.3215 | 1449.6785 | 903.9616 | 717.7109 | 850.8113 | 800.0796 | 413.6725 | 269.5694 | 256.9056 | 258.2891 | 354.3835 | 1063.3039 | 1040.4306 | 1003 | 462.1563 | 727.8968 | 865.0944 | 668.9792 | 383.8201 | 834.944 | 1002.4366 | 1051.9144 | 504.0236 | 517.3599 | 543.1416 | 508.8584 | 300.1563 | 846.7965 | 1037.9528 | 936.6165 | 662.3687 | 751.2331 | 949.8908 | 1240.3894 | 665.407 | 815.9676 | 1074.1711 | 904.4778 | 733.9616 | 782.469 | 1031.5782 | 921.3039 | 804.7729 | 553.118 | 632.9676 | 558.8525 | 492.2095 | 612.3746 | 582.0944 | 540.5694 | 366.8908 | 558.2891 | 714.9056 | 703.9912 | 459.7965 | 1527.2803 | 2099.0532 | 1825.7021 | 1157.882 | 1618.593 | 1813.3451 | 1570.2891 | 1248.525 | 247.8289 | 172.6401 | 169.6018 | 247.8289 | 306.7581 | 319.9144 | 287.8289 | 304.1563 | 308.1799 | 337.062 | 332.7729 | 281.0856 | 187.7345 | 194.3363 | 205.5782 | 226.0856 | 228.2891 | 241.1799 | 227.7817 | 165.1092 | 1018.3127 | 1096.1711 | 1027.1563 | 520.0236 | 750.0856 | 996.9676 | 936.3599 | 567.1475 | 1521.8201 | 1396.9912 | 1326.9764 | 971.8261
548 | 8.0224 | 728.2321 | 1064.1612 | 1128.2135 | 564.3038 | 697.9916 | 891.2321 | 794.9907 | 504.8397 | 455.5182 | 394.0624 | 381.3122 | 380.4287 | 881.8127 | 685.3392 | 699.3933 | 620.7409 | 756.1426 | 713.357 | 691.4464 | 561.6253 | 752.1789 | 1506.5443 | 1381.9021 | 539.6608 | 1079.6337 | 1251.8658 | 1162.8481 | 780.3933 | 519.3392 | 632.4641 | 602.5713 | 401.8127 | 684.9823 | 795.5713 | 768.0447 | 251.4287 | 1392.919 | 1657.8397 | 1446.9198 | 907.2945 | 715.1072 | 848.1966 | 789.6785 | 416.5713 | 247.2861 | 238.9646 | 283.3122 | 364.8034 | 1056.643 | 1028 | 1036.9907 | 463.1696 | 751.8928 | 861.3392 | 709.2135 | 388.4464 | 834.643 | 982.1789 | 1037.5983 | 516.4641 | 534.8658 | 575.8388 | 515.2143 | 315.6337 | 889.6067 | 1061.5266 | 1003.5797 | 648.7857 | 700.0186 | 956.4017 | 1262.5089 | 680.3122 | 831 | 1087.616 | 935.0084 | 761.1603 | 770.3215 | 1049.8034 | 932 | 816.3392 | 538.6608 | 638.5983 | 564.357 | 494.3663 | 592.411 | 614.4996 | 519.0093 | 356.0447 | 568.8481 | 725.1696 | 661.3131 | 420.1612 | 1618.1063 | 2147.9283 | 1794.1612 | 1155.7055 | 1591.3933 | 1842.2852 | 1553.3663 | 1282.9637 | 238.7232 | 186.2321 | 171.3392 | 238.7232 | 291.4287 | 340.3122 | 323.562 | 339.8835 | 292.7055 | 326.643 | 324.0895 | 250.5629 | 188.8928 | 197.2321 | 202.1072 | 227.384 | 213.5359 | 241.616 | 237.1873 | 172.3392 | 1031.1873 | 1045.9739 | 1023.8304 | 483.2861 | 774.5266 | 1010.5536 | 959.6515 | 547.9021 | 1487.3933 | 1452.6692 | 1302.2414 | 951.2684
549 | 8.0334 | 729.529 | 1077.1554 | 1071.6823 | 589.8329 | 711.6264 | 890.3225 | 812.413 | 523.3619 | 472.2134 | 394.529 | 368.1996 | 392.9164 | 864.9675 | 731.1299 | 678.3805 | 645.8584 | 762.7935 | 689.1091 | 681.703 | 544.8515 | 765.0069 | 1474.1346 | 1429.5753 | 565.2134 | 1076.4385 | 1254.3805 | 1131.6961 | 816.0138 | 518.058 | 599.7541 | 612.884 | 441.6008 | 684.413 | 758.8121 | 745.8771 | 264.6519 | 1422.884 | 1785.951 | 1469.4199 | 881.8446 | 685.3156 | 774.0697 | 832.7169 | 426.1485 | 245.0325 | 240.0325 | 266.5221 | 383.768 | 1024.9931 | 1063.3039 | 1083.8329 | 483.8584 | 711.8121 | 878.4454 | 730.645 | 392.6775 | 802.9931 | 970 | 1074.4524 | 522 | 490.3087 | 557.6961 | 513 | 320.529 | 852.5152 | 1042.5221 | 1038.5035 | 642.9095 | 741.6658 | 924.6961 | 1234.8446 | 715.0069 | 819.9675 | 1074.29 | 965.5035 | 794.1229 | 783.1809 | 1076.1229 | 897.4316 | 770.1927 | 565.6844 | 630.587 | 584.9745 | 470.3481 | 612.4199 | 627.058 | 533.2134 | 353.7355 | 581.826 | 684.0766 | 669.4199 | 452.2783 | 1580.3342 | 2159.7355 | 1766.703 | 1181.8909 | 1546.1091 | 1857.7935 | 1545.529 | 1251.7797 | 279.4524 | 180.174 | 165.645 | 279.4524 | 306.8584 | 335.29 | 310.1091 | 328.2576 | 329.9234 | 350.1554 | 323.884 | 264.9489 | 190 | 204.413 | 186.29 | 222.3225 | 227.1229 | 246.6775 | 235.645 | 178.413 | 1039.2065 | 1071.5428 | 1065.6589 | 514.8654 | 810.6844 | 979.8446 | 904.8052 | 521.29 | 1542.1368 | 1432.9606 | 1368.4917 | 965.3619
550 | 8.0444 | 736.4779 | 1058.8362 | 1120.2656 | 610.9735 | 711.0354 | 864.323 | 837.9292 | 518.5487 | 492.469 | 379.792 | 361.4956 | 411.7124 | 840.8097 | 665.491 | 679.7389 | 664.2168 | 768.7301 | 669.0265 | 721.3717 | 555.7124 | 759.0442 | 1465.4779 | 1412.115 | 537.6106 | 1080.4779 | 1231.8097 | 1156.3894 | 842.9735 | 522.2345 | 630.1195 | 620.2345 | 463.9823 | 671.792 | 750.7301 | 767.416 | 264.2699 | 1417.5221 | 1751.7521 | 1474.7655 | 853.0531 | 680.2345 | 801.5974 | 814.0973 | 410.8009 | 248.2522 | 239.5133 | 271.9558 | 361.8274 | 1019.7389 | 1087.2168 | 1077.3053 | 490.2522 | 676.8097 | 823.4291 | 745.9469 | 389.5133 | 809.7036 | 1014.1195 | 1071.8009 | 533.9646 | 492.6859 | 557.7124 | 540.6682 | 317.0088 | 851.9469 | 1049.4513 | 1072.1638 | 640.7478 | 741.323 | 917.9823 | 1226.9912 | 688.8539 | 785.3407 | 1071.9912 | 1008.885 | 760.6194 | 797.2257 | 1026.9159 | 970.9957 | 744.7743 | 568.5221 | 620.7743 | 598.9735 | 475.2168 | 617.0177 | 622.2611 | 511.5929 | 365.9646 | 585 | 725.8319 | 680.7478 | 473.7389 | 1588.1549 | 2092.699 | 1772.7213 | 1157.8539 | 1605.2921 | 1846.531 | 1512.8451 | 1264.4071 | 246.1415 | 168.0265 | 174.9646 | 246.1415 | 316.2434 | 332.2434 | 300.4956 | 332.4602 | 322.5663 | 330.3407 | 323.0088 | 288.4602 | 188.5044 | 214.9735 | 195.208 | 199.3141 | 234.7478 | 240.5221 | 225.5221 | 161.3053 | 1031.7743 | 1036.6548 | 1063.8009 | 510.8097 | 800.8097 | 954.792 | 932.5974 | 533.9469 | 1584.1991 | 1440.1903 | 1384.7832 | 924.6548
551 | 8.0554 | 721.5182 | 1080.9233 | 1115.0767 | 579.0363 | 708.5202 | 860.5605 | 803.4355 | 512.2399 | 480.7984 | 417.5645 | 384.0423 | 403.8387 | 841.6008 | 676.004 | 684.8004 | 622.3952 | 773.2802 | 702.4838 | 732.1996 | 542.5182 | 755.5202 | 1464.1996 | 1434.7237 | 510.5585 | 1073.119 | 1231.0806 | 1181.1613 | 808.7561 | 546.0423 | 632.8387 | 580.1956 | 443.9577 | 693.6028 | 765.1613 | 780.8004 | 268.0806 | 1409.6794 | 1768.6048 | 1501.883 | 892.6048 | 665.2782 | 787.0767 | 804.5202 | 419.4415 | 265.4818 | 243.3206 | 263.3186 | 389.4838 | 1041.5222 | 1125.2036 | 1096.0826 | 460.3569 | 682.9214 | 803.7601 | 739.8387 | 400.9214 | 805.879 | 981.115 | 1032.5565 | 529.6391 | 532.6834 | 539.1976 | 493.7541 | 310.6794 | 843.5585 | 1080.0826 | 1027.2743 | 648.6008 | 747.4415 | 936.7218 | 1250.0423 | 684.6008 | 799.3225 | 1109.4838 | 958.9134 | 778.1633 | 796.9597 | 1026.2419 | 982.5182 | 773.9233 | 574.3609 | 601.2782 | 581.4778 | 494.2016 | 606.6391 | 617.1996 | 538.2439 | 403.4435 | 543.1956 | 706.996 | 670.3589 | 461.3186 | 1545.7944 | 2117.885 | 1852.768 | 1144.4798 | 1633.5605 | 1852.881 | 1529.3629 | 1284.121 | 269.5241 | 173.3609 | 178.2399 | 269.5241 | 322.3206 | 340.6008 | 307.8407 | 315.7178 | 323.3609 | 366.3649 | 326.5605 | 274.7581 | 194.8407 | 212.6794 | 213.6814 | 206.4415 | 252.4818 | 212.9174 | 232.121 | 172.4818 | 1021.3992 | 1085.1269 | 1078.5222 | 561.006 | 774.2379 | 984.9637 | 931.7581 | 536.1996 | 1537.5545 | 1464.4818 | 1381.7601 | 947.2842
552 | 8.0664 | 719.0895 | 1097.179 | 1125.854 | 546.3736 | 706.6829 | 871.2684 | 800.0408 | 495.7802 | 472.1381 | 403.967 | 388.6829 | 415.4474 | 831.6421 | 681.0487 | 682.1381 | 591.0078 | 766.2763 | 728.675 | 673.0722 | 591.8383 | 797.708 | 1544.0989 | 1448.3171 | 521.4474 | 1070 | 1307.1476 | 1166.4631 | 803.9513 | 547.5934 | 650.6264 | 575.4961 | 457.0816 | 634.0314 | 766.6829 | 788.9513 | 280.0408 | 1355.4788 | 1785.3171 | 1466.9749 | 868.3815 | 672.3579 | 780.0895 | 840.5291 | 425.5447 | 272.5447 | 253.4961 | 271.3579 | 363.9262 | 1126.0094 | 1169.2119 | 1063.8367 | 476.4882 | 710.9435 | 831.8053 | 705.8775 | 406.5447 | 809.179 | 1021.6107 | 1004.2355 | 489.9262 | 507.4709 | 530.4553 | 494.5369 | 325.2198 | 805.0157 | 1058.6499 | 1006.1381 | 664.1303 | 722.6499 | 938.1381 | 1234.6012 | 630.6169 | 827.0816 | 1092.4223 | 1006.9686 | 736.2512 | 811.4474 | 1069.0738 | 949.1947 | 804.854 | 556.146 | 653.1554 | 599.7159 | 478.9184 | 580.0565 | 641.4882 | 566.9922 | 403.9592 | 549.3093 | 735.9356 | 669.3171 | 494.8462 | 1546.7645 | 2105.967 | 1796.6734 | 1181.0738 | 1567.8037 | 1827.4223 | 1583.5699 | 1236.0235 | 296.675 | 190.675 | 178 | 296.675 | 317.0487 | 351.4961 | 318.4961 | 319.0408 | 331.4066 | 353.1947 | 315.6421 | 290.6264 | 194.6829 | 215.6342 | 215.6829 | 225.675 | 258.7724 | 234.3501 | 223.4144 | 167.1868 | 1025.9513 | 1102.9105 | 1070.325 | 551.1947 | 783.4474 | 1038.4804 | 968.4804 | 576.708 | 1541.854 | 1458.4144 | 1402.0816 | 933.7394
553 | 8.0774 | 672.1356 | 1093.5074 | 1144.5546 | 576.8791 | 693.4454 | 899.8938 | 831.2478 | 483.938 | 481.2006 | 388.938 | 415.4632 | 402.7374 | 849.1859 | 677.4307 | 696.6932 | 615.8171 | 753.0147 | 691.4129 | 699.1564 | 563.2743 | 785.6755 | 1515.4277 | 1415.2595 | 547.1859 | 1047.2448 | 1257.5958 | 1169.6313 | 854.649 | 538.938 | 659.354 | 624.295 | 445.7374 | 626.5546 | 779.3393 | 776.0914 | 256.3215 | 1356.4779 | 1740.4896 | 1426.5368 | 848.3687 | 726.2184 | 825.7257 | 828.2448 | 428.354 | 264.3687 | 258.354 | 260.8761 | 380.4632 | 1047.7786 | 1146.8289 | 1032.3835 | 490.2773 | 703.8761 | 876.0944 | 737.0178 | 400.7227 | 792.1681 | 1016.8141 | 1037.4484 | 492.3393 | 547.2184 | 565.3098 | 509.4159 | 319.7994 | 817.7552 | 1031.9528 | 1055.2184 | 676.6313 | 765.003 | 921.3068 | 1239.7699 | 683.8349 | 794.5516 | 1072.2301 | 979.1356 | 751.6018 | 804.2301 | 1011.7344 | 944.9233 | 783.5368 | 535.0914 | 644.1062 | 631.3245 | 487.1239 | 586.3393 | 674.8938 | 547.6755 | 409.6313 | 567.5546 | 722.1062 | 637.8289 | 498.1534 | 1561.4159 | 2109.7699 | 1852.2508 | 1217.8938 | 1580.9558 | 1808.0147 | 1575.8141 | 1275.1417 | 281.3835 | 184.0147 | 155.2448 | 281.3835 | 290.6755 | 324.1829 | 303.7374 | 328.2773 | 344.7699 | 319.8908 | 316.708 | 301.708 | 199.4926 | 184.8289 | 193.0295 | 235.4926 | 223.6902 | 243.7847 | 212.938 | 171.062 | 1046.8319 | 1087.0914 | 1115.4337 | 503.7669 | 784.8614 | 1021.1829 | 962.1681 | 576.4454 | 1504.0589 | 1458.1386 | 1371.9056 | 936.2006
554 | 8.0884 | 670.1573 | 1135.8324 | 1121.7006 | 622.0659 | 697.9695 | 916.5634 | 840.5939 | 509.8933 | 513.4873 | 375.8427 | 371.995 | 387.6396 | 847.0305 | 725.6141 | 721.7208 | 651.8933 | 713.3402 | 715.8629 | 681.5127 | 550.203 | 776.6091 | 1525.7055 | 1430.7055 | 533.0762 | 1043.3909 | 1303.3501 | 1113.0255 | 862.4213 | 514.6853 | 622.5432 | 610.7006 | 433.0305 | 672.2385 | 779.0152 | 783.9543 | 249.797 | 1384.1116 | 1805.3044 | 1502.68 | 889.0355 | 701.7463 | 797.3554 | 899.3044 | 424.2183 | 236.4975 | 238.2792 | 241.0609 | 377.6396 | 1054.269 | 1087.792 | 1027.797 | 503.1573 | 711.9543 | 873.2488 | 763.9391 | 371.9036 | 815.6903 | 1014.1878 | 1087.0507 | 491.2183 | 544.264 | 544.7158 | 508.8122 | 324.1726 | 883.7715 | 1082.7867 | 1012.4163 | 718.8476 | 698.5078 | 887.5127 | 1250.9695 | 696.6244 | 815.0812 | 1059.4366 | 1004.2537 | 767.1726 | 779.4823 | 1051.9745 | 932.4518 | 799.9086 | 557.2994 | 641.7817 | 629.2335 | 491 | 606.736 | 649.3249 | 550.5634 | 362.5889 | 536.731 | 743.6903 | 660.8781 | 456.9493 | 1550.4518 | 2136.9086 | 1894.7208 | 1181.9645 | 1595.5787 | 1798.6244 | 1597.0964 | 1371.2892 | 264.0457 | 157.8884 | 160.1573 | 264.0457 | 294.3604 | 356.6446 | 309.3604 | 325.2183 | 320.9036 | 347.269 | 328.3604 | 293.4366 | 202.5939 | 176 | 193.3756 | 237 | 238.7055 | 216.9036 | 213.3909 | 161.0457 | 1001.792 | 1094.9543 | 1094.731 | 520.0964 | 783.203 | 1003.4366 | 964.1726 | 610.4568 | 1504.7512 | 1513.9897 | 1389.0964 | 935.0152
555 | 8.0994 | 711.8437 | 1054.7462 | 1169.2193 | 622.9076 | 709.7109 | 891.4857 | 835.3353 | 550.9882 | 523.4277 | 368.9538 | 380.8496 | 416.5605 | 913.7857 | 707.2487 | 711.6244 | 646.0521 | 737.9882 | 678.8731 | 712.8437 | 566.9941 | 755.7689 | 1562.7513 | 1471.7975 | 503.7227 | 1099.8378 | 1278.7286 | 1182.9706 | 851.2891 | 500.9076 | 675.3118 | 597.526 | 491.6933 | 658.7227 | 739.1563 | 727.9252 | 245.9538 | 1356.821 | 1786.1563 | 1484.3471 | 886.2429 | 687.9538 | 846.0748 | 808.9428 | 441.6126 | 257.5143 | 246.7571 | 257.2311 | 387.9479 | 1072.7109 | 1090.3756 | 1070.8899 | 500.3353 | 706.9076 | 860.2891 | 729.9655 | 430.5488 | 856.179 | 976.9655 | 1074.3412 | 515.8958 | 497.1101 | 569.3698 | 506.3815 | 326 | 849.6361 | 1043.5899 | 1069.2135 | 689.3471 | 716.0345 | 952.0227 | 1214.1563 | 694.1908 | 752.5958 | 1056.1908 | 1000.2429 | 798.1328 | 817.6992 | 1109.8899 | 984.0286 | 769.584 | 547.8151 | 632.4798 | 667.4622 | 496.6647 | 609.3815 | 666.0866 | 597.5084 | 396.9361 | 570.0807 | 716.2025 | 693.2773 | 462.3756 | 1541.3353 | 2149.474 | 1862.7748 | 1197.705 | 1588.9076 | 1793.763 | 1653.3639 | 1268.1857 | 281.2311 | 166.5664 | 171.9017 | 281.2311 | 314.8034 | 334.6302 | 324.9479 | 353.1328 | 335.0403 | 346.2891 | 347.9941 | 310.4219 | 194.9076 | 201.8958 | 197.4277 | 245.9017 | 217.4857 | 215.6647 | 228.5664 | 159.6185 | 1028.6529 | 1139.2714 | 1093.2832 | 541.5664 | 797.5664 | 993.7168 | 998.3698 | 544.7403 | 1525.8034 | 1510.1966 | 1404.9017 | 957.4681
556 | 8.1104 | 740.7168 | 1072.4337 | 1163.9262 | 603.7478 | 719.3938 | 890.7507 | 858.646 | 515.4586 | 507.5693 | 390.1814 | 398.323 | 408.3908 | 879.8864 | 709.0369 | 692.3908 | 643.8215 | 756.68 | 635.1386 | 730.2153 | 598.9322 | 728.8186 | 1597.9322 | 1456.997 | 521.003 | 1081.7817 | 1247.6401 | 1260.5074 | 823.5324 | 527.7537 | 632.4926 | 615.7168 | 460.8156 | 657.1077 | 757.1107 | 797.7966 | 258.1446 | 1356.3938 | 1803.2891 | 1532.5074 | 889.5723 | 723.9691 | 827.9601 | 888.978 | 448.4646 | 262.3569 | 230.2832 | 242.2832 | 381.1416 | 1125.1136 | 1141.649 | 1103.1814 | 520.3599 | 699.2493 | 903.1844 | 726.7507 | 359.7388 | 853.32 | 1014.8275 | 1088.7168 | 531.0369 | 513.2891 | 581.1077 | 491.2124 | 314.4985 | 810.0678 | 1084.3998 | 1067.9262 | 716.1475 | 719.2493 | 957.4985 | 1304.4057 | 700.5723 | 822.2612 | 1066.68 | 971.5324 | 768.8525 | 810.7478 | 1050.9911 | 974.8186 | 756.8923 | 567.8245 | 600.4247 | 643.3171 | 510.323 | 624.6092 | 675.2861 | 560.351 | 401.4277 | 551.2463 | 708 | 694.0708 | 472.5723 | 1553.1446 | 2176.4676 | 1874.5384 | 1181.8186 | 1593.5723 | 1741.2433 | 1625.5663 | 1322.7966 | 275.32 | 175.7507 | 176.4646 | 275.32 | 293.5324 | 334.3938 | 348.5384 | 318.1017 | 303.0309 | 331.677 | 348.7139 | 296.9262 | 178.2124 | 206.3569 | 194.7139 | 220.0678 | 236.4676 | 217 | 240.2153 | 166.5723 | 1042.9292 | 1079.1695 | 1112.2522 | 540.0708 | 780.4616 | 1008.4307 | 993.677 | 584.5074 | 1565.326 | 1518.323 | 1411.1077 | 935.8893
557 | 8.1214 | 701.641 | 1093.5074 | 1153.3294 | 605.003 | 686.8131 | 923.6854 | 871.5044 | 548.5251 | 465.644 | 404.1721 | 382.6558 | 409.1691 | 866.4985 | 724.3412 | 699.8397 | 624.822 | 765.6706 | 678.0296 | 737.8368 | 610.8368 | 757.353 | 1522.2851 | 1472.0118 | 558.5192 | 1052.4867 | 1237.1632 | 1219.6351 | 807.9941 | 503.1484 | 640.8456 | 586.647 | 466.0089 | 677.178 | 762.1662 | 807.6617 | 249.3264 | 1384.6824 | 1744.7954 | 1528.3235 | 888.4985 | 683.6381 | 840.178 | 876.9793 | 403.1396 | 248.6588 | 256.0177 | 217.9882 | 418.1898 | 1112.6529 | 1096.9675 | 1064.641 | 551.6824 | 673.819 | 886.3176 | 763.0207 | 385.1928 | 805.9734 | 986.644 | 1083.8279 | 521.3264 | 502.3235 | 568.6588 | 499.6736 | 322.8397 | 836.5192 | 1058.4778 | 1098.1869 | 683.1425 | 748.0177 | 987.5192 | 1310.1573 | 712.8397 | 757.6174 | 1069.8338 | 954.3264 | 790.1839 | 831.181 | 1056.0118 | 1034.2046 | 791.8545 | 583.0059 | 658.2046 | 602.647 | 515.5015 | 576.3028 | 670.997 | 547.4985 | 428.3501 | 550.8368 | 697.9941 | 716.3471 | 435.644 | 1534.3205 | 2245.3708 | 1877.3323 | 1209.5192 | 1599.1691 | 1811.7149 | 1596.1544 | 1290.9704 | 271.3323 | 182.8368 | 190.3412 | 271.3323 | 304.6765 | 334.3323 | 322.9823 | 334.8486 | 361.7061 | 343.175 | 334.6588 | 293 | 188.3412 | 196.8279 | 208.175 | 209.8309 | 227.825 | 242.0148 | 239.4985 | 180.5074 | 1054.0059 | 1109.0266 | 1100.1573 | 523.9911 | 831.0325 | 1055.359 | 960.1484 | 612.8427 | 1580.5044 | 1512.6617 | 1425.3412 | 955.0148
558 | 8.1324 | 675.2315 | 1091.7694 | 1223.9199 | 588.2315 | 671.5388 | 934.2306 | 844.232 | 535.8471 | 540.0734 | 373.0014 | 375.6155 | 400.6927 | 901.5369 | 684.6941 | 742.6136 | 645.5374 | 779.6918 | 664.3087 | 706.0014 | 590.8471 | 778.384 | 1475.5398 | 1499.6913 | 554.8466 | 1052.2306 | 1244.4612 | 1188.6932 | 796.6927 | 529.1524 | 607.7709 | 633.3054 | 463.0772 | 677.6155 | 772.1534 | 784.8471 | 266.4607 | 1366.3087 | 1753.999 | 1574.9209 | 873.616 | 657.0777 | 842.3078 | 889.4607 | 415.1529 | 276.4602 | 238.3087 | 274.9204 | 409.9238 | 1108.8461 | 1105.4607 | 1121.151 | 540.0777 | 701.1524 | 922.4597 | 761.5388 | 384.6927 | 840.152 | 1006.9219 | 1100.6146 | 539.3068 | 530.3063 | 566 | 509.6151 | 331.7689 | 818.4626 | 1083.1524 | 1143.0753 | 718.152 | 743.8466 | 975.3854 | 1291.7699 | 698.9238 | 795.1515 | 1073.3845 | 973.999 | 785.8466 | 857.1529 | 1025.3092 | 1023.3087 | 794.7694 | 578.2311 | 640.5398 | 593.4616 | 541.3835 | 631.151 | 693.6913 | 564.7685 | 395.9247 | 565.5379 | 744.2287 | 682.9247 | 477.9209 | 1595.151 | 2170.8499 | 1861.7699 | 1216.8461 | 1567.8476 | 1827.1539 | 1641.0748 | 1300.4607 | 276.0767 | 146.7709 | 183.6927 | 276.0767 | 311.3845 | 342.4612 | 319.5384 | 311.232 | 360.616 | 322.3087 | 344.6918 | 277.7699 | 195.2306 | 193.3078 | 189.001 | 237.768 | 258.8447 | 241.0772 | 234.7694 | 166.9238 | 960.3888 | 1123.0767 | 1135.0753 | 526.0767 | 823.3854 | 1074.1534 | 961.6151 | 597.3854 | 1567.616 | 1472.9247 | 1438.9995 | 956.6155
559 | 8.1434 | 677.1504 | 1128.77 | 1197.5132 | 619.3363 | 666.5664 | 874.9114 | 864.7522 | 532.8584 | 489.761 | 397.0443 | 383.5841 | 429.0443 | 864.2212 | 674.4248 | 715.6637 | 664.5929 | 731.3539 | 671.1593 | 700.8584 | 582.708 | 768.1239 | 1585.8851 | 1476.5309 | 545.2743 | 1065.0177 | 1249.4336 | 1185 | 830.1947 | 528.9911 | 634.4779 | 629.2655 | 485.177 | 677.8584 | 780.8673 | 838.5133 | 264.8496 | 1353.4159 | 1782.8938 | 1577.9911 | 895.8938 | 649.708 | 799.0796 | 880.1239 | 441.3186 | 264.8318 | 235.7168 | 246.5132 | 413.0089 | 1118.4425 | 1121.8761 | 1136.4336 | 516.3982 | 666.6548 | 886.2212 | 787.4691 | 337.5044 | 841.1327 | 1036.0354 | 1064.5132 | 563.6018 | 510.2478 | 582.3098 | 541.9027 | 324.4159 | 805.4159 | 1092.4336 | 1095.0619 | 725.1416 | 788.354 | 1004.6195 | 1306.1682 | 651.3628 | 787.8318 | 1081.7257 | 1025.2124 | 797.7345 | 781.168 | 1077.3718 | 1024.1504 | 829.1947 | 558.115 | 642.7257 | 605.0177 | 525.3982 | 572.6105 | 689.4159 | 607.4868 | 404.4513 | 534.5221 | 794.2036 | 678.5752 | 482.708 | 1593.2655 | 2166.1593 | 1946.5576 | 1208.4159 | 1622.947 | 1915.416 | 1674.0354 | 1293.6991 | 261.5487 | 194.0797 | 170.8407 | 261.5487 | 308.5664 | 347.4336 | 341.4602 | 313.7257 | 358.8584 | 354.0531 | 320.3893 | 311.9115 | 203.7257 | 205.8761 | 209.8938 | 236.1327 | 222.938 | 224.5487 | 227.9911 | 185.4602 | 1053.7346 | 1126.5752 | 1126.5487 | 551.0354 | 777.0796 | 1027.0707 | 980.1682 | 592.2832 | 1543.5398 | 1579.3099 | 1456.4513 | 975.7434
560 | 8.1544 | 696.2847 | 1140.2242 | 1233.6643 | 599.7497 | 696.4745 | 891.121 | 826.3357 | 505.1376 | 487.707 | 398.5172 | 392.8363 | 406.1376 | 862.2242 | 676.6121 | 775.3024 | 672.2242 | 736.9312 | 670.3879 | 699.2586 | 573.293 | 789.5089 | 1566.6892 | 1506.828 | 586.6643 | 1128.8196 | 1241.293 | 1209.3796 | 882.1471 | 505.3618 | 621.7153 | 640.0605 | 458.5255 | 679.7414 | 797.6726 | 799.2408 | 264.8707 | 1345.9051 | 1780.9051 | 1559.586 | 885.1981 | 656.8363 | 751.9478 | 887.5777 | 445.8707 | 259.3879 | 255.1554 | 243.4828 | 419.2242 | 1154.828 | 1153.6038 | 1166.6038 | 512.1293 | 682.6382 | 912.0866 | 759.7841 | 391.8196 | 840.8707 | 1009.5255 | 1078.8968 | 518.2408 | 485.1032 | 591.9656 | 505.2064 | 335.1898 | 832.7331 | 1064.2669 | 1075.5516 | 714.5516 | 735.9218 | 1004.7758 | 1355.1471 | 656.1898 | 815.4745 | 1071.6809 | 1006.879 | 766.0427 | 842.0094 | 1040.8529 | 1001.4911 | 860.2503 | 510.5943 | 642.2586 | 640.0866 | 535.9312 | 620.2075 | 663.6204 | 600.0688 | 393.9395 | 580.3712 | 768.7841 | 715.5694 | 491.5777 | 1562.2669 | 2205.4401 | 1939.2325 | 1225.2847 | 1616.4567 | 1857.732 | 1657.9739 | 1367.7508 | 293.828 | 166.4306 | 176.8363 | 293.828 | 277.5255 | 337.5516 | 336.293 | 345.4745 | 349.4223 | 336.4911 | 345.6038 | 310.1637 | 197.1637 | 211.4828 | 198.3274 | 258.5089 | 245.6038 | 264.6643 | 236.5777 | 190.7414 | 1011.9218 | 1092.172 | 1175.3712 | 549.7758 | 777.8363 | 1038.1554 | 1014.3452 | 562.3962 | 1551.3191 | 1503.0938 | 1486.8624 | 1000.7675
561 | 8.1654 | 663.6804 | 1093.3185 | 1164.0628 | 567.7443 | 664.7974 | 880.8722 | 872.9794 | 551.3304 | 473.9892 | 391.8191 | 361.3294 | 423.1918 | 906.2665 | 683.1809 | 739.0845 | 663.2871 | 791.9794 | 760.065 | 672.5103 | 557.8722 | 762.9783 | 1577.7768 | 1437.8288 | 583.2871 | 1069.1268 | 1292.9794 | 1205.0531 | 851.9144 | 539.9686 | 673.7454 | 582.8397 | 466.3619 | 642.9144 | 751.4355 | 763.7443 | 263.234 | 1354.7129 | 1800.3088 | 1603.7985 | 918.3196 | 684.4897 | 817.5222 | 888.117 | 430.9892 | 262.532 | 243.8722 | 258.1278 | 397.9252 | 1096.4247 | 1206.5645 | 1187.7768 | 507.585 | 677.1701 | 863.1376 | 746.1701 | 402.766 | 835.7021 | 1018.2449 | 1172.065 | 507.468 | 548.2243 | 565.6273 | 537.8516 | 329.9361 | 804.3294 | 1105.0325 | 1088.1278 | 720.9469 | 739.3619 | 979.2763 | 1334.6273 | 718.0433 | 848.2557 | 1056.7551 | 960.6165 | 811.3304 | 778.8288 | 1061.3727 | 1010.3619 | 856.9361 | 531.3727 | 670.2557 | 661.7768 | 487.6696 | 627.234 | 670.5959 | 551.2015 | 404.3619 | 574.7551 | 739.2763 | 690.9783 | 474.4572 | 1547.4041 | 2156.2546 | 1994.2774 | 1249.1918 | 1628.1278 | 1875.2557 | 1739.7671 | 1366.6381 | 316.6598 | 204.2665 | 162.9892 | 316.6598 | 318.9155 | 336.883 | 345.5959 | 307.6165 | 344.468 | 355.9577 | 335.8722 | 321.3619 | 201.2979 | 217.2979 | 186.2871 | 230.2124 | 255.2979 | 230.3825 | 244.1809 | 191 | 1041.8516 | 1145.2774 | 1155.6273 | 564.8938 | 839.9263 | 1063.0748 | 1009.2871 | 557.117 | 1599.7985 | 1514.6067 | 1490.117 | 1003.117
562 | 8.1764 | 669.7434 | 1087.8953 | 1205.031 | 602.4971 | 696.7065 | 878.1047 | 847.7699 | 551.733 | 497.0679 | 417.8584 | 399.9735 | 408.0944 | 888.7227 | 686.6858 | 753.5915 | 672.7434 | 755.1312 | 705.7492 | 697.649 | 581.9632 | 736.618 | 1569.2566 | 1512.1564 | 609.7537 | 1097.6018 | 1341.1829 | 1159.2359 | 821.9321 | 487.702 | 682.7906 | 619.7641 | 468 | 696.1933 | 778.1254 | 782.382 | 236.1416 | 1328.2463 | 1752.8642 | 1604.6283 | 939.1151 | 699.6387 | 786.7123 | 866.5132 | 474.6593 | 249.5708 | 218.7227 | 251.0472 | 415.5915 | 1039.6548 | 1175.3982 | 1213.2773 | 519.5339 | 693.0103 | 900.7641 | 763.8009 | 411.9528 | 859.1726 | 1014.6283 | 1139.2359 | 552.6593 | 507.7595 | 563.7906 | 553.7434 | 345.1151 | 872.5178 | 1145.916 | 1157.1461 | 701.4085 | 748.1623 | 983.1623 | 1357.8584 | 722.4189 | 818.8746 | 1060.3717 | 987.2301 | 777.7123 | 863.8998 | 1083.8009 | 971.7123 | 762.8907 | 553.8009 | 683.8481 | 607.5973 | 481 | 607.3038 | 687.2198 | 578.1254 | 402.4189 | 574.7906 | 718.9897 | 704.0103 | 512.2877 | 1682.978 | 2177.649 | 1993.9425 | 1244.8377 | 1610.3038 | 1903.1726 | 1687.435 | 1305.9114 | 282.2935 | 178.6652 | 163.6858 | 282.2935 | 310.6755 | 350.4292 | 317.4557 | 309.1623 | 351.1623 | 353.6283 | 337.5811 | 306.8849 | 233.3348 | 185.7699 | 213.649 | 238.5339 | 239.8849 | 238.4292 | 236.9425 | 197.267 | 1028.1991 | 1164.6387 | 1177.0679 | 532.084 | 782.6445 | 1081.2198 | 1038.4396 | 577.5915 | 1533.4822 | 1593.2036 | 1481.9425 | 1012.8481
563 | 8.1874 | 657.3864 | 1087.0924 | 1193.7561 | 601.5546 | 675.588 | 886.1682 | 873.0423 | 526.4955 | 484.5713 | 410.1091 | 384.1258 | 447.7483 | 885.0924 | 711.5045 | 713.3441 | 697.5969 | 789.0256 | 679.8486 | 696.4621 | 614.9499 | 734 | 1551.6637 | 1569.1938 | 577.6047 | 1048.4532 | 1260.6882 | 1173.9666 | 853.4877 | 513.6726 | 622.1927 | 627.7227 | 450.7561 | 708.4454 | 758.4031 | 788.6303 | 240.2606 | 1394.8831 | 1874.0601 | 1533.2094 | 863.8564 | 724.5969 | 821.9332 | 882.1514 | 415.5624 | 230.7561 | 242.3196 | 243.647 | 405.294 | 1109.3285 | 1210.9332 | 1235.059 | 540.9666 | 699.5379 | 906 | 798.6726 | 429.3363 | 872.8909 | 1055.7483 | 1217.4967 | 588.8575 | 540.118 | 555.8318 | 576.7817 | 350.6303 | 860.1258 | 1159.0757 | 1129.6047 | 732.5802 | 722.6804 | 1032.1013 | 1324.6971 | 673.8987 | 811.3697 | 1139.0512 | 1005.5212 | 813.8408 | 887.706 | 1101.4287 | 988.7817 | 784.6726 | 581.412 | 674.1091 | 599.1849 | 497.3363 | 638.5802 | 665.4031 | 573.8318 | 422.8742 | 607.6726 | 740.5969 | 708.7227 | 488.8653 | 1596.4443 | 2219.118 | 1987.5546 | 1245.8151 | 1631.5969 | 1895.1091 | 1750.7906 | 1348.9165 | 288.8909 | 196.7817 | 172.1682 | 288.8909 | 311.7227 | 355.6303 | 332.1514 | 320.8909 | 327.4955 | 398.3786 | 344.353 | 282.3107 | 188.8987 | 206.5045 | 196.1258 | 237.2773 | 244.353 | 214.588 | 294.0846 | 191.647 | 1073.1938 | 1146.0334 | 1211.765 | 536.1682 | 843.0679 | 1117.4877 | 1059.2439 | 589.0757 | 1638.4465 | 1489.461 | 1505.5045 | 1033.9666
564 | 8.1984 | 662.439 | 1111.8363 | 1221.4356 | 608.3589 | 717.1534 | 904.4774 | 945.9096 | 496.4041 | 496.798 | 387.8431 | 405.9164 | 402.3274 | 851.8849 | 725.0384 | 735.676 | 677.9233 | 795.7596 | 730.4322 | 740.1534 | 604.202 | 803.9096 | 1544.4808 | 1575.8397 | 576.7596 | 1053.1185 | 1329.2685 | 1281.7843 | 853.3205 | 545.5158 | 660.1534 | 633.5192 | 440.7212 | 673.1253 | 802.913 | 810.1569 | 255.7178 | 1331.1705 | 1770.1774 | 1568.3137 | 936.0281 | 688.3657 | 814.9616 | 895.0384 | 420.9582 | 240.9582 | 203.6061 | 249.439 | 399.4007 | 1132.6377 | 1227.8781 | 1146.8534 | 565.0767 | 719.3171 | 946.474 | 778.0836 | 413.5226 | 846.4041 | 980.8918 | 1167.1287 | 554.2856 | 544 | 584.4356 | 521.0486 | 373.9966 | 898.474 | 1176.5575 | 1143.4774 | 742.439 | 764.1534 | 992.8466 | 1368.8329 | 750.8678 | 812.8397 | 1106.526 | 991.3623 | 755.4493 | 797.9335 | 1089.202 | 1027.7945 | 833.9932 | 580.3205 | 700.5959 | 592.561 | 534.8747 | 674.1952 | 683.237 | 651.1884 | 447.9966 | 579.7246 | 721.8431 | 672.2055 | 478.6411 | 1556.5644 | 2296.5891 | 2060.5891 | 1261.6377 | 1586.1671 | 1938.1534 | 1724.8849 | 1304.3274 | 284.4808 | 179.6829 | 158.2822 | 284.4808 | 294.5226 | 354.1603 | 358.8363 | 322 | 318.561 | 350.5678 | 333.9616 | 286.439 | 203.3171 | 225.5575 | 205.9582 | 233.3205 | 257.8781 | 233.1569 | 253.087 | 217.676 | 1043.9651 | 1205.6308 | 1172.6863 | 536.0801 | 804.9267 | 1163.3137 | 1091.3555 | 634.1534 | 1606.7664 | 1579.185 | 1532.8363 | 1033.2404
565 | 8.2094 | 627.5781 | 1154.0826 | 1199.764 | 598.7463 | 703.2891 | 889.2212 | 942.6785 | 514.5074 | 491.4749 | 381.3392 | 416.3894 | 424.1003 | 886.2862 | 733.4572 | 773.4219 | 703.0325 | 719.5633 | 721.9498 | 705.7817 | 592.7463 | 745.6813 | 1571.9646 | 1603.0325 | 586.3215 | 1089.4219 | 1266.0884 | 1245.3244 | 801.7315 | 523.764 | 698.4897 | 618.1534 | 463.3039 | 674.6608 | 772.5103 | 815.7286 | 271.9145 | 1393.0473 | 1838.301 | 1591.5752 | 918.764 | 667.2891 | 858.7434 | 836.3421 | 416.4071 | 280.2183 | 222.3717 | 264.9145 | 434.4219 | 1116.2891 | 1254.1682 | 1241.537 | 576.3215 | 749.8968 | 899.6637 | 799.3039 | 401.7463 | 853.3215 | 1027.1328 | 1153.6106 | 566.8466 | 544.9322 | 580.4749 | 557.0148 | 348.9675 | 811.581 | 1155.6283 | 1143.1357 | 728.0855 | 738.1711 | 1011.3717 | 1357.1534 | 735.6283 | 729.1061 | 1074.1032 | 1028.2183 | 801.2685 | 779.7463 | 1053.5103 | 1014.2212 | 853.8466 | 541.7817 | 689.0177 | 592.9322 | 510.0354 | 677.9322 | 709.236 | 612.3244 | 444.4071 | 573.2714 | 706.0177 | 669.9322 | 491.9823 | 1626.708 | 2247.0707 | 2037.1711 | 1293.7611 | 1585.7286 | 1921.4926 | 1814.2833 | 1355.9293 | 250.4424 | 180.7965 | 177.5074 | 250.4424 | 290.2035 | 366.118 | 328.3746 | 344.3717 | 344.1003 | 370.236 | 332.0678 | 326.1505 | 200.3392 | 177.5958 | 205.1357 | 237.6608 | 245.9498 | 265.7611 | 243.4071 | 182.7138 | 1062.4396 | 1135.4955 | 1090.699 | 580.7434 | 735.7492 | 1165.1357 | 1067.8997 | 611.8997 | 1657.065 | 1576.8141 | 1535 | 1025.5428
566 | 8.2204 | 649.5556 | 1135.2778 | 1272.6112 | 576.2778 | 721.8334 | 999.4446 | 942 | 544.3334 | 515.5556 | 423.5 | 381.1111 | 423.1667 | 869.1666 | 711.3333 | 774.1111 | 692.7222 | 776.3334 | 745.5556 | 692.6111 | 576.8889 | 780.6667 | 1624.0556 | 1576.6666 | 586.0556 | 1069.3333 | 1280.8334 | 1290.1667 | 865.0556 | 527.6667 | 663.2222 | 622.6667 | 430.0555 | 692.9445 | 846.5001 | 847.1667 | 227.6666 | 1334.7222 | 1782.611 | 1626.0556 | 879.2222 | 632 | 861.0556 | 906.6112 | 465.1112 | 239.5555 | 231.5556 | 250.8889 | 412.4444 | 1178.2778 | 1233.3333 | 1141.3332 | 532.6111 | 698.1666 | 860.1111 | 816.1111 | 378.3333 | 888.9445 | 1017.7778 | 1144.5 | 579.3333 | 571.4445 | 580.9444 | 580.7778 | 339.4444 | 815.3889 | 1109.6111 | 1177 | 693.9444 | 810.6112 | 984.6666 | 1417.389 | 724.5555 | 808.9445 | 1123.0001 | 979.0555 | 814.4445 | 832.8334 | 1092.5556 | 1047.9445 | 812.5 | 587.1667 | 670.0555 | 627 | 541.0556 | 697.8334 | 735.5556 | 589.1666 | 437.3889 | 549.3889 | 776.7779 | 713.4445 | 538.3334 | 1614.0555 | 2203.3333 | 2062.3889 | 1241.2222 | 1598.2778 | 2026.7223 | 1769.9999 | 1295.7777 | 285.7778 | 185.7222 | 174.2778 | 285.7778 | 336.2778 | 333.9444 | 330.7222 | 329.9444 | 315.7777 | 352.1666 | 360.3334 | 301.6111 | 213.2222 | 214.6112 | 228.6111 | 237.0556 | 259.1667 | 238.7222 | 244.8889 | 182.8333 | 1079.1111 | 1131.8889 | 1214.389 | 592.5 | 843.389 | 1115.8888 | 1060.3333 | 635.5 | 1563.7221 | 1654.389 | 1593.5556 | 985.3333
567 | 8.2314 | 631.8653 | 1125.3894 | 1250.2114 | 668.7602 | 731.6106 | 893.1051 | 832.932 | 574.7021 | 503.6057 | 409.7355 | 382.8269 | 418.2163 | 882.351 | 745.3993 | 778.7837 | 669.9951 | 744.6007 | 724.9951 | 768.5389 | 631.4907 | 746.644 | 1629.8702 | 1586.4808 | 594.6106 | 1085.2213 | 1306.8751 | 1246.1199 | 862.3028 | 509.822 | 620.8171 | 635.4376 | 511.2361 | 653.8171 | 773.5043 | 814.5575 | 245.0915 | 1330.0433 | 1765.6057 | 1609.822 | 912.3993 | 672.0964 | 859.0865 | 935.8751 | 439.2979 | 256.1347 | 239.6539 | 255.7404 | 418.6539 | 1081.5427 | 1219.5624 | 1200.0581 | 557.7454 | 734.2262 | 902.9666 | 802.649 | 436.3177 | 865.1681 | 1065.7936 | 1192.7936 | 591.4808 | 573.9567 | 570.4759 | 560.9518 | 379.1829 | 887.7552 | 1162.4907 | 1112.9852 | 757.0581 | 775.7738 | 968.649 | 1358.8121 | 749.8319 | 821.6539 | 1113.5624 | 1011.3993 | 829.351 | 796.7738 | 1095 | 1005.0334 | 849.2262 | 546.9469 | 676.6539 | 606.9951 | 567.8751 | 693.2596 | 717.8653 | 651.1446 | 446.5674 | 578.6156 | 758.0383 | 663.3795 | 535.2596 | 1680.9283 | 2268.9283 | 1957.8022 | 1279.1397 | 1679.3659 | 1964.115 | 1853.1063 | 1324.529 | 293.7404 | 198.4376 | 149.1249 | 293.7404 | 339.9567 | 354.0049 | 339.6106 | 325.1731 | 308.2596 | 327.0816 | 365.8269 | 307.6539 | 198.6922 | 192.1249 | 208.9518 | 224.5624 | 234.1681 | 240.8269 | 246.9135 | 182.0433 | 1103.9184 | 1151.1347 | 1201.9085 | 579.6057 | 810.7738 | 1057.5093 | 1102.0964 | 639.8702 | 1569.4808 | 1568.11 | 1523.3313 | 1060.4957
568 | 8.2424 | 679.4514 | 1136.6283 | 1280.0089 | 635.2079 | 707.7743 | 922.885 | 865.7921 | 527.5486 | 514.6283 | 420.6283 | 397.5354 | 423.8142 | 829.7035 | 754.7522 | 780.9381 | 688.3806 | 765.2876 | 743.3806 | 727.4247 | 651.4425 | 776.0089 | 1578.6415 | 1617.0399 | 578.5265 | 1042.3938 | 1272.146 | 1258.5354 | 890.1018 | 513.8451 | 701.3673 | 631.1549 | 431.6636 | 678.1637 | 830.0797 | 854.6682 | 232.4336 | 1356.1637 | 1853.1815 | 1663.2655 | 934.3496 | 738.9248 | 833.8053 | 900.177 | 476.7611 | 236.6504 | 221.5885 | 286.0399 | 439.3496 | 1124.4824 | 1248.0708 | 1252.4204 | 565.7832 | 691.4247 | 946.6682 | 792.3097 | 433.1858 | 796.1681 | 1068.969 | 1175.6194 | 530.9513 | 554.6194 | 593.2567 | 570.6593 | 375.1858 | 908.4425 | 1205.6992 | 1179.77 | 744.4956 | 760.4336 | 974.7832 | 1374.4115 | 729.6814 | 848.1637 | 1132.3806 | 1057.5753 | 845.5044 | 824.0708 | 1114.3806 | 1060.1726 | 884.916 | 526.5885 | 669.2478 | 625.3806 | 527.3318 | 661.0221 | 752.854 | 611.3628 | 444.0929 | 581.9381 | 767.6593 | 703.6372 | 508.8363 | 1645.2389 | 2186.7256 | 2020.8319 | 1301.3496 | 1715.9469 | 1941.6194 | 1798.8583 | 1357.9779 | 301.7522 | 192.2168 | 179.9779 | 301.7522 | 317.7124 | 334.6504 | 336.1239 | 329.8451 | 364.2036 | 369.6062 | 334.0221 | 292.4956 | 221.2567 | 194.8761 | 200.2478 | 262.7611 | 243.6593 | 235.1858 | 271.2257 | 192.6593 | 1038.1371 | 1131.6504 | 1187.4336 | 567.3717 | 880.708 | 1166.4382 | 1134.0399 | 625.4646 | 1586.4735 | 1666.7169 | 1562.6372 | 1028.146
569 | 8.2534 | 673.1495 | 1129.1495 | 1276.0934 | 625.1681 | 699.1495 | 905.3549 | 938.6363 | 550.533 | 536.589 | 357.2143 | 426.4582 | 429.8879 | 871.178 | 700.0462 | 827.122 | 656.6165 | 771.8879 | 707.6912 | 756.4209 | 652.9813 | 799.5703 | 1646.6736 | 1620.944 | 588.7945 | 1024.3176 | 1295.533 | 1299.2341 | 907.6824 | 511.056 | 674.5604 | 654.5516 | 501.6176 | 648.5791 | 817.2802 | 808.8967 | 264.3835 | 1376.6264 | 1854.0374 | 1612.9901 | 904.5791 | 709.5978 | 823.1868 | 877.411 | 478 | 248.7571 | 272.0286 | 239.8967 | 417.4297 | 1115.2055 | 1266.6637 | 1241.2429 | 589.5516 | 709.6264 | 895.0088 | 798.8692 | 433.9813 | 916.6648 | 992.4571 | 1263.3187 | 579.0473 | 533.3923 | 576.3363 | 566.0934 | 364.2055 | 894.2802 | 1164.8033 | 1137.8407 | 767.5516 | 752.1495 | 1027.9912 | 1371.0747 | 763.3462 | 805.822 | 1109.4484 | 1044.2802 | 870.533 | 753.3637 | 1054.1582 | 1068.8692 | 843.8033 | 580.9538 | 674.8879 | 630.9066 | 558.3835 | 680.6077 | 714.7473 | 720.889 | 450.8692 | 594.7571 | 768 | 674.5791 | 518.7945 | 1626.3363 | 2365.5439 | 2112.3 | 1341.2527 | 1593.3539 | 1936.0934 | 1751.8593 | 1399.2341 | 281.3923 | 173.3549 | 172.1681 | 281.3923 | 323.8692 | 350.6824 | 349.7385 | 339.8132 | 347.3549 | 352.3549 | 352.6264 | 296.9066 | 222 | 216.589 | 222.5703 | 236.5231 | 265.589 | 243.8319 | 259.2429 | 195.944 | 1042.8692 | 1173.1967 | 1204.6637 | 584.6637 | 873.1868 | 1183.7385 | 1113.411 | 637.7571 | 1639.0099 | 1606.2143 | 1529.6538 | 1000.5044
570 | 8.2644 | 636.2364 | 1126.0192 | 1258.115 | 630.9617 | 695.0256 | 949.7124 | 938.0128 | 571.8658 | 558.8594 | 380.8402 | 386.262 | 446.8914 | 938.5718 | 743.7124 | 753.4793 | 701.706 | 820.6869 | 745.7508 | 705.3323 | 624.1853 | 826.8275 | 1644.0256 | 1625.968 | 597.9425 | 1087.591 | 1309.9105 | 1291.0575 | 897.0703 | 532.8594 | 683.9361 | 671.8914 | 471.2045 | 721.5271 | 832.8978 | 850.7252 | 281.8914 | 1382.9617 | 1827.1725 | 1592.1278 | 886.115 | 733.8402 | 794.1853 | 876.0064 | 430.3067 | 259.9297 | 237.2301 | 246.9489 | 402.0959 | 1179.5846 | 1350.4632 | 1220.1342 | 606.8914 | 758.6869 | 842.3323 | 845.6997 | 428.0383 | 853.4218 | 1045.6485 | 1170.6071 | 560.1278 | 544.9233 | 650.5207 | 605.7444 | 336.1789 | 835.377 | 1163.0064 | 1163.8275 | 764.0256 | 785.7827 | 967.3962 | 1368.0192 | 714.3195 | 906.3482 | 1067.2684 | 1021.147 | 885.9041 | 895.0798 | 1085.7891 | 1027.2684 | 871.8147 | 583.9872 | 676.9872 | 646.8978 | 587.8147 | 681.9936 | 735.8594 | 615.6902 | 471.8658 | 608.9105 | 735.2109 | 688.9041 | 536.885 | 1644.8786 | 2263.6774 | 2065.3131 | 1403.6038 | 1620.8083 | 1980.7124 | 1760.9361 | 1408.9425 | 288.9489 | 187.9041 | 186.9041 | 288.9489 | 332.9425 | 300.3259 | 359.9361 | 356.8914 | 357.9297 | 385.7827 | 354.9872 | 279.115 | 232.9297 | 212.032 | 201.1406 | 221.0959 | 231.2237 | 242.0128 | 250.0575 | 191.032 | 1062.8722 | 1214.738 | 1228.8466 | 566.1214 | 879.9553 | 1191.9489 | 1169.6358 | 622.1022 | 1555.5432 | 1632.8211 | 1553.8402 | 1110.2907
571 | 8.2754 | 691.2301 | 1166.7522 | 1272.0708 | 666.8525 | 695.9233 | 937.2242 | 905.236 | 584.6873 | 545.0531 | 414.2065 | 411.8348 | 494.8643 | 919.8525 | 735.8171 | 754.1475 | 745.9646 | 740.3422 | 734.0177 | 778.8466 | 624.295 | 778.2714 | 1629.1593 | 1623.1121 | 620.8997 | 995.4662 | 1317.236 | 1300.7699 | 897.9705 | 533.9115 | 689.8879 | 649.2301 | 504.8702 | 650.3422 | 816.1711 | 831.8348 | 246.2832 | 1366.6696 | 1844.1357 | 1691.0118 | 949.1062 | 705.0885 | 838.6637 | 884.8761 | 467.1062 | 249.1652 | 234.0413 | 226.118 | 432.7994 | 1226.6873 | 1270.2183 | 1292.764 | 620.0413 | 703.1357 | 988.3627 | 816.0472 | 429.059 | 875.7522 | 1095.7404 | 1175.9587 | 610.9351 | 573.941 | 652.0295 | 569.9174 | 400.9705 | 834.0354 | 1201.7699 | 1215.7109 | 755.9823 | 740.3658 | 980.1711 | 1408.8466 | 785.8289 | 895.5869 | 1134.8525 | 1027.177 | 841.1475 | 797.4662 | 1077.2301 | 1027.4072 | 865.0531 | 617.7699 | 713.764 | 654.9174 | 545.3599 | 748.581 | 765.9174 | 541.5016 | 470.2301 | 590.8879 | 772.9292 | 771.5043 | 500.9882 | 1633.764 | 2165.7435 | 2026.3658 | 1345.0059 | 1612.1711 | 1980.4158 | 1831.0236 | 1389.9941 | 295.1121 | 209.8702 | 172.9764 | 295.1121 | 308.1239 | 353 | 329.8348 | 366.0413 | 321.9646 | 341.059 | 362.8997 | 308.9174 | 213.0708 | 190.1652 | 207.9351 | 244.9469 | 239.0472 | 232.0472 | 234.1357 | 187.0472 | 1136.7286 | 1135.1298 | 1252.0649 | 517.1947 | 822.4721 | 1241.7286 | 1151.9705 | 678.3186 | 1645.941 | 1646.9115 | 1594.7376 | 1088.0531
572 | 8.2864 | 729.0176 | 1125.2911 | 1283.909 | 641.6725 | 683.5094 | 975.382 | 945.0364 | 532.0546 | 554.0728 | 448.2359 | 384 | 472.0546 | 895.1637 | 705.855 | 779.4912 | 738.2354 | 797.8545 | 736.3451 | 752.9818 | 674.2365 | 822.6902 | 1656.1819 | 1641.0722 | 604.5639 | 1074.4002 | 1356.3087 | 1262.9278 | 893.2365 | 519.0364 | 670.5452 | 687.2905 | 483.2001 | 707.7453 | 843.9449 | 805.6918 | 293.3087 | 1311.4912 | 1867.3093 | 1691.7812 | 966.618 | 719.3451 | 782.6549 | 864.2001 | 485.1092 | 276.6725 | 240.7453 | 245.8181 | 399.9278 | 1176.0557 | 1307.5639 | 1254.9283 | 626.8181 | 725.5634 | 882.5827 | 824.5276 | 438.6362 | 834.1092 | 1052.5821 | 1170.237 | 599.8545 | 564.3093 | 656.4906 | 556.8186 | 395.4724 | 839.9636 | 1163.946 | 1168.7282 | 753.5821 | 802.0364 | 1008.0358 | 1383.6367 | 757.1819 | 826.6367 | 1110.1273 | 1056.618 | 866.3638 | 874.29 | 1115.4542 | 1095.8363 | 874.1092 | 578.9454 | 674.2001 | 641.946 | 604.5992 | 678.5458 | 752.1637 | 626.5094 | 509.2911 | 571.7635 | 760.9272 | 688.3093 | 499.0728 | 1596.0011 | 2289.4537 | 2087.145 | 1346.1273 | 1641.8004 | 1884.0557 | 1834.5816 | 1389.7277 | 313.6362 | 188.2911 | 169.0546 | 313.6362 | 329.3275 | 353.4002 | 302.9642 | 371.9449 | 316.6185 | 350.7817 | 346.5458 | 295.2365 | 224.7453 | 218.2729 | 197.3093 | 236.4548 | 246.4542 | 240.3093 | 256.5816 | 194.6725 | 1091.3275 | 1157.473 | 1263.9642 | 551.3644 | 901.7999 | 1196.9096 | 1146.9454 | 730.9995 | 1635.6544 | 1633.0187 | 1717.9631 | 1097.1455
573 | 8.2974 | 675.2074 | 1140.6647 | 1280.1888 | 624.8535 | 710.2989 | 995.4818 | 944.4394 | 536.2193 | 558.0914 | 406.7866 | 385.5546 | 475.5182 | 901.7748 | 751.2989 | 805.3903 | 645.9873 | 788.4454 | 701.8535 | 751.4818 | 687.2134 | 752.884 | 1666.5487 | 1594.1397 | 634.6647 | 1096.3048 | 1318.939 | 1314.2193 | 905.5428 | 520.177 | 646.7621 | 649.8535 | 493.5123 | 693.5732 | 788.9814 | 894.7138 | 256.3963 | 1340.0728 | 1884.0305 | 1627.8595 | 945.177 | 684.3963 | 818.5792 | 875.7011 | 490.116 | 258.5428 | 227.6097 | 235.8476 | 448.8112 | 1285.4149 | 1338.0914 | 1356.5309 | 617.0914 | 703.5851 | 967.4208 | 852.5732 | 419.3353 | 841.2803 | 1084.3963 | 1235.7807 | 591.5123 | 581.884 | 629.3353 | 599.5063 | 367.5487 | 839.4022 | 1212.5004 | 1261.9695 | 785.3048 | 803.4208 | 958.4691 | 1416.7502 | 768.4327 | 860.8476 | 1117.4877 | 1035.939 | 884.0491 | 786.1152 | 1088.6825 | 1087.884 | 880.0914 | 576.061 | 684.8171 | 691.4149 | 528.8476 | 708.6401 | 762.9204 | 654.6097 | 523.9636 | 561.0119 | 756.3294 | 705.823 | 504.5851 | 1703.5123 | 2208.9264 | 2044.0483 | 1353.0914 | 1684.0245 | 1992.6528 | 1813.2803 | 1426.3785 | 293.4513 | 203.1465 | 171.3903 | 293.4513 | 346.2074 | 374.9086 | 355.4572 | 317.0424 | 350.6097 | 338.6952 | 375.0855 | 307.4513 | 211.0305 | 232.7866 | 213.6037 | 259.6283 | 218.3108 | 257.3658 | 234.7316 | 177.4572 | 1108.4818 | 1185.1033 | 1316.1219 | 624.0245 | 891.6401 | 1243.9272 | 1146.0119 | 677.3658 | 1617.1219 | 1686.4454 | 1661.3294 | 1106.1583
574 | 8.3084 | 651.2404 | 1131.711 | 1317.8289 | 652.4705 | 687.8717 | 976.3908 | 865.9942 | 574.9572 | 561.1711 | 400.4705 | 434.059 | 491.0162 | 834.7213 | 728.2301 | 787.0266 | 740.3584 | 741.3098 | 729.6416 | 735.0428 | 694.3422 | 783.0104 | 1682.2035 | 1691.733 | 623.7861 | 1107.4971 | 1319.3098 | 1350.8185 | 892.8821 | 495.7965 | 669.2035 | 674.2625 | 480.0531 | 681.3688 | 850.1607 | 823.8821 | 268.8289 | 1404.6312 | 1883.2035 | 1688.1121 | 921.2242 | 697.385 | 836.7861 | 898.3422 | 463.9248 | 243.9572 | 247.5988 | 232.326 | 411.2566 | 1236.0693 | 1342.4543 | 1277.3156 | 617.3052 | 754.9838 | 949.8555 | 838.5295 | 430.2994 | 880.8879 | 1094.5619 | 1199.6948 | 576.6844 | 610.1711 | 638.289 | 553.8393 | 384.3156 | 880.9631 | 1135.8659 | 1260.0162 | 794.7006 | 784.615 | 1035.9734 | 1381.7479 | 814.8555 | 854.674 | 1132.059 | 1035.1549 | 824.5237 | 914.1769 | 1168.6105 | 1113.4763 | 883.5561 | 578.0855 | 679.4705 | 641.6844 | 526.8231 | 729.8555 | 821.3052 | 673.615 | 490.5295 | 622.1769 | 737.267 | 731.7006 | 553.7596 | 1688.2566 | 2300.5885 | 2141.0428 | 1353.1341 | 1651.9734 | 1924.63 | 1855.2832 | 1346.2508 | 277.8393 | 177.1977 | 152.9838 | 277.8393 | 321.6416 | 370.7168 | 369.4867 | 380.9734 | 369.6578 | 330.5826 | 345.4867 | 291.8821 | 212.385 | 225.5723 | 201.5133 | 217.0693 | 258.8451 | 267.5457 | 258.1283 | 192.2566 | 1091.4439 | 1252.385 | 1317.8185 | 594.0693 | 914.2242 | 1180.5237 | 1210.0428 | 691.0531 | 1622.9248 | 1639.1815 | 1642.267 | 1142.1873
575 | 8.3194 | 678.4494 | 1170.2203 | 1311.2468 | 663.4405 | 684.8348 | 922.4052 | 931.1829 | 574.4956 | 566.7621 | 409.8987 | 412.6696 | 468.8348 | 916.7335 | 716.947 | 809.0088 | 728.7532 | 792.8348 | 743.8899 | 735.7797 | 702.7709 | 825.6873 | 1643.8722 | 1640.6696 | 621.6982 | 1138.9075 | 1365.6036 | 1303.1013 | 937.5881 | 537.5418 | 632.7444 | 616.6894 | 521.0442 | 713.2114 | 808.7797 | 866.2753 | 265.4405 | 1372.7335 | 1844.1013 | 1710.8525 | 970.3392 | 708.3127 | 828.9714 | 905.2291 | 505.1278 | 244.9273 | 260.2842 | 261.7247 | 417.4956 | 1263.5242 | 1375.826 | 1265.6343 | 559.9184 | 771.934 | 967.5683 | 828.6057 | 438.9823 | 859.414 | 1040.2753 | 1258.8987 | 592.2753 | 614.1652 | 604.2951 | 596.8171 | 371.0374 | 815.4229 | 1199.0639 | 1238.5506 | 786.7335 | 776.3304 | 1011.6234 | 1485.4119 | 831.6873 | 826.0462 | 1111.4956 | 1061 | 879.9538 | 850.4494 | 1054.0021 | 1057.2203 | 895.9449 | 580.881 | 689.2291 | 658.2114 | 588.9538 | 750.3216 | 761.0551 | 667.2577 | 481.0462 | 560.6519 | 782.9891 | 724.0639 | 525.4494 | 1690.6961 | 2277.3699 | 2138.0904 | 1296.5439 | 1632.414 | 2071.6234 | 1890.8831 | 1422.9647 | 316.9075 | 201.9626 | 173.9538 | 316.9075 | 333.7423 | 341.119 | 356.3481 | 358.3304 | 359.4119 | 365.8899 | 335.9823 | 331.3569 | 220.0639 | 217.4868 | 213.3855 | 242.5418 | 234.8722 | 235.5506 | 259.5683 | 198.0551 | 1080.3216 | 1257.9714 | 1268.6145 | 612.5486 | 959.0442 | 1235.6234 | 1213.0925 | 734.7974 | 1668.3678 | 1685.2468 | 1692.2842 | 1123.0639
576 | 8.3304 | 668.8653 | 1179.5905 | 1265.6165 | 671.4041 | 683.6165 | 995.0363 | 902.9637 | 583.5388 | 594.1037 | 392.6839 | 406.3938 | 465.5959 | 894.3264 | 812.938 | 793.4818 | 774.5182 | 792.4248 | 740.7876 | 734.2227 | 713.7773 | 785.5959 | 1673.912 | 1634.3264 | 665.5752 | 1102.4612 | 1277.5802 | 1319.171 | 983.0103 | 563.4612 | 693.9016 | 678.0467 | 439.3004 | 731.7099 | 804.1917 | 869.4508 | 274.3471 | 1348.3731 | 1863.3368 | 1680.3315 | 959.8083 | 747.2591 | 780.3678 | 891.1244 | 472.6632 | 277.1244 | 247.5285 | 254.912 | 431.0517 | 1320.4975 | 1390.6062 | 1330.4922 | 613.8393 | 687.9637 | 942.7979 | 839.2694 | 471.0467 | 910.1867 | 1049.0414 | 1241.7513 | 597.8083 | 612.9586 | 677.4612 | 626.0467 | 407.9533 | 859.6942 | 1188.9947 | 1251.5025 | 770.9947 | 783.6839 | 988.6062 | 1444.031 | 794.088 | 844 | 1121.0103 | 1061.0674 | 863.5492 | 842.2798 | 1156.2591 | 1062.4818 | 897.1554 | 594.9223 | 677.0103 | 681.6269 | 575.5129 | 775.1554 | 763.886 | 706.2694 | 498.1244 | 588.6632 | 698.3471 | 709.0777 | 517.6165 | 1637.0984 | 2416.0157 | 2080.0257 | 1429.6372 | 1691.0674 | 2043.9586 | 2000.7826 | 1491.2694 | 281.3368 | 168.4818 | 157.3471 | 281.3368 | 299.5129 | 326.5388 | 329.3471 | 365.0103 | 317.9637 | 364.2021 | 369.1917 | 287.0103 | 203.8653 | 240.9793 | 220.3368 | 269.0674 | 265.3835 | 245.8083 | 238.3004 | 198.7306 | 1098.6892 | 1209.7046 | 1264.5595 | 522.1657 | 879.9273 | 1214.9637 | 1252.3004 | 700.4352 | 1707.9223 | 1638.7409 | 1678.7202 | 1108.9533
577 | 8.3414 | 666.6018 | 1245.6901 | 1294.6107 | 676.5221 | 703.8495 | 952.4071 | 942.0088 | 591.2389 | 553.7611 | 414.4867 | 398.7168 | 461.4336 | 889.062 | 729.1593 | 814.4779 | 751.6106 | 820.9822 | 726.6284 | 765.1593 | 680.6106 | 780.3982 | 1653.2833 | 1625.0796 | 632.7168 | 1097.3894 | 1342.4336 | 1286.6107 | 993.3362 | 552.9734 | 652.1947 | 653.9469 | 486.8319 | 699.6018 | 791.2478 | 847.8407 | 290.5309 | 1319.2479 | 1862.1858 | 1762.3451 | 971.3186 | 730.6282 | 831.5399 | 881.6284 | 466.646 | 265.1593 | 240.7611 | 227.7522 | 494.7079 | 1277.2567 | 1398.6814 | 1367.708 | 663.9734 | 731.3098 | 937.4513 | 842.6018 | 444.2389 | 808.2125 | 1103.5664 | 1266.7965 | 610.4336 | 561.3806 | 672.2301 | 600.354 | 430.0619 | 894.5309 | 1269.9292 | 1294.5044 | 855.354 | 804.6106 | 1001.2213 | 1487.7522 | 824.7522 | 846.2301 | 1133.1681 | 1057.3805 | 887.4691 | 860 | 1144.8849 | 1086.6638 | 927.8761 | 584.2213 | 692.8319 | 717.2035 | 632.584 | 800.938 | 790.9558 | 707.2124 | 488.2301 | 586.4691 | 724.1682 | 724.5575 | 539.9204 | 1680.2389 | 2347.9558 | 2202.6106 | 1475.6991 | 1695.3452 | 1997.9559 | 1892.8939 | 1496.6726 | 285.2832 | 189.4779 | 177.6726 | 285.2832 | 358.5752 | 336.5487 | 347.2035 | 377.1681 | 359 | 367.4779 | 357.9204 | 302.9912 | 203.4336 | 214.4336 | 222.8495 | 269.5221 | 242.2832 | 260.8231 | 281.0531 | 195.3186 | 1045.4514 | 1261.8761 | 1301.5929 | 569.3452 | 961.2832 | 1258.5487 | 1290.3539 | 749.5929 | 1688.8584 | 1650.1947 | 1663.6284 | 1134.0885
578 | 8.3524 | 664.3904 | 1124.5616 | 1365.2998 | 667.5064 | 684.884 | 939.6082 | 932.0552 | 594.4597 | 553.2065 | 397.7355 | 407 | 478.5516 | 940.2292 | 729.6209 | 819.3564 | 732.8261 | 766.8614 | 754.1499 | 765.6209 | 655.6662 | 786.3791 | 1743.024 | 1629.3098 | 643.5743 | 1123.8967 | 1355.966 | 1362.0806 | 953.0226 | 521.4017 | 665.4484 | 687.3564 | 504.3225 | 695.3678 | 775.505 | 844.8967 | 259.9774 | 1405.1146 | 1800.0566 | 1699.2985 | 976.471 | 663.0679 | 881.058 | 914.2985 | 457.0679 | 266.2645 | 236.8967 | 248 | 424.4823 | 1341.7695 | 1400.3324 | 1400.3791 | 634.5163 | 756.7129 | 914.4357 | 838.9194 | 448.6549 | 880.3678 | 1085.8967 | 1278.2985 | 625.4371 | 609.058 | 695.058 | 616.1839 | 376.034 | 870.0453 | 1246.7355 | 1239.1825 | 868.9774 | 782.884 | 1044.116 | 1506.9887 | 842.6096 | 873.9307 | 1121.2857 | 1011.1712 | 907.5417 | 858.1612 | 1132.9194 | 1122.5643 | 934.3225 | 634.6549 | 710.6902 | 705.9774 | 570.9194 | 754.9194 | 798.3225 | 674.8161 | 518.3211 | 611.7822 | 799.7369 | 729.0693 | 538.0806 | 1687.505 | 2351.2772 | 2182.447 | 1436.6648 | 1734.7355 | 2109.5064 | 1958.9774 | 1511.3112 | 283.3437 | 194.3564 | 195.8967 | 283.3437 | 318.4484 | 365.4484 | 339.1033 | 357.5629 | 356.6096 | 371.3451 | 356.3564 | 326.4597 | 218.1612 | 228.2419 | 199.1952 | 264.4597 | 235.2871 | 300.587 | 246.3564 | 197.7129 | 1117.2419 | 1271.3564 | 1318.4371 | 609.8388 | 952.9081 | 1285.1499 | 1213.86 | 767.4484 | 1649.7341 | 1668.0453 | 1697.4017 | 1115.8614
579 | 8.3634 | 685.0816 | 1216.707 | 1312.648 | 634.0423 | 739.352 | 994.1662 | 1050.8898 | 600.6677 | 574.7915 | 408.1269 | 405.3323 | 485.8761 | 975.3293 | 716.3127 | 813.3127 | 814.1632 | 805.5211 | 721.6224 | 760.0256 | 686.3776 | 799.1662 | 1688.4819 | 1672.6314 | 671.8142 | 1110.4985 | 1327.355 | 1363.8398 | 971.8535 | 557.0846 | 659.4789 | 676.3097 | 472.6903 | 702.8565 | 817.1858 | 842.8595 | 245.8368 | 1426.7296 | 1851.355 | 1718.0393 | 988.8112 | 720.0816 | 847.8308 | 936.7915 | 513.9774 | 256.1858 | 227.4592 | 245.0816 | 446.29 | 1373.4759 | 1447.997 | 1415.9804 | 688.0619 | 746.71 | 982.0649 | 852.8988 | 459.4759 | 882.4366 | 1075 | 1302.7719 | 608.2281 | 588.0256 | 668.3973 | 619.5634 | 408.1888 | 900.3067 | 1249.1012 | 1319.7915 | 854.003 | 830.994 | 987.7719 | 1501.5861 | 818.6873 | 884.7659 | 1207.1632 | 1081.8958 | 832.5468 | 842.1888 | 1145.3353 | 1067.0256 | 906.1299 | 650.2704 | 681.2311 | 685.2704 | 583.2311 | 769.6284 | 767.8368 | 680.6087 | 555.5801 | 575.8308 | 730.8142 | 716.0453 | 533.4819 | 1731.7915 | 2290.009 | 2260.1692 | 1537.1012 | 1738.4562 | 2066.4532 | 1942.3353 | 1464.3746 | 337.349 | 187.2704 | 186.4592 | 337.349 | 313 | 360.7296 | 350 | 375.293 | 336.4396 | 359.2734 | 350.4169 | 329.9577 | 200.5211 | 203.3973 | 210.1662 | 270.5634 | 249.1465 | 259.6057 | 240.2085 | 187.0846 | 1084.3716 | 1260.2055 | 1307.5861 | 621.9547 | 951.6873 | 1262.3157 | 1332.4592 | 755.9547 | 1731.7266 | 1709.9804 | 1733.293 | 1154.9381
580 | 8.3744 | 657.5531 | 1195.8348 | 1342.2522 | 644.9808 | 709.3982 | 982.2227 | 962.9513 | 612.6696 | 574.0487 | 432.6696 | 397.0591 | 499.4366 | 918.3496 | 725.2921 | 814.7183 | 764.5826 | 810.4661 | 692.2729 | 823.8731 | 715.494 | 794.3304 | 1726.6504 | 1754.4453 | 690.7566 | 1097.9721 | 1337.7183 | 1449.7757 | 966.1947 | 573.6018 | 657.4469 | 643.2243 | 512.767 | 736 | 799.1756 | 911.5427 | 297.2817 | 1427.2905 | 1861.1357 | 1701.8276 | 967.9026 | 696.0487 | 813.5148 | 934.1844 | 491.0295 | 232.1165 | 238.3687 | 222.0974 | 426.146 | 1330.8938 | 1403.8348 | 1435.2139 | 680.1165 | 770 | 1007.6784 | 892.9026 | 420.506 | 878.9042 | 1081.5251 | 1312.5235 | 596.0295 | 649.8923 | 686.087 | 634.8643 | 424.7861 | 822.826 | 1212.4765 | 1314.1548 | 899.3304 | 750.826 | 1006.6121 | 1535.0575 | 822.7582 | 819.0104 | 1161.6608 | 1079.7183 | 904.1357 | 863.913 | 1164.0958 | 1132.7374 | 973.4366 | 641.087 | 713.4278 | 680.9808 | 611.5826 | 816.3687 | 825.3408 | 745.767 | 511.0974 | 544.6608 | 755 | 764.9705 | 571.767 | 1740.4868 | 2412.5914 | 2300.8938 | 1497.4469 | 1713.593 | 2009.6624 | 1975.7773 | 1469.9529 | 269.0591 | 182.1652 | 195.0383 | 269.0591 | 316.6121 | 367 | 348.9513 | 390.8053 | 366 | 396.6991 | 355.5148 | 322.7478 | 204.767 | 226.5634 | 205.3304 | 282.3687 | 259.233 | 277.3392 | 242.233 | 205 | 1040.6608 | 1220.6313 | 1347.1165 | 575.1269 | 953.1947 | 1310.3304 | 1346.3982 | 712.506 | 1698.0295 | 1732.3599 | 1758.593 | 1166.9617
581 | 8.3854 | 639 | 1185.1966 | 1331.1898 | 666.6982 | 723.4779 | 943.0373 | 956.0305 | 593.8407 | 585.0915 | 419.766 | 443.5152 | 484.766 | 915.2035 | 772.2271 | 806.1593 | 768.2271 | 810.7797 | 760.1525 | 786.2881 | 654.2203 | 813.2881 | 1713.5663 | 1689.2134 | 649.1966 | 1165.3559 | 1332.3796 | 1393.4543 | 1004.5694 | 563.1288 | 678.3186 | 694.9085 | 511.2576 | 735.2271 | 837.4985 | 862.5999 | 306.0983 | 1394.1761 | 1846.3559 | 1816.7766 | 955.3491 | 705.2881 | 823.3255 | 924.7797 | 517.3491 | 237.122 | 230.2508 | 241.3491 | 451.8034 | 1368 | 1391.6509 | 1466.8133 | 683.4474 | 775.2813 | 940.8643 | 876.8712 | 480.7492 | 949.1052 | 1133.5762 | 1277.827 | 627.888 | 597.8034 | 668.8948 | 648.0983 | 405.939 | 875.888 | 1261.0846 | 1289.5084 | 906.9527 | 797.8338 | 1038.766 | 1492.9627 | 869.4916 | 862.0236 | 1187.7256 | 1068.5831 | 895.9253 | 883.0373 | 1103.6845 | 1111.3491 | 954.6441 | 626.4711 | 746.0983 | 696.6441 | 618.5762 | 809.5389 | 876.4916 | 739.7492 | 532.7965 | 581.4169 | 761.4406 | 747.0236 | 569.6136 | 1797.458 | 2358.4611 | 2318.0343 | 1515.6136 | 1751.6204 | 2136 | 2006.9764 | 1550.295 | 311.9085 | 190.0983 | 162.4238 | 311.9085 | 347.8643 | 365.9695 | 341.2576 | 361.9627 | 367.2881 | 398.5221 | 365.8407 | 336.5152 | 203.1288 | 241.7492 | 222.7424 | 271.0305 | 258.8102 | 244.4474 | 245.9322 | 205 | 1087.2066 | 1234.9322 | 1340.2713 | 618.5084 | 979.4611 | 1331.8643 | 1349.4001 | 772.6204 | 1721.5152 | 1770.0373 | 1796.1052 | 1189.2713
582 | 8.3964 | 641.2576 | 1170.3191 | 1350.3742 | 712.5644 | 743.4724 | 925.9631 | 962.1534 | 583.6013 | 600.9203 | 451.5767 | 442.6258 | 517.5644 | 878.2576 | 763.8957 | 816.2576 | 769.2085 | 789.1411 | 788.5093 | 801.0674 | 681.503 | 817.7669 | 1640.0369 | 1737.7178 | 630.9326 | 1156.4847 | 1362.9877 | 1379.7546 | 1042.3682 | 559.4847 | 692.7301 | 675.8343 | 516.5276 | 729.7178 | 790.1594 | 911.3191 | 299.2945 | 1454.5339 | 1851.4539 | 1693.6869 | 980.2576 | 719.9263 | 816.3742 | 899.4724 | 535.908 | 261.1902 | 256.178 | 265.2699 | 459.6258 | 1369.411 | 1382.4047 | 1555.0063 | 698.589 | 808.4602 | 960.086 | 879.8098 | 454.9631 | 919.5276 | 1144.086 | 1319.7669 | 691.5276 | 610.8466 | 699.6013 | 639.6013 | 396.6564 | 923.1597 | 1309.8466 | 1314.6381 | 846.6195 | 816.4356 | 1061.9695 | 1520.0551 | 838.8589 | 891.2025 | 1134.595 | 1064.2208 | 936.7055 | 868.6441 | 1194.675 | 1135.8343 | 960.1288 | 677.2822 | 736.4724 | 690.4173 | 640.3742 | 848.411 | 859.5458 | 717.6318 | 573.1288 | 590.4356 | 803.3068 | 775.7792 | 561.7792 | 1651.914 | 2325.7975 | 2200.3556 | 1513.1411 | 1723.6932 | 2132.7546 | 1983.0183 | 1536.3742 | 293.1165 | 188.6564 | 152.6195 | 293.1165 | 366.8038 | 361.1165 | 343 | 382.2822 | 373.7424 | 377.822 | 360.2576 | 340.1411 | 204.1411 | 216.6687 | 218.8835 | 278.2822 | 248.2331 | 263.1165 | 254.3436 | 205.4233 | 1160.6013 | 1246.3068 | 1297.3128 | 649.5644 | 953.1471 | 1357.0123 | 1311.1225 | 743 | 1717.6627 | 1761.006 | 1779.3679 | 1146.0305
583 | 8.4074 | 646.8702 | 1167.652 | 1382.062 | 716.4602 | 728.1357 | 940.7787 | 967.8909 | 574.5222 | 575.9823 | 444.5516 | 412.1652 | 523.1475 | 895.9882 | 759.5724 | 824.4778 | 800.2507 | 794.1416 | 753.4484 | 828.7964 | 716.6873 | 790.6195 | 1637.4927 | 1728.1769 | 607.0384 | 1126.5457 | 1370.9941 | 1359.5398 | 1002.6136 | 532.9204 | 684.5398 | 685.3127 | 533.0914 | 738.5841 | 850.4661 | 889.6312 | 295.4602 | 1408.062 | 1906.8525 | 1786.0502 | 1004.2774 | 748.3303 | 842.3864 | 886.2832 | 515.7552 | 243.8289 | 240.8407 | 271.7729 | 439.5959 | 1378.4602 | 1459.8584 | 1493.3511 | 691.3805 | 785.1769 | 927.2036 | 897.1711 | 467.7846 | 928.7285 | 1094.3303 | 1288.4778 | 701.6489 | 621.6756 | 684.0796 | 633.7434 | 429.6018 | 855.7375 | 1302.115 | 1307.0031 | 900.8231 | 818.2154 | 1014.8231 | 1567.9086 | 834.9321 | 864.4543 | 1183.7964 | 1098.4543 | 939.9262 | 900.6991 | 1148.7552 | 1143.3186 | 964.8525 | 678.233 | 714.3864 | 637.1977 | 662.2448 | 860.0679 | 937.705 | 762.6018 | 582.7611 | 588.5339 | 797.1416 | 744.1534 | 536.4425 | 1703.8172 | 2352.5457 | 2311.2154 | 1509.8584 | 1729.5516 | 2114.0738 | 2048.4484 | 1561.0059 | 307.8407 | 215.0059 | 194.2448 | 307.8407 | 323.9204 | 373.3864 | 343.3068 | 380.3186 | 362.4543 | 395.767 | 377.2213 | 342.0738 | 204.8466 | 239 | 209.2213 | 280 | 272.0059 | 279.528 | 231.2154 | 206.0059 | 1142.6253 | 1238.4543 | 1349.8584 | 647.7846 | 1020.5398 | 1353.6077 | 1378.115 | 754.9647 | 1677.4484 | 1827.3864 | 1878.233 | 1200.0796
584 | 8.4184 | 612.6038 | 1231.8991 | 1407.5147 | 714.8579 | 756.8638 | 952.6575 | 1025.8461 | 606.1598 | 606.6804 | 420.8284 | 392.1598 | 539.3078 | 927.7754 | 828.8991 | 801.1244 | 769.9587 | 819.9823 | 777.8343 | 781.6451 | 732.7931 | 797.5802 | 1716.0413 | 1705.8762 | 674.5265 | 1131.2253 | 1389.7872 | 1364.4558 | 1006.1657 | 537.5029 | 690.2842 | 674.0118 | 514.2842 | 783.0825 | 845.148 | 859.5029 | 296.0118 | 1435.1716 | 1897.651 | 1847.7636 | 1056.0295 | 702.852 | 834.8579 | 940.4499 | 547.9646 | 266.3903 | 249.3432 | 264.3373 | 484.7813 | 1390.7754 | 1446.6097 | 1548.2128 | 697.0118 | 747.3019 | 970.3785 | 869.1421 | 442.148 | 865.3078 | 1059.1244 | 1267.6097 | 644.9646 | 660.2901 | 676.349 | 678.0236 | 449.201 | 903.4021 | 1236.6333 | 1427.7688 | 942.6745 | 841.6981 | 1055.0177 | 1595.9941 | 821.9764 | 875.5147 | 1146.4204 | 1110.1775 | 934.9941 | 890.8166 | 1184.1126 | 1130.6215 | 964.7636 | 680.4617 | 716.6333 | 699 | 642.9587 | 876.5029 | 919.4676 | 785.8461 | 600.1539 | 593.6863 | 812.8756 | 749.4735 | 563.2187 | 1758.3314 | 2358.8821 | 2322.4374 | 1486.5029 | 1710.3019 | 2121.6568 | 2070.0177 | 1558.6156 | 314.0236 | 205.1598 | 173.6333 | 314.0236 | 334.9646 | 370.1657 | 346.9882 | 367.7872 | 381.9646 | 393.3726 | 391.1893 | 347.6686 | 205.9882 | 236.8461 | 227.8284 | 279.3373 | 262.4912 | 292.5383 | 257.0177 | 192.5147 | 1142.2253 | 1248.1893 | 1342.9057 | 624.9646 | 1019.6568 | 1361.0413 | 1325.0596 | 818.5147 | 1707.4676 | 1824.3314 | 1857.5972 | 1199.1421
585 | 8.4294 | 668.1357 | 1167.1012 | 1394.1101 | 690.6185 | 732.7798 | 1082.9566 | 1011.6185 | 618.5851 | 599.6274 | 428.7375 | 403.6952 | 559.0423 | 966.4994 | 781.7208 | 833.0845 | 801.4405 | 829.4828 | 788.5684 | 799.822 | 760.4149 | 858.3637 | 1674.4149 | 1810.1101 | 657.8732 | 1208.762 | 1430.7798 | 1403.9577 | 1013.941 | 543.7208 | 727.0167 | 670.0512 | 548.703 | 715.3137 | 863.3137 | 864.475 | 285.4661 | 1431.178 | 1914.7375 | 1892.762 | 1037.4071 | 704.6196 | 828.416 | 972.3215 | 564.6518 | 222.9833 | 240.1101 | 264.2036 | 517.2536 | 1447.6529 | 1487.703 | 1524.1279 | 691.8053 | 780.4071 | 935.0256 | 881.4572 | 460.6696 | 903.7297 | 1089.4828 | 1322.5851 | 672.5083 | 685.0756 | 663.356 | 662.0934 | 422.6696 | 856.5684 | 1262.1524 | 1349.8988 | 934.7542 | 818.9155 | 1043.9155 | 1607.7631 | 847.3393 | 880.0856 | 1207.7709 | 1106.8732 | 940.7119 | 905.7119 | 1214.5161 | 1174.3393 | 1012.6096 | 701.9655 | 744.644 | 703.0934 | 687.0768 | 880.0512 | 945.5762 | 779.2714 | 617.4071 | 574.1524 | 783.7631 | 773.2458 | 529.3226 | 1763.0256 | 2452.9488 | 2255.713 | 1492.8988 | 1747.1446 | 2117.8977 | 2050.1947 | 1595.5929 | 299.6952 | 209.5762 | 193.9911 | 299.6952 | 359.6607 | 375.0934 | 357.8899 | 415.8988 | 409.1524 | 364.2714 | 370.2202 | 350.6274 | 219.2036 | 228.3137 | 229.3304 | 278.6696 | 270 | 266.7631 | 242.8899 | 183.5595 | 1226.3471 | 1232.9155 | 1412.8554 | 655 | 1037.7887 | 1332.0512 | 1418.4917 | 810.8053 | 1732.8643 | 1811.406 | 1789.9499 | 1212.3471
586 | 8.4404 | 680.2832 | 1141.6019 | 1372.4955 | 683.1815 | 743.2832 | 999.1859 | 1000.7567 | 645.1947 | 624.6151 | 445.2389 | 433.6151 | 548.6681 | 915.146 | 807.9203 | 807.5796 | 775.1061 | 812.7523 | 847.8983 | 797.4779 | 732.854 | 815.2213 | 1662.6461 | 1785.4512 | 687.0929 | 1142.4292 | 1427.0132 | 1409.3849 | 1047.7124 | 583.4292 | 704.7611 | 692.1017 | 513.8584 | 723.7168 | 869.4336 | 892.3407 | 282.8097 | 1433.7124 | 1926.6725 | 1834.0399 | 1067.5708 | 770.9689 | 883.9645 | 932.5133 | 511.9425 | 245.6195 | 232.3849 | 274.6239 | 468.8983 | 1471.9645 | 1482.54 | 1595.5221 | 728.27 | 804.4823 | 958.7036 | 856.8408 | 471.8584 | 975.8628 | 1080.9337 | 1361.7523 | 669.3717 | 635.1372 | 665.3805 | 687.2788 | 437.6637 | 906.3849 | 1248.6195 | 1391.1903 | 940.7965 | 833.1947 | 1054.9601 | 1632.3407 | 881.5133 | 955.7655 | 1192.6992 | 1127.7212 | 941.146 | 909.7611 | 1147.5664 | 1188.9159 | 966.8496 | 655.1859 | 752.3231 | 723.7124 | 708.4512 | 897.9159 | 943.9469 | 805.6593 | 639.3893 | 565.7567 | 810.6239 | 784.1372 | 570.1504 | 1774.3363 | 2434.8939 | 2350.6725 | 1530.5752 | 1796.4735 | 2078.9337 | 2047.4072 | 1559.3407 | 325.4292 | 210.8009 | 162.3275 | 325.4292 | 339.7567 | 388.6725 | 359.146 | 446.1548 | 398.0929 | 386.2832 | 362.7611 | 371.6195 | 236.6637 | 236.7168 | 215.6681 | 294.0443 | 270.1903 | 294.3849 | 252.3275 | 201.7124 | 1181.6769 | 1256.708 | 1375.7036 | 657.0929 | 1087.7567 | 1372.1771 | 1425.6239 | 818.73 | 1728.1372 | 1753.3231 | 1868.5133 | 1192.0311
587 | 8.4514 | 694.8614 | 1212.0512 | 1322.3069 | 718.4051 | 751.3795 | 1023.0767 | 1015.3284 | 631.0256 | 639.5949 | 446.0512 | 442.113 | 552.9253 | 924.066 | 741.3156 | 784.3049 | 743.1279 | 843.5182 | 856.7954 | 799.3049 | 741.1514 | 871.29 | 1738.7591 | 1721.6992 | 688.5054 | 1149.7228 | 1381.6716 | 1404.1279 | 1056.823 | 574.7207 | 701.3156 | 671.0512 | 513.8614 | 717.8486 | 840.0747 | 866.4307 | 288.6844 | 1450.3176 | 1909.2665 | 1853.467 | 1060.2772 | 710.0384 | 843.0767 | 952.5417 | 547.1642 | 247.8486 | 226.7228 | 256.2026 | 492.9872 | 1427.0256 | 1585.3647 | 1606.0512 | 779.2772 | 774.7207 | 996.354 | 910.8849 | 459.5054 | 951.1642 | 1154.2153 | 1392.113 | 709.0512 | 658.4179 | 681.5841 | 707.5182 | 448.7974 | 900.1151 | 1246.7974 | 1395.8486 | 987.5565 | 821.2537 | 1043.2302 | 1597.113 | 911.8721 | 929.8465 | 1230.9468 | 1105.8614 | 943.5841 | 916.8486 | 1160.5182 | 1152.887 | 1002.0767 | 666.7207 | 791.5949 | 735.0512 | 640.0404 | 867.9637 | 966.6205 | 832.9872 | 604.7719 | 594.71 | 793.6205 | 796.0747 | 538.5693 | 1752.467 | 2472.1386 | 2334.887 | 1531.5969 | 1789.1749 | 2176.9274 | 2128.5821 | 1540.5202 | 327.0512 | 237.5182 | 185 | 327.0512 | 352.9105 | 366.6077 | 349.2281 | 399.5949 | 400.9233 | 397.2153 | 359.9233 | 364.9361 | 252.0512 | 237.533 | 212.0256 | 303.7591 | 264.5182 | 281.6332 | 282.8997 | 209.6077 | 1158.8123 | 1268.644 | 1419.4307 | 665.5949 | 1109.0128 | 1433.3176 | 1424.0148 | 767.5202 | 1738.8593 | 1821.5606 | 1919.9381 | 1256.2026
588 | 8.4624 | 708.8436 | 1235.3049 | 1412.6288 | 724.297 | 752.9494 | 1034.1091 | 1004.6018 | 643.297 | 632.4139 | 458.5624 | 405.7615 | 586.6288 | 1002.2733 | 727.2464 | 826.6682 | 792.0236 | 816.0664 | 791.4961 | 839.9494 | 734.3791 | 862.1373 | 1713.7536 | 1803.8279 | 669.7424 | 1194.2306 | 1419.684 | 1438.4139 | 1077.9652 | 554.297 | 683.8088 | 679.6951 | 540.9494 | 728.0585 | 813.8909 | 889.7503 | 310.4882 | 1476.2576 | 1928.0664 | 1810.0664 | 1037.2543 | 726.9415 | 871.7267 | 884.0315 | 534.2812 | 266.8673 | 265.4297 | 262.3712 | 490.5039 | 1439.5118 | 1513.2891 | 1614.2654 | 756.2543 | 751.9336 | 982.8594 | 883.1642 | 440.9573 | 932.2654 | 1136.6557 | 1374.0236 | 717.2654 | 665.9336 | 726.7536 | 685.0079 | 452.5118 | 929.5545 | 1248.793 | 1418.9494 | 968.4139 | 839.9336 | 1103.613 | 1597.2858 | 879.5197 | 886.297 | 1161.883 | 1124.7852 | 986.6051 | 915.2418 | 1147.8909 | 1177.4218 | 1029.8673 | 653.387 | 794.9415 | 750.3555 | 705.7615 | 920.0742 | 969.5624 | 834.1564 | 603.09 | 622.7694 | 793.7694 | 757.4297 | 535.8752 | 1706.2733 | 2470.1879 | 2364.5782 | 1616.2037 | 1737.9067 | 2276.207 | 2131.3791 | 1588.5276 | 339.1327 | 217.8009 | 181.7773 | 339.1327 | 333.1642 | 373.297 | 361.0079 | 408.7424 | 387.5276 | 398.5624 | 347.8167 | 356.1721 | 260.6951 | 267.613 | 221.0742 | 287.9921 | 251.2464 | 295.0079 | 299.1754 | 218.4455 | 1217.91 | 1227.7346 | 1423.6288 | 666.5782 | 1111.0664 | 1481.1721 | 1484.7694 | 814.883 | 1721.4264 | 1902.8469 | 1959.3476 | 1258.9336
589 | 8.4734 | 703.1741 | 1246.1002 | 1402.9941 | 749.0856 | 777.0029 | 983.3658 | 1040.6401 | 663.9971 | 616.9971 | 443.7286 | 389.944 | 567.6814 | 991.8142 | 773.59 | 864.0443 | 835.4484 | 824.8998 | 793.7228 | 864.4572 | 767.4985 | 919.7019 | 1669.6814 | 1788.6668 | 705.5929 | 1201.3274 | 1464.9616 | 1422.9971 | 1121.0501 | 565.0029 | 711.0029 | 667.7641 | 563.6401 | 710 | 853.5044 | 942.0443 | 293.3186 | 1458.3511 | 1925.0885 | 1810.2685 | 1088.0029 | 751.8142 | 883.3215 | 929.59 | 522.5486 | 273.2772 | 259.5044 | 280.41 | 501.4514 | 1477.5783 | 1540.9411 | 1612.8025 | 800.1887 | 762.6342 | 978.0501 | 913.4543 | 480.0914 | 917.8555 | 1075.2154 | 1399.0501 | 679.233 | 689.3541 | 684.2714 | 715.9883 | 470.1357 | 904.6372 | 1286.0296 | 1441.1858 | 955.7228 | 848.3628 | 1082.4426 | 1688.1416 | 918.8113 | 898.82 | 1198.2271 | 1134.2213 | 936.0443 | 872.8526 | 1182.7345 | 1159.1416 | 1044.9086 | 704.5399 | 800.9616 | 752.9971 | 686.0827 | 926.2714 | 967.9027 | 854.8171 | 636.0413 | 610.9027 | 784.8555 | 769.4484 | 563.7257 | 1706.4898 | 2499.3833 | 2376.7257 | 1680.7846 | 1820.1741 | 2249.6459 | 2169.7228 | 1635.8584 | 328.6342 | 232.8672 | 174.7699 | 328.6342 | 353.9144 | 393.3157 | 368.82 | 446.8643 | 419.8672 | 393.2685 | 391.6814 | 390.8672 | 232.6372 | 236.6756 | 226.8171 | 275.41 | 296 | 309.4071 | 260.4455 | 237.2301 | 1194.4985 | 1277.5783 | 1385.6017 | 670.3215 | 1107.18 | 1504.9645 | 1482.6697 | 850.5457 | 1812.5486 | 1805.3541 | 1961.8998 | 1265.0914
590 | 8.4844 | 736.2679 | 1190.8028 | 1452.563 | 780.4789 | 748.2679 | 969.324 | 1050.3654 | 664.7187 | 637.5915 | 435.7606 | 438.789 | 579.8167 | 997.5216 | 780.9154 | 862.1411 | 864.2959 | 886.6618 | 808.2679 | 841.8171 | 783.1829 | 813.7744 | 1691.3939 | 1855.8878 | 695.4366 | 1157.0984 | 1436.4504 | 1451.7321 | 1102.6195 | 562.1264 | 702.3801 | 717.0423 | 536.7325 | 716.2252 | 839.9715 | 948.2817 | 295.7886 | 1509.2118 | 1948.3378 | 1835.7183 | 1073.6337 | 774.2817 | 865.1972 | 929.2537 | 509.028 | 247.4789 | 246.9293 | 278.5915 | 520.4789 | 1541.4228 | 1615.1691 | 1689.1553 | 761.676 | 789.1264 | 950.5211 | 908.2959 | 472.2114 | 968.352 | 1138.2537 | 1378.4646 | 666.3516 | 728.0004 | 700.6195 | 773.9439 | 479.2675 | 913.6476 | 1377.7606 | 1426.8589 | 992.2955 | 858.2675 | 1123.5077 | 1675.9573 | 968.8167 | 890.8732 | 1202.1126 | 1163.7325 | 942.2817 | 937.2394 | 1154.887 | 1151.7459 | 1042.7325 | 731.4931 | 769.0984 | 775.9858 | 728.2537 | 931.6057 | 1004.0423 | 858.1268 | 662.3943 | 631.4793 | 823.3805 | 811.9435 | 570.1268 | 1770.3098 | 2387.8732 | 2387.676 | 1643.8443 | 1852.3102 | 2197.352 | 2183.789 | 1656.4228 | 335.2537 | 211.5211 | 191.2394 | 335.2537 | 333.2955 | 399.3805 | 358.4789 | 439.648 | 401.6337 | 409.1411 | 398.7886 | 368.0846 | 222.2537 | 260.3102 | 232.7606 | 280.0561 | 294.324 | 316.9297 | 302.465 | 219.5634 | 1200.3663 | 1360.0984 | 1349.098 | 657.7041 | 1108.4228 | 1469.7459 | 1544.3805 | 862.1687 | 1782.2679 | 1856.2114 | 2022.9435 | 1273.9715
591 | 8.4954 | 703.7827 | 1181.9656 | 1479.0933 | 772.4346 | 716.0344 | 976.3344 | 1112.7452 | 635.1515 | 646.7345 | 439.531 | 415.355 | 608.2311 | 955.3413 | 796.3137 | 839.7896 | 852.8172 | 889.1002 | 792.6481 | 830.0758 | 795.176 | 854.7965 | 1753.7552 | 1794.5172 | 705.2862 | 1199.0895 | 1482.293 | 1496.5517 | 1129.4966 | 597.4346 | 733.7827 | 709.469 | 526.0413 | 751.355 | 875.4828 | 943.7345 | 299.9242 | 1447.2999 | 1978.7552 | 1835.2311 | 1089.8034 | 758.1553 | 866.7345 | 926.1033 | 528.5241 | 264.355 | 272.3932 | 284.9656 | 518.7276 | 1525.3481 | 1613.824 | 1682.0238 | 783.6136 | 821.9174 | 965.3619 | 901.3481 | 461.2793 | 978.4966 | 1140.8929 | 1409.5723 | 742.8309 | 698.6687 | 736.3068 | 744.5172 | 498.0138 | 955.4897 | 1367.4346 | 1458.2724 | 1049.1653 | 862.1622 | 1061.1928 | 1733 | 992.9518 | 901.7621 | 1219.2999 | 1175.948 | 959.6343 | 936.2379 | 1195.9793 | 1204.6174 | 1039.3413 | 721.8034 | 804.5448 | 793.9931 | 721.0758 | 927.7758 | 1018.6205 | 855.7965 | 685.6068 | 587.8103 | 805.8172 | 821.4139 | 563.0138 | 1770.9105 | 2408.8309 | 2385.7001 | 1715.8003 | 1780.2831 | 2232.6687 | 2183.5134 | 1627.5103 | 323.5448 | 212.0138 | 190.9931 | 323.5448 | 363.9862 | 378.293 | 364.0275 | 434.0344 | 408.4897 | 396.1033 | 414 | 375.3413 | 210.7965 | 225.0344 | 225.2035 | 321.062 | 285.9793 | 299.469 | 267.9931 | 220.0344 | 1210.8791 | 1372.3864 | 1414.3795 | 681.2655 | 1086.5585 | 1539.4828 | 1545.6963 | 906.0138 | 1743.4897 | 1859.3757 | 2034.176 | 1308.7276
592 | 8.5064 | 705.8554 | 1189.1771 | 1416.2803 | 748.7522 | 730.472 | 1016.6962 | 1025.2331 | 701 | 643.9602 | 458.0162 | 451.7595 | 608.1844 | 994.472 | 811.5914 | 859.9041 | 871.1755 | 834.8245 | 852.5118 | 850.8554 | 762.472 | 856.5191 | 1718.2462 | 1816.2567 | 716.5118 | 1225.0634 | 1498.6239 | 1515.0398 | 1142.5207 | 568.208 | 737.4395 | 693.6239 | 552.1844 | 769.8067 | 864.3038 | 929.6076 | 274.6564 | 1471.1121 | 1954.3347 | 1827.528 | 1109.5678 | 729.9366 | 860.264 | 964.1357 | 535.0398 | 299.7035 | 228.3761 | 262.6076 | 512.1682 | 1562.7197 | 1653.3924 | 1788.3274 | 827.4557 | 786.792 | 1021.9204 | 945.5841 | 486.4086 | 982.5443 | 1182.143 | 1439.4159 | 751.8627 | 751.143 | 748.0634 | 755.4322 | 507.6962 | 946.7758 | 1333.7197 | 1454.5353 | 1091.9586 | 843.0089 | 1135.7595 | 1738.8083 | 975.4322 | 909.9041 | 1232.8156 | 1240.0869 | 1002.6873 | 938.1283 | 1185.4159 | 1155.7124 | 1086.1283 | 729.6873 | 822.6799 | 790.208 | 757.9602 | 960.385 | 1072.0561 | 882.6003 | 726.9838 | 612.4159 | 830.5118 | 794.5605 | 584.8408 | 1746.5207 | 2451.6564 | 2364.528 | 1678.1064 | 1711.0812 | 2308.6401 | 2283.7433 | 1645.4647 | 335.6076 | 223.5443 | 189.5841 | 335.6076 | 363.8083 | 402.5443 | 383.6564 | 454.5443 | 403.208 | 448.9204 | 413.208 | 412.6239 | 238.9204 | 238.4159 | 206.5841 | 341.6637 | 281.2242 | 292.3363 | 261.0398 | 229.1917 | 1246.671 | 1332.2405 | 1398.6978 | 693.0162 | 1112.528 | 1543.0885 | 1620.1755 | 915.6962 | 1762.4971 | 1911.6239 | 1989.0634 | 1296.0959
593 | 8.5174 | 676.1446 | 1223.3696 | 1437.6313 | 763.1052 | 742.8152 | 1020.645 | 1077.29 | 707.6313 | 677.1043 | 469.7896 | 431.5261 | 632.6578 | 1007.3678 | 773.1317 | 851.6981 | 846.8024 | 870.5261 | 852.4337 | 824.4602 | 759.8948 | 818.2626 | 1636.4858 | 1825.8829 | 710.6313 | 1215.8015 | 1482.3294 | 1507.3156 | 1176.9213 | 562.9607 | 712.9735 | 705.5124 | 583.0128 | 709.1967 | 862.8417 | 909.5526 | 304.8417 | 1486.3943 | 1897.4071 | 1859.3934 | 1111.8811 | 773.1839 | 881.4465 | 935.6057 | 533.3943 | 279.8417 | 241.6313 | 243.1052 | 513.8161 | 1568.7228 | 1657.9744 | 1745.5398 | 820.5654 | 805.3019 | 1021.6194 | 946.9607 | 514.645 | 993.3028 | 1114.0531 | 1446.0522 | 691.4337 | 703.2232 | 737.6963 | 763.0659 | 524.6313 | 952.1308 | 1334.473 | 1443.5252 | 1026.7493 | 886.3294 | 1100.3294 | 1751.408 | 972.0137 | 917.4474 | 1217.4602 | 1187.5654 | 963.8683 | 957.1711 | 1170.7768 | 1183.3156 | 1108.2104 | 706.065 | 829.4995 | 793.5261 | 786.4071 | 1017.237 | 1087.9863 | 907.6313 | 718.5261 | 601.645 | 801.1455 | 824.2763 | 610.4876 | 1783.1317 | 2468.0265 | 2403.8536 | 1775.2122 | 1751.7118 | 2321.2498 | 2245.3815 | 1695.9863 | 324.118 | 228.5005 | 195.3815 | 324.118 | 389.3943 | 405.29 | 398.3687 | 461.1583 | 404.592 | 479.2891 | 407.6844 | 399.0924 | 254.645 | 243.6706 | 214.592 | 331.1839 | 289.6844 | 311.3815 | 266.5782 | 211.8948 | 1182.5398 | 1356.8948 | 1500.4474 | 694.5663 | 1124.7759 | 1589.1967 | 1597.4602 | 909.9744 | 1828.2763 | 1904.7237 | 1973.9991 | 1279.9872
594 | 8.5284 | 689.3088 | 1174.6648 | 1438.2615 | 778.4096 | 790.8324 | 1015.5438 | 1071.9937 | 717.4965 | 735 | 470.0605 | 437.3422 | 637.8658 | 1030.5703 | 769.3491 | 832.9534 | 855.6509 | 900.0069 | 876.6307 | 841.4026 | 776.3359 | 846.0737 | 1703.1947 | 1771.242 | 730.3088 | 1280.5167 | 1460.7517 | 1507.6307 | 1163.0202 | 560.7719 | 727.1544 | 732.644 | 590.1474 | 749.8658 | 890.3289 | 912.7316 | 323.0939 | 1515.9061 | 1969.5507 | 1919.6641 | 1144.7316 | 775.5501 | 913.0668 | 926.8727 | 546.4562 | 302.7114 | 260.443 | 249.1745 | 493.557 | 1628.7719 | 1651.4972 | 1762.363 | 851.3693 | 821.4493 | 960.6912 | 946.7921 | 495.9798 | 1003.2218 | 1088.5772 | 1466.6307 | 732.6578 | 732.0202 | 791.1474 | 774.4499 | 535.3623 | 990.0737 | 1368.0737 | 1529.1613 | 1082.1412 | 870.235 | 1068.6509 | 1760.1815 | 935.1141 | 897.8185 | 1228.6307 | 1219.5236 | 991.3157 | 940.8588 | 1161.0271 | 1222.5911 | 1120.6377 | 758.8859 | 859.8859 | 813.4833 | 811.7984 | 1008.2684 | 1120.2684 | 948.0876 | 756.3762 | 603.4701 | 767.9666 | 820.3825 | 564.8859 | 1783.1008 | 2466.2552 | 2496.3623 | 1689.7587 | 1705.3025 | 2302.3888 | 2251.8859 | 1742.4096 | 350.8859 | 213.121 | 200.1544 | 350.8859 | 394.0869 | 387.0069 | 386.7719 | 440.5974 | 415.1544 | 484.5501 | 398.7114 | 389.6578 | 238.8658 | 266.6912 | 222.2684 | 339.8992 | 285.0403 | 317.3088 | 305.3825 | 219.3894 | 1183.2016 | 1349.6711 | 1493.8053 | 658.5369 | 1168.0069 | 1637.3693 | 1618.443 | 886.4701 | 1842.8526 | 1916.9868 | 2018.6106 | 1279.2016
595 | 8.5394 | 710.947 | 1236.9115 | 1403.3186 | 812.3097 | 792.6194 | 1022.6371 | 1116.0265 | 696.8938 | 739.2124 | 480.4867 | 420.8142 | 646.9204 | 1000.1151 | 773.5663 | 882.3716 | 861.062 | 933.0885 | 854.4956 | 832.5929 | 802.6371 | 818.7699 | 1699.4513 | 1815.9557 | 748.9381 | 1269.4071 | 1471.8053 | 1489.4071 | 1180.0354 | 574.0973 | 714.053 | 694.1416 | 573.7523 | 760.7257 | 886.5575 | 898.3451 | 316.1151 | 1515.3628 | 2033.1947 | 1884.6372 | 1163.7434 | 736.5044 | 862.0089 | 956.6725 | 528.2301 | 304.708 | 268.8142 | 257.9027 | 510.7611 | 1645.1062 | 1727.5663 | 1864.3363 | 865.3805 | 772.5487 | 979.9735 | 964.1947 | 491 | 1060.4071 | 1099.3275 | 1466.7611 | 773.2567 | 746.6284 | 765.7257 | 829.7434 | 519 | 973.9027 | 1362.4337 | 1563.8849 | 1102.2566 | 875.8495 | 1068.947 | 1770.9999 | 943.3452 | 874.062 | 1226.354 | 1233.0708 | 1048.4425 | 914.4248 | 1207.5133 | 1276.6371 | 1149.9381 | 761.1858 | 863.6903 | 841.1416 | 767.1151 | 1051.54 | 1136.1594 | 989.6548 | 814.6018 | 640.8761 | 805.2301 | 810.8053 | 581.9292 | 1819.1328 | 2485.5752 | 2488.4248 | 1753.7523 | 1774.8849 | 2228.3186 | 2279.7611 | 1793.4603 | 367.0266 | 246.9115 | 200.5929 | 367.0266 | 363.6549 | 423.0973 | 388.9646 | 443.3894 | 427.0266 | 458.4425 | 405.2213 | 408.292 | 241.3009 | 262.2035 | 227.0266 | 313.7965 | 301.9204 | 326.6106 | 306.6372 | 233.2654 | 1249.5487 | 1326.0619 | 1496.947 | 664.4425 | 1200.4867 | 1665.2213 | 1633.4337 | 930.4956 | 1847.2123 | 1952.3097 | 2010.1505 | 1351.5665
596 | 8.5504 | 722.0063 | 1275.1255 | 1425.8294 | 818.4341 | 748.417 | 992.417 | 1064.2635 | 681.6652 | 750.8791 | 480.3348 | 434.8186 | 670.8294 | 1032.9829 | 756.0821 | 867.8294 | 903.4233 | 923.0605 | 862.2744 | 845.1643 | 772.417 | 815.3565 | 1680.7302 | 1804.1751 | 754.5722 | 1280.0108 | 1473.4946 | 1503.4558 | 1197.9783 | 590.4774 | 701.1146 | 731.6977 | 599.9675 | 793.2139 | 868.6589 | 876.6092 | 324.5551 | 1505.417 | 2002.0668 | 1906.1706 | 1215.8511 | 750.3348 | 874.324 | 931.2419 | 551.0605 | 296.5939 | 268.7256 | 263.1318 | 538.6435 | 1652.8186 | 1716.972 | 1903.3348 | 858.1426 | 802.5442 | 1020.7753 | 955.9287 | 503.8403 | 1056.9179 | 1155.1363 | 1507.4837 | 805.4729 | 763.0496 | 778.1706 | 859.3023 | 547.4991 | 976.3132 | 1404.6589 | 1566.093 | 1075.2031 | 839.4449 | 1105.6868 | 1719.2464 | 979.6264 | 925.8186 | 1260.8574 | 1208.4341 | 1060.704 | 936.1922 | 1210.1706 | 1215.0993 | 1172.4837 | 766.2852 | 855.6868 | 870.8682 | 782.7581 | 1160.7473 | 1177.676 | 971.2247 | 812.5551 | 610.4124 | 815.2202 | 821.2635 | 592.1597 | 1857.5271 | 2483.8511 | 2508 | 1793.7473 | 1773.3348 | 2269.9332 | 2310.4621 | 1824.4016 | 382.676 | 251.6155 | 200.6372 | 382.676 | 411.0713 | 419.7473 | 361.6264 | 428.1101 | 451.4449 | 500.769 | 433.2744 | 400.4341 | 250.1426 | 268.1318 | 244.5551 | 315.6977 | 319.2527 | 327.7969 | 325.9783 | 222.5984 | 1282.3628 | 1319.8015 | 1514.8962 | 690.4837 | 1184.0496 | 1693.9007 | 1689.4386 | 940.7798 | 1809.9441 | 1976.5984 | 2082.8836 | 1385.2094
597 | 8.5614 | 694.9431 | 1304.1749 | 1476.5232 | 805.8869 | 789.6981 | 1033.3727 | 1084.5094 | 678.5281 | 756.6273 | 487.0521 | 445.882 | 692.882 | 1082.2547 | 777.6273 | 904.5281 | 905.5989 | 910.585 | 882.7689 | 892.4858 | 828.0187 | 850.8869 | 1734.0895 | 1859.882 | 718.5753 | 1292.9528 | 1486.2262 | 1513.8633 | 1196.179 | 625.467 | 756.2783 | 751.764 | 569.9292 | 806.9813 | 914.0659 | 929.533 | 319.6037 | 1544.4199 | 1956.9951 | 1923.4996 | 1233.8446 | 782.764 | 904.2262 | 922.0236 | 562.6697 | 269.5566 | 259.297 | 269.6745 | 503.1131 | 1675.9243 | 1749.7217 | 1934.1367 | 888.3678 | 790.5094 | 988.533 | 950.5801 | 525.1652 | 1044.4858 | 1211.1228 | 1506.297 | 800.1318 | 768.2262 | 776.297 | 864.9479 | 609.0236 | 965.9341 | 1430.8633 | 1615.2262 | 1098.8112 | 868.1554 | 1100.118 | 1783.2311 | 1004.1082 | 918.0049 | 1250.1603 | 1188.0757 | 1018.2311 | 938.3963 | 1190.0708 | 1201.0423 | 1176.179 | 797.7689 | 857.2783 | 850.3303 | 807.9243 | 1153.7689 | 1170.2311 | 1029.83 | 820.297 | 617.2213 | 813.9667 | 845.0895 | 566.9292 | 1863.1936 | 2537.3206 | 2523.9479 | 1783.8397 | 1841.5655 | 2332.2547 | 2293.1554 | 1781.4906 | 365.1416 | 242 | 222.4386 | 365.1416 | 396.2408 | 427.7404 | 375.3678 | 433.5045 | 465.8348 | 515.118 | 449.5378 | 409.3678 | 278.0895 | 273.6981 | 235.6981 | 332.1839 | 322.3255 | 330.5517 | 304 | 245.0049 | 1265.118 | 1383.0659 | 1456.0757 | 672.0472 | 1195.0846 | 1729.4386 | 1749.118 | 983.6933 | 1869.585 | 1998.3539 | 2151.4199 | 1361.2498
598 | 8.5724 | 744.8407 | 1247.1992 | 1524.0457 | 796.795 | 798.1461 | 1066.478 | 1099.6888 | 723.3909 | 758.9867 | 480.3452 | 444.5162 | 708.4041 | 1088.3245 | 786.4041 | 929.4631 | 928.1784 | 878.1917 | 864.205 | 895.8407 | 860.7419 | 841.795 | 1729.3303 | 1880.1328 | 759.0531 | 1292.2714 | 1542.239 | 1502.0664 | 1248.1328 | 605.2183 | 765.0531 | 759.5029 | 575.3245 | 772.9071 | 941.0531 | 933.2315 | 319.9336 | 1514.3701 | 1988.0266 | 1950.4158 | 1192.3643 | 805.7021 | 939.6357 | 929.795 | 595.3909 | 276.2979 | 277.4896 | 270.0133 | 535.1726 | 1737.683 | 1757.1593 | 1980.7228 | 925.8216 | 816.8348 | 1020.6033 | 950.9071 | 512.6888 | 1063.0398 | 1183.1461 | 1545.7685 | 833.1652 | 796.205 | 809.0133 | 899.6888 | 587.0722 | 952.1195 | 1422.7817 | 1655.3643 | 1141.0075 | 903.8938 | 1109.2523 | 1813.2788 | 1053.6888 | 895.0797 | 1275.4705 | 1155.3643 | 1026.6622 | 944.1461 | 1193.1195 | 1214.1784 | 1212.9336 | 787.9867 | 869.4307 | 822.3378 | 852.4573 | 1146.351 | 1212.1004 | 1089.5502 | 864.1593 | 655.966 | 843.9394 | 847.0855 | 578.7419 | 1841.3054 | 2565.9867 | 2571.8614 | 1767.3112 | 1918.0988 | 2341.3643 | 2319.7743 | 1786.239 | 354.351 | 232.205 | 237.2714 | 354.351 | 359.1593 | 448.3112 | 404.3776 | 484.7817 | 468.8539 | 504.3245 | 435.0398 | 424.8672 | 287.1784 | 280.795 | 242.795 | 343.9867 | 326.0398 | 362.9203 | 316.1593 | 220.7286 | 1261.0797 | 1368.171 | 1432.4838 | 692.6564 | 1222.0324 | 1781.4249 | 1758.9278 | 997.5029 | 1853.0664 | 1963.6091 | 2138.9336 | 1418.2788
599 | 8.5834 | 735.3521 | 1262.6479 | 1483.4504 | 836.653 | 814.7492 | 1106.4955 | 1115.6018 | 748.4524 | 793.7553 | 546.002 | 441.2528 | 719 | 1080.349 | 798.4002 | 911.6018 | 920.4022 | 873.8024 | 850.6008 | 862.8987 | 873.7503 | 855.7492 | 1698.9529 | 1910.1515 | 769.8004 | 1290.3511 | 1555.5968 | 1524.9519 | 1253.2979 | 582.3992 | 764.5988 | 740.2508 | 579.2508 | 758.6499 | 931.4975 | 930.6018 | 319.8014 | 1495.5537 | 1989.1996 | 1873.5015 | 1214.8526 | 802.2487 | 954.2528 | 964.7523 | 590.6981 | 279.4012 | 268.5025 | 270.2497 | 560.2999 | 1756.7452 | 1803.0532 | 2025.348 | 957.4012 | 830.7984 | 1032.996 | 935.9498 | 520.9007 | 1045.5466 | 1199.0491 | 1566.6008 | 818.5015 | 807.1043 | 803.6489 | 920.1525 | 562.7021 | 964.0491 | 1462.0532 | 1640.3972 | 1171.6971 | 908.653 | 1128.7974 | 1876.3531 | 1083.9539 | 904.5496 | 1289.6991 | 1167.001 | 1016.5486 | 963.1996 | 1213.4504 | 1176.6479 | 1207.5507 | 797.5517 | 891.7011 | 857.3049 | 884.652 | 1163.3521 | 1276.6499 | 1127.6008 | 896.0512 | 599.5998 | 783.2497 | 817.3511 | 594.5507 | 1868.7974 | 2595.1545 | 2599.2497 | 1785.7031 | 1899.1023 | 2365.9529 | 2322.7553 | 1824.1003 | 371.3521 | 236.8034 | 223.7994 | 371.3521 | 391.4012 | 441.4995 | 409.9498 | 506.5507 | 453.651 | 508.6008 | 444.5005 | 422.5025 | 264.7001 | 300.7001 | 267.2508 | 340.5998 | 338.651 | 368.1525 | 355.4022 | 229.2999 | 1281.4012 | 1317.1033 | 1483.004 | 729.4483 | 1201.1535 | 1847.4504 | 1824.2538 | 1007.655 | 1857.7492 | 1964.1013 | 2145.4524 | 1463.8506
600 | 8.5944 | 751.4921 | 1248.6819 | 1482.2025 | 865.2168 | 804.6377 | 1037.8544 | 1135.3766 | 764.5063 | 857.2896 | 544.0997 | 468.6662 | 736.3924 | 1062.9272 | 816.5807 | 914.7089 | 935.5934 | 917.3196 | 876.943 | 882.9588 | 894.7547 | 832.231 | 1742.1598 | 1910.1583 | 770.2896 | 1284.0712 | 1527.0443 | 1567.8417 | 1240.856 | 564.0285 | 755.5221 | 748.1883 | 606.3924 | 789.1614 | 889.5791 | 949.6519 | 330.7532 | 1515.8244 | 1986.8987 | 1893 | 1242.9272 | 771.5063 | 970.0712 | 984.9415 | 572.5221 | 288.1155 | 303.174 | 269.8987 | 556.4636 | 1733.7705 | 1881.3212 | 2063.4097 | 982.7832 | 804.9272 | 991.8702 | 955.0158 | 534.3481 | 1030.9873 | 1200.3196 | 1590.4066 | 851.4066 | 864.4494 | 821.5237 | 971.2041 | 604.856 | 983.0743 | 1473.288 | 1628.2769 | 1161.1756 | 949.5364 | 1112.7547 | 1885.201 | 1126.0285 | 914.1313 | 1248.6504 | 1170.9415 | 1018.5664 | 975.7547 | 1240.8275 | 1224.6423 | 1195.3465 | 837.1171 | 909.9858 | 912.6662 | 912.8987 | 1214.6392 | 1315.1329 | 1159.3781 | 911.7247 | 607.1456 | 806.4509 | 849.8417 | 619.3054 | 1869.4224 | 2665.6819 | 2597.087 | 1829.6234 | 1898.5633 | 2362.7801 | 2394.2611 | 1818.5791 | 404.1598 | 272.6662 | 217.087 | 404.1598 | 416.0585 | 441.5221 | 429.7405 | 520.0728 | 469.9858 | 500.1583 | 471.2911 | 455 | 268.8987 | 302.7247 | 278.4494 | 346.6962 | 355.3481 | 399.2753 | 366.97 | 226.5506 | 1292.6519 | 1353.1155 | 1539.3481 | 710.8987 | 1271.0015 | 1840.4051 | 1867.9272 | 1056.057 | 1833.5063 | 2000.2041 | 2157.5206 | 1464.4779
601 | 8.6054 | 728.5044 | 1284.8614 | 1489.9824 | 838.6342 | 796.1327 | 1038.357 | 1139.4897 | 729.2596 | 851.9971 | 505.8643 | 462.9912 | 769.2478 | 1056.9971 | 836.882 | 889.1151 | 905.6312 | 930.0088 | 899.0059 | 921.0265 | 899.0265 | 828.2301 | 1741.3864 | 1870.2566 | 762.3746 | 1257.6165 | 1561.6136 | 1605.121 | 1261.3805 | 565.7346 | 780.1121 | 752.118 | 618.643 | 809.9085 | 890.2331 | 933.0088 | 322.4956 | 1509.2096 | 1982.6342 | 1949.5695 | 1231.3776 | 761.3599 | 941.7434 | 973.4956 | 583.2507 | 280.8702 | 301.7551 | 277.9971 | 553.8732 | 1793.3835 | 1926.0206 | 2170.879 | 1005.9941 | 801.9941 | 1020.7404 | 979.6371 | 542.3746 | 1076.0059 | 1193.5279 | 1652.2125 | 894.4809 | 861.7551 | 872.879 | 1007.9971 | 648.6077 | 1024.761 | 1452.0945 | 1668.0029 | 1202.6312 | 924.404 | 1134.2596 | 1879.3276 | 1110.1268 | 921.7669 | 1190.6312 | 1168.1121 | 1054.8732 | 985.2713 | 1261.5191 | 1348.1327 | 1182.2272 | 875.7434 | 936.4809 | 915.2331 | 905.2625 | 1262.8702 | 1384 | 1210.4868 | 922.7404 | 624.2507 | 855.9971 | 869.1386 | 640.5074 | 1900.404 | 2664.7728 | 2620.354 | 1841.6195 | 1844.6254 | 2333.5842 | 2437.4632 | 1842.0857 | 401.8879 | 263.2449 | 226.4926 | 401.8879 | 417.3893 | 458.2449 | 467.2507 | 521.5074 | 483.7434 | 498.0945 | 488.643 | 462.118 | 278.4956 | 291.6371 | 273.882 | 372.8643 | 369.7434 | 398.3746 | 359.8526 | 240.4838 | 1296.9882 | 1339.8761 | 1538.0088 | 736.2301 | 1286.1533 | 1866.4544 | 1897.9618 | 1018.2596 | 1847.3364 | 2023.8849 | 2154.2213 | 1442.1357
602 | 8.6164 | 735.3162 | 1271.6292 | 1543.6023 | 833.4469 | 782.6131 | 1073.6431 | 1162.8039 | 719.5777 | 853.1308 | 537.5423 | 463.2969 | 784.6077 | 1066.6431 | 814.94 | 917.2208 | 920.9239 | 933.1854 | 902.5423 | 906.5669 | 881.8585 | 860.8392 | 1733.9946 | 1891.1146 | 749.4577 | 1277.8039 | 1560.88 | 1611.6131 | 1274.0546 | 588.5177 | 788.8446 | 784.8939 | 585.2915 | 781.6323 | 921.7085 | 939.2808 | 330.3869 | 1572.1961 | 1992.15 | 2038.97 | 1234.7385 | 793.9646 | 945.2615 | 1003.0546 | 587.03 | 280.1008 | 284.1008 | 279.1308 | 572.09 | 1795.6377 | 1915.3762 | 2149.6838 | 1003.6185 | 836.7631 | 1037.7738 | 982.6023 | 535.2615 | 1065.2262 | 1165.7331 | 1673.0515 | 905.3323 | 886.6623 | 872.5777 | 1038.15 | 671.8092 | 1014.5123 | 1502.8392 | 1688.15 | 1222.5669 | 895.5069 | 1147.0193 | 1932.5284 | 1110.0954 | 885.6485 | 1219.4662 | 1217.0846 | 1081.2154 | 978.7931 | 1244.8639 | 1326.1008 | 1226.4169 | 881.97 | 972.0954 | 933.94 | 880.5177 | 1284.5423 | 1414.5669 | 1244.2208 | 998.9731 | 624.5477 | 849.0054 | 877.3408 | 647.7331 | 1877.3108 | 2620.1961 | 2663.3515 | 1900.9131 | 1847.8692 | 2387.5231 | 2488.3923 | 1918.2862 | 386.9646 | 270.4523 | 245.5777 | 386.9646 | 410.0246 | 470.4823 | 448.9754 | 510.5477 | 483.0054 | 551.5477 | 480.4416 | 457.7492 | 290.2562 | 273.7085 | 273.6077 | 380.5531 | 381.3869 | 399.7738 | 384.8446 | 264.5177 | 1295.7192 | 1369.3408 | 1531.1254 | 749.1061 | 1242.2562 | 1912.2723 | 1964.5477 | 1017.09 | 1924.2208 | 1982.2369 | 2216.5423 | 1457.2754
603 | 8.6274 | 750.4149 | 1233.5615 | 1574.2134 | 869.9607 | 792.7788 | 1057.0471 | 1164.0157 | 748.3639 | 850.1937 | 526.2644 | 470.7395 | 822.1623 | 1095.1584 | 801.3953 | 943.7827 | 969.1859 | 979.9647 | 922.0039 | 958.5929 | 912.0275 | 847.0236 | 1763.7788 | 1905.6715 | 739.5772 | 1266.2408 | 1570.5183 | 1615.7906 | 1314.589 | 581.8024 | 766.8024 | 783.6676 | 596.5694 | 799.4542 | 945.5615 | 968.6086 | 330.6008 | 1568.9882 | 2041.5811 | 2012.6361 | 1268.1584 | 795.0196 | 940.1859 | 1021.233 | 592.2055 | 270.7906 | 273.5929 | 270.2055 | 584.822 | 1817.6086 | 1942.0236 | 2174.1073 | 1005.1584 | 836.0786 | 1028.6204 | 1012.4149 | 536.9725 | 1080.767 | 1204.9843 | 1601.4071 | 870.6008 | 947.9843 | 900.9647 | 1088.3796 | 671.5811 | 1026.4031 | 1501.3992 | 1742.3717 | 1256.6283 | 928 | 1167.8495 | 1918.9411 | 1096.6361 | 889.1701 | 1275.2134 | 1249.2212 | 1141.1544 | 1024.4031 | 1285.5694 | 1296.4306 | 1214.2408 | 877.1937 | 982.589 | 936.3639 | 916.9175 | 1306 | 1451.0236 | 1305.9136 | 1062.8613 | 635.3835 | 846.3678 | 926.2016 | 654.2487 | 1887.4621 | 2615.7906 | 2707.7356 | 1923.3115 | 1857.1937 | 2414.0995 | 2452.8338 | 1906.2526 | 415.5654 | 274.7866 | 242.8259 | 415.5654 | 424.6244 | 453.2448 | 459.5262 | 534.1584 | 491.1466 | 565.178 | 507.2173 | 447.5694 | 299.3953 | 295.9647 | 302.3796 | 362.4031 | 393.178 | 405.3914 | 374.3796 | 260.5969 | 1255.2134 | 1434.1701 | 1572.7434 | 717.9961 | 1288.123 | 1879.9647 | 2015.7042 | 1048.9843 | 1887.3076 | 1967.3521 | 2247.9804 | 1506.4071
604 | 8.6384 | 729.1239 | 1264.2345 | 1595.6859 | 883.1283 | 829.1638 | 1051.0399 | 1185.7435 | 753.7787 | 868.1682 | 520.2081 | 505.0663 | 838.2433 | 1133.5841 | 821.0487 | 943.0088 | 982.0575 | 1000.0044 | 896.7787 | 973.8761 | 910.3408 | 859.8629 | 1766.8318 | 1875.2788 | 787.8629 | 1266.6859 | 1622.3009 | 1650.5709 | 1321.7655 | 601.3983 | 748.1283 | 741.8761 | 610.1239 | 769.5222 | 949.4867 | 973.5222 | 325.4734 | 1586.6416 | 2049.8274 | 2000.2257 | 1278.1902 | 802.2213 | 942.7699 | 983.6548 | 594.9956 | 289.1106 | 269.5354 | 280.4027 | 571.292 | 1824.1682 | 1937.5797 | 2239.2389 | 1040.292 | 810.155 | 1021.2301 | 1027.7479 | 548.9469 | 1105.177 | 1214.9425 | 1608.3452 | 868.354 | 942.3053 | 910.7168 | 1121.2832 | 700.8142 | 1031.8186 | 1508.1726 | 1770.1726 | 1239.1151 | 936.6416 | 1153.0443 | 1972.6416 | 1093.6903 | 927.6372 | 1237.8406 | 1246.5797 | 1130.3229 | 1031.4646 | 1282.2521 | 1313.3275 | 1209.3364 | 882 | 1012.3983 | 977.5797 | 940.4955 | 1315.8761 | 1456.0443 | 1380.6416 | 1052.2125 | 645.531 | 859.0663 | 894.9601 | 624.4071 | 1836.708 | 2645.6328 | 2778.3983 | 1849.1106 | 1883.3983 | 2493.469 | 2482.9648 | 1937.1905 | 442 | 285 | 255.9204 | 442 | 405.5885 | 446.5709 | 504.0663 | 558.1814 | 532.4115 | 591.4646 | 506.7567 | 464.0044 | 304.646 | 325.469 | 299.4823 | 377.2832 | 412.8805 | 439.1594 | 410.1638 | 261.7655 | 1272.9824 | 1433.1327 | 1613.1195 | 746.5133 | 1308.3185 | 1904.531 | 2072.9513 | 1042.4823 | 1812.8186 | 2025.3983 | 2293.8629 | 1487.0088
605 | 8.6494 | 718.8142 | 1275.7473 | 1629.5575 | 871.2901 | 876.647 | 1074.9292 | 1222.1337 | 756.3275 | 882.3382 | 564.2005 | 484.1858 | 837.3757 | 1123.9292 | 841.881 | 923.8476 | 1019.3275 | 976.8623 | 859.9144 | 1010.885 | 947.3235 | 872.6391 | 1708.1631 | 1877.7915 | 819.7099 | 1329.3757 | 1591.3195 | 1680.2527 | 1329.7473 | 638.3569 | 731.2045 | 737.7767 | 598.119 | 785.1524 | 921.4759 | 993.3904 | 323.6618 | 1588.8289 | 2072.7847 | 1999.5388 | 1249.9666 | 803.4866 | 906.8436 | 975.7473 | 604.1524 | 305 | 261.0856 | 287.8623 | 578.9666 | 1838.762 | 1954.4572 | 2278.0414 | 1046.2714 | 831.0294 | 1032.1524 | 1059.6431 | 545.3235 | 1103.1524 | 1234.885 | 1646.766 | 879.9853 | 928.8663 | 919.4799 | 1153.2045 | 732.1711 | 1045.9666 | 1517 | 1780.2714 | 1252.6283 | 940.5241 | 1174.9813 | 1981.6096 | 1126.4572 | 945.3382 | 1215.9519 | 1271.0856 | 1077.2901 | 1021.2674 | 1258.8329 | 1377.3569 | 1259.2901 | 914.2085 | 1045.9666 | 1024.5468 | 911.123 | 1334.238 | 1521.2085 | 1420.5468 | 1120.1043 | 630.1337 | 842.8476 | 849.4572 | 626.9479 | 1848.5281 | 2670.3717 | 2795.0147 | 1903.9893 | 1885.1818 | 2465.6391 | 2562.5762 | 2029.6765 | 430.1337 | 285.8476 | 299.6096 | 430.1337 | 424.9184 | 488.5428 | 481.9144 | 563.9519 | 536.8142 | 601.1858 | 531.7807 | 469.2567 | 319.7139 | 336.6283 | 287.8329 | 396.9144 | 440.3088 | 478 | 444.5094 | 247.2861 | 1305.7955 | 1405.7807 | 1604.762 | 771.5575 | 1264.8663 | 1954.3984 | 2070.3422 | 1045.242 | 1816.3717 | 2046.6765 | 2337.1524 | 1470.3904
606 | 8.6604 | 724.8722 | 1283.3309 | 1624.9022 | 883.203 | 877.5791 | 1059.9022 | 1207.3083 | 794.2782 | 896.06 | 562.6766 | 499.4961 | 890.2179 | 1084.9174 | 849.03 | 906.6843 | 1042.8874 | 941.1204 | 866.5187 | 1004.4965 | 952.5865 | 836.1052 | 1698.917 | 1869.7592 | 823.9322 | 1338.173 | 1644.1349 | 1674.8496 | 1364.3683 | 643.9474 | 738.8196 | 730.8796 | 603.6165 | 791.4135 | 931.2556 | 993.203 | 346.1352 | 1612.0674 | 2082.4213 | 2015.2253 | 1253.1278 | 785.5487 | 892.7518 | 1000.3383 | 601.2556 | 297.5865 | 268.4361 | 287.8796 | 591.7218 | 1796.3686 | 2002.0223 | 2297.5191 | 1051.9248 | 799.1052 | 1040.5939 | 1044.4587 | 561.4887 | 1108.1052 | 1279.9548 | 1658.8648 | 919.4961 | 944.1278 | 954.5339 | 1184.8044 | 762.5713 | 1049.1278 | 1518.3083 | 1794.6465 | 1269.5413 | 915.0452 | 1183.9848 | 1992.2631 | 1133.03 | 942.0526 | 1248.5939 | 1271.0226 | 1057.8047 | 969.8121 | 1278.1278 | 1380.3309 | 1261.173 | 941.8648 | 1051.3083 | 1059.8196 | 962.827 | 1359.1879 | 1564.5639 | 1509.4583 | 1112.346 | 616.6165 | 856.6465 | 883.5035 | 638.6766 | 1833.0604 | 2654.7669 | 2778.6165 | 1987.09 | 1846 | 2483.6162 | 2587.7292 | 2040.5487 | 434.4961 | 298.3383 | 295.8722 | 434.4961 | 462.03 | 493.3083 | 488.0826 | 602.6992 | 538.5113 | 614.7518 | 541 | 501.7744 | 334.3835 | 359.2105 | 322.3909 | 399.5939 | 466.3309 | 494.1352 | 445.8948 | 236.9774 | 1282.0452 | 1419.797 | 1600.3083 | 754.2556 | 1326.7892 | 2024.8196 | 2130.7066 | 1080.203 | 1836.9622 | 2036.1804 | 2280.1808 | 1493.3157
607 | 8.6714 | 733.1741 | 1273.3097 | 1589.23 | 880.1387 | 844.6549 | 1077.587 | 1191.9616 | 773.584 | 926.0354 | 586.6195 | 524.6549 | 919.4808 | 1059.9292 | 883.9706 | 904.0738 | 1038.5576 | 920.6903 | 893.416 | 978.4189 | 938.5162 | 854.7611 | 1732.7876 | 1903.4454 | 822.3038 | 1305.5486 | 1707.2684 | 1708.3157 | 1384.8259 | 613.0973 | 727.4808 | 752.2065 | 606.5516 | 790.2389 | 935.2773 | 985.2065 | 348.6165 | 1635.3392 | 2060.5222 | 2072.7581 | 1294.596 | 798.5132 | 935.3835 | 1021.7257 | 611.1062 | 294.2773 | 285.1416 | 295.3067 | 586.41 | 1731.0649 | 2052.7551 | 2289.3217 | 1090.18 | 833.006 | 1029.2359 | 1045.416 | 574.5192 | 1120.0354 | 1304.2773 | 1642.0354 | 939.2743 | 950.1741 | 956.2449 | 1213 | 811.3511 | 1063.6933 | 1548.2479 | 1820.8643 | 1291.8997 | 920.7641 | 1210.5162 | 2021.416 | 1179.6284 | 952.2773 | 1295.5252 | 1278.59 | 1045.5252 | 943.6549 | 1288.6578 | 1367.6195 | 1261.2803 | 929.9676 | 1079.4543 | 1079.8673 | 977.9646 | 1397.8997 | 1563.1032 | 1580.9322 | 1080.2773 | 627.6224 | 870.7581 | 902.8554 | 672.4838 | 1774.1357 | 2644.8613 | 2790.9676 | 2002.2359 | 1789.9527 | 2525.6813 | 2622.3451 | 2084.8997 | 472.2095 | 317.4838 | 300.1387 | 472.2095 | 467.3776 | 494.1032 | 518.003 | 609.997 | 552.3835 | 643.0354 | 567.006 | 511.1711 | 324.7227 | 350.7876 | 323.9616 | 398.1003 | 462.587 | 523.5192 | 456.8319 | 262.59 | 1273.416 | 1399.0973 | 1608.2773 | 750.6549 | 1342.5073 | 1978.0589 | 2171.2065 | 1056.0649 | 1901.6608 | 2065.006 | 2252.3895 | 1481.5781
608 | 8.6824 | 755.3594 | 1286.1234 | 1568.0783 | 915.7862 | 853.1347 | 1112.8649 | 1187.5841 | 779.9661 | 921.3259 | 582.3033 | 513.7866 | 896.5619 | 1065.2021 | 899.4833 | 944.7414 | 1056.1799 | 925.382 | 917.4381 | 1018.4381 | 916.8766 | 844.0452 | 1715.8988 | 1934.9883 | 806.3481 | 1291.3033 | 1635.6745 | 1736.0787 | 1384.2921 | 583.6854 | 752.9322 | 774.4155 | 622.1234 | 776.0787 | 943.4607 | 969.1799 | 332.91 | 1597.1686 | 2087 | 2119.2917 | 1358.9774 | 791.2808 | 975.0448 | 1023.09 | 626 | 307.528 | 289.7979 | 285.6741 | 599.6289 | 1735.2582 | 2039.528 | 2379.0222 | 1119.0226 | 846.1125 | 1068.1791 | 1080.4942 | 594.6741 | 1112.1012 | 1289.4272 | 1664.0448 | 930.5393 | 937.8092 | 980.0674 | 1238.7975 | 822.4272 | 1107.4155 | 1583 | 1849.5841 | 1293.3372 | 917.7531 | 1210.9887 | 2033.9213 | 1204.8766 | 943.4159 | 1311.2812 | 1292.09 | 1075.5619 | 967.7975 | 1293.7866 | 1364.6854 | 1272.1012 | 936.854 | 1090.8092 | 1135.1795 | 963.3259 | 1406.7079 | 1565.2247 | 1551.9891 | 1114.7188 | 638.8427 | 885.0448 | 872.9887 | 672.472 | 1809.449 | 2581.5736 | 2860.9657 | 1987.7414 | 1769.831 | 2493.4381 | 2595.4385 | 2084.0339 | 484.618 | 320.6967 | 310.4494 | 484.618 | 464.4494 | 505.8988 | 550.1234 | 572.8092 | 575.9213 | 622.6632 | 611.4381 | 504 | 327.1347 | 345.4155 | 322.8088 | 399.8875 | 454.5619 | 533.0787 | 499.3481 | 281.1573 | 1288.6854 | 1337.4385 | 1615.0787 | 721.3598 | 1302.3594 | 1906.0339 | 2172.2247 | 992.5397 | 1885.0117 | 2058.3037 | 2321.7414 | 1473.4716
609 | 8.6934 | 766.2704 | 1303.1082 | 1594.1897 | 958.2026 | 861.2163 | 1161.297 | 1200.8515 | 804.7837 | 883.9872 | 589.2566 | 505.3648 | 881.527 | 1076.8515 | 883.6352 | 956.3245 | 1034.7974 | 954.9047 | 942.4051 | 1033 | 916.4455 | 813.6352 | 1692.2172 | 1982.9056 | 802.5142 | 1284.5408 | 1615.5536 | 1754.8644 | 1400.1485 | 581.8515 | 767.0953 | 812.4592 | 633.4326 | 788.2429 | 956.297 | 965.297 | 345.7296 | 1591.3107 | 2085.1082 | 2133.0146 | 1350.7708 | 812.2566 | 996.4189 | 1011.6893 | 649.648 | 300.2841 | 268.6893 | 326.3236 | 608.2575 | 1794.8506 | 2111.7415 | 2424.824 | 1112.0944 | 812.6489 | 1079.4207 | 1079.1494 | 567.2172 | 1103.1485 | 1263 | 1628.1906 | 907.2979 | 923.8781 | 1010.121 | 1260.9597 | 792.2163 | 1118.9734 | 1557.9331 | 1862.8515 | 1265.5408 | 862.9734 | 1231.2429 | 2035.473 | 1168.9734 | 931.5674 | 1282.8378 | 1262.2438 | 1067.1485 | 964.8927 | 1266.9193 | 1374.2026 | 1245.6489 | 909.2979 | 1097.2154 | 1139.5683 | 966.6618 | 1389.4326 | 1572.7837 | 1567.2823 | 1112.8927 | 622.5545 | 875.2035 | 891.824 | 658.9056 | 1806.3529 | 2537.5399 | 2926.2163 | 1992.2301 | 1846.8644 | 2494.2841 | 2641.4574 | 2096.1073 | 474.0137 | 340.7021 | 332.1888 | 474.0137 | 435.582 | 514.3785 | 560.4867 | 563.6077 | 578.8918 | 626.1073 | 625.527 | 501.6352 | 317.2438 | 356.9734 | 363.0266 | 405.9193 | 447.0944 | 518.7571 | 545.1751 | 284.2026 | 1286.8515 | 1316.7699 | 1613.0541 | 697.0403 | 1320.3648 | 1889.2566 | 2147.1494 | 960.0807 | 1835.6755 | 2062.6609 | 2396.7012 | 1522.5674
610 | 8.7044 | 774.1018 | 1302.941 | 1580.2198 | 943.6194 | 838.0752 | 1121.137 | 1209.5281 | 834.7183 | 859.3673 | 589.6327 | 524.9838 | 851.8849 | 1062.236 | 882.941 | 925.2522 | 1041.6165 | 980.177 | 951.0885 | 1013.1047 | 898.6651 | 811.4853 | 1644.6327 | 1990.5281 | 781.7345 | 1303.6004 | 1617.9867 | 1733.3273 | 1420.1018 | 570.1475 | 738.4263 | 802.1342 | 594.9144 | 781.1637 | 981.646 | 983.8525 | 322.8554 | 1589.177 | 2082.5147 | 2086.9115 | 1371.6137 | 776.7241 | 995.5737 | 1058.4662 | 659.5014 | 285 | 272.2788 | 341.0457 | 611.3644 | 1758.0942 | 2162.1932 | 2480.2169 | 1130.1903 | 813.5871 | 1042.7907 | 1051.8525 | 570.0989 | 1118.2493 | 1226.1209 | 1622.7745 | 906.2626 | 903.9734 | 1029.236 | 1238.8982 | 801.5871 | 1125.7183 | 1539.705 | 1867.646 | 1291.8791 | 850.3083 | 1252.3083 | 2038.4263 | 1126.2227 | 934.177 | 1265.6194 | 1248.1903 | 1111.8496 | 924.118 | 1249 | 1332.9306 | 1250.4691 | 885.941 | 1090.2683 | 1125.4986 | 962.1342 | 1378.0885 | 1503.2417 | 1590.6946 | 1127.9219 | 609.764 | 855.5871 | 915.1903 | 640.8097 | 1754.5871 | 2598.0266 | 2878.3273 | 1970.236 | 1835.0324 | 2528.9809 | 2694.5766 | 2143.1903 | 469.823 | 354.8392 | 327.1342 | 469.823 | 427.9971 | 517.6755 | 544.2065 | 583.4263 | 560.1342 | 649.4129 | 644.8953 | 494.1475 | 301.7345 | 361.7773 | 349.0029 | 406.4396 | 473.4396 | 516.1313 | 564.6917 | 282.236 | 1281.941 | 1328.0885 | 1536.786 | 704.5442 | 1271.5633 | 1840.6223 | 2106.7507 | 992.0723 | 1785.2255 | 2118.4853 | 2421.6946 | 1515.9572
611 | 8.7154 | 799.941 | 1325.8968 | 1625.3255 | 911.5516 | 817.5025 | 1070.9656 | 1217.5418 | 845.5811 | 825.1917 | 598.9312 | 524.5909 | 817.5172 | 1036.0885 | 875.5467 | 926.3649 | 1066.4435 | 960.5663 | 911.1818 | 1018.8673 | 843.0737 | 789.113 | 1651.4435 | 1973.6646 | 809.4091 | 1313.1081 | 1606.473 | 1722.8231 | 1424.5467 | 562.5123 | 714.6302 | 784.4779 | 563.9951 | 752.527 | 974.5958 | 1001.9361 | 316.8821 | 1617.828 | 2119.8133 | 2091.4926 | 1395.6204 | 761.8624 | 991.5074 | 1080.1573 | 632.5467 | 292.4631 | 285.473 | 315.5319 | 633.0885 | 1668.5811 | 2138.9951 | 2471.2654 | 1113.6351 | 796.6548 | 1071.0688 | 1074.9115 | 586.1622 | 1075.2555 | 1241.7297 | 1657.6253 | 923.5418 | 894.3796 | 1044.3845 | 1170.2703 | 791.6204 | 1131.6057 | 1565.4238 | 1861.0934 | 1296.172 | 846.5516 | 1249.0491 | 1964.8759 | 1114.9312 | 897.6499 | 1206.6843 | 1247.059 | 1110.7039 | 922.4877 | 1197.2555 | 1291.9558 | 1257.059 | 851.1818 | 1017.1425 | 1147.9754 | 912.6351 | 1318.2801 | 1402.113 | 1578.355 | 1108.8661 | 606.059 | 881.941 | 906.5958 | 639.4533 | 1753.5762 | 2511.0971 | 2858.3698 | 2005.7838 | 1785.0688 | 2543.6892 | 2714.715 | 2106.7334 | 487.4435 | 350.4631 | 314.4533 | 487.4435 | 456.5319 | 527.0098 | 531.527 | 584.0098 | 534.0197 | 676.7936 | 657.0442 | 483.527 | 304.5319 | 377.0491 | 328.4287 | 440.4287 | 478.0737 | 540.4435 | 557.5025 | 265.0442 | 1265.0934 | 1324.5025 | 1478.8919 | 719.9165 | 1207.5467 | 1784 | 2076.1081 | 980.1327 | 1747.887 | 2146.8624 | 2379.0049 | 1456.2162
612 | 8.7264 | 800.7837 | 1356.6868 | 1595.7413 | 937.708 | 843.5114 | 1099.9229 | 1227.354 | 781.0756 | 770.1951 | 633.3934 | 529.4524 | 784.9395 | 1079.6899 | 900.6687 | 978.472 | 1047.9198 | 947.5098 | 859.7246 | 1010.3116 | 852.9819 | 760.3918 | 1676.7655 | 1958.354 | 814.7443 | 1333.3162 | 1598.7443 | 1717.7035 | 1420.6082 | 565.0983 | 724.1982 | 772.6853 | 570.6082 | 749.0393 | 948.3722 | 1008.882 | 330.8033 | 1613.233 | 2180.9425 | 2085.8623 | 1379.6672 | 793.059 | 986.941 | 1079.6097 | 609.2345 | 289.8033 | 295.5885 | 311.5492 | 600.7246 | 1631.0968 | 2127.764 | 2325.2299 | 1069.1755 | 762.9607 | 1054.4508 | 1040.9969 | 572.3737 | 1044.413 | 1286.2935 | 1605.4886 | 901.2345 | 884.3313 | 1068.5098 | 1132.3344 | 778.2163 | 1129.3737 | 1563.6656 | 1820.587 | 1263.0393 | 818.1558 | 1228.8033 | 1904.3147 | 1141.236 | 878.2163 | 1188.3737 | 1251.3147 | 1087.354 | 941.3147 | 1149.5885 | 1315.7852 | 1260.8049 | 829.295 | 962.5083 | 1115.2723 | 881.4312 | 1244.6853 | 1364.7246 | 1553.4478 | 1005.4493 | 599.0983 | 815.4856 | 889.0393 | 657.6672 | 1705.8805 | 2404.8049 | 2840.1346 | 1999.0166 | 1808.7277 | 2554.4539 | 2760.2542 | 2116.6112 | 489.823 | 340.1558 | 312.7837 | 489.823 | 465.295 | 514.8033 | 516.823 | 586.5885 | 519.8033 | 684.3313 | 636.2738 | 473.9213 | 282.705 | 366.4902 | 343 | 416.7625 | 451.2738 | 541.8033 | 563.118 | 250.3918 | 1226.1165 | 1302.587 | 1510.6868 | 696.2133 | 1166.3903 | 1744.233 | 2024.9395 | 944.8427 | 1775.0787 | 2173.4705 | 2415.2179 | 1419.1377
613 | 8.7374 | 775.9026 | 1337.2831 | 1526.7345 | 965.6902 | 881.0089 | 1113.6902 | 1198.0087 | 766.3275 | 737.7699 | 635.7699 | 581.7611 | 723.584 | 1127.5132 | 896.407 | 1006.177 | 1036.2301 | 916.1504 | 850.6991 | 1000.0178 | 844.6282 | 756.0443 | 1641.9645 | 1925.3539 | 788.9469 | 1321.2654 | 1596.4425 | 1683.2213 | 1432.7876 | 547.0265 | 717.7521 | 769.531 | 577.5664 | 740.5575 | 971.239 | 1006.6549 | 328.3098 | 1559.8584 | 2137.2122 | 2117.6373 | 1386.9557 | 777.7256 | 985.5664 | 1064.5751 | 591.823 | 289.9204 | 306.7876 | 289.4601 | 618.3275 | 1536.8582 | 2120.177 | 2244.3098 | 1024.8052 | 771.4425 | 1037.3805 | 980.0354 | 555.4955 | 1061.3451 | 1282.7876 | 1541.3627 | 867.6371 | 843.1681 | 1052.8141 | 1073.814 | 771.7699 | 1153.2213 | 1553.2655 | 1762.407 | 1226.3627 | 797.8672 | 1202.8141 | 1898.5929 | 1147.7787 | 859.2389 | 1172.0176 | 1271.177 | 1157.7346 | 953.3451 | 1165.4868 | 1298.6725 | 1232.1415 | 805.3627 | 928.6017 | 1073.7787 | 858.1592 | 1196.6282 | 1311.8406 | 1495.2566 | 988.9646 | 601.9115 | 750.4425 | 866.4601 | 642.5044 | 1629.5751 | 2360.4777 | 2768.8052 | 1948.4336 | 1806.8849 | 2623.9293 | 2761.5753 | 2176.6991 | 479.9557 | 330.8319 | 306.6991 | 479.9557 | 440.8406 | 521.1859 | 520.531 | 597.7876 | 511.0443 | 642.1239 | 632.8319 | 481.4868 | 282.0974 | 358.6902 | 337.7787 | 367.9911 | 424.8584 | 501.7167 | 534.5397 | 258.5753 | 1188.2035 | 1258.5044 | 1479.2742 | 693.5488 | 1104.1504 | 1648.0441 | 1967.761 | 899.407 | 1715.8229 | 2132.8406 | 2427.5044 | 1459.416
614 | 8.7484 | 775.2916 | 1338.2552 | 1551.4617 | 921.0565 | 898.413 | 1090.4331 | 1135.761 | 778.4331 | 679.7811 | 641.3766 | 577.4695 | 669.4857 | 1086.0039 | 909.0286 | 1019.0851 | 1019.935 | 906.239 | 837.3643 | 1008.1052 | 832.1013 | 751.6558 | 1588.9149 | 1846.4006 | 773.7935 | 1277.2753 | 1529.0441 | 1667.1578 | 1457.7046 | 535.0851 | 689.6195 | 764.4331 | 569.9149 | 744.3604 | 942.4207 | 989.1578 | 345.6883 | 1559.3604 | 1994.6999 | 2117.1941 | 1428.7533 | 735.348 | 969.8987 | 1051.3241 | 576.7773 | 284.5708 | 284.6759 | 296.9273 | 617.6071 | 1434.263 | 2159.8059 | 2200.2305 | 1026.7896 | 767.1052 | 1025.0688 | 975.8786 | 507.5019 | 1005.6272 | 1299.6883 | 1481.4169 | 843.1214 | 781.4896 | 1008.6071 | 985.2266 | 739.9675 | 1168.6883 | 1578.7734 | 1673.4896 | 1178.174 | 772.5545 | 1204.5669 | 1855.4857 | 1073.5259 | 818.1052 | 1152.826 | 1239.1941 | 1150.4207 | 920.6071 | 1130.0402 | 1257.7084 | 1182.9149 | 740.0727 | 905.5708 | 1083.9636 | 819.6922 | 1094.8538 | 1235.3844 | 1430.1454 | 951.9714 | 550.2993 | 742.8987 | 853.0851 | 627.9474 | 1543.6435 | 2304.9312 | 2720.5545 | 1904.2467 | 1768.2591 | 2630.6396 | 2783.0851 | 2200.7734 | 446.935 | 329.4493 | 308.3279 | 446.935 | 413.5507 | 518.3643 | 506.8824 | 554.2993 | 483.1415 | 621.174 | 631.9838 | 495.2065 | 275.2266 | 344.5182 | 336.2065 | 373.5669 | 373.5909 | 434.4532 | 486.4331 | 240.5383 | 1134.1253 | 1218.8299 | 1419.0201 | 698.1415 | 1058.9675 | 1542.0928 | 1962.3766 | 885.3078 | 1655.8623 | 2099.6721 | 2420.4292 | 1467.0201
615 | 8.7594 | 802.9479 | 1376.7611 | 1511.3017 | 921.3491 | 851.5161 | 1126.8438 | 1107.5192 | 762.9066 | 594.3431 | 644.1455 | 584.4425 | 670.3216 | 1033.3048 | 936.3323 | 1048.3216 | 1006.2144 | 929.0138 | 820.9066 | 999.2802 | 806.5437 | 707.0904 | 1558.0383 | 1810.8928 | 762.7994 | 1277.7198 | 1520.4977 | 1649.5467 | 1503.5222 | 532.6126 | 705.1072 | 732.5054 | 541.2251 | 700.2359 | 915.8683 | 1001.6156 | 327.1179 | 1569.9066 | 2003.3659 | 2067.412 | 1433.1976 | 734.3216 | 974.1347 | 1049.7581 | 586.1347 | 283.4808 | 282.9617 | 315.8928 | 620.9755 | 1380.4916 | 2118.288 | 2116.6371 | 958.6065 | 705.7964 | 1034.7473 | 948.7443 | 506.5743 | 917.3568 | 1307.3599 | 1445.2389 | 775.1287 | 723.6126 | 1049.1516 | 916.6095 | 668.5958 | 1210.2557 | 1517.2742 | 1663.7228 | 1156.6126 | 737.6922 | 1207.5192 | 1804.9311 | 1006.9862 | 804.9066 | 1126.2772 | 1222.2144 | 1095.9862 | 885.706 | 1073.5988 | 1259.0276 | 1171.1731 | 682.291 | 845.4365 | 1039.049 | 798.1731 | 972.9556 | 1149.1562 | 1321.1425 | 896.2389 | 528.8683 | 738.9341 | 849.5192 | 618.1455 | 1443.2221 | 2247.1425 | 2667.6509 | 1861.5329 | 1698.8484 | 2603.6922 | 2816.1485 | 2204.3323 | 425.0138 | 312.1593 | 299.4395 | 425.0138 | 402.412 | 515.0276 | 489.3737 | 526.8545 | 456.9066 | 593.0521 | 576.3431 | 478.8653 | 258 | 321.0521 | 301.2772 | 370.5054 | 336.6263 | 406.5605 | 428.8101 | 222.2006 | 1106.3737 | 1148.3293 | 1372.7581 | 695.9617 | 994.1562 | 1476.8653 | 1845.9633 | 885.5054 | 1569.8071 | 2100.3599 | 2391.4503 | 1368.2742
616 | 8.7704 | 803.025 | 1364.4041 | 1392.8332 | 905.9571 | 783.9631 | 1158.938 | 1047.1312 | 747.466 | 547.292 | 595.7831 | 606.292 | 643.2801 | 992.9941 | 927.795 | 1059.292 | 969.3672 | 897.1061 | 778.6341 | 954.4852 | 766.292 | 670.087 | 1492.0442 | 1795.497 | 708.4852 | 1243.1371 | 1556.702 | 1604.1622 | 1497.2551 | 518.969 | 665.1622 | 723.031 | 534.621 | 632.7581 | 881.4542 | 1039.857 | 332.5089 | 1536.0191 | 2080.062 | 2091.6889 | 1425.5649 | 703.3672 | 991.677 | 997.8141 | 582.4351 | 259.4601 | 283.5531 | 305.528 | 613.6961 | 1288.3302 | 1993.1562 | 2024.7462 | 914.9809 | 661.646 | 1000.6341 | 933.18 | 505.82 | 889.826 | 1323.3289 | 1409.5221 | 755.4219 | 678.6651 | 1050.8952 | 835.9881 | 603.1121 | 1231.969 | 1457.031 | 1596.6163 | 1152.472 | 706.5531 | 1167.2551 | 1718.5102 | 967.9011 | 824.1239 | 1098.354 | 1204.3731 | 1041.2492 | 839.4291 | 1028.6961 | 1236.0191 | 1123.2801 | 625.2551 | 777.4911 | 935.6592 | 755.8701 | 896.6091 | 1042.7152 | 1229.8391 | 814.1252 | 522.9631 | 733.466 | 820.4351 | 573.6961 | 1424.4219 | 2121.1813 | 2566.7712 | 1824.5781 | 1680.3349 | 2614.4779 | 2914.1417 | 2172.3171 | 390.3111 | 305.677 | 293.441 | 390.3111 | 356.8141 | 517.733 | 477.907 | 523.2301 | 442.025 | 558.8761 | 546.062 | 424.9881 | 241.2301 | 289.1121 | 301.3039 | 356 | 335.795 | 372.8701 | 400.826 | 231.236 | 1066.9571 | 1063.8701 | 1355.472 | 664.6902 | 947.528 | 1383.8583 | 1726.174 | 823.4852 | 1480.9941 | 2100.118 | 2400.0429 | 1274.4911
617 | 8.7814 | 791.4287 | 1356.9971 | 1342.4287 | 870.2861 | 802.1338 | 1124.7207 | 1028.709 | 737.7139 | 543.1426 | 601.7061 | 611.8545 | 604.001 | 978.5713 | 906.0059 | 1069.9961 | 929.8574 | 862.2871 | 768.8516 | 883.8652 | 742.1475 | 668.1426 | 1456.7139 | 1746.8672 | 677.8555 | 1202.5752 | 1488.7285 | 1583.4248 | 1435.2949 | 507.002 | 637.5693 | 717.2881 | 552.7119 | 606.8594 | 862.8516 | 1032.8613 | 365.2832 | 1508.8574 | 2066.291 | 2056.1592 | 1476.4199 | 657.002 | 958.7226 | 988.2822 | 553.8603 | 258.708 | 280.9971 | 311.1406 | 579.8594 | 1209.8623 | 1924.8613 | 1919.7285 | 905.5801 | 650.4307 | 982.2822 | 906.5801 | 483.2891 | 839.4424 | 1342.1416 | 1330.4433 | 708.4463 | 631.1475 | 986.1484 | 798.5732 | 554.7226 | 1250.2803 | 1392.4453 | 1551.5635 | 1093.1582 | 703.4258 | 1081.8721 | 1663.7158 | 970.707 | 801.7236 | 1085.0039 | 1149.5791 | 1012.8545 | 770.1523 | 987.4326 | 1186.0059 | 1106.2813 | 581.0049 | 757.2852 | 842.2978 | 714.1455 | 846.4355 | 968.5771 | 1144.7265 | 714.4394 | 506.2852 | 695.1494 | 798.1445 | 535.2891 | 1386.0156 | 2047.0029 | 2550.1328 | 1745.2988 | 1611.3095 | 2662.1358 | 2941.2939 | 2181.1367 | 362.4277 | 293.2891 | 290.1436 | 362.4277 | 352.9951 | 484.5781 | 430.0088 | 481.4355 | 392.7236 | 521.291 | 524.8642 | 389.2871 | 231.4277 | 264.1455 | 317.4307 | 349.7158 | 317.4336 | 322.0049 | 375.5791 | 214.1475 | 1040.4268 | 1035.2852 | 1332.7207 | 631.8555 | 922.291 | 1332.7119 | 1665.1562 | 750.0088 | 1470.5703 | 2149.5596 | 2366.5879 | 1262.2832
618 | 8.7924 | 809.6746 | 1354.912 | 1323.909 | 874.8358 | 792.0673 | 1141.7656 | 1039.7478 | 729.6627 | 535.2463 | 608.909 | 642.7596 | 568.9031 | 1004.094 | 899.6716 | 1058.0732 | 884.0643 | 829.9881 | 734.0614 | 863.6686 | 724.4985 | 652.0762 | 1448.6627 | 1704.912 | 687.0851 | 1176.4105 | 1369.7181 | 1581.5777 | 1355.0614 | 456.6479 | 624.7419 | 712.9179 | 537.8239 | 586.5777 | 845.5688 | 980.3165 | 334.4837 | 1475.3195 | 2014.9852 | 2031.6775 | 1494.2463 | 638.7448 | 937.0851 | 980.909 | 539.3314 | 268.6627 | 295.8358 | 305.1612 | 573 | 1150.8151 | 1847.2226 | 1850.3225 | 887.7537 | 617.7389 | 973.7419 | 856.6568 | 484.6716 | 798.1672 | 1300.4807 | 1261.6568 | 691.9327 | 583.7359 | 963.7478 | 752.1494 | 523.1642 | 1296.179 | 1259.8823 | 1474.457 | 1029.6598 | 671.1494 | 1031.3314 | 1630.4046 | 950.9822 | 750.8239 | 1046.1553 | 1164.5163 | 998.2404 | 729.3284 | 982.5865 | 1183.7567 | 1087.9031 | 542.0732 | 733.3225 | 804.4164 | 666.3165 | 784.733 | 900.2285 | 1074.3195 | 672.2463 | 470.0673 | 707.0969 | 778.4105 | 560.5984 | 1312.3225 | 2020.4926 | 2416.4304 | 1714.3373 | 1544.4194 | 2634.3957 | 2924.7537 | 2232.6805 | 361.4985 | 274.2463 | 261.7389 | 361.4985 | 345.8239 | 464 | 398.3314 | 442.4926 | 360.4985 | 489.8269 | 509.4194 | 373.3284 | 228.7478 | 259.5015 | 293.4926 | 316.9881 | 266.4016 | 333.8447 | 326.3225 | 197.0821 | 1005.733 | 1034.0821 | 1288.4896 | 628.2463 | 848.8091 | 1273.4718 | 1541.7152 | 750.4283 | 1434.9001 | 2101.9644 | 2331.7567 | 1265.9149
619 | 8.8034 | 814.6578 | 1368.4514 | 1340.1977 | 832.522 | 753.0826 | 1083.9821 | 1045.5339 | 702.7404 | 489.6725 | 527.4748 | 673.3009 | 524.5486 | 939.7698 | 934.8142 | 1070.1977 | 862.7257 | 802.8495 | 688.0619 | 833.0353 | 712.4661 | 640 | 1456.8968 | 1671.9734 | 678.0826 | 1124.289 | 1390.4839 | 1536.2477 | 1303.6991 | 432.1505 | 591.3362 | 689.9734 | 535.705 | 599.0472 | 826.9587 | 983.5812 | 317.7257 | 1406.1179 | 1986.4248 | 1990.4129 | 1509.4514 | 601.994 | 947.7463 | 962.6372 | 539.5752 | 248.8908 | 295.8289 | 293.0206 | 544.3982 | 1077.0973 | 1760.4807 | 1784.3303 | 852.289 | 586.4867 | 910.5427 | 852.5133 | 465.5899 | 804.3628 | 1229.8642 | 1231.8908 | 701.7197 | 569.5339 | 920.0766 | 706.3156 | 493.9734 | 1266.3509 | 1215.1092 | 1359.0973 | 998.9321 | 616.3775 | 1021.4454 | 1565.3097 | 915.6372 | 729.0413 | 1018.6578 | 1192 | 955.7817 | 708.4867 | 953.4395 | 1189.6372 | 1045.3362 | 504.1445 | 700.6991 | 800.4248 | 648.7257 | 750.6372 | 849.6991 | 1043.8495 | 626.6725 | 465.0472 | 675.1798 | 735.2271 | 573.4867 | 1291.0206 | 1968.6725 | 2364.5458 | 1700.1445 | 1502.522 | 2648.4573 | 2936.9587 | 2291.1564 | 365.5546 | 242.9734 | 250.7463 | 365.5546 | 311.528 | 450.295 | 398.5752 | 431.1917 | 337.7404 | 439.0766 | 483.0147 | 354.2743 | 211.8908 | 254.2537 | 271.4661 | 285.0826 | 239.8083 | 295.7964 | 318.7257 | 209.3009 | 968.0619 | 1017.6991 | 1237.0353 | 591.6106 | 791.3156 | 1195.0147 | 1471.6784 | 731.2684 | 1341.4542 | 1944.873 | 2283.4129 | 1205.2005
620 | 8.8144 | 840.0163 | 1397.5959 | 1307.698 | 810.1224 | 710.9429 | 1074.4367 | 999.698 | 656.1592 | 496.4489 | 525.3959 | 700.6367 | 514.2571 | 890.0816 | 928.6735 | 1060.2 | 846.8898 | 750.5633 | 689.9061 | 799.0531 | 676.2 | 622.9715 | 1417.7388 | 1662.9061 | 656.6204 | 1076.5388 | 1346.7348 | 1440.2368 | 1312.853 | 433.351 | 591.9877 | 643.8082 | 529.1878 | 634.6775 | 781.1307 | 942.9674 | 314.6612 | 1347.0245 | 1972.4449 | 1946.9592 | 1491.1592 | 570.8368 | 919.5511 | 965.6898 | 518.2817 | 258.2449 | 291.1755 | 312.2857 | 514.6612 | 1022.4572 | 1653.3307 | 1663.2083 | 775.9552 | 584.0816 | 865.4449 | 847.6204 | 451.8653 | 744.7511 | 1197.9184 | 1212.0531 | 659.5919 | 557.7551 | 852.2531 | 670.6735 | 467.8776 | 1232.6082 | 1202.1062 | 1356.151 | 964.7143 | 616.0285 | 962.8735 | 1499.7266 | 916.2571 | 742.0285 | 965.5633 | 1146.996 | 938.6082 | 693.3102 | 906.8898 | 1149.5103 | 988.8205 | 500.3388 | 637.4817 | 788.8776 | 626.8082 | 688.6164 | 810.8082 | 980.6164 | 584.7959 | 452.0245 | 651.3796 | 773.8162 | 520.6041 | 1266.498 | 1893.347 | 2335.4124 | 1599.0328 | 1407.7511 | 2557.3267 | 2901.4694 | 2320.4735 | 348.1062 | 215.0531 | 249.9184 | 348.1062 | 314.2041 | 411.8082 | 402.8245 | 442.9469 | 331.9061 | 404.7021 | 448.0123 | 331.1878 | 208.4735 | 247.1755 | 254.6612 | 274.5674 | 211.0245 | 264.9061 | 313.2286 | 212.3102 | 944.9715 | 972.7266 | 1215.8245 | 581.3388 | 783.649 | 1194.0694 | 1445.7551 | 729.3265 | 1297.6082 | 1866.2286 | 2265.502 | 1183.0285
621 | 8.8254 | 833.941 | 1388.3864 | 1251.6637 | 807.9341 | 671.0069 | 1059.7069 | 948.6273 | 645.5477 | 467.7433 | 503.1228 | 690.9046 | 507.0727 | 843.7433 | 926.65 | 1036.0364 | 846.8909 | 732.4091 | 686.3137 | 743.6046 | 639.6273 | 618.825 | 1361.1819 | 1679.1682 | 652.3432 | 1012.5024 | 1256.8387 | 1394.6205 | 1285.291 | 443.4091 | 571.8387 | 646.0954 | 506.5909 | 578.5092 | 729.7659 | 916.65 | 297.4523 | 1318.4523 | 1887.4433 | 1909.6342 | 1495.4386 | 546.8682 | 866.3569 | 948.2841 | 489.9705 | 253.2478 | 285.0364 | 302.4228 | 507.3795 | 982.4228 | 1540.5751 | 1629.5772 | 710.2546 | 556.3569 | 825.7364 | 812.941 | 453.1386 | 723.2341 | 1178.4887 | 1208.1023 | 655.5113 | 524.9046 | 802.4228 | 641.9705 | 469.4454 | 1249.7522 | 1152.6342 | 1323.9842 | 929.9046 | 604.6637 | 928 | 1457.1455 | 893.5614 | 723.2182 | 942.4454 | 1128.5477 | 929.0727 | 683.4159 | 914.9568 | 1139.0295 | 958.5182 | 466.6705 | 567.6342 | 764.3864 | 622.2704 | 624.3205 | 759.7364 | 951.0659 | 578.7227 | 432.7591 | 671.1682 | 773.9114 | 492 | 1220.9773 | 1811.1228 | 2309.0954 | 1566.5772 | 1407.9499 | 2521.6726 | 2926.8772 | 2362.4681 | 304.2182 | 188.1455 | 237.9341 | 304.2182 | 286.7728 | 381.2114 | 370.4955 | 432.6273 | 296.6705 | 401.2409 | 417.2841 | 340.8545 | 197.1091 | 226.1455 | 256.6863 | 268.0364 | 210.9931 | 272.4818 | 305.6205 | 199.9341 | 902.3569 | 977.9204 | 1161.7796 | 593.5841 | 791.5841 | 1139.991 | 1382.5024 | 737.5546 | 1264.4955 | 1810.8455 | 2164.5957 | 1168.5614
622 | 8.8364 | 852.8009 | 1347.261 | 1216.4867 | 849.8231 | 672.0221 | 990.2389 | 922.1814 | 631.7522 | 400.9336 | 478.3274 | 722.2921 | 495.938 | 804.1416 | 897.6681 | 1045.6549 | 877.5133 | 701.4026 | 624.6857 | 724.7567 | 620.1416 | 599.5885 | 1344.1018 | 1635.3804 | 639.1814 | 1001.2478 | 1237.1018 | 1395 | 1241.3849 | 430.1194 | 560.3274 | 652.0398 | 483.8141 | 549.2876 | 726.8584 | 902.2212 | 309.9823 | 1280.3628 | 1778.6238 | 1856.3628 | 1510.3673 | 554.0841 | 814.73 | 880.584 | 452.8938 | 261.3496 | 265.5486 | 288.6327 | 538.0045 | 964.2035 | 1471.3628 | 1597.1769 | 713.0443 | 535.1018 | 835.739 | 792.5708 | 423.5265 | 721.7124 | 1146.2212 | 1129.0087 | 621.9734 | 489.2212 | 757.6283 | 618.8141 | 463.8761 | 1283.4115 | 1110.1194 | 1264.8938 | 891.6902 | 565.6902 | 922.938 | 1476.6372 | 860.2832 | 713.8584 | 962.8186 | 1128.0398 | 914.1416 | 656.5885 | 917.6106 | 1073.6016 | 951.938 | 464.1461 | 571.3098 | 715.0353 | 586.0751 | 607.3673 | 711.5265 | 903.4822 | 544.3008 | 417.4469 | 650.1592 | 759.4912 | 506.5531 | 1175.4247 | 1790.1239 | 2271.3804 | 1508.8671 | 1381.7477 | 2544.5708 | 2880.8052 | 2351.1991 | 291.062 | 188.0221 | 203.8938 | 291.062 | 277.9204 | 382.2257 | 358.8584 | 394.792 | 306.1682 | 367.1991 | 419.1859 | 329.5885 | 172.8141 | 231.7168 | 250.6106 | 251.7124 | 205.5486 | 271.469 | 281.9557 | 180.4469 | 879.2035 | 943.115 | 1125.469 | 610.2876 | 755.1371 | 1033.8273 | 1353.531 | 710.8539 | 1262.9823 | 1770.6106 | 2104.1239 | 1105.5442
623 | 8.8474 | 846.5535 | 1369.3472 | 1176.2635 | 848.6184 | 629.3323 | 965.355 | 869.1376 | 614.1947 | 389.0305 | 451.0383 | 713.8131 | 481.3894 | 808.0305 | 871.4848 | 1026.5535 | 810.6606 | 685.4504 | 586.8397 | 705.1986 | 589.8436 | 594.3855 | 1284.9468 | 1547.3018 | 629.355 | 996.8741 | 1210.8436 | 1380.8092 | 1176.4621 | 423.8053 | 552.3894 | 670.5762 | 466.6145 | 546.0344 | 732.645 | 874.8092 | 305.5191 | 1252.3245 | 1694.6223 | 1807.0383 | 1469.3628 | 516.9124 | 803.0954 | 835.4543 | 468.8968 | 240.3933 | 241.8092 | 295.4504 | 515.3628 | 916.2025 | 1407.8475 | 1499.6911 | 713.4543 | 556.5762 | 803.3323 | 798.3855 | 409.5801 | 675.2979 | 1067.2064 | 1081.355 | 634.4121 | 474.7099 | 681.9468 | 605.4848 | 426.3933 | 1280.2291 | 1078.0039 | 1180.2713 | 886.6755 | 547.7748 | 869.0422 | 1411.1454 | 837.0993 | 648.0461 | 950.7138 | 1124 | 886.4238 | 637.1947 | 874.233 | 1057.4465 | 934.1642 | 429.9124 | 573.7443 | 669.1337 | 556.1298 | 590.2291 | 704.0305 | 882.6755 | 506.2291 | 397.3894 | 624.9695 | 721.1376 | 501.4543 | 1171.6106 | 1744.3972 | 2232.9695 | 1480.5457 | 1351.355 | 2496.2025 | 2806.9734 | 2331 | 292.0649 | 175.6489 | 206.9961 | 292.0649 | 269.2291 | 348.9429 | 346.584 | 378.355 | 289.878 | 347 | 397.6833 | 300.5191 | 176.2557 | 197.4925 | 238.1947 | 259.4809 | 215.9961 | 245.4886 | 254.4543 | 182.3206 | 865.3894 | 913 | 1104.0039 | 581.233 | 742.3206 | 1008.2557 | 1333.7138 | 673.4886 | 1234.0078 | 1739.4886 | 2019.695 | 1062.5191
624 | 8.8584 | 833.972 | 1368.6509 | 1111.0432 | 761.3514 | 609.2306 | 901.8966 | 838.4267 | 562.3578 | 372.9655 | 439.9159 | 667.3923 | 443.4483 | 770.9311 | 921.5861 | 981.763 | 777.5172 | 673.2242 | 572.8254 | 689.6853 | 603.6077 | 575.3362 | 1225.3923 | 1540.6357 | 597.4483 | 944.6164 | 1177.2802 | 1370.3707 | 1148.3147 | 427 | 529.5668 | 670.5108 | 454.7694 | 555.1185 | 672.526 | 870.944 | 294.3707 | 1176.0087 | 1741.6206 | 1834.6982 | 1493.6702 | 511.3771 | 786.6229 | 807.6229 | 460.0496 | 246.6638 | 225.4548 | 294.3987 | 468.0216 | 927.8103 | 1332.3859 | 1472.8319 | 690.2242 | 537.4483 | 773.3707 | 783.2802 | 411.6573 | 633.6229 | 1030.3427 | 1065.2242 | 615.5604 | 474.6573 | 640.7974 | 587.8815 | 404.3987 | 1247.9936 | 1080.8039 | 1172.5797 | 867.0216 | 571.6357 | 834.4267 | 1357.1121 | 828.028 | 658.9568 | 915.3362 | 1067.4699 | 892.4051 | 668.8383 | 811.9311 | 1065.4267 | 887.2177 | 412.5452 | 568.3147 | 649.4267 | 560.5733 | 554.0496 | 659.0432 | 815.0367 | 493.7134 | 402.1746 | 611.4548 | 699.6293 | 473.6229 | 1116.8686 | 1703.5108 | 2191.1897 | 1475.5108 | 1303.0152 | 2480.8599 | 2668.7373 | 2282.3578 | 292.3147 | 184.0625 | 198.2802 | 292.3147 | 251.4548 | 338.5172 | 329.1121 | 353.0216 | 281.1746 | 332.5388 | 386.9159 | 305.8039 | 169.8535 | 192.0625 | 244.8599 | 237.763 | 220.4267 | 233.972 | 265.4051 | 182.0841 | 862.944 | 889.3362 | 1121.9224 | 580.0625 | 700.6725 | 967.0152 | 1318.056 | 652.7694 | 1192.9375 | 1745.0625 | 1977.8879 | 1108.5581
625 | 8.8694 | 839.0177 | 1363.6962 | 1103.7493 | 748.1062 | 605.6342 | 867.6608 | 818.9292 | 560.4366 | 388.1151 | 406.8406 | 657.6873 | 429.6785 | 716.8584 | 895.4306 | 966.6696 | 772.9557 | 635.5368 | 545.5722 | 717.7936 | 638.0885 | 581.0797 | 1221.0443 | 1507.1268 | 610.4632 | 884.5014 | 1159.6342 | 1297.3509 | 1104.1356 | 417.6254 | 533.7227 | 646.5899 | 420.8672 | 546.2596 | 659.0797 | 869.6519 | 298.3569 | 1133.6431 | 1611.2564 | 1756.2181 | 1437.6459 | 529.3658 | 782.6962 | 808.3835 | 443.9911 | 227.528 | 251.8024 | 277.938 | 440.938 | 849.9054 | 1287.2507 | 1420.0736 | 667.2684 | 490.1976 | 710.3952 | 783.7139 | 422.0443 | 608.2684 | 992.5014 | 1087.8024 | 595 | 446.2064 | 656.7581 | 582.0089 | 407.3569 | 1204.528 | 992.9143 | 1137.1091 | 813.8229 | 574.2861 | 819.6165 | 1353.6696 | 809.5899 | 666.9203 | 882.2419 | 1023.938 | 905.3481 | 620.7079 | 811.4278 | 980.9763 | 853.6076 | 429.0531 | 536.8584 | 656.7139 | 542.5722 | 565.4455 | 641.705 | 805.1239 | 467.2241 | 415.3658 | 627.7581 | 718.4101 | 473.0443 | 1142.5871 | 1629.3686 | 2111.7344 | 1418.1091 | 1260.5899 | 2362.7904 | 2608.3481 | 2293.829 | 298.3569 | 180.6076 | 185.9911 | 298.3569 | 238.295 | 319.2241 | 318.3038 | 358.0797 | 266.2419 | 324.3304 | 380.6254 | 311.6608 | 174.3835 | 220.7493 | 248.6608 | 240.7493 | 217.6696 | 235.6696 | 264.2861 | 160.9203 | 798.0382 | 850.2153 | 1137.6696 | 577.2773 | 665.2861 | 939.3215 | 1204.1739 | 633.5988 | 1146.8406 | 1696.7433 | 1931.7787 | 1135.6608
626 | 8.8804 | 836.2266 | 1291.3081 | 1093.9066 | 755.9085 | 576.7694 | 828.8131 | 818.501 | 577.4553 | 385.3161 | 370.5885 | 649.0895 | 437.8191 | 696.9085 | 848.5427 | 961.5447 | 780.9563 | 613.499 | 533.6362 | 677.4473 | 600.9483 | 555.9046 | 1179.6302 | 1490.1372 | 609.998 | 831.9046 | 1117.1312 | 1214.2664 | 1066.3161 | 402.0895 | 516.8151 | 635.9543 | 448.9602 | 525.7256 | 687.0934 | 814.8111 | 307.1372 | 1099.6322 | 1551.1849 | 1710.7734 | 1389.1352 | 509.1332 | 763.497 | 805.8171 | 434.8171 | 216.0915 | 266.3181 | 281.9105 | 432.6362 | 802.6819 | 1250.7237 | 1349.4036 | 645.6799 | 487.9563 | 696.3658 | 758.996 | 403.8151 | 601 | 944.6779 | 1051.1749 | 561.5865 | 413.5885 | 600.1729 | 571.4076 | 373.8588 | 1219.004 | 956.9105 | 1103.8171 | 783.0895 | 556.7256 | 803.7714 | 1293.3101 | 779.1789 | 666.8648 | 856.3161 | 1000.6342 | 880.4056 | 595.501 | 809.0437 | 979.8688 | 891.3241 | 436.3638 | 547.5944 | 657.2266 | 566.0954 | 539.1312 | 617.3598 | 803.998 | 476.8211 | 406.0437 | 592.994 | 718.1352 | 459.5885 | 1109.6282 | 1627.3678 | 2172.0119 | 1384.1352 | 1241.0895 | 2238.5825 | 2528.171 | 2236.3081 | 293.499 | 186.9563 | 178.1809 | 293.499 | 234.6819 | 305.6362 | 336.8211 | 355.4076 | 268.2744 | 311.0437 | 378.7276 | 295.3161 | 176.5905 | 207.497 | 220.0417 | 269.0934 | 201.6342 | 233.9543 | 229.6779 | 171.7754 | 792.2763 | 824.0437 | 1116.8608 | 562.8171 | 649.7714 | 919.5885 | 1172.1849 | 666.2326 | 1128.3181 | 1617.5388 | 1876.996 | 1111.8151
627 | 8.8914 | 819.4218 | 1266.413 | 1108.6028 | 741.5044 | 579.5782 | 777.6784 | 792.1494 | 548.0668 | 355.0914 | 371.7188 | 638.4464 | 401.4306 | 695.6942 | 828.587 | 928.4552 | 765.4798 | 588.4798 | 556.6028 | 645.058 | 567.1986 | 517.7856 | 1171.884 | 1463.8436 | 605.7768 | 839.6028 | 1015.696 | 1164.9262 | 1133.2794 | 402.5536 | 496.2812 | 591.9596 | 430.819 | 529.0246 | 690.4464 | 770.5624 | 295.4218 | 1083.6116 | 1496.051 | 1735.9086 | 1371.8102 | 466.761 | 706.4886 | 777.703 | 452.5202 | 226.3304 | 234.0668 | 279.3638 | 448.6608 | 768.9842 | 1201.2074 | 1325.8348 | 639 | 493.3058 | 678.3146 | 745 | 386.058 | 594.7522 | 931.6942 | 1033.942 | 519.6204 | 401.529 | 562.058 | 550.7276 | 363.2478 | 1176.7698 | 929.3726 | 1111.8014 | 814.0984 | 557.2478 | 778.8682 | 1295.5448 | 751.645 | 653.645 | 839.2812 | 1007.5782 | 849.2566 | 579.5624 | 778.703 | 970.232 | 895.3392 | 416.8682 | 505.8524 | 636.2566 | 545.2074 | 525.1652 | 593.587 | 721.3322 | 441.1828 | 404.8594 | 569.9174 | 672.348 | 469.7434 | 1055.3146 | 1617.703 | 2118.8946 | 1350.1494 | 1183.935 | 2108.341 | 2488.6116 | 2173.9262 | 272.645 | 171.4798 | 184.7188 | 272.645 | 257.9086 | 308.471 | 326.8682 | 353.471 | 244.8436 | 287.645 | 352.232 | 290.7768 | 156.9508 | 186.587 | 216.7188 | 231.4886 | 171.0914 | 203.8436 | 246.5448 | 173.3638 | 797.3638 | 824.942 | 1015.3656 | 534.703 | 624.174 | 872.819 | 1164.9508 | 630.8524 | 1084.9596 | 1530.1336 | 1825.5132 | 1090.587
628 | 8.9024 | 799.5767 | 1239.9144 | 1070.2521 | 730.3481 | 561.2728 | 798.3245 | 786.0649 | 563.6726 | 338.6416 | 391.5428 | 595.73 | 406.6077 | 700.2389 | 799.9793 | 864.8392 | 739.7507 | 582.5325 | 533.0884 | 639.174 | 590.9661 | 488.9247 | 1123.6002 | 1419.2079 | 594.2832 | 804.7846 | 983.9454 | 1216.5842 | 1095.3509 | 393.6961 | 514.6077 | 604.7921 | 411.174 | 494.8495 | 681.6416 | 789.5531 | 293.2389 | 1034.2521 | 1453.5767 | 1673.2314 | 1380.1298 | 453 | 700.8363 | 755.6961 | 446.4572 | 223.6416 | 224.9454 | 273.174 | 451.7611 | 725.0339 | 1145.4365 | 1287.969 | 620.6312 | 472.5118 | 677.5974 | 731.5767 | 369.5767 | 577.8702 | 891.0235 | 955.0471 | 534.8467 | 413.4233 | 542.0442 | 546.1195 | 362.4675 | 1142.1091 | 844.3509 | 1078.8495 | 785.4911 | 524.6755 | 790.4882 | 1244.0028 | 763.0753 | 627.6312 | 821.5767 | 1017.5325 | 877.031 | 554.0442 | 761.6416 | 902.8937 | 845.6106 | 367.0235 | 491.1844 | 645.6623 | 476.7197 | 481.7846 | 580.5221 | 733.9219 | 437.9558 | 371.6755 | 573.2389 | 672.3688 | 465.2832 | 1019.2728 | 1440.9747 | 2094.6623 | 1363.8467 | 1080.46 | 2056.9351 | 2506.3688 | 2032.0057 | 262.174 | 171.8909 | 203.8363 | 262.174 | 231.969 | 311.413 | 317.2935 | 334.5118 | 230.587 | 292.0103 | 347.0649 | 272.2183 | 164.5428 | 175.6416 | 199.0988 | 219.4779 | 189.9661 | 212.1947 | 241.6312 | 184.8363 | 785.5221 | 803.9793 | 985.5974 | 560.7375 | 607.2286 | 838.3377 | 1101.1872 | 616.8909 | 1088.6077 | 1506.2389 | 1836.5531 | 1073.9897
629 | 8.9134 | 795.4376 | 1237.0629 | 1026.999 | 695.6548 | 540.9371 | 756.2162 | 760.2487 | 576.7188 | 348.9381 | 378.5929 | 548.0924 | 403.7807 | 692.6558 | 756.936 | 840.4061 | 764.6264 | 571.0619 | 507.4995 | 614.2802 | 575.6863 | 488.1878 | 1043.4975 | 1396.6558 | 573.9056 | 775.9371 | 1001.8132 | 1200.5919 | 1015.6853 | 378.9371 | 514.6558 | 623.7503 | 422.2193 | 475.4061 | 637.3107 | 785.9056 | 271.2802 | 1003.2183 | 1457.3442 | 1607.0609 | 1306.934 | 446.531 | 676.0924 | 729.4365 | 421.3117 | 208.0619 | 224.1249 | 255.468 | 427.7178 | 720.1878 | 1064.3096 | 1270.8437 | 629.532 | 469.0315 | 658.2802 | 677.1228 | 362.5624 | 570.5624 | 881.9066 | 955.5015 | 537.3746 | 390.2487 | 535 | 520.4051 | 359.5624 | 1110.6863 | 809.9685 | 1034.9675 | 746.1553 | 492.0305 | 742.2162 | 1174.5919 | 749.4365 | 604.9056 | 783.6548 | 958.6223 | 865.4051 | 570.001 | 725.936 | 936.065 | 835.9066 | 379.4701 | 497.8751 | 588.3096 | 491.6579 | 457.9685 | 543.936 | 710.5919 | 427.0305 | 354.8437 | 551.2802 | 679.2812 | 451.3746 | 979.5299 | 1428.065 | 2004.2457 | 1316.0599 | 1105.2213 | 2003.6853 | 2387.4944 | 1987.3442 | 272.5005 | 173.8437 | 207.8437 | 272.5005 | 227.0629 | 280.3736 | 302.6244 | 340.3756 | 256.5949 | 299.1563 | 330.5929 | 289.8761 | 163.8122 | 183.7817 | 198.6254 | 223 | 189.7807 | 222.4376 | 227.531 | 183.0934 | 742.467 | 749.4355 | 1043.9086 | 577.4376 | 584.5929 | 792.1858 | 1039.9046 | 577.8731 | 1077.1553 | 1505.8437 | 1740.1827 | 1017.9665
630 | 8.9244 | 813.5458 | 1149.6159 | 995.3367 | 683 | 488.0181 | 726.6893 | 789.209 | 569.6893 | 355.4621 | 354.1853 | 545.5039 | 413.6215 | 661.9503 | 714.7571 | 814.1435 | 692.1955 | 555.0339 | 511.0418 | 621.0836 | 554.765 | 512.9322 | 1074.3683 | 1307.4645 | 584.0836 | 786.6215 | 914.4227 | 1143.0599 | 1012.8147 | 369.1514 | 482.4882 | 601.1435 | 427 | 463.0339 | 628.3107 | 774.8825 | 276.6215 | 1021.6633 | 1412.0181 | 1586.4203 | 1300.8565 | 473.2429 | 639.8746 | 727.5797 | 389.1435 | 178.1435 | 236.1593 | 269.47 | 436.5458 | 735.4282 | 1070.4361 | 1191.7752 | 576.052 | 454.9164 | 666.5458 | 674.8147 | 354.6893 | 554.6475 | 852.8328 | 940.4882 | 548.6215 | 392.8904 | 507.2192 | 508.5379 | 353.8825 | 1003.4984 | 834.3186 | 1036.4282 | 745.3107 | 439.6735 | 739.739 | 1185.0497 | 734.4203 | 606.3107 | 822.1751 | 960.5876 | 835.9921 | 566.5718 | 683.026 | 941.1853 | 811.9503 | 337.6317 | 495.3446 | 568.8486 | 490.6475 | 469.8904 | 550.0079 | 716.6633 | 420.7311 | 374.47 | 531.765 | 642.4464 | 474.3186 | 954.8407 | 1441.9582 | 1990.8565 | 1311.2768 | 1090.7153 | 2005.9322 | 2297.0599 | 1946.4046 | 263.8407 | 154.7232 | 197.4961 | 263.8407 | 252.9322 | 275.3107 | 292.6136 | 332.8407 | 272.1175 | 301.4621 | 325.9243 | 260.5221 | 171.9661 | 178.4961 | 211.5039 | 242.739 | 180.6136 | 205.4542 | 225.7311 | 179.2689 | 735.7729 | 781.713 | 1005.052 | 577.2689 | 584.3107 | 822.3265 | 1019.8825 | 591.7051 | 1036.8328 | 1445.7831 | 1714.8147 | 1008.235
631 | 8.9354 | 817.0265 | 1113.7699 | 953.5486 | 677.0531 | 511.6284 | 698.7256 | 723.7167 | 577.4071 | 338.1593 | 314.5486 | 470.2034 | 409.3363 | 620.0353 | 732.2478 | 813.177 | 669.177 | 521.2654 | 474.6017 | 606.9292 | 559.177 | 483.0884 | 1047.1681 | 1284.177 | 572.9026 | 776.3894 | 946.9028 | 1150.3009 | 995.9557 | 358.823 | 469.7699 | 567.469 | 407.6725 | 464.2035 | 610.9292 | 758.8761 | 276.7965 | 1007.6991 | 1341.2212 | 1596.6372 | 1226.7432 | 434.938 | 645.1239 | 745.6106 | 366.6194 | 202.4514 | 222.4159 | 271.7965 | 442.2566 | 698.3716 | 1025.5043 | 1154.823 | 559.2035 | 454.2035 | 612.7875 | 660.9292 | 358.6903 | 550.4867 | 810.5221 | 925.5398 | 531.6991 | 380.1593 | 481.3894 | 504.2832 | 337.8761 | 953.1327 | 792.7079 | 1017.2124 | 749.4867 | 483.9558 | 672.6636 | 1190.5398 | 711.9292 | 609 | 783.0176 | 932.0884 | 807.4424 | 549.3628 | 668.2832 | 935.7434 | 795.3097 | 358.885 | 465.7522 | 562.823 | 489.4602 | 457.9026 | 522.8053 | 733.9204 | 418.0265 | 380.5133 | 527.2566 | 647.5841 | 463.9292 | 982.708 | 1363.9202 | 1931.6105 | 1320.9735 | 1097.8142 | 1882.4246 | 2260.4424 | 1895.8053 | 250.0796 | 155.4336 | 192.5133 | 250.0796 | 224.5752 | 296.5841 | 282.823 | 317.5929 | 241.292 | 316.1239 | 324.0265 | 284.6549 | 149.9823 | 176.4867 | 216.4867 | 224.7256 | 167.1062 | 216.8407 | 223.0265 | 166.3628 | 726.3628 | 743.5043 | 988.2035 | 559.9026 | 566.9292 | 818.1858 | 1014.2832 | 568.8053 | 974.4512 | 1394.2654 | 1701.6725 | 983.0088
632 | 8.9464 | 806.1765 | 1035.1675 | 957.8688 | 665.9321 | 467.0589 | 691.5113 | 739.828 | 555.3077 | 325.1991 | 328.6923 | 461.8914 | 384.8417 | 593.8869 | 691.3711 | 807.6878 | 693.1583 | 480.7738 | 478.3801 | 606.8009 | 530.2625 | 472.267 | 987.9277 | 1302.1131 | 586.647 | 767.2217 | 913.3259 | 1173.8688 | 973.6426 | 350.7104 | 463.7104 | 537.086 | 394.1991 | 447.8643 | 599.4661 | 755.5113 | 297.6697 | 1003.0226 | 1282.9503 | 1569.2625 | 1193.4661 | 436.1357 | 628.8417 | 729.086 | 404.3167 | 182.2625 | 207.6878 | 267.733 | 410.2625 | 647.7286 | 931.9458 | 1131.5973 | 560.2443 | 452.9774 | 595.7783 | 657.0226 | 379.8914 | 529.8417 | 794.7104 | 892.2625 | 521.733 | 375.5113 | 487.3348 | 461.439 | 326.1991 | 935.9095 | 764.4208 | 984.0634 | 683.5024 | 488.3756 | 697.6289 | 1210.914 | 702.733 | 557.6155 | 744.1086 | 909.9321 | 800.7557 | 498.1494 | 671.5339 | 936.7557 | 814.1583 | 367.7104 | 427.2851 | 577.3801 | 504.6018 | 411.1946 | 483.7964 | 703.7512 | 388.2851 | 330.1268 | 505.086 | 610.9277 | 425.2625 | 933.5476 | 1364.4479 | 1929.914 | 1297.819 | 1015.8327 | 1771.258 | 2234.8643 | 1813.7241 | 244.733 | 154.2217 | 165.5521 | 244.733 | 199.9095 | 271.2625 | 284.5339 | 321.5566 | 252.6697 | 266.5928 | 360.7828 | 289.6878 | 167.4479 | 169.4434 | 207.9321 | 218.267 | 173.5792 | 204.1086 | 240.8914 | 163.5113 | 727.2896 | 728.5339 | 927.2806 | 559.2896 | 540.353 | 770.1946 | 944.991 | 558.5113 | 1021.9863 | 1356.0408 | 1627.7467 | 994.1583
633 | 8.9574 | 765.3706 | 1053.005 | 933.0501 | 650.7129 | 470.2704 | 697.3756 | 728.6578 | 511.6745 | 352.182 | 337 | 439.586 | 400.2704 | 595.1436 | 671.1603 | 762.763 | 643.172 | 502.4858 | 457.8898 | 600.5526 | 556.0935 | 461.4808 | 1035.4357 | 1346.9332 | 576.8731 | 713.0117 | 908.0551 | 1130.6194 | 946.7296 | 352.0718 | 428.9783 | 552.2704 | 401.9833 | 443.5359 | 594.1603 | 736.8013 | 281.1937 | 951.0117 | 1246.9616 | 1482.9733 | 1168.1937 | 452.7512 | 602.0334 | 718.1603 | 392.6578 | 176.6077 | 207.3038 | 247.0334 | 422.2704 | 674.3973 | 881.1219 | 1062.0284 | 559.2321 | 454.3038 | 590.8564 | 688.7178 | 372.177 | 519.9282 | 722.349 | 868.9449 | 534.823 | 362.9449 | 466.4257 | 474.8781 | 314.7846 | 902.818 | 731.0501 | 967.5526 | 718.828 | 484.5359 | 674.8347 | 1184.9783 | 738.0935 | 573.2537 | 744.9115 | 962.2921 | 801.7679 | 505.2704 | 673.7679 | 927.0167 | 796.4257 | 336.818 | 439.5025 | 575.3205 | 475.9783 | 434.8614 | 488.0551 | 608.5259 | 383.6077 | 328.591 | 479.5693 | 603.9115 | 477.9716 | 900.177 | 1349.9616 | 1906.2822 | 1293.8397 | 1003.8948 | 1747.6962 | 2197.5142 | 1867.7946 | 255.5192 | 164.5192 | 159.3038 | 255.5192 | 212.8948 | 291.7178 | 278.3205 | 304.8013 | 245.409 | 258.2154 | 359.9449 | 278.5526 | 156.8013 | 163.1603 | 196.5526 | 215.9282 | 166.2488 | 210.2871 | 225.4975 | 156.3205 | 701.3539 | 756.1102 | 899.3205 | 543.3372 | 512.4975 | 743.9449 | 977.5242 | 558.2321 | 957.758 | 1314.7463 | 1597.5526 | 943.404
634 | 8.9684 | 768.8245 | 1002.8008 | 867.8229 | 636.0767 | 477.7802 | 660.7684 | 713.1976 | 540.0119 | 345.3953 | 326.8569 | 415.3953 | 396.0767 | 610.264 | 648.9336 | 735.1755 | 681.3968 | 461.2846 | 442.9779 | 552.7462 | 560.7581 | 467.3628 | 1005.6253 | 1223.2387 | 569.6593 | 708.7036 | 871.4277 | 1083.23 | 923.615 | 332.7138 | 419.7802 | 532.4719 | 395.6372 | 443.2198 | 583.6372 | 708.3731 | 231.8672 | 889.7462 | 1223.615 | 1463.1209 | 1216.3968 | 449.7581 | 556.9881 | 656.9217 | 385 | 202.9676 | 198.6372 | 255.0443 | 425.8791 | 645.4277 | 877.6814 | 1014.1312 | 551.1976 | 430.0324 | 585.8791 | 637.9217 | 368 | 532.264 | 730.649 | 846.2743 | 489.0648 | 343.3953 | 452.7581 | 481.4395 | 329.9455 | 808.9542 | 737.6047 | 926.7684 | 709.4719 | 431.9439 | 646.615 | 1126.6253 | 706.087 | 574.4174 | 735.2964 | 938.4277 | 813.7036 | 494.0545 | 667.7581 | 876.4055 | 843.6166 | 335.0221 | 398.4055 | 535.5486 | 432.4498 | 437.4174 | 492.7802 | 604.0664 | 394.3628 | 347.0443 | 488.8245 | 612.2419 | 450.966 | 901.4617 | 1296.1858 | 1913.385 | 1298.1209 | 965.3067 | 1715.7905 | 2106.0751 | 1837.9439 | 241.0545 | 157.8569 | 163.1209 | 241.0545 | 233.6047 | 257.3067 | 304.8688 | 309.1431 | 230.0767 | 262.5605 | 324.0103 | 263.5162 | 138.5162 | 174.4838 | 190.0988 | 204.8569 | 156.7581 | 214.7802 | 240.0664 | 162.5826 | 685.1976 | 703.1415 | 910.264 | 509.1312 | 506.7802 | 723.615 | 943.2846 | 532.2522 | 967.8909 | 1287.0545 | 1581.736 | 942.385
635 | 8.9794 | 741.2989 | 936.6558 | 934.4229 | 640.1328 | 469.2822 | 645.2448 | 655.523 | 508.2035 | 342.5851 | 333.5103 | 407.0374 | 363.6765 | 583.0914 | 632.6971 | 729.4149 | 666.8839 | 502.477 | 406.9213 | 578.8339 | 566.1328 | 437.5063 | 990.4149 | 1236.5143 | 549.1868 | 702.4897 | 863.9626 | 1020.1081 | 905.527 | 334.1328 | 402.5644 | 526.5851 | 389.8299 | 437.4523 | 554.8466 | 688.527 | 243.7758 | 865.1121 | 1192.054 | 1452.9046 | 1152.7472 | 417.8839 | 555.5103 | 646.3402 | 347.7512 | 223.473 | 217.3983 | 232.054 | 422.6224 | 635.5851 | 865.527 | 1017.0954 | 526.8092 | 410.3195 | 554.8839 | 618.6224 | 359.2822 | 483.6931 | 720.7719 | 836.2448 | 509.8712 | 358.8132 | 474.7758 | 458.8092 | 300.1288 | 806.3983 | 710.2989 | 886.6765 | 755.0994 | 467.7218 | 626.942 | 1082.0914 | 686.8672 | 556.942 | 735.9626 | 898.469 | 762.3156 | 479.4897 | 636.6765 | 877.2655 | 778.9547 | 337.7925 | 397.6805 | 508.3569 | 428.5477 | 408.054 | 496.1701 | 593.3569 | 361.3362 | 359.7178 | 460.5063 | 617.9626 | 395.996 | 851.4857 | 1290.1328 | 1814.1788 | 1306.1328 | 949.0374 | 1714.1328 | 2077.4523 | 1816.6598 | 227.2822 | 153.4149 | 156.8672 | 227.2822 | 220.5644 | 263.2281 | 308.2448 | 295.3569 | 217.4897 | 273.3029 | 305.4897 | 244.1494 | 130.2448 | 165.3195 | 187.4149 | 206.7552 | 175.6057 | 190.4316 | 209.5437 | 158.6598 | 652.0914 | 684.4149 | 894.1868 | 519.2281 | 491.942 | 686.5063 | 903.6391 | 507.5644 | 921.7305 | 1246.3362 | 1547.8839 | 926.3943
636 | 8.9904 | 708.8554 | 948.1446 | 916.5639 | 661.3157 | 444.4651 | 666.5349 | 670.7783 | 439.8578 | 323.6843 | 283.6867 | 417.2675 | 369.6819 | 567.7566 | 632.4145 | 703.2458 | 659 | 536.3157 | 486.5301 | 529.4434 | 522.1253 | 433.0241 | 969.0747 | 1157.4458 | 553.6578 | 693.5614 | 799.0048 | 1012.4386 | 890.7325 | 336 | 419.7301 | 550.3398 | 364.0506 | 403.0024 | 573.7542 | 698.6819 | 268.5108 | 857.1711 | 1217.9976 | 1404.1253 | 1103.0265 | 401.9518 | 563.6337 | 671.3398 | 363.7542 | 218.9518 | 202.0747 | 231.4386 | 389.8072 | 596.5639 | 851.5373 | 1032.8771 | 504.0988 | 375.612 | 562.2916 | 642.9494 | 382.7542 | 478.853 | 669.3205 | 809.2458 | 481.1735 | 335.8313 | 428.6867 | 449.7807 | 302.2675 | 774.9783 | 699.5855 | 895.0964 | 675.641 | 433.5157 | 592.4169 | 1045.8313 | 687.4145 | 545.7566 | 708.8313 | 932.8506 | 785.8265 | 465.7325 | 591.9783 | 827.0771 | 807.8988 | 346.0482 | 405.0241 | 519.2916 | 411.4892 | 396.9759 | 469.6361 | 613.9494 | 360.0482 | 334.6361 | 467.2916 | 609.3422 | 405.1205 | 848.4627 | 1274.294 | 1757.2217 | 1213.8361 | 926.2699 | 1658.053 | 2030.9301 | 1765.906 | 220.1711 | 136.9036 | 143.7325 | 220.1711 | 211.9759 | 284.9012 | 293.3181 | 287.7807 | 240.7542 | 247.0265 | 317.4867 | 266.559 | 159.5831 | 161.1952 | 207.1205 | 216.0482 | 161.6843 | 186.4145 | 231.388 | 161.0241 | 640.7807 | 684 | 866.4651 | 530.4386 | 485.5855 | 698.9253 | 886.5373 | 486.0988 | 893.3181 | 1211.2458 | 1524.7084 | 958.0217
637 | 9.0014 | 702.1829 | 918.8643 | 845.7167 | 616.9734 | 450.4395 | 588.6547 | 653.4867 | 467.6726 | 300.2065 | 306.1357 | 396.3038 | 392.6106 | 557.8289 | 619.9262 | 716.6106 | 637.7551 | 483.8023 | 445.0825 | 566.4779 | 510.1829 | 453.6106 | 947.6578 | 1135.8171 | 532.3038 | 691.1829 | 824.6726 | 1018.9027 | 914.9646 | 310.6696 | 445.4278 | 534.7551 | 389.0502 | 417.062 | 543.23 | 696.2802 | 248.4867 | 868.2567 | 1203.1209 | 1396.2684 | 1113.9764 | 397.5487 | 552.7433 | 661.4749 | 356.9262 | 205.5605 | 176.5722 | 224.8289 | 390.1711 | 595.1711 | 819.584 | 995.9616 | 469.7669 | 372.0856 | 580.708 | 622.8525 | 360.4011 | 497.3422 | 674.3422 | 768.6814 | 472.8171 | 346.1593 | 429.8909 | 400.7905 | 304.1829 | 709.6194 | 663.0471 | 858.1445 | 641.5605 | 416.2802 | 571.9262 | 1034.9144 | 637.3392 | 478.6312 | 691.3776 | 878.4483 | 760.6814 | 460.5487 | 639.3806 | 827.0738 | 786.4985 | 321.8525 | 396.1947 | 468.2536 | 456.0266 | 376.3894 | 459.7316 | 596.3038 | 333.4011 | 325.5487 | 473.4513 | 548.9852 | 417.354 | 853.4513 | 1240.584 | 1709.1445 | 1174.6578 | 911.1947 | 1588.4365 | 2069.0266 | 1694.3509 | 242.6962 | 139.5369 | 172.0502 | 242.6962 | 204.4631 | 262.8525 | 289.1829 | 275.5605 | 243.7316 | 232.646 | 307.9262 | 237.8643 | 160.4631 | 174.0738 | 181.7669 | 211.4631 | 160.2684 | 173.9262 | 207.1327 | 146.4749 | 625.292 | 682.3658 | 831.584 | 488.6932 | 479.2802 | 665.5958 | 828.4365 | 513.8673 | 913.6962 | 1155.1563 | 1462.1681 | 971.9027
638 | 9.0124 | 702.8294 | 945.8349 | 866.8349 | 578.6298 | 429.7768 | 555.0708 | 604.8711 | 508.5173 | 300.147 | 301.559 | 361.971 | 349.7241 | 547.706 | 562.2595 | 681.1887 | 629.6824 | 446.118 | 400.3122 | 514.1361 | 484.2886 | 445.559 | 905.8475 | 1131.0236 | 509.5826 | 639.5771 | 765.9891 | 985.1651 | 863.7713 | 344.8113 | 422.6298 | 502.6298 | 413.4174 | 414.5354 | 532.5118 | 693.3412 | 254.6116 | 796.3539 | 1173.118 | 1392.853 | 1070.7241 | 393.6824 | 545.0236 | 666.294 | 369.7586 | 203.8294 | 206.0055 | 227.147 | 387.853 | 620.2232 | 818.8058 | 968.7532 | 488.7114 | 360.559 | 566.5826 | 580.5063 | 354.8294 | 484.412 | 652.2886 | 782.5644 | 473.8294 | 333.412 | 414.7532 | 438.9347 | 310.6352 | 720.2232 | 603.5771 | 840.706 | 657.2468 | 424.9528 | 570.6588 | 1001.6533 | 670.7877 | 502.1525 | 690.6588 | 794.3303 | 735.5826 | 460.8294 | 612.1887 | 800.971 | 701.0363 | 342.5408 | 403.1234 | 487.5173 | 458.706 | 391.0762 | 458.1706 | 569.4356 | 340.2704 | 320.0236 | 439.1651 | 556.735 | 415.6824 | 778.5245 | 1211.6062 | 1720.735 | 1144.4592 | 898.2178 | 1532.0417 | 1983.7894 | 1641.1887 | 205.7005 | 132.706 | 158.265 | 205.7005 | 201.3412 | 226.3122 | 288.1706 | 297.882 | 215.6298 | 257.5408 | 287.5826 | 252.3938 | 124.9945 | 169.5354 | 148.4592 | 185.9474 | 164.3176 | 165.1942 | 207.4646 | 143.8294 | 614.5354 | 647.1651 | 855.6878 | 483.9764 | 492.9292 | 640.412 | 827.6116 | 530.9528 | 912.3648 | 1148.147 | 1499.5935 | 912.4537
639 | 9.0234 | 680.2743 | 844.4217 | 820.9734 | 616.768 | 440.9922 | 584.826 | 634.7179 | 470.3903 | 305.2085 | 293.9498 | 342.5329 | 340.8417 | 521.5909 | 566.1505 | 709.1928 | 622.2665 | 484.768 | 409.9922 | 503.5251 | 498.3589 | 433.7414 | 889.5831 | 1128.3166 | 525.3589 | 619.7414 | 758.7335 | 983.8918 | 887.3511 | 315.9655 | 415.3166 | 523.0925 | 363.1317 | 420.5752 | 516.0078 | 677.8495 | 251.1082 | 820.5595 | 1107.3981 | 1359.174 | 1034.0658 | 399.7335 | 520.4326 | 648.6411 | 361.2163 | 189.6912 | 189.4326 | 217.8997 | 377.7414 | 558.5063 | 805.6912 | 905.2398 | 486.4248 | 369.7837 | 570.5752 | 594.0423 | 363.4169 | 445.6489 | 665.5172 | 733.1317 | 487.4671 | 335.8918 | 419.4169 | 433.8495 | 314.5251 | 701.4326 | 632.5595 | 842.3668 | 703.9263 | 424.4749 | 587.8338 | 1000.0502 | 676.6332 | 520.1003 | 702.7837 | 838.9185 | 752.2006 | 457.6332 | 574.5408 | 767.2242 | 760.7523 | 332.8495 | 410.8918 | 478.8495 | 439.3245 | 382.3746 | 439.4828 | 562.4749 | 331.2163 | 312.2665 | 444.6254 | 571.1003 | 389.7492 | 800.8762 | 1191.8495 | 1615.58 | 1205.4937 | 903.5752 | 1493.5408 | 1919.3903 | 1695.2852 | 240.768 | 152.0423 | 154.8417 | 240.768 | 216.1505 | 225.0502 | 264.4326 | 284.4828 | 200.7414 | 233.5408 | 268.0078 | 247.7414 | 147.4592 | 168.8417 | 174.9843 | 207.9342 | 184.3589 | 173.2586 | 187.116 | 159.9922 | 614.6834 | 614.7492 | 844.3245 | 490.0502 | 458.9655 | 643.7335 | 861.1426 | 534.6834 | 853.7649 | 1197.8182 | 1388.7963 | 880.6411
640 | 9.0344 | 634.1548 | 893.1727 | 827.2257 | 589.9867 | 442.292 | 564.5265 | 613.8185 | 495.3054 | 290.6283 | 268.2345 | 390.385 | 329.0442 | 488.8185 | 545.9424 | 665.615 | 642.3496 | 463.9646 | 427.5177 | 547.5532 | 478.9424 | 407.2345 | 872.6283 | 1070.7831 | 521.3141 | 636.7876 | 815.5089 | 976.4602 | 836.7831 | 325.2257 | 372.3008 | 515.1902 | 399.9691 | 410.0442 | 472.0088 | 690.3717 | 246.5841 | 831.4159 | 1116.4956 | 1345.3141 | 1005.9424 | 388.1902 | 530.5177 | 646.708 | 322.2787 | 195.5398 | 194.3938 | 217.708 | 393.9337 | 589.5532 | 737.2433 | 892.292 | 448.2787 | 410.4293 | 501.1194 | 552.7389 | 324.0088 | 449.2478 | 647.6504 | 758.8673 | 468.6504 | 343.8319 | 408.1902 | 446.3717 | 300.4823 | 646.615 | 629.7522 | 836.1681 | 661.615 | 412.0442 | 544.0309 | 979.6504 | 638.4247 | 502.3584 | 684.5044 | 850 | 754.292 | 476.6416 | 559.6062 | 767.9779 | 741.6946 | 293.2787 | 377.8407 | 497.3496 | 441.1239 | 381.708 | 418.9203 | 540.6504 | 323.8761 | 307.5841 | 423.0885 | 543.1106 | 352.5486 | 774.6946 | 1131.7831 | 1595.854 | 1183.8407 | 918.6637 | 1464.9424 | 1880.2566 | 1626.7256 | 227.6504 | 159.4159 | 131.9424 | 227.6504 | 244.6195 | 227.708 | 251.4602 | 292.1018 | 228.8894 | 229.708 | 279.5177 | 271.6195 | 147.8761 | 160.4602 | 191.2478 | 184.8185 | 161.5265 | 186.9558 | 188.1239 | 168.1239 | 608.1681 | 646.7213 | 774.3893 | 491.854 | 408.4468 | 611.6946 | 813.199 | 489.7389 | 846.4159 | 1113.0618 | 1425.7788 | 891.5177
641 | 9.0454 | 611.4071 | 853.8898 | 773.6922 | 590.9302 | 423.8082 | 565.1976 | 622.8604 | 454.2212 | 282.8024 | 268.3314 | 346.1573 | 332.1976 | 486.5988 | 519.4769 | 614.5525 | 618.2792 | 442.5408 | 396.2153 | 509.0875 | 450.7443 | 414.2616 | 866.5349 | 1072.2616 | 468.8898 | 600.1514 | 713.2506 | 978.4651 | 820.0698 | 300.2792 | 411.7788 | 511.2675 | 393.2733 | 391.5408 | 474.529 | 604.64 | 239.0698 | 808.6106 | 1053.2969 | 1268.6341 | 987.2733 | 379.0698 | 480.1573 | 629.6745 | 318.5988 | 196.1337 | 196.8663 | 221.4651 | 336.3608 | 586.6047 | 739.8604 | 872.942 | 478.3835 | 384.0816 | 498.529 | 563.1918 | 308.3373 | 444.9361 | 647.4651 | 715.6224 | 502.2498 | 295.6224 | 383.4769 | 397.8898 | 287.6047 | 634.5349 | 594.2153 | 770.8957 | 616.6165 | 383.1455 | 543.7965 | 914.4945 | 606.878 | 471.2792 | 613.4306 | 862.9941 | 667.3725 | 432.3549 | 528.5467 | 749.942 | 725.6047 | 325.9824 | 383.2616 | 481.942 | 417.7443 | 340.4188 | 422.9302 | 534.4012 | 286.6165 | 292.2733 | 419.8663 | 527.6047 | 380.7847 | 744.7443 | 1083.0176 | 1583.0059 | 1157.2094 | 820.5122 | 1393.4306 | 1816.9596 | 1604.3373 | 211.8722 | 119.2851 | 142.7267 | 211.8722 | 211.7502 | 247.058 | 245.6686 | 275.8082 | 225.3373 | 236.0639 | 276.8024 | 274.2675 | 140.9361 | 155.5349 | 196.4651 | 209.4533 | 154.4012 | 159.5467 | 194.1976 | 136.0816 | 585.3431 | 614.8839 | 800.2498 | 477.2733 | 403.5988 | 642.3835 | 771.9478 | 514.0522 | 807.1514 | 1032.8957 | 1367.5643 | 870.6106
642 | 9.0564 | 601.9715 | 852.15 | 774.6642 | 584.9715 | 405.1858 | 571.2714 | 559.9861 | 433.8215 | 260.0359 | 275.15 | 320.4287 | 334.7571 | 487 | 497.5501 | 585.1573 | 646.5069 | 437.3643 | 360.2502 | 473.0073 | 486.5355 | 405.4572 | 846.6716 | 1040.6145 | 447.8215 | 585.2143 | 717.9641 | 907.8362 | 818.1214 | 317.9641 | 366.2861 | 511.8786 | 370.793 | 384.6072 | 502.357 | 651.621 | 217.793 | 732.0791 | 989.4075 | 1278.0856 | 1026.2926 | 376.2429 | 504.5069 | 637.5428 | 353.2641 | 166.1288 | 152.1932 | 216.7286 | 348.9641 | 599.057 | 719.1573 | 884.057 | 438.3146 | 381.6357 | 482.4287 | 526.4645 | 325.4499 | 436.9715 | 597.0432 | 678.1288 | 475.4931 | 300.2999 | 389.6642 | 426.0212 | 273.7001 | 657.7213 | 600.4214 | 733.7645 | 654.9283 | 420.2926 | 558.1785 | 886.5501 | 604.7571 | 485.4499 | 693.4923 | 841.2787 | 718.1353 | 446.9641 | 526.6357 | 782.1426 | 701.1573 | 326.7286 | 390.2714 | 460.5501 | 416.6357 | 364.7498 | 400.2787 | 534 | 323.1711 | 284.7286 | 403.3073 | 540.9356 | 378.0929 | 734.85 | 1075.2429 | 1567.8215 | 1133.793 | 859.4711 | 1355.7645 | 1834.357 | 1572.2502 | 224.057 | 140.2355 | 150.2714 | 224.057 | 191.943 | 202.5575 | 255.5428 | 295.8427 | 203.793 | 258.0856 | 299.7213 | 234.4645 | 148.7857 | 157.6357 | 182.943 | 228.9356 | 150.4857 | 171.6927 | 189.7286 | 161.7498 | 555.643 | 637.1141 | 773.4931 | 457.4287 | 416.2999 | 604.0717 | 725.7074 | 517.2429 | 815.9356 | 1021.2429 | 1320.1002 | 772.1149
643 | 9.0674 | 623.2714 | 852.1091 | 743.0383 | 550.1474 | 390.528 | 508.7492 | 553.6726 | 436.4543 | 280.1623 | 265.3097 | 348.2891 | 331.4366 | 454.9292 | 532.416 | 617.6343 | 572.6046 | 420.0737 | 357.7817 | 481.472 | 472.4012 | 395.0914 | 780.7492 | 1045.7994 | 511.9234 | 643.6874 | 698.7286 | 877.5103 | 811.764 | 308.528 | 367.9086 | 534.2714 | 347.5103 | 373.3097 | 464.1297 | 643.0737 | 253.3068 | 722.8909 | 1014.0708 | 1249.82 | 974.0943 | 354.6194 | 461.7846 | 580.2035 | 343.7463 | 190.5074 | 197.6697 | 199.0737 | 365.3628 | 541.3126 | 683.0383 | 790.6783 | 459.6166 | 378.4366 | 464.8554 | 545.944 | 335.1269 | 429.764 | 614.944 | 691.8171 | 474.5634 | 289.528 | 352.6932 | 426.5457 | 267.5457 | 612.0029 | 602.8909 | 748.708 | 597.7492 | 400.1652 | 515.4572 | 910.7257 | 618.3628 | 502.2537 | 644.5309 | 815.7286 | 710.0737 | 416.1474 | 535.9086 | 746.0206 | 700.6726 | 318.8731 | 357.1474 | 433.9469 | 399.1829 | 360.9823 | 404.1269 | 534 | 309.4012 | 277.764 | 411.6903 | 515.3834 | 400.1623 | 757.1623 | 1052.7286 | 1546.4012 | 1089.1297 | 805.5309 | 1360.0177 | 1863.8348 | 1512.7669 | 234.0177 | 130.6372 | 148.3274 | 234.0177 | 201.5811 | 220.944 | 241.8554 | 264.2566 | 191.2006 | 233.3834 | 273.6017 | 236.1269 | 130.4012 | 156.2183 | 189.9086 | 185.5663 | 156.236 | 181.1269 | 186.3274 | 150.8554 | 569.8171 | 579.5309 | 782.3628 | 471.9263 | 385.9292 | 595.3274 | 750.2891 | 462.6577 | 784.1474 | 984.4749 | 1291.8377 | 819.5783
644 | 9.0784 | 579.9395 | 786.9734 | 788.6731 | 561.3535 | 422.4165 | 512.7409 | 566.6441 | 446.9346 | 282.0969 | 249.5496 | 353.8063 | 344.5472 | 484.4165 | 518.937 | 570.5206 | 593.707 | 427.0315 | 388.707 | 418.8765 | 446.5182 | 403.0315 | 773.9031 | 1030.7433 | 476.6489 | 575.1356 | 698.7094 | 881.322 | 808.2906 | 289.8402 | 387.9661 | 466.5545 | 329.6465 | 419.8668 | 473.4504 | 620.2276 | 238.1308 | 758.2228 | 996.3245 | 1219.8087 | 1005.8353 | 377.2881 | 476.7724 | 577.5157 | 344.7094 | 179.5496 | 169.6804 | 193.3874 | 372.4189 | 506.9056 | 649.1961 | 805.1913 | 466.6126 | 384.322 | 478.3535 | 489.3923 | 335.0969 | 401.9056 | 552.9734 | 663.293 | 433.4552 | 301.5472 | 347.0969 | 381.7458 | 291.385 | 614.1283 | 569.5835 | 729.3245 | 617.0944 | 383.4528 | 531.6755 | 894.1308 | 561.2954 | 511.2252 | 627.0654 | 791.3245 | 694.4528 | 429.1598 | 573.1259 | 698.5521 | 692.8717 | 325.2252 | 338.5496 | 458.0944 | 393.3874 | 332.9056 | 405.9031 | 506.9056 | 329.5787 | 297.7724 | 412.0969 | 500.2591 | 368.6804 | 717.5521 | 996.7143 | 1490.7143 | 1096.6441 | 800.7094 | 1426.0266 | 1769.4601 | 1512.322 | 225.0654 | 144.3535 | 162.4504 | 225.0654 | 175.9056 | 233.0315 | 273.4165 | 279.8692 | 216.1913 | 221.8717 | 257.3559 | 231.5811 | 137.9346 | 149.678 | 174.7433 | 181.8063 | 142.5496 | 163.937 | 186.9031 | 141.7748 | 569.2906 | 583.7409 | 732.5206 | 449.615 | 396.4504 | 561.5835 | 744.0654 | 468.6441 | 829.6731 | 995.3535 | 1280.8717 | 854.9976
645 | 9.0894 | 564.9302 | 753.4523 | 730.8347 | 537.3677 | 382.0589 | 563.4337 | 522.2281 | 439.7611 | 284.7463 | 233.353 | 303.6509 | 354.2389 | 462.3677 | 473.0589 | 590.6323 | 557.6361 | 399.6214 | 380.0993 | 441.294 | 468.7168 | 388.4375 | 790.4779 | 974.9891 | 467.4228 | 571.5772 | 719.1397 | 902.1397 | 791.5221 | 294.4081 | 339.6509 | 463.5772 | 349.9705 | 368.2428 | 433.8053 | 616.1691 | 215.8603 | 730.8751 | 972.9449 | 1230.7316 | 999.9302 | 332.3972 | 459.7758 | 535.8898 | 309.2979 | 204.5477 | 177.9007 | 188.4228 | 354.6912 | 493.9302 | 658.8161 | 805.2537 | 433.1288 | 379.5074 | 449.7905 | 525.1102 | 296.5516 | 431.0404 | 526.7758 | 680.1397 | 468.3639 | 265.467 | 328.6912 | 396.2242 | 279.353 | 601.2684 | 534.8751 | 699.5368 | 586.1288 | 380.1691 | 494.6361 | 836.1582 | 613.5881 | 464.3972 | 548.1877 | 726.7502 | 693 | 437.4081 | 558.6912 | 670.1986 | 664.5368 | 298.5368 | 337 | 438.114 | 414.9705 | 336.4081 | 387.6912 | 482.9449 | 303.6214 | 289.9302 | 401.0147 | 479.7758 | 383.3088 | 756.0256 | 1041.3491 | 1528.0256 | 1065.9596 | 772.6214 | 1247.1659 | 1706.82 | 1523.0698 | 228.5772 | 126.7758 | 162.1691 | 228.5772 | 180.3235 | 227.5919 | 229.3972 | 276.5074 | 223.5772 | 197.1986 | 244.0993 | 233.7463 | 128.9302 | 162.7316 | 158.353 | 172.8456 | 163.886 | 149.1839 | 178.7611 | 134.5919 | 547.0295 | 564.8603 | 671.1582 | 470.8014 | 397.0846 | 561.6618 | 687.1582 | 465.4228 | 810.2832 | 962.2133 | 1263.5221 | 832.3677
646 | 9.1004 | 559.3614 | 762.1328 | 716.6504 | 561.9042 | 370.5782 | 482.6489 | 515.2168 | 426.9395 | 294.2773 | 259.8319 | 297.1445 | 313.2522 | 471.1328 | 483.8437 | 541.9749 | 551.2168 | 381.2286 | 349.3126 | 458.8437 | 469.1445 | 351.7463 | 773.4454 | 955.1563 | 439.1681 | 568.2891 | 674.6135 | 889.1563 | 769.5899 | 313.5546 | 335.9277 | 463.0723 | 340.8672 | 382.4823 | 459.6874 | 600.2286 | 207.5059 | 652.8539 | 906.9866 | 1202.3126 | 930.348 | 339.1328 | 455.2168 | 557.0487 | 320.8437 | 167.1076 | 171.5782 | 194.4941 | 359.4941 | 499.4941 | 674.8437 | 779.0236 | 425.3614 | 382.7109 | 448.8555 | 529.9277 | 303.205 | 406.1681 | 572.3142 | 662.5177 | 417.2639 | 241.5899 | 339.9882 | 407.2773 | 275.2168 | 546.1917 | 556.121 | 738.7478 | 623.8201 | 387.4218 | 516.9764 | 872.7478 | 584.674 | 490.6151 | 622.6402 | 708.9395 | 705.9882 | 420.3731 | 579.2655 | 719.9529 | 697.2537 | 305.2773 | 346.2773 | 423.0118 | 398.4454 | 345.3496 | 402.6991 | 454.0958 | 318.6269 | 300.1328 | 404.6386 | 508.6151 | 339.5412 | 716.3967 | 1034.8672 | 1442.0102 | 1054.6504 | 790.4101 | 1246.6991 | 1639.8421 | 1498.0236 | 217.8672 | 154.6874 | 142.5177 | 217.8672 | 172.6504 | 222.3614 | 254.6874 | 268.5782 | 186.8908 | 253.447 | 247.6386 | 224.7227 | 123.3614 | 151.0118 | 156.0723 | 189.6269 | 160.4336 | 175.8319 | 211.3983 | 139.5664 | 536.6504 | 561.1445 | 724.447 | 487.8437 | 396.0723 | 560.1445 | 721.8924 | 443.6622 | 725.432 | 923.7463 | 1202.6253 | 780.8303
647 | 9.1114 | 556.1799 | 694.3785 | 723.5202 | 523.5791 | 377.5202 | 504.2006 | 533.8004 | 437.2803 | 260.2192 | 234.7394 | 310.1603 | 313.7601 | 457.8997 | 455.8594 | 548.3402 | 556.6401 | 427.001 | 380.8408 | 435.5595 | 447.3795 | 389.4209 | 780.4602 | 941.7797 | 484.9411 | 531.3392 | 642.7994 | 803.3981 | 797.1406 | 291.4995 | 346.3402 | 467.7001 | 334.3599 | 352.0393 | 407.4788 | 586.7797 | 223.9204 | 690.241 | 901.06 | 1179.3195 | 936.2803 | 327.7797 | 427.7394 | 559 | 334.2203 | 160.2399 | 163.4798 | 202.5202 | 366.5801 | 498.12 | 682.5801 | 750.6794 | 425.94 | 363.2596 | 463.1003 | 495.2192 | 261.6991 | 402.12 | 562.8397 | 677.9204 | 410.1799 | 274.7808 | 330.6598 | 360.0589 | 284.4002 | 562.6805 | 565.5202 | 722.2596 | 563.0786 | 370.1396 | 502.0796 | 853.4395 | 581.06 | 458.2192 | 556.6184 | 698.5998 | 675.0393 | 419.94 | 503.9183 | 672.2989 | 650.1789 | 260.879 | 328.1996 | 470.8811 | 439.3009 | 352.5801 | 375.7994 | 465.1603 | 310.5998 | 285.9597 | 355.1789 | 523.2203 | 344.4602 | 644.3785 | 1044.3402 | 1382.3588 | 1022.9793 | 756.2792 | 1263.9804 | 1646.2803 | 1438.6587 | 237.6805 | 130.6794 | 142.88 | 237.6805 | 188.9204 | 206.0196 | 246.6598 | 276.4602 | 189.6401 | 206.2989 | 261.1603 | 243.7404 | 128.6401 | 139.6598 | 149.4199 | 175.0196 | 175.0403 | 150.7394 | 191.4395 | 130.5998 | 563.2606 | 599.5409 | 675.4188 | 441.999 | 405.4002 | 535.5595 | 705.2596 | 455.1603 | 747.2006 | 961.4209 | 1195.1799 | 776.06
648 | 9.1224 | 519.8121 | 711.9032 | 634.4827 | 506.7153 | 365.6199 | 516.4754 | 500.7168 | 418.0015 | 259.9046 | 230.1431 | 306.2384 | 309.2384 | 461.7616 | 423.526 | 580.4263 | 575.0939 | 400.4783 | 375.3815 | 438.7616 | 484.0925 | 371.0491 | 690.5304 | 953.3801 | 459.4306 | 536.6185 | 636.2384 | 798.9523 | 759.9552 | 295.7139 | 347.9523 | 435.6214 | 348.2847 | 350.9523 | 444.9494 | 540.289 | 265.9494 | 695.8569 | 922.9032 | 1181.8092 | 917.9538 | 335.5708 | 412.6676 | 558.0477 | 265.4812 | 178.0939 | 158.2384 | 183.9538 | 398.4263 | 458.0029 | 622.052 | 791.854 | 424.0954 | 317.2413 | 440.1922 | 457.7645 | 309.4725 | 377.2399 | 556.2861 | 643.7645 | 431.9032 | 258.9061 | 325.2384 | 367.4754 | 247.8598 | 554.4769 | 512.6705 | 685.7645 | 594.2355 | 356.6199 | 491.4769 | 826.2876 | 541.0029 | 431.2399 | 595.8063 | 700.8569 | 623.4798 | 422.8569 | 476.1445 | 731.8526 | 623.1922 | 288.474 | 324.1431 | 430.1937 | 372.4812 | 345.3815 | 368.2861 | 426.9552 | 308.0954 | 285.9523 | 371.9985 | 471.6228 | 353.5708 | 671.4263 | 906.9148 | 1382.8092 | 993.383 | 714.9552 | 1144.0564 | 1658.4277 | 1464.5217 | 213.2876 | 141.3801 | 145.8569 | 213.2876 | 194.7616 | 232.617 | 224.0968 | 244.6214 | 187.1431 | 214.4277 | 239.1445 | 198.3367 | 117.5723 | 149.4754 | 140.4292 | 172.0954 | 145.526 | 182.3309 | 178.5723 | 151.9032 | 549.763 | 556.289 | 728.1864 | 465.6647 | 383.1445 | 506.383 | 661.146 | 451.2384 | 695.6705 | 910.6705 | 1122.6243 | 817.9017
649 | 9.1334 | 521.8584 | 697.5664 | 681.124 | 495.3894 | 354.3894 | 489.0265 | 496.1062 | 399.6371 | 295.6903 | 230.9646 | 325.2921 | 340.8319 | 489.0089 | 460.5841 | 525.0884 | 572.1416 | 364.2743 | 347.0265 | 424.531 | 413.6548 | 407.6195 | 654.1681 | 834.3892 | 454.1416 | 509.9911 | 620.5664 | 819.2567 | 762.823 | 271.8849 | 355.7168 | 424.354 | 339.354 | 339.4248 | 422.8849 | 534.1416 | 207.23 | 684.4248 | 851.6548 | 1077.8229 | 898.6725 | 312.8495 | 410.0708 | 530.9911 | 289.9735 | 188.646 | 172.469 | 192.646 | 356.5929 | 460.823 | 591.9911 | 754.4513 | 428.823 | 333.3275 | 421.6371 | 470.469 | 327.4336 | 349.9557 | 499.0884 | 606.3097 | 389.5929 | 263.7876 | 363.5841 | 369.9292 | 248.8938 | 522.1681 | 501.3186 | 670.4956 | 585.3894 | 338.6371 | 504.5044 | 797.0265 | 562.1505 | 404.9203 | 561.3451 | 666.2743 | 680.8054 | 380.5575 | 460.531 | 639.5043 | 653.8319 | 286.1416 | 334.6106 | 387.4867 | 346.8141 | 340.177 | 372.823 | 463.5841 | 302.2124 | 292.7522 | 370.1062 | 496.0089 | 360.7522 | 684.5752 | 991.6373 | 1382.0354 | 957.2743 | 721.6814 | 1171.7611 | 1522.991 | 1410.053 | 209.1062 | 122.7079 | 156.6106 | 209.1062 | 169.9203 | 209.8849 | 254.8319 | 240.1062 | 192.7876 | 202.4602 | 263.0797 | 201.7876 | 112.177 | 124.9203 | 144.823 | 197.0797 | 157.5044 | 173.3894 | 200.1859 | 124.0619 | 498.8406 | 551.1062 | 664.4424 | 438.0619 | 394.5398 | 503.0708 | 603.053 | 438.4602 | 689.1416 | 872.3097 | 1103.5664 | 770.8052
650 | 9.1444 | 491.7163 | 710.6765 | 626.3401 | 526.2606 | 387.2144 | 439.1553 | 525.3068 | 453.7061 | 249.1322 | 228.0693 | 299.6239 | 347.8614 | 454.8318 | 431.7163 | 514.208 | 517.2939 | 383.5148 | 376.2837 | 439.6303 | 435.4223 | 322.0564 | 678.3992 | 897.4057 | 418.8318 | 507.0462 | 673.7292 | 792.647 | 751.2773 | 272.9538 | 329.6239 | 429.8614 | 324.3466 | 349.7458 | 416.1386 | 561.353 | 228.4455 | 684.9769 | 777.6867 | 1095.4917 | 877.4852 | 333.4917 | 468.6136 | 515.3466 | 318.353 | 161.647 | 168.1155 | 188.1155 | 363.792 | 481.5148 | 547.0398 | 709.0398 | 418.2542 | 366.2375 | 397.5545 | 475.8845 | 300.647 | 419.3364 | 498.9538 | 622.5841 | 435.8678 | 284.5148 | 326.9011 | 353.3928 | 271.4686 | 460.4326 | 485.3697 | 672.9307 | 569.3697 | 365.353 | 482.5314 | 762.7856 | 565.9307 | 412.792 | 532.647 | 684.5379 | 625.3633 | 395.6072 | 482.4686 | 670.1913 | 614.9473 | 304.561 | 312.5314 | 418.2375 | 380.1913 | 308.7394 | 341.7394 | 472.8152 | 319.5841 | 279.3235 | 358.2773 | 479.4159 | 393.2375 | 638.1091 | 888.5186 | 1361.4852 | 1003.8678 | 725.9076 | 1066.5186 | 1535.5841 | 1362.086 | 209 | 114.1848 | 138.439 | 209 | 169.9769 | 207.0462 | 227.6701 | 235.1155 | 184.208 | 233.2606 | 212.2246 | 196.1386 | 128.6072 | 130.8383 | 147.9307 | 175.5314 | 154.0924 | 153.4621 | 173.647 | 138.6303 | 528.2606 | 538.3004 | 692.2837 | 469.2375 | 382.3004 | 503 | 626.3992 | 462.4223 | 704.6303 | 828.0167 | 1114.7227 | 754.3466
651 | 9.1554 | 463.3028 | 679.3461 | 634.8919 | 481.4975 | 347.4434 | 459.7621 | 514.1298 | 430.2704 | 278.6647 | 267.5674 | 303.9459 | 312.3893 | 430.2596 | 428.0324 | 497.1839 | 516 | 383.0108 | 360.1839 | 414.2812 | 402.3677 | 370.4484 | 748.243 | 808.9842 | 479.3295 | 527.7729 | 631.4759 | 742.5407 | 785.6215 | 275.9676 | 329 | 451.7621 | 330.9243 | 305.4867 | 397.2055 | 569.9135 | 199.3244 | 643.4542 | 856.124 | 1083.1406 | 890.8486 | 296.411 | 410.6489 | 524.8919 | 295.2596 | 186.7188 | 177.8919 | 164.2596 | 337.292 | 397.9193 | 574.6864 | 712.9459 | 431.8486 | 333.3677 | 433.5998 | 444.3461 | 296.0433 | 379.4542 | 477.2379 | 600.2487 | 381.6056 | 284.0108 | 327.9784 | 346.0757 | 269.0324 | 500.5458 | 513.6864 | 621.5624 | 568.0108 | 332.3677 | 444.411 | 770.9027 | 549.1839 | 436.7404 | 545.8486 | 695.881 | 622.0216 | 360.3893 | 462.2271 | 607.6922 | 629.827 | 265.4326 | 302.1081 | 370.5299 | 326.5948 | 301.0757 | 359.7945 | 470.0324 | 311.0973 | 266.1406 | 340.1947 | 426.5732 | 349.4867 | 648.8702 | 928.535 | 1282.8544 | 1029.7296 | 683.465 | 1058.0649 | 1470.7138 | 1395.6215 | 204.0541 | 113.0108 | 141.9567 | 204.0541 | 175.9351 | 213.9243 | 232.9351 | 228.0757 | 191.9135 | 213.2271 | 258.4809 | 220.7296 | 135.9243 | 133.9676 | 145.0324 | 176.9784 | 144.1081 | 190.589 | 183.881 | 156.8053 | 524.0541 | 591.416 | 647.4975 | 484.8378 | 348.3677 | 486.1839 | 636.8919 | 443.2163 | 703.0216 | 819.0865 | 1005.2005 | 737.1839
652 | 9.1664 | 462.9779 | 665.9291 | 655.9808 | 491.0177 | 320.031 | 440.0295 | 503.9852 | 393.0251 | 227.0487 | 247.9646 | 304.9483 | 314.9705 | 453.9336 | 424.9793 | 531.9587 | 523.9705 | 384.9321 | 330.9823 | 408.9469 | 390.0325 | 368.9734 | 666.0856 | 852.9262 | 441.972 | 548.8863 | 575.0177 | 747.0266 | 717.0103 | 261.0162 | 304.0103 | 421.0059 | 322.969 | 365.9188 | 412.941 | 558.9055 | 208.003 | 650.8642 | 772.0428 | 1053.9572 | 839.9793 | 309.0133 | 401.034 | 487.9749 | 310.9528 | 168 | 157.031 | 199.9675 | 333.0207 | 436 | 589.9616 | 702.9852 | 425.9469 | 343.9646 | 414.9587 | 459.9306 | 287.9911 | 384.0251 | 513.9926 | 614.9439 | 422.9365 | 271.9897 | 332.9424 | 361.9823 | 271.9985 | 476.9956 | 492.9616 | 658.9646 | 547.0709 | 369.0177 | 458.9985 | 710.0635 | 536.9926 | 409.9395 | 537.0162 | 682.0487 | 639.9454 | 377.0354 | 477.9882 | 605.941 | 634.9793 | 285.0089 | 283.0649 | 412.9424 | 390.9956 | 302.0177 | 331.0133 | 424.997 | 296.9483 | 286.966 | 346.9779 | 442.9838 | 366.9616 | 657.9675 | 873.028 | 1268.9395 | 1008.9454 | 723.9572 | 1083.941 | 1450.9055 | 1371.9985 | 219.9705 | 131.9764 | 146.9749 | 219.9705 | 164.0059 | 223.9572 | 243.9646 | 271.9542 | 190.0177 | 187.0148 | 219.0546 | 212.997 | 114.9985 | 126.0207 | 171.9705 | 170.9793 | 154.0044 | 169.9572 | 180.9705 | 127.0207 | 506 | 534.9852 | 622.9779 | 459.0325 | 381.9114 | 501.9469 | 598.0044 | 462.9616 | 723.9144 | 817.0103 | 1006.9616 | 772.0207
653 | 9.1774 | 448.0551 | 617.9587 | 642.8348 | 503.1927 | 340.6283 | 459.6834 | 493.3529 | 410.1101 | 259.821 | 224.8398 | 270.6884 | 295.3442 | 409.2203 | 411.3992 | 503.8073 | 503.7797 | 339.7297 | 319.1652 | 373.3029 | 412.0964 | 350.8761 | 723.6421 | 803.5232 | 422.4355 | 471.766 | 586.7384 | 764.0225 | 723.4631 | 271.6834 | 310.6283 | 424.2015 | 301.9587 | 311.3442 | 373.2891 | 495.2753 | 210.2203 | 559.0826 | 801.3029 | 1025.7159 | 826.0138 | 317.8761 | 423.766 | 471.3029 | 279.6195 | 167.7522 | 177.821 | 177.8899 | 346.7247 | 436.0551 | 564 | 692.9587 | 389.8486 | 319.4906 | 387.592 | 413.2891 | 282.0688 | 400.174 | 508.5732 | 577 | 380.1927 | 264.7384 | 293.9174 | 350.0964 | 270.8761 | 474.0551 | 466.9036 | 634.8073 | 594.1877 | 380.0638 | 457.8348 | 753.1652 | 532.3029 | 369.4681 | 547.5732 | 713.995 | 603.2203 | 400.5044 | 469.9174 | 566.3166 | 620.6145 | 290.7797 | 326.1189 | 374.179 | 387.5457 | 314.2616 | 339.7109 | 422.4768 | 262.3304 | 264.234 | 332.3855 | 432.2616 | 341.3855 | 635.4355 | 891.4631 | 1227.1051 | 972.0138 | 694.1189 | 1043.1189 | 1386.9174 | 1370.9587 | 200.4681 | 116.1514 | 130.1239 | 200.4681 | 168.234 | 194.9587 | 219.7384 | 241.1101 | 201.821 | 197.179 | 256.1377 | 211.0138 | 114.5094 | 140.2203 | 152.0275 | 157.5094 | 157 | 141.0138 | 161.3029 | 140.9587 | 505.9449 | 524.5595 | 608.1377 | 480.5732 | 323.0188 | 466.3992 | 600.7109 | 436.587 | 666.0138 | 823.4493 | 981.5232 | 785.7935
654 | 9.1884 | 450.9577 | 614.6612 | 631.3648 | 516.7655 | 314.7557 | 436.7915 | 447.2345 | 417.9479 | 247.0521 | 283.8534 | 319.6091 | 319.0098 | 424.0098 | 439.4267 | 489.4528 | 488.2345 | 391.2183 | 330.9479 | 395.1824 | 418.6612 | 342.3648 | 698.2606 | 841.8339 | 382.0521 | 455.9642 | 567.8437 | 695.8762 | 684.9218 | 249.912 | 284.1042 | 368.2769 | 298.8176 | 336.2866 | 393.8697 | 513.8534 | 225.5831 | 564.9218 | 822.6612 | 1074.0815 | 827.2606 | 308.4267 | 406.4007 | 492.8176 | 323.7134 | 150.4951 | 164.6612 | 170.2085 | 327.2866 | 439.9739 | 564.3127 | 689.1401 | 378.2964 | 283.4951 | 429.0619 | 433.9739 | 286.7915 | 342.1466 | 478.2085 | 576.3746 | 394.4169 | 247.0944 | 289.0684 | 355.3844 | 260.8795 | 476.9577 | 459.4528 | 619.9056 | 536.1042 | 314.0684 | 446.1042 | 763.3323 | 554.2606 | 403.3388 | 517.5993 | 642.2085 | 618.7394 | 364.1922 | 462.6971 | 588.6352 | 592.7394 | 274.6352 | 264.6156 | 387.6515 | 355.0782 | 331.9577 | 319.912 | 385.7296 | 285.9218 | 280.557 | 358.9316 | 450.6612 | 367.8013 | 595.3127 | 852.2704 | 1163.1042 | 971.9056 | 631.7557 | 979.8437 | 1379.1238 | 1364.8991 | 233.2964 | 127.3388 | 138.9218 | 233.2964 | 184.8958 | 192.9902 | 200.8958 | 248.8697 | 188.6873 | 209.9739 | 264.7753 | 212.1824 | 150.4267 | 154.9837 | 153.531 | 193.5049 | 156.8697 | 142.8599 | 183.2085 | 137.8437 | 501.6091 | 492.5831 | 617.8437 | 449.1922 | 395.4267 | 494.2704 | 579.0098 | 407.7557 | 666.4528 | 783.7394 | 1018.114 | 769.8534
655 | 9.1994 | 412.8437 | 602.0383 | 643.6961 | 506.5811 | 342.6549 | 428.885 | 454.7728 | 414.8495 | 249.3068 | 240.233 | 304.1563 | 282.4218 | 388.8024 | 417.3481 | 468.9233 | 496.5015 | 362.9971 | 328.6549 | 401.5782 | 405.7316 | 356.115 | 704.4335 | 871.2743 | 385.9942 | 492.0383 | 563.4573 | 764.466 | 682.2301 | 281.584 | 288.4985 | 414.466 | 291.6165 | 345.1209 | 388.5398 | 471.6519 | 210.3451 | 562.6519 | 808.0058 | 965.2301 | 837.2684 | 287.4985 | 384.2301 | 485.5782 | 313.1534 | 169.4985 | 152.0767 | 178.3451 | 336.9646 | 439.4985 | 572.5486 | 654.5073 | 352.1534 | 302.8437 | 394 | 434.0738 | 280.0738 | 384.7316 | 485.2714 | 553.9587 | 408.5811 | 286.1976 | 326.4306 | 297.0325 | 221.1092 | 437.1563 | 440.534 | 578.6549 | 539.1563 | 353.1947 | 447.4689 | 702.8791 | 561.8524 | 416.5369 | 538.6194 | 650.5369 | 608.7316 | 335.0738 | 415.8024 | 575.2684 | 583.4602 | 261 | 324.1947 | 410.2743 | 357.8466 | 294.0738 | 352.9646 | 413.9587 | 283.6903 | 264.3068 | 319.1917 | 437.5782 | 324.0354 | 607.9204 | 824.5782 | 1166.5015 | 930.3481 | 656.9321 | 973.3097 | 1308.9233 | 1247.6194 | 207.9204 | 140.5369 | 136.3835 | 207.9204 | 180.7316 | 228.3893 | 197.4985 | 244.5752 | 178.7257 | 208.2714 | 218.1947 | 218.4631 | 128.5782 | 117.1917 | 136.3068 | 174.3097 | 151.7699 | 174.0796 | 190.0029 | 131.8466 | 485.5044 | 476.4248 | 612.6903 | 419.4602 | 372.8495 | 466.2714 | 543.0738 | 435.7699 | 645.9233 | 770.4335 | 982.5073 | 725.8879
656 | 9.2104 | 433.6962 | 603.3038 | 609.734 | 471.3545 | 333.5949 | 425.3923 | 425.8737 | 383.2654 | 254.7718 | 222.2154 | 282.4936 | 292.4936 | 432.6834 | 400.291 | 467.0506 | 484.0506 | 386.025 | 320.7596 | 389.2276 | 402.241 | 358.8987 | 615.8102 | 827.7468 | 433.9744 | 489.7084 | 598.7846 | 722.9109 | 689.266 | 223.8865 | 301.3545 | 376.9622 | 283.8737 | 298.9115 | 376.291 | 487.0378 | 217.734 | 578.5949 | 756.6834 | 970.7468 | 842.5314 | 299.3923 | 389.3417 | 473.4808 | 315.3289 | 180.1769 | 154.1519 | 186.6962 | 311.8609 | 450.6834 | 488.9371 | 596.0512 | 356.6583 | 325.1141 | 393.7468 | 458.6199 | 305.4808 | 376.4808 | 466.1391 | 577.5442 | 374.266 | 242.6711 | 265.7474 | 347.8987 | 271.1135 | 416.9622 | 478.3417 | 571.975 | 516.1769 | 332.0885 | 384.6077 | 748.3289 | 508.9622 | 429.4936 | 503.7974 | 662.1769 | 602.709 | 362.1519 | 458.6705 | 581.8987 | 596.975 | 262.1141 | 301.8731 | 364.7212 | 354.5064 | 320.7974 | 326.8481 | 438.0506 | 301.0506 | 270.6327 | 324.3038 | 426.8987 | 349.5442 | 629.3289 | 812.8353 | 1149.0879 | 911.9744 | 579.8987 | 956.8609 | 1306.3417 | 1217.9121 | 230.6455 | 152.6327 | 144.9872 | 230.6455 | 175.0128 | 189.0634 | 210.5064 | 257.4302 | 220.025 | 189.8987 | 199.5698 | 204.291 | 118.5064 | 122.6583 | 144.4558 | 158.1647 | 147.1141 | 151.8609 | 165.9115 | 128.1013 | 449.8231 | 462.6077 | 629.6455 | 430.0378 | 343.4936 | 448.9115 | 570.709 | 430.1519 | 642.3795 | 685.1525 | 922.0256 | 698.2532
657 | 9.2214 | 426.8043 | 606.6086 | 583.9233 | 480.1397 | 325.4336 | 414.3147 | 460.5595 | 334.5173 | 215.2026 | 242.3569 | 270.7345 | 284.1957 | 408.1259 | 388.8879 | 462.7767 | 485.7552 | 349.5035 | 333.3009 | 354.3638 | 457.5879 | 353.8534 | 631.5035 | 819.9793 | 398.3285 | 429.6991 | 573.7345 | 665.3147 | 708.2172 | 274.2241 | 305.4198 | 393.6784 | 316.6086 | 317.6362 | 362.7483 | 468.8112 | 195.0069 | 572.3845 | 729.7974 | 964.4896 | 812.7974 | 287.056 | 378.5733 | 444.1819 | 284.7552 | 145.9371 | 156.1819 | 179.9931 | 325.9862 | 425.8741 | 542.6155 | 667.2801 | 369.0629 | 342.9164 | 386.294 | 413.3707 | 275.9302 | 346.3009 | 449.5595 | 566.0422 | 398.5035 | 271.0422 | 333.6715 | 344.3009 | 219.5173 | 435.3147 | 465.056 | 606.1681 | 485.65 | 315.0629 | 398.0767 | 719.0138 | 525.4267 | 416.9793 | 498.1121 | 629.0698 | 616 | 364.4336 | 416.2586 | 579.4336 | 633.2448 | 280.1681 | 263.9509 | 327.1604 | 361.7974 | 315.5526 | 325.3845 | 435.4129 | 298.1612 | 244.4965 | 330 | 423.6784 | 341.6922 | 598.5035 | 790.2448 | 1060.0207 | 873.3078 | 582.5311 | 968.5948 | 1293.7483 | 1305.7207 | 225.1328 | 126.7483 | 125.119 | 225.1328 | 191.7345 | 208.7483 | 219.056 | 229.9509 | 181.175 | 188.7552 | 228.9371 | 190.6224 | 127.8112 | 135.0629 | 153.3147 | 178.175 | 166.4267 | 166.7414 | 184.0698 | 129.6224 | 482.9931 | 476.2026 | 621.8672 | 409.5457 | 332.9931 | 463.294 | 580.0983 | 432.3707 | 616.2793 | 759.2241 | 958.8672 | 703.3147
658 | 9.2324 | 405.4645 | 573.031 | 555.1815 | 509.1415 | 317.5266 | 417.4203 | 449.7123 | 369.8938 | 246.8717 | 204.0531 | 239.7832 | 303.354 | 406.9911 | 414.2168 | 388.3098 | 495.4203 | 357 | 307.1726 | 333.385 | 378.8053 | 310.7611 | 642.6107 | 772.8185 | 433.6018 | 455.677 | 538.7964 | 670.3761 | 639.7832 | 220.4071 | 269.9337 | 372.4956 | 283.031 | 340.7964 | 361.1815 | 467.6549 | 210 | 590.2168 | 699.1815 | 941.0973 | 777.1415 | 275.1062 | 401.0221 | 434.4071 | 294.9469 | 146.4292 | 144.4513 | 164.2478 | 296.7522 | 426.4071 | 516.6592 | 594.4424 | 369.1726 | 297.7522 | 349.6107 | 402.4956 | 261.354 | 324.0841 | 443.9027 | 524.708 | 403.3672 | 229.4071 | 300.6549 | 318.0973 | 255.0442 | 437.6681 | 450.4734 | 564.677 | 535.8628 | 316 | 426.5929 | 741.7034 | 503.1283 | 374.9337 | 469.3982 | 642.292 | 614.4956 | 355.9248 | 435.4513 | 568.7433 | 618.6902 | 238.3761 | 295.1062 | 398.4335 | 327.7964 | 296.9337 | 344.7212 | 382.2257 | 245.5133 | 237.3009 | 327.969 | 401.3672 | 350.4956 | 602.0884 | 775.9159 | 1100.4114 | 866.0841 | 646.3893 | 903.1504 | 1288.1194 | 1240.8717 | 237.7123 | 122.6991 | 114.1062 | 237.7123 | 159.8053 | 202.969 | 203.8717 | 257.4956 | 156.4071 | 200.3009 | 227.0973 | 187.1062 | 125 | 134.5708 | 158.4513 | 152.354 | 144.1283 | 149.5044 | 198.4203 | 125.4292 | 468.0752 | 507.7212 | 601.8407 | 375.4823 | 319.0531 | 427.4381 | 522.3761 | 423 | 675.3362 | 703 | 940.0442 | 706.8717
659 | 9.2434 | 367.5133 | 593.3618 | 582.8751 | 469.5368 | 324 | 398.175 | 397.7876 | 349.9509 | 233.7876 | 217.2625 | 275.5752 | 317.6371 | 369.2124 | 382.0128 | 427.7748 | 476.175 | 355.6873 | 298.4877 | 373.6873 | 398.7246 | 358.2252 | 684.0245 | 747.3127 | 436.0117 | 465.5752 | 525.2252 | 672.3746 | 672.2497 | 249.5624 | 304.3372 | 355.0502 | 306.7748 | 326.2625 | 386.5123 | 489.8378 | 210.3501 | 557.7503 | 728.1878 | 927.5496 | 741.4749 | 300.2999 | 388.2999 | 462.2497 | 278.587 | 164 | 149.2999 | 157.4504 | 309.9381 | 453.3746 | 451.3501 | 562.2999 | 361.8004 | 306.6499 | 391.2871 | 388.3757 | 277.7375 | 361.1249 | 450.7994 | 555.8122 | 372.7503 | 256.8997 | 316.2743 | 308.4877 | 230.0374 | 411.1878 | 442.5998 | 571.7748 | 486.7503 | 316.4376 | 393.937 | 652.7001 | 519.9755 | 408.1121 | 462.9626 | 607.8623 | 598.5379 | 354.825 | 436.9744 | 540.6499 | 578.175 | 241.3372 | 319.3372 | 330.7876 | 314.6628 | 332.3874 | 328.7375 | 386.8378 | 298.1121 | 238.8122 | 303.2625 | 375.8261 | 330.7748 | 557.9381 | 740.1878 | 1049.3383 | 903.2124 | 604.4504 | 907.0128 | 1264.825 | 1226.4749 | 187.8004 | 118.825 | 139.0374 | 187.8004 | 182.5251 | 177.7375 | 191.9253 | 237.9498 | 184.8623 | 203.3874 | 216.9626 | 213.8751 | 125.175 | 118.1377 | 164.3501 | 170.9253 | 155.1121 | 168.2625 | 178.2124 | 140.8997 | 470.4002 | 492.5251 | 565.9755 | 407.7503 | 332.2625 | 474.1238 | 525.4248 | 423 | 619.0502 | 700.1121 | 909.524 | 692.825
660 | 9.2544 | 421.8063 | 546.1977 | 589.1057 | 428.5891 | 322.0039 | 400.5988 | 407.3973 | 408.5068 | 240.7025 | 218.1037 | 283.317 | 295.2955 | 362.8963 | 397.8102 | 416.8963 | 477.2016 | 342.6986 | 310.407 | 363.0939 | 376.6967 | 365.6106 | 650.9041 | 755.0117 | 382.1037 | 476.1115 | 536.4031 | 645.998 | 651.6047 | 242.9041 | 289.5949 | 363.2055 | 296.8943 | 326.0059 | 370.9002 | 504.6086 | 212.1037 | 572.1076 | 750.6008 | 899.9041 | 743.1859 | 291.9022 | 375.91 | 441.9041 | 247.3914 | 163.6008 | 141.9002 | 179.7065 | 348.1057 | 423.5049 | 453.3033 | 558.593 | 385.5108 | 299.2074 | 364.5969 | 431.9041 | 276.3973 | 354.5949 | 411.1918 | 530.4051 | 391.998 | 237.1957 | 267.499 | 324.3992 | 225.4892 | 430.7065 | 429.9961 | 559.5988 | 505.3992 | 319.1037 | 363.2975 | 656.407 | 544.5127 | 375.6047 | 473.8982 | 592.8063 | 625 | 348.0098 | 395.2994 | 536.4012 | 581.0978 | 227.2935 | 301.501 | 337.4031 | 335.9922 | 323.9041 | 324.2035 | 401.2094 | 265.7045 | 217.3973 | 307.4971 | 437.5186 | 320.0959 | 594.409 | 762.9002 | 1096.9158 | 892.6047 | 624.6106 | 919.317 | 1256.9119 | 1213.2152 | 214.8043 | 118.0978 | 130.9961 | 214.8043 | 188.1996 | 176.3973 | 211.3014 | 228.8963 | 172.9002 | 197.5988 | 227.0998 | 220.4051 | 124.6047 | 130.6986 | 165.9041 | 186.3092 | 126.5969 | 168.0059 | 172.0959 | 118.3014 | 477.9139 | 494.8063 | 596.3014 | 422.2074 | 333.1037 | 393.7965 | 512.1937 | 420.7045 | 630.2994 | 668.5029 | 845.4932 | 692.1977
661 | 9.2654 | 393.7847 | 551.5073 | 561.7669 | 473.8436 | 310.8378 | 403.7581 | 417.9823 | 379.0177 | 230.0089 | 201.4484 | 212.1741 | 316.2242 | 383.475 | 348.8702 | 436.8024 | 467.1976 | 348.1032 | 283.7169 | 392.2331 | 392.7847 | 316.2153 | 629.5339 | 700.5516 | 369.4838 | 423.8968 | 515.6106 | 656.2331 | 630.4484 | 223.3274 | 305.7138 | 340.0177 | 322.1032 | 301.6903 | 372.4661 | 463.2065 | 193.3186 | 540.2508 | 738.1444 | 870.2388 | 799.5928 | 283.4572 | 331.0089 | 423 | 282.292 | 160.7847 | 139.5428 | 149.5516 | 319.9823 | 400.3274 | 439.4572 | 587.525 | 340.3717 | 269.1475 | 379 | 408.0678 | 289.7847 | 374.8613 | 445.9734 | 507.3363 | 397.1887 | 255.879 | 274.4661 | 329.0089 | 271.7492 | 403.0177 | 444.1887 | 558.3864 | 510.2331 | 303.121 | 377.354 | 627.8113 | 483.4218 | 356.354 | 483.2331 | 565.0089 | 621.3008 | 304.0089 | 398.6726 | 529.5428 | 590.2153 | 260.475 | 295.879 | 324.0089 | 370.2065 | 304.6637 | 311.3629 | 358.2242 | 246.6814 | 232.1298 | 319.3186 | 353.7935 | 335.525 | 553.4484 | 757.4041 | 1024.4484 | 871.1121 | 580.3717 | 844.475 | 1199.1976 | 1143.2242 | 194.6637 | 130.6903 | 146.5339 | 194.6637 | 187.0855 | 192.9233 | 203.7669 | 243.5339 | 170.9911 | 203.6726 | 226.4307 | 198.9056 | 104.3451 | 137.4484 | 145.9911 | 146.354 | 141.4484 | 141.3363 | 187.525 | 114.0177 | 418.8201 | 461.7404 | 580.9203 | 389.7935 | 317.5693 | 407.0855 | 539.0944 | 396.8613 | 631.2065 | 653.7669 | 877.6726 | 698.7404
662 | 9.2764 | 397.3879 | 513.2242 | 549.3879 | 436.9804 | 358.0343 | 382.144 | 401.8756 | 394.0049 | 238.3732 | 206.617 | 267.383 | 315.6367 | 410.8756 | 336.8511 | 450.9149 | 451.9657 | 339.1342 | 323.6514 | 396.2782 | 401.4877 | 311.4926 | 609.7512 | 701.4779 | 404.2537 | 431.8707 | 468.7316 | 661.2733 | 634.7463 | 223.473 | 251.4828 | 354.5074 | 313.8805 | 318.0245 | 385.6563 | 446.8854 | 185.473 | 557.7905 | 672.7267 | 788.4828 | 739.4681 | 293.1342 | 337.383 | 424.3683 | 246.617 | 165.2586 | 130.6121 | 146 | 306.6121 | 391.8903 | 456.0999 | 560.1244 | 371.7659 | 300.1342 | 378.3781 | 370.7414 | 287.9149 | 349.3781 | 419.8854 | 508.5074 | 367.7463 | 250.6072 | 294 | 337.1244 | 242.5172 | 419.2488 | 410.7659 | 512.9657 | 516.5172 | 312.1244 | 396.7463 | 647.9247 | 463.3732 | 379.6023 | 491.6317 | 569.6416 | 590.1342 | 314.239 | 404.3732 | 515.0147 | 576.7856 | 280.0393 | 285.1342 | 331.0049 | 358.1146 | 298.0196 | 332.1538 | 365.473 | 258.8805 | 246.8854 | 303.6367 | 364.5172 | 305.0147 | 558.8658 | 719.2586 | 1015.6857 | 867.8952 | 608.4075 | 862.671 | 1175.6317 | 1139.4026 | 192.6219 | 147.6465 | 126.8756 | 192.6219 | 165.1195 | 218.6514 | 193.6268 | 223.2537 | 165.8609 | 207.0098 | 214.9951 | 213.8756 | 115 | 140.5025 | 133.5221 | 163.1342 | 144.3781 | 145.617 | 158.7561 | 132.1146 | 456.7512 | 424.0196 | 515.9657 | 403.2537 | 328.2586 | 388.9755 | 520.2684 | 376.6023 | 610.0343 | 644.1244 | 873.671 | 664.8756
663 | 9.2874 | 379.8201 | 542.7561 | 532.5035 | 457.6194 | 308.3166 | 357.0502 | 402.9135 | 380.1661 | 236.763 | 219.2301 | 252.8564 | 296.8132 | 411.0934 | 364.5761 | 397.9637 | 498.6765 | 330.3737 | 283.4031 | 354.4965 | 414.763 | 318.173 | 608.5467 | 736.9637 | 397.2232 | 434.9931 | 497.5969 | 632.9204 | 635.6263 | 253.7993 | 277.5035 | 347.9637 | 306.1367 | 293.1141 | 334.9862 | 430.2232 | 217.5761 | 504.1003 | 707.7768 | 814.7768 | 778.9862 | 279.7266 | 325.59 | 424.8426 | 255.4031 | 152.59 | 147.9066 | 148.1868 | 319.5329 | 376.9706 | 483.526 | 564.9637 | 350.7768 | 283.4464 | 369.0796 | 381.6332 | 237.5605 | 345.8201 | 404.7266 | 499.0934 | 375.1868 | 269.5761 | 293.4533 | 334.173 | 220.7336 | 421.6834 | 390.5969 | 564.1868 | 488.4464 | 315.7336 | 399.4031 | 586.5674 | 464.4965 | 408.0363 | 480.0865 | 546.9567 | 574.4031 | 329.1799 | 406.8633 | 497.4169 | 529.1436 | 227.4169 | 273.2301 | 322.3599 | 367.5329 | 274.1003 | 297.0138 | 399.4896 | 253.1868 | 232.5467 | 289.5969 | 345.3304 | 284.41 | 565.7128 | 711.7837 | 939.9342 | 843.5104 | 560.9066 | 793.7197 | 1151.3755 | 1098.9066 | 197.41 | 119.7336 | 125.4533 | 197.41 | 171.4965 | 185.1003 | 191.41 | 217.1799 | 182.7197 | 192.8633 | 218.3097 | 210.2664 | 114.59 | 135.7699 | 105.2803 | 150 | 138.9066 | 156.59 | 152.1868 | 147.1367 | 458.1436 | 405.4308 | 565.41 | 396.4965 | 318.0502 | 421.2664 | 490.9931 | 408.59 | 556.8997 | 639.8062 | 803.0796 | 666.4602
664 | 9.2984 | 381.4247 | 499.0856 | 537.6799 | 412.4897 | 300.2758 | 367.6165 | 412.6165 | 340.5753 | 215.9588 | 223.7227 | 231.5103 | 287.767 | 407.2742 | 329.8098 | 414.6577 | 481.3186 | 356.8082 | 272.3629 | 342.7655 | 388.2979 | 298.1711 | 601.2742 | 754.8495 | 386.149 | 415.851 | 513.4041 | 655.9557 | 609.6387 | 200.2124 | 286.851 | 363.615 | 303.7227 | 350.1474 | 305.1933 | 429.7258 | 177.2979 | 532.2345 | 702.9114 | 822.7227 | 729.9763 | 279.6387 | 322.7021 | 378.1727 | 237.9572 | 148.5103 | 140.851 | 159.3186 | 290.5118 | 411.7639 | 432.5325 | 581.2108 | 355.5943 | 266.6608 | 341.4263 | 372.4041 | 275.5737 | 346.9778 | 398.382 | 502.2758 | 384.084 | 228.851 | 290.149 | 308.5103 | 238.7448 | 418.9985 | 406.2552 | 567.1253 | 466.7448 | 323.9129 | 383.5959 | 642.8289 | 455.7448 | 378.5309 | 462.7861 | 554.0428 | 560.2345 | 325.1918 | 399.2964 | 511.7005 | 547.9572 | 247.6593 | 278.3186 | 315.5753 | 336.1283 | 300 | 331.7227 | 368.3186 | 267 | 234.9572 | 302.7227 | 375.8717 | 288.3407 | 527.0428 | 733.0619 | 998.8701 | 870.9572 | 554.2979 | 776.1283 | 1059.149 | 1099.4484 | 198.5103 | 135.6593 | 122 | 198.5103 | 163.7876 | 207.2758 | 197.5753 | 217.0649 | 162.8938 | 192.8938 | 197.6814 | 192.7021 | 110.9572 | 124.7227 | 127.6387 | 154.916 | 134.0856 | 151.4675 | 165.1062 | 142.0412 | 469.3598 | 464.4041 | 560.9985 | 390.7242 | 332.7876 | 407.8732 | 477.3629 | 406 | 534.8732 | 610.4263 | 776.916 | 675.5943
665 | 9.3094 | 366.6441 | 502.937 | 509.4523 | 423.4189 | 284.9666 | 356.7414 | 396.5811 | 360.355 | 245.9685 | 207.0964 | 224.2901 | 313.5487 | 349.709 | 356.2271 | 377.9657 | 467.4513 | 348.2261 | 304.1288 | 331.2901 | 394.6765 | 322.3225 | 541.9351 | 709.9685 | 385.0649 | 419.6765 | 505.4847 | 610 | 610.7757 | 197.418 | 283.2586 | 308.7099 | 286.4513 | 306.0009 | 351.355 | 475.3578 | 181.7414 | 543.7099 | 609.7099 | 806.4189 | 668.4513 | 285.4523 | 321.645 | 441.1937 | 237.4828 | 142.0964 | 143.0639 | 148.0315 | 319.4198 | 364.4828 | 445.8712 | 536.3216 | 318.5162 | 295.4523 | 357.9361 | 372.3865 | 258.2576 | 329.6441 | 375.709 | 480.355 | 355.1937 | 233.8054 | 283.7433 | 301.1288 | 241.0649 | 378.6775 | 397.9685 | 500.8072 | 463.582 | 273.6126 | 386.0324 | 620.2901 | 459.5162 | 358.1603 | 442.6756 | 553.8721 | 573.3225 | 333.2901 | 360.8387 | 476.7729 | 542.645 | 240 | 266.0639 | 339.3865 | 353.6775 | 298.5487 | 306.7109 | 356.8702 | 261.8397 | 229.3225 | 280.7748 | 355.5811 | 298.5802 | 529.6135 | 684.0964 | 945.354 | 874.8368 | 561.4828 | 794 | 1055.1622 | 1128.2929 | 194.5153 | 126.3874 | 120.2261 | 194.5153 | 175.1288 | 187.9351 | 212.1946 | 242.4523 | 170.4189 | 199.2901 | 210.4198 | 191.9676 | 106.6126 | 109.709 | 138.7739 | 179.4523 | 143.2586 | 139.2576 | 161.6126 | 111.3865 | 427.7405 | 459.5487 | 518.4198 | 404.582 | 336.3883 | 417.9055 | 500.2595 | 409.8702 | 555.0649 | 628.7099 | 795.1632 | 637.0649
666 | 9.3204 | 390.7935 | 472.1073 | 497.5949 | 440.6859 | 319.471 | 371.1736 | 378.967 | 362.3801 | 221.0083 | 217.8264 | 232.0413 | 301.8348 | 372.471 | 323.0826 | 431.529 | 471.1819 | 335.7025 | 309.3058 | 338.5207 | 409.1076 | 318.6199 | 554.0826 | 689.1736 | 377.4297 | 439.3141 | 482.3471 | 600.1076 | 578.7602 | 237.2149 | 266.9174 | 312.9337 | 290.529 | 281.0413 | 343.3141 | 396.0826 | 196 | 532.8348 | 614.8015 | 819.1736 | 681.9007 | 268.9091 | 325 | 435.2645 | 268.6033 | 152.1322 | 154.1406 | 167.7851 | 304.0909 | 389.529 | 438.2645 | 558.7935 | 313.9174 | 276.6529 | 344.7768 | 386.8761 | 262.5373 | 356.2232 | 393.0909 | 478.2149 | 349.7851 | 256.5373 | 248.6033 | 306.6529 | 231.1736 | 377.5207 | 383.2478 | 484.3801 | 431.0743 | 290.1569 | 372.3225 | 622.3141 | 459.4297 | 386.562 | 478.7688 | 521.1322 | 565.8018 | 336.8761 | 360.3471 | 505.7522 | 550.1652 | 232.3638 | 272.9754 | 355.0909 | 346.967 | 288.9174 | 276.9504 | 367.2812 | 239.1652 | 230.8264 | 268.5123 | 342.1322 | 312.6942 | 518.2478 | 695 | 966.438 | 909.3391 | 580.1076 | 785.8431 | 1028.7851 | 1066.8924 | 216.0909 | 115.7438 | 111.5207 | 216.0909 | 181.6942 | 194.3554 | 186.6859 | 224.8678 | 181.0909 | 203.0496 | 200.1239 | 193.8761 | 115.3471 | 127.6116 | 144.7025 | 157.7025 | 132.1239 | 148.0496 | 169.1322 | 124.6612 | 456.5456 | 446.4464 | 508.1239 | 374.504 | 300.2645 | 364.5453 | 467.4627 | 427.7438 | 541.1819 | 614.8844 | 736.5373 | 642.2725
667 | 9.3314 | 361.9115 | 519.2389 | 523.9823 | 459.6548 | 328.938 | 367.354 | 358.1593 | 374.0619 | 204.4159 | 214.2124 | 232.7256 | 281.0708 | 386.3982 | 331.6991 | 413.6991 | 451.3274 | 321.1151 | 295.2212 | 341.9292 | 364.4159 | 297.8319 | 569.5753 | 687.0265 | 380.2654 | 413.0973 | 480.0973 | 559.5045 | 633.1593 | 238.9115 | 262.9469 | 363.6991 | 276.0443 | 291.7611 | 326.2036 | 402.8407 | 197.3009 | 511.1416 | 663.1946 | 816.469 | 710.1858 | 271.6814 | 327.6018 | 423.8407 | 274.6814 | 146.5133 | 131.9292 | 158.2832 | 287.7522 | 375.7876 | 430.7434 | 530.469 | 309.5753 | 274.6283 | 352.7964 | 368.8407 | 239.6373 | 340.9557 | 379.5398 | 486.0443 | 347.7168 | 223.6018 | 260.8142 | 303.8849 | 229.7699 | 385.7611 | 407.0531 | 510.7434 | 456.469 | 313.3893 | 331.5575 | 605.0177 | 471.1858 | 385.6371 | 412.3009 | 515.5133 | 523.6726 | 324.0443 | 365.7168 | 476.7434 | 571.1151 | 201.7611 | 240.4071 | 338.1946 | 320.3982 | 271.3805 | 285.7964 | 327.1328 | 262.1593 | 226.6548 | 286.6548 | 329.823 | 314.8849 | 536.292 | 692.2124 | 949.3362 | 829.0708 | 537.2744 | 750.7168 | 1015.3805 | 1112.5133 | 205.6991 | 105.3717 | 115.3009 | 205.6991 | 181.6548 | 178.4867 | 207.3274 | 223.4247 | 173.8584 | 188.2566 | 213.6991 | 181.0443 | 115.5133 | 121.1858 | 132.9027 | 146.2744 | 147.9292 | 139.9469 | 161.469 | 110.3186 | 405.9027 | 418.5133 | 526.531 | 420.7611 | 291.0708 | 415.5753 | 492.7787 | 411.2389 | 519.0973 | 588.9912 | 704.3452 | 696.7875
668 | 9.3424 | 343.6037 | 495.1888 | 493.8765 | 443.6131 | 300.7739 | 356.9066 | 367.6411 | 312.2075 | 223.3216 | 208.1421 | 211.6224 | 279.7739 | 374.6037 | 318.0747 | 412.5571 | 441.6224 | 324.2168 | 250.3776 | 337.2821 | 371.4336 | 316.028 | 589.2448 | 673.6037 | 341.0093 | 402.5384 | 468.5477 | 585.4243 | 638.8765 | 223.7739 | 283.4523 | 349.8765 | 282.8205 | 309.0747 | 342.6597 | 389 | 201.0187 | 500.4243 | 624.6411 | 800.4056 | 703.4709 | 249.944 | 336.2261 | 408.2168 | 248.7832 | 144.9533 | 135.6037 | 145.1701 | 261.1515 | 381.6597 | 437.3683 | 514.0093 | 331.4243 | 273.5944 | 334.1608 | 354.2075 | 290.0467 | 333.9813 | 362.7739 | 495.5944 | 363.8019 | 230.4429 | 262.1795 | 301.1795 | 222.5757 | 405.056 | 377.8299 | 501.5176 | 438.8205 | 253.916 | 341.1328 | 616.499 | 465.2635 | 356.1328 | 422.5571 | 505.6317 | 533.1795 | 323.6878 | 376.056 | 474.2541 | 568.0747 | 229.3776 | 281.4429 | 305.1888 | 309 | 250.6805 | 277.0674 | 358.9533 | 265.3008 | 211.0093 | 271.8019 | 299.5477 | 310.5944 | 488.7645 | 672.6691 | 906.4056 | 822.2261 | 552.4336 | 759.4709 | 978.2355 | 1098.0747 | 203.3683 | 109.3776 | 120.0093 | 203.3683 | 165.2168 | 184.4056 | 195.6411 | 203.3589 | 182.3963 | 183.6131 | 205.8205 | 189.8019 | 114.1515 | 123.9813 | 145.6411 | 161.0373 | 150.6131 | 164.0187 | 149.3683 | 136.6131 | 419.0373 | 414.1795 | 535.4616 | 432.3288 | 295.916 | 432.6691 | 447.8019 | 392.9347 | 509.3309 | 616.8672 | 721.5757 | 618.7832
669 | 9.3534 | 336.9498 | 496.4228 | 464.1564 | 430.0039 | 306.1081 | 385.0058 | 358.8378 | 315.5231 | 251.1081 | 222.8996 | 210.417 | 289.7375 | 369.2124 | 298.415 | 413.7452 | 437.6814 | 312.1602 | 265.471 | 299.4749 | 365.417 | 298.3186 | 570.2104 | 679.3648 | 347.942 | 432.3668 | 479.8475 | 575 | 598.8456 | 230.1602 | 255.1081 | 317.2104 | 275.0521 | 280.7876 | 320.3668 | 395.944 | 191.3166 | 492.1042 | 597.3205 | 797.5792 | 669.0019 | 263.4788 | 319.1081 | 395.3186 | 250.7394 | 162.8958 | 135.8938 | 155.8958 | 290.2065 | 358.3147 | 446.7915 | 507.8958 | 320.7896 | 273.4749 | 354.7876 | 345.7394 | 262.2143 | 340.7375 | 377.1564 | 488.7413 | 363.8417 | 214.3147 | 269.1062 | 309.1583 | 233.6313 | 373.8456 | 373.417 | 461.5212 | 441.9942 | 291.2625 | 369.1583 | 578.4189 | 430.4247 | 380.5811 | 438.6853 | 488.8436 | 542.4208 | 302.6235 | 353.6834 | 453.1042 | 534.9479 | 246.5753 | 260.2645 | 308.9479 | 316.7857 | 294.8456 | 327.8977 | 369.1101 | 222.2104 | 209.7355 | 275.6293 | 325.7876 | 313.6313 | 503.1583 | 645.1042 | 883.5889 | 813.1042 | 546.8378 | 730.473 | 970.7857 | 1078.6254 | 215.1062 | 114.3707 | 114.9479 | 215.1062 | 156.1062 | 178.8436 | 182.2085 | 220.1583 | 182.3166 | 175.2645 | 195.3166 | 189.4208 | 130.7915 | 128.2124 | 129.0521 | 149.6293 | 143.1062 | 155.7896 | 160.6853 | 129.9479 | 409.7336 | 422.1583 | 506.0541 | 372.7915 | 330.527 | 397.9498 | 454.5753 | 410.4788 | 539.2104 | 581.4769 | 732.6313 | 620.1081
670 | 9.3644 | 314.4366 | 493.913 | 478.6726 | 412.0266 | 284 | 344.4543 | 382.8687 | 337.1047 | 231.4499 | 188.7862 | 239.9956 | 283.8864 | 351.6593 | 333.8642 | 350.559 | 458.7773 | 300.4587 | 276.646 | 295.7906 | 385.1961 | 287.913 | 565.8776 | 695.8642 | 379.2005 | 438.3185 | 440.6726 | 568.764 | 576.677 | 219.5723 | 236.118 | 313.323 | 276.9956 | 293.2183 | 335.8953 | 413.5368 | 185.4454 | 497.5501 | 568.5723 | 782.4144 | 666.2493 | 228.8997 | 302.1224 | 377.5634 | 240.7995 | 155.2227 | 138.9911 | 148.2227 | 310.5324 | 375.6859 | 424.5413 | 503.7862 | 321.1136 | 270.6815 | 355.4144 | 329.118 | 233.6681 | 329.764 | 378.3097 | 458.2316 | 363.882 | 223.8909 | 268.0133 | 296.8687 | 229.9867 | 345.8864 | 402.5501 | 490.7551 | 472.7462 | 288.4277 | 358.205 | 589.6548 | 411.5767 | 362.441 | 439.5767 | 485.0177 | 540.6593 | 348.2964 | 365.1224 | 450.9778 | 529.8864 | 277.0044 | 249.3319 | 305.8909 | 341.323 | 270.4499 | 303.2183 | 329.1047 | 227.0089 | 223.7862 | 287.6548 | 335.1003 | 310.2094 | 511.5812 | 648.323 | 817.913 | 821 | 568.6415 | 723.0825 | 975.7241 | 1106.6238 | 208.4454 | 96.1091 | 114.3407 | 208.4454 | 152.118 | 167.4454 | 192.6548 | 216.1091 | 179.5634 | 169.0089 | 186.9956 | 196.3452 | 115.6681 | 115.3363 | 137.677 | 169.4499 | 134.8864 | 152.7728 | 152 | 127.7817 | 434.3319 | 425.0133 | 494.1091 | 370.3629 | 314.5324 | 383.4543 | 479.882 | 377.9042 | 542.2272 | 548.9823 | 734.3319 | 604.236
671 | 9.3754 | 310.2252 | 515.5496 | 492.2301 | 439.1946 | 284 | 353.4198 | 353.5851 | 335.6952 | 240 | 203.88 | 235.5851 | 278.825 | 348.2252 | 303.4602 | 357.88 | 457.06 | 317.8299 | 254.3403 | 309.6548 | 342.355 | 320.5949 | 546.355 | 656.765 | 348.0551 | 412.6499 | 455.4051 | 550.5251 | 591.355 | 236.1297 | 243.7099 | 289.2448 | 266.7099 | 283.2399 | 339.6548 | 388.2301 | 187.705 | 495.3452 | 599.4651 | 722.3097 | 690.9045 | 239.1848 | 321.41 | 383.6548 | 271.2399 | 152.7099 | 129.2301 | 148.06 | 273.53 | 393.0747 | 395.6548 | 511.5949 | 332.9902 | 287.4198 | 308 | 344.9351 | 236 | 311.5251 | 348.2852 | 468.0649 | 347.3599 | 226.765 | 281.7148 | 266.41 | 209.94 | 342.235 | 393.06 | 453.59 | 437.5704 | 265.7601 | 333.5851 | 564.4198 | 443.355 | 353.7099 | 463.8348 | 503.8348 | 535.8152 | 297.7552 | 383 | 422.5802 | 520.1248 | 273.1897 | 247.235 | 316.2852 | 311.3696 | 274.7699 | 296.53 | 317.825 | 239.825 | 231.8299 | 271.5851 | 323.6401 | 288.355 | 541.4247 | 624.2448 | 852.2399 | 816.7699 | 534.1701 | 674.2399 | 886.2999 | 1031.7148 | 205.7699 | 95.3501 | 130.7001 | 205.7699 | 162.295 | 175.1102 | 171.1799 | 210.4149 | 186.8299 | 177.3599 | 180.705 | 209.5949 | 117.295 | 117.1799 | 149.5349 | 168.1297 | 129.355 | 143.59 | 156.4651 | 133.53 | 430.355 | 434.9547 | 490.295 | 415.59 | 285.0502 | 395.7099 | 468.765 | 397.7699 | 528.8494 | 526.1701 | 719.3097 | 626
672 | 9.3864 | 332.7959 | 463.2365 | 484.9797 | 388.0324 | 285.9784 | 338.1824 | 364.473 | 357.8176 | 233.8176 | 212.7203 | 233.6135 | 268.323 | 372.2797 | 324.8176 | 362.2689 | 438.6676 | 294.8176 | 288.9892 | 296.6662 | 337 | 297.0649 | 532.3446 | 645.8608 | 364.677 | 403.2689 | 465.8284 | 558.0432 | 580.0649 | 208.9459 | 235.2257 | 269.1932 | 258.9676 | 278.6662 | 315.5378 | 397.7959 | 200.1392 | 510.5595 | 618.6986 | 682.9784 | 634.9676 | 228.1284 | 323.527 | 366.9568 | 260.4838 | 137.5487 | 145.2257 | 140.5487 | 279.9135 | 355.6986 | 383.6554 | 492.5162 | 344.1622 | 259.5703 | 311.9568 | 359.5162 | 239.4622 | 316.5703 | 370.2689 | 452.4946 | 328.0108 | 223.7743 | 260.7203 | 274.9568 | 224.3757 | 337.0649 | 388.5162 | 453.4513 | 476.8392 | 280.7203 | 339.7743 | 546.7095 | 438.7419 | 338.3014 | 435.7311 | 473.2581 | 557.2581 | 319.7851 | 373.1081 | 441.527 | 503.9676 | 239.7311 | 258.3865 | 322.9365 | 274.6662 | 260.5054 | 299.4513 | 322.9027 | 228.8284 | 227.3649 | 279.0108 | 336.0649 | 281.0216 | 505.5162 | 590.5919 | 836.0432 | 787.9149 | 547.2905 | 674.3649 | 901.9121 | 1013.6878 | 187.0541 | 107.6986 | 137.0973 | 187.0541 | 152.0865 | 192.2797 | 164.7203 | 201.5054 | 167.7743 | 165.677 | 183.7419 | 189.2797 | 111.7851 | 112.6986 | 128.0757 | 141.1932 | 129.1932 | 145.677 | 159.8716 | 126.3122 | 418.0757 | 402.9568 | 491.2149 | 408.7743 | 306.5162 | 385.9892 | 459.0973 | 383.5054 | 482.0094 | 546.7095 | 683.4405 | 619.5703
673 | 9.3974 | 326.4897 | 470.1151 | 421 | 385.7493 | 290.9616 | 346.8761 | 366.5486 | 334.0678 | 218.1151 | 216.3274 | 209.0089 | 254.4897 | 360.6254 | 307.8171 | 366.0856 | 404.2006 | 280.6726 | 294.8555 | 317.4041 | 341.9322 | 286.5281 | 524.1712 | 605.1356 | 388.1151 | 406.3068 | 438.0295 | 536.3952 | 566.6726 | 225.0384 | 247.1918 | 275.3952 | 261.5103 | 291.0973 | 303.7021 | 382.6637 | 223.7315 | 486.7315 | 620.2802 | 690.0384 | 635.4336 | 264.3363 | 315.4808 | 371.2124 | 265 | 132.0472 | 146.5486 | 130.6342 | 294.1356 | 361.9528 | 399.8761 | 478.9144 | 303.413 | 238.9616 | 320.3658 | 341.761 | 254.9322 | 300.3746 | 359.0295 | 454 | 342.1328 | 218.0384 | 266.6637 | 285.7021 | 249.3952 | 322.8938 | 377.5103 | 468.8171 | 445.7699 | 286.1445 | 336.1151 | 548.1356 | 439.1829 | 331.2685 | 430.6165 | 483.2006 | 546.9528 | 311.2979 | 354.4219 | 432.7021 | 508.8466 | 235.1356 | 275.7226 | 272.0384 | 293.8466 | 254.587 | 309.6254 | 342.2124 | 207.7788 | 253.3185 | 261.9822 | 322.413 | 279.1534 | 505.413 | 589.4013 | 819.8466 | 758.3541 | 525.0973 | 708.6254 | 948.2358 | 1025.2891 | 180.0472 | 116.2891 | 109.7404 | 180.0472 | 133.9705 | 179.8466 | 176.3746 | 196.1062 | 159.4425 | 182.6254 | 182.3658 | 185.1534 | 103.2979 | 125.4425 | 104.6637 | 157.5192 | 147.8555 | 161.587 | 135.3452 | 112.3746 | 398.0384 | 400.4631 | 498.9233 | 399.9233 | 297.0678 | 385.1062 | 437.1918 | 377.3274 | 551.0767 | 558.7788 | 691.6932 | 601.0384
674 | 9.4084 | 334.3967 | 465.406 | 421.8156 | 402.2812 | 276.8529 | 322.3501 | 343.1155 | 316.4656 | 222.1062 | 202.7281 | 231.4469 | 267.0187 | 351.8156 | 304.8436 | 372.9684 | 448.5158 | 299.9813 | 291.3092 | 307.5531 | 357.6313 | 297.6686 | 569.203 | 579.9907 | 384.7654 | 412.8716 | 430.5158 | 515.797 | 577.2812 | 229.9032 | 258.1937 | 252.6219 | 244.9032 | 262.7188 | 301.9032 | 373.5531 | 200.5531 | 454.3092 | 571.4936 | 684.2999 | 613.7188 | 263.6593 | 305.5717 | 347.3687 | 259.0187 | 148.203 | 130.7281 | 139.0968 | 267.9032 | 340.9032 | 377.5251 | 457.8938 | 308.6593 | 232.7374 | 302.1155 | 314.8752 | 254.3185 | 312.4469 | 347.9813 | 450.4656 | 374.1435 | 221.2719 | 261.9032 | 274.9311 | 221.4562 | 334 | 358.4562 | 466.3874 | 424.1937 | 275.5531 | 345.5438 | 514.0187 | 438.0782 | 359.8845 | 406.7095 | 501.4154 | 517.203 | 312.5531 | 381.8343 | 433.0782 | 504.3407 | 215.7001 | 246.9627 | 280.9813 | 289.3407 | 239.5531 | 294.5624 | 315.3781 | 212.7188 | 220.5438 | 234.7783 | 326.5717 | 288.2812 | 508.2123 | 651.3874 | 801.7468 | 801.1937 | 498.8938 | 693.2905 | 816.6219 | 1003.5624 | 205.175 | 104.8938 | 106.2812 | 205.175 | 149.9125 | 167.4562 | 180.5624 | 186.9032 | 157.4469 | 173 | 179.0689 | 192.6499 | 108.6313 | 122.9032 | 104.5251 | 159.5438 | 151.1062 | 154.1657 | 158.4656 | 124.9907 | 399.9125 | 370.6908 | 490.0093 | 404.0595 | 283.5438 | 374.5438 | 443.5717 | 366.175 | 540.9591 | 545.7748 | 650.3278 | 605.9032
675 | 9.4194 | 315.6618 | 432.9784 | 440.4819 | 396.1367 | 264.6186 | 336.8417 | 332.5143 | 315.2409 | 216.6726 | 198.59 | 231.4641 | 260.3776 | 347.7483 | 280.1583 | 336.2085 | 463.2015 | 313.1583 | 255.1691 | 297.8741 | 362 | 274.2626 | 549.8633 | 588.8525 | 351.1475 | 379.7877 | 452.59 | 530.0216 | 566.0216 | 235.2842 | 244.8741 | 270.3992 | 257.3884 | 272.5575 | 298.7591 | 364.4425 | 187.4641 | 449.4278 | 554.1977 | 665.8309 | 617.0216 | 244.2842 | 277.6008 | 335.3274 | 242.4317 | 132.8417 | 133.6942 | 139.6834 | 277.8309 | 343.7266 | 395.2842 | 472.4209 | 293.2626 | 235.9352 | 285.5467 | 338.6008 | 221.8201 | 312.1799 | 357.1799 | 444.9784 | 319.3382 | 216.0324 | 260.7483 | 258.5575 | 212.6116 | 333.7158 | 363.2517 | 444.0108 | 417.9784 | 270.705 | 343.0216 | 505.3884 | 446.7266 | 367.9282 | 418.0216 | 458.4209 | 494.7374 | 305.7158 | 368 | 456.5035 | 525.705 | 234.1691 | 278.1475 | 288.7591 | 306.7266 | 234.1367 | 276.0216 | 305.2734 | 220.2842 | 222.5683 | 261.0432 | 309.9676 | 293.5035 | 487.3992 | 616.7915 | 788.1259 | 797.5359 | 507.4533 | 666.0648 | 829.2842 | 991.8417 | 201.9065 | 112.6008 | 103.8309 | 201.9065 | 154.3884 | 167.4209 | 165.1475 | 194.8417 | 162.0108 | 173.5683 | 189.4749 | 179.295 | 116.1259 | 123.4533 | 116.6008 | 153.8957 | 128.3382 | 167.4317 | 147.0108 | 129.3058 | 397 | 403.4209 | 480.0108 | 416.2231 | 282.4425 | 381.1151 | 425.2626 | 379.1043 | 495.4425 | 515.7985 | 603 | 589.9714
676 | 9.4304 | 334.3672 | 441.8142 | 468.0663 | 416.4912 | 299.5132 | 326.2123 | 352.7743 | 324.7345 | 236 | 185.2566 | 215.8894 | 273.6239 | 333.7788 | 273.5929 | 370.5177 | 441.9292 | 297.1062 | 252.1903 | 280.8097 | 353.4026 | 287.6682 | 543.7389 | 588.0354 | 349.8894 | 402.2168 | 445.4823 | 512.8849 | 554.8142 | 231.2566 | 233.7389 | 279.8495 | 274.5177 | 282.2611 | 277.7035 | 360.0752 | 176.0398 | 500.2875 | 577.8052 | 681.854 | 607.8894 | 243.8142 | 283.27 | 322.5222 | 238.9248 | 138.8142 | 131.7345 | 153.7035 | 290.2965 | 348.3761 | 401.9292 | 463.8451 | 294.8097 | 259.7743 | 292.5531 | 319.9602 | 235.7389 | 291.7389 | 345.9292 | 449.3672 | 300.6328 | 205.7434 | 244.1106 | 274.7832 | 200.8186 | 338.3363 | 372.7743 | 440.7788 | 434.2257 | 270.2566 | 334.1858 | 508.584 | 445.1505 | 332.0354 | 403.2566 | 464.0752 | 497.9337 | 302.6283 | 360.292 | 453.314 | 522.5885 | 219.3318 | 270.6637 | 268 | 299.5177 | 238.1858 | 278.4514 | 315.4469 | 228.708 | 218.0708 | 250.7832 | 316.9955 | 322.6283 | 486.4734 | 575.6637 | 761.0176 | 810.7389 | 506.3761 | 644.5974 | 819.9203 | 982.6946 | 183.4469 | 98.1106 | 116.5929 | 183.4469 | 164.1062 | 167.146 | 166.854 | 209.708 | 167.3761 | 169.6637 | 166.5177 | 170.2566 | 113.9203 | 117.0398 | 113.6726 | 138.6682 | 115.8585 | 161.5531 | 147.3363 | 122.0752 | 399.9646 | 409.3717 | 480.9292 | 383.7434 | 282.2257 | 391.9248 | 426.8097 | 383.9557 | 503.5266 | 551.0045 | 615.7478 | 553.8983
677 | 9.4414 | 313.8987 | 437.765 | 438.0875 | 428.9773 | 272.6775 | 301.2901 | 342.6313 | 305.1013 | 225.1937 | 181.7188 | 215.8387 | 253.0138 | 346.3363 | 291.9814 | 352.4886 | 442.1061 | 280.3225 | 269.9862 | 269.4976 | 344.7326 | 288.1337 | 527.6637 | 571.1199 | 350.765 | 379.6175 | 442.5162 | 501.8663 | 551.6912 | 214.1337 | 230.3225 | 270.3088 | 262.0462 | 284.7512 | 266.1937 | 359.7512 | 183.9862 | 464.5438 | 537 | 676.4886 | 621.7326 | 243.1613 | 323.7512 | 350.0413 | 228.7512 | 135.6912 | 106.8525 | 153.3088 | 282.9724 | 367.7512 | 403.3411 | 437.1613 | 276.705 | 251.4838 | 290.9538 | 303.6775 | 229.8525 | 293.5713 | 345.4886 | 433.8387 | 320.7926 | 230.2764 | 249.1013 | 295.2212 | 220.3963 | 327.5348 | 365.1013 | 439.4425 | 443.1337 | 260.5438 | 327.7374 | 488.3639 | 457.7788 | 328.705 | 402.8063 | 465.6037 | 527.295 | 298.235 | 342 | 380.2626 | 504.0737 | 210.2212 | 255.2212 | 266.1475 | 292.295 | 251.8063 | 304.9125 | 325.6912 | 239.5899 | 205.5438 | 273.3825 | 304.7926 | 314.53 | 464.0413 | 570.4101 | 698.0689 | 781.3873 | 515.5624 | 662.53 | 782.1061 | 948.3225 | 191.2212 | 97.8525 | 121.3963 | 191.2212 | 136.825 | 153.6775 | 173.53 | 237.2626 | 185.5162 | 160.0875 | 167.94 | 167.9538 | 90 | 123.7512 | 129.7374 | 153.0276 | 137.825 | 145.7512 | 151.3549 | 118.9724 | 407.6175 | 405.1199 | 477.4011 | 400.2488 | 295.1475 | 374.3411 | 406.5438 | 350.2212 | 525.9587 | 556.4149 | 624.6961 | 598.94
678 | 9.4524 | 306.0261 | 426.0108 | 435.3687 | 368.1736 | 265.5683 | 331.0108 | 339.8372 | 318.7527 | 214.8049 | 194.1843 | 213.2473 | 257.8264 | 352.4793 | 328.0737 | 315.8264 | 408.7527 | 289.679 | 260.0369 | 288.8264 | 363.295 | 272.7527 | 504.8264 | 558.9523 | 356.0261 | 381.6421 | 424.1475 | 499.911 | 544.9002 | 198.7527 | 236.1475 | 268.4317 | 238.1475 | 271.0737 | 250.0261 | 345.7527 | 163.4424 | 474.7419 | 546.6313 | 635.6529 | 630.6637 | 245.7527 | 311.679 | 373.7896 | 230.1628 | 127.616 | 111.6421 | 149.8264 | 265.6421 | 352.7896 | 377.8894 | 450.0261 | 284.0369 | 266.6421 | 304.5791 | 309.089 | 233.0369 | 297.9847 | 326.4838 | 425.1475 | 337.9263 | 249.7266 | 252.1583 | 283.8633 | 226.3948 | 295.8678 | 362.7005 | 430.8525 | 422.616 | 264.321 | 321.1259 | 516.0108 | 462.7158 | 337 | 423.1475 | 456.0737 | 524.7788 | 292.6529 | 326.5899 | 404.2104 | 500.3948 | 210.4209 | 248.3579 | 264.5683 | 289.1367 | 276.6421 | 289.0108 | 326.2842 | 217.8633 | 211.5683 | 268.1475 | 299.4532 | 298.7266 | 476.8741 | 592.1367 | 728.232 | 734.9263 | 483.5683 | 635.4901 | 766.089 | 952.5423 | 192.063 | 103.9263 | 120.3318 | 192.063 | 132.3103 | 151.063 | 168 | 239.3795 | 171 | 165.7158 | 175.7896 | 181.2581 | 95.4577 | 110.3948 | 113.1736 | 161.3687 | 122.0737 | 137.8525 | 136.8157 | 104.2104 | 398.4056 | 394.5576 | 448.2734 | 406.5423 | 291 | 351.7788 | 416.4209 | 342.0737 | 496.1106 | 529.5315 | 583.4209 | 599.0845
679 | 9.4634 | 318 | 419 | 422.3334 | 380.6667 | 282.3334 | 313 | 319 | 308 | 233.6666 | 182.3333 | 213 | 257.6667 | 328 | 310.3333 | 308.3333 | 396 | 295 | 254 | 289.3334 | 348.3333 | 282.0001 | 511.6667 | 570.6666 | 361.3333 | 372.6666 | 410.6666 | 483.3334 | 536.6667 | 201 | 226.6667 | 266.3333 | 230.3333 | 267.3333 | 273 | 344 | 151 | 480.3333 | 557.3333 | 625.3334 | 597.6667 | 237.6667 | 298.3333 | 344.3332 | 248.3333 | 122.3334 | 120.3334 | 137.3333 | 264 | 341.3333 | 383 | 465.3333 | 288.6667 | 270 | 291.3333 | 338.3333 | 244 | 273.6667 | 347.3333 | 419 | 347.3334 | 224.6667 | 237 | 263.3333 | 221.6667 | 339 | 326 | 438.3333 | 412 | 269 | 340.3333 | 496.3333 | 457.6667 | 333 | 408.6666 | 454 | 527.6666 | 278.3333 | 306.3334 | 409.6667 | 505.3334 | 221 | 259 | 275.3334 | 290.3333 | 267.3333 | 273 | 323.6666 | 204.3333 | 217.3333 | 262 | 268.3334 | 283 | 439.6666 | 590 | 695.3333 | 742.3333 | 485 | 594 | 786.3333 | 936.6667 | 204.3333 | 110 | 92.6666 | 204.3333 | 147.3333 | 160.3333 | 174.6667 | 199 | 157.6666 | 160.3333 | 177.6667 | 164.3333 | 103.6667 | 113.3334 | 130.6667 | 142 | 119.6667 | 144.3333 | 145.6667 | 115.6667 | 389.6667 | 401.6666 | 450.9999 | 386.3333 | 285.6666 | 351 | 433.6667 | 349.6667 | 485.6667 | 531.6666 | 595.3333 | 580
680 | 9.4744 | 302.8156 | 430.5575 | 434.235 | 382.1244 | 297.0506 | 309.1244 | 333.6913 | 311.2075 | 213.3456 | 182.1106 | 211.295 | 270.1612 | 331.3087 | 283.6913 | 311.3456 | 385.1844 | 310.1106 | 247.3825 | 311.8019 | 348.9262 | 299.9769 | 538.5069 | 548.5437 | 339.0369 | 365.1338 | 408.295 | 486.0043 | 546.6913 | 217.0738 | 229.8756 | 267.1244 | 233.0738 | 269.0369 | 281.1981 | 351.0369 | 175.3456 | 469.9125 | 532.1244 | 632.0506 | 620.8988 | 241.3688 | 280.6038 | 312.1338 | 243.5163 | 137.8156 | 126.0137 | 126.5437 | 254.8894 | 335.4563 | 381 | 447.8293 | 297.8525 | 271.9262 | 276.7925 | 319.8662 | 259.2856 | 295.2719 | 332.6913 | 415.3962 | 334.1518 | 219.0369 | 254.4331 | 244.3456 | 221.0738 | 320.7556 | 339.3319 | 426.9769 | 419.9262 | 281.1475 | 315.5437 | 484.2856 | 440.2119 | 336.7513 | 400.8388 | 454.1981 | 486.1887 | 287.4194 | 326.1612 | 404.9262 | 515.4331 | 206.705 | 267.7925 | 283.235 | 276.1844 | 251.8756 | 273.1244 | 307.47 | 196.9262 | 219.0369 | 259.4331 | 290.5806 | 287.0506 | 433.6406 | 561.5806 | 687.8525 | 771.12 | 484.2581 | 598.8894 | 761.6175 | 940.7881 | 200.3825 | 112.2719 | 81.9494 | 200.3825 | 147.6775 | 155.9125 | 183.8525 | 199.765 | 135.8388 | 157.765 | 186.705 | 174.2075 | 102.8756 | 122.8894 | 138.3225 | 140.5437 | 126.1106 | 127.5206 | 131.5944 | 120.1612 | 406.9769 | 381.5163 | 415.765 | 378.3319 | 276.3825 | 329.4931 | 442.7281 | 363.6681 | 491.1707 | 510.9769 | 578.175 | 583.8525
681 | 9.4854 | 293.3579 | 421.3167 | 421.8417 | 374.9371 | 269.8308 | 301.2213 | 334.2842 | 328.5467 | 216.8633 | 190.8633 | 218.7679 | 265.6529 | 323.1996 | 278.1996 | 318.8004 | 394.2842 | 307.0521 | 238.5467 | 297.8829 | 353.9479 | 280.2842 | 512.4621 | 547.3688 | 344.2213 | 387.1367 | 430.4425 | 449.1063 | 541.9154 | 211.6942 | 233.1258 | 262.8004 | 236.6421 | 259.5467 | 271.1367 | 349.4208 | 176 | 481.2104 | 522.4317 | 609.4837 | 651 | 237.0196 | 286.41 | 335.5683 | 255.9675 | 131.5792 | 107.2104 | 124.295 | 255.4425 | 338.4208 | 390.3058 | 421 | 296.0846 | 256.1887 | 248.3579 | 292.8525 | 233.9892 | 295.9371 | 325.0521 | 410.6638 | 302.6008 | 227.0846 | 237.3471 | 255.7375 | 210.6833 | 307.5379 | 343.7571 | 421.6008 | 423.5163 | 280.7679 | 310.0738 | 475.0954 | 407.41 | 354.705 | 408.9262 | 446.2842 | 470.9175 | 303.0954 | 316.2842 | 399.5683 | 503 | 198.1583 | 242.2213 | 279.0738 | 282.7787 | 260.8525 | 273.0954 | 298.0738 | 203.0217 | 219.2104 | 241.6312 | 282.7787 | 280.59 | 452.5054 | 557 | 701.475 | 787.5879 | 475.8092 | 587.9892 | 753.9892 | 936.7246 | 196.1996 | 107.9262 | 105.59 | 196.1996 | 164.2104 | 169 | 168.8417 | 207.1475 | 155.3796 | 173.0217 | 174.5792 | 186.8937 | 108.6312 | 117.3579 | 117.1367 | 146.1692 | 134.5054 | 116.6746 | 131.1171 | 118.5163 | 372.9675 | 399.6942 | 418.8525 | 401.7267 | 276.1367 | 344.4425 | 434.1887 | 354.9479 | 496.8829 | 490.5683 | 558.4425 | 572.4208
682 | 9.4964 | 306.2169 | 433.3023 | 414.9233 | 372.0044 | 257.9956 | 315.1446 | 329.6681 | 312.4425 | 219.4085 | 194.149 | 205.851 | 263.6342 | 313.2212 | 270.0723 | 328.7788 | 407.0679 | 281.0044 | 218.7404 | 262.9617 | 348.8982 | 280.4808 | 487.6637 | 551.4469 | 342.5958 | 368.3023 | 419.7875 | 489.1062 | 538.0339 | 203.885 | 221.9617 | 257.6002 | 245.2552 | 254.1785 | 270.8127 | 342.2596 | 170.4469 | 463.5619 | 516.4852 | 604.5148 | 636.1917 | 206.2169 | 265.4469 | 318.9321 | 220.8554 | 138.3702 | 106.9617 | 138.6637 | 258.115 | 328.2979 | 384.4173 | 415.4469 | 299.9661 | 243.1446 | 247.149 | 278.6342 | 221.6681 | 295.9661 | 310.1106 | 413.5619 | 328.4808 | 228.7448 | 232.3319 | 276.1106 | 190.9233 | 337.8894 | 321.851 | 422.6769 | 409.9748 | 274.1446 | 302.3746 | 490.1873 | 401.9956 | 347 | 399.9661 | 462.77 | 516.1873 | 316.3363 | 327.2169 | 398.851 | 485.6002 | 211.3702 | 248 | 279.8894 | 284.7744 | 247.3746 | 275.9704 | 306.2935 | 219.8894 | 204.8938 | 232 | 281.8127 | 284.5236 | 445.1534 | 558.851 | 729.6342 | 726.4513 | 428.2256 | 581.9617 | 736.4852 | 905.2508 | 188.8127 | 109.7021 | 116.5575 | 188.8127 | 156.1873 | 164.5575 | 168.5871 | 204.3363 | 164.0044 | 172.8598 | 175.4469 | 178.3614 | 98.6298 | 116.5192 | 114.5915 | 153.1917 | 137.8894 | 141.149 | 161.2256 | 129.0383 | 354.885 | 373.3746 | 431.2891 | 404.5958 | 271.7404 | 356 | 390.7875 | 373.2212 | 473.8083 | 482.4469 | 565.1873 | 564.5192
683 | 9.5074 | 315.525 | 401.9675 | 419.2625 | 370.6824 | 261.6125 | 315.23 | 324.765 | 308.1425 | 217.295 | 187.5575 | 213.2075 | 241.7925 | 317.3825 | 281.2075 | 326.1475 | 419.1425 | 275.0925 | 238.7424 | 269.9125 | 318.7375 | 269.995 | 502.4425 | 553.8575 | 327.3225 | 354.5625 | 386.59 | 491.1701 | 533.5526 | 207.9075 | 211.7 | 232.59 | 255.2625 | 271.2025 | 283.59 | 349.415 | 168.65 | 433.2625 | 513.8575 | 592.6076 | 590.3451 | 240.3875 | 264.0325 | 309.9175 | 216.945 | 138.6175 | 118.885 | 137.76 | 254.475 | 324.6775 | 347.1475 | 409.4425 | 280.35 | 252.7925 | 243.6175 | 264.825 | 224.0325 | 285.53 | 313.1475 | 388.235 | 336.355 | 217.2075 | 234.5575 | 281.06 | 196.41 | 323.3775 | 321.175 | 397.9175 | 384.6549 | 278.4375 | 292.5025 | 482.3825 | 399.1101 | 336.29 | 373.0825 | 461.2526 | 509.1475 | 307.6775 | 345.705 | 397.0275 | 457.1475 | 216.9725 | 226.9626 | 261.9351 | 302.59 | 250.5074 | 268.0374 | 315.47 | 219.9125 | 190.9725 | 235.06 | 287.705 | 277.6225 | 425.765 | 540.1976 | 684.4601 | 732.5124 | 415 | 580.115 | 749.1574 | 928.4375 | 195.47 | 110.2625 | 109 | 195.47 | 142.645 | 156.6175 | 185.0275 | 194.1475 | 155.7975 | 145.825 | 172.5025 | 192.3451 | 103.355 | 113.235 | 114.7925 | 143.6824 | 126.82 | 142.9725 | 148 | 125.1475 | 376.12 | 361.9725 | 454.7925 | 390.0875 | 256.9351 | 353.3225 | 361.0325 | 362.465 | 489.055 | 472.235 | 560.4425 | 562.765
684 | 9.5184 | 289.1042 | 394.1323 | 401.7094 | 386.9198 | 246.7114 | 294.9739 | 337.4489 | 303.6072 | 224.9479 | 176.4729 | 205.7876 | 246.0541 | 314.8417 | 285.2365 | 326.4208 | 389.3407 | 284.3146 | 259.1303 | 276.1583 | 329.9739 | 269.1343 | 505.2365 | 580.8958 | 334.0541 | 366.4469 | 384.525 | 475.4489 | 516.475 | 202.6333 | 207.9239 | 250.2645 | 257.4489 | 245.9459 | 281.1303 | 347.2084 | 176.6573 | 419.6573 | 517.6032 | 552.7635 | 560.9479 | 245.8136 | 279.3687 | 341.8958 | 231.0261 | 140.7635 | 123.6834 | 116.4208 | 263.9198 | 323.8677 | 341.8938 | 426.0541 | 271.5531 | 244.0261 | 258.3968 | 278.501 | 231.0782 | 277.7355 | 305.1303 | 390.5531 | 332.7615 | 222.8156 | 227.4208 | 278.8938 | 205.5812 | 302.2104 | 320.9218 | 414.499 | 412.0782 | 257.0782 | 304.1844 | 479.4469 | 400.1623 | 326.501 | 358.0802 | 419.8697 | 517.7114 | 294.6292 | 329.7615 | 384.6573 | 458.6052 | 220.2625 | 212.3447 | 235.6052 | 298.5511 | 271.7355 | 280.6813 | 290.497 | 228.1323 | 195.4469 | 235.6573 | 263.4709 | 296.8156 | 448.3187 | 534.6092 | 642.7635 | 757.7595 | 422.501 | 561.499 | 766.5491 | 884.5751 | 179.5771 | 104.1583 | 112.9479 | 179.5771 | 140.3166 | 155.2104 | 176.6052 | 198.7635 | 161.7094 | 160.6854 | 171.9459 | 160.9739 | 109.6313 | 107.2625 | 113.1323 | 153.998 | 110.8156 | 152.1844 | 133.7876 | 116.7355 | 376.9198 | 380.2645 | 430.2345 | 394.501 | 246.3968 | 339.1303 | 379.1323 | 343.7375 | 480.501 | 479.6854 | 558.1042 | 557.2886
685 | 9.5294 | 276.5929 | 410.5487 | 394.2921 | 370.1504 | 253.7434 | 294.7434 | 341.9734 | 321.5221 | 216.3451 | 176.9558 | 203.5399 | 262.5929 | 304.3009 | 294.0266 | 321.5575 | 388.062 | 268.9646 | 238.3009 | 281.8142 | 327.7079 | 282.4955 | 496.5221 | 560.8406 | 352.6283 | 350.4159 | 377.3275 | 477.938 | 533.0354 | 209.7876 | 237.9646 | 265.5929 | 271.7434 | 229.8142 | 276.9912 | 318.115 | 167.9646 | 422.7699 | 489.115 | 578.9824 | 558.8584 | 235.2478 | 263.8672 | 329.5751 | 225.9646 | 132.7876 | 111.1504 | 124.177 | 260.9912 | 322.1062 | 333.0354 | 424.2743 | 291.2478 | 249.1416 | 268.1592 | 289.5929 | 217.5221 | 265.6637 | 298.1416 | 401.2921 | 316.0354 | 212.4513 | 227.8496 | 278.5841 | 217.9734 | 301.8142 | 327.5575 | 405.8496 | 401.7788 | 258.5841 | 300.3009 | 472.3717 | 430.2389 | 332.3009 | 368.7434 | 433.3717 | 506.8318 | 284.6903 | 314.6637 | 373.9292 | 453.115 | 217.4779 | 227.3451 | 248.0266 | 281.8142 | 258.8496 | 261.4337 | 267.8053 | 233.5575 | 190 | 226.9646 | 243.8053 | 297.8496 | 465.1238 | 566.5663 | 649.0354 | 716.4071 | 447.4337 | 556.5133 | 706.053 | 836.9292 | 167.9558 | 109.4071 | 113.3009 | 167.9558 | 147.1504 | 163.3629 | 173.5575 | 203.4071 | 145.7434 | 168.7079 | 165.177 | 170.1062 | 108.4425 | 101.6283 | 118.1504 | 129.7788 | 109.8142 | 145.4513 | 120.9558 | 118.5045 | 379.6903 | 401.6991 | 415.2124 | 399.8938 | 256.1592 | 335.3983 | 387.4071 | 344.4867 | 472.053 | 476.3097 | 532.9734 | 550.6637
686 | 9.5404 | 284.8068 | 399.0958 | 409.5487 | 371.8712 | 252.355 | 295.0325 | 323 | 307.4513 | 210.9361 | 178.2576 | 214.8058 | 268.71 | 316.1299 | 316.355 | 318.4194 | 404.5482 | 278.1618 | 227.4838 | 281.3225 | 327.9681 | 258.6775 | 482.0319 | 532.8717 | 342.8382 | 353.033 | 390.0969 | 454.7419 | 547.7424 | 199.5487 | 231.645 | 248.6445 | 268.2581 | 247.7749 | 272.0314 | 315.6131 | 168.7743 | 423.9995 | 478.645 | 573.2565 | 561.2586 | 245.0963 | 249.4199 | 307.8393 | 224.6775 | 125.4843 | 109.9356 | 145.2906 | 269.8707 | 341.9037 | 345.2262 | 397.4194 | 285.5796 | 238.7733 | 258.4843 | 281.8707 | 215.6131 | 266.3869 | 311.6775 | 401.4513 | 329.0649 | 215.0974 | 237.8712 | 280.5801 | 199.8387 | 293.7738 | 326.0969 | 410 | 394.1288 | 269.3869 | 294.0969 | 469.5806 | 401.6775 | 312.677 | 367.355 | 409.1927 | 473.9031 | 295.1607 | 313.29 | 365.9031 | 446.8387 | 215.6126 | 218.5812 | 266.1613 | 291.3874 | 248.7414 | 267.1607 | 281.645 | 230 | 194.1937 | 237.4199 | 252.1932 | 302.4194 | 433.6126 | 539.0319 | 636.9026 | 717.4194 | 453.1607 | 568.0969 | 661.4832 | 856.5817 | 188.5487 | 104.5482 | 101.645 | 188.5487 | 150.1293 | 160.6445 | 184.2586 | 194.7738 | 154.0005 | 164.7743 | 173.29 | 167.2576 | 106.5482 | 109.8712 | 101.8382 | 131.355 | 114.7743 | 140.5487 | 129.8063 | 130.5801 | 398.1288 | 387.7733 | 430.9031 | 381.3544 | 249.0005 | 343.6126 | 377.9351 | 325.3544 | 463.1299 | 444.8707 | 525.7424 | 563.9681
687 | 9.5514 | 290.5251 | 371.6116 | 412.2301 | 369.2084 | 251.82 | 302 | 306.1651 | 298.295 | 223.9351 | 175.3816 | 201.7916 | 278.8633 | 330.6184 | 310.6401 | 316.41 | 384.4818 | 298.4533 | 238.6765 | 270.3667 | 343.3166 | 270.9282 | 467.5467 | 540.4385 | 347.1515 | 359.8702 | 391.2733 | 446.0934 | 551.0501 | 196.1868 | 205.8485 | 233.7699 | 260.5034 | 266.6834 | 252.0649 | 319.41 | 164.9351 | 415.4966 | 480.0433 | 524.0216 | 575.3884 | 226.9784 | 278.7916 | 312.7551 | 231.3166 | 134.3884 | 116.59 | 150.115 | 259.5683 | 324.1936 | 350.6617 | 401.0216 | 267.9932 | 220.0866 | 266.0934 | 282.3599 | 218.115 | 269.9282 | 315.7267 | 388.41 | 331.7768 | 233.41 | 233.4818 | 253.7984 | 205.7483 | 290.4966 | 330.295 | 397.0501 | 399.8132 | 250.4818 | 297 | 457.7768 | 415.2232 | 289.6617 | 360.3451 | 397.9932 | 471 | 272.115 | 313.9282 | 366.4533 | 446.705 | 203.5683 | 221.7551 | 242.9852 | 289.1868 | 222.9567 | 270.8781 | 272.2517 | 220.0718 | 183.5968 | 239.5968 | 240 | 289.1868 | 421.1367 | 536.6333 | 619.885 | 717.5683 | 446.5182 | 561.0718 | 660.2882 | 875.5968 | 206.3383 | 110.3816 | 104.7699 | 206.3383 | 146.1868 | 153.5399 | 187.5968 | 190.6333 | 162.5034 | 163.9567 | 170.4749 | 163.9499 | 103.7483 | 123.18 | 112.1948 | 138.59 | 123.0216 | 142.3667 | 135.2232 | 112.8633 | 379.2084 | 370.3816 | 421.0934 | 377.4966 | 262.6834 | 348.8349 | 366.1583 | 329.2665 | 470.7118 | 451.8349 | 539.8417 | 565.9351
688 | 9.5624 | 275.3407 | 376.3363 | 408.2154 | 342.5693 | 256.5428 | 298.0044 | 298.2065 | 300 | 216.4395 | 192.556 | 212.4484 | 264.9056 | 336.9735 | 280.1297 | 320.1033 | 370.6637 | 285.2418 | 251.7846 | 263.8835 | 335.2374 | 275.5649 | 470.5472 | 531.0944 | 363.9011 | 366.6328 | 374.7935 | 438.3319 | 519.7935 | 184.2154 | 199.7537 | 244.444 | 234.2374 | 250.3451 | 265.3319 | 317.3319 | 166.7626 | 439.5428 | 478.233 | 547.9779 | 562.6726 | 222.1121 | 270.3583 | 297.1077 | 227.677 | 135.8791 | 123.5516 | 134.7891 | 259 | 295.6505 | 349.4351 | 397.8967 | 302.767 | 229.7846 | 262.3274 | 289.3407 | 217.8835 | 279.8923 | 306.9011 | 386.3319 | 312.1165 | 221.5649 | 223.2154 | 245.2198 | 218.6593 | 298.1165 | 319.1253 | 383.5516 | 410.8967 | 263.7449 | 300.5516 | 454.9868 | 426.4484 | 294.6505 | 337.0088 | 410.1253 | 476.7714 | 266.556 | 329.6637 | 365.6726 | 442.3363 | 215.8747 | 200.3363 | 221.9912 | 285.6505 | 224.767 | 281.91 | 269.9956 | 217.2109 | 185.0856 | 218 | 243.5516 | 270.1121 | 422.2198 | 520.4748 | 626.3319 | 736.09 | 447.4616 | 548.444 | 675.4572 | 830.0265 | 218.6637 | 124.4484 | 116.7758 | 218.6637 | 141.3186 | 177.5516 | 180.6505 | 193.5649 | 153.9956 | 154.6681 | 173.6726 | 183.2198 | 98.4572 | 114.4616 | 125.8023 | 131.7891 | 130.5516 | 143.8747 | 134.0177 | 116.2198 | 366.3319 | 367.1342 | 412.8879 | 373.5737 | 258.7758 | 349.6902 | 371.2198 | 336.0177 | 463.09 | 471.3363 | 504.5914 | 531.8023
689 | 9.5734 | 278.4061 | 368.8938 | 416 | 337.4936 | 261.4189 | 291.175 | 308.2438 | 291.3313 | 214.2438 | 184.2438 | 202.5251 | 244.175 | 352.4002 | 269.9499 | 312.5752 | 371.2625 | 264.3687 | 252.6687 | 277.2999 | 331.2311 | 265.4562 | 492.7749 | 534.9439 | 337.2251 | 373.5191 | 361.1062 | 452.3186 | 509.2999 | 189.2625 | 221.0688 | 232.6814 | 219.2812 | 245.6687 | 267.4562 | 290.2566 | 168.9439 | 433.0127 | 453.2251 | 570.7188 | 550.0688 | 221 | 250.2124 | 299.175 | 239.2685 | 131.0314 | 117.9626 | 132.9499 | 261.7375 | 311.7188 | 353.6126 | 391.7375 | 298.5752 | 222.4562 | 254.4002 | 277.8063 | 218.0688 | 286.9499 | 293.4562 | 397.1249 | 294.5939 | 223.7749 | 226.4375 | 253.0187 | 229.2812 | 285.6126 | 321.706 | 372.9439 | 386.9439 | 274.5378 | 293.1376 | 447.1877 | 418.8063 | 317.1062 | 328.825 | 407.2311 | 501.3373 | 268.2812 | 316.5752 | 342.5752 | 447.2124 | 220.5939 | 213.8811 | 244.3186 | 290.7689 | 242.9312 | 269.4061 | 288.2311 | 227.2625 | 204.4189 | 222.5625 | 236.1376 | 283.1436 | 425.4562 | 486 | 640.3186 | 749.5064 | 432.1062 | 546.7188 | 644.5752 | 805.4749 | 202.8377 | 114.5251 | 110.7876 | 202.8377 | 144.3126 | 173.3313 | 176.1877 | 181.175 | 155.35 | 147.5251 | 169.2812 | 173.2251 | 92.6687 | 107.7749 | 101.1937 | 129.4936 | 120.8564 | 143.5752 | 131.7936 | 125.8437 | 361.1563 | 346.9312 | 397.1189 | 341.1376 | 271.4936 | 315.7001 | 358.9439 | 325.5811 | 460.0815 | 451.1189 | 505.5183 | 527.7247
690 | 9.5844 | 282.6293 | 364 | 398.1956 | 361.7134 | 241.5663 | 292.7483 | 297.5944 | 283.3427 | 192.8181 | 184.3707 | 201.6224 | 256.0561 | 342.028 | 278.7832 | 283.1607 | 381.119 | 272.5595 | 239.5732 | 295.8393 | 340.4405 | 277.2449 | 489.7271 | 510.909 | 328.5595 | 345.3707 | 374.9022 | 436.5452 | 504.881 | 193.9651 | 220.091 | 236.7414 | 231.8393 | 260.6854 | 256.2866 | 283.1751 | 158.9651 | 409.4056 | 448.3078 | 537.0766 | 530.3495 | 221 | 258.5944 | 300.7483 | 251.5383 | 115.7483 | 107.4056 | 130.0698 | 261.2517 | 314.6224 | 331.4754 | 402.028 | 277.5944 | 214.5663 | 265.1122 | 275.4056 | 206.7832 | 271.4193 | 305.7134 | 395.8181 | 293.6505 | 220.2586 | 237.3988 | 241.6642 | 220.2866 | 280.8112 | 332.2866 | 354.0629 | 365.7202 | 236.909 | 289.7763 | 418.5944 | 411.7202 | 316.8461 | 336.1539 | 415.972 | 511.2237 | 272.8741 | 319.4897 | 320.1888 | 457.9371 | 209.8112 | 218.5383 | 245.4125 | 284.8393 | 238.2866 | 277.3776 | 288.5383 | 215.5663 | 185.5034 | 220.972 | 243.5527 | 292.5034 | 432.5595 | 477.0978 | 629.2305 | 712.6361 | 437.4685 | 539.7832 | 643.7414 | 834.0212 | 183.7483 | 106.1259 | 116.9302 | 183.7483 | 141.119 | 159.7202 | 150.4056 | 185.091 | 160.8393 | 148.4966 | 169.6573 | 158.4685 | 101.1259 | 109.881 | 101.7202 | 147.6224 | 111.4966 | 120.7202 | 139.4125 | 134 | 341.7551 | 344.6293 | 392.4617 | 345.2729 | 260.5732 | 321.0561 | 335.8461 | 337.4685 | 425.1888 | 432.8741 | 523.6081 | 566.3078
691 | 9.5954 | 275.3658 | 368.8083 | 382.8083 | 369.7109 | 235.8466 | 282.5752 | 303.2124 | 287.4425 | 210.2538 | 195 | 198.8643 | 259.0177 | 324.997 | 265.3451 | 285.944 | 403.6932 | 280 | 231.2714 | 296.9026 | 325.3068 | 274.1327 | 465.1534 | 489.0944 | 337.9233 | 342.056 | 386.4425 | 418.3481 | 503.944 | 199.1534 | 222.6726 | 246.4601 | 252.1357 | 246.2861 | 252.6165 | 304.3835 | 162.2301 | 401.8643 | 469.1357 | 528.9263 | 541.003 | 215.2301 | 247.8643 | 291.5368 | 223.9026 | 137.7906 | 107.5575 | 121.6165 | 251.2301 | 326.7699 | 332.7906 | 409.0383 | 255.3244 | 230.9647 | 287.8643 | 257.2861 | 225.0797 | 258.7139 | 301.2094 | 386.8083 | 294.5368 | 205.2508 | 239.171 | 243.1564 | 215.1741 | 293.6165 | 329.0974 | 357.8083 | 378.829 | 257.888 | 291.4218 | 435.7523 | 391.171 | 307.1534 | 365.0797 | 394.056 | 514.8673 | 276.4425 | 310.3804 | 326.1357 | 455.6342 | 213.4808 | 203.4041 | 227.115 | 295.0383 | 237.0207 | 258.4985 | 268.115 | 207.9233 | 172.6726 | 204.8259 | 253.0944 | 260.4395 | 428.4601 | 476.6165 | 594.2301 | 713.8496 | 427.9026 | 538.3658 | 671.7316 | 839.2861 | 190.4041 | 107.8466 | 127.3068 | 190.4041 | 154.0767 | 166.0974 | 158.7316 | 198.7316 | 156.1327 | 171.4248 | 164.5959 | 179.6756 | 97.5575 | 98.8466 | 103.7699 | 149.1917 | 125.7699 | 115.5575 | 137.4632 | 124.8643 | 349.6756 | 335.9233 | 406.6342 | 375.6165 | 251.3098 | 327.8643 | 329.0383 | 346.6549 | 422 | 419.6135 | 506.2714 | 567.4218
692 | 9.6064 | 271.1101 | 365.617 | 368.2753 | 348.7133 | 240 | 278.3417 | 310.0138 | 296.8899 | 222.7959 | 190.5619 | 197.383 | 243.5206 | 313.7798 | 256.9862 | 285.7133 | 411.0138 | 264.2202 | 240.6445 | 276.6308 | 319.9587 | 237.7684 | 459.5206 | 496.6697 | 337.5344 | 341.3004 | 378.6308 | 425.1514 | 506.672 | 196.9725 | 208.1376 | 223.1514 | 252.6308 | 232.3968 | 256.0963 | 295.4794 | 164.8761 | 400.383 | 479.4931 | 543.6996 | 540.3854 | 214.4243 | 251.8073 | 280.4518 | 216.9449 | 145.6583 | 102.5482 | 135.4518 | 255.8486 | 336.4518 | 344.1101 | 395.1789 | 253.711 | 245.6858 | 280.4656 | 236.9862 | 237.672 | 268.1652 | 291.3692 | 388.5482 | 272.6032 | 198.1927 | 198.8097 | 243.8097 | 220.5482 | 294.6308 | 331.1239 | 375.328 | 392.5619 | 277.6445 | 292.8211 | 445.1514 | 400.1215 | 290.1789 | 382.1101 | 400.2041 | 494.523 | 256.7959 | 290.7523 | 331.0688 | 443.617 | 212.0275 | 216.3555 | 228.9312 | 280.1927 | 254.4381 | 260.7798 | 273.383 | 221.8348 | 176.383 | 196.9862 | 242.9174 | 239.9312 | 423.3968 | 523.4908 | 556.4404 | 738.0963 | 426.3692 | 524.7409 | 653.3579 | 811.0963 | 192.0138 | 103.617 | 114.0688 | 192.0138 | 152 | 177 | 148.7409 | 206.9862 | 141.9587 | 172.6996 | 162.9862 | 191.6445 | 101.9174 | 107.9312 | 119.8624 | 136.1101 | 128.5482 | 124.8486 | 145.0963 | 111.5482 | 374.4656 | 333.0688 | 400.0413 | 385.5069 | 257.672 | 331.8073 | 326.0275 | 340.7133 | 422.9862 | 424.6972 | 504.7959 | 560.9312
693 | 9.6174 | 265.0216 | 349.4189 | 366.1947 | 340.5811 | 253.6461 | 284.4405 | 302.4297 | 309.0433 | 204.0325 | 177.408 | 197.9135 | 249.5811 | 320.3972 | 266.592 | 279.6028 | 404.9459 | 237.3864 | 223.9351 | 267 | 327.5487 | 228.6136 | 462.0433 | 509.3756 | 316.8053 | 354.4838 | 370.0108 | 420.5703 | 502.592 | 195.4621 | 195.473 | 219.0757 | 237.4405 | 234.9459 | 257.0865 | 294.9784 | 171.9784 | 399.3972 | 464.8377 | 507.2674 | 516.0433 | 211.408 | 254.881 | 293.0974 | 219.4838 | 120.8594 | 97.9892 | 146.5811 | 266.4946 | 312.7079 | 328.9243 | 381.4946 | 278.0325 | 232.5054 | 264.8053 | 265.2921 | 226.0108 | 262.5703 | 293.9243 | 365.7944 | 283.1947 | 210.7867 | 199.2488 | 241.2164 | 204.3648 | 287.0216 | 309.8702 | 397.6028 | 375.8702 | 263.4621 | 291.8485 | 436.0216 | 422.9459 | 283.5703 | 357.3215 | 391.17 | 456 | 249.6569 | 313.5487 | 341.6677 | 432.9784 | 197.3648 | 211.3323 | 252.1947 | 280.6569 | 268.0974 | 253.7512 | 269.8702 | 215.8053 | 172.8702 | 195.9784 | 245.4621 | 257.6352 | 437.5703 | 522.5889 | 551.7975 | 736.0541 | 424.3756 | 513.5703 | 600.2133 | 810.0649 | 184.9351 | 103.0865 | 108.4946 | 184.9351 | 153.0108 | 174.9784 | 144.6461 | 197.3864 | 150.5595 | 148.3972 | 170.5703 | 167.8594 | 106.4838 | 110.473 | 133.0325 | 133.5595 | 129.0433 | 134.4838 | 140.0216 | 108 | 371.9459 | 339.119 | 390.4297 | 371.8594 | 266.7326 | 325.7836 | 337.6461 | 342.1839 | 420.9675 | 435.881 | 477.4405 | 556.3972
694 | 9.6284 | 267 | 339.4115 | 372.6106 | 327.8097 | 251.469 | 267.6106 | 299.6239 | 303.1637 | 211.1416 | 164.8584 | 188.9646 | 248.7168 | 308.4115 | 250.1504 | 291.5885 | 403.6239 | 245.6372 | 225.2478 | 285.6372 | 342.354 | 252.9779 | 458.7522 | 499.0354 | 306.7655 | 358.6947 | 368.4115 | 438.9071 | 520.3186 | 211.1549 | 199.2124 | 224.9646 | 239.7655 | 240.8717 | 251.5398 | 293 | 159.1283 | 417.4381 | 448.9646 | 468.9159 | 544.8496 | 203 | 246.0708 | 294.2345 | 221.1062 | 118.8717 | 115.6372 | 136.3982 | 276.354 | 306.1903 | 335.9779 | 384.1062 | 273.2345 | 214.3628 | 241.3053 | 281.1283 | 233.7301 | 278.8363 | 295.2832 | 363.0487 | 298.4115 | 229.2345 | 227.6947 | 252.1991 | 188.8938 | 301.9425 | 312.4956 | 380.0796 | 367.6239 | 258.9646 | 268.1637 | 427.1283 | 401.9513 | 313.2965 | 350.3186 | 374.9779 | 458.5885 | 265.5531 | 318 | 356.4823 | 450.1549 | 188.1062 | 228.0974 | 244.115 | 297.0708 | 259.3982 | 245.4956 | 274.0487 | 204.7301 | 174.9779 | 198.6593 | 251.3186 | 257.5752 | 440.3761 | 492.7655 | 566.5752 | 707.8672 | 409.3761 | 502.9159 | 587.4602 | 797.8805 | 175.8938 | 107.8938 | 100.7522 | 175.8938 | 166.4248 | 159.5398 | 150.2345 | 182.8584 | 157.5531 | 150.3894 | 172.3407 | 163.8009 | 113.2832 | 100.2345 | 132.8938 | 140.5531 | 128.8584 | 126.7876 | 146.6593 | 116.8009 | 351.9867 | 347.4115 | 403.6726 | 388.5089 | 268.8451 | 316.8717 | 348.9292 | 352.2699 | 414.8938 | 432.2478 | 463.1991 | 541.823
695 | 9.6394 | 262.2301 | 349.7198 | 360.94 | 320.1298 | 247.5998 | 272.3698 | 300.35 | 301.9499 | 209.7001 | 173.7198 | 192.2399 | 234.1101 | 312.8899 | 242.3698 | 295.59 | 399.5801 | 249.2202 | 252.1397 | 288.6902 | 326.5604 | 269.1799 | 458.3599 | 510.0698 | 324.5998 | 356.5801 | 385.6097 | 438.8702 | 519.9901 | 218.4002 | 209.2399 | 226.65 | 252.8299 | 250.47 | 253.3098 | 286.1101 | 155.3599 | 420.7404 | 441.6401 | 499.1093 | 553.1602 | 225.2596 | 241.6401 | 287.53 | 221.88 | 122.6401 | 116.0403 | 117.88 | 252.6105 | 317.5801 | 357.4799 | 362.6204 | 274.4799 | 232.5094 | 234.41 | 272.06 | 232.5801 | 289.59 | 303.53 | 359.4503 | 298.65 | 203.4503 | 224.5201 | 236.5801 | 202.4297 | 291.2104 | 301.6204 | 390.5094 | 398.0296 | 251.6401 | 278.6097 | 402.6902 | 392.8299 | 315.9203 | 348.9302 | 380.5801 | 476.8997 | 270.1799 | 306.3402 | 345.4002 | 466.41 | 194.7099 | 252.7001 | 227.9499 | 298.47 | 241.41 | 260.06 | 293.7699 | 201.4602 | 195.4199 | 218.8997 | 252.0501 | 243.35 | 430.6401 | 477.7404 | 542.2802 | 682.2999 | 417.1298 | 489.12 | 602.12 | 768.8102 | 173 | 118.2498 | 108.8398 | 173 | 155.7404 | 168.7296 | 149.8899 | 181.65 | 145.7502 | 159.41 | 158.4602 | 157.1602 | 114.1101 | 101.4799 | 120.9901 | 136.7001 | 112.8102 | 119.41 | 152.06 | 114.9302 | 356.5497 | 341.8201 | 409.2802 | 388.4404 | 238.2802 | 313.7502 | 367.1397 | 349.1799 | 404.0501 | 419.3903 | 448.1101 | 540.1799
696 | 9.6504 | 265.592 | 359.3732 | 370.3034 | 350.6067 | 262.9651 | 262.6819 | 297.4577 | 286.2241 | 207.7114 | 183.9154 | 175.9355 | 249.149 | 310.8658 | 233.2241 | 297 | 403.8456 | 257.0645 | 248.851 | 272.204 | 296.408 | 263.8658 | 455.8658 | 497.3087 | 324.7812 | 355.4228 | 381.3087 | 436.592 | 510.9154 | 210.6268 | 217.3383 | 231.1691 | 242.2738 | 248.3732 | 261.6617 | 279.4577 | 155.5772 | 383.1047 | 465.2833 | 523.7315 | 558.4376 | 235.2389 | 250.0993 | 273.3584 | 229.2188 | 126.2188 | 128.6564 | 123.0497 | 251.1987 | 306.1195 | 336.259 | 354.0148 | 281.4577 | 240.1396 | 251.4376 | 271.9154 | 221.1195 | 291 | 310.5074 | 337.5772 | 306.4228 | 194.0497 | 209.408 | 229.4577 | 203.9852 | 279.6765 | 297.353 | 390.5476 | 411.1544 | 255.2188 | 284.0698 | 401.3879 | 411.5571 | 300.0846 | 351.9302 | 373.4577 | 476.3584 | 281.1342 | 290.0349 | 337.6268 | 454.6966 | 194.204 | 244.204 | 250.7262 | 281.1892 | 237.8309 | 265.3383 | 280.1047 | 198.9651 | 201.4577 | 219.4429 | 243.3732 | 252.9302 | 410.3584 | 471.5571 | 557.0296 | 694.0497 | 427 | 467.6819 | 569.8362 | 763.9651 | 162.6966 | 109.3933 | 109.5275 | 162.6966 | 143.0497 | 168.4778 | 165.7611 | 189.4228 | 132.3732 | 165.0497 | 167.8953 | 141.2886 | 103.6617 | 102.4926 | 111.9154 | 134.7114 | 115.0148 | 114.204 | 137.8161 | 107.6268 | 363.7812 | 335.7463 | 397.4577 | 360.204 | 245.9799 | 295.4926 | 359.5127 | 353.0846 | 406.2188 | 397.6617 | 441.4577 | 531.0698
697 | 9.6614 | 242.053 | 343.0265 | 372.8997 | 350.2536 | 264.6726 | 260.1918 | 284.2448 | 278.0915 | 224.4189 | 189.1003 | 172.8643 | 253.6902 | 296.236 | 240.065 | 294.7817 | 375.8436 | 264.0177 | 232.6637 | 246.8171 | 321.6107 | 260.882 | 447.3363 | 482.3274 | 330.3098 | 346.6902 | 350.2448 | 423.5899 | 517.2095 | 212.5546 | 221 | 215.8082 | 240.0915 | 264.7464 | 238.0354 | 277.8909 | 157.6549 | 374.6549 | 464.4896 | 493.9174 | 542.9439 | 213.6902 | 242.5899 | 252.6814 | 241.4366 | 132.8909 | 133.1445 | 122.3451 | 271.118 | 298.118 | 319.2095 | 346.6991 | 277.6726 | 220 | 266.4454 | 264.3451 | 232.5281 | 265.4896 | 294.9262 | 339.6549 | 317.6549 | 190.0177 | 204.1091 | 230.1091 | 199.6549 | 283.118 | 325.8643 | 359.5634 | 378.7079 | 247.4719 | 269.5634 | 425.9735 | 403.5899 | 291.0177 | 340.9262 | 373 | 478.4189 | 270.2536 | 286.6637 | 312.935 | 453.2095 | 207.6372 | 232.6814 | 241.8171 | 271.4366 | 266.5015 | 271.7729 | 265 | 192.354 | 198.7817 | 205.7817 | 218.7079 | 268.5457 | 407.9823 | 468.0265 | 565.6902 | 688.354 | 413.6902 | 459.646 | 592.5753 | 779.5369 | 172.8555 | 91.4454 | 105.1003 | 172.8555 | 149 | 164.2006 | 160.6902 | 193.4454 | 149.8555 | 148.8171 | 189.6549 | 158.9646 | 117.7464 | 95.8909 | 107.118 | 129.236 | 130.9912 | 131.5192 | 127.2183 | 96.7994 | 346.0177 | 339.1003 | 415.3009 | 373.6372 | 266.4366 | 306.6372 | 329.1268 | 346.236 | 409.0088 | 424.5015 | 455.4189 | 527.6549
698 | 9.6724 | 226.503 | 320.7964 | 364.5988 | 341.1018 | 247.8284 | 277.4012 | 276.2675 | 289.2675 | 221.2615 | 178.6947 | 179.4651 | 233.3633 | 306.4391 | 264.4012 | 292.4331 | 356.7704 | 266.2036 | 235.1337 | 246.5728 | 326.026 | 280.4391 | 437.497 | 486.7006 | 320.026 | 336.5669 | 340 | 388.7266 | 493.523 | 211.8663 | 200.5929 | 220.1397 | 230.2935 | 240.4532 | 267.3492 | 277.5669 | 156.7645 | 393.0379 | 465.5409 | 486.6686 | 521.4012 | 209.8024 | 247.4071 | 253.3692 | 233.0958 | 128.5988 | 121.9681 | 123.2355 | 266.2994 | 312.5728 | 329.5349 | 369.9799 | 252.3254 | 227.9361 | 237.0899 | 238.5929 | 253.503 | 245.5669 | 287.6367 | 351.2355 | 298.6248 | 202.974 | 204.4331 | 250.8403 | 208.4012 | 277.1657 | 301.8544 | 352.0319 | 389.477 | 244.0698 | 268.2675 | 422.6947 | 395.9361 | 295.471 | 342.7065 | 385.471 | 510.974 | 243.529 | 296.503 | 310.7065 | 472.0379 | 210.2296 | 221.4651 | 215.497 | 273.8663 | 265.4911 | 265.497 | 251.3952 | 198.3053 | 189.6308 | 188.6947 | 220.974 | 249.7586 | 405.7964 | 478.6746 | 535.7266 | 718.1137 | 405.8343 | 479.7704 | 600.555 | 795.3692 | 193.1018 | 92.7006 | 107.1657 | 193.1018 | 144.4651 | 167.7645 | 155.6686 | 183.9302 | 165 | 144.0379 | 167.2236 | 169.8982 | 120.6627 | 104.6367 | 119.3053 | 133.2036 | 124.7964 | 128.7266 | 127.2994 | 94.3692 | 346.503 | 341.1657 | 403.9242 | 388.7006 | 262.0639 | 307.529 | 346.3752 | 346.2355 | 411.2355 | 420.0899 | 457.3633 | 518.2615
699 | 9.6834 | 224.8918 | 306.6293 | 374.7414 | 348.4749 | 228.8918 | 291.5831 | 283.7916 | 279.996 | 208.8417 | 180.3206 | 204.9578 | 238.1623 | 293.5172 | 263.5251 | 286.2084 | 361.5251 | 282.1082 | 233.1042 | 240.3087 | 320.4248 | 268.0963 | 442.5831 | 485.6834 | 313.2665 | 332.3668 | 357.3747 | 389.058 | 465.7876 | 212.1583 | 193.1082 | 232.2625 | 232.2164 | 233.0079 | 259.7216 | 284.9499 | 168.7955 | 401.2625 | 446.0924 | 492.7375 | 522.2625 | 217.8958 | 242.1504 | 251.471 | 222.2625 | 131.7916 | 118.6293 | 119.3127 | 255.7335 | 323.1042 | 334.7375 | 391.471 | 237.8958 | 235.7375 | 250.9076 | 217.7876 | 226.409 | 257.5831 | 288.6293 | 339.7836 | 300.2704 | 209.7335 | 199.9459 | 245.7296 | 200.5752 | 280.2124 | 297.0079 | 345.5251 | 396.942 | 260.2665 | 277.529 | 394.7836 | 399.6834 | 293.4169 | 341.3628 | 383.996 | 510.7836 | 242.1082 | 291.996 | 312.8377 | 481.4208 | 204.3166 | 215.1042 | 220.5831 | 274.7375 | 246 | 242.2045 | 262.4288 | 216.9499 | 184 | 190.3206 | 226.5752 | 254.5871 | 395.1042 | 471.0963 | 518.1042 | 711.5132 | 414.3707 | 483.3668 | 543.1464 | 799.8417 | 174.9921 | 96.8958 | 122.3747 | 174.9921 | 147.3707 | 163.5792 | 146.1003 | 186.8456 | 152.8377 | 163.2665 | 153.529 | 159.6293 | 106.7876 | 113.7375 | 125.2084 | 137.5251 | 131.4789 | 127.3206 | 132.3707 | 109.2665 | 353.5792 | 329.7335 | 377.0501 | 386.5251 | 258.3166 | 313.058 | 355.7335 | 342.8918 | 400.942 | 393.9459 | 437.8377 | 507
700 | 9.6944 | 249.7552 | 318.5605 | 360.9336 | 343.3112 | 245.475 | 281.6652 | 286.2257 | 283.2375 | 220.4204 | 165.1592 | 187.556 | 252.9145 | 292.152 | 233.202 | 297.3776 | 350.1283 | 273.8053 | 251.7006 | 247.8835 | 306.1165 | 247.0428 | 440.354 | 489.9145 | 329.4631 | 327.2257 | 359.9454 | 405.9145 | 485.6579 | 197.6224 | 200.7743 | 220.9454 | 237.4941 | 254.9145 | 234.323 | 260.4277 | 177.4513 | 400 | 457.8526 | 457.9218 | 540.5177 | 217.6342 | 216.354 | 260.7862 | 213.3112 | 136 | 112.2257 | 115.9572 | 260.2375 | 300.2994 | 320.0309 | 378.3112 | 266.0237 | 221.0309 | 250.3849 | 240.0237 | 216.1947 | 264.8171 | 268.6224 | 327.4086 | 299.6224 | 195.9026 | 202.323 | 230.4086 | 202.4631 | 286.3658 | 310.6342 | 347.1401 | 392.5059 | 264.6342 | 272.354 | 406.0664 | 411.0118 | 299.1947 | 332.9572 | 367.7197 | 483.0309 | 288.2184 | 284 | 308.1401 | 494.6033 | 212.5059 | 215.9572 | 234.323 | 280.1401 | 235.9454 | 255.4322 | 272.0855 | 216.677 | 174.5369 | 222.475 | 216.6342 | 275.7316 | 420.798 | 456.5487 | 530.1947 | 693.2803 | 439.7006 | 468.7625 | 522.3658 | 777.7079 | 166.6888 | 99 | 119.6224 | 166.6888 | 140.7625 | 149.2138 | 148.4204 | 179.8053 | 154.0546 | 155.8053 | 170.2375 | 163.8717 | 105.9572 | 121.5059 | 117.4513 | 132.0428 | 137.2684 | 128.7197 | 128.1283 | 113.6342 | 335.6652 | 343.1092 | 387.9691 | 348.5132 | 245.8053 | 313.9454 | 324.7507 | 339.3658 | 403.6033 | 385.677 | 433.1401 | 504.6342
701 | 9.7054 | 249.8692 | 323.9626 | 353.4523 | 333.1701 | 246.5103 | 280.471 | 291.6411 | 303.0748 | 208.4729 | 148 | 171.8673 | 240.0953 | 306.415 | 235.7346 | 279.6243 | 347.0187 | 273.0561 | 238.8318 | 252.7365 | 295.2262 | 246.2075 | 442.0561 | 477.6991 | 308.8318 | 317.5477 | 358.4336 | 418.4523 | 486.3028 | 194.2449 | 202.6037 | 215.8112 | 235.415 | 238.4729 | 241.6224 | 260.5271 | 165.3402 | 406.0374 | 458.643 | 468.6205 | 532.8692 | 231.6972 | 216.2449 | 254.0953 | 198.9439 | 132.9813 | 117.6411 | 109.5477 | 244.4542 | 287.2075 | 325.3009 | 368.1701 | 267.0953 | 202.1514 | 254.3383 | 258.6037 | 205.4729 | 267.0187 | 270.6785 | 334.2449 | 297.4523 | 203.0748 | 203.585 | 214.3028 | 178.8131 | 277.1327 | 304.7738 | 347.585 | 395.1888 | 250.3215 | 267.415 | 408.114 | 399.4729 | 307.1888 | 355.5271 | 352.3402 | 484.0748 | 291.8505 | 298.4897 | 311 | 462.9458 | 217.6037 | 211.3589 | 229.5477 | 260.6617 | 251.3383 | 264.5477 | 287.9233 | 207.5664 | 174.6411 | 203.5869 | 206.5477 | 269.9439 | 410.4169 | 439.6804 | 532.1514 | 716.5084 | 445.5477 | 464.0187 | 538.4897 | 770.2449 | 189.5084 | 96.585 | 109.6037 | 189.5084 | 145.6785 | 142.6224 | 155.1888 | 194.7532 | 155.5664 | 145.3963 | 153.8505 | 157.9252 | 114.6411 | 121.7738 | 104.7365 | 133.6224 | 134 | 121.1888 | 134.0748 | 112 | 340.5084 | 339.4916 | 392.3589 | 334.2635 | 249.2822 | 310.6224 | 320.0748 | 322.8879 | 434.7346 | 385.6224 | 442.0374 | 506.0187
702 | 9.7164 | 234.7679 | 371.7464 | 369.4725 | 331.2321 | 225.9115 | 310.7931 | 287.9916 | 310.384 | 189.4557 | 155.3923 | 176.5359 | 217.9115 | 287.055 | 250.616 | 271.4725 | 361.3206 | 283.9282 | 221 | 243.5359 | 286.2955 | 263.0167 | 447.384 | 426.3421 | 309.4809 | 336.6412 | 377.4007 | 427.0801 | 494.4809 | 186.0634 | 188.8313 | 212.6878 | 256.0969 | 223.1519 | 241.5359 | 258.5993 | 156.5359 | 405.0718 | 421.055 | 480.9115 | 522.6962 | 221.671 | 229.6244 | 238.0718 | 196.6962 | 127.9199 | 117.0718 | 123.0969 | 240.7045 | 304.0167 | 319.3672 | 376.0885 | 259.0801 | 220.4091 | 245.5909 | 252.2237 | 206.1687 | 252.9833 | 269.1435 | 338.384 | 301.7679 | 182.6627 | 202.616 | 218.1687 | 174.4007 | 266.9199 | 295.2237 | 342.3038 | 373.671 | 254.3206 | 295.4892 | 387.2955 | 406.945 | 304.7679 | 359.7596 | 341.0718 | 503.0885 | 258.1352 | 297.5275 | 315.3122 | 454.4809 | 206.9115 | 227.329 | 222.7679 | 247.8481 | 253.6794 | 262.0801 | 288.0634 | 204.616 | 171.6077 | 192.7045 | 209.6244 | 247.3672 | 375.5993 | 422.0718 | 551.6412 | 683.0299 | 413.5108 | 446.2871 | 542.4557 | 775 | 194.8397 | 101.7763 | 128.4809 | 194.8397 | 152.7679 | 150.5443 | 161.3923 | 183.8947 | 149.5359 | 143.7679 | 148.4725 | 154.9282 | 94.3588 | 109.7596 | 112.1687 | 137.2321 | 122.9115 | 124.3122 | 129.0634 | 122.4725 | 345.8397 | 337.2404 | 381.8397 | 350.8565 | 259.8481 | 316.6962 | 345.8648 | 329.4809 | 422.5108 | 391.0801 | 424.439 | 504.9199
703 | 9.7274 | 226.4602 | 346.8229 | 344.584 | 333.885 | 229.0531 | 290.4424 | 306.2478 | 303.7168 | 191.0266 | 171.9381 | 171.2301 | 241.1594 | 286.5664 | 239.0619 | 270.4602 | 380.3098 | 244.2742 | 237.3363 | 231.3186 | 299.1062 | 243.469 | 416.8406 | 426.7611 | 319.7434 | 358.2832 | 366.8938 | 407.0088 | 487.1504 | 182.3982 | 183.1416 | 217.5398 | 246.6371 | 228.2832 | 219.2654 | 264.0797 | 146.8318 | 396.9734 | 427.4779 | 490.3363 | 496.0973 | 214.6549 | 221.8053 | 256.9912 | 231.0443 | 122.2301 | 97.6637 | 127.4602 | 241.0885 | 292.0088 | 322.1947 | 376.7168 | 278.593 | 211.1238 | 225.3451 | 242.9734 | 204.9469 | 237.9734 | 284.3629 | 341.7699 | 301.6283 | 182.593 | 206.1416 | 236.4248 | 199.0797 | 277.5664 | 289.115 | 359.4779 | 374.823 | 251.9469 | 284.3539 | 393.823 | 420.6283 | 325.9912 | 338.292 | 346.1682 | 486.1238 | 245.7434 | 285.9734 | 301.6637 | 461.6018 | 189.3186 | 224.292 | 220.7434 | 257.1682 | 240.3451 | 241.3805 | 256.1238 | 210.6549 | 188.9912 | 208.1682 | 199.9203 | 238.885 | 384.8496 | 433.4514 | 536.2123 | 713.4337 | 395.115 | 444.6814 | 533.3451 | 770.6018 | 184.7168 | 109.7699 | 106.0707 | 184.7168 | 157.0266 | 144.5752 | 160.3451 | 192.7611 | 153.0266 | 164.3629 | 163.1682 | 163.0266 | 85.9115 | 100.2301 | 119.7434 | 129.2035 | 118.5133 | 123.8584 | 130.4248 | 120.8318 | 352.6814 | 355.5664 | 378 | 363.2832 | 272.3098 | 297.0088 | 333.7522 | 346.0266 | 394.0619 | 391.115 | 430.4779 | 509.2832
704 | 9.7384 | 214.9533 | 341.1869 | 366.3585 | 338.2031 | 233.0782 | 256.5472 | 273.876 | 289.7502 | 211.9371 | 163.6253 | 180.531 | 259.6406 | 304.8904 | 249.9371 | 264.7188 | 386.3594 | 259.3585 | 227.1411 | 234.6092 | 307.7969 | 238.8122 | 420.1402 | 470.1869 | 331.8122 | 329.9695 | 369.7341 | 416.4214 | 482.4061 | 167.0629 | 178.5939 | 218.7969 | 226.3127 | 256.9838 | 223.6559 | 274.2812 | 152.8904 | 406.8122 | 437.7188 | 485.2659 | 502.062 | 214.1563 | 231.2965 | 262.3127 | 225.0162 | 111.0314 | 97.6092 | 130.0467 | 234.4376 | 300.7808 | 324.2345 | 375.5624 | 259.5319 | 215.5777 | 236.7341 | 243.2031 | 206.6873 | 265.7493 | 261.1258 | 329.9066 | 288.5472 | 200.0467 | 220.8122 | 236.8751 | 206.0782 | 271.5472 | 282.2345 | 367.7969 | 370.8284 | 222.938 | 274.4061 | 392.3908 | 412.0314 | 328.1096 | 331.5624 | 344.5939 | 473.2812 | 261.0153 | 279.7969 | 296.4843 | 438.5005 | 190.6873 | 199.6253 | 220.6406 | 265.8437 | 238.9218 | 255.5463 | 277.2336 | 201.1878 | 190.469 | 205.3127 | 215.9838 | 236.1563 | 390.2345 | 461.2183 | 527.6092 | 735.8284 | 399.1249 | 433.7035 | 528.0782 | 757.1096 | 178.4376 | 111.3594 | 93.6873 | 178.4376 | 144.6253 | 161.4214 | 179.4214 | 199.0314 | 160.4843 | 154.5786 | 150.7035 | 164.7188 | 95.7655 | 110.8122 | 112.5939 | 152.8275 | 121.2812 | 124.5624 | 127.6721 | 121.1249 | 351.9533 | 326.4852 | 384.4061 | 375.9686 | 250.8131 | 307.062 | 313.5939 | 327.8598 | 399.1716 | 391.2812 | 425.3441 | 510.4376
705 | 9.7494 | 231.8929 | 344.5949 | 374.6362 | 342.6116 | 226.7768 | 258.1406 | 255.8348 | 297.7109 | 214.1652 | 168.058 | 173.2891 | 251.5123 | 305.1239 | 249.5536 | 272.4877 | 359.8839 | 259.8013 | 233.6284 | 237.3884 | 296.8594 | 257.7522 | 418.942 | 475.5536 | 327.9007 | 317.2232 | 377.3471 | 406.1484 | 501.0167 | 198.1328 | 180.8761 | 205.2478 | 262.4799 | 245.5368 | 217.9833 | 276.9587 | 163.5703 | 406.8181 | 430.471 | 450.231 | 516.9587 | 217.8761 | 243.6116 | 259.3058 | 229.9342 | 120.3638 | 118.0167 | 129.4297 | 243.4464 | 297.3304 | 332.0993 | 388.0993 | 246.2645 | 224.0413 | 237.8181 | 252.8348 | 222.7522 | 260.1484 | 289.3147 | 326.5703 | 299.9342 | 187.0245 | 218.2065 | 228.7768 | 214.1406 | 266.6116 | 285.529 | 379.058 | 373.1406 | 210 | 266.2478 | 395.4877 | 416.7935 | 293.9252 | 301.6607 | 342.3058 | 477.2645 | 261.5123 | 282.5703 | 300.9587 | 427 | 193.0413 | 217.1161 | 235.3638 | 270.6116 | 234.1239 | 252.7187 | 272.1897 | 195.9587 | 174.8594 | 192.5123 | 203.8839 | 241.1819 | 381.7768 | 462.8594 | 502.9665 | 674.1562 | 399.3884 | 432.4877 | 524.3884 | 755.5703 | 172.4297 | 92.7187 | 104.529 | 172.4297 | 141.2232 | 163.5536 | 158.0491 | 195.9587 | 142.1071 | 148.529 | 169.0748 | 160.7355 | 100.3058 | 114.7355 | 114.8761 | 133.7433 | 123.3058 | 124.0413 | 127.5703 | 115.5123 | 350.6116 | 326.24 | 365.1484 | 364.0245 | 233.0413 | 297.8013 | 316.529 | 336.24 | 398.8181 | 394.6116 | 408.1652 | 497.9007
706 | 9.7604 | 233.6873 | 346.6475 | 380.9691 | 323.379 | 222.0044 | 300.2508 | 264.6564 | 295.6873 | 206.674 | 181.6519 | 161.3436 | 228.3746 | 302.6652 | 226.705 | 278.9956 | 349.326 | 274.6077 | 233.6873 | 237.9956 | 293.3304 | 250.6961 | 450.5812 | 457.3613 | 307.3658 | 316.0088 | 362.3658 | 395.9956 | 485.0531 | 202.0398 | 188.6564 | 215.6298 | 248.0796 | 225.3481 | 206.0088 | 254.7182 | 156.6873 | 381.7138 | 458.2641 | 444.6475 | 518 | 201.7094 | 224.379 | 270.6431 | 210.7315 | 128 | 117.3525 | 120.3481 | 232.3702 | 310.2862 | 328.0221 | 380.6961 | 276.6033 | 221.0044 | 232.6696 | 239.705 | 221.6829 | 259.3083 | 302.3613 | 340.3083 | 311.9956 | 185.9823 | 218.9867 | 218.0177 | 197.0486 | 247.379 | 289 | 369.3702 | 390.6387 | 241.2641 | 277.9602 | 389.3569 | 412.0221 | 294.2994 | 316.2641 | 341.6696 | 455.0531 | 247.6873 | 277.0177 | 302.6652 | 453.6077 | 177.3658 | 224.3481 | 246.9912 | 254.705 | 227.0088 | 219.0531 | 266.3171 | 187.0221 | 175.9867 | 186.0044 | 193.9912 | 259.6342 | 395.6298 | 465.3171 | 506.2906 | 663.6254 | 393.3436 | 463.6077 | 517.0133 | 752.6785 | 185.9646 | 79.674 | 105.3392 | 185.9646 | 130.6961 | 156.6696 | 159.2906 | 188.3525 | 133.6608 | 153.3304 | 166.705 | 167.6475 | 111.6431 | 112.3348 | 116.0044 | 130.2994 | 136.6387 | 146.2818 | 141.3083 | 113.6608 | 361.9779 | 356.2906 | 358.9867 | 378.9469 | 237.9867 | 302.6298 | 314.6785 | 353.6519 | 402.9823 | 379.3702 | 399.3436 | 497.9867
707 | 9.7714 | 218.1603 | 353.7099 | 354.8033 | 338.7444 | 236.5821 | 267.5112 | 269.7099 | 303.1947 | 196.1927 | 181.6126 | 190.8418 | 230.5172 | 293.5152 | 220.7099 | 264.4179 | 355.0649 | 271.0629 | 219.5152 | 239.6775 | 285.8702 | 267.0345 | 447.643 | 461.8397 | 312.8722 | 310.6126 | 360.4199 | 395.645 | 459.4503 | 190.2901 | 181.5152 | 218.5801 | 238.1298 | 232.8397 | 221.6146 | 246.3874 | 147.2576 | 366.9351 | 445.5457 | 453.7424 | 523.4199 | 189.2901 | 215.355 | 270.5801 | 212.9371 | 135.4523 | 117.0649 | 112.2576 | 230.7748 | 312.1927 | 327.0649 | 371.2901 | 269.998 | 225.4199 | 222.5152 | 242.5172 | 220.0325 | 253.4828 | 276.353 | 340.5801 | 296.0629 | 183.2252 | 220.645 | 209.9351 | 186.6775 | 260.7119 | 283.5801 | 354.2252 | 376.6755 | 262.4199 | 280.9026 | 368.4179 | 409.0325 | 310.8073 | 310.3205 | 378.2616 | 477.5517 | 261.2921 | 290.6146 | 302.3225 | 453.4503 | 203.1967 | 231.8397 | 234.7728 | 235.8377 | 227.0325 | 221.2272 | 273.3874 | 186.0649 | 187.1298 | 173.4828 | 200.0649 | 264.9675 | 384.353 | 433.7708 | 487.5457 | 694.002 | 396.4199 | 450.5781 | 507.2252 | 695.8013 | 187.2252 | 87.4848 | 104.6775 | 187.2252 | 134.1623 | 132.9655 | 162.9026 | 173.8377 | 144.4848 | 149.2576 | 172.2272 | 163.1927 | 92.6106 | 101.8377 | 115.6775 | 148.1623 | 124.0304 | 138.353 | 143.6126 | 102.1278 | 350.7404 | 354.4828 | 362.6775 | 379.4828 | 234.9026 | 309.645 | 324.8722 | 349.5477 | 402.2576 | 378.4523 | 401.0649 | 505.7424
708 | 9.7824 | 214.3796 | 326.0294 | 356.9342 | 333.0658 | 235.0329 | 260.9342 | 271.6898 | 296.584 | 185.792 | 177.2409 | 188.7556 | 243.5875 | 284.1716 | 252.3501 | 278.383 | 351.4818 | 242.9965 | 232.9342 | 255.8648 | 297.1751 | 244.8908 | 443.8978 | 468.3796 | 336.2444 | 313.8284 | 354.7227 | 375.6863 | 452.3102 | 170.3761 | 183.8978 | 218.7591 | 236.4818 | 224.2045 | 238.2773 | 266.6239 | 150.5182 | 364.3102 | 414.0658 | 473.2444 | 518.4125 | 197.6569 | 202.8943 | 241.7885 | 234.1057 | 123.8249 | 112.1022 | 122.416 | 231.2409 | 294.8943 | 324.8613 | 376.8978 | 274.4853 | 225.2409 | 220.0693 | 241.1022 | 218.9307 | 236.9636 | 276.6569 | 326.9636 | 320.4194 | 193.1057 | 205.5147 | 249.3865 | 197.3466 | 250.617 | 269.9636 | 335.1352 | 373.8978 | 245.6863 | 263.5147 | 366.5182 | 393.4454 | 288.0988 | 322.7626 | 364.2703 | 482.9636 | 241.0294 | 288.6534 | 299.9307 | 434.584 | 198.9965 | 228.7227 | 223.1716 | 264.1092 | 222.4818 | 230.7591 | 251.5476 | 205.2444 | 173.7556 | 194.9012 | 182.6863 | 250.8943 | 369.4818 | 416.3102 | 473.3102 | 725.383 | 394.1716 | 418.6863 | 504 | 691.383 | 180.5511 | 89.2409 | 102.2409 | 180.5511 | 143.8284 | 149.591 | 164.1387 | 179.3466 | 125.5476 | 135.9636 | 170.7227 | 171.416 | 95.4853 | 103.8978 | 113.9307 | 129.5476 | 125.3466 | 139.3466 | 137.8613 | 117.383 | 345.7591 | 352.4489 | 355.4125 | 356.7556 | 237.5182 | 302.1022 | 319.9636 | 350.8284 | 404.1387 | 346.8215 | 407.1387 | 521.1057
709 | 9.7934 | 215.7021 | 331.2537 | 376.2331 | 344.6579 | 205.7256 | 266.8938 | 274.8083 | 270.6401 | 189.3186 | 183.7227 | 162.7463 | 234.3628 | 304.4661 | 287.062 | 265.5339 | 360.9351 | 268.4012 | 208.705 | 232.8112 | 299.7876 | 244.0413 | 440.6814 | 457.0649 | 318.7256 | 330.7434 | 341.1711 | 374.0207 | 480.7847 | 178.8496 | 209.4661 | 196.8318 | 230.4897 | 234.8496 | 251.8289 | 275 | 157.2124 | 402.6136 | 429.8703 | 462.0442 | 469.3657 | 207.6165 | 214.5517 | 229.2979 | 217.5339 | 119.1062 | 101.2773 | 113.9587 | 218.0649 | 298.1269 | 313.1711 | 360.3421 | 269.7669 | 214.8731 | 237.1475 | 234.4897 | 219.4041 | 264.295 | 273.9793 | 343.062 | 314.6401 | 200.4041 | 203.9145 | 228.6843 | 198.4897 | 246.6165 | 260.0855 | 335.7227 | 358.7463 | 226.469 | 244.3628 | 366.1917 | 378.9793 | 287.9351 | 340.7227 | 331.0442 | 469.5752 | 238.531 | 282.5959 | 285.6607 | 457.7847 | 176.6607 | 211.6607 | 218.8938 | 249.5132 | 238.9558 | 239.0207 | 254.3393 | 211.5959 | 170.2124 | 178.9174 | 209.1033 | 242.1917 | 383.8496 | 422.3186 | 468.0855 | 730.0855 | 384.9793 | 409.2979 | 496.9793 | 735.4012 | 181.8083 | 97.8289 | 106.6165 | 181.8083 | 152.3186 | 140.2359 | 140.7256 | 163.6401 | 123.4248 | 137.3186 | 164.8938 | 165.0649 | 101.2979 | 106.2979 | 111.5959 | 130.2331 | 132.1062 | 129.9587 | 120.5545 | 105.938 | 315.407 | 345.5752 | 366.0413 | 356.0207 | 231.5752 | 282.8525 | 303.7669 | 338.2566 | 397.5752 | 367.5074 | 397.7669 | 515.0649
710 | 9.8044 | 223.1436 | 323.2841 | 358.1406 | 325.1401 | 205.7154 | 307.4327 | 270.9995 | 260.5713 | 192 | 164.1406 | 171.4302 | 229.7144 | 306.8564 | 257.4247 | 219.5678 | 350.9985 | 270.1416 | 230.0035 | 230.7154 | 290.1421 | 228.5693 | 430.1421 | 449.1426 | 325.1446 | 311.2831 | 346.2866 | 399.8594 | 456.5678 | 186 | 204.7129 | 207.7164 | 223.2851 | 261.2876 | 257.4287 | 254.998 | 153.9995 | 387.5683 | 403.5678 | 452.5713 | 480.003 | 202.1421 | 212.7134 | 259.003 | 206.5713 | 128.5723 | 86.5703 | 126.5733 | 215.1431 | 292.7134 | 315.4292 | 352.7144 | 265.7144 | 223.1441 | 236.8564 | 236.5718 | 214.2851 | 250.1401 | 269.5713 | 324.8544 | 316.7154 | 214.5728 | 210.2861 | 212 | 181.2841 | 241.8564 | 268.7154 | 374.0035 | 354.5718 | 214.1421 | 258.2876 | 365.7144 | 388.1441 | 319.4312 | 315.4257 | 362.2896 | 468.8574 | 238.7139 | 246.9965 | 312.1461 | 459.999 | 186.0015 | 185.998 | 222.2861 | 265.5743 | 241.2851 | 223.4267 | 252.1421 | 199.5703 | 185.5728 | 182.0015 | 216.1421 | 243.1431 | 398.1436 | 422.8569 | 479.5728 | 704.4262 | 377.7139 | 441.8604 | 503.2866 | 718.5683 | 194.4297 | 111.2866 | 116.8579 | 194.4297 | 154.2856 | 144.5728 | 159.2886 | 187.8604 | 134.8579 | 140.7144 | 144.7124 | 144.2841 | 87.4272 | 106 | 110.2856 | 128.5708 | 120.1416 | 125.4287 | 132.5733 | 101.2861 | 318.4302 | 351.2866 | 360.5703 | 363.2861 | 231.5718 | 298.8594 | 308.2866 | 352.7164 | 383.2846 | 380.1426 | 405.1441 | 517.8579
711 | 9.8154 | 210.0137 | 311.1868 | 338.8269 | 327.0797 | 238.3393 | 262.2349 | 250.1071 | 272.353 | 202.1731 | 146.2802 | 172.1868 | 237.2665 | 300.3668 | 236.1868 | 263.8585 | 368.9794 | 279.533 | 236.0069 | 250.9863 | 316.0659 | 240.3462 | 448.0728 | 436.3736 | 313.8338 | 322.0728 | 353.6332 | 394.467 | 475.4258 | 178.0069 | 177.2006 | 204.3736 | 229.7198 | 262.4602 | 228.9341 | 244.0934 | 159.2665 | 374.2734 | 391.1799 | 457.8132 | 484.0069 | 222.9794 | 222.8063 | 249.2006 | 214.7198 | 124.7335 | 89.2665 | 115.1071 | 239.2527 | 304.2596 | 321.6332 | 345.7335 | 234.7541 | 215.647 | 232.5467 | 259.0728 | 212 | 260.7994 | 274.8132 | 319.8132 | 284.6676 | 194.5673 | 209.8201 | 220.7198 | 180.8132 | 261.5261 | 278.0865 | 332.7747 | 370.533 | 220.4464 | 268.1799 | 365.2734 | 396.6332 | 285.6745 | 317.0797 | 340.2143 | 458.3736 | 258.7994 | 256.9794 | 290.1209 | 461.8132 | 180.3736 | 199.7994 | 226.9066 | 265.3736 | 229.5536 | 227.6264 | 261.353 | 205.1731 | 176.467 | 182.9135 | 189.0206 | 220.7473 | 361.761 | 429.2665 | 491.5398 | 675.3805 | 407.2459 | 416.4876 | 499.7335 | 724.8929 | 164.1209 | 108.4602 | 108.3736 | 164.1209 | 145.2802 | 149.2734 | 146.294 | 174.8407 | 140.9066 | 138.0934 | 163.1593 | 156.1662 | 78.3668 | 93.647 | 112.9066 | 128.9066 | 113.5467 | 118.7335 | 130.5536 | 110.2665 | 319.1868 | 325.9341 | 368.353 | 368.6332 | 230.8201 | 305.0934 | 309.0934 | 339.2006 | 391.353 | 373.9135 | 416.5398 | 497.0206
712 | 9.8264 | 203.2611 | 303.0885 | 326.8717 | 326.8274 | 234.2212 | 239 | 273.2965 | 276.2611 | 187.5265 | 164.4292 | 186.9956 | 222.2655 | 281.2655 | 217.7433 | 269.0044 | 354.3539 | 267.2655 | 227.8274 | 271.0398 | 295.9645 | 239.1327 | 439.0044 | 444.562 | 305.5221 | 310.0486 | 332.8318 | 394.1726 | 467.3097 | 173.5221 | 189.6903 | 205.1726 | 234.4779 | 254.8274 | 221.6947 | 242.2611 | 149.438 | 371.7832 | 403.823 | 405.3185 | 486.1726 | 212.0486 | 219.1327 | 243.9558 | 192.8761 | 117.5664 | 101.6062 | 110.2168 | 233.9602 | 283.3982 | 303.7876 | 349.6505 | 254.0355 | 194.7433 | 240.1283 | 246.3097 | 200.9159 | 249.0486 | 277 | 330.8673 | 309.4248 | 184.2611 | 200.1327 | 216.6106 | 194.823 | 270.7389 | 274.8274 | 350.4248 | 397.4292 | 214.3938 | 265.3053 | 381.2567 | 397.2611 | 268.2611 | 319.0442 | 340.7788 | 438.4823 | 238.1814 | 267.4779 | 293.2567 | 453.6548 | 213.6859 | 225.7345 | 228.7389 | 249.177 | 217.8717 | 239.3894 | 254.177 | 211.9558 | 162.8717 | 199.4735 | 189.6062 | 237.8629 | 373.6018 | 412.7876 | 470.3539 | 725.6373 | 395.3539 | 452.9868 | 497.7389 | 711.3097 | 167.2567 | 108.9558 | 124.6903 | 167.2567 | 140.5221 | 151.2168 | 151.7788 | 177.562 | 144.2168 | 139.2168 | 161.9159 | 165.9558 | 112.4691 | 105.9956 | 104.3938 | 127.0442 | 124.823 | 129.3009 | 129.9558 | 96.7433 | 323.6505 | 329.0398 | 364.8717 | 367.7832 | 231.4779 | 285.5265 | 319.0841 | 324.3495 | 411.5177 | 373.4779 | 416.0442 | 496.1283
713 | 9.8374 | 223.2832 | 310.766 | 333.0148 | 335.5172 | 225.2438 | 263.7906 | 282.4975 | 284.2635 | 192.2685 | 165.9902 | 186.2389 | 208.4877 | 278.0049 | 212.7512 | 241.2143 | 331.7266 | 264.7562 | 212.4779 | 279.7562 | 296.2685 | 236 | 417.2242 | 441.4877 | 283.2143 | 307.5074 | 355.0492 | 404.2635 | 501.3177 | 185.0197 | 197.7512 | 192.7266 | 222.9803 | 237.2242 | 234.2685 | 248.7611 | 137.4877 | 352.2192 | 423.0246 | 429.5714 | 515.7955 | 191.4779 | 209.2389 | 256.2685 | 218.5566 | 133.2783 | 108.7562 | 101.9852 | 233.5074 | 308.5566 | 309.7709 | 361.766 | 258.9902 | 225.0591 | 236.2389 | 228.4828 | 204.5123 | 230.7315 | 256.7168 | 309.9606 | 324.5025 | 187.7562 | 190.2389 | 234.2832 | 204.2586 | 253.7217 | 286.5221 | 358.7414 | 370.4434 | 217.7611 | 248.9754 | 368.9705 | 405.2635 | 283.7758 | 312.7414 | 332.4779 | 493.0984 | 221.2389 | 240.2045 | 294.4926 | 444.7414 | 184.9311 | 220.7315 | 221.4877 | 246.5025 | 214.2488 | 237.4926 | 247.7463 | 198.7266 | 171.2685 | 200.7414 | 183.234 | 263.5271 | 384.5025 | 426.2832 | 445.4729 | 691.9114 | 371.9754 | 446.9606 | 525.0443 | 718.2734 | 169.9951 | 110.7512 | 110.2143 | 169.9951 | 139.2488 | 146.7414 | 154.7463 | 170.7315 | 148.0049 | 171.5517 | 148.234 | 161.7414 | 103.2143 | 115.0049 | 107.0098 | 129.7562 | 107.2143 | 116.7217 | 122.7365 | 98.5123 | 332.0098 | 326.4877 | 391.298 | 370.7562 | 264.3029 | 307.5468 | 303.468 | 322.7512 | 378.6872 | 360.4779 | 376.6872 | 490.7365
714 | 9.8484 | 218.5472 | 330.034 | 314.6214 | 345.8717 | 219.4189 | 250.6214 | 265.966 | 252.6415 | 195.2365 | 189.1962 | 177.1283 | 238.8314 | 279.7635 | 224.4528 | 259.1962 | 319.3648 | 230.1145 | 220.2163 | 275.6553 | 277.0944 | 243.6352 | 405.8918 | 484.0478 | 271.4189 | 337.2503 | 360.4189 | 369.5874 | 479.1145 | 202.7434 | 198 | 194.1082 | 220.527 | 241.9258 | 235.7094 | 244.1283 | 141.1082 | 355.1623 | 447.088 | 444.7635 | 463.9183 | 209.1421 | 216.1623 | 264.5811 | 213.2025 | 131.3648 | 119.1623 | 129.5409 | 237.2906 | 303.966 | 311.7094 | 333.6956 | 285.2503 | 211.0403 | 221.7837 | 224.2365 | 194.7837 | 260.5949 | 268.3245 | 329.4868 | 281.4793 | 192.8176 | 197.1623 | 221.1484 | 192.2566 | 236.5472 | 270.3849 | 347.8377 | 353.6553 | 214.6553 | 243.2365 | 346.2025 | 398.8377 | 285.1824 | 307.1824 | 318.0742 | 461.8441 | 238.8516 | 291.3182 | 299.1082 | 436.8918 | 184.7434 | 184.1686 | 209.8377 | 238.6013 | 234.6151 | 255.088 | 245.473 | 186.3648 | 155.1484 | 184.4931 | 182.2906 | 252.2025 | 352.9321 | 452.088 | 481.9937 | 717.5207 | 364.7094 | 410.7497 | 496.5874 | 739.7975 | 165.9459 | 94.966 | 103 | 165.9459 | 130.6013 | 157.9799 | 147.8918 | 166.2365 | 139.0742 | 161.3849 | 129.7296 | 165.3447 | 112.7975 | 106.0742 | 102.8918 | 144.7434 | 127.4868 | 111.7635 | 129.1623 | 112.4528 | 356.9056 | 321.7094 | 384.966 | 326.1887 | 221.5535 | 310.1283 | 309.2163 | 334.4528 | 386.6151 | 359.8176 | 398.3585 | 499.4528
715 | 9.8594 | 225.8614 | 346.6372 | 325.8437 | 327.0531 | 233.5163 | 265.7228 | 253.2419 | 255.9646 | 195 | 182.2596 | 194.3953 | 254.4307 | 293.1888 | 211.7079 | 270.8791 | 327.0856 | 227.5339 | 231.7581 | 239.0884 | 306.4632 | 244.4484 | 428.0502 | 466.1917 | 291.7228 | 303.3303 | 342.7079 | 378.9469 | 511.4455 | 200.0177 | 173.9498 | 176.6047 | 206.2596 | 224.0531 | 238.8791 | 267.6018 | 143.7758 | 384.3776 | 426.6224 | 410.8642 | 435.6902 | 212.8967 | 232.1888 | 243.5014 | 211.8791 | 134.4307 | 112.6902 | 122.7079 | 209.2949 | 312.1888 | 317.2065 | 342.6195 | 227.2801 | 211.5339 | 213.3451 | 238.7404 | 189.4484 | 262.4661 | 297.2744 | 327.1386 | 283.5163 | 182.3628 | 185.2596 | 218.8791 | 193.4307 | 238.4307 | 300.4632 | 322.5014 | 392.3423 | 232.3953 | 218.174 | 374.3599 | 371.9498 | 284.7758 | 306.7758 | 321.2065 | 463.8437 | 241.8967 | 260.348 | 297.1209 | 456.7228 | 208.3953 | 184.0856 | 208.5516 | 263.1534 | 227.8112 | 247.8112 | 248.1033 | 189.4307 | 160.6372 | 180 | 193.8614 | 254.7581 | 384.1181 | 440.1563 | 455.4336 | 704.6224 | 375.6372 | 404.3274 | 450.8642 | 703.1061 | 184.3953 | 107.0679 | 98.3451 | 184.3953 | 124.8967 | 147.2596 | 163.8437 | 176.0856 | 148.413 | 148.0177 | 162.239 | 168.5516 | 102.4837 | 116.9646 | 91.6902 | 131.9321 | 115.8289 | 121.3098 | 121.9144 | 122.9823 | 320.5545 | 337.2921 | 382.3274 | 316.6549 | 260.0827 | 294.0354 | 299.8112 | 348.0856 | 377.4837 | 346.2596 | 364.7787 | 512.3098
716 | 9.8704 | 220.3309 | 316.8997 | 313.6617 | 359.6171 | 220.6617 | 272 | 262.8216 | 260.7881 | 198.9405 | 156.7212 | 165.3234 | 256 | 297.7881 | 208.5762 | 255.4498 | 341.0334 | 241.0334 | 221.3903 | 252.6431 | 294.145 | 240.0595 | 423.1784 | 445.1784 | 281.4498 | 313.855 | 361.6431 | 372.3903 | 505.0855 | 181.4498 | 188.2788 | 193.0669 | 205.1524 | 222.7286 | 220.2974 | 238.7472 | 166.0669 | 353.3829 | 407.1784 | 421.4907 | 492.8959 | 220.6691 | 206.8402 | 264.5836 | 224.8216 | 137.5762 | 100.5428 | 133.7621 | 248.2862 | 286.0521 | 345.7955 | 355.881 | 238.736 | 201.3903 | 233.2788 | 229.6022 | 200.8216 | 224.5353 | 292.9666 | 333.4572 | 287.2119 | 193.1859 | 167.6022 | 249.1598 | 208.3978 | 283.3457 | 297.6022 | 326.2453 | 387.4498 | 223.8141 | 237.0074 | 367.4498 | 380.7621 | 283.4238 | 337.736 | 332.4572 | 449.2974 | 265.4312 | 264.9145 | 276.2974 | 457.4833 | 216.3643 | 180.6952 | 189.2974 | 254.4498 | 227.9405 | 212.4759 | 277.3717 | 179.1784 | 156.9071 | 166.6022 | 196.2119 | 268.8216 | 369.9926 | 408.9926 | 444 | 706.4572 | 361.6617 | 397.119 | 461.4907 | 695.7286 | 186.0595 | 101.7547 | 106.4572 | 186.0595 | 127.9405 | 147.7286 | 154.8141 | 164.026 | 149.6357 | 127.0855 | 158.026 | 166.6357 | 94.8476 | 105.2379 | 106.3383 | 132.5167 | 108.4238 | 132.6691 | 126.881 | 125 | 356.0743 | 310.4759 | 371.1784 | 327.4572 | 241.3234 | 265.5688 | 319.6431 | 327.3569 | 390.3383 | 357.7621 | 392.1933 | 477.171
717 | 9.8814 | 228.4051 | 316.8043 | 341.8162 | 324.3775 | 217.6008 | 243.1858 | 259.5968 | 273.8063 | 195.1976 | 159.0039 | 183.2134 | 225.585 | 281.1917 | 212.2016 | 259.8043 | 357.6067 | 247.2016 | 181.9822 | 233.3874 | 291.2016 | 238.1996 | 422.4012 | 445.2016 | 289.0059 | 285.583 | 330.3815 | 361.7964 | 459.9803 | 177 | 175.5909 | 182.1917 | 228.411 | 237.6067 | 235.8103 | 271.4209 | 171.1996 | 342.1996 | 395.1957 | 399.7866 | 485.7886 | 206.9921 | 193.3972 | 240.7846 | 220.7964 | 123.5929 | 95.5988 | 113.9882 | 235.3874 | 286.8043 | 324.1858 | 341.9921 | 255.0039 | 224.413 | 214.1878 | 228.4012 | 198.3972 | 229.4071 | 245.9783 | 335.1996 | 281.3972 | 196.1996 | 192.8142 | 232.9882 | 201.5949 | 273.3893 | 315.6107 | 343.4071 | 387.002 | 223.2016 | 258.4071 | 392.6146 | 412.2134 | 279.7984 | 312.3834 | 323.7945 | 439.998 | 232.7807 | 235.583 | 268.5988 | 434.3893 | 181.7826 | 183.002 | 216.8162 | 248.3992 | 215.3933 | 258.4268 | 243.3795 | 171.1976 | 167.8063 | 165.4012 | 213.6087 | 243.9862 | 374.2055 | 432.415 | 456.0059 | 715.4032 | 373.8083 | 393.3992 | 455.7945 | 712.2075 | 184.1996 | 94.1976 | 103.3972 | 184.1996 | 162.6166 | 136.1937 | 163.8063 | 166.4032 | 150.6008 | 151.6146 | 162.8043 | 161.1976 | 87.5968 | 111.4051 | 106.998 | 140.4032 | 140.8162 | 126.9961 | 138.6047 | 113.7945 | 332.1818 | 322.0099 | 378.4051 | 340.4051 | 208.7886 | 291.0158 | 317.9961 | 340.2095 | 355.7807 | 337.9882 | 411.0039 | 515.8241
718 | 9.8924 | 229.3746 | 295.4321 | 328.0575 | 334.9425 | 216.3746 | 232.7493 | 264.5015 | 281.0634 | 183.4351 | 176.441 | 177.8097 | 243.1933 | 299.7552 | 226.003 | 257.9366 | 348.8097 | 222.8067 | 204.6947 | 230.2507 | 290.3746 | 221.7463 | 427.0634 | 467.9425 | 253.8038 | 305.444 | 347.0059 | 377.8791 | 513.3895 | 173.7493 | 195.3806 | 175.5619 | 224.2478 | 235.3112 | 230.4351 | 247.8672 | 147.4321 | 368.8186 | 379.997 | 429.5709 | 506.0059 | 198.9366 | 205.003 | 232.1873 | 223.8761 | 115.9366 | 99.0634 | 144.5709 | 232.2507 | 305.2537 | 312.9366 | 338.8127 | 237.4955 | 237.5015 | 205.5619 | 257.444 | 205.9395 | 241.1269 | 239.8761 | 316.3082 | 267.8097 | 204.1269 | 208.1269 | 222.1239 | 182.7463 | 253.3717 | 307.997 | 353.5015 | 364.4321 | 230.5015 | 251.4351 | 380.4955 | 402.7463 | 300.9425 | 298.3112 | 316.1239 | 467.444 | 249.8186 | 271.6977 | 252.559 | 430.6254 | 190.0664 | 188.8761 | 194.9306 | 235.8097 | 218.5015 | 248.4321 | 273.6343 | 193.5679 | 161.2478 | 174.9395 | 190.3687 | 263.8186 | 371.3112 | 400.9911 | 468.7522 | 675.553 | 373.1239 | 423.8821 | 459.5015 | 706.2478 | 179.9366 | 92.1873 | 97.9366 | 179.9366 | 144.9941 | 129.7493 | 165.3746 | 179.3776 | 137.997 | 163.0634 | 152.8097 | 148.6224 | 107.1298 | 103.4351 | 118.1903 | 127.3717 | 127.8702 | 124.1873 | 138.5619 | 115.8761 | 319.7493 | 323.7493 | 386.6888 | 338.1239 | 209.5015 | 314.441 | 324.1269 | 310.0545 | 394.7611 | 340.9395 | 352.0486 | 524.7493
719 | 9.9034 | 227.35 | 317.2242 | 372.4986 | 302.876 | 215.175 | 261.6992 | 268.4749 | 250.6509 | 196.6746 | 191.5497 | 171.7001 | 227.5506 | 290.9754 | 223.2252 | 262.7748 | 344.35 | 231.8496 | 206.2252 | 240.0747 | 304.0246 | 250.1741 | 418.9253 | 437.526 | 282.9489 | 314.4749 | 327.4257 | 396.8496 | 511.5005 | 190.3245 | 184.5005 | 189.0246 | 213.7502 | 223.2753 | 218.9253 | 225.9754 | 166.7493 | 353.5506 | 391.8496 | 402.526 | 474.876 | 224.3993 | 198.1003 | 232.825 | 205.2006 | 135.6244 | 97.5251 | 126.6007 | 232.175 | 317.2498 | 304.5752 | 343.1249 | 242.8997 | 194.4513 | 213.2498 | 215.3264 | 193.9754 | 257.0246 | 262.4494 | 293.8505 | 286.4494 | 188.6755 | 207.5251 | 210.2753 | 187.2498 | 255.7748 | 310.7748 | 338.5005 | 347.4503 | 210.5506 | 249.825 | 384.2498 | 393.2252 | 280.4257 | 315.1495 | 300.9754 | 431.9262 | 227.9508 | 253.1258 | 299.3236 | 431.825 | 183.2753 | 191.65 | 202.8496 | 241.2498 | 229.0747 | 238.8751 | 234.3264 | 194.0501 | 160.65 | 175.35 | 167.5005 | 260.1003 | 356.8004 | 414.2744 | 443.7758 | 689.9243 | 374.4749 | 420.2753 | 453.5752 | 698.2252 | 175.7001 | 112.6244 | 90.4002 | 175.7001 | 138.175 | 167.7739 | 163.35 | 168.8004 | 145.7247 | 140.0757 | 147.5251 | 160.8496 | 90.5506 | 94.4002 | 97.0757 | 138.0246 | 126.2999 | 133.8997 | 147.8997 | 114.5251 | 312.4002 | 339.4995 | 371.5005 | 326.2753 | 215.9499 | 289.1258 | 335.8997 | 303.65 | 366.4011 | 360.3245 | 369.3491 | 504.2006
720 | 9.9144 | 223.6509 | 306.2546 | 346.9975 | 328.4907 | 209.9764 | 273.0236 | 278.2099 | 294.2361 | 216.7454 | 178.9292 | 163.4646 | 243.0944 | 266.231 | 227.8609 | 277.3963 | 348.1863 | 239.1863 | 199.9764 | 240.3255 | 301.1391 | 216.8111 | 409.4646 | 429.1627 | 277.6037 | 310.8137 | 364.7007 | 363.9975 | 471.1602 | 177.2546 | 194.3963 | 202.0472 | 233.769 | 215.9764 | 210.3019 | 251.4671 | 161.9528 | 341.4646 | 402.5354 | 428.4907 | 484.5826 | 184.7875 | 194.3255 | 266.4907 | 210.2099 | 141.6745 | 116.2572 | 138.5826 | 235.3491 | 289.6956 | 289.6037 | 346.5118 | 254.2099 | 200.0708 | 225.0472 | 211.6981 | 206.0708 | 244.0919 | 261.1391 | 338.5615 | 298.5354 | 181.6509 | 193.6037 | 213.8609 | 188.1627 | 261.1863 | 272.6484 | 350.9081 | 354.2099 | 221.0708 | 267.5826 | 363.3938 | 393.6745 | 285.8845 | 290.5329 | 334.8398 | 454.8162 | 257.1652 | 240.3019 | 270.6484 | 432.8373 | 185.1863 | 186.9764 | 199.3019 | 261.4199 | 225.1391 | 222.0919 | 254.1416 | 177.9292 | 162.6745 | 176.6745 | 183.2572 | 257.1627 | 365.7217 | 408.9528 | 412.0447 | 697.5118 | 377.5118 | 398.7428 | 443.6273 | 676.9056 | 173.3255 | 91.0447 | 106.5826 | 173.3255 | 122.9292 | 140.8348 | 140.3938 | 170.1863 | 137.9528 | 131.6509 | 154.5354 | 154.7901 | 105.2572 | 103.8845 | 98.6981 | 119.231 | 123.6509 | 129.3019 | 134.9292 | 115.6745 | 356.2125 | 327.0919 | 313.5776 | 340.7454 | 231.2336 | 301.4199 | 318.7428 | 309.8609 | 391.4907 | 340.5565 | 391.9081 | 486.6037
721 | 9.9254 | 226.3982 | 287.7079 | 334.9026 | 312.0619 | 208.1504 | 225.5751 | 255.3628 | 284.4601 | 197.0619 | 172.6018 | 163.6991 | 220.6637 | 281.5399 | 217.9557 | 246.8672 | 354.0974 | 228.1062 | 216.8407 | 270.5841 | 283.0088 | 214.9469 | 438.5841 | 445.1416 | 275 | 275.1681 | 337.3185 | 372.2921 | 492.0354 | 181.646 | 180.8584 | 195.5044 | 232.9026 | 223.4956 | 212.3982 | 237.4601 | 146.4071 | 368.0354 | 398.0531 | 401.0176 | 482.9026 | 213.3806 | 205.8938 | 250.0619 | 225.5929 | 114.8141 | 115.7522 | 121.6106 | 219.0088 | 311.1859 | 315.885 | 335.1062 | 241.5575 | 192.8053 | 232.9469 | 204.5044 | 185.2124 | 244.3982 | 283.7876 | 299.5751 | 282.1593 | 198.8407 | 189.3009 | 220.9469 | 180.354 | 280.6903 | 285.3894 | 357.3982 | 356 | 197.6637 | 242.115 | 342.8584 | 379.5575 | 294.7965 | 285 | 317.3628 | 453.354 | 224.9203 | 229.6549 | 291.0354 | 433.8496 | 197.0443 | 213.1859 | 214.1416 | 244.6106 | 236.7434 | 236.8407 | 243.0088 | 171.6018 | 161.3009 | 207.5841 | 189.5487 | 244.2566 | 377.3451 | 398.5044 | 440.1328 | 667.4159 | 361.0088 | 417.0885 | 461.5399 | 664.5044 | 204.4337 | 101.2921 | 115.9469 | 204.4337 | 128.4956 | 154.3894 | 153.8407 | 174.3982 | 149.5929 | 124.2035 | 144.9557 | 143.6549 | 117.4956 | 118.7434 | 109.3451 | 122.646 | 112.8053 | 127.1504 | 138.7965 | 106.6549 | 333.5663 | 351.1859 | 343.7788 | 346.5487 | 262.885 | 312.646 | 316.6991 | 309.3009 | 381.8584 | 370.8319 | 362.7168 | 466.1593
722 | 9.9364 | 220.1052 | 315.1647 | 334 | 310.5855 | 208.8618 | 242.8555 | 246.6908 | 299.9607 | 182.6578 | 162.5197 | 164.8618 | 225.4803 | 274.6578 | 206.5197 | 256.5133 | 377.4081 | 242.3751 | 202.763 | 239.8023 | 290.3422 | 236.6844 | 418.1445 | 449.7237 | 287.9277 | 331.0532 | 323.2434 | 351.7301 | 461.6642 | 169.2104 | 207.3029 | 180.2104 | 225.1052 | 219.8289 | 206.9671 | 234.8618 | 166.4081 | 336.8023 | 407.3422 | 408.7896 | 468.2104 | 224.3092 | 204.5526 | 227.9012 | 212.4867 | 122.9277 | 115 | 115.4145 | 228.0659 | 282.3879 | 276.1838 | 321.796 | 288.1254 | 204.7896 | 232.2763 | 207.3092 | 179.2763 | 233.796 | 270.763 | 335.8162 | 273.8289 | 193.3815 | 201.9277 | 228.0329 | 191.0659 | 249.526 | 264.8682 | 344.2104 | 350.8289 | 221.441 | 230.9671 | 354.6514 | 395.0988 | 247.7364 | 289.3092 | 347.474 | 489.0595 | 242.1318 | 266.7832 | 285.6578 | 427.9671 | 184.3486 | 193.0064 | 210.9671 | 262.5462 | 243.3092 | 260.6844 | 243.4474 | 182.204 | 176.5133 | 204.3815 | 203.7896 | 211.8353 | 395.3751 | 429.7503 | 421.8682 | 729.2243 | 383.8555 | 403.763 | 433.1116 | 677.6514 | 191.0393 | 105.7237 | 102.3486 | 191.0393 | 136.8948 | 149.3815 | 144.0723 | 181.0329 | 142.5197 | 147.9936 | 141.2763 | 137.6908 | 118.1382 | 106.3486 | 101.5197 | 126.5855 | 135.9936 | 122.6908 | 127.0723 | 110.1711 | 327.1382 | 337.9012 | 356.1711 | 362.5133 | 230.9405 | 309.6908 | 295.4538 | 321.9277 | 385.8948 | 337.3549 | 375.0988 | 491.441
723 | 9.9474 | 226.8673 | 323.4966 | 324.3844 | 304.0069 | 210.7483 | 261.8604 | 257.3638 | 282.0206 | 198.4828 | 169.7414 | 168.4966 | 234.8673 | 290.4828 | 209.3707 | 222.2861 | 371.3844 | 254.6156 | 202.6224 | 253.2311 | 304.238 | 210.2792 | 430.6087 | 468.357 | 272.5172 | 301.6637 | 341.2311 | 355.8673 | 463.8673 | 173.119 | 191.8948 | 191.9863 | 223.1259 | 227.7414 | 194.6362 | 241.9931 | 168.2517 | 354.6018 | 378.4051 | 410.1259 | 492.2242 | 233.7414 | 208.3707 | 239.8604 | 239.7208 | 131.9931 | 91.3982 | 109.7552 | 222.1327 | 290.9863 | 292.6018 | 335.7345 | 266.2792 | 226.2311 | 218.8879 | 206.2517 | 213.0915 | 233.7483 | 273.2448 | 289.6774 | 298.3501 | 189.3776 | 234.5949 | 252.6018 | 165.0275 | 254.4897 | 273.238 | 335.881 | 364.8604 | 239.1121 | 256.2242 | 349.1327 | 389.2586 | 271.4691 | 283.881 | 319.7826 | 467.9016 | 237.2586 | 274.7483 | 290.9931 | 417.3844 | 199.4828 | 189.8741 | 228.357 | 269.4966 | 233.5103 | 276.238 | 243.1259 | 181.3776 | 190.3638 | 176.7758 | 180.6499 | 255.9519 | 352.5447 | 406.1533 | 409.2586 | 720.7689 | 349.5378 | 392.2586 | 435.8673 | 682.6224 | 161.7758 | 90.2655 | 116.6087 | 161.7758 | 137.1259 | 151.4966 | 135.0069 | 185.4966 | 131.3844 | 140.6362 | 152.3638 | 158.8535 | 95.2723 | 97.0069 | 109.6156 | 123.5034 | 121.643 | 117.6293 | 124.1259 | 117.9931 | 331.3707 | 323.6362 | 341.2655 | 345.7689 | 262.5881 | 300.2586 | 306.8604 | 331.8673 | 374.762 | 339.7414 | 387.6156 | 485.5103
724 | 9.9584 | 228 | 308.0442 | 323.8864 | 324.2744 | 214.5457 | 247.1578 | 251.0221 | 269.2492 | 200.1136 | 153.2714 | 177.8643 | 217.3849 | 296.5457 | 217.9779 | 241.8201 | 362.9085 | 249.795 | 216.2965 | 275.5015 | 285.6121 | 242.3437 | 455.1608 | 430.2241 | 287.7286 | 302.205 | 328.9307 | 354.3407 | 481.8422 | 180.205 | 175.7035 | 170.0663 | 224.7729 | 230.7729 | 209.8422 | 264.2744 | 174.205 | 356.2271 | 386.41 | 416.205 | 526.1387 | 180.9276 | 190.3849 | 222.4985 | 206.7699 | 137.4322 | 86.2271 | 117.8643 | 240.5015 | 247.792 | 296 | 356.6151 | 292.1387 | 216.59 | 198.3849 | 213.0915 | 193.1799 | 226.9085 | 276.6593 | 315.6844 | 305.5457 | 193.4322 | 208.8613 | 212.5648 | 198.2301 | 256.8864 | 266.1357 | 329.6814 | 374.0915 | 209.0884 | 221.8834 | 355.9779 | 408.3879 | 291.0693 | 287.4322 | 343.3658 | 463.1136 | 248.41 | 274.1136 | 252.1106 | 451.4573 | 193.1357 | 217.4794 | 199.0884 | 262.0221 | 223.1357 | 258.4985 | 256.2965 | 191.6372 | 169.8392 | 173 | 205.3658 | 250.59 | 338.0221 | 409.9779 | 426.6151 | 710.9085 | 382.1166 | 382.1357 | 420.1578 | 703.3879 | 167.7508 | 104.8422 | 102.1578 | 167.7508 | 131.6814 | 128.9528 | 146.41 | 188.6593 | 145.9558 | 146.9779 | 162.8643 | 161.1136 | 103.5236 | 97.7729 | 100.3628 | 131.8643 | 116.3407 | 114.3407 | 125.7729 | 111.9085 | 334.6593 | 326.4322 | 336.3407 | 358.0693 | 270.6593 | 303.4322 | 316.0915 | 353.3879 | 390.7286 | 349.8643 | 365.0663 | 453.8613
725 | 9.9694 | 219.9115 | 312.2911 | 332.9872 | 338.6834 | 221.2911 | 266.5694 | 258.9872 | 266.2026 | 199.1013 | 171.6706 | 159.2281 | 226.6834 | 278.1268 | 228.886 | 221.6332 | 315.2664 | 205.8613 | 226.0885 | 245.646 | 290.1898 | 258.6834 | 440.9243 | 453.7591 | 293.5949 | 311.9872 | 330.5949 | 345.9115 | 468.7217 | 175.6077 | 160.5192 | 190.3668 | 226.7974 | 193.2536 | 180.5447 | 232.8485 | 137.2536 | 353.3038 | 369.1268 | 415.2026 | 482.3677 | 191.9745 | 204.177 | 243.3668 | 202 | 119.1268 | 102.177 | 137.8732 | 241.2026 | 275.2528 | 290.6077 | 325.7472 | 233.0894 | 202.4179 | 216.6706 | 232.8732 | 218.7591 | 244.8732 | 274.3038 | 315.5064 | 251.1779 | 179.6204 | 198.7089 | 219.5821 | 210.1898 | 268.6834 | 310.8349 | 315.5192 | 354.3294 | 196.0128 | 278.1132 | 316.5575 | 400.2153 | 275.9243 | 299.6834 | 329.0255 | 480.0757 | 257.1898 | 265.9115 | 278.4553 | 424.5447 | 201.886 | 188.646 | 227.354 | 239.4306 | 222 | 234.4306 | 233.7345 | 194.7974 | 161.6077 | 181.9872 | 173.0511 | 258.886 | 367.5566 | 405.6077 | 456.8604 | 718.0885 | 369.9243 | 409.7591 | 417.1013 | 697.9115 | 183.3796 | 95.3166 | 101.7974 | 183.3796 | 143.5821 | 146.6706 | 144.4051 | 164.7345 | 146.2026 | 148.8987 | 178.3796 | 133.1396 | 111.2911 | 105.1898 | 90.9115 | 134.7974 | 126.7847 | 127.4808 | 126 | 93.9243 | 311.6332 | 346.7719 | 326.114 | 341.1268 | 228.76 | 271.646 | 307.114 | 308.3677 | 353.4562 | 342.9115 | 381.7719 | 501.2272
726 | 9.9804 | 219.911 | 293.8687 | 317.6017 | 350.0212 | 217.4449 | 263.5339 | 248.1568 | 276.0212 | 213.5762 | 159.4238 | 187.9745 | 212.5128 | 275.089 | 214.5128 | 263.6397 | 326.3983 | 259.3049 | 233.3771 | 234.7119 | 286.4449 | 238.1357 | 412.5806 | 477.0423 | 305.8432 | 307.5339 | 331 | 376.8855 | 422.3603 | 176.822 | 177.2203 | 184.8009 | 230.6441 | 225.4406 | 223.4617 | 263.6185 | 160.3304 | 353.911 | 354.2458 | 434.1313 | 481.5551 | 198.5551 | 203.267 | 205.0044 | 185.6017 | 147.0635 | 111.2881 | 125.4238 | 210.9365 | 298.1313 | 300.9322 | 338.3983 | 247.8643 | 201 | 196.2246 | 227.7119 | 201.0467 | 216.0255 | 244.8475 | 338.6864 | 285.9956 | 192.5762 | 230.7965 | 236.4872 | 202.8009 | 218.9832 | 234.0087 | 325.8432 | 352.911 | 246.0168 | 222.14 | 342.0635 | 372.5806 | 299.5084 | 287.3348 | 306.0467 | 424.6061 | 294.4406 | 258.6229 | 266.5128 | 432.8432 | 196.6229 | 215.9745 | 220.0678 | 245.1991 | 223.822 | 239.2881 | 227.3559 | 171.3136 | 159.178 | 163.8687 | 215.4617 | 249.0678 | 368.267 | 412.2881 | 404.4281 | 729.9322 | 374.3771 | 366.5383 | 435.2203 | 697 | 183.178 | 113.1313 | 80.1357 | 183.178 | 144.089 | 130.7797 | 173.1525 | 185.6864 | 130.5128 | 140.8009 | 151.7586 | 158.2414 | 92.8687 | 95.9788 | 101.8432 | 137.733 | 116.1568 | 118.0678 | 145.1313 | 102.0212 | 315.3771 | 325.3136 | 333.1991 | 332.6229 | 226.733 | 281.6652 | 345.1736 | 323.0423 | 379.0635 | 325.6017 | 362.1357 | 479.6696
727 | 9.9914 | 228.3097 | 308.6195 | 330.7729 | 323.3008 | 200.3805 | 276.8496 | 275.6225 | 269.6136 | 160.525 | 161.6932 | 160.5309 | 213.7699 | 250.9941 | 200.0737 | 239.3775 | 352.9292 | 243.764 | 224.7669 | 233.0767 | 294.3097 | 211.9941 | 433.0826 | 460.5339 | 319.9263 | 331.0059 | 302.3775 | 367.0737 | 499.2509 | 175.1534 | 192.8496 | 194.1564 | 220.8436 | 209.6106 | 207.6873 | 256.8436 | 160.2301 | 340.1504 | 372.3894 | 411.9941 | 470.9203 | 210.0797 | 215.9263 | 228.6992 | 215.3924 | 131.5339 | 87.9941 | 115.6903 | 202.4602 | 302.7699 | 298.3068 | 318.764 | 273.0826 | 202.8466 | 197.6932 | 198.3775 | 196.2301 | 223.1564 | 264.1593 | 333.6136 | 267.8407 | 181.0737 | 212.764 | 226.9203 | 189.0737 | 244.4691 | 299.8644 | 342.6962 | 343.7669 | 245.4602 | 197.5339 | 331.1504 | 377.3864 | 289.997 | 301.6962 | 304 | 467.9352 | 293.3835 | 230.3008 | 259.4602 | 425.6903 | 196.9233 | 182.0678 | 214.3835 | 228.4572 | 233.2331 | 227.997 | 275.0119 | 185.6195 | 152.5369 | 185.0826 | 206.1504 | 235.0737 | 369.8466 | 412.0767 | 381.4572 | 665.4453 | 372.2301 | 378.6195 | 411.1475 | 729.3157 | 174.6903 | 89.1475 | 94.6195 | 174.6903 | 136.6136 | 141.003 | 149.2242 | 170.4572 | 129 | 146.4631 | 181.3157 | 148.997 | 111.3127 | 108.8496 | 99.3068 | 148.1564 | 116.8466 | 142.8525 | 132.2271 | 99.3068 | 330.7729 | 331.3097 | 348.7729 | 354.1593 | 235.3097 | 278.3835 | 323.1475 | 341.6195 | 370.9203 | 368.3186 | 335.0708 | 516.7021
728 | 10.0024 | 196.2541 | 294.0949 | 333.8712 | 322.8712 | 207.4204 | 242.4473 | 262.0949 | 260.5796 | 186.8747 | 160.1288 | 173.9051 | 222.4204 | 255.5492 | 203.678 | 237.9356 | 357.8068 | 262.5831 | 202.4813 | 248.9051 | 287.5152 | 200.644 | 410.6745 | 451.5152 | 286.3829 | 317.0949 | 321.5187 | 370.678 | 458.2846 | 179.678 | 194.9356 | 198.7424 | 245.2611 | 222.0339 | 238.7459 | 276.5831 | 157.1932 | 332.4508 | 401.1323 | 417.4848 | 470 | 196.0305 | 224.4848 | 227.2576 | 224.5492 | 118.7728 | 86.9356 | 129.9695 | 228.1967 | 296.4508 | 291.4508 | 331.9695 | 276.8712 | 183.3525 | 241.973 | 220.3255 | 195.0644 | 233.356 | 240.7389 | 249.7319 | 251.0305 | 197.7763 | 196.9661 | 238.1628 | 188 | 198.349 | 280.7389 | 303.7694 | 360.7763 | 246.8712 | 243.7154 | 375.8442 | 377.0644 | 297.4204 | 276.8033 | 319.9051 | 457.0305 | 243.4134 | 263.5527 | 292.6815 | 397.8677 | 188.5796 | 191.1628 | 186.8677 | 242.9051 | 204.9965 | 229.8068 | 245.3185 | 205.7119 | 184.7459 | 190.7424 | 211.5492 | 224.644 | 374.678 | 400.7728 | 405.2611 | 657.1932 | 354.2237 | 372.5152 | 425.8407 | 649.6675 | 165.5796 | 97.2916 | 110.9695 | 165.5796 | 140.678 | 143.8712 | 149.8068 | 181.1628 | 144.9051 | 143.2576 | 149.3829 | 162.9695 | 97.0949 | 96.9017 | 112.0983 | 124.6745 | 132.9051 | 141.2576 | 141.2916 | 100.8712 | 331.0644 | 310.4813 | 346.2576 | 346.644 | 241.6136 | 294.8407 | 334.0983 | 339.2576 | 345.6745 | 320.5422 | 384.4578 | 473.2202
729 | 10.0134 | 218.6451 | 290.1563 | 336.8437 | 315.4169 | 197.5732 | 269.3845 | 248.6775 | 268.531 | 195.6352 | 150.5211 | 155.0944 | 211.6254 | 283.4887 | 268.4563 | 259.8014 | 356.1042 | 236.5113 | 221.8535 | 225.3549 | 304.062 | 257.8211 | 445.0197 | 435.8338 | 315.2803 | 300.8338 | 308.7817 | 357.7296 | 450.2606 | 187.5831 | 164.6676 | 224.593 | 225.1986 | 216.3648 | 215.407 | 240.0845 | 139.938 | 365.1761 | 362.2408 | 404.7296 | 508.8634 | 192.1563 | 219.3127 | 220.3648 | 219.3127 | 106.6254 | 94.5831 | 104.4592 | 219.5732 | 279.8859 | 306.1662 | 347.2183 | 258.0423 | 197.1662 | 217.5113 | 215.3648 | 196.8958 | 218.8338 | 239.9479 | 332.1535 | 274.6451 | 195.2085 | 229.1761 | 221.938 | 176.6254 | 255.6648 | 264.7817 | 333.2282 | 334.5113 | 196.762 | 228.0423 | 337.293 | 371.3127 | 299.8958 | 303.4366 | 344.6972 | 438.938 | 223.8859 | 238.5113 | 276.0423 | 422.5408 | 188 | 214.7493 | 230.4986 | 262.9577 | 213.4268 | 235.6873 | 247.7394 | 185.1986 | 173.7296 | 180.5732 | 191.1465 | 228.7394 | 362.6775 | 401.8958 | 425.9577 | 702.4986 | 386.1761 | 350.1986 | 404.2507 | 675.2803 | 206.707 | 104.6352 | 91.1465 | 206.707 | 127.7296 | 148.7394 | 157.5831 | 173.469 | 136.5211 | 157.2183 | 170.6972 | 164 | 95.0521 | 110.2183 | 85.5113 | 138.1662 | 137.7915 | 129.6254 | 118.3028 | 89.6254 | 322.469 | 311.8437 | 343.1563 | 343.1563 | 207.8761 | 296.9479 | 314.1465 | 318.1465 | 362.9577 | 330.2704 | 362.407 | 488.9577
730 | 10.0244 | 215.1991 | 284.2389 | 316.8363 | 314.0398 | 202.7611 | 266.1991 | 247.0398 | 275.7212 | 177.7566 | 166.323 | 169.3629 | 214.8407 | 259.0752 | 199.9866 | 238.9159 | 327.1947 | 231.1593 | 218.1991 | 224.9602 | 273.3141 | 215.8716 | 434.5177 | 459.0045 | 316.0398 | 297.1195 | 324.323 | 399.2478 | 450.9602 | 192.8009 | 187.0045 | 231.7611 | 218.2389 | 200.6371 | 192.8761 | 247.6018 | 150.5221 | 305.5486 | 366.7212 | 412.6416 | 465.8716 | 204.4823 | 188.2743 | 243.0443 | 205.5575 | 127.1239 | 115.1637 | 120.2832 | 219 | 269.3982 | 311.8009 | 304.792 | 246.438 | 161.5132 | 223.6814 | 236.1239 | 194.1195 | 228.562 | 263.0443 | 325.4779 | 264.4779 | 185.3982 | 206.9955 | 231.562 | 183.6814 | 231.1548 | 269.7611 | 364.7655 | 368.5266 | 213.2036 | 241.4027 | 340.7611 | 388.2832 | 295.1991 | 305.9602 | 302.792 | 438 | 269.0886 | 259.0841 | 275 | 415.3584 | 207.2036 | 212.1593 | 222.438 | 281.2832 | 229.3629 | 214.8761 | 263.3629 | 185.9204 | 170.1195 | 167.5177 | 182.3186 | 233.8009 | 358.1593 | 388.5575 | 453.885 | 705 | 362.0752 | 368.2036 | 411.6416 | 621.3097 | 151.3893 | 102.1195 | 88.0796 | 151.3893 | 142.3629 | 135.5575 | 130.1548 | 182.6018 | 157.1239 | 145.5177 | 157.5973 | 168.8009 | 109.4027 | 103.3186 | 109.9248 | 136.1195 | 135.1195 | 126.1195 | 135.2434 | 97.6416 | 331.6018 | 340.8053 | 347.8009 | 333.3982 | 226.1637 | 273.9557 | 315.8805 | 342.9248 | 345.7566 | 350.2036 | 335.0752 | 471.7566
731 | 10.0354 | 230.5595 | 315.119 | 297.5231 | 319.8348 | 227.3117 | 213.4867 | 240.1927 | 271.1377 | 180.8899 | 168.9449 | 175.8348 | 213.0551 | 260.9174 | 196.0275 | 214.6607 | 347.3943 | 259.2016 | 207.3028 | 256.119 | 295.3393 | 223.7247 | 396.0737 | 437.6332 | 283.9085 | 311.587 | 325 | 352.3765 | 418.9085 | 194.9449 | 175.3579 | 193.1012 | 225.7798 | 258.3482 | 214.3668 | 230.4956 | 146.1377 | 325.3668 | 404.9263 | 412.0275 | 507.7611 | 174.8534 | 193.8073 | 211.9085 | 209.8623 | 117.3028 | 103.3579 | 119.0551 | 223.8623 | 248.5781 | 313.9449 | 338.9814 | 246 | 220.2931 | 225.9449 | 215.6057 | 215.3943 | 256.2291 | 241.6332 | 328.8899 | 269.8348 | 185 | 216.6972 | 200.881 | 186.9174 | 249.4494 | 250.5506 | 317.3765 | 377.7798 | 216.9174 | 227.413 | 379.8988 | 322.8721 | 286.2478 | 310.8623 | 320.4494 | 465.2291 | 221.4041 | 243.468 | 281.8073 | 419.8623 | 197.3028 | 201.3028 | 258.9538 | 242.1288 | 236.8073 | 243.1741 | 244.5506 | 174.3304 | 171.9449 | 167.9725 | 158.6607 | 246.6421 | 355.0826 | 403.5595 | 421.936 | 670.9636 | 357.1101 | 354.413 | 416.8623 | 707.4947 | 167.4769 | 94.2202 | 111.3393 | 167.4769 | 134.2478 | 137.9174 | 178.5959 | 162.5781 | 137.5781 | 142.0826 | 129.7709 | 144.6883 | 97.3579 | 107.8623 | 94.468 | 130.1652 | 134.0275 | 134.7522 | 117.5231 | 108.6972 | 326.1652 | 320.6057 | 346.0551 | 341.7522 | 212.413 | 297.3117 | 320.8623 | 335.2478 | 359.587 | 389.8988 | 337.8899 | 475.8623
732 | 10.0464 | 212.2896 | 292.3657 | 321.619 | 305.2286 | 197.4724 | 242.5276 | 256.7409 | 294.6343 | 190.8476 | 174.9086 | 163.1981 | 222.8476 | 282.6647 | 226.5276 | 216.9543 | 356.8628 | 232.4267 | 237.5276 | 246.1676 | 274.3353 | 227.939 | 426.5123 | 432.0762 | 280.0457 | 305.1067 | 357.4971 | 382.5123 | 458.3752 | 169.3962 | 206.5123 | 217.6038 | 206.3048 | 203.8687 | 206.1372 | 244.7714 | 156.8324 | 350.619 | 396.1524 | 394.2743 | 489.3048 | 157.2591 | 194.9848 | 222.8171 | 217.8781 | 139.6495 | 103 | 113.0914 | 200.3657 | 276.5581 | 311.0457 | 329.1676 | 246 | 210.1829 | 216.1524 | 235.68 | 183.5029 | 280.6343 | 244.939 | 329.9848 | 281.8171 | 200.7562 | 204.1981 | 231.5123 | 190.939 | 239.1676 | 285.4514 | 344.5581 | 337.6248 | 223.8933 | 237.8324 | 322.8991 | 362.3599 | 290.9238 | 295.2438 | 308.1981 | 430.5486 | 248.5581 | 256.7866 | 273.1372 | 439.6952 | 170.4115 | 194.1067 | 205.8382 | 257.7409 | 223.2134 | 247.939 | 240.061 | 164.1524 | 163.1372 | 196.5581 | 179.6647 | 245.0305 | 377.6495 | 374.4572 | 404.2591 | 693.6343 | 344.1981 | 375.6647 | 413.061 | 682.4267 | 153.2286 | 82.1829 | 107.0762 | 153.2286 | 124.1524 | 148.8324 | 154.3962 | 172.8324 | 151.7714 | 127.2286 | 140.8171 | 156.8019 | 96.0152 | 108.9848 | 104.8324 | 134.9238 | 120.2134 | 115.3048 | 138.6647 | 116.8781 | 339.7866 | 309.1676 | 364.7104 | 358.7409 | 226.7714 | 284.2134 | 295.3962 | 311.3657 | 364.9238 | 296.463 | 344.8933 | 464.1829
733 | 10.0574 | 210.0059 | 304.9617 | 324.9912 | 324.941 | 207.9676 | 230.0383 | 241.0472 | 242.1563 | 191 | 174.0029 | 182.941 | 201.0649 | 287.9853 | 224.0088 | 216.0029 | 330.0796 | 234.9912 | 242.9853 | 234.0354 | 292.944 | 215.0383 | 446.941 | 467.8939 | 292.9617 | 293.0354 | 367.9705 | 368.0442 | 459.9971 | 175.9794 | 182.0737 | 197.0619 | 223.9469 | 232.9115 | 189.0501 | 265.9381 | 135.0649 | 353.9912 | 368.0826 | 394.9971 | 440.1445 | 205.8555 | 211.9499 | 249.9204 | 203.0442 | 112.0826 | 105.9912 | 127.9558 | 228.9145 | 269.0236 | 293.0531 | 329.9971 | 227.056 | 193.0501 | 233.9469 | 196.1179 | 184.9941 | 237.1297 | 291.8614 | 328.0059 | 286.9853 | 183.0531 | 188.0472 | 222.0295 | 191.9971 | 214.0737 | 278.0236 | 330.0442 | 319.0531 | 201.0678 | 206.0944 | 364.8732 | 382.941 | 292.9941 | 286.0265 | 319.9646 | 436.9794 | 230.056 | 244.0383 | 282.9705 | 422.0531 | 191.9351 | 195.9941 | 217.9617 | 245.0383 | 243.9381 | 248 | 240 | 181.9469 | 150.0383 | 154.1268 | 180.9971 | 250.9823 | 365.0383 | 378.9853 | 405.9941 | 694.9971 | 330.0413 | 392.9499 | 411.0059 | 650.0944 | 179.9204 | 81.0029 | 93.0413 | 179.9204 | 132.9735 | 133.0472 | 171.9469 | 164.0265 | 104.1415 | 142.9528 | 155.9558 | 176.941 | 96 | 109 | 99.0177 | 135 | 126.9794 | 160.8644 | 114.0737 | 123.9794 | 330.0295 | 341.9027 | 350.0442 | 343.0472 | 214.0383 | 292.9735 | 286.0265 | 320.9705 | 390.9233 | 322.9174 | 366.9351 | 474.9676
734 | 10.0684 | 198.9346 | 284.3924 | 361.7104 | 314.972 | 216.9813 | 274.7104 | 225.5792 | 230.4671 | 215.6637 | 171.8131 | 168.8225 | 205.0934 | 270.9346 | 220.8599 | 219.1775 | 332.8879 | 261.8505 | 231.9066 | 243.8318 | 309.7384 | 206.0467 | 420.9253 | 422.0093 | 272.4111 | 317.7851 | 328.0187 | 357.1869 | 442.9159 | 161.0093 | 200.8599 | 193.1869 | 265.5983 | 210.0747 | 194.972 | 249.9159 | 156.7664 | 340.0934 | 374.2709 | 402 | 460.5325 | 208.2616 | 198.8225 | 206.2616 | 200.028 | 120.9066 | 96.8879 | 106.0467 | 195.1869 | 268.0374 | 310.7104 | 351.1682 | 230.1308 | 209.8785 | 221.7478 | 212.0561 | 207.8225 | 267.7104 | 295.6356 | 315.9907 | 283.2149 | 196 | 209.8131 | 219.6543 | 188.8038 | 232.9066 | 296.1869 | 314.7851 | 373.7291 | 208.1775 | 231.9907 | 337.9907 | 360.9626 | 278.9159 | 290.1495 | 314.1588 | 465.8412 | 217.1962 | 209.383 | 264.4017 | 414.1588 | 190.8972 | 201.8038 | 213 | 208.4484 | 214.2803 | 228.0374 | 272.9066 | 176.1775 | 169.8692 | 171.0467 | 188.9253 | 273.6637 | 327.1495 | 432.4301 | 424.5702 | 694.3176 | 358.5516 | 380.4301 | 451.2619 | 665.6917 | 155 | 85.0561 | 110.8879 | 155 | 135.9626 | 124.2616 | 163.673 | 175.0841 | 145.9346 | 138.9253 | 159.9159 | 124.3083 | 97.0747 | 109.8879 | 90.3363 | 122.9533 | 131.9066 | 104.1962 | 148.9066 | 101.9907 | 335.9626 | 369.0561 | 364.1028 | 386.701 | 239.8879 | 301.4862 | 293.3737 | 296.1775 | 358.028 | 340.1869 | 370.8318 | 527.617
735 | 10.0794 | 192.1082 | 325.3943 | 331.7572 | 312.5841 | 213.8967 | 243.762 | 285.4856 | 278.9616 | 180.3029 | 152.7139 | 150.1947 | 214.4808 | 264.5408 | 206.411 | 238.7139 | 322.1033 | 245.8486 | 221.6322 | 226.3461 | 282.6057 | 211.1298 | 412.9784 | 422.8918 | 315.2645 | 295.5625 | 330 | 376.762 | 434.411 | 162.3894 | 186.3461 | 212.1996 | 222.9135 | 217.4375 | 191.8486 | 240.7188 | 132.8004 | 349.5457 | 402.3943 | 401.6539 | 515.8967 | 235.7188 | 180.4327 | 234.6706 | 202.5241 | 111.0433 | 85.3678 | 111.4543 | 215.3029 | 271.8053 | 280.2163 | 368.351 | 243.3077 | 196.9784 | 195.649 | 217.6971 | 188.8486 | 236.8702 | 257.8004 | 315.4327 | 304.7453 | 195.6971 | 189.7404 | 183.5192 | 168.8004 | 223.5408 | 315.3294 | 292.0865 | 344.9784 | 226.2645 | 231.238 | 336.9135 | 357.3029 | 270.2812 | 306.0216 | 330.7188 | 449.411 | 237.5024 | 249.4592 | 306.5024 | 430.0265 | 180.4759 | 181.5192 | 212.9567 | 255.4808 | 243.8053 | 231.9784 | 262.4592 | 194.4808 | 156.6706 | 175.8702 | 180.8702 | 237.4808 | 343.238 | 373.5793 | 380.6009 | 726.053 | 311.6057 | 321.1898 | 373.411 | 633.3245 | 155.5192 | 90.6755 | 99 | 155.5192 | 132.4543 | 152.0649 | 129.4543 | 183.2428 | 138.7837 | 131.2812 | 150.7404 | 156.8269 | 104.4375 | 98.1082 | 125.3943 | 118.1947 | 122.2812 | 125.0216 | 138.3726 | 101.2163 | 330.7453 | 373.5938 | 374.9351 | 355.4759 | 227.8053 | 248.06 | 332.4592 | 314.4808 | 360.5673 | 359.4808 | 352.7404 | 486.2861
736 | 10.0904 | 197.1696 | 298.5089 | 365.4911 | 338.5251 | 165.3231 | 232.3554 | 218.5266 | 233.2891 | 193.7625 | 184.0162 | 159.6106 | 191.8481 | 289.0339 | 223.8466 | 271.0339 | 370.4395 | 238.3554 | 205.0339 | 241.3894 | 308.3716 | 217.2375 | 412.3392 | 417.4572 | 281.5929 | 318.5574 | 331.0517 | 366.5089 | 452.1519 | 179.5929 | 200.1681 | 176.882 | 218.8304 | 193.1195 | 185.4071 | 229.3231 | 168.2876 | 329.2035 | 374.3894 | 387.0177 | 465.4233 | 221.8466 | 199.9322 | 263.8126 | 182.0517 | 112.9661 | 102.0339 | 131.3554 | 228.3215 | 264.5266 | 290.1018 | 338.7965 | 213.9337 | 196.6785 | 224.9661 | 205.2891 | 182.3053 | 231.6446 | 293.0501 | 334.2876 | 248.7124 | 182.5767 | 178.5089 | 206.9661 | 204.2197 | 247.5929 | 285.0678 | 298.0694 | 344.0678 | 193.6785 | 241.0162 | 333.6785 | 370.8643 | 282.6268 | 306.8982 | 316.7787 | 468.2714 | 215.0339 | 226.357 | 283.5929 | 387.1534 | 200.6091 | 204.4233 | 211.8481 | 232.5428 | 234.0501 | 231.4071 | 238.0339 | 171.882 | 185.8126 | 170.7803 | 175.5089 | 214.3732 | 352.8466 | 443.2698 | 451.2359 | 639.4588 | 337.8805 | 375.5767 | 392.441 | 646.8466 | 178.4233 | 76.8142 | 99.1018 | 178.4233 | 151.9144 | 154.4233 | 150.5767 | 150.5605 | 129.5767 | 144.1357 | 139.2035 | 149 | 79.1357 | 102.6268 | 98.1018 | 127 | 134.118 | 126.2714 | 110.5767 | 111.7124 | 274.1018 | 311.357 | 371.2197 | 375.3038 | 218.9661 | 295.7286 | 308.2375 | 291.5428 | 341.7463 | 336.1018 | 341.0339 | 453.6268
737 | 10.1014 | 201.2605 | 312.4916 | 350.7227 | 325.7395 | 203.9243 | 214.5084 | 263.5713 | 267.4581 | 186.7227 | 156.8319 | 175.4749 | 216.3235 | 289.7227 | 192.4327 | 269.5968 | 326.1386 | 219.6303 | 205.2143 | 224.1849 | 262.3235 | 224.1849 | 421.353 | 404.2646 | 270.2311 | 250.2478 | 358.7354 | 381.1386 | 428.2773 | 169.2478 | 150.5419 | 201.9538 | 212.3362 | 223.6135 | 197.6008 | 266.8151 | 147.6303 | 335.6932 | 356.9538 | 415.2605 | 450.0167 | 189.0924 | 198.0462 | 230.1849 | 211.4454 | 112.1849 | 102.3992 | 112.1219 | 209.9243 | 306.5673 | 292.2143 | 335.1721 | 267.3362 | 214.9833 | 225.8486 | 240.937 | 191.5546 | 249.5841 | 265.6303 | 314.3235 | 267.8908 | 199.4622 | 193.7857 | 206.9243 | 183.1554 | 236.2311 | 286.6765 | 356.7227 | 345.2143 | 213.1386 | 212.9538 | 351.6597 | 366.3992 | 272.3235 | 304.2311 | 282.9243 | 476.1386 | 217.0167 | 263.706 | 270.706 | 418.1514 | 162.2016 | 187.3068 | 234.6135 | 247.5378 | 207.8781 | 244.0167 | 239.5546 | 196.0295 | 152.2311 | 191.7522 | 190.1386 | 226.4789 | 321.3865 | 392.6597 | 400.9538 | 680.2143 | 307.3403 | 389.643 | 406.6175 | 616.0335 | 161.7227 | 99.4454 | 101.9076 | 161.7227 | 122.5251 | 138.2311 | 165.7984 | 193.3825 | 145.7227 | 147.4454 | 145.7857 | 148.2143 | 83.6008 | 92.3697 | 101.0462 | 127.3697 | 109.0924 | 132.6597 | 126.0757 | 131.0757 | 277.9243 | 348.8908 | 349 | 327.2311 | 219.4789 | 287.4916 | 313.5211 | 310.1889 | 362.8151 | 338.2605 | 342.6932 | 443.6932
738 | 10.1124 | 187.5797 | 303.2286 | 346.8722 | 338.3087 | 223.2925 | 225.9361 | 292.8991 | 218.4519 | 180.3564 | 173.1809 | 144.5266 | 223.6436 | 282.9469 | 222.5266 | 219.883 | 329.2925 | 212.8191 | 189.8191 | 228.9946 | 269.7606 | 227.883 | 409.0531 | 452.5797 | 275.4681 | 277.6967 | 312.1755 | 382.8299 | 434.9361 | 193.3618 | 202.4734 | 201.0585 | 180.5158 | 196.1701 | 209.883 | 262.0639 | 139.4149 | 349.4734 | 357.8137 | 399.766 | 471.7021 | 191.7606 | 197.8299 | 233.4149 | 200.8776 | 115.5319 | 90.6436 | 94.7606 | 227.5373 | 277.7021 | 278.9838 | 377.6597 | 232.9415 | 194.7021 | 261.4842 | 218.8245 | 203.4095 | 240.9415 | 257.4734 | 321.7606 | 247.3988 | 209.351 | 209.1809 | 225.3564 | 205.8884 | 239.0693 | 280.4681 | 349.7714 | 328.5905 | 215.1809 | 213.6382 | 324.9946 | 353.2394 | 277.8884 | 307.6543 | 302.351 | 474.846 | 237.7075 | 240.1594 | 245.3457 | 379.8245 | 185.8299 | 173.9946 | 206.2925 | 237.9415 | 225.2394 | 265.9361 | 250.8245 | 175.4149 | 157 | 166.7606 | 192.0639 | 256.0639 | 349.1224 | 366.5797 | 400 | 663.0639 | 332.4257 | 342.1116 | 439.2394 | 657.5319 | 156.9361 | 88.2925 | 100.351 | 156.9361 | 152.1278 | 143.2925 | 141.5266 | 161.4042 | 139.4149 | 135.7075 | 162.9361 | 132.5266 | 95.883 | 99.2979 | 101.7075 | 134.883 | 111.9946 | 105.5851 | 107.6436 | 111.0054 | 299.8083 | 326.8776 | 348.766 | 333.9892 | 248.3033 | 274.5427 | 285.1062 | 372.4896 | 359.9946 | 322.2979 | 355.8299 | 459.9892
739 | 10.1234 | 213.354 | 322.7876 | 374.5221 | 296.1328 | 226.8672 | 239.4424 | 242.3628 | 271.3097 | 170.9204 | 159.5752 | 155.0531 | 231.6637 | 266.708 | 231.354 | 218.708 | 332.7256 | 227.1328 | 201.2301 | 245.2124 | 282.708 | 223.8761 | 425.7168 | 474.7433 | 282.5752 | 306.062 | 314.5752 | 365.2036 | 447.1681 | 167.4248 | 192.8672 | 201.8584 | 221.5664 | 216.062 | 207.292 | 247.0708 | 130.9912 | 339.8672 | 387.8496 | 397.2744 | 480.3097 | 202.4424 | 180.3452 | 223.5752 | 214.9292 | 108.4248 | 100.292 | 106.5752 | 205.1328 | 289.4956 | 324.7345 | 338.0885 | 235.1151 | 206.2124 | 221.2832 | 219.4691 | 209.0796 | 239.8584 | 250.1328 | 332.0177 | 285.4336 | 214.8584 | 196.2124 | 216.3452 | 185.8761 | 209.2036 | 288 | 330.7788 | 322.0708 | 202.708 | 239.2212 | 341.7788 | 342.062 | 258.7964 | 287.416 | 306.4424 | 409.8407 | 235.0531 | 289.0885 | 267.2124 | 380.3983 | 167.4248 | 190.354 | 211.4956 | 237.2124 | 214.2036 | 277.1769 | 248.1416 | 166.708 | 157.2832 | 180.9115 | 177.2832 | 239.938 | 334.0796 | 391.8584 | 398.938 | 651.2567 | 296.3098 | 359.5044 | 425.938 | 650.2124 | 172.2832 | 92.7168 | 105.9292 | 172.2832 | 123.3452 | 147.3628 | 150.7788 | 181.7345 | 129.3628 | 132.416 | 177.646 | 142.1328 | 94.9204 | 90.1239 | 96.292 | 131.3009 | 127.9292 | 114.1504 | 117.7876 | 96.3452 | 341.7433 | 340.292 | 343.3009 | 361.9645 | 220.646 | 239.3895 | 317.0885 | 315.8584 | 373.5929 | 311.4956 | 336.7168 | 489.0265
740 | 10.1344 | 218.914 | 318.5875 | 344.1603 | 309.172 | 210.5875 | 219.5816 | 237.3236 | 239.3236 | 183.1662 | 155.4956 | 180.9198 | 203.4096 | 277.3294 | 233.2581 | 228.4154 | 317.0802 | 241.2522 | 179.4956 | 246.9198 | 292.4154 | 198.825 | 424.7419 | 433.7449 | 278.3294 | 318.7537 | 309.8309 | 378.9228 | 413.9883 | 173 | 177.3353 | 198.5875 | 229.7478 | 230.4985 | 198.3324 | 245.9228 | 144.1662 | 322.0919 | 397.8368 | 413.8338 | 445.4985 | 181.6676 | 197.4184 | 221.8221 | 212.089 | 114.1662 | 91.5816 | 100.9169 | 217.7566 | 291.5132 | 296.9941 | 315.7419 | 275.2522 | 208.0029 | 224.5015 | 261.261 | 196.7507 | 237.6676 | 262.0919 | 304.5044 | 276.2551 | 211.6706 | 200.828 | 231.9228 | 161.9883 | 225.4985 | 289.6618 | 337.4272 | 323.914 | 212.5816 | 227.6706 | 351.5015 | 354.089 | 243.6647 | 303.175 | 289.4897 | 432.4213 | 258.3441 | 261.3412 | 270 | 425 | 173.7478 | 195.8309 | 217.086 | 243.2405 | 229.5846 | 226.1691 | 249.5015 | 177.4125 | 161.2493 | 205.1691 | 180.086 | 225.3353 | 323.911 | 391.7449 | 383.4213 | 692.5044 | 362.6764 | 354.7449 | 413.2463 | 653.5816 | 175.086 | 88.0029 | 105.3324 | 175.086 | 141.3324 | 140.4125 | 160.8338 | 154.4096 | 122.2463 | 148.3412 | 170.7566 | 156.6676 | 105.9198 | 115.7566 | 87.8309 | 109.3324 | 125.6706 | 103.3324 | 115.7478 | 113.5015 | 302.3236 | 330.7507 | 325.4037 | 311.0772 | 215.5846 | 295.4213 | 293.7449 | 313.8338 | 345.4066 | 315.3412 | 332.9169 | 453.8221
741 | 10.1454 | 227.4739 | 301.7601 | 363.6125 | 289.0954 | 195.4769 | 231.2459 | 270.9478 | 273.4247 | 185.4769 | 170.4739 | 170.2399 | 226.3756 | 291.0462 | 204.0492 | 233 | 327.4739 | 233.9538 | 196.4769 | 235.763 | 297.763 | 228.8554 | 451.2802 | 452.6185 | 288.4218 | 305.2861 | 319.763 | 355.4277 | 456.2832 | 172.7139 | 172.2891 | 183.0954 | 234.0894 | 233.9016 | 202.1908 | 224.0492 | 144.0924 | 295.1505 | 384.5231 | 410.9508 | 450.237 | 180.0984 | 189.9508 | 261.3294 | 191.6677 | 117.7168 | 99.3353 | 100.2861 | 193.4769 | 246.3875 | 321.9568 | 342.9478 | 264.0432 | 196.7139 | 217.6647 | 219.6185 | 192.1878 | 231.4247 | 230.4769 | 290.766 | 257.763 | 197.2861 | 217.5663 | 210.1446 | 204.9508 | 229.7601 | 306.0924 | 298 | 335.7168 | 219.763 | 214.5261 | 346 | 337.9598 | 255.5782 | 274.1937 | 326.9508 | 419.2921 | 216.2832 | 236.2921 | 268.8554 | 422.8063 | 181.5231 | 205.763 | 207.4769 | 275.8525 | 222.0432 | 216.6185 | 244 | 191.7601 | 163.6185 | 193.8063 | 172.9568 | 218.6677 | 344.8092 | 409.5693 | 371.3413 | 672.3756 | 326.2861 | 372.2832 | 423.0894 | 660.1908 | 165.0954 | 78.1446 | 107.4739 | 165.0954 | 145.0954 | 155.5231 | 159.0954 | 179.0924 | 137.2399 | 121.6706 | 145.237 | 153.9538 | 92.6647 | 90.8092 | 98.5261 | 114.5261 | 111.9538 | 109.0984 | 123.7139 | 106.0924 | 335.1848 | 325.5201 | 371.3294 | 331.9016 | 212.3353 | 278.763 | 310.8063 | 314.003 | 386.1475 | 287.9076 | 334.5753 | 490.468
742 | 10.1564 | 212.8923 | 291.1844 | 325.6608 | 290.969 | 198.9233 | 278.3082 | 242.5384 | 247.1696 | 191.3997 | 155.8923 | 181.5693 | 191.7847 | 282.1077 | 211.2006 | 234.0767 | 315.1534 | 239.5545 | 201.7537 | 244.6151 | 304.2463 | 213.9085 | 422.3526 | 446.5236 | 246.6151 | 310.7994 | 329.9071 | 353.8762 | 439.4616 | 170.323 | 192.1696 | 186.7994 | 197.4307 | 211.8466 | 205.2921 | 232.0619 | 127.2921 | 341.0929 | 381.2921 | 398.3082 | 443.969 | 197.2006 | 181.0767 | 238.7994 | 197.5693 | 130.6622 | 111.2463 | 105.0457 | 200.1534 | 286.8156 | 345.2625 | 319.1682 | 234.6003 | 197.8762 | 204.5693 | 217.9381 | 174.4616 | 210.2301 | 239.8452 | 315.1696 | 266.6151 | 196.5545 | 185.1077 | 221.323 | 192.5236 | 217.031 | 288.2153 | 301.3378 | 352.4307 | 225.4926 | 227.6003 | 347.0767 | 381.3392 | 295.3392 | 292.0929 | 316.2463 | 459.8923 | 203.0148 | 271.7241 | 259.7227 | 401.8923 | 179.1534 | 210.2006 | 212.323 | 245.8923 | 199.9219 | 214.3997 | 245.9381 | 179.2463 | 162.5841 | 174.9381 | 200.031 | 225 | 341.1696 | 395.354 | 419.5236 | 645.3216 | 324.5855 | 356.2153 | 383.0929 | 658.4469 | 165.5693 | 88.8923 | 91.9233 | 165.5693 | 144.4926 | 148.7389 | 160.8614 | 164.0148 | 146.4159 | 145.2773 | 140.7994 | 163.4307 | 81.8304 | 92.1225 | 111.7079 | 127.2773 | 121.7537 | 128.7847 | 123.2611 | 92.5222 | 302.5531 | 303.2921 | 349.2301 | 316.7375 | 222.7389 | 288.7994 | 289.5693 | 328.1548 | 411.3701 | 309.4455 | 354.9085 | 446.506
743 | 10.1674 | 209.9606 | 299.6804 | 354.9203 | 297.0806 | 191.1996 | 253.6008 | 244.9606 | 234.0394 | 200.6008 | 157.3992 | 176.0806 | 191.6795 | 282.1603 | 191.4789 | 241.0806 | 326.8819 | 213.2793 | 208.9597 | 252.7217 | 296.9203 | 188.2774 | 468.5221 | 423.9606 | 261.6795 | 327.2821 | 346.5624 | 382.761 | 430.2015 | 171.2006 | 176.1603 | 208.0806 | 198.3608 | 198.6785 | 215.1603 | 219.4789 | 138 | 320.3589 | 390.6804 | 383.5576 | 451.4002 | 177.7188 | 189.7601 | 259.3608 | 191.9606 | 104.5595 | 105 | 121.4002 | 217.4002 | 276.2802 | 307.1996 | 359.3608 | 223.5998 | 225.3215 | 205.0787 | 233.5202 | 171.7994 | 237.12 | 277.4808 | 298.3205 | 278.9203 | 173.6382 | 188.9991 | 226.0394 | 172.4798 | 213.1987 | 290.2399 | 328.1603 | 351.3215 | 211.4405 | 217.4395 | 354.8004 | 351.5998 | 262.6008 | 272.9184 | 310.5998 | 453.7217 | 226.4405 | 230.6392 | 274.7198 | 402.4395 | 197.3599 | 188.5595 | 212.9606 | 245 | 244.6814 | 225.1603 | 262.8397 | 174.5595 | 181.1209 | 188.2409 | 191.6401 | 224.7601 | 327.4395 | 386.6411 | 402.119 | 699.0422 | 295.9175 | 359.0797 | 361.9991 | 629.7198 | 160.9203 | 87.4002 | 83.2399 | 160.9203 | 135.3589 | 129.88 | 168.4798 | 184.2015 | 126.3196 | 131.8397 | 163.6401 | 165.6804 | 97.7198 | 115.2409 | 99.4405 | 113.5998 | 127.7601 | 126.4002 | 140.8004 | 125.3215 | 326.8416 | 315.6795 | 373.3608 | 356.3599 | 208.4386 | 309.8407 | 286.8397 | 291.8388 | 371.0806 | 340.3992 | 323.3992 | 509.3224
744 | 10.1784 | 197.5128 | 294.0998 | 346.8677 | 289.7449 | 203.1741 | 239.5708 | 236.0836 | 249 | 185.5128 | 175.9258 | 162.3968 | 203.3387 | 270.7099 | 218.058 | 227.9258 | 287.8191 | 229.8677 | 213.7517 | 215.8515 | 289.7935 | 244.3711 | 426.4224 | 414.1579 | 268.8421 | 280.6451 | 293.058 | 365.3481 | 404.7867 | 157.8677 | 172.116 | 198.8934 | 182.1741 | 225.9906 | 209 | 244.8677 | 140.116 | 346.587 | 386.4224 | 430.4036 | 444.7517 | 201.4292 | 190.3806 | 237.4872 | 180.5708 | 114.8097 | 103.2807 | 114.884 | 215.3806 | 267.6937 | 309.2551 | 338.5452 | 227.8677 | 196.4224 | 235 | 227.2227 | 184.2065 | 234.1647 | 257.8097 | 291.0256 | 269.0162 | 211.5614 | 208.3549 | 238.2227 | 179.4386 | 239.5034 | 292.6613 | 321.6032 | 318.7193 | 205.9514 | 228 | 346.2645 | 351.6356 | 246.7193 | 306.2227 | 315.1323 | 424.7867 | 216.8515 | 249.2065 | 276.9582 | 412.471 | 198.942 | 198.413 | 196.5452 | 245 | 213.1323 | 220.1903 | 264.2389 | 186 | 160.2645 | 164.4292 | 187.0162 | 221.6775 | 334.9582 | 362.529 | 424.8677 | 650.6451 | 352.8097 | 364.2227 | 386.6451 | 634.6032 | 158.6869 | 82.2065 | 88.0418 | 158.6869 | 155.7679 | 130.3225 | 169.7517 | 149.7935 | 136 | 136.6775 | 161.1323 | 156 | 102.471 | 94.3387 | 91.4386 | 120.9096 | 126.1323 | 120.6775 | 131.8677 | 91.7935 | 290.587 | 327.6032 | 357.3063 | 355.0324 | 239.942 | 292.2807 | 294.9838 | 318.2807 | 352.9002 | 358.6613 | 335.9744 | 458.2901
745 | 10.1894 | 216.0236 | 318.5427 | 345.5664 | 329.5191 | 221.9764 | 263.5782 | 271.8082 | 251.6018 | 197.9528 | 182.4336 | 165.2891 | 205.2154 | 242.0709 | 221.6755 | 233.7109 | 322.6755 | 222.6401 | 199.8673 | 228.5782 | 287.1946 | 218.1918 | 465.18 | 446.9882 | 236.0118 | 288.4601 | 301.59 | 402.3746 | 463.941 | 156.7109 | 183.9764 | 231.5782 | 206.3245 | 180.7227 | 209.2891 | 244.2891 | 151.2537 | 352.23 | 421.5664 | 381.9499 | 434.1918 | 182.9145 | 204.2655 | 224.6136 | 205.59 | 108.0118 | 92.5782 | 102.59 | 226.5191 | 250.9145 | 273.0945 | 331.7463 | 224.9764 | 226.0737 | 233.6991 | 214.059 | 178.8555 | 207.9882 | 251.0118 | 327.3245 | 258.4573 | 188.4218 | 192.3009 | 217.2537 | 198.82 | 213.6136 | 299.1682 | 317.5546 | 328.8436 | 237.4837 | 228.7227 | 348.4336 | 329.0709 | 260.6018 | 288.4336 | 317.4454 | 483.2182 | 227.41 | 244.5782 | 258.3009 | 404.6637 | 190.9882 | 186.6991 | 191.1918 | 242.1091 | 217.9027 | 223.8082 | 232.5664 | 184.8436 | 163.0118 | 145.9145 | 174.7227 | 214.59 | 321.5045 | 371.7818 | 416.0501 | 669.7345 | 345.1445 | 347.1564 | 399.8082 | 632.7227 | 192.5309 | 77.4336 | 107.4218 | 192.5309 | 130.3127 | 138.5664 | 155 | 154.8555 | 136.5782 | 125.9764 | 157.23 | 153.8318 | 97.8436 | 92.1682 | 111.8318 | 134.3628 | 125.8436 | 114.8909 | 128.9764 | 101.6254 | 302.7345 | 323.8436 | 373.3628 | 330.1564 | 233 | 284.7581 | 311.7581 | 303.9646 | 325.4218 | 353.3155 | 313.5517 | 492.1327
746 | 10.2004 | 223.1381 | 305.5684 | 342.059 | 303.1568 | 210.941 | 262.9223 | 250.9209 | 263.7065 | 181.0576 | 181.8633 | 168.5684 | 239.6475 | 276.9022 | 210.6259 | 235.764 | 304.5684 | 190.1756 | 206.7252 | 237.3324 | 243.5094 | 236.0993 | 475.8446 | 440.2158 | 245.9209 | 244.2331 | 310.5295 | 368.5698 | 436.3525 | 158.9209 | 171.8432 | 228.0993 | 221.9223 | 189.2345 | 207.5496 | 246.1568 | 135.1568 | 323.2158 | 415.3928 | 422.4129 | 458.8432 | 204.3151 | 188.9813 | 240.5108 | 207.3151 | 113.8432 | 98.6662 | 112.3137 | 202.9799 | 276.236 | 315.3539 | 344.5496 | 219.7439 | 195.3914 | 222.236 | 236.9424 | 178 | 205.1367 | 258.0978 | 345.9022 | 270.8043 | 183.4316 | 202.1957 | 205.7051 | 179.6863 | 235.3137 | 309.177 | 313.5482 | 323.2547 | 195.5094 | 232.059 | 350.2158 | 366.5684 | 268.7856 | 293.5094 | 318.1582 | 449.7842 | 220.0389 | 245.6475 | 266.1568 | 383.0777 | 185 | 186.0576 | 221.1756 | 225.4705 | 237.8633 | 206.5281 | 232.0389 | 176.7453 | 166.6475 | 177.6662 | 180.5684 | 226.0389 | 359.2158 | 407.5684 | 380.2921 | 678.4316 | 348.3525 | 353.8432 | 377.1957 | 641.3914 | 169.3338 | 82.1957 | 104.3137 | 169.3338 | 141.177 | 135.3727 | 157.9799 | 156.9799 | 137.8043 | 120.4302 | 134.3324 | 140.5295 | 91.7842 | 104.6863 | 100.7252 | 104.2935 | 120.882 | 130.2748 | 116.6863 | 125.8633 | 312.059 | 318.5684 | 344.0777 | 329.6287 | 235.6662 | 306.5295 | 327.2561 | 287.0978 | 324.5094 | 296.4705 | 368.6287 | 480.8834
747 | 10.2114 | 199.5369 | 311.8014 | 337.6765 | 305.5221 | 212.6175 | 245.2282 | 264.5584 | 249.3825 | 211.2202 | 168.2134 | 174.8014 | 225.647 | 273.5369 | 247.6981 | 252.1101 | 308.6027 | 217.8309 | 210.588 | 255.6833 | 262.412 | 195.4483 | 429.9705 | 444.9637 | 269.0295 | 312.706 | 308.5074 | 339.7503 | 452.8899 | 174.9262 | 166.6027 | 186.2645 | 205.2429 | 213.5221 | 190.0148 | 245.8161 | 136.1691 | 322.8899 | 393.6322 | 389.8751 | 461.2134 | 168.5812 | 165.2134 | 216.0738 | 170.5664 | 111.4779 | 109.1101 | 112.9852 | 220.8604 | 262.4268 | 287.0295 | 326.6765 | 253.6322 | 212.0658 | 204.706 | 203.9047 | 183.412 | 217.4336 | 267.0443 | 340.6765 | 259.1691 | 179.059 | 210.9557 | 222.0363 | 174.6175 | 236.8309 | 297.2134 | 333.0953 | 328.647 | 206.4631 | 227 | 346.6765 | 370.0953 | 233.3825 | 303.7866 | 274.1396 | 457.7208 | 227.6175 | 232.1544 | 263.9557 | 396.8819 | 183.9852 | 217.235 | 244.941 | 226.9852 | 216.7718 | 246.4926 | 243.5074 | 169.6617 | 154.8456 | 187.2645 | 187.4779 | 237.3383 | 349.0806 | 412.4483 | 430.5142 | 666.4483 | 359.8161 | 348.2645 | 386.9705 | 652.992 | 156.6913 | 92.647 | 103.9705 | 156.6913 | 129.5517 | 132.6765 | 172.8161 | 169.2792 | 126.1691 | 155.5074 | 162.4926 | 138.5074 | 92.4483 | 101.1396 | 107.4631 | 118.2792 | 128.8014 | 119.1396 | 119.7355 | 109 | 309.0295 | 327.3383 | 366.6765 | 301.1623 | 242.3893 | 304.5074 | 293.1986 | 296.2134 | 343.5812 | 300.6913 | 335.5959 | 446.3087
748 | 10.2224 | 209.823 | 298.4602 | 344.0841 | 312.2743 | 195.0929 | 264.3672 | 237.7965 | 261.8274 | 178.6239 | 180.5531 | 167.9912 | 215.5531 | 279.1061 | 203.2699 | 235.5398 | 293.8939 | 222.9868 | 191.0753 | 208.6283 | 288.4557 | 182.1726 | 424.8983 | 489 | 276.4602 | 316.022 | 314.4469 | 363.8274 | 458.5751 | 153.6283 | 147.3584 | 166.6283 | 227.7345 | 220.0929 | 194.8186 | 240 | 139.9027 | 340.9115 | 383.6328 | 405.9956 | 476.2743 | 194.5487 | 179.7301 | 241 | 191.1858 | 104 | 89.6372 | 108.1814 | 224.3761 | 291.823 | 298.6328 | 334.354 | 242.5443 | 182.354 | 216.0044 | 231.0841 | 199.6592 | 200.6239 | 287.2611 | 335.1106 | 263.8098 | 195.0088 | 195.4557 | 187.7965 | 163.6239 | 234.3672 | 305.7433 | 316.6195 | 326.8983 | 195.2699 | 230.8098 | 354.7168 | 347.4513 | 253.4469 | 291.6283 | 270.6239 | 467.1106 | 206.4646 | 235.4381 | 254.8054 | 364.2567 | 176.6416 | 188.354 | 220.1017 | 229.2567 | 201.8054 | 241.6416 | 246.5443 | 173.2611 | 157.3672 | 163.2743 | 175.646 | 233.1946 | 323.2655 | 396.9956 | 383.7168 | 645.1902 | 350.9159 | 326.9956 | 374.6416 | 603.6151 | 163.3672 | 86.9071 | 94.3628 | 163.3672 | 145.3672 | 141.9868 | 163.7345 | 158.0753 | 130.2655 | 152.9203 | 158.0044 | 146.4425 | 77.5398 | 92.7257 | 97.7213 | 105.6239 | 118.9071 | 117.4381 | 139.0088 | 107.7301 | 316.4602 | 327.5487 | 370.3628 | 347.1106 | 212.9824 | 303.3717 | 302.5487 | 308.9159 | 364.2876 | 313.5354 | 363.8274 | 433.1017
749 | 10.2334 | 200.1937 | 288.7434 | 358.9685 | 305.707 | 194.293 | 262.2301 | 277.1573 | 246.6805 | 202.644 | 165.7748 | 183.8692 | 209.298 | 255.942 | 215.7798 | 247.1937 | 328.293 | 246.9007 | 225.3245 | 227.5182 | 284.644 | 204.4868 | 445.0629 | 486.4818 | 262.0944 | 277.3924 | 331.356 | 345 | 406.3295 | 167.2881 | 157.8063 | 188.8113 | 213 | 216.5811 | 194.7748 | 235.5447 | 150.4818 | 339.6805 | 393.3245 | 409.5762 | 466.0265 | 186.8377 | 179.4238 | 240.4189 | 184.5497 | 104.7748 | 90.5182 | 108.2252 | 196.2881 | 278.3195 | 300.1887 | 355.5497 | 238.8013 | 199.2881 | 209.1308 | 241.707 | 159.1987 | 222.9007 | 303.6391 | 303.7169 | 281.1937 | 180.293 | 194.356 | 234.9056 | 182.4139 | 228.9371 | 281.7169 | 355.4553 | 345.3195 | 204.0993 | 246.2566 | 374.4503 | 349.351 | 270.9371 | 302.9636 | 301.6175 | 426.6126 | 193.5546 | 268.2252 | 283.1308 | 396.4189 | 166.9685 | 204.3195 | 201.4868 | 251.9272 | 226.644 | 232.9371 | 246.6755 | 196.2252 | 156.1987 | 161.356 | 155.4189 | 211.649 | 340.7119 | 402.9007 | 403.2566 | 633.6175 | 330.4139 | 344.9106 | 364.9685 | 646.0315 | 158.7119 | 88.2881 | 94.8377 | 158.7119 | 137.8063 | 163.3825 | 152.0993 | 181.8642 | 145 | 133.8742 | 147.8377 | 168.8692 | 91.5182 | 98.7119 | 109.0315 | 129.8377 | 121.4503 | 146.9321 | 121.7748 | 102.5811 | 303.8377 | 318.8692 | 365.995 | 306.0315 | 244.9321 | 284.8692 | 297.1623 | 295.1937 | 335.1623 | 328.9636 | 347.5182 | 416.6175
750 | 10.2444 | 199.558 | 301.768 | 369.18 | 289.734 | 208.618 | 274.112 | 252.468 | 256.85 | 194.734 | 169 | 173.558 | 240.172 | 280.966 | 236.584 | 251.708 | 331.906 | 234.206 | 235.206 | 234.91 | 278.382 | 214.262 | 432.056 | 470.644 | 254.738 | 302.352 | 337.618 | 343.97 | 435.438 | 161.974 | 157.618 | 207.026 | 215.678 | 214.056 | 194.91 | 216.382 | 140.618 | 346.966 | 399.91 | 387.824 | 442.768 | 186.326 | 204.266 | 237.382 | 188.322 | 111.914 | 97.91 | 104.382 | 191.18 | 266.914 | 287.948 | 353.966 | 220.622 | 185.116 | 219.03 | 231.502 | 188.296 | 217.416 | 269.91 | 340.618 | 271.082 | 193.176 | 198.352 | 235.258 | 160.768 | 232.262 | 313.056 | 369.056 | 329.588 | 216.794 | 234.468 | 365.116 | 331.442 | 281.884 | 286.326 | 319.438 | 421.498 | 226.236 | 260.262 | 285.82 | 394.412 | 169.352 | 192.09 | 214.352 | 214.532 | 219.764 | 235.232 | 239.678 | 189.498 | 176.644 | 166.176 | 162.682 | 228.116 | 349.528 | 390 | 395.588 | 655.764 | 309.386 | 373.76 | 365.292 | 636.644 | 168.352 | 75.764 | 92.266 | 168.352 | 138.03 | 137.94 | 166.854 | 154.592 | 146.236 | 149.794 | 149.592 | 155.468 | 102.412 | 105.056 | 102.734 | 126.03 | 113.764 | 126.292 | 125 | 105 | 299 | 313.502 | 351.3 | 312.712 | 228.824 | 280.532 | 302.618 | 297.854 | 341.236 | 311.502 | 350.172 | 438.558
751 | 10.2554 | 199.6756 | 322.2537 | 380.994 | 287.9853 | 210.5634 | 241.8938 | 234.5634 | 243.1121 | 185.3451 | 166.1622 | 168 | 209.7817 | 260.4779 | 207.7817 | 262.0708 | 294.0915 | 234.1268 | 239.0561 | 230.4219 | 274.2537 | 201.8938 | 426.9706 | 451.9646 | 273.2183 | 290.9793 | 338.472 | 350.0414 | 415.9646 | 178.233 | 160.2183 | 188.8024 | 224.2537 | 197.8378 | 186.9292 | 224.4779 | 140.1622 | 314.0354 | 389.0915 | 402.7877 | 468.4926 | 197.5427 | 206.6696 | 238.0561 | 178.6549 | 123.2891 | 87.0915 | 105.7109 | 211.7256 | 277.1976 | 333.5634 | 322.5634 | 246.7256 | 178.8024 | 221.0354 | 238.6696 | 195.5427 | 237.2684 | 264.3304 | 334.9233 | 237.7316 | 188.472 | 192 | 195.6756 | 180.1622 | 216.8378 | 296.6195 | 349.5634 | 329.0915 | 220.5841 | 222.0207 | 352.6903 | 337 | 287.5427 | 290.994 | 297.3451 | 413.239 | 229.6903 | 238.5074 | 258.9439 | 394.6903 | 165.6195 | 205.3097 | 204.0708 | 231.5634 | 218.9292 | 219.767 | 247.8171 | 172.9439 | 154.0354 | 160.8171 | 189.7403 | 223.3304 | 333.233 | 388.9085 | 397.7109 | 658.8585 | 340.767 | 343.528 | 350.5634 | 619.0561 | 157.8525 | 75.3658 | 102.9646 | 157.8525 | 138.0708 | 134.5841 | 165.4867 | 179.1268 | 145.761 | 152.0561 | 166.4867 | 140.1829 | 102.2537 | 86.4366 | 95.528 | 128.0354 | 112.7109 | 113.9646 | 122.5988 | 103.2537 | 296.5988 | 335.7316 | 381.8318 | 344.8318 | 229.5988 | 293.8525 | 301.5074 | 301.2891 | 344.0354 | 319.5427 | 325.0207 | 435.8378
752 | 10.2664 | 216.7719 | 312.3107 | 349.9902 | 308.8495 | 205.7719 | 253.5462 | 240.3786 | 260.619 | 184.1529 | 158.0752 | 166.3859 | 207.1553 | 294.0801 | 212.0728 | 242.466 | 300.3058 | 238.8398 | 245.3884 | 251.0049 | 275.6092 | 220.6942 | 469.1602 | 458.3083 | 272.3859 | 277.6116 | 335.9976 | 381.3884 | 427.6116 | 159.9951 | 158.9248 | 204.8471 | 220.0752 | 198.7767 | 203.767 | 256.6966 | 134.8422 | 308.6141 | 392.3083 | 434.3933 | 469.9321 | 183.7645 | 182.6165 | 248.3083 | 179.1553 | 112.3058 | 94.2281 | 117.8471 | 203.9199 | 264.8422 | 323.2379 | 329.7621 | 237.3058 | 201.7645 | 214 | 217.3835 | 180.381 | 209.9199 | 277.3156 | 310.6116 | 263.9272 | 185.767 | 192.4612 | 212.0801 | 173.9199 | 229.9976 | 296.5462 | 341.5437 | 326.7743 | 232.3884 | 242.0752 | 345.6942 | 342.9951 | 269.6141 | 257.2233 | 296.3835 | 426.466 | 222.233 | 227.3835 | 242.466 | 387.6942 | 177.5364 | 207.6942 | 197.3786 | 230.4612 | 230.6942 | 244.2306 | 241.8447 | 166.1505 | 149.9976 | 153.2306 | 163.8373 | 237.9296 | 309.2306 | 383.6165 | 411 | 675.7015 | 357.3908 | 349.9223 | 363.7572 | 620.3156 | 151.9927 | 91.4612 | 105.8495 | 151.9927 | 127.2281 | 149.3859 | 134.0777 | 174.3859 | 127.4612 | 155.3908 | 136.2306 | 153.0728 | 93.4636 | 90.5364 | 100.3083 | 121.4612 | 126.9223 | 118.4636 | 118.8422 | 96.0777 | 292.1505 | 360.3131 | 355.9174 | 317.3034 | 223.9976 | 282.2281 | 302.1431 | 305.2936 | 340.4587 | 301.8447 | 339.5413 | 447.6141
753 | 10.2774 | 206.9715 | 292.471 | 373.9704 | 305.6864 | 197.9145 | 217.6283 | 269.1288 | 250.8575 | 185.8141 | 160.3854 | 161 | 203.6716 | 282.2144 | 224.8712 | 234.6864 | 310.6716 | 257.9567 | 226.3569 | 235.6146 | 284.4847 | 217.8575 | 435.7138 | 456.2429 | 262.1425 | 289.3569 | 337.4995 | 367.943 | 446.4002 | 178.057 | 162.2006 | 204.1573 | 224.0855 | 180.8722 | 210.5713 | 237.5428 | 149.5428 | 318.1288 | 390.9715 | 412.358 | 445.3865 | 198.1425 | 174.8427 | 250.1288 | 179.3147 | 119.7571 | 102.5713 | 116.4287 | 219.9145 | 281 | 297.1573 | 351.7138 | 239.4139 | 215.6568 | 214 | 220.5428 | 189.8997 | 223.7286 | 244.2714 | 327.7001 | 250.1573 | 189.1711 | 191.0855 | 208.2292 | 189.9145 | 237.3284 | 270.3432 | 326.0718 | 310.3717 | 215.7856 | 247.5428 | 329.0137 | 360.3284 | 271.114 | 289.943 | 305.3717 | 416.7434 | 212.0285 | 231.7571 | 234.9293 | 387.0433 | 191.1003 | 202.4287 | 217.3854 | 234.9145 | 223.2429 | 247.6716 | 250.9145 | 172.7138 | 158.057 | 158.8575 | 211.415 | 211.9715 | 314.8575 | 383.1288 | 408.0855 | 636.057 | 342.4435 | 350.5713 | 408.6283 | 585.0855 | 178.5998 | 94.943 | 94.6716 | 178.5998 | 138.7286 | 138.6568 | 137.3147 | 168.5143 | 129 | 130.2429 | 144.2862 | 161.7423 | 87.1859 | 97.7856 | 98.2429 | 127.8575 | 125.3854 | 114.1288 | 140.8289 | 95.6716 | 310.8575 | 338.4857 | 374.3854 | 339.3717 | 227.9282 | 298.8289 | 331.6705 | 352.7994 | 342.2851 | 309.4572 | 329.4572 | 446.4424
754 | 10.2884 | 202.6003 | 278.9514 | 340.8408 | 316.0353 | 206.4897 | 220.8408 | 282.7655 | 259.6209 | 184.3584 | 146.7655 | 161 | 215.3171 | 276.3038 | 210.9793 | 269.531 | 320.531 | 232.938 | 219.0826 | 244.3864 | 245.7582 | 236.5723 | 421.9307 | 452.6622 | 245.2139 | 277.2345 | 316.3584 | 375.531 | 449.3451 | 166.6416 | 185.3244 | 208.8967 | 213.9793 | 215.6416 | 206.3171 | 236.5723 | 149.5929 | 324.8761 | 394.2552 | 437.1592 | 469.2005 | 181.469 | 168.3171 | 253.3038 | 196.4484 | 117.469 | 96.0207 | 123.4897 | 226.4484 | 274.8761 | 295.5176 | 331.5516 | 223.7861 | 199.2552 | 211.1932 | 220.5929 | 179.8688 | 228.0413 | 242 | 320.2418 | 259.4897 | 173.2965 | 194.2479 | 237.7655 | 194.6622 | 229.2552 | 294.3997 | 340.9174 | 326.2965 | 210.7242 | 238.9174 | 317.3305 | 347.1519 | 256.1932 | 292.6829 | 310.8348 | 439.1519 | 211.5516 | 232.531 | 263.8967 | 395.7315 | 201.8967 | 217.3997 | 212.1859 | 243.4897 | 222.469 | 249.6209 | 255.9174 | 149.4897 | 154.8068 | 170.1725 | 221.2285 | 219.5929 | 325.6622 | 398.0413 | 397.9793 | 627.1932 | 354.8421 | 344.2758 | 399.5929 | 585.6962 | 168.5103 | 105.8481 | 99.6829 | 168.5103 | 141.7655 | 125.3171 | 152.6622 | 167.2552 | 134.3584 | 135.8481 | 163.4277 | 137.0207 | 91.4484 | 95.0207 | 97.9793 | 138.9174 | 108.1932 | 132.1033 | 157.469 | 102.9793 | 329.3171 | 338.2139 | 361.0207 | 345.3451 | 216.3997 | 318.531 | 300.7788 | 314.6209 | 322.2065 | 319.1033 | 325.8348 | 430.8481
755 | 10.2994 | 189.826 | 308 | 379.3018 | 261.0355 | 201.5231 | 268.9302 | 286.6047 | 241.5349 | 194.9184 | 150.0698 | 157.7905 | 214.8367 | 262.71 | 199.9764 | 275.0698 | 320.9882 | 220.7905 | 225.9184 | 225.058 | 273.9065 | 237.9065 | 452.8142 | 455.7798 | 248.0236 | 279.8142 | 333.3373 | 373.5811 | 421.9184 | 155.8142 | 167.4887 | 191.3255 | 211 | 201.3373 | 211.1858 | 253.9538 | 159.2556 | 316.6284 | 402.7562 | 392.2793 | 436.9656 | 184.2213 | 179.3373 | 222.5929 | 193.8142 | 114.8722 | 104.884 | 120.6627 | 212.0462 | 254.5929 | 275.1053 | 335.5231 | 222.3136 | 195.1858 | 207.8142 | 234.2675 | 189.7787 | 225.4415 | 237.4533 | 293.4307 | 254.2556 | 179.7325 | 227.5929 | 230.3716 | 184.4071 | 232.1396 | 309.3136 | 333.0118 | 328.1858 | 206.1976 | 235.0236 | 362.058 | 340.4887 | 240.2438 | 285.4071 | 290.2095 | 431.6864 | 226.2213 | 235.3953 | 251.1396 | 363.9538 | 187 | 226.9645 | 229.6391 | 247.6165 | 213.4533 | 229.9302 | 250.1515 | 154.4189 | 162.7325 | 165.1515 | 158.5704 | 226.5811 | 316.2095 | 400.5231 | 399.8142 | 629.9656 | 313.9195 | 353.6165 | 405.7787 | 598.2202 | 169.1976 | 109.0344 | 90 | 169.1976 | 145.3373 | 136.0698 | 143.7444 | 162.116 | 144.9184 | 141.9764 | 150.5929 | 144.0118 | 89.3491 | 96.9302 | 97.942 | 133.6864 | 104.8142 | 138.9764 | 150.8604 | 100.2675 | 329.6391 | 341.5585 | 368.0118 | 332.6284 | 230.7787 | 318.9882 | 329.7669 | 285.3018 | 360.2911 | 322.7669 | 310.0236 | 442.3255
756 | 10.3104 | 209.8038 | 305.2025 | 360.0708 | 294.8392 | 194.5177 | 257.3279 | 295.4759 | 249.9936 | 185.4759 | 155.7975 | 155.4342 | 187.8392 | 291.2507 | 187.5114 | 278.2797 | 298.8102 | 215.0772 | 213.6785 | 215.0418 | 246.4823 | 214.6785 | 447.257 | 492.2861 | 263.5658 | 287.4051 | 340.6367 | 358.5177 | 426.743 | 170.6785 | 180.4469 | 191.5531 | 209.6013 | 208.3569 | 195.2025 | 264.0481 | 145.0836 | 336.5658 | 415.0481 | 411.5531 | 460.0064 | 194.4886 | 186.0772 | 210.4051 | 198.8874 | 116.9228 | 110.3697 | 114.4823 | 193.8392 | 247.1608 | 308.762 | 316.4886 | 207.6367 | 187.8747 | 213.4469 | 240.8747 | 172.238 | 203.3987 | 234.2316 | 321.7203 | 247.6367 | 186.8329 | 215.9646 | 218.5468 | 189.0418 | 230.1671 | 301.0708 | 332.5304 | 326.1898 | 201.2025 | 253.6431 | 349.8038 | 363.7975 | 231.3088 | 290.881 | 305.7139 | 452.238 | 237.3279 | 236.0354 | 257.2797 | 382.4405 | 183.6431 | 197.6367 | 199.5241 | 240.6203 | 201.5595 | 239.5886 | 258.8874 | 170.6013 | 163.6785 | 175.5658 | 192.7203 | 210.6785 | 335.6303 | 392.119 | 411.881 | 663.6367 | 354.0481 | 361.4469 | 382.0836 | 572.9037 | 170.0772 | 83.1962 | 88.881 | 170.0772 | 144.8038 | 140.9582 | 155.3987 | 158.8683 | 135.7557 | 121.9582 | 149.3152 | 150.5241 | 100.3633 | 95.3987 | 100.6848 | 143.0481 | 113.5241 | 114.762 | 140.8038 | 102.119 | 298.1253 | 352.9709 | 382.6367 | 355.6431 | 213.5177 | 308.2797 | 300.119 | 273.0645 | 368.1735 | 293.119 | 329.762 | 427.4469
757 | 10.3214 | 202.2124 | 298.8761 | 375.7787 | 288.823 | 193.407 | 237.9204 | 281.8848 | 261.2832 | 194.3717 | 168.2567 | 169.9557 | 194.9646 | 264.5044 | 209.9557 | 266.7345 | 283.7257 | 220.0796 | 218 | 225.1859 | 249.9293 | 222.5044 | 415.6461 | 483.7611 | 243.4956 | 270.3717 | 341.5309 | 367.9204 | 455.1415 | 171.7876 | 161 | 205.7433 | 215.9293 | 211.1593 | 186.8319 | 243.0443 | 142.3009 | 324.6726 | 397.8407 | 428.3717 | 445.5044 | 171.4602 | 190.7876 | 205.9293 | 184.4248 | 103.1593 | 87.6726 | 113.2567 | 191.9115 | 249.6726 | 302.708 | 308.3541 | 214.0796 | 208.5487 | 194.8761 | 250.7433 | 178.8761 | 216.6372 | 261.2124 | 323.3363 | 253.7876 | 196.1946 | 207.3806 | 243.4424 | 186.0443 | 206.4956 | 317.3628 | 305.6283 | 331.9291 | 192.2478 | 245 | 341.3363 | 360.4867 | 265.6283 | 283.9115 | 325.0443 | 442.8141 | 218.2124 | 244.0354 | 263.8407 | 386.2567 | 174.708 | 206.1239 | 193.8407 | 189.7168 | 209.1328 | 259.1504 | 240.6283 | 167.8761 | 172.0885 | 160.1681 | 191.4159 | 210.0354 | 358.3274 | 395 | 411.9204 | 670.0796 | 331.8761 | 348.1328 | 372.8761 | 612.5752 | 172.4513 | 93.4159 | 91.5487 | 172.4513 | 134 | 135.4867 | 160.1681 | 186.4513 | 138.3628 | 127.8761 | 164.7522 | 137.8319 | 92.4602 | 100.1681 | 96.0531 | 127.593 | 100.5398 | 116.593 | 135.8407 | 105.5841 | 297.6018 | 332.8054 | 383.823 | 344.0796 | 219.708 | 303.1593 | 296.2832 | 322.8407 | 333.7965 | 302.7168 | 322.2478 | 419.6815
758 | 10.3324 | 203.6091 | 305.5649 | 364.9991 | 270.5216 | 181.9995 | 250.7822 | 227.3909 | 240.1308 | 187.3913 | 178.5654 | 171.3909 | 214.3039 | 260.8269 | 210.4779 | 247.9558 | 307.8702 | 212.6957 | 218.3043 | 252.5658 | 263.3485 | 229.4784 | 456.913 | 508.3476 | 241.6952 | 268.8692 | 323.1303 | 376.2173 | 403.0437 | 163.6957 | 163.4346 | 194.2173 | 226.3052 | 204.6082 | 190.6952 | 252.2606 | 149.3485 | 324.7399 | 403.7831 | 415.3048 | 444.5658 | 172.8697 | 185.739 | 237.9125 | 199.087 | 95.0866 | 101.4346 | 118.0875 | 186.2606 | 268.3039 | 305.6519 | 337.3485 | 210.3476 | 211.6529 | 200.6519 | 231.3048 | 181.3043 | 233.9134 | 266.5654 | 322.3048 | 249.9563 | 170.913 | 214.0437 | 210.261 | 187.6524 | 208.9558 | 299.7822 | 320.8255 | 277.826 | 197.1298 | 234.0442 | 346.087 | 338.3476 | 270.174 | 284.9558 | 329.3913 | 422.4337 | 225.6957 | 244.4774 | 278 | 401.739 | 177.0433 | 198.8264 | 211.0433 | 214.5212 | 231.3039 | 224.3481 | 246.087 | 162.0875 | 177.4351 | 165.1303 | 188.7827 | 204.6952 | 343.9567 | 385.8702 | 423.8692 | 643.2187 | 335.826 | 362.3913 | 385.9558 | 597.6087 | 163.2606 | 95.9563 | 102.8697 | 163.2606 | 132.174 | 113.3476 | 159.0437 | 176.6519 | 119.5649 | 126.6524 | 167.7827 | 132.261 | 91.1308 | 104.2173 | 110.3918 | 136.2615 | 107.6515 | 126.1745 | 143.9134 | 100.9134 | 323.5658 | 357.174 | 366.4346 | 342.1736 | 225.3909 | 290.826 | 284.5649 | 326.3485 | 339.9567 | 325.3909 | 322.8692 | 450.739
759 | 10.3434 | 186.1504 | 317.2665 | 387.1848 | 278.116 | 207.1819 | 263.0501 | 242.9499 | 230.8997 | 187.2665 | 170.7837 | 177.1347 | 227.5172 | 245.0315 | 226.7994 | 268.2507 | 287 | 219.2822 | 216.467 | 233.2163 | 252.6146 | 227.8997 | 454.5673 | 513.1504 | 252.884 | 283.4828 | 330.533 | 383.4327 | 397.7994 | 169.0158 | 166.1504 | 188.7679 | 192.1504 | 213.4513 | 198.4011 | 262.6332 | 133.8653 | 304.6991 | 377.4011 | 397.467 | 432.4986 | 168.9499 | 191.6834 | 250.6175 | 189.4011 | 104.884 | 103.5172 | 96.5172 | 196.3166 | 277.0845 | 308.6175 | 319.0158 | 224.3324 | 188 | 204.5673 | 225.1819 | 188.3324 | 221.2665 | 264.4828 | 307.3166 | 262.7994 | 175.8496 | 209.7493 | 207.7994 | 176.2006 | 228.3009 | 307.6175 | 339.4169 | 294.0315 | 210.4011 | 221.3481 | 347.7994 | 330.8023 | 259.9842 | 298.2851 | 317.235 | 456.8338 | 217.4011 | 267.6175 | 278.3166 | 402.9341 | 180.8338 | 187 | 219.8997 | 244.765 | 243.2507 | 221.6991 | 245.2665 | 156.3481 | 160.351 | 165.5673 | 180.9341 | 211.1347 | 329.7178 | 370.6991 | 425.1848 | 608.4484 | 336.5172 | 363.8496 | 400.8682 | 593.6175 | 176.7994 | 103.4169 | 96.4169 | 176.7994 | 125.467 | 125.4327 | 148.4169 | 185.3166 | 125.2507 | 130.3983 | 157.0845 | 132.3324 | 86.9656 | 107.9499 | 92.6834 | 111.467 | 129.1504 | 115.1819 | 131.2665 | 88.5831 | 294.0845 | 346.9842 | 381.1819 | 348.9341 | 232.0845 | 293.1003 | 302.5989 | 305.4828 | 343.4484 | 328.6017 | 327.6676 | 452.884
760 | 10.3544 | 187.236 | 317.0398 | 363.6976 | 279.1062 | 213.8436 | 254.0929 | 262.4322 | 247.4986 | 180.7905 | 165.3422 | 164.2493 | 211.7773 | 269.6844 | 229.0796 | 249.7773 | 300.8142 | 225.4616 | 208.0398 | 236.3555 | 279.6976 | 227.7242 | 451.9071 | 499.1062 | 247.9469 | 294.2891 | 331.3953 | 395.6313 | 419.8142 | 175.4218 | 158.6844 | 178.236 | 189.9469 | 199.9336 | 198.4455 | 255.1194 | 141.8171 | 311.7374 | 387.3128 | 391.0133 | 442.146 | 172.9735 | 179.4882 | 238.3422 | 182.8673 | 100.6047 | 96.9867 | 89.3289 | 196.6711 | 269.2095 | 291.7374 | 329.6976 | 236.3555 | 187.6711 | 197.9602 | 230.528 | 196.7374 | 218.7374 | 261.4749 | 321.4853 | 256.8569 | 176.0796 | 214.4085 | 224.8806 | 171.618 | 219.2891 | 311.1299 | 345.8673 | 320.6578 | 196.6313 | 245.0398 | 358.6313 | 315.7611 | 253.5782 | 279.2493 | 312.09 | 440.7905 | 221.3924 | 256.3289 | 289.5251 | 379.8303 | 175.3156 | 184.0398 | 220.0531 | 249.91 | 223.1327 | 242.5517 | 257.2095 | 180.6976 | 153.5649 | 158.6313 | 175.2626 | 190.3555 | 314 | 378.3953 | 398.4085 | 631.3289 | 335.9071 | 358.1592 | 382.2891 | 597.7877 | 179.4085 | 92.4351 | 102.2095 | 179.4085 | 124.6047 | 130.0664 | 150.5915 | 176.4616 | 119.9336 | 154.6711 | 144.9336 | 150.6047 | 100.3555 | 110.9867 | 90.6844 | 110.9336 | 119.7109 | 133.3555 | 128.4085 | 94.9602 | 282.2626 | 341.8938 | 389.1592 | 335.6976 | 211.7109 | 296.8938 | 318.0133 | 304.1194 | 359.7507 | 306.3024 | 320.9071 | 443.0133
761 | 10.3654 | 197.5644 | 308.6116 | 358.3176 | 278.6952 | 202.6952 | 246.1416 | 294.588 | 262.4228 | 174.9656 | 161.2704 | 173.9292 | 192.2704 | 265.294 | 218.0236 | 235.0472 | 322.176 | 205.3176 | 200.9764 | 237.0944 | 272.2704 | 221.3884 | 467.706 | 490.5064 | 246.1052 | 295.882 | 324.7296 | 396.2232 | 433.3284 | 157.5644 | 154.6352 | 197.0944 | 204.4828 | 205.5644 | 210.5043 | 232.5644 | 126.8476 | 307.1416 | 421.8927 | 395.4828 | 424.7768 | 176.6588 | 159.412 | 222.6696 | 205.3884 | 108.6352 | 99 | 93.7532 | 193.9528 | 292.4356 | 293.2832 | 317.8348 | 224.8112 | 187.3412 | 212.53 | 213.2832 | 178.8348 | 212.4356 | 259.7897 | 350.706 | 237.8948 | 172.1888 | 208.8004 | 234.8112 | 170.1288 | 222.588 | 330.058 | 375.8948 | 335.9892 | 195.8584 | 246.8476 | 351.7168 | 331.8691 | 272.8004 | 291.6588 | 329.2103 | 414.8348 | 242.8691 | 260.0708 | 299.058 | 361.1416 | 172.1996 | 184.4828 | 217.118 | 223.9184 | 210.8476 | 244.9872 | 261.3284 | 176.3412 | 161.8348 | 165.0236 | 182.9292 | 198.2596 | 303.4228 | 365.588 | 388.3648 | 624.4936 | 352.3884 | 348.3304 | 394.8004 | 636.4828 | 165.6116 | 83.4828 | 109.3992 | 165.6116 | 136.7296 | 126.4828 | 155.5172 | 160.7532 | 131.3648 | 148.5408 | 155 | 152.8348 | 102.4592 | 103.4464 | 91.412 | 122.3648 | 113 | 119.0816 | 123.1416 | 105.7768 | 299.1416 | 337.1867 | 375.5772 | 331.6824 | 208.412 | 308.9056 | 322.1416 | 294.8712 | 353.1416 | 319.8712 | 335.6824 | 439.294
762 | 10.3764 | 196.7768 | 299.7581 | 361.5954 | 297.6883 | 220.2743 | 247.0417 | 292.2419 | 243.0605 | 195.0511 | 164.1303 | 164.1116 | 200.7862 | 265.5349 | 227.7768 | 241.8279 | 310.0605 | 202.2325 | 202.8884 | 240.7581 | 259.5767 | 214.6883 | 469.9395 | 500.0511 | 234.7768 | 282.4651 | 325.7581 | 387.707 | 405.707 | 145.8186 | 153.7675 | 202.5256 | 211.6976 | 199.4744 | 183.8467 | 224.707 | 142.3535 | 326.0698 | 384.293 | 400.9302 | 436.8279 | 170.6976 | 173.7768 | 200.3024 | 205.7581 | 107.0605 | 97.9395 | 105.5954 | 187.879 | 292.8697 | 314.879 | 313.8978 | 213.9489 | 191.1814 | 221.2232 | 233.8186 | 172.7768 | 220.293 | 295.1721 | 337.7394 | 247.8092 | 179.8697 | 212.2232 | 221.121 | 148.0511 | 216.707 | 311.2419 | 382.2743 | 356.2837 | 191.0698 | 251.7487 | 331.1814 | 293.3535 | 277.6371 | 295.2419 | 303.3722 | 409.1303 | 206.4744 | 256.8092 | 273.5256 | 373.3535 | 171.9583 | 193.4651 | 221.6976 | 230.5068 | 215.7487 | 223.7768 | 246.7862 | 170.2837 | 143.7581 | 160.2837 | 183.7163 | 215.4046 | 304.9165 | 380.563 | 404.0792 | 619.5443 | 351.3441 | 359.3859 | 384.7905 | 620.0137 | 156.4046 | 91.0511 | 83.4651 | 156.4046 | 140.2325 | 136.879 | 143.707 | 167.6465 | 139.6559 | 140.8279 | 155.3535 | 148.5443 | 108.7581 | 89.5954 | 101.1814 | 122.1721 | 113.707 | 105.0698 | 123.3348 | 111.1116 | 300.3952 | 316.7956 | 392.1116 | 335.121 | 210.4046 | 310.1908 | 325.1627 | 308.1116 | 355.1022 | 319.3254 | 338.7675 | 451.9071
763 | 10.3874 | 212.4631 | 296.3894 | 369.6342 | 285.7788 | 193.708 | 243.0945 | 295.9764 | 239.1475 | 182.3422 | 169.8554 | 169.4366 | 217.9528 | 256.6371 | 233.587 | 257.9735 | 304.6843 | 207.413 | 214.3894 | 234.0973 | 245.6342 | 191.4395 | 461.4159 | 483.3186 | 247.1711 | 274.1711 | 321.6578 | 375.466 | 400.7817 | 145.8525 | 157.7316 | 203.4867 | 197.2448 | 201.8053 | 193.4424 | 231.4867 | 134.587 | 345.5841 | 376.7817 | 405.5576 | 430.2949 | 168.6814 | 179.587 | 222.4366 | 193.6106 | 105.3422 | 92.3422 | 106.3186 | 184 | 276.1711 | 314.292 | 331.6342 | 224.4631 | 200.292 | 220.1682 | 231.292 | 186.2684 | 216.0737 | 293.6578 | 314.2183 | 237.9499 | 163.9027 | 196.5369 | 223.1711 | 144.121 | 217.2684 | 308.0265 | 365.9499 | 341.8053 | 208.3894 | 230.8053 | 346.5104 | 299.8525 | 266.3658 | 292.7581 | 327.4631 | 421.8053 | 211.7316 | 234.5133 | 271.9263 | 380.5841 | 180.2949 | 192.8525 | 217.121 | 224.1181 | 203.5841 | 230.6843 | 262.1239 | 166.7788 | 137.0973 | 158.6077 | 199.6578 | 216.5104 | 343.5369 | 411.2212 | 422.879 | 622.3687 | 340.2684 | 338.8761 | 343.8053 | 559.3186 | 152.3894 | 85.292 | 80.6578 | 152.3894 | 132.6106 | 141.7788 | 150.121 | 160.4159 | 139.3923 | 146 | 146.1239 | 160.5133 | 101 | 92.5133 | 108.8289 | 115.5369 | 120.4867 | 115.8053 | 122.4839 | 105.4631 | 287.2183 | 337.6371 | 384.6342 | 341.1947 | 204.5605 | 316.0265 | 303.5133 | 307.5841 | 337.3658 | 299.0708 | 334.6843 | 458.6872
764 | 10.3984 | 199.877 | 309.6711 | 366.3535 | 292.6824 | 197.4519 | 249.4498 | 290.1342 | 226.7316 | 175.2439 | 161.5615 | 192.5123 | 205.8904 | 233.4027 | 229.5861 | 269.4385 | 296.1588 | 217.8412 | 221.2439 | 236.3781 | 246.5123 | 196.4519 | 442.0604 | 486.2572 | 250.0492 | 280.0738 | 335.5615 | 355.7807 | 399.0738 | 151.4631 | 159.7561 | 202.4139 | 197.6957 | 198.3781 | 172.0738 | 241.3781 | 134.7449 | 352.085 | 380.3668 | 432.2193 | 405.4631 | 175.5727 | 176.7203 | 239.8412 | 189.7807 | 103.915 | 90.5369 | 104.3422 | 193.8299 | 266.9508 | 316.2193 | 326.085 | 217.1701 | 207.5123 | 236.1947 | 235.4877 | 182.7316 | 208.3535 | 292.4385 | 319.707 | 244.3289 | 175.9887 | 193.6219 | 226.0492 | 169.6353 | 215.6219 | 312.8187 | 367.0358 | 327.0358 | 206.5481 | 239.4764 | 349.7674 | 303.9508 | 262.8412 | 286.3422 | 326.5973 | 422.5369 | 209.5973 | 224.2439 | 282.293 | 384.0604 | 167.5615 | 195.8166 | 222.9754 | 243.499 | 213.8658 | 228.9642 | 252.5615 | 182.7561 | 136.7316 | 168.3043 | 191.2551 | 227.707 | 332.3043 | 400.915 | 404.9262 | 599.2684 | 328.792 | 352.5973 | 340.7561 | 561.8791 | 156.5973 | 87.5973 | 91.1588 | 156.5973 | 128.4027 | 151.3289 | 155.9754 | 149 | 124.7807 | 139.1947 | 137.6957 | 147.9754 | 102.8904 | 85.2684 | 103.3043 | 111.4877 | 120.9262 | 115.4027 | 139.4744 | 105.7316 | 297.6219 | 327.6353 | 385.8904 | 340.8412 | 206.7316 | 289.0604 | 300.4273 | 311.8166 | 334.9754 | 317.3535 | 322 | 432.8545
765 | 10.4094 | 194.1279 | 310.412 | 372.9302 | 282.1229 | 200.8987 | 235.0797 | 278.3372 | 232.3688 | 175.9518 | 175.9518 | 186.0216 | 209.7807 | 241.3422 | 239.1927 | 257.7807 | 312.7591 | 223.1013 | 218.4385 | 233.4867 | 245.8771 | 213.1711 | 435.6362 | 500.5083 | 243.6047 | 286.8023 | 331.6047 | 375.6628 | 411.6578 | 143.6096 | 162.1711 | 194.7591 | 205.7542 | 195.0963 | 186.2193 | 243.5615 | 148.4601 | 339.706 | 392.6794 | 441.6146 | 398.3904 | 170.0747 | 187.8023 | 238.0747 | 194.4385 | 111.6096 | 102.6844 | 120.6578 | 207.2674 | 247.2409 | 306.4867 | 337.1279 | 205.294 | 207.2674 | 222.6578 | 231.0482 | 179 | 207.5133 | 265.5565 | 325.8339 | 233.8289 | 192.7542 | 198.8555 | 232.8771 | 189.5831 | 216.9518 | 344.5831 | 335.7276 | 317.588 | 198.4169 | 248.6794 | 343.6146 | 318.593 | 260.6844 | 293.289 | 326.4651 | 413.2143 | 219.6146 | 220.6578 | 285.5349 | 373.3422 | 180.7807 | 190.6362 | 205.1179 | 228.9302 | 223.3206 | 251.294 | 264.6096 | 179.706 | 131.4385 | 163.3688 | 166.6096 | 217.0482 | 313.706 | 395.3372 | 390 | 601.8289 | 320.9784 | 335.6794 | 343.1711 | 586.0482 | 169.7375 | 85.5133 | 100.3422 | 169.7375 | 132.0482 | 142 | 149.0482 | 153.294 | 121.6312 | 133.8555 | 152.7807 | 145.7326 | 102.0963 | 91.7326 | 96.0266 | 117.1977 | 111.4651 | 126.0747 | 107.0482 | 100.0482 | 300.1229 | 353.0482 | 384.706 | 328.1445 | 198.706 | 275.2193 | 305.5349 | 302.7326 | 325.706 | 307.1445 | 319.2674 | 446.1927
766 | 10.4204 | 212.323 | 300.6239 | 381.1416 | 284.5841 | 190.8584 | 259.5442 | 271.8983 | 243.3761 | 180.208 | 170.9469 | 174.4159 | 211.8053 | 253.8584 | 227.4027 | 268.6638 | 322.3628 | 239.9867 | 214.3894 | 233.2345 | 243.0133 | 212.9867 | 443.5841 | 495.4956 | 223.2345 | 282.4558 | 312.8452 | 378.0752 | 406.792 | 147.4292 | 164 | 202.7522 | 185.1283 | 189 | 186.208 | 231.907 | 132.4027 | 336.6239 | 385.4691 | 451.5044 | 404.6372 | 175.5841 | 172.9867 | 252.4425 | 185.5575 | 102.9469 | 98.7256 | 110.1548 | 199.7787 | 238.4425 | 307.8452 | 348.0752 | 215.6239 | 203.4027 | 214.5708 | 225.1814 | 178.5973 | 213 | 255.0797 | 333.6991 | 237.6372 | 184.6106 | 197.9336 | 230.8186 | 180.8186 | 212.3495 | 310.8539 | 322.2478 | 324.9602 | 203.9602 | 247.5089 | 372.4292 | 329.6194 | 258.7389 | 272.9867 | 340.6372 | 379.7522 | 233.1283 | 219.8186 | 283.4425 | 360.0884 | 176.7655 | 196.5708 | 189.0266 | 237.5442 | 229.7256 | 245.5177 | 262.792 | 168.1681 | 139.0664 | 173.1681 | 176.4691 | 214.4027 | 305.792 | 373.1947 | 407.3142 | 618.9248 | 333.792 | 339.7434 | 349.4292 | 572.1283 | 186.115 | 83.7522 | 98.7655 | 186.115 | 123.7655 | 134.7522 | 149.6239 | 151.5442 | 117.4691 | 133.7389 | 147.9602 | 150.4027 | 88.7522 | 100.4292 | 99.7522 | 127.1814 | 120 | 128.3628 | 119.3009 | 98.6106 | 309.0266 | 350 | 371.5575 | 331.0797 | 200.0531 | 286.4823 | 292.9734 | 318.2744 | 324.2345 | 300.8186 | 322.6505 | 422.323
767 | 10.4314 | 178.8171 | 306.4149 | 365.0954 | 290.1288 | 201.0954 | 254.4691 | 292.3402 | 233.0206 | 186.5644 | 155.2655 | 164.3067 | 212.1701 | 257.8712 | 226.3402 | 276.3608 | 321.0541 | 230.3608 | 204.5309 | 233.5309 | 250.9794 | 206.6804 | 443.3196 | 504.9665 | 238.4691 | 285.4562 | 320.0954 | 358.7139 | 397.9459 | 153.1701 | 168.1495 | 193.2448 | 186.4278 | 201.8634 | 191.7345 | 232.9922 | 142.1288 | 329.9665 | 388.5516 | 417.5516 | 425.0335 | 169.9253 | 173.3196 | 252.0413 | 182.2242 | 100.9587 | 89.884 | 95.8505 | 192.9253 | 256.2989 | 303.4768 | 324.5644 | 213.116 | 195.2861 | 204.2655 | 219.3402 | 182.9794 | 235.8221 | 275.9046 | 318.884 | 253.0541 | 187 | 198.3531 | 228.7758 | 183.3402 | 205.1495 | 290.0335 | 352.5025 | 303.1366 | 203.7139 | 261.6263 | 367.7964 | 301.4484 | 232.9665 | 270.8299 | 336.1366 | 395.5566 | 215.3402 | 224.8299 | 278.0619 | 347.4484 | 173.9794 | 182.5309 | 192.9253 | 221.2655 | 212.5851 | 224.5103 | 263.9046 | 168.0541 | 149.4356 | 161.0206 | 175.8171 | 217.0747 | 323.9175 | 384.8634 | 419.7217 | 613.3893 | 316.232 | 350.4227 | 365.9587 | 556.4149 | 162.0206 | 82.1288 | 84.7758 | 162.0206 | 129.2783 | 134.7886 | 150.0206 | 142.7345 | 137.5644 | 125.8093 | 135.2655 | 154.3196 | 77.5851 | 97.8712 | 87.3402 | 128.3943 | 113.7758 | 124.5644 | 131.6263 | 98.5103 | 311.6804 | 350.4149 | 374.8634 | 347.7552 | 218.6392 | 295.4562 | 301.9381 | 320.8918 | 344.0335 | 316.6185 | 345.6185 | 412.4691
768 | 10.4424 | 177.3619 | 306.5728 | 361.8545 | 285.9105 | 198.709 | 247.4832 | 286.1549 | 232.8358 | 185.3097 | 148.291 | 149.6996 | 196.0467 | 249.6996 | 238.9533 | 256.7463 | 326.7276 | 216.3004 | 207.6716 | 227.6996 | 259.709 | 216.9533 | 458.681 | 481.5914 | 256.2631 | 282.681 | 325.8265 | 370.1455 | 390.5634 | 162.681 | 170.5821 | 194.1455 | 201.6008 | 218.291 | 189.7369 | 254.6008 | 149.8731 | 319.8358 | 385.0989 | 427.3433 | 424.9105 | 166.1455 | 181.8451 | 239.709 | 193.1362 | 112.4366 | 98.4366 | 101.9627 | 188.291 | 259.8825 | 290.2631 | 331 | 207.5541 | 182.5728 | 199.8545 | 226.3992 | 184.4459 | 254.3284 | 279.5728 | 354.7798 | 254.8825 | 179.7369 | 208.4646 | 210.6008 | 167.7556 | 208.4366 | 324.944 | 366.3284 | 293.5448 | 205.319 | 236.319 | 357.1269 | 324.5541 | 240.7798 | 275.4179 | 308.1735 | 418.0467 | 222.8265 | 235.3992 | 274.5168 | 362.8638 | 178.0093 | 177.1269 | 204.0989 | 225.3992 | 218.8358 | 226.5541 | 256.0373 | 166.4646 | 147.2724 | 146.3097 | 161.5448 | 205.0467 | 348.1549 | 381.2108 | 408.6903 | 576.2537 | 310.2351 | 330.2537 | 377.8638 | 574.5168 | 158.4179 | 94.5728 | 84.972 | 158.4179 | 142.4459 | 141.028 | 156.2444 | 152.2724 | 132.8918 | 133.5728 | 137.2631 | 159.4272 | 80.4179 | 88.8451 | 95.681 | 127.8825 | 99.8731 | 128.0093 | 121.2724 | 93.291 | 305.7183 | 340.319 | 385.3097 | 351.1269 | 228 | 285.4179 | 305.056 | 316.3806 | 353.3097 | 325.7369 | 345.7649 | 412.319
769 | 10.4534 | 204.6814 | 307.3186 | 364.3186 | 264.7581 | 204.0767 | 239.3835 | 283.6372 | 254.7464 | 167.528 | 163.1416 | 165.2301 | 187.4278 | 259.5162 | 244.4955 | 260.5812 | 313.528 | 218.3511 | 226.4395 | 225.2094 | 259.8024 | 223.8141 | 463.3304 | 487.944 | 251.9351 | 296.5605 | 319.2212 | 368.1209 | 393.5605 | 166.4513 | 158.0885 | 190.3628 | 196.0325 | 205.0118 | 191.8673 | 242.4395 | 147.8348 | 317.5722 | 404.4395 | 452.5722 | 406.3953 | 170.7139 | 187 | 249.9115 | 180.6165 | 111.6372 | 88.8466 | 111.8466 | 190.8348 | 245.0442 | 304.8348 | 339.3511 | 206.6489 | 180.6814 | 201 | 225.3746 | 181.3953 | 234.6814 | 277.2419 | 369.1002 | 240.4838 | 170 | 185.0885 | 198.8791 | 150.4395 | 217.7464 | 326.584 | 348.4395 | 294.8908 | 194.5162 | 216.7256 | 349.9351 | 345.1859 | 255.5398 | 297.5812 | 303.3511 | 392.2861 | 226.3304 | 226.0235 | 293.1652 | 356.9233 | 163.8908 | 189.2744 | 228.2744 | 228.7699 | 238.1092 | 235.7581 | 232.528 | 157.1534 | 148.1652 | 142.5929 | 185.7464 | 192.9115 | 312.6725 | 361.2301 | 414.1652 | 597.0325 | 327.1652 | 331.6932 | 370.1652 | 601.0767 | 153.7699 | 95.3186 | 94.3628 | 153.7699 | 129.2861 | 141.7464 | 166.0885 | 160.6372 | 126.7906 | 143.1092 | 146.5605 | 153.4071 | 94.6696 | 90.9233 | 103.4071 | 113.9233 | 113.6579 | 123.1209 | 126.1209 | 91.8791 | 288.6165 | 330.8348 | 387.7464 | 350.0885 | 217.8908 | 285.6047 | 297.2626 | 322.4955 | 328.4955 | 321.2744 | 327.5162 | 402.3953
770 | 10.4644 | 203 | 312.1627 | 373.2291 | 281.9101 | 199.8672 | 269.8437 | 290.1627 | 249.7644 | 179.6874 | 184.5482 | 173.8373 | 202.4518 | 260.3854 | 221.03 | 273.5118 | 295.8672 | 222.2227 | 219.3191 | 225.7473 | 258.3555 | 225.0728 | 468.1692 | 496.1563 | 235.3555 | 289.2227 | 308.4583 | 384.621 | 402.4882 | 158.4218 | 147.0964 | 194.2291 | 200.03 | 184.2227 | 192.5417 | 249.7773 | 147.6745 | 306.6146 | 398.6745 | 435.7409 | 413.3555 | 183.8737 | 188.8073 | 248.7045 | 164.6146 | 106.4154 | 86.4218 | 100.8373 | 198.8073 | 254.0064 | 319.5846 | 348.6445 | 205.9401 | 179 | 211.8437 | 212.97 | 176.6081 | 232.5482 | 289.4583 | 327.379 | 245.0364 | 180.3919 | 193.5246 | 217.621 | 159.5846 | 232.6445 | 307.0364 | 336.8008 | 302.3319 | 203.5182 | 213.1627 | 349.1692 | 344.379 | 254.1392 | 296.0535 | 317.1627 | 388.6146 | 218.97 | 213.3619 | 290.1627 | 344.8672 | 159.1328 | 194.1927 | 222.8008 | 221.97 | 225.6981 | 238.4518 | 234.7473 | 173.5846 | 141.0964 | 154.1627 | 190.7045 | 196.2227 | 289.8137 | 382.9401 | 403.03 | 590.1863 | 322.3555 | 327.7473 | 357.6745 | 596.4154 | 154.1992 | 101.5182 | 99.1328 | 154.1992 | 135.1028 | 148.5118 | 150.5782 | 159.03 | 144.7773 | 138.8308 | 145.5482 | 147.7109 | 99.3191 | 89.7109 | 105.8437 | 111.8073 | 131.4154 | 119.2891 | 121.8373 | 100.2291 | 267.6445 | 341.97 | 400.8373 | 345.8737 | 209.5182 | 287.2291 | 312.5782 | 309.8737 | 328.5246 | 321.2227 | 319.349 | 411.1627
771 | 10.4754 | 196.9666 | 319.2488 | 369.469 | 290.2537 | 200.6746 | 267.6843 | 293.0718 | 231.8899 | 199.0383 | 175.646 | 156.5408 | 199.2871 | 244.8613 | 223.8181 | 261.0718 | 298.531 | 214.4641 | 214.6411 | 239.0718 | 270.6746 | 232.1484 | 477.3972 | 489.9617 | 243.0334 | 276.3589 | 314.8613 | 399.5024 | 422.354 | 168.2488 | 149.2488 | 207.177 | 188.8948 | 185.2822 | 174.0718 | 245.9331 | 127.0049 | 306.823 | 400.2822 | 434.3923 | 425.2105 | 186.3589 | 179.8613 | 231.7512 | 170.3923 | 96.4641 | 89.7512 | 101.177 | 206.1052 | 280.7463 | 330.4641 | 345.1436 | 222.177 | 177.1436 | 227.7847 | 236.6362 | 174.2105 | 217.6125 | 294.9331 | 306.9282 | 234.6509 | 185.1101 | 209.5024 | 236.2153 | 170.9282 | 221.2537 | 306.8613 | 330.8181 | 322.3589 | 210.8564 | 224.8899 | 357.0049 | 327.6411 | 274.0383 | 278.1387 | 311.7178 | 411.5645 | 227.3206 | 227.3589 | 296.3206 | 352.6362 | 166.6794 | 187.3589 | 208 | 223.3589 | 208.9233 | 232.9666 | 248.5359 | 193.7463 | 133.0383 | 162.177 | 179.7847 | 182.8948 | 316.4258 | 393.6077 | 401.177 | 591.4209 | 327.7129 | 333.646 | 356.9617 | 574.8613 | 165.2871 | 106.0718 | 93.6843 | 165.2871 | 143.5742 | 133.7512 | 148.177 | 155.7847 | 136.7561 | 126.8181 | 146.8564 | 147.2871 | 95.1052 | 94.3923 | 110.646 | 107.0383 | 115.4307 | 112.7512 | 122.177 | 98.7895 | 278.0669 | 340.1101 | 395.1436 | 353.4641 | 215.9282 | 288.1101 | 307.3923 | 309.5742 | 331.9714 | 310.2153 | 306.3206 | 409.4307
772 | 10.4864 | 178.5663 | 326.8112 | 356.9056 | 269.4719 | 188.7035 | 241.4248 | 280.0899 | 224.3259 | 193.8584 | 173.146 | 152.9101 | 195.9528 | 239.6224 | 220.6607 | 245.7079 | 306.3304 | 225.0045 | 216.6652 | 228.9484 | 270.6135 | 227.3865 | 446.3731 | 501.3348 | 255.2404 | 272.4292 | 292.9484 | 380.0899 | 426.7596 | 165.3304 | 161.0988 | 208.1888 | 209.2021 | 200.2876 | 179.6696 | 243.9101 | 116.764 | 321.0561 | 405.2832 | 428.854 | 426.9955 | 178.618 | 178.9101 | 241.3393 | 165.3304 | 94.1416 | 84.5708 | 105.5236 | 202.4719 | 286.2832 | 321.9484 | 344.4292 | 220.8068 | 174.0472 | 229.0944 | 228.8922 | 186.9528 | 211.9573 | 279.0944 | 333.2493 | 219.5281 | 186.4809 | 202 | 228.7124 | 171.5236 | 214.2876 | 316.3909 | 342.4292 | 323.1932 | 205.854 | 228.7596 | 362.9617 | 319.6607 | 279.3393 | 280.0427 | 300.0472 | 411.0855 | 221.9012 | 213.9012 | 298.5236 | 365.3776 | 168.2876 | 195.3393 | 215.6224 | 222.764 | 223.0944 | 236.0045 | 244.6652 | 177.8451 | 137.8629 | 156.0427 | 176.8068 | 198.4381 | 315.9484 | 375.3259 | 409.3348 | 600.7079 | 316.7551 | 329.2404 | 364.0472 | 552.4719 | 157.1888 | 101.1888 | 82.1932 | 157.1888 | 133.0944 | 137.6224 | 146.8068 | 156.618 | 128.4809 | 142.7168 | 149.9528 | 141.5708 | 97.6652 | 106.0045 | 101 | 107.5752 | 111.382 | 106.618 | 121.7596 | 100.8157 | 286.2832 | 332.9056 | 404.9101 | 343.0427 | 222.7168 | 274.236 | 310.4292 | 304.3348 | 321.9101 | 316.0516 | 299.4719 | 397.7596
773 | 10.4974 | 181.6382 | 315.3618 | 367.7965 | 260.9548 | 175.6834 | 236.8869 | 257.7148 | 217.7286 | 204.3304 | 188.3304 | 166 | 192.6018 | 260.706 | 202.0905 | 256.8781 | 295.5791 | 224.76 | 213 | 221.9322 | 257.9774 | 244.9095 | 442.6382 | 501.5791 | 259.0452 | 265.2035 | 299.6156 | 363.5791 | 416.1131 | 155.0678 | 171.4661 | 212.3078 | 224.3166 | 217.706 | 189.9322 | 261.8869 | 119.0678 | 335.1583 | 406.8191 | 423.4435 | 433.1043 | 172.0905 | 184.6696 | 244.76 | 170.2173 | 96.8643 | 94.7286 | 93.76 | 185.1583 | 259.4749 | 298.3166 | 346 | 203.7374 | 185.2173 | 221.6244 | 210.2173 | 185.0678 | 227.4209 | 280.9095 | 359.0452 | 227.5565 | 203.8643 | 203.9548 | 227.4209 | 168.0678 | 217.0452 | 313.6332 | 347.4209 | 309.9636 | 193.1131 | 235.2173 | 370.3618 | 312.3304 | 269.0766 | 276.3078 | 322.4573 | 384.0678 | 214.4661 | 218.6834 | 293.1131 | 376.0591 | 171.5339 | 206.0905 | 216.1809 | 220.6696 | 218.0678 | 247.4887 | 248.8191 | 154.8417 | 143.2035 | 150.8417 | 169.5113 | 214.6696 | 318.706 | 376.0591 | 420.3304 | 584.5113 | 286.3166 | 339.8869 | 378.1495 | 543.4661 | 160.8191 | 96.0226 | 83.1357 | 160.8191 | 124.647 | 146 | 148.3078 | 147.647 | 143.4209 | 140.2035 | 151.9774 | 139.0226 | 89.1131 | 107.7148 | 103.4435 | 116.5113 | 114 | 109.8643 | 120.8869 | 100.8505 | 295.1495 | 326.2035 | 399.9183 | 349.0817 | 234.8643 | 276.3304 | 318.8417 | 298.2261 | 345.7513 | 308.407 | 299.7513 | 412.0364
774 | 10.5084 | 187.9803 | 314.0295 | 391.0276 | 260.4951 | 193.0059 | 244.0039 | 266.5364 | 226.9941 | 198.4734 | 182.4734 | 157.4833 | 200.0236 | 274.501 | 217.5384 | 262.9567 | 307.0295 | 214.002 | 223.0197 | 231.0118 | 261.5049 | 218.9409 | 460.002 | 516.5364 | 256.499 | 268.0236 | 304.4754 | 380.0394 | 414.0059 | 151.499 | 157.4695 | 217.4931 | 203.9882 | 225.9902 | 190.9961 | 267 | 118.5049 | 319.4833 | 421.0118 | 422.9941 | 439.4774 | 176.5167 | 175.998 | 239.5128 | 183.501 | 107.5089 | 99.4852 | 93.0217 | 173.4911 | 256.0354 | 295.5226 | 344.497 | 208.0295 | 193.9921 | 203.9764 | 222.499 | 172.9823 | 206.9527 | 280.4911 | 361.5089 | 219.9902 | 195.9724 | 206 | 232.002 | 160.4911 | 203.9783 | 312.0532 | 353.0039 | 298.0177 | 207.0374 | 231.9685 | 343.9803 | 323.5069 | 245.5049 | 298.0256 | 325.4577 | 390.5187 | 213.4951 | 225.4852 | 272.4695 | 375.9606 | 165.4911 | 203.002 | 205.4951 | 223.0197 | 220.0098 | 252.5089 | 239.4655 | 153.4833 | 144.0197 | 163.0098 | 166.4951 | 219.5246 | 312.9626 | 371.9527 | 405.4557 | 551.436 | 301.5581 | 320.9527 | 373.9606 | 578.065 | 173.0079 | 95 | 81.5089 | 173.0079 | 123.0098 | 143.9961 | 158.503 | 146.0098 | 134.4754 | 135.0079 | 155.5049 | 130.4852 | 88.5089 | 96.497 | 102.9941 | 110.9902 | 115.002 | 117.002 | 138.0236 | 99.0276 | 300.9803 | 317.501 | 387.0118 | 353.9724 | 215.9508 | 264.4616 | 308.9665 | 290.5049 | 355.497 | 307.5345 | 309.998 | 415.4655
775 | 10.5194 | 180.0531 | 322.8407 | 380.3628 | 266.2124 | 192.4071 | 228.0353 | 275.761 | 220.4071 | 191.1593 | 174.1328 | 143.8672 | 212.5133 | 250.3628 | 212.3628 | 256.3983 | 315.8407 | 231.9381 | 223.761 | 233.4071 | 249.115 | 202.8584 | 452.2743 | 529.354 | 247.2743 | 278.4602 | 303.8053 | 383.5752 | 417.5133 | 149.9734 | 150.2124 | 204.761 | 196.9734 | 215.354 | 188.4867 | 275.2124 | 121 | 307.4071 | 421.8672 | 423.0797 | 396.6902 | 188.5929 | 181.1593 | 232.1416 | 170.6548 | 104.8141 | 97.646 | 106.0531 | 177.7257 | 272.4602 | 309.5664 | 331.7079 | 217.8672 | 200.7788 | 201.239 | 217.3805 | 167.0797 | 205.5841 | 265.7345 | 359.8407 | 216.5398 | 174.8141 | 204.4602 | 214.5221 | 170.885 | 206.8584 | 310.2566 | 360.1328 | 319.3186 | 209.0619 | 228.3186 | 349.3983 | 322.3805 | 244.4071 | 285.3362 | 310.1593 | 420.0178 | 225.3717 | 229.8053 | 248.2743 | 364.2124 | 170.7522 | 187.0619 | 210.6991 | 233 | 224.4867 | 229.7964 | 226.1062 | 153.7257 | 145.7876 | 165.4336 | 170.6726 | 223.2743 | 306.3186 | 373.1505 | 391.7257 | 551.3364 | 320.7345 | 309.3186 | 366.8319 | 567.3716 | 162.115 | 87.8141 | 92.1593 | 162.115 | 125.9469 | 147.1328 | 157.9469 | 144.8407 | 118.4071 | 135.9203 | 158 | 140.9647 | 99.1593 | 92.9469 | 100.5133 | 111.1328 | 116 | 117.4867 | 132.5486 | 105.3009 | 280.7345 | 316.4602 | 388.3805 | 315.876 | 193.0531 | 256.8053 | 283.2743 | 313.531 | 326.7964 | 315.761 | 309 | 392.354
776 | 10.5304 | 188.3068 | 302.2841 | 357.5256 | 274.5256 | 201.6136 | 225.1903 | 266.4744 | 231.1903 | 197.5256 | 165.3353 | 152.1392 | 215.1023 | 235.9346 | 208.9715 | 284.1392 | 309.4744 | 230.6563 | 237.0738 | 226.8466 | 239.7301 | 216 | 453.4602 | 521.8977 | 235.8466 | 275.9489 | 303.9631 | 386.3948 | 396.9773 | 151.1023 | 161.679 | 196.5256 | 204.4091 | 202.6421 | 185.8977 | 264.0796 | 135.1903 | 288.233 | 410.6932 | 434.4091 | 400.1107 | 192.5256 | 180.6932 | 228.4602 | 169.0369 | 102.7301 | 104.5767 | 106.4233 | 184.9489 | 268.3722 | 304.6421 | 357.2641 | 218.7812 | 202.5909 | 214.2045 | 231.9204 | 175.2557 | 210.7074 | 277.0227 | 327.7217 | 221.679 | 167.4744 | 194.0654 | 220.6505 | 173.9631 | 224.2045 | 296.1392 | 344.5029 | 315.7586 | 188.2699 | 236.8466 | 356.1165 | 320.6278 | 245.2045 | 262.0511 | 313.8977 | 426.912 | 234.7955 | 236.2699 | 268.3806 | 357.2841 | 184.2045 | 176.2557 | 234.2926 | 225.6421 | 217.1676 | 216.088 | 236.3068 | 159.3722 | 144.8324 | 157.7443 | 180.679 | 219.7301 | 315.3722 | 388.0654 | 402.6278 | 572.5398 | 315.7301 | 328.8835 | 394.767 | 555.9573 | 154.8324 | 92.5625 | 93.7955 | 154.8324 | 127.679 | 150.4233 | 156 | 158.9715 | 127.088 | 130.8977 | 149.5909 | 146.4375 | 94.4887 | 93.1023 | 97.321 | 123.8835 | 105.4887 | 111.7443 | 116 | 101.679 | 278.3579 | 338.1249 | 366.1307 | 311.3948 | 216.0227 | 273.2557 | 315.4686 | 311.4518 | 312.5625 | 295.4375 | 322.6647 | 385.4233
777 | 10.5414 | 181.6293 | 292.765 | 368.2193 | 269.0836 | 196.8643 | 270.2715 | 260.6214 | 227.3264 | 206.0679 | 150.8486 | 158.0836 | 197.0993 | 254.2193 | 244.7493 | 278.2507 | 302.0078 | 204.2428 | 232.2585 | 227.765 | 253.1436 | 200.94 | 460.9243 | 527.53 | 231.9243 | 270.1593 | 316.59 | 387.4021 | 390.3707 | 148.6971 | 170.3786 | 203.9922 | 224.9086 | 192.235 | 181.8486 | 243.235 | 146.9243 | 299.8172 | 407.6893 | 430.705 | 420.3186 | 189.7729 | 176.0757 | 227.8564 | 171.47 | 114.53 | 91.4778 | 103.9243 | 181.3107 | 263.8486 | 288.8564 | 375.4778 | 203.0235 | 190.1593 | 217.4621 | 239.3186 | 178.9243 | 198.1514 | 283.6293 | 299.1593 | 229.8407 | 169.1514 | 202.6736 | 222.6371 | 179.06 | 216.1671 | 311.765 | 353.9686 | 297.6971 | 207.1279 | 221.6293 | 344.6971 | 312.7807 | 266.7493 | 274.295 | 299.6293 | 386.4178 | 212.175 | 228.235 | 294 | 357.4464 | 181.5457 | 191.7572 | 224.2585 | 205.5536 | 221.2193 | 229.1514 | 237.6971 | 159.1514 | 141.3186 | 154.0757 | 184 | 225.0757 | 338.2793 | 367.6293 | 378.1071 | 563.4621 | 345.2793 | 329.0078 | 407.3864 | 564.175 | 169.0679 | 99.235 | 91.0757 | 169.0679 | 131 | 143.0836 | 152.235 | 172.6971 | 136.9243 | 135.9922 | 143.0757 | 145.6814 | 89.3029 | 98.2271 | 95.0757 | 124.5457 | 110.5222 | 119.9086 | 103.6293 | 100.1593 | 296.8329 | 325.6528 | 368.2872 | 324.235 | 223.1671 | 274.7729 | 327.7964 | 311.3629 | 320.3107 | 284.4621 | 316.7128 | 423.8015
778 | 10.5524 | 173.2006 | 290.4985 | 353.3436 | 252.4469 | 177.4985 | 264.9911 | 278.006 | 210.2493 | 191.5442 | 157.2522 | 150.2493 | 200.9057 | 269.0516 | 237.994 | 259.6991 | 318.006 | 191.149 | 203.2979 | 234.3009 | 269.2522 | 192.4012 | 445.6961 | 524.6475 | 238.1519 | 267.1003 | 312.3436 | 350.3436 | 404.3009 | 153.2522 | 161.7964 | 202.8481 | 209.5929 | 198.9027 | 189.9027 | 245.5015 | 137.1976 | 313.1976 | 396.2463 | 422.6504 | 414.2006 | 192.5015 | 178.6504 | 207.8967 | 179.3039 | 103.9454 | 97.7051 | 104.6504 | 180.1003 | 254.2979 | 295.3039 | 350.8967 | 191.6018 | 183.2493 | 212.0487 | 228.2493 | 172.4985 | 196.149 | 269.149 | 305.4528 | 232.8997 | 171.5501 | 212.0487 | 218.1033 | 167.7934 | 195.5472 | 306.1976 | 385.1519 | 295.1003 | 208.8451 | 206.5988 | 347.0516 | 288.0457 | 261.6445 | 288.4012 | 317.0576 | 386.3068 | 195.4499 | 233.8024 | 291.2493 | 360.1976 | 172.6991 | 199.8997 | 203.5501 | 191.2493 | 211.2949 | 216.6961 | 231.2493 | 182.4558 | 140.1519 | 158.851 | 186.2006 | 210.0457 | 323.6902 | 359.7507 | 356.6504 | 543.7448 | 327.3893 | 325.7507 | 403.7994 | 538.0973 | 179.3009 | 95.4499 | 88.149 | 179.3009 | 135.4012 | 136.3496 | 142.9484 | 160.1976 | 131.0487 | 139.7994 | 142.3496 | 151.7994 | 95.2006 | 93.2979 | 107.003 | 122.851 | 113.6475 | 126.0487 | 102.9027 | 99.851 | 272.8908 | 312.4042 | 416.4588 | 327.6018 | 217.4012 | 264.2979 | 304.7994 | 343.4042 | 313.5988 | 294.4528 | 319.7051 | 407.3863
779 | 10.5634 | 163.7512 | 296.6864 | 366.0591 | 250.3117 | 183.1239 | 254.3108 | 290.3756 | 217.5615 | 174 | 155.5634 | 162.6234 | 210.6883 | 275.6873 | 224.4995 | 252.5005 | 310.1278 | 205.9981 | 203.7493 | 225.1888 | 269.8132 | 197.6873 | 398.5663 | 522.062 | 228.8141 | 270.2498 | 297.4995 | 359.6844 | 382.8151 | 158.3127 | 162.8112 | 186.3136 | 192.4995 | 206.4376 | 192.938 | 242.1258 | 136.1868 | 286.3146 | 391.7493 | 439.1859 | 417.1249 | 182.3766 | 190.6864 | 213.1859 | 166.2526 | 104.6234 | 104.5634 | 104.3127 | 173.1258 | 254.1868 | 292.3766 | 330.8761 | 202.6244 | 186.062 | 221.4986 | 228.8122 | 174.1868 | 208.7483 | 276.1239 | 325.2488 | 235.3746 | 185.4986 | 208.5624 | 238.1239 | 167.123 | 193.7493 | 308.5615 | 371.877 | 297.1249 | 198.4995 | 210.8742 | 366.061 | 314.3708 | 235.4385 | 256.0038 | 331.5634 | 401.938 | 214.6854 | 236.5005 | 274.3766 | 353.5624 | 172.8122 | 185.5014 | 191.0639 | 209.8103 | 222.4347 | 215.1239 | 225.0629 | 168.5034 | 146 | 144.5644 | 173.939 | 191.938 | 312.5605 | 358.6254 | 388.3718 | 596.0552 | 310.1859 | 326.8751 | 413.8112 | 519.8761 | 174.6883 | 93.3127 | 88.9371 | 174.6883 | 146.8742 | 131.0629 | 145.3117 | 149.0629 | 130.3746 | 131.2507 | 147.7493 | 143.2507 | 91.938 | 94.8742 | 114.3127 | 125.4376 | 100.938 | 119.1878 | 118.3117 | 95.1258 | 284.3708 | 341.8732 | 424.4395 | 335.8122 | 210.8761 | 290.0591 | 299.6254 | 313.817 | 313.9371 | 304.6873 | 317.5644 | 404.2459
780 | 10.5744 | 171.0925 | 313.621 | 401.9199 | 268.0694 | 189.2758 | 254.9306 | 271.057 | 226.1495 | 182.0462 | 151.2989 | 168.2527 | 209.5978 | 282.7473 | 216.5054 | 262.2188 | 288.8505 | 211.9538 | 223.3683 | 214.8505 | 245.4591 | 209.3452 | 424.45 | 495.5393 | 240.564 | 290.3914 | 310.7704 | 398.6672 | 381.8167 | 167.9199 | 160.5054 | 199.8398 | 175.3097 | 200.8274 | 179.8043 | 228.5285 | 124.3328 | 299.1618 | 386.6548 | 468.2419 | 403.6317 | 153.1833 | 193.8274 | 210.6317 | 162.5177 | 121.7473 | 94.5516 | 101.8505 | 168.1495 | 267.621 | 312.0355 | 329.8968 | 202.9769 | 185.4022 | 240.621 | 229.8505 | 175.5516 | 231.3452 | 284.5747 | 332.2758 | 226.5054 | 185.6548 | 217.0462 | 248.8737 | 175.678 | 206.4715 | 313.7011 | 356.4253 | 300.8737 | 214.0694 | 218.7242 | 376.4253 | 317.4129 | 239.4946 | 255.0124 | 333.0462 | 392.2527 | 223.1032 | 222.0801 | 270.4715 | 339.6317 | 186.4946 | 172.1263 | 186.1726 | 219.678 | 215.9876 | 228.7473 | 224.8737 | 162.9661 | 149.4484 | 140.7704 | 160.1263 | 184.5516 | 310.9075 | 365.1957 | 401.0801 | 581.1817 | 320.2758 | 329.4484 | 378.7457 | 532.1156 | 171.2989 | 91.4253 | 86.8274 | 171.2989 | 135.7581 | 126.8505 | 147.5516 | 145.4253 | 119.7812 | 139.2188 | 145.5285 | 142.5978 | 90.8737 | 104.4484 | 107.2527 | 120.9769 | 104.4715 | 108.379 | 130.322 | 98.7704 | 312.1263 | 322.8151 | 413.4484 | 313.8613 | 208.1726 | 301.2065 | 290.1032 | 298.2419 | 325.621 | 303.1263 | 322.3914 | 416.1833
781 | 10.5854 | 173.6076 | 325.2832 | 382.41 | 278.7611 | 190.9351 | 246.587 | 266.4572 | 216.4336 | 192.1092 | 172.9588 | 166.3481 | 189.5191 | 277.0206 | 225.6107 | 269.6519 | 283.3038 | 220.6756 | 238.1092 | 216.3481 | 240.9794 | 212.8908 | 442.8672 | 473.587 | 251.5428 | 285.2153 | 326.8053 | 388.6932 | 385.0206 | 168.3687 | 152.587 | 196.0412 | 177.6107 | 181.1504 | 165.3687 | 247.0679 | 133.7198 | 308.2596 | 396.0236 | 466.3244 | 399.4572 | 158.7847 | 171.8023 | 219.3717 | 174 | 121.3038 | 91.413 | 95.7168 | 174.8702 | 263.4336 | 319.1504 | 338.5221 | 188.2596 | 193.7404 | 227.0412 | 226.0649 | 175.3481 | 232.5428 | 286.174 | 310.4542 | 241.4809 | 180.9351 | 230.0443 | 259.2183 | 168.6519 | 219.0443 | 291.4542 | 381.2419 | 294.7817 | 200.6932 | 213.5428 | 382.4572 | 324.59 | 245.0649 | 273.2389 | 333.7168 | 400.9145 | 215.0649 | 226.9145 | 264.2596 | 327.826 | 176.8023 | 175.8702 | 211.1328 | 200.3244 | 197.174 | 227.7168 | 238.7404 | 171.7168 | 145.5428 | 153.2832 | 162.1092 | 202.5458 | 305.4572 | 343.8229 | 397.6962 | 510.4542 | 333.0885 | 331.413 | 377.6992 | 532.9115 | 154.8023 | 79.2596 | 94.1534 | 154.8023 | 123.587 | 118.9557 | 152.0443 | 146.7611 | 115.8702 | 144.3038 | 137.6519 | 139.5428 | 81.8466 | 94.6725 | 103 | 104.4985 | 118.8053 | 114.9145 | 110.9321 | 99.5428 | 300.6076 | 308.9794 | 411.8908 | 325.3511 | 203.1947 | 272.8023 | 282.0649 | 303.3687 | 320.8466 | 304.5221 | 315.4542 | 400.0649
782 | 10.5964 | 177.7876 | 318.2124 | 357.6062 | 275.2055 | 193 | 252.3938 | 276.9931 | 237.767 | 186.6096 | 181.6062 | 168.5993 | 194.5753 | 264.0069 | 222.4179 | 279.9863 | 287.7911 | 209.4247 | 214.6302 | 222.7945 | 244.4041 | 208.8014 | 466.9554 | 497.3732 | 214.0378 | 267.2055 | 308.4282 | 359.2158 | 381.5959 | 152.2124 | 157.7945 | 211.9657 | 197.1918 | 194.5718 | 157.0034 | 254.411 | 141.2089 | 311.3869 | 402.0069 | 459.5959 | 398.6062 | 155.8185 | 171.5856 | 233.5993 | 175.1986 | 103.0172 | 88.0034 | 96.7945 | 189.7876 | 260.7945 | 305.6027 | 329.0137 | 203.9725 | 182.822 | 213 | 213.2124 | 162.0172 | 223.2055 | 278.6096 | 312.7808 | 250.2089 | 184.1986 | 229.0069 | 248.8185 | 168.7979 | 218.6062 | 306.9657 | 361.2433 | 283.6062 | 205.3766 | 219.7876 | 397.7876 | 320.0275 | 238.2055 | 272.5993 | 319.2124 | 386.0275 | 201.6131 | 225.8117 | 263.7911 | 354.5684 | 164 | 199.178 | 200.8288 | 196.5924 | 200.9966 | 250.9691 | 242.2055 | 175.7911 | 137.4041 | 144.4144 | 176.3869 | 206.8117 | 295.6165 | 333.589 | 390.8117 | 522.9691 | 317.2295 | 315.4179 | 377.8254 | 511.0137 | 150.9897 | 76.9931 | 85.2192 | 150.9897 | 116.2089 | 126.5856 | 160.5959 | 159.3869 | 139.178 | 135.6062 | 143.7911 | 135.5993 | 72.2021 | 93.1918 | 104.1986 | 101.5924 | 108.2192 | 119.8048 | 116.1746 | 95 | 297.5959 | 298.6199 | 397.0137 | 331.0172 | 200.5959 | 252.8082 | 280 | 309.3869 | 298.0172 | 299.8082 | 323.1746 | 387.2124
783 | 10.6074 | 205.1829 | 313.4464 | 368.0669 | 280.5625 | 185.0491 | 257.8348 | 292.6205 | 231.3126 | 178.5536 | 171.884 | 151.2635 | 211.4464 | 260 | 250.5311 | 273.3215 | 296.058 | 206.6205 | 217.7945 | 215.6027 | 261.5713 | 222.0669 | 460.585 | 516.6696 | 214.6294 | 272.5625 | 319.5222 | 353.058 | 383.3884 | 160.2901 | 172.0089 | 217.3215 | 212.3973 | 201.8259 | 148.884 | 239.4375 | 131.7188 | 298.8171 | 388.2143 | 419.7993 | 399.2812 | 166.4061 | 179.3884 | 216.8751 | 174.7768 | 100.3393 | 87.8348 | 96.3304 | 183.5447 | 265.8348 | 300.942 | 332.0089 | 218.1652 | 195.6874 | 198.9331 | 226.7945 | 180.7454 | 220.6116 | 276.6696 | 318.4956 | 243.7768 | 167.8751 | 214.6027 | 213.5358 | 171.2232 | 212.5536 | 308.0403 | 358.0178 | 292.2321 | 204.9331 | 224.5536 | 383.5939 | 304.6116 | 263.1338 | 288.2901 | 345.0266 | 388.9597 | 209.9017 | 212.2724 | 267.7768 | 353.4287 | 177.4553 | 197.3215 | 188.2812 | 205.8928 | 202.3884 | 235.9734 | 245.116 | 182.8348 | 136.8348 | 144.5625 | 165.0403 | 189.6027 | 295.1741 | 362.0178 | 396.2321 | 530.6027 | 303.058 | 325.1829 | 387.4642 | 511.5625 | 155.1652 | 91.3973 | 86.2321 | 155.1652 | 125.6785 | 130.1072 | 177.6785 | 157.2143 | 142.2143 | 135.6696 | 155.7276 | 134.7768 | 88.1249 | 94.942 | 108.6696 | 125.5713 | 116.5713 | 128.0089 | 127.3304 | 96.8348 | 293.2724 | 324.3079 | 381.0491 | 321.6116 | 195.6607 | 259.0089 | 292.8437 | 310.2724 | 320.415 | 283.9911 | 334.3304 | 414.8614
784 | 10.6184 | 203.0265 | 306.2655 | 353.292 | 253.5574 | 189.3584 | 257.7522 | 260.6946 | 247.4426 | 167.6416 | 160.5133 | 154.9735 | 206.1372 | 264.9912 | 242.5574 | 262.7522 | 299.8717 | 207.761 | 223.5133 | 239.5709 | 265.2655 | 204.7964 | 456.8452 | 509.6416 | 224.0088 | 269.8894 | 314.5398 | 354.3761 | 399.2213 | 173.1195 | 169.6416 | 214.239 | 198.4115 | 181.6504 | 189.2966 | 240.8628 | 133.9912 | 299.7257 | 415.6903 | 414.2124 | 375.7964 | 193.7257 | 160.907 | 219.7257 | 174 | 99.3849 | 89 | 93.3761 | 173.7522 | 262.6327 | 291.5133 | 342.7434 | 200.7787 | 188.292 | 199.3584 | 243.7434 | 190.8894 | 219.7522 | 295.2213 | 311.2566 | 228.6504 | 173.8452 | 185.5398 | 222.3319 | 183.2301 | 215.3673 | 292.8805 | 350.1637 | 322.4602 | 194.1283 | 211.7699 | 355.6504 | 300.6327 | 252.3008 | 295.5044 | 351.6504 | 378.5398 | 201.907 | 213.6151 | 238.3008 | 325.4026 | 162.9159 | 183.0088 | 197.8628 | 206.5044 | 195.761 | 224.9823 | 243.3849 | 167.7787 | 131.1372 | 148.1283 | 155.4956 | 191.7345 | 312.2301 | 358.5309 | 396.5133 | 525.8717 | 309.9912 | 346.1106 | 365.5663 | 536.3408 | 157.1195 | 98.6239 | 88.3849 | 157.1195 | 126.8894 | 141.9735 | 165.7876 | 161.6062 | 149.7257 | 132.3849 | 153.1372 | 129.0088 | 91.5133 | 88.6327 | 115.9912 | 120.5309 | 107.1637 | 125.0177 | 131.8628 | 86.7699 | 269.6592 | 301.2079 | 388.4779 | 309.5221 | 198.4867 | 275.9823 | 310.9823 | 307.8717 | 334.761 | 303.2036 | 330.1283 | 401.438
785 | 10.6294 | 206.7237 | 303.5447 | 349.8152 | 235.2724 | 203.9066 | 257 | 259.9027 | 240.9164 | 182.358 | 143.2763 | 169.7257 | 209.6342 | 245.0973 | 231.2704 | 267.7257 | 293.3658 | 204 | 213.2743 | 221.7374 | 249.821 | 208.7218 | 468.9086 | 526.2665 | 225.4533 | 264.3638 | 303.5447 | 361.6342 | 394.3677 | 161.0039 | 169.7257 | 202.0953 | 183.4553 | 174.9086 | 195.642 | 240.5467 | 134.4553 | 312.4533 | 425.1848 | 416.1848 | 375.9047 | 180.3697 | 176.0836 | 238.179 | 174 | 109.9961 | 89.6362 | 96.8171 | 179.3619 | 260 | 293.998 | 317.644 | 201.8152 | 184.179 | 194.821 | 223.7335 | 158.644 | 217.7276 | 261.1031 | 331.2665 | 223.1809 | 162.3697 | 172.6362 | 210.9164 | 169.642 | 223.7257 | 300.5428 | 361.1751 | 314.3697 | 201.2704 | 218.3599 | 333.0039 | 317.0856 | 248.3599 | 282.5486 | 331.5486 | 406.9883 | 215.1751 | 224.6342 | 269.3502 | 336.358 | 169.358 | 191.4514 | 215.3599 | 212.6342 | 189.4553 | 225.2743 | 244.4533 | 164.998 | 137.8152 | 147 | 155.0914 | 183.6401 | 309.4572 | 337.1848 | 372.9144 | 501.5525 | 309.1829 | 335.0953 | 372.9008 | 529.5506 | 149.4572 | 93.9105 | 81.1829 | 149.4572 | 127.0895 | 136.3677 | 142.1848 | 147.642 | 146.5486 | 130.9086 | 161.0875 | 129.8171 | 92.7257 | 80.2743 | 99.9144 | 108.0914 | 125.8093 | 115.8191 | 118.823 | 104.8113 | 284.7198 | 315.6245 | 371.8249 | 305.8171 | 210.6342 | 275.6381 | 289.644 | 307.0914 | 321.4572 | 290.3716 | 341.0875 | 385.0895
786 | 10.6404 | 189.3579 | 314.0787 | 364.4302 | 239.2424 | 201.8667 | 261.5393 | 274.9454 | 223.4061 | 185.5698 | 158.6967 | 160.6789 | 212.0546 | 227.4607 | 228.2183 | 249.6002 | 289 | 218.915 | 203.5152 | 216.5635 | 239.4607 | 214.1637 | 461.5698 | 486.4733 | 234.4848 | 265.297 | 319.2665 | 369.2424 | 389.8909 | 154.9454 | 161.3274 | 188.4607 | 186.5393 | 200.6421 | 198.915 | 255.5089 | 122.6244 | 322.1333 | 396.3033 | 427.5089 | 394.0787 | 182.2665 | 184.5152 | 219.3579 | 162.3274 | 111.5152 | 96.4848 | 105.4848 | 168.085 | 251.5698 | 301.2424 | 323.3998 | 191.1396 | 178.3274 | 205.8604 | 217.4848 | 155.9695 | 231.915 | 235.3515 | 341.4061 | 222.4061 | 175.9391 | 187.915 | 195.3515 | 167.7272 | 218.5698 | 320.915 | 360.7335 | 304.5939 | 207.297 | 222.7576 | 336.6726 | 324.1091 | 234.6002 | 277.297 | 334.7272 | 414.7881 | 206.9518 | 221.8667 | 284.0305 | 333.085 | 150.5216 | 185.0305 | 211.9759 | 192.3579 | 207.4543 | 212.2728 | 232.085 | 154.085 | 149.8363 | 144.4061 | 180.5876 | 192.2119 | 318.915 | 319.9759 | 359.4061 | 533.0419 | 304.4061 | 320.1637 | 388.703 | 545.5876 | 131.6789 | 82.5698 | 78.3515 | 131.6789 | 135.1878 | 118.9213 | 147.0241 | 155.4543 | 132.8667 | 142.3756 | 162.8122 | 145.618 | 86.9213 | 86.7272 | 101.9695 | 104.4061 | 121.9518 | 114.6485 | 113.5393 | 101.4911 | 276.0063 | 332.2183 | 343.7335 | 330.6967 | 205.2728 | 275.8909 | 290.2119 | 293.6789 | 321.8363 | 279.1878 | 349.297 | 393.1878
787 | 10.6514 | 206.3953 | 317.6785 | 344.5516 | 254.8761 | 187.4277 | 258.7139 | 270.7139 | 213.41 | 191.5723 | 159.7669 | 167.8761 | 191.1769 | 256.7169 | 218.7139 | 243.9469 | 281.7315 | 218.41 | 194.0531 | 220.3746 | 254.18 | 185.8908 | 430.5693 | 487.7346 | 213.5516 | 279.2154 | 325.0177 | 352.4985 | 373.4985 | 159.3038 | 160.2861 | 184.6785 | 191.6431 | 204.7492 | 200.3923 | 253.7669 | 116.3392 | 320.7139 | 397.2154 | 460.1269 | 373.23 | 191.3215 | 190.9115 | 193.4454 | 165.2508 | 112.6254 | 95.3746 | 107.6785 | 161.9823 | 256.2508 | 286.4808 | 332.3569 | 202.4838 | 166.0531 | 200.4631 | 221.6608 | 164.3215 | 232.7315 | 258.1269 | 328.1062 | 219.0177 | 180.0885 | 207.8938 | 204.2508 | 170.3569 | 222.59 | 333.6254 | 374.1446 | 296.7492 | 202.0531 | 222.3215 | 362.1623 | 304.82 | 240.18 | 278.6608 | 325.4631 | 396.7492 | 200.2861 | 226.59 | 285.59 | 326.3215 | 162.413 | 173.3746 | 182.8731 | 196.18 | 202.1416 | 208.9823 | 217.3923 | 177.7169 | 151.6785 | 164.8054 | 163.9439 | 200.3392 | 307.8377 | 341.7523 | 371.8761 | 516.3716 | 306.3038 | 292.5516 | 380.0708 | 543.4808 | 147.4661 | 83.2861 | 82.6254 | 147.4661 | 134.6962 | 129.8584 | 141.1062 | 161.3746 | 133.6254 | 122.23 | 150.7492 | 147.0531 | 89.2685 | 93.9823 | 104.3746 | 112.9115 | 108.6785 | 129.5369 | 108.7315 | 85.3923 | 278.8584 | 306.8554 | 357.1446 | 319.2123 | 208.59 | 267.4277 | 274.5516 | 302.1977 | 320.3746 | 273.41 | 332.1593 | 395.3392
788 | 10.6624 | 197.0968 | 302.1927 | 348.1534 | 268.0575 | 192.0959 | 270.8073 | 287.5187 | 228.5187 | 172.7699 | 162.0767 | 172.9808 | 215.9991 | 278.3845 | 224.0767 | 259.1151 | 303.5763 | 216.0192 | 207.1534 | 213.9616 | 273.769 | 204.6338 | 449.9799 | 484.8274 | 203.6922 | 253.0201 | 299.0968 | 341.6539 | 357.2694 | 150.231 | 156.2694 | 166.5005 | 198.3845 | 182.5005 | 183.5388 | 245.9808 | 120.7114 | 299.1543 | 401.9808 | 452.4813 | 374.8073 | 157.674 | 181.8658 | 196.4228 | 174.7114 | 106.9233 | 102.4228 | 93.5388 | 156.2694 | 244.8466 | 311.653 | 334.3653 | 205.5963 | 179.1534 | 197.7114 | 248.9224 | 155.5772 | 222.2694 | 257.2119 | 320.6539 | 230.1151 | 181.0575 | 198.5196 | 220.442 | 171.0192 | 220.2694 | 296.2895 | 378.2694 | 294.0192 | 207.7498 | 231.7498 | 365.9425 | 310.1342 | 253.7114 | 270.9233 | 320.6922 | 375.1735 | 213.769 | 226.9616 | 271.8274 | 298.0584 | 165.558 | 171 | 194.5571 | 200.9616 | 192.6539 | 203.2694 | 209.9616 | 178.5196 | 151 | 149.7507 | 149.3461 | 188.558 | 286.558 | 365.4228 | 401.211 | 528.653 | 313.3845 | 300.8648 | 357.1543 | 544.4228 | 146.2119 | 90.7114 | 78.9425 | 146.2119 | 132.3269 | 130.5963 | 135 | 162.3653 | 140.7114 | 128.5187 | 158.1151 | 139.9616 | 85.5963 | 102.4037 | 95.9425 | 115.9808 | 103.9616 | 108.7315 | 105 | 87.3845 | 276.9041 | 283.8849 | 360.5963 | 327.1918 | 208.2886 | 266.0384 | 266.7114 | 299.231 | 332.1342 | 287.8457 | 298.7699 | 408.4612
789 | 10.6734 | 193.9096 | 333.3796 | 358.0561 | 267.5732 | 194.9439 | 246.4739 | 268.9006 | 222.9223 | 179.5045 | 159.1465 | 160.349 | 202.8446 | 252.2153 | 201.2713 | 263.6293 | 296.8102 | 214.2586 | 206.0904 | 231.8752 | 253.642 | 193.2713 | 439.2713 | 494.5045 | 228.3019 | 253.4484 | 301.3924 | 336.2025 | 382.7847 | 166.2459 | 178.358 | 174.4484 | 197.5732 | 187.707 | 201.6726 | 242.2586 | 128.4828 | 313.9873 | 389.349 | 392.2841 | 375.1465 | 171.1554 | 162.6637 | 211.2803 | 177.6854 | 116.0217 | 111.1121 | 97.9656 | 172.1898 | 244.5389 | 305.3834 | 322.293 | 200.6293 | 172.6076 | 208.9096 | 237.3274 | 148.9439 | 214.8879 | 259.2242 | 324.7975 | 231.8879 | 186.0561 | 204.707 | 219.1465 | 158.293 | 192.3274 | 296.8752 | 354.4395 | 256.6204 | 181.8446 | 231.2025 | 347.2369 | 300.5516 | 245.0344 | 254.6637 | 322.6854 | 357.4051 | 213.5172 | 207.1809 | 260.9439 | 327.1771 | 162.0561 | 165.5172 | 185.4395 | 185.293 | 196.1121 | 217.1338 | 227.1898 | 165.5172 | 144.8318 | 155.5045 | 154.1121 | 186.4268 | 294.0217 | 352.8662 | 392.8102 | 520.3274 | 292.0127 | 312.3707 | 356.2242 | 542.1465 | 152.3363 | 77.2369 | 68.4611 | 152.3363 | 132 | 133.1682 | 142.5389 | 153.7197 | 137.5172 | 127.0344 | 155.0904 | 146.2242 | 90.2242 | 93.6637 | 91.6293 | 110.2025 | 102 | 115.5605 | 112.5389 | 107.1338 | 295.3019 | 289.651 | 367.2803 | 328.0904 | 194.349 | 274.8535 | 271.0561 | 300.8535 | 341.0561 | 283.293 | 328.8064 | 383.4739
790 | 10.6844 | 202.1858 | 297.2816 | 374.3481 | 250.6519 | 212.1387 | 250.8363 | 265.767 | 211.1165 | 200.5575 | 167.1623 | 161.9764 | 201.5575 | 272.4868 | 212.7198 | 235.0943 | 285.2094 | 210.2094 | 202 | 213.792 | 257.3481 | 181 | 430.4882 | 507.1858 | 238.3024 | 272.1623 | 308.6976 | 345.1623 | 375.7684 | 183.6741 | 185.1165 | 186.1858 | 198.7906 | 182.8377 | 204.6283 | 262.6269 | 119.8377 | 331.9764 | 365.8613 | 411.5561 | 355.9542 | 182.9071 | 168.8599 | 213.9071 | 179.3953 | 116.4189 | 98.3495 | 93.9307 | 182.3953 | 236.1401 | 290.1165 | 297.1637 | 202.7906 | 189.0222 | 217.1858 | 221.8142 | 170.3245 | 207.4189 | 249.0472 | 324.2094 | 222.3259 | 164.6755 | 208.9071 | 232.0458 | 157.5811 | 227.0443 | 317.1623 | 373.6033 | 269.6962 | 203.5797 | 217.8377 | 327.4425 | 291.6047 | 230.9307 | 274.1151 | 306.9542 | 406.0208 | 201.233 | 219.2316 | 262.7906 | 358.8599 | 165.7906 | 188.1151 | 161.3495 | 199.2316 | 200.0929 | 226.3953 | 233.9071 | 168.5811 | 132.233 | 158.4189 | 156.6976 | 190.7906 | 309.0693 | 356.1623 | 355.3967 | 532.7198 | 291.4646 | 313 | 357.2094 | 520.1637 | 161.4882 | 83.2552 | 80.3481 | 161.4882 | 139.6741 | 124.8377 | 134.1401 | 150.3953 | 135.6976 | 124.7906 | 162.8599 | 152.7906 | 90.5118 | 92.8835 | 98.6741 | 121.9528 | 129.208 | 115.9307 | 118.0929 | 101.7448 | 277.3967 | 296.3953 | 356.6519 | 318.4189 | 211.3245 | 266.1401 | 301.9986 | 297.0236 | 330.8377 | 265.8377 | 349.3953 | 380.8599
791 | 10.6954 | 173.4729 | 294.588 | 367.5113 | 248.2596 | 187.3431 | 236.2822 | 274.2596 | 257.9853 | 198.9007 | 159.9311 | 172.8093 | 207.7099 | 253.6332 | 230.8093 | 235.7788 | 278.3205 | 226.7483 | 216.9086 | 202 | 271.8093 | 193.7788 | 428.4503 | 502.6106 | 233.7404 | 285.5192 | 300.4808 | 338.6411 | 353.5113 | 162.1524 | 143.9537 | 214.9774 | 197.8702 | 191.4887 | 197.4503 | 235.7934 | 114.2901 | 338.5497 | 375.0384 | 425.0305 | 349.7099 | 175.6106 | 166.1907 | 212.2901 | 172.1907 | 104.0609 | 88.4503 | 93.5497 | 179.4503 | 233.8397 | 288 | 308.8781 | 208.9695 | 169.1828 | 212.6106 | 211.4808 | 189.9391 | 211.3894 | 238.7404 | 308.0914 | 223.2596 | 162.0993 | 195.2212 | 214.1524 | 161.4199 | 207.9537 | 289.3431 | 353.3431 | 298.6185 | 187.763 | 210.8702 | 352.5271 | 286.7404 | 231.2596 | 274.351 | 315.6185 | 387.8239 | 218.298 | 218.7709 | 261.8702 | 345.5418 | 170.5497 | 193.3205 | 187.6569 | 199.4808 | 213.0689 | 213.5113 | 237.9695 | 166.7404 | 137.2291 | 170.1987 | 152.7404 | 172.1219 | 313 | 362.4199 | 394.1151 | 508.2133 | 312.3284 | 316.5497 | 368.0689 | 531.1682 | 155.9007 | 85.4503 | 88.4199 | 155.9007 | 135.1907 | 120.7099 | 143.1987 | 160.2291 | 133.8702 | 120.3205 | 150.2517 | 134.8318 | 84.7404 | 101.3894 | 91.351 | 111.6716 | 121.8318 | 116.9695 | 120.4199 | 91.0305 | 284.1682 | 296.2901 | 402.535 | 327.3589 | 198.9921 | 273.0689 | 304.351 | 289.7404 | 330.2596 | 309.9853 | 337.9311 | 388.8397
792 | 10.7064 | 174.7222 | 316 | 368.5 | 263.7222 | 194.9445 | 252.2778 | 264.4444 | 238.2777 | 182.2778 | 159.3333 | 163.7222 | 196.4444 | 259.5 | 218.8333 | 246.7778 | 265.8889 | 223.8889 | 213.6111 | 210.6667 | 267.0555 | 206.2222 | 432.7778 | 468.2222 | 232.7222 | 273.8333 | 292.6667 | 350.6111 | 352.3333 | 154.1667 | 150.3889 | 207.2222 | 195.5556 | 187.6111 | 201.0556 | 219.5556 | 118.3333 | 327 | 410.1667 | 410.7222 | 336.2778 | 163.6111 | 175.2778 | 210.5556 | 179.8334 | 92.7778 | 78.3333 | 107.2778 | 155.6111 | 251.6111 | 312.5556 | 327.8334 | 208.8333 | 175.7778 | 206.3889 | 212.3333 | 179.5555 | 203.8889 | 240.6111 | 319.3334 | 213.4444 | 170.0556 | 179.8889 | 210.5 | 165.6111 | 195.6111 | 286.8334 | 302.4444 | 289.8333 | 175.7222 | 246.1112 | 359.2222 | 293.6667 | 227.9444 | 270 | 309 | 381.8334 | 197.3888 | 220.7778 | 261.7222 | 320.6666 | 174.1667 | 173.6666 | 184.3889 | 200.3333 | 212.9444 | 228.2223 | 229.8889 | 153.4444 | 151.8334 | 175.2778 | 165.4445 | 195.0556 | 286.9999 | 338.4444 | 386.1111 | 508.4445 | 307.4444 | 326.6667 | 359.2778 | 492.7777 | 167.4445 | 80.3889 | 82.5 | 167.4445 | 139.9445 | 124.6111 | 143.9444 | 161.1111 | 134.4444 | 132.4445 | 150.2222 | 140.7222 | 90.9445 | 91 | 95.6667 | 104.2778 | 108.2222 | 119.7222 | 109.4444 | 97.6667 | 304.2778 | 306.8334 | 392.2222 | 321.8889 | 176.2778 | 293.1667 | 268.2222 | 295.9445 | 322.6111 | 319.8889 | 294.7221 | 399.3889
793 | 10.7174 | 180 | 315.8584 | 368.0619 | 252.8407 | 184.6371 | 257.5929 | 243.8407 | 235.4868 | 181.4071 | 149.9823 | 155.3274 | 183.1858 | 264.7965 | 214.469 | 221.823 | 280.3629 | 232.4868 | 190.9026 | 200.0442 | 251.5132 | 194.3097 | 433.531 | 481.708 | 210.23 | 273.8761 | 296.1416 | 338.6371 | 365.7522 | 135.9026 | 154.5929 | 203.6726 | 190.5929 | 186.9381 | 185.3716 | 227.8142 | 134.6903 | 296.292 | 351.9556 | 425.8319 | 392.6992 | 155.5929 | 161.6371 | 204.8584 | 186.9381 | 99.5487 | 79.4071 | 102.4513 | 155.0797 | 271.2213 | 301.4336 | 337.8761 | 204.3274 | 185.9381 | 205.7345 | 247.7877 | 162.9823 | 188.9823 | 244.9381 | 319.3894 | 219.2832 | 160.2478 | 186.2832 | 215.9381 | 171.4071 | 197.7345 | 286.5929 | 341.6196 | 275.9203 | 178.2035 | 222.5397 | 355.531 | 295.531 | 250.9735 | 275.1416 | 301.0619 | 339.7256 | 194.8142 | 235.4868 | 247.3097 | 321.3451 | 161.0442 | 171.4071 | 202.239 | 213.0177 | 204.3894 | 236.7345 | 223.7965 | 164.4248 | 167.0177 | 174.2655 | 177.6106 | 168.0707 | 315.9293 | 347.3629 | 345.6194 | 491.9646 | 300.7965 | 309.4336 | 376.77 | 483.0797 | 159.0442 | 93.6903 | 86.6106 | 159.0442 | 124.6371 | 118.6549 | 144.2035 | 149.7168 | 134.2655 | 131.3894 | 153 | 144.531 | 106.4868 | 91.1416 | 90.9203 | 106.2035 | 104.531 | 110.4513 | 110.1416 | 101.7345 | 305.3274 | 320.5487 | 367.5132 | 302.5752 | 199.6461 | 268.8849 | 288.3186 | 307.0797 | 313.8584 | 293.23 | 303.5045 | 359.6636
794 | 10.7284 | 207.6318 | 322.2149 | 350.5703 | 244.0128 | 200.4042 | 247.7851 | 272.3531 | 236.266 | 177.0256 | 152.7213 | 156.2404 | 198.6701 | 271.4681 | 215.7468 | 229.9233 | 280.2787 | 198.9129 | 177.2787 | 210.6829 | 246.2532 | 208.4169 | 441.9617 | 487.266 | 229.6318 | 262.5575 | 277.8362 | 333.4936 | 369.2532 | 143.1893 | 155.2404 | 199.7724 | 191.9872 | 193.9744 | 182.734 | 213.0767 | 113.8618 | 310.151 | 359.1126 | 378.4832 | 384.381 | 172.6701 | 158.734 | 216.4425 | 190.9872 | 101.5064 | 89.2149 | 105.7213 | 140.8235 | 270.0256 | 271.5958 | 304.9001 | 200.7596 | 198.9489 | 200.7724 | 245.0639 | 162.734 | 210.3914 | 248.9872 | 324.4681 | 223.7468 | 166.4553 | 170.5831 | 211.0256 | 191.6701 | 198.7468 | 287.2404 | 325.1533 | 303.619 | 216.3403 | 231.4042 | 349.7724 | 256.9129 | 245.0639 | 265.0511 | 287.3043 | 355.1126 | 209.2021 | 230.2915 | 224.8235 | 309.8107 | 167.9489 | 176.734 | 191.5831 | 212.5192 | 209.4681 | 234.0128 | 193.1278 | 155.8107 | 134.4065 | 160.5575 | 156.1022 | 181.6318 | 305.3554 | 330.849 | 345.4553 | 497.4425 | 286.5575 | 319.1765 | 366.3299 | 489.734 | 144.2915 | 88.5447 | 88.2532 | 144.2915 | 135.1765 | 126.4553 | 159.1893 | 140.0256 | 120.5575 | 121.5319 | 164.9489 | 133.5447 | 102.0383 | 83.2915 | 100.6957 | 101.7724 | 117.4425 | 103.266 | 109.7596 | 90.0511 | 296.5319 | 311.3043 | 366.734 | 313.4297 | 212.2404 | 263.9744 | 291.7851 | 298.7979 | 338.885 | 264.8362 | 329.9233 | 346.2404
795 | 10.7394 | 197.2635 | 318.9272 | 345.7591 | 243 | 204.9636 | 205.527 | 268.059 | 217.5408 | 187.9046 | 161.0728 | 140.2999 | 171.5997 | 247.4316 | 222.0728 | 228.409 | 274.2045 | 198.9046 | 197.0138 | 202.3363 | 245.2409 | 211.9636 | 409.3224 | 457.1543 | 215.468 | 270.1455 | 278.591 | 318.8181 | 369.7591 | 154.8319 | 190.1594 | 175.9862 | 170.2271 | 198.2773 | 190.8319 | 209.2773 | 111.0728 | 295.9862 | 345.7089 | 403.2685 | 367.9272 | 171.409 | 169.8683 | 220.2409 | 173.7817 | 97.2045 | 89.7227 | 92.8181 | 150.1819 | 252.8181 | 314.8595 | 297.5546 | 200 | 180.2271 | 183.8181 | 240 | 167.0364 | 212.1681 | 241.6499 | 317.8908 | 214.8908 | 167.7227 | 186.7729 | 228.7365 | 180.5408 | 186.8545 | 281.9272 | 309.2045 | 280.5997 | 200.9134 | 220.7817 | 354.0728 | 266.7729 | 242.2773 | 249.6135 | 299.7001 | 366.7591 | 216.7955 | 220.9272 | 228.1092 | 301.2045 | 161.3726 | 158.2635 | 185 | 189.7453 | 200.6135 | 227.6863 | 223.2321 | 169.9774 | 149.3275 | 156.7591 | 151.0364 | 178.0954 | 300.7955 | 345.0138 | 380.123 | 510.3501 | 290.3501 | 291.5997 | 369.8683 | 508.4592 | 163.7729 | 89.5546 | 94.8319 | 163.7729 | 122.0226 | 131.5182 | 134.3952 | 130.409 | 121.3137 | 121.2773 | 142.4316 | 139.1092 | 96.7227 | 92.1455 | 97.409 | 115.941 | 128.0728 | 111.8683 | 111.2773 | 105.7365 | 287.1681 | 302.4454 | 379.3865 | 331.1455 | 208.4454 | 269.7955 | 250.2861 | 302.591 | 336.6135 | 271.6637 | 324.6135 | 362.941
796 | 10.7504 | 194.0856 | 312.3717 | 337.5147 | 226.8009 | 203.2286 | 219.9986 | 251.4292 | 207.3717 | 173.4867 | 156.8289 | 172.798 | 191.0841 | 228.2006 | 217.8289 | 205.9439 | 279.9425 | 213.7994 | 209.3997 | 219.5989 | 240.3717 | 204.8289 | 418.8275 | 457.2567 | 222.6564 | 263.972 | 289.4853 | 340.2272 | 353.8009 | 153.1431 | 188.6578 | 190.5989 | 182.2847 | 202.8569 | 186.8289 | 210.7714 | 120.7139 | 294.3997 | 362.6844 | 433.1991 | 379.885 | 173.6283 | 173.7714 | 220.7714 | 179.5708 | 85.972 | 88.2286 | 99.5708 | 163.4853 | 232.5722 | 289.3451 | 306.7139 | 201.5428 | 181.4853 | 193.6564 | 228.4292 | 162.6003 | 194.5722 | 253.6564 | 324.2567 | 221.2567 | 172.3997 | 167.0014 | 227.2861 | 165.7433 | 206.913 | 269.972 | 336.5414 | 283.1136 | 203.5708 | 220.3997 | 339.0295 | 290.1991 | 237.6003 | 256.028 | 307.3142 | 370.8569 | 189.4586 | 222.8569 | 257.9986 | 313.1136 | 162.6283 | 175.1416 | 193.4853 | 199.9705 | 207.7994 | 221.3717 | 229.0575 | 157.4867 | 144.115 | 160.8569 | 160.4853 | 184.028 | 291.2006 | 327.3156 | 355.2875 | 489.087 | 296.0856 | 306.4558 | 368.3717 | 515.5428 | 170.2286 | 91.7714 | 91.6003 | 170.2286 | 116 | 118.115 | 141.1991 | 154.9986 | 133.028 | 122.7714 | 141.7139 | 134.2861 | 99.8569 | 96 | 100.3997 | 120.2286 | 121.5147 | 112.6858 | 123.5708 | 96.5722 | 281.1431 | 274.0014 | 374.5147 | 294.8878 | 203.9144 | 262.5147 | 257.2847 | 298.0575 | 306.0014 | 269.8289 | 330.2567 | 391.913
797 | 10.7614 | 195.7837 | 297.6775 | 349.1062 | 262.7513 | 177.1386 | 239.7552 | 267.592 | 226.3757 | 183.6736 | 146.3795 | 166.7591 | 189.8122 | 257.9145 | 216 | 234.4819 | 267.8938 | 228.7552 | 196.8938 | 189.9508 | 242.9184 | 210.0531 | 414.8122 | 457.649 | 232.4858 | 277.4573 | 286.5142 | 339.3795 | 353.7021 | 150.4326 | 160.706 | 193.0816 | 184.0816 | 193.8122 | 181.0816 | 222.7552 | 104.1917 | 299.1347 | 362.9469 | 412.1386 | 360.4896 | 181.2694 | 165.3795 | 203.7591 | 186.9184 | 101.8083 | 92.7021 | 88.8938 | 148.7591 | 265.6166 | 279.351 | 316.0531 | 209.8368 | 191.8368 | 204.2694 | 215.5959 | 126.5181 | 215.8614 | 245.4611 | 314.4611 | 227.1347 | 182.6205 | 170.7552 | 214.0285 | 151.2448 | 197.5427 | 279.5389 | 323.057 | 288.5674 | 202.2979 | 224.351 | 338.7021 | 318.5104 | 252.4573 | 268.4041 | 317.4041 | 376.7021 | 199.0246 | 224.7837 | 246.408 | 296.6243 | 188.294 | 183.5674 | 196 | 198.7306 | 198.4611 | 209.8122 | 232.4858 | 163.7552 | 143.9184 | 149.4611 | 169.2694 | 190.9184 | 280.1632 | 329.4041 | 384.9676 | 485.9184 | 300.1347 | 291.2733 | 360.7306 | 481.5181 | 159.0285 | 87.2979 | 82.9469 | 159.0285 | 133.2409 | 160.2371 | 138.1632 | 155.1632 | 139.9184 | 119.8653 | 137.7306 | 136.7021 | 91.5959 | 81.1101 | 83.1917 | 100.408 | 111.1632 | 115.1347 | 104.2733 | 88.0816 | 283.1347 | 267.5674 | 380.6205 | 304.943 | 200.649 | 268.6205 | 266.9184 | 298.351 | 321.5104 | 276.6205 | 309.4896 | 384.1101
798 | 10.7724 | 176.1008 | 294.796 | 326.733 | 251.7128 | 199.4509 | 231.0605 | 233.2015 | 232.796 | 179.2443 | 166.2872 | 148.4685 | 181.4282 | 265.4081 | 200.0806 | 246.3879 | 252.0605 | 214.4887 | 197.7758 | 195.9194 | 271.063 | 207.2242 | 428.7154 | 470.5315 | 226.8363 | 252.5491 | 275.4484 | 340.9798 | 356.5919 | 150.796 | 146.8363 | 204.7355 | 177.4282 | 197.3678 | 189.5516 | 226 | 126.063 | 315.9194 | 337.1209 | 387.4887 | 385.0428 | 172.6524 | 157.4282 | 201.3879 | 176.8564 | 100.6322 | 90.8161 | 98.5315 | 161.5113 | 279.1839 | 276.0202 | 300.4887 | 208.0202 | 158.1814 | 184.5088 | 221.7557 | 138.4912 | 211.0605 | 219.7128 | 304.6322 | 210.4887 | 168.2846 | 177.9798 | 228.5113 | 154.3678 | 197.7758 | 277.4282 | 306.6524 | 309.6952 | 181.8967 | 208.2846 | 320.8967 | 256.5466 | 247.4484 | 245.529 | 312.0403 | 393.9194 | 176.937 | 210.6725 | 263.2872 | 301.3476 | 186.2443 | 190.3678 | 175.3048 | 218.4912 | 187.8363 | 210.9798 | 211.7128 | 181.3275 | 127.4887 | 146 | 147.1209 | 170.5088 | 291.5315 | 332 | 377.6927 | 473.4685 | 306.5718 | 305.6952 | 391.6348 | 461.6524 | 152.0202 | 100.3275 | 82.5919 | 152.0202 | 122.8766 | 153.1008 | 135.204 | 146.6322 | 129.8564 | 129.3476 | 137.5919 | 147.5516 | 75.4685 | 83.3678 | 96.3073 | 110.1234 | 124.1234 | 116.796 | 114.7154 | 88.5919 | 277.6322 | 292.6751 | 367.8766 | 316.5718 | 194.4282 | 263.8363 | 274.3678 | 279.8967 | 313.6725 | 237.6096 | 299.8161 | 380.796
799 | 10.7834 | 174.233 | 284.4926 | 315.1504 | 238.7257 | 194.8761 | 229.6165 | 280.3865 | 213.6017 | 216.6048 | 162.3009 | 145.8083 | 188.0826 | 241.5604 | 226.7139 | 243.767 | 256.2743 | 225.357 | 211.9322 | 206.4661 | 239.2035 | 218.1239 | 411.177 | 469.9587 | 228.233 | 278.947 | 266.5339 | 335.5339 | 363.4661 | 141.3009 | 142.5752 | 203.9587 | 187.3157 | 187.6843 | 175.8348 | 209.0265 | 125.7257 | 302.2183 | 332.6165 | 399.1652 | 358.2448 | 174.0413 | 168.9322 | 197.9587 | 202.2891 | 107.8909 | 103.7404 | 97.1504 | 182.9735 | 251.7109 | 262.8761 | 306.5074 | 217.5074 | 182.947 | 184.6578 | 218.3422 | 169.8644 | 225.7817 | 260.8791 | 254.5043 | 197.1091 | 164 | 196.7817 | 203.9026 | 170.5487 | 206.2743 | 290.5487 | 324.2065 | 287.5191 | 182.6578 | 212.0826 | 328.1239 | 261.6313 | 241.767 | 257.59 | 297.0678 | 365.6696 | 214.4544 | 207.8083 | 239.0943 | 301.5752 | 153.2861 | 175.0265 | 180.5074 | 228.0413 | 188.4248 | 228.1652 | 214.8909 | 175.3009 | 145.6313 | 160.5487 | 167.6726 | 182.7817 | 317.6313 | 310.177 | 361.6843 | 474.0413 | 312.8496 | 290.7935 | 345.0382 | 509.1122 | 142.9174 | 71.6696 | 83 | 142.9174 | 119.8083 | 139.9174 | 139.0413 | 147.4248 | 143.1652 | 123.1091 | 137.1917 | 151.6165 | 87.357 | 104.3983 | 103.4248 | 102.6843 | 110.2183 | 109.7257 | 110.9174 | 85.767 | 288.9322 | 283.643 | 326.82 | 339.823 | 197.0413 | 246.643 | 278.4248 | 290.357 | 347.18 | 265.7965 | 300.6165 | 358.3687
800 | 10.7944 | 170.0767 | 297.5905 | 311.5383 | 238.6411 | 188.7178 | 241.4877 | 265.2807 | 205.7178 | 179.2285 | 145.2301 | 150.1028 | 196.5644 | 239.2822 | 214.3067 | 247.1028 | 294.9248 | 202.7423 | 210.0767 | 228.5138 | 218.5123 | 202.1273 | 405.1794 | 445.204 | 231.4617 | 251.7163 | 272.385 | 307.7423 | 377.3083 | 137.3589 | 146.1028 | 210.385 | 199.0261 | 198.9494 | 162.9739 | 218.9494 | 124.8206 | 293.0767 | 346.9494 | 389.8712 | 382.0015 | 179.1028 | 184.3083 | 199.4617 | 167.9724 | 124.7699 | 92.2301 | 97.6411 | 164.0245 | 258.9494 | 278.0521 | 299.1534 | 208.5123 | 179.5123 | 175.3328 | 211.2561 | 155.4862 | 204.5629 | 250.6656 | 297.977 | 204.0261 | 185.3344 | 196.8972 | 197.8206 | 168.2561 | 194.8712 | 276.7684 | 329 | 290.0261 | 185.6411 | 201.6917 | 295.7163 | 286.6933 | 246.7439 | 251.3328 | 328.4632 | 344.0506 | 209.4095 | 221.9494 | 233.6411 | 307.5644 | 178.0015 | 176.7439 | 165.7684 | 195.3574 | 185.7178 | 213.9479 | 249.8221 | 149.204 | 136.2301 | 159.8972 | 165.7945 | 165.6656 | 321.3589 | 327.1549 | 378.6933 | 488.1288 | 285.2807 | 326.2071 | 361.5399 | 494.7423 | 141 | 80.411 | 75.615 | 141 | 122.4617 | 127.3328 | 121.1273 | 157.0261 | 110.0752 | 118.5383 | 127.1534 | 142.9739 | 96.7439 | 89.3067 | 94.1534 | 119.6933 | 119.9494 | 107.1794 | 101.615 | 92.385 | 285.4356 | 303.796 | 353.2837 | 310.5368 | 198.8206 | 243.8206 | 244.5368 | 275.1273 | 315.7929 | 281.5644 | 319.0521 | 342.3328
801 | 10.8054 | 198.1495 | 295.1701 | 316.8299 | 238.1672 | 178.8387 | 239.8358 | 249.8387 | 224.9882 | 178.1613 | 149.4956 | 142.6716 | 194.6686 | 254.1583 | 185.0147 | 245.5015 | 291.3402 | 194.5015 | 217.3284 | 233 | 261.8064 | 188.0059 | 408.3314 | 478.3108 | 229.5015 | 249.8299 | 299.8181 | 302 | 390.827 | 159.4868 | 141.1701 | 187.0147 | 179.3461 | 187.8417 | 172.6598 | 226.9971 | 134.1613 | 287.8358 | 339.173 | 412.8181 | 395.6627 | 190.827 | 170.3431 | 191.6716 | 174.824 | 106.3461 | 88.1672 | 90.5044 | 174.824 | 246.176 | 314.4809 | 314.4897 | 203.5015 | 177 | 207.9794 | 213.3314 | 151 | 188.173 | 233.5074 | 303.3373 | 201.8358 | 180.8387 | 176.0118 | 193.8358 | 172.8299 | 194.4985 | 252.1789 | 304.8475 | 287.8358 | 185.1672 | 219.8211 | 334.6392 | 310.1554 | 264.6569 | 275.651 | 296.8564 | 372.648 | 193.5074 | 182.525 | 253.1554 | 308.1672 | 182.5015 | 159.6775 | 205.3078 | 205.4897 | 214.9824 | 203.3373 | 219.5221 | 176.4809 | 136.3314 | 139.0118 | 157.3373 | 163.4985 | 294.349 | 347.824 | 363.0118 | 498.4956 | 264.8417 | 301.6863 | 347.1789 | 460.6833 | 170.1495 | 85.6657 | 74.8328 | 170.1495 | 132.9941 | 146.6539 | 139.4868 | 142.3431 | 120.3225 | 137.1554 | 134.9941 | 155.1583 | 104.6627 | 83.3343 | 82.0059 | 117.3373 | 121.3343 | 116.1613 | 124.1525 | 100.6627 | 284.8328 | 289.0118 | 341.0118 | 284.6775 | 192.3373 | 269.8181 | 287.8034 | 301.8152 | 318.6598 | 276.3373 | 288.8534 | 359.1554
802 | 10.8164 | 185.407 | 295.6903 | 331.5221 | 228.7035 | 183.7611 | 208.5752 | 257.2965 | 226.4646 | 192.677 | 147.6195 | 162.9735 | 194 | 228.2654 | 193.5221 | 228.9425 | 276.323 | 194 | 205.4779 | 198.3495 | 232.969 | 210.5089 | 412.3805 | 450.5044 | 218.8584 | 265.3673 | 284.7168 | 340.031 | 384.5487 | 144.5619 | 140.8451 | 171.8584 | 187.677 | 201.9027 | 189.3673 | 217.0133 | 128.3938 | 325.8762 | 371.6505 | 389.2654 | 371.646 | 171.8716 | 172.9159 | 200.9867 | 167.8584 | 102.8451 | 98.1416 | 94.0708 | 151.8008 | 276.8054 | 296.4911 | 285.8849 | 219.0575 | 171.9292 | 188.8008 | 232.593 | 157.7611 | 215.5797 | 241.1416 | 281.7168 | 191.7035 | 177.3097 | 185.5221 | 202.2965 | 164.7035 | 222.8894 | 237.0133 | 305.9159 | 287.8451 | 173.1681 | 232.4513 | 332.1681 | 290.3362 | 262.0841 | 239.5884 | 301.677 | 361.2522 | 187.6195 | 204.4248 | 268.8319 | 301.2389 | 182 | 180.5089 | 215.6903 | 215.7611 | 217.6195 | 207.0708 | 236.5089 | 177.9292 | 163.1992 | 133.3097 | 157.6903 | 165.6903 | 266.1814 | 328.1814 | 341.2522 | 496.6195 | 300.8762 | 293.3097 | 342.1549 | 445.7035 | 160.7876 | 86 | 80.0708 | 160.7876 | 135.8451 | 153.5354 | 116.9557 | 141.5354 | 157.8054 | 134.2389 | 145.4513 | 137.7168 | 94.1681 | 82.1549 | 86.7611 | 116.8451 | 119.3097 | 105.323 | 122.2389 | 113.8319 | 289.2257 | 297.677 | 341.2257 | 300.4381 | 202.8319 | 271.6195 | 252.3628 | 284.3362 | 304.9425 | 298.6638 | 324.2567 | 370.6062
803 | 10.8274 | 188.002 | 265.1327 | 306.5624 | 231.2871 | 193.5743 | 208.1446 | 226.4178 | 215.7109 | 157.2733 | 166.7208 | 170.4297 | 215.4356 | 223 | 187.4257 | 220.8554 | 282.5743 | 226.5822 | 217.5762 | 214.2931 | 245.7208 | 193.5644 | 405.2832 | 434.5683 | 210.1406 | 238.8475 | 254.4198 | 305.8436 | 350.4178 | 154.7188 | 131.5683 | 184.5762 | 169.4218 | 186.1366 | 182.5683 | 211.5703 | 124.4277 | 284.9842 | 356.5644 | 392.5743 | 336.1327 | 157.7109 | 176.5723 | 173.8475 | 184.0059 | 110.7168 | 99.1426 | 106.1465 | 166.7208 | 230.6971 | 306.5762 | 269.7109 | 208.2812 | 181.2891 | 199.4337 | 208.5624 | 169.2891 | 202.9941 | 215.5624 | 266.8535 | 191.7149 | 162.4238 | 189.7149 | 194.5683 | 159.5703 | 196.2753 | 258.1505 | 316.4317 | 285.4277 | 177.8594 | 199.703 | 319.7109 | 268.8515 | 241.2792 | 246.5762 | 278.2772 | 373.4337 | 197.2891 | 216.002 | 258.996 | 318.8634 | 178.5703 | 205.5782 | 194.5644 | 213.5703 | 201.5663 | 230.2931 | 216.1347 | 177.8574 | 154.2812 | 139.002 | 169.1465 | 208.0138 | 283.4356 | 315.4257 | 380.0138 | 470.2772 | 285.7069 | 296.4297 | 341.1426 | 468.8653 | 155.4277 | 74.8535 | 80.1426 | 155.4277 | 123.996 | 131.7069 | 122.2891 | 142.8574 | 157.1406 | 123.5683 | 137.5683 | 125.4257 | 92 | 89.7168 | 88.8574 | 98.9941 | 114.7129 | 105.5723 | 133.004 | 94.5644 | 302.8614 | 287.996 | 347.1446 | 323.7208 | 211.8594 | 240.9901 | 270.5802 | 311.7247 | 330.295 | 271.2753 | 315.7089 | 341.1327
804 | 10.8384 | 195.088 | 292.1795 | 306.3486 | 229.3909 | 191.5211 | 220.3063 | 238.3943 | 215.7394 | 160.5669 | 140.4296 | 150.1268 | 223.3486 | 240.3943 | 198.176 | 221.7394 | 290.9577 | 213.736 | 188.6902 | 199.736 | 234.2148 | 205.6549 | 375.2994 | 419.9543 | 249.8767 | 239.2183 | 288.2675 | 281.6057 | 345.8697 | 147.4331 | 164.7887 | 173.9543 | 172.088 | 202.1338 | 186.2183 | 197.9543 | 117.912 | 293.5246 | 335.6057 | 343.5565 | 323.1726 | 172.5246 | 157.8662 | 197.7006 | 176.5634 | 100.6937 | 78.9965 | 110.6091 | 150.8662 | 231.5669 | 301.1726 | 290.6126 | 232.0915 | 176.0423 | 188.0845 | 203.1303 | 140.5599 | 200 | 239.7006 | 274.5669 | 195.4789 | 166.9577 | 195.2183 | 200.8274 | 177.264 | 210.1338 | 270.6972 | 310.1726 | 284.1303 | 180.7394 | 190.5211 | 344.0915 | 271.2183 | 256.264 | 269.0035 | 278.3486 | 369.912 | 219.8732 | 203.9543 | 231.7782 | 317.6514 | 164.9543 | 167.2536 | 161.4296 | 213 | 226.8309 | 221.824 | 218.088 | 168.4331 | 145.0423 | 155.6549 | 151.8662 | 197.6057 | 287 | 296.6057 | 353.9508 | 481.6549 | 302.8732 | 290.0423 | 320.1268 | 449.5177 | 165.4366 | 83.4366 | 81.7394 | 165.4366 | 129.8274 | 130.6091 | 144.8732 | 153.4366 | 122.9508 | 141.1338 | 138.6091 | 147.4823 | 93.7394 | 91.8697 | 87.2606 | 116.8732 | 100.9543 | 111.2183 | 128.912 | 113.6126 | 283.2571 | 280.7817 | 334.9543 | 300.9085 | 204.3028 | 229.912 | 241.0811 | 287.4296 | 308.0388 | 266.8697 | 283.4296 | 362.9612
805 | 10.8494 | 181.0059 | 319.0502 | 338.7523 | 219.2979 | 183.0619 | 232.5841 | 232.1799 | 195.7138 | 196.3983 | 171.2803 | 157.5841 | 217.8259 | 263.2862 | 204.41 | 238.7581 | 272.5958 | 192.4778 | 206.0502 | 174.9498 | 224.0619 | 200.9439 | 407.1623 | 431.2301 | 235.7138 | 267.3422 | 267.5398 | 339.8585 | 354.8201 | 143.354 | 156.7699 | 192.2862 | 177.41 | 182.0678 | 178.1799 | 204.8201 | 116.118 | 296 | 355.9322 | 369.5163 | 351.9883 | 168.8259 | 152.354 | 186.1239 | 185.5841 | 99.882 | 102.4602 | 90.7138 | 148 | 262.9883 | 275.3038 | 271.9498 | 196.3097 | 165.2979 | 200.1121 | 180.0678 | 160.6962 | 214.9941 | 234.2979 | 273.354 | 202.1741 | 175.9381 | 181.0059 | 222.2862 | 165.0059 | 207.708 | 287.8761 | 288.7138 | 276.9439 | 174.8259 | 199.7021 | 321.5398 | 242.0117 | 232.5398 | 292.2862 | 274.59 | 356.6519 | 219.472 | 209.9381 | 216.5339 | 294.9498 | 146.2419 | 207.7464 | 210.8024 | 195.3599 | 186.8996 | 220.882 | 214.59 | 159.9439 | 141.354 | 145.6519 | 197.5104 | 187.9439 | 323.1623 | 315.1682 | 365.7581 | 467.2419 | 328.0502 | 296.0561 | 365.5104 | 405.4277 | 146.7138 | 84.118 | 85.528 | 146.7138 | 112.4778 | 126.59 | 148 | 141.7699 | 128.5841 | 129.8879 | 150.4661 | 128.0678 | 99.292 | 93.764 | 94.0561 | 127.0561 | 130.7523 | 110.236 | 106.8318 | 99.3599 | 261.4778 | 265.0059 | 336.528 | 306.7021 | 205.646 | 228.118 | 240.41 | 268.0059 | 316.3481 | 245.8318 | 281.646 | 374.0561
806 | 10.8604 | 179 | 306.7969 | 308.1224 | 234.0974 | 196.3088 | 243.8373 | 261.4061 | 196.5772 | 175.9596 | 158.114 | 154.5285 | 199.114 | 263.3171 | 221.9917 | 200.7566 | 264.6342 | 237.3978 | 176.8053 | 193.4632 | 223 | 210.7316 | 374.4395 | 450.886 | 224.9513 | 305.8776 | 271.1544 | 292.5535 | 331.8539 | 167.1461 | 134.4311 | 182.4798 | 172.6342 | 209.4061 | 172.5285 | 192.5855 | 141.0404 | 319.2518 | 366.1544 | 399.0404 | 341.6912 | 192.1461 | 168.0974 | 208.1461 | 173.5855 | 117.886 | 89.0083 | 94.2601 | 144.4228 | 246.4311 | 282.7316 | 319.975 | 194.5772 | 168.4715 | 186.7969 | 184.1544 | 145.2197 | 205.3741 | 222.2684 | 275.6829 | 226.2518 | 166.2684 | 180.7886 | 204.4311 | 154.057 | 198.057 | 295.3658 | 281.5285 | 290.3088 | 184.7316 | 208.1544 | 321.5772 | 288.975 | 231.6829 | 262.8053 | 295.4632 | 369.3088 | 220.7886 | 202.9513 | 214.1057 | 289.3171 | 167.2518 | 191.6425 | 171.4965 | 207.3088 | 192.6259 | 229.0487 | 207.7399 | 165.2601 | 151.7316 | 126.114 | 186.114 | 168.2197 | 357.5118 | 326.943 | 362.6342 | 489.1461 | 319.3741 | 321.1461 | 375.5772 | 432.1947 | 160.9917 | 93.8373 | 100.3088 | 160.9917 | 114.4715 | 141.2031 | 134.5855 | 156.0974 | 127.3171 | 132.4715 | 128.7482 | 148.2518 | 103.5772 | 116.3575 | 94.1057 | 124.4228 | 109.0653 | 118.0487 | 110.2601 | 93.4228 | 265.2601 | 260.3171 | 322.6912 | 311.5772 | 223.886 | 232.4715 | 268.7233 | 286.5689 | 291.171 | 256.4145 | 268.5855 | 340.1224
807 | 10.8714 | 197.1318 | 295.0275 | 303.0934 | 236 | 189.8407 | 217.8023 | 241.4287 | 219.6647 | 163.0275 | 152.3736 | 139.4946 | 197 | 269.3461 | 200.4287 | 210.5054 | 272.1593 | 222.1484 | 183.8791 | 208.6923 | 235.6923 | 232.8516 | 425.302 | 458.4395 | 209.4946 | 275.5496 | 272 | 323.1702 | 363.4504 | 139.176 | 140.1593 | 182.8132 | 165.6539 | 191.2419 | 171.0934 | 214.5713 | 113.176 | 312.0275 | 366.0934 | 386.588 | 338.1868 | 191.3736 | 150.9616 | 191.055 | 153.8682 | 100.9616 | 87.9066 | 91.3736 | 164.8516 | 264.8516 | 264.055 | 325.0934 | 204.9725 | 196.1977 | 172.3077 | 223.0768 | 177.4504 | 229.3845 | 220.0934 | 283.2527 | 194.5496 | 161.3736 | 187.3461 | 194.7473 | 147.5605 | 203.3461 | 269.7089 | 292.7857 | 296.5329 | 195.0659 | 217.1593 | 334.6923 | 301.3461 | 214.7748 | 243.588 | 288.9341 | 351.9616 | 209.2143 | 201.0934 | 232.1318 | 334.3295 | 166.3736 | 180.8407 | 188.6647 | 187.2419 | 208.5054 | 221.8407 | 219.6923 | 182.3186 | 133.9616 | 144.8516 | 149.5496 | 185.0384 | 280.3135 | 347.0384 | 348.4012 | 447.5771 | 285.3628 | 316.7473 | 332.4837 | 437.8132 | 134.8957 | 76.8682 | 75.7089 | 134.8957 | 115 | 132.1209 | 123.0275 | 129.8957 | 135.1593 | 115.7748 | 141.412 | 123.8023 | 92.2143 | 99.055 | 88.5605 | 105.8682 | 110.5329 | 116.2802 | 104.6539 | 107.5054 | 277.7857 | 261.8132 | 350.9175 | 275.7364 | 187.0167 | 257.4779 | 281.0659 | 288.0934 | 314.2911 | 256.1868 | 284.2252 | 358.6647
808 | 10.8824 | 173.2713 | 253.5693 | 329.6475 | 238.7566 | 166.028 | 227.8643 | 201.3259 | 185.2448 | 173.9454 | 183.2419 | 139.8378 | 210.7832 | 220.3805 | 195.2434 | 237.7287 | 283.1077 | 207.1357 | 186.8378 | 186.1091 | 216.7846 | 196.407 | 381.3805 | 393.7595 | 214.4322 | 267.4056 | 265.5678 | 319.649 | 353.2168 | 158.972 | 156.621 | 169.2168 | 158.5678 | 181.649 | 182.0266 | 219.7566 | 121.9454 | 295.379 | 363.2434 | 367.5413 | 339.8378 | 179.9734 | 182.0797 | 171.5413 | 159.351 | 96.2434 | 98.1077 | 88.2434 | 155.0546 | 246.7846 | 292.3231 | 290.0825 | 199.5678 | 185.2168 | 189.3776 | 205.8657 | 146.0825 | 198.9203 | 236.5398 | 272.0546 | 197.4322 | 174.7832 | 194.4322 | 186.649 | 163.5398 | 195.7301 | 230.2448 | 284.8112 | 281.379 | 210.7021 | 187.6769 | 328.649 | 288.2168 | 240.5664 | 259.4587 | 277.8923 | 343.5678 | 203.4056 | 203.7566 | 238.5944 | 290.2993 | 166 | 180.9189 | 180.8923 | 184.0811 | 194.379 | 208.1357 | 230.1888 | 147.2448 | 123.7301 | 142.4056 | 163.4587 | 181.4867 | 314.2685 | 325.1091 | 313.9203 | 412.6769 | 308.6475 | 285.6769 | 343.621 | 466.4853 | 150.3776 | 97.972 | 91.3776 | 150.3776 | 120.5133 | 137.4322 | 112.8112 | 133.4322 | 111.1902 | 130.5398 | 131.9734 | 143.972 | 87.3245 | 83.2168 | 104.5398 | 98.4867 | 124.7832 | 95.7846 | 115.0266 | 82.3525 | 296.4587 | 258.3245 | 308.9749 | 297.7287 | 211.4853 | 240.7035 | 283.8378 | 269.6224 | 302.2979 | 263.351 | 247.407 | 372.0266
809 | 10.8934 | 195.2105 | 265.8299 | 313.3766 | 228.7571 | 171.4494 | 230.8623 | 254.8015 | 216.4533 | 189.8987 | 146.8908 | 151.1741 | 196.1701 | 261.6274 | 204.3117 | 232.5506 | 267.2389 | 235.7975 | 162.7895 | 180.2753 | 257.8339 | 191.1377 | 382.587 | 423.3845 | 227.1052 | 276.3117 | 267.7935 | 302.2389 | 339.8948 | 145.1701 | 150.5506 | 201.5222 | 179.4169 | 181 | 179.2753 | 200.4454 | 101.5831 | 289.3441 | 314.5791 | 366 | 364.2105 | 171.5506 | 155.2025 | 204.4533 | 160 | 90.413 | 83.1701 | 105.6923 | 175.4169 | 260.8299 | 267.0649 | 270.2389 | 192.4818 | 178.413 | 189.1377 | 205.8623 | 167.2105 | 211.8299 | 234.2753 | 292.4169 | 205.4494 | 150.8584 | 170.7895 | 198.1052 | 189.2105 | 208.0364 | 262.3845 | 289.587 | 279.0688 | 202.6883 | 188.7247 | 308.4454 | 253.4778 | 258.8299 | 247.0324 | 289.1052 | 350.4494 | 201.1377 | 188.1701 | 214.7895 | 309.2793 | 167.8623 | 188.4494 | 181.8623 | 184 | 180.8948 | 224.6923 | 201.2025 | 151.4494 | 125.7935 | 150.3805 | 145.4454 | 178.2065 | 323.587 | 318.3441 | 336.1416 | 475.182 | 280.2713 | 306.2793 | 355.2429 | 418.7167 | 171.5546 | 96.2753 | 89.2753 | 171.5546 | 121 | 160.3481 | 164.1456 | 156.3481 | 127.6234 | 117.1013 | 145.8987 | 142.2753 | 75.8259 | 83.8623 | 84.5831 | 98.9312 | 123.2065 | 111.6923 | 108.5506 | 102.3481 | 284.9636 | 268.2429 | 353.4209 | 305.587 | 200.0324 | 216.6519 | 245.822 | 284.7611 | 292.6195 | 275.1741 | 311.0444 | 376.7247
810 | 10.9044 | 178.1307 | 291.53 | 307.2827 | 245.9258 | 175.7739 | 238.5477 | 225.9788 | 191.6396 | 188.1696 | 134.5654 | 146.3392 | 168.583 | 236.6961 | 199.3392 | 196.1484 | 304.682 | 207.8092 | 190.2473 | 178.1131 | 243.0742 | 207.9823 | 427.3427 | 402.4134 | 228.9435 | 260.9611 | 284.9823 | 298.1696 | 309.7527 | 162.8693 | 148.1131 | 173.8092 | 171.5654 | 181.9435 | 153.5265 | 204.6608 | 125.4735 | 280.5088 | 335.53 | 380.1519 | 348.0742 | 189.8693 | 151.1131 | 177.7527 | 154.3392 | 90.9435 | 98.0389 | 98.5088 | 154.3569 | 278.9823 | 279.1519 | 303.9081 | 199.5477 | 200.6431 | 177.6784 | 198.4523 | 154.8481 | 212.0565 | 232.1131 | 251.5442 | 189.9611 | 166.9258 | 183.1519 | 222.5866 | 179.6784 | 202.3958 | 237.6396 | 261.6961 | 290.3216 | 194.4523 | 188.0565 | 320.2085 | 259.4912 | 246.7915 | 249.7739 | 262.6396 | 331.1873 | 183.0742 | 203.9823 | 229.9823 | 302.5088 | 159.5088 | 187.1131 | 161.2438 | 190.6042 | 202.6431 | 217.5088 | 187.6784 | 183.1343 | 157.1343 | 145.3392 | 150.6042 | 194.9823 | 301.3569 | 307.6219 | 366.3039 | 430.9399 | 271.3958 | 309.8869 | 335.2438 | 440.4735 | 182.4346 | 70.5265 | 93.7173 | 182.4346 | 119.1131 | 141.2438 | 158.5654 | 148.5654 | 125.2261 | 121.6608 | 153.6042 | 124.0742 | 77.8304 | 76.4523 | 90.5477 | 110.3216 | 126.7739 | 106.3958 | 128.7562 | 98.3392 | 262.3004 | 308.6254 | 323.9788 | 310.7173 | 205.6042 | 242.3604 | 267.53 | 299.2085 | 313.6996 | 259.9611 | 297.1307 | 336.4311
811 | 10.9154 | 200.8938 | 289.177 | 288.8407 | 244.1327 | 181.7345 | 213.1947 | 235.469 | 197.646 | 164.1062 | 157.8938 | 147.9115 | 187.0708 | 267.4956 | 235.3186 | 186.354 | 240.0972 | 225.1151 | 192 | 179.9115 | 227.6637 | 194.6637 | 379.3451 | 456.4337 | 218.4867 | 258.0885 | 261.1504 | 296.0885 | 340.4956 | 129.5929 | 144.177 | 148.1062 | 165.2655 | 203.9823 | 186.4071 | 192.5752 | 121.2655 | 287.646 | 328.3982 | 359.0177 | 343.177 | 168.0619 | 146.2212 | 192.2478 | 168.3363 | 113.9381 | 111.4248 | 94.177 | 144.3982 | 262.7964 | 276.177 | 309.823 | 194.2655 | 178.1062 | 164.5752 | 212.3363 | 175.9823 | 219.646 | 216.7079 | 276.7169 | 189 | 175.646 | 161.0619 | 210.6194 | 160.8407 | 208.6903 | 261.8053 | 319.2567 | 312.0266 | 187.3097 | 186.0885 | 285.6371 | 263.823 | 219.2389 | 262.4248 | 270.5575 | 377.7877 | 215.4514 | 188.7522 | 223.354 | 302 | 182.8938 | 181.2655 | 180.0708 | 216.8053 | 193.4867 | 218.9115 | 232.8762 | 168.7522 | 141.7964 | 130.6637 | 162.469 | 164.4601 | 300 | 317.5133 | 344.1062 | 412.7079 | 292.9823 | 294.7079 | 355.9823 | 421.9292 | 165.7964 | 90.9823 | 85.3982 | 165.7964 | 126.646 | 135.2212 | 139.8407 | 149.9115 | 135.5133 | 133.469 | 133.9292 | 138.2921 | 83.7345 | 95.1151 | 76.6637 | 95.7079 | 99.2831 | 100.2655 | 103.2389 | 87.4867 | 288.7169 | 283.2831 | 302.8849 | 312.9115 | 194.531 | 260.2478 | 228.8583 | 286.6194 | 307.354 | 239.8849 | 286.4425 | 330.177
812 | 10.9264 | 179.735 | 289 | 303.4887 | 228.5113 | 197.4887 | 199.4154 | 198.2463 | 215.4248 | 175.5846 | 153.1917 | 179.9455 | 161.8628 | 239.9587 | 188.5978 | 196.6485 | 267.0094 | 182.438 | 181.3515 | 189.6804 | 214.4154 | 207.5526 | 354.735 | 408.6617 | 231.5526 | 246.3835 | 264.8402 | 291.1598 | 341.032 | 148.3289 | 147.8722 | 156.6804 | 156.2876 | 193.3835 | 175.4154 | 175.5432 | 118.0959 | 262.8308 | 311.5432 | 346.3835 | 328.4793 | 165.0639 | 174.0733 | 190.0959 | 160.2876 | 107.2556 | 84.8947 | 93.032 | 141.0959 | 245.5432 | 291.4887 | 305.1598 | 183.3515 | 171.1917 | 194.0094 | 191.703 | 163.4474 | 191.9267 | 196.6391 | 260.5752 | 194.8083 | 180.8402 | 176.4568 | 194.5113 | 159.032 | 235.1372 | 217.5019 | 280.3741 | 240.3966 | 190.8722 | 193.7444 | 320.7857 | 264.968 | 235.4248 | 229.1184 | 290.3609 | 381.9361 | 217 | 213.1692 | 222.032 | 281.6711 | 186.9041 | 179.0639 | 163.5752 | 207.3515 | 207.5207 | 215.1278 | 225.3196 | 158.3196 | 130.3515 | 139.6804 | 161.0639 | 167.8402 | 269.9906 | 306.3835 | 352.6804 | 470.0827 | 298.8402 | 294.968 | 348.2876 | 440.3609 | 172.7444 | 98.7763 | 91.7763 | 172.7444 | 124.0959 | 121.4474 | 125.4474 | 128.703 | 110.8308 | 123.3515 | 129.1278 | 141.9041 | 94.6485 | 83.4154 | 91.4887 | 123.0733 | 125.1053 | 109.6804 | 99.0959 | 90.8722 | 285.1598 | 274.2556 | 291.3515 | 314.9361 | 208.5207 | 259.0639 | 265.7218 | 284.0639 | 288.6072 | 288.3703 | 309.2331 | 351.297
813 | 10.9374 | 179 | 271.354 | 285.3736 | 247.6067 | 196.0393 | 211.7444 | 222.4887 | 244.4297 | 200.5084 | 156.9213 | 151.5899 | 183.5477 | 233.118 | 183.0787 | 215.6264 | 270.941 | 244.7218 | 203.5477 | 156.6686 | 215.9607 | 176.6293 | 386.3511 | 400.1376 | 226.118 | 261.6854 | 270.882 | 264.5309 | 319.4326 | 158.8034 | 146.0393 | 163.8624 | 178.5477 | 178.295 | 178.9213 | 197.5477 | 105.2556 | 306.1151 | 321.7837 | 339.1376 | 320.1573 | 170.882 | 177.941 | 171.3736 | 155.0983 | 106.0197 | 88.9017 | 85.1573 | 189.0365 | 233.236 | 281.2163 | 282.4523 | 208.4887 | 168.059 | 179.3146 | 206.6854 | 139.4719 | 226.2921 | 211.6854 | 272.7444 | 205.7837 | 168.2556 | 162.295 | 204.7837 | 152.1376 | 210.5113 | 231.6854 | 288.8034 | 249.764 | 193.941 | 177.3343 | 275.9242 | 267.941 | 270.3118 | 263.2921 | 293.941 | 353.5703 | 211.118 | 209.0983 | 259.2528 | 263.354 | 162.4916 | 192.7247 | 196.3314 | 207.9803 | 182.5113 | 220.882 | 209.3146 | 149.177 | 133.9213 | 146.8624 | 144.3343 | 182.705 | 280.764 | 316.7837 | 315.7472 | 463.177 | 291.1573 | 298.9213 | 324.4719 | 452.764 | 130.8456 | 99 | 105.7247 | 130.8456 | 138.705 | 126.882 | 145.5871 | 149.5674 | 112.941 | 113.1966 | 125.0787 | 124.354 | 98.9213 | 88.882 | 89.059 | 104.3933 | 96.5899 | 106.0787 | 116.646 | 94.9213 | 306.5674 | 274.9803 | 317.4691 | 297.354 | 222.7247 | 264.882 | 262.0983 | 279.0983 | 307.6067 | 289.0197 | 295.295 | 345.1376
814 | 10.9484 | 184.9558 | 247.177 | 299.8894 | 219.2138 | 201.9558 | 205.0516 | 208.1106 | 195.3687 | 192.0664 | 146.0811 | 145.0442 | 177.0516 | 248.882 | 183 | 220.9631 | 246.1843 | 219.1991 | 190.1032 | 179.823 | 215.0074 | 209.7493 | 384.0221 | 406.9484 | 230.9631 | 257.0369 | 259.0885 | 327.5281 | 328.9263 | 138.1548 | 151.9558 | 137.1991 | 170.0664 | 183.9558 | 171.059 | 206.9336 | 145.6977 | 307.9926 | 300.1622 | 351.9041 | 335.882 | 186.882 | 151.1991 | 170.0074 | 182.7936 | 106 | 88.0074 | 103.8599 | 177.0959 | 257.8157 | 293.9041 | 299.8673 | 185.177 | 157.0811 | 185.9484 | 193.1032 | 159.8452 | 232.9558 | 219.941 | 281.9336 | 206 | 150.1327 | 185.823 | 217.9041 | 183.7641 | 212.9779 | 256.8157 | 273.118 | 249.0074 | 207.8968 | 209.7567 | 303.7862 | 261.0516 | 219.3834 | 257.0516 | 290.0295 | 358.9558 | 216.9558 | 197.0885 | 213.3465 | 277.8894 | 180.8599 | 173.1475 | 171.1917 | 200.059 | 219.7198 | 215.0442 | 203.0442 | 160.9115 | 144.9189 | 154.941 | 167.823 | 185.9779 | 289.9336 | 306.0811 | 322.941 | 441.1622 | 300.9263 | 279.1475 | 355.7641 | 462.9263 | 156.8009 | 85.1032 | 94.0885 | 156.8009 | 129.0737 | 131.9631 | 131.1106 | 146.0295 | 121.9336 | 163.624 | 130.9558 | 149.8083 | 100.9853 | 73.118 | 72.1253 | 122.8599 | 92.0295 | 108.9779 | 115.0147 | 103.9336 | 262.3318 | 250.1843 | 312.0442 | 313.8747 | 189.2507 | 270.9558 | 262 | 352.4544 | 303.0369 | 270.1401 | 282.0959 | 380.7346
815 | 10.9594 | 186.0344 | 279.9361 | 245.1721 | 224.0295 | 181.0836 | 215.9508 | 185.1869 | 203.0148 | 149.2066 | 142.0393 | 131.0246 | 202.9705 | 239.1131 | 202.9115 | 221.9016 | 233.2213 | 230.7934 | 160.0492 | 199.8623 | 223.0738 | 189 | 379.1574 | 397.9361 | 217.118 | 276.7885 | 277 | 306.8328 | 317.9311 | 151.9213 | 160.9852 | 191.8672 | 156.118 | 176.059 | 170.9607 | 201.0344 | 115.0049 | 277.0295 | 305.9852 | 345.9311 | 365.0148 | 186.8771 | 143.0885 | 173.0344 | 198.7984 | 115.8475 | 103.0197 | 96.9852 | 133.0738 | 266.7639 | 230.1623 | 298.0197 | 190.0787 | 181.9311 | 171.0738 | 182.0885 | 151.9164 | 207.0787 | 240.8328 | 288.7885 | 190.9754 | 167.0344 | 186.0393 | 196.0836 | 153.0344 | 191.0836 | 264.9607 | 286.0098 | 311.8426 | 174.0639 | 189.1475 | 301.8918 | 263.1377 | 245.0934 | 239.0492 | 280.9262 | 318.1525 | 197.9754 | 203.941 | 200.0885 | 281.0541 | 165.1475 | 200.7344 | 167.0934 | 177.0787 | 188.1672 | 181.1475 | 222.9803 | 127.1721 | 146.9016 | 150.9311 | 143.0934 | 167.9361 | 297.9066 | 339.9705 | 332.9016 | 422.1574 | 300.8574 | 289.9361 | 339.0246 | 425.0344 | 156.9066 | 88.0984 | 78.0689 | 156.9066 | 123.9951 | 131 | 130.9115 | 146.8771 | 111.0836 | 128.9361 | 150.8918 | 115.0295 | 79.0443 | 91.0098 | 117.818 | 125.9951 | 113.9607 | 105.0934 | 123.8574 | 82.0885 | 268.0344 | 308.9066 | 256.3885 | 286.118 | 222.0098 | 271.941 | 285.8771 | 313.0246 | 276.1623 | 229.2115 | 264.2066 | 344.9754
816 | 10.9704 | 193.2065 | 266.9828 | 279.914 | 229.8451 | 197.8451 | 205.6214 | 223.0172 | 205.5526 | 190.9484 | 149.8451 | 136.2065 | 196.9828 | 261.6386 | 185.0516 | 201.8107 | 277.7935 | 189.2237 | 170.2237 | 172.1205 | 237.6214 | 188.8967 | 410.1568 | 384.4321 | 240.3977 | 234.6023 | 276.5009 | 273.5851 | 304.2409 | 136.3442 | 157.9312 | 165.3098 | 179.8107 | 188.3958 | 162.9312 | 207.8623 | 116.2926 | 283.6884 | 303.7916 | 332.2926 | 367.0019 | 162.4474 | 161.327 | 180.1205 | 157.8623 | 85.3786 | 106.673 | 93.6558 | 147.8107 | 220.0153 | 262.673 | 301.8967 | 206.0344 | 167.6902 | 185.4149 | 200.0688 | 135.5507 | 222.6214 | 207.5335 | 246.4646 | 186.2753 | 174.1205 | 193.6214 | 212.2428 | 159.8279 | 207.9656 | 256.6042 | 288.086 | 279.8279 | 186.8967 | 218.7935 | 280.1721 | 290.3977 | 263.0019 | 249.0172 | 266.9465 | 349.2753 | 192.9312 | 192.1033 | 217.4837 | 292.2581 | 194.587 | 147.4819 | 185.4321 | 193.4302 | 221.6214 | 211.2926 | 219.3614 | 161.5181 | 127.6884 | 136.8279 | 161.8795 | 155.0516 | 279.4819 | 333.4665 | 312.9828 | 454.2065 | 272.4474 | 277.6884 | 343.4837 | 431.4665 | 138.0516 | 107.4149 | 91.7591 | 138.0516 | 122.9312 | 130.673 | 113.4819 | 122.2926 | 128.0688 | 116.2065 | 128.8107 | 121.5163 | 88.0688 | 92.8279 | 81.327 | 124.7247 | 105.8279 | 123.7935 | 95.327 | 99.9484 | 275.2237 | 289.8279 | 334.4149 | 309.914 | 223.6558 | 259.673 | 261.3098 | 317.5354 | 309.6023 | 271.9484 | 305.3633 | 340.0344
817 | 10.9814 | 203.938 | 265.764 | 274.764 | 221.708 | 188.7345 | 184.413 | 224.3245 | 180.0885 | 187.9115 | 141.295 | 147.2625 | 195.1445 | 240.941 | 187.9705 | 191.354 | 265.823 | 202.8555 | 183.1475 | 179.059 | 216.531 | 183.8555 | 362.118 | 351.6755 | 206.7375 | 268.941 | 248.7965 | 306.41 | 319.298 | 155.4985 | 153.174 | 182.0855 | 168.823 | 210.8525 | 159.59 | 200.0295 | 132.587 | 322.469 | 347.466 | 349.5015 | 310.708 | 187.764 | 180.236 | 186.764 | 150.649 | 106.882 | 88.295 | 74.708 | 137.413 | 277.056 | 245.3275 | 296.0295 | 208.059 | 150.531 | 153.416 | 203.9705 | 166.528 | 200.9705 | 236.879 | 271.8495 | 201.2625 | 180.6755 | 172.177 | 169.6785 | 149.5575 | 205.9705 | 234.826 | 291.2005 | 270.7375 | 180.4985 | 206.646 | 289.6165 | 256.0885 | 207.5635 | 250.531 | 320.115 | 363.8495 | 189.5015 | 198.3835 | 188.767 | 306.587 | 170.528 | 174.3215 | 153.531 | 217.41 | 200.2065 | 228.295 | 238.8495 | 135.2095 | 166.6755 | 127.295 | 154.8525 | 158 | 305.3185 | 302.7935 | 312.0295 | 464.879 | 297.587 | 316.0265 | 314.649 | 402.0915 | 141.295 | 74.708 | 77.705 | 141.295 | 118.6755 | 112.59 | 140.233 | 138.8525 | 132.767 | 127.823 | 118.7375 | 150.056 | 92.0885 | 83.118 | 99.7345 | 108.5575 | 97.121 | 111.9705 | 113.6165 | 97.6785 | 289.0915 | 279.056 | 301.6195 | 305.413 | 203.469 | 241.177 | 278.882 | 290.4395 | 342.9675 | 268.2035 | 268.6755 | 341.8525
818 | 10.9924 | 168.9582 | 258.0836 | 269.3403 | 243.4537 | 179.3313 | 197.4567 | 234.0388 | 184.0866 | 184.3731 | 150.7075 | 123.6687 | 168.7552 | 238.7075 | 187.209 | 202.7493 | 260.6269 | 230.206 | 187.4985 | 180.5821 | 233.7493 | 212.1672 | 365.7075 | 343.8 | 229.4537 | 266.206 | 275.5433 | 287.4179 | 362 | 139.8776 | 127.3791 | 152.9194 | 162.9164 | 204.203 | 179.1254 | 200.5403 | 118.4985 | 303.6627 | 297.1672 | 365.4985 | 333.3731 | 179.9164 | 187.3313 | 178.6657 | 171.2478 | 103.2507 | 97.6657 | 97.3313 | 151.0418 | 247.8806 | 289.2507 | 296.4985 | 208.6209 | 168.8776 | 198.6627 | 203.2507 | 150.5403 | 200.0836 | 200.6687 | 234.4597 | 178.4627 | 170.209 | 177.3313 | 191.5821 | 136.0448 | 204.4985 | 261.0806 | 233.5463 | 294.4985 | 165.5463 | 194.5821 | 275.5791 | 259.209 | 257.4119 | 266.203 | 291.794 | 327.4209 | 205.8746 | 211.0836 | 213.2896 | 293.5851 | 155.3343 | 152.4597 | 171.2507 | 197.1642 | 206.2478 | 236.9552 | 201.5851 | 173.7463 | 155.3731 | 137.1672 | 150.3343 | 158.794 | 251.4209 | 296.209 | 313.2507 | 427.206 | 283.8328 | 285.8806 | 335.3313 | 437.2478 | 150.5403 | 97.2478 | 68.6269 | 150.5403 | 108.6269 | 130.8299 | 116.1701 | 134.6687 | 156.4119 | 123.1701 | 142.9164 | 119.9612 | 95.2507 | 88.2119 | 90.9582 | 94.8776 | 133.6657 | 110.4149 | 101.4179 | 118.6627 | 324.5403 | 248 | 320.5373 | 320.6299 | 187.5045 | 246.9164 | 273.4955 | 272.4179 | 309.5851 | 243.0866 | 259.7134 | 336.8746
819 | 11.0034 | 169.3521 | 261.0276 | 319.8092 | 207.7837 | 164.2704 | 185.9735 | 211.2969 | 208.5673 | 170.4867 | 144.3786 | 139.649 | 208.2429 | 233.8929 | 192 | 197.5949 | 272.5122 | 212.1082 | 175.8378 | 171.1622 | 228.5408 | 215.7296 | 358.5673 | 407.1347 | 194.1082 | 250.1092 | 286.8102 | 295.9184 | 361.1888 | 160.2704 | 158.4051 | 174.2704 | 161 | 164.1357 | 181.5673 | 189.9459 | 107.5949 | 274.0276 | 301.3245 | 351.2959 | 321.001 | 177.4592 | 171.5673 | 171 | 154.2163 | 108.7296 | 89.0265 | 82.4327 | 152.4867 | 289.8102 | 291.9173 | 286.2439 | 178.4602 | 188.5673 | 168.8653 | 207.8643 | 141.1898 | 202.5949 | 215.5673 | 245.5949 | 210.9184 | 173.9184 | 161.6214 | 184.2714 | 160.4327 | 193.4867 | 239.0806 | 270.4061 | 281.702 | 200.2429 | 184.7296 | 245.2714 | 262.4857 | 220.8918 | 226.947 | 309.5133 | 360.0541 | 202.5673 | 212.0806 | 197.5408 | 306.9459 | 164.1357 | 163.1082 | 175.8643 | 177.7296 | 190.5684 | 213.1622 | 215.5408 | 146.1898 | 140.9459 | 140.6755 | 156.7561 | 176.5673 | 285.1357 | 300.2429 | 317.5939 | 409.5949 | 280.5408 | 328.0806 | 320.1082 | 420.7031 | 141.5143 | 80.2704 | 82.5133 | 141.5143 | 122.4051 | 106.1092 | 140.1877 | 151.1357 | 119.7296 | 148.2153 | 140.8378 | 141.8378 | 100.8378 | 114.8102 | 90.1082 | 114.1347 | 123.5663 | 98.4061 | 111.1622 | 87.5133 | 312.4857 | 251.7857 | 287.5408 | 356.053 | 221.5133 | 245.0541 | 237.9724 | 281.351 | 322.8918 | 266.9724 | 299.0541 | 335.2439
820 | 11.0144 | 192.0708 | 278.5354 | 265.5266 | 206.2567 | 169.7301 | 205.1903 | 201.7877 | 202.1947 | 177.0752 | 151.2655 | 148.4778 | 194.8009 | 266.5354 | 191.8673 | 206.6062 | 229.7965 | 214.3319 | 172.7345 | 176.1239 | 236.5398 | 209.1416 | 350.0044 | 391.469 | 196.9292 | 285.2079 | 265.8009 | 277.0664 | 346.4026 | 164.8009 | 149.6593 | 178.4026 | 159.8053 | 184.0664 | 174.7965 | 189.1991 | 116.7389 | 294.1903 | 307.531 | 305.9912 | 354.938 | 167.2035 | 164.6637 | 172.3275 | 158.0664 | 103.7345 | 72.531 | 89.3363 | 159.8717 | 269.2655 | 238.7965 | 306.938 | 203.0044 | 180.1371 | 183.2699 | 188.1991 | 160.4734 | 214.593 | 208.7301 | 256.1327 | 192 | 155.2655 | 155.9956 | 220.0176 | 167.6681 | 202 | 225.3894 | 293.3407 | 258.0044 | 185.938 | 180.1991 | 286.1283 | 237.3938 | 220.062 | 260.7345 | 298.4114 | 360.0044 | 195.5974 | 196.5974 | 207.2655 | 304.0088 | 183.8673 | 165.531 | 157.1947 | 172.7345 | 218.5974 | 216.3319 | 223.3407 | 166.7345 | 139.8009 | 135.3982 | 136.1947 | 168.6681 | 304.4026 | 288.8584 | 292.1371 | 421.2611 | 290.531 | 311.6018 | 322.1991 | 431.2079 | 167.0708 | 84.5354 | 74.3982 | 167.0708 | 113.1947 | 142.469 | 93.1903 | 168.6769 | 115.2655 | 118.2611 | 138 | 138.3363 | 97.8009 | 92.2079 | 93.062 | 100.6593 | 82.3938 | 121.4734 | 112.3407 | 81.1239 | 286.531 | 314.7477 | 296.0044 | 322.0664 | 213.4646 | 247.3938 | 221.1947 | 270.7965 | 319.6062 | 249.9956 | 298.938 | 355.1416
821 | 11.0254 | 178.6067 | 269.4041 | 285.7414 | 238.3147 | 180 | 234.2694 | 239.3481 | 218.6067 | 154.0119 | 154.6853 | 116.9332 | 192.0787 | 255.9881 | 189.764 | 188.416 | 242.6519 | 218.1347 | 172.7759 | 201.2015 | 216.5506 | 185.2026 | 334.1907 | 386.2026 | 207.9547 | 261.1013 | 263.472 | 277.292 | 337.2134 | 160.2694 | 171.3481 | 168.4267 | 147.4839 | 184.4494 | 182.3028 | 190.9774 | 100.7306 | 320.9095 | 314.6067 | 331.6401 | 340.0787 | 157.7306 | 174 | 189.3481 | 157.3481 | 100.3147 | 80.236 | 81.18 | 144.708 | 270.0894 | 248.4267 | 291.764 | 190.9666 | 166.6627 | 174.7306 | 191.0787 | 139.6293 | 236.2134 | 220.3373 | 260.1239 | 194.1239 | 160.9666 | 170.0787 | 166.3599 | 163.708 | 202.708 | 255.0668 | 266.6174 | 247.2252 | 175.1692 | 185.0453 | 300.0334 | 255.584 | 234.8427 | 254.6401 | 262 | 345.764 | 203.292 | 205.6293 | 209.1907 | 276.7866 | 179.4827 | 172.4494 | 172.9774 | 170.1013 | 223.9321 | 220.764 | 202.9892 | 162.6067 | 135.82 | 140.8427 | 152.9213 | 163.6853 | 296.7866 | 312.9547 | 283.9332 | 438.8427 | 298.6293 | 303.0894 | 325.236 | 402.4601 | 155.2586 | 78.472 | 78.82 | 155.2586 | 128.584 | 133.8987 | 125.5733 | 137.6519 | 120.18 | 134.5054 | 138.3933 | 129.3147 | 94.4494 | 69.2813 | 107.292 | 124.3933 | 102.8653 | 102.9106 | 89.8653 | 109.6627 | 291.4827 | 266.9547 | 282.3933 | 321.2694 | 218.3481 | 265.584 | 237.764 | 281.764 | 300.5506 | 264.7866 | 282.2694 | 327.9774
822 | 11.0364 | 173.2724 | 239.9096 | 271.3638 | 242.5457 | 182.2739 | 213.0904 | 217.7261 | 211.2713 | 198.0894 | 147.9075 | 159.9985 | 193.4548 | 210.6372 | 185.3628 | 203.3628 | 261.6362 | 209.0915 | 211.6336 | 146.5462 | 223.7276 | 212.0005 | 318.9121 | 406.1789 | 184 | 272.5447 | 267.1809 | 271.8196 | 327.0904 | 141.2734 | 149.7261 | 150.4548 | 198.7266 | 180.092 | 143.8206 | 181.9106 | 120.8181 | 273.6367 | 311.2734 | 307.9111 | 336.8161 | 178.9096 | 173.909 | 169.7271 | 139.7282 | 102.4538 | 85.2739 | 94.1814 | 153.1819 | 235.0904 | 235.1845 | 289.5457 | 209.8161 | 153.9106 | 192.7261 | 190.7266 | 147.4548 | 223.6347 | 231.6341 | 284.0904 | 218.5452 | 183.6362 | 171.091 | 193.5452 | 169.0894 | 209.9995 | 210.2734 | 231.5478 | 284.909 | 221.908 | 208.6362 | 278.4553 | 240.0005 | 232.4543 | 228.0005 | 262.091 | 343.1819 | 195.3638 | 212.3633 | 191.4563 | 288.6377 | 154.4568 | 168.6377 | 163.092 | 183.9096 | 188.4548 | 215.6352 | 233.1794 | 158.8186 | 125.0925 | 139.2729 | 150.9085 | 161.0005 | 305.5452 | 290.7282 | 327.4533 | 433.1799 | 306.6367 | 271.3648 | 326.3628 | 365.4568 | 168.9085 | 85.7282 | 65.8186 | 168.9085 | 112.1819 | 122.4553 | 144.2729 | 157.1819 | 133.3633 | 117.5457 | 140.5442 | 135.3648 | 88.091 | 94.3628 | 99 | 128.5452 | 113.3638 | 134.8171 | 101.9101 | 94.6367 | 263.3643 | 247.8206 | 289.092 | 302.8191 | 200.6372 | 247.9995 | 234.8181 | 279 | 307.8186 | 275.5462 | 262.3638 | 320.093
823 | 11.0474 | 166.3097 | 251.6844 | 275.2065 | 246.0383 | 202.9351 | 203.6903 | 194.174 | 183.0708 | 166.8318 | 118.6519 | 132.9941 | 199.9617 | 228.239 | 169.6195 | 188.3422 | 257.2773 | 218.761 | 164.298 | 165.5104 | 230.4838 | 218.4896 | 377.0708 | 349.4454 | 186.7877 | 255.5221 | 251.7877 | 298.3481 | 316.6578 | 152.8643 | 126.6903 | 153.761 | 186.413 | 201.413 | 190.5546 | 208.2123 | 117.2448 | 283.2006 | 323.4838 | 346.7935 | 297.8909 | 185.9351 | 172.8968 | 166.8968 | 157.3097 | 88.5487 | 106.8643 | 87.6519 | 154.8968 | 230.8526 | 288.239 | 295.826 | 166.5546 | 183.413 | 171.7552 | 177.8968 | 152.7227 | 194.5428 | 187.0708 | 274.4838 | 214.9292 | 179.3805 | 171.0708 | 187.6578 | 142.4779 | 201.0325 | 221.9675 | 277.413 | 282.1416 | 197.7994 | 202.9351 | 293.9292 | 249.587 | 229.9941 | 239.0325 | 264.4454 | 342.9351 | 197.8643 | 202.5929 | 222.1799 | 311.4896 | 196.4188 | 192.1091 | 186.3747 | 191.1416 | 194.0325 | 192.1032 | 191.0915 | 165.2448 | 153.4188 | 143.0325 | 140.5162 | 169.9675 | 295.2182 | 305.9351 | 303.6519 | 396.239 | 312.8968 | 294.4838 | 311.0325 | 408.2448 | 158 | 101.0383 | 75.239 | 158 | 116.5812 | 134.9351 | 144.1091 | 154.7669 | 126.4838 | 124.1357 | 117.8584 | 156.9351 | 89.413 | 80.4779 | 95.9026 | 125.7552 | 117 | 114.239 | 119.8643 | 99.5546 | 275.8643 | 295.587 | 306.9026 | 321.826 | 214.7286 | 235.9351 | 238.2656 | 280.4454 | 315.174 | 283.876 | 265.4838 | 355.3215
824 | 11.0584 | 170.6175 | 271.765 | 275.6136 | 225.8805 | 184.3068 | 201.2311 | 188.7331 | 178.3506 | 151.6972 | 133.8088 | 158.1912 | 211.8008 | 242.3506 | 175.3466 | 201.7331 | 249.498 | 206.7291 | 215.8805 | 195.0757 | 226.5777 | 191.6972 | 369.3466 | 364.5418 | 210.7689 | 238.8526 | 275.6534 | 286.4263 | 308.3506 | 143.8088 | 126.0797 | 143.6932 | 187.1116 | 201.4183 | 176.1514 | 183.6534 | 101.7689 | 306.1474 | 317.6136 | 343.0359 | 320.3426 | 169.1554 | 170.498 | 164.9602 | 158.6136 | 102.6932 | 94.8048 | 103.2709 | 155.502 | 272.3068 | 295.765 | 298.3785 | 155.9641 | 187.4621 | 189.6574 | 177.3466 | 158.1912 | 229.6096 | 180.7331 | 265.494 | 221.0359 | 174.5777 | 164.2709 | 175.1912 | 162.8446 | 209.7689 | 213.3466 | 279.6136 | 266.5777 | 171.2709 | 187.0797 | 297.9562 | 251.6613 | 250.6853 | 244.3028 | 273.1873 | 324.7689 | 188.8088 | 174.8486 | 191.8884 | 281.6932 | 173.7689 | 169.4263 | 200.761 | 176.3864 | 202.7689 | 194.1554 | 237.6454 | 151.1554 | 131.5777 | 151.0757 | 145.1155 | 161.1155 | 251.0079 | 288.8088 | 314.8805 | 408.2709 | 309.2271 | 286.6494 | 321.1554 | 397.8526 | 157.5379 | 84.6932 | 85.0757 | 157.5379 | 136.3426 | 115.6136 | 119.3466 | 122.506 | 123.6175 | 133.8845 | 132.8446 | 139.2311 | 92.5379 | 101.5379 | 72.004 | 139.6136 | 116.8845 | 122.1116 | 110.6932 | 88.7331 | 268.6574 | 286.3387 | 284.506 | 330.3864 | 196.9641 | 219.1554 | 277.6454 | 289.1873 | 307.6534 | 227.4303 | 262.502 | 311.1992
825 | 11.0694 | 184.4061 | 249.6844 | 263.8495 | 204.6844 | 178.3609 | 203.767 | 200.8495 | 203.3609 | 178.5939 | 140.6391 | 153.2783 | 173.6618 | 263.0147 | 182.474 | 212.5713 | 237.3609 | 200.8349 | 194.3009 | 186.8495 | 231.3835 | 217.0599 | 370.0826 | 385.2783 | 207.5939 | 276.3835 | 273.767 | 301.0373 | 333.3609 | 149.233 | 143.1505 | 150.1505 | 166.4061 | 176.0453 | 164.3009 | 183.9401 | 102.1731 | 270.7896 | 306.233 | 327.8722 | 302.8122 | 175.9547 | 161.0453 | 161.7297 | 149.6618 | 105.8269 | 79.9174 | 112.233 | 164.1878 | 262.3982 | 270.4887 | 264.4061 | 172.9174 | 189.4661 | 205.0147 | 180.2557 | 150.4661 | 197.0226 | 196.8122 | 233.5566 | 207.9174 | 180.1505 | 175.5339 | 172.3235 | 152.3382 | 206.4661 | 217.2783 | 266.5713 | 268.3156 | 182.0226 | 203.1278 | 270.3235 | 292.3156 | 214.2183 | 222.7749 | 244 | 325.1731 | 197.1731 | 191.0453 | 214.2557 | 287.1278 | 172.767 | 185.4434 | 151.4661 | 188.2783 | 197.0373 | 200.9547 | 188.6165 | 157.3156 | 135.1052 | 142.2104 | 146.5113 | 160.7217 | 294.0599 | 293.5939 | 292.4061 | 414.6765 | 289.5339 | 263.9401 | 329.1052 | 428.8642 | 156.6844 | 90 | 78 | 156.6844 | 113.9547 | 107.4287 | 123.2783 | 157 | 134.3382 | 132.6165 | 125.2783 | 138.3156 | 90.7896 | 96.8495 | 95.3835 | 127.3382 | 116.8948 | 103.0679 | 116.3835 | 101.1052 | 287.5939 | 244.6618 | 314.526 | 335.2477 | 211.1052 | 229.0226 | 232.8349 | 266.1358 | 306.7896 | 269.8948 | 274.5113 | 353.0453
826 | 11.0804 | 192.879 | 263.5162 | 257.121 | 224.4012 | 175.5413 | 210.4012 | 192.7198 | 199.2802 | 171.5222 | 141.9174 | 158.2168 | 183.1784 | 230.2611 | 208.9365 | 195.8599 | 237.3437 | 222.3761 | 209.1593 | 178.5796 | 233.7198 | 197.1401 | 355.3186 | 389.0958 | 198.8407 | 270.8731 | 274.3761 | 272.0383 | 328.4395 | 144.1401 | 152.6814 | 157.2994 | 174.879 | 187.618 | 180.1401 | 201.3377 | 115.1784 | 289.4455 | 303.5222 | 323.6371 | 323.121 | 164.382 | 175.8407 | 191.3377 | 161.4204 | 93.8024 | 90.4778 | 104.4778 | 147.2419 | 236.7449 | 264.3377 | 276.1018 | 184.879 | 177.879 | 172.2611 | 181.8599 | 140.8407 | 207.7832 | 215.8599 | 248.618 | 220.5796 | 184.9174 | 185.1593 | 191.7773 | 142.6814 | 197.4012 | 224.3186 | 248.5988 | 252.9425 | 187.879 | 200.3569 | 292.0192 | 273.8599 | 241.4587 | 266.3377 | 248.6239 | 341.9617 | 211.5796 | 198.6946 | 215.1593 | 288.2802 | 176.1784 | 170.382 | 142.6814 | 197.2802 | 174.9042 | 187.7006 | 186.9808 | 141.6623 | 129.2802 | 130 | 147.3377 | 155.6431 | 277.6431 | 284.7006 | 303.6814 | 391.1593 | 299.1593 | 282.3186 | 327.4838 | 384.618 | 173.879 | 89.5796 | 81.503 | 173.879 | 105.3245 | 124.4204 | 133.2611 | 153.9174 | 122.0192 | 125.5162 | 131.618 | 121.2611 | 101.3186 | 88.5796 | 95.618 | 118.6623 | 123.8407 | 123.4395 | 120.5413 | 93.4587 | 279.2611 | 260.0635 | 286.2419 | 294.879 | 206.5413 | 235.7198 | 256.618 | 307.7198 | 315.6371 | 272.0767 | 275.1976 | 361.3952
827 | 11.0914 | 184.1711 | 236.7621 | 266.4386 | 234.2193 | 185.6096 | 220.5241 | 190.6952 | 199.9331 | 183.647 | 131.2302 | 135 | 174.2862 | 239.2193 | 184.1524 | 195.7621 | 264.3048 | 193.8476 | 203.0855 | 172.647 | 226.9705 | 198.7621 | 342.1042 | 361.591 | 201.2566 | 230.1446 | 250.8771 | 269.4495 | 328.6579 | 145.6096 | 161.1711 | 154.8289 | 163.4277 | 170.5614 | 179.0187 | 186.0669 | 104.3048 | 322.7512 | 316.2566 | 311.0778 | 323.2939 | 177.3235 | 180.3904 | 174.0855 | 164.3048 | 111.1524 | 78.2006 | 87.4759 | 158.2566 | 267.4277 | 245.5614 | 284.8849 | 179.9813 | 166.1229 | 175.7325 | 182.0669 | 145.5428 | 228.5614 | 208.7512 | 238.4199 | 214.647 | 168.8958 | 179.3904 | 174.1337 | 153.7621 | 207.2193 | 220.4012 | 240.4572 | 270.3904 | 183.7434 | 179.0669 | 279.4463 | 271.4759 | 228.0373 | 248.6283 | 274.7138 | 351.409 | 206.409 | 167.0109 | 206.7994 | 286.3421 | 166.8289 | 185.6096 | 155.1337 | 195.647 | 205.1524 | 193.6765 | 199.8585 | 157.6952 | 134.6579 | 129.8476 | 132.3717 | 176.7807 | 296.647 | 288.3904 | 310.3421 | 386.4572 | 292.3235 | 279.1633 | 345.3981 | 377.1633 | 163.647 | 87.3048 | 100.409 | 163.647 | 137.5428 | 125.9331 | 137.9518 | 135 | 121.4868 | 102.1151 | 118.2193 | 133.8771 | 88.8663 | 89.0482 | 84.0482 | 132.2566 | 127.0187 | 118.3235 | 127.5614 | 88.4199 | 284.7138 | 277.2271 | 297.2566 | 293.4868 | 217.9813 | 235.3717 | 236.2084 | 304.4759 | 301.2488 | 248 | 256.0187 | 326.5428
828 | 11.1024 | 178.6116 | 243.6352 | 278.5177 | 240.3412 | 189.9882 | 223.2354 | 191.7999 | 195.9764 | 168.4587 | 172.8825 | 135.1647 | 185.5177 | 245.0118 | 191.4234 | 203.4587 | 269.4587 | 193.1647 | 194.2119 | 159.2705 | 204.5999 | 203.8235 | 329.9528 | 372.706 | 190.2822 | 278.2001 | 278.8235 | 314.9178 | 342.0825 | 148.3412 | 152.4823 | 168.7881 | 147.388 | 159.9646 | 169.9764 | 188.3766 | 110.941 | 291.9528 | 300.1765 | 345.1297 | 290.1175 | 169.647 | 174.2001 | 169.8235 | 164.1883 | 112.1647 | 97.3176 | 82.4587 | 143 | 242.3294 | 234.3058 | 262.3294 | 188.5295 | 142.294 | 154.7999 | 190.1411 | 142.8353 | 216.4823 | 180.2587 | 257.8943 | 197.8117 | 179.8943 | 173.8589 | 187.3176 | 158.1647 | 216.1411 | 257.8117 | 245.1411 | 271.4469 | 200.9764 | 188.1293 | 248.1057 | 257.7295 | 212.5881 | 244.6234 | 261.5059 | 335.353 | 192.4941 | 202.5413 | 192.4587 | 273.2469 | 175.6824 | 185.8707 | 163.706 | 179.1411 | 208.7999 | 204.4823 | 230.5884 | 157.3176 | 151.2119 | 131.1411 | 140 | 170.8235 | 279.4823 | 283.353 | 289.6706 | 394.1057 | 283.3294 | 315.5413 | 297.4352 | 413.706 | 145.9882 | 87.5177 | 83.8589 | 145.9882 | 135.8235 | 123.4587 | 123.1411 | 138.6234 | 151.1765 | 125.9882 | 123.8471 | 153.9178 | 79.4705 | 102.706 | 97.0472 | 115.0236 | 112.706 | 106.5295 | 113.1765 | 110.8589 | 274.6352 | 238.6116 | 281.0118 | 320.0472 | 222.5766 | 242.3412 | 211.0822 | 297.3176 | 335.5059 | 253.9293 | 242.6942 | 328.7763
829 | 11.1134 | 197.7522 | 260.5399 | 269.2301 | 233.6371 | 191.9911 | 195.5929 | 201.3982 | 203.8761 | 185.292 | 150.7611 | 139.8938 | 179.062 | 241.4159 | 217.8053 | 217.1062 | 288.062 | 201.0797 | 175.177 | 187.4159 | 230.6991 | 203.5752 | 345.9292 | 355.469 | 210.0531 | 270.7257 | 280.3451 | 276.8761 | 305.4071 | 147.4779 | 154.8672 | 181.9822 | 191.6371 | 173.5841 | 178.0531 | 170.885 | 128.9203 | 309.1681 | 299.1859 | 293.0089 | 314.3629 | 174.4601 | 157.6991 | 171.3451 | 155.354 | 111.2301 | 99.7522 | 97.6991 | 146.5399 | 251.0797 | 240.3362 | 257.8053 | 180.4248 | 157.938 | 172.3629 | 196.3982 | 141.115 | 219.2212 | 215.1859 | 235.1859 | 205.5841 | 155.062 | 163.354 | 192.5841 | 157.938 | 221.6902 | 260.2743 | 249.9822 | 286.0353 | 208.5221 | 204.885 | 269.2212 | 234.0089 | 248.3009 | 258.2212 | 254.3982 | 333.7257 | 194.115 | 187.8319 | 205.2212 | 299.5574 | 169.708 | 168.938 | 150.8938 | 187.8761 | 219.9911 | 212.531 | 201.6726 | 163.469 | 136.4248 | 143.2389 | 136.2832 | 176.062 | 293.3717 | 272.1681 | 289.4159 | 421.7699 | 277.5663 | 295.5221 | 321.2832 | 399.6549 | 153.8319 | 83.1859 | 76.5399 | 153.8319 | 136.6371 | 136.9292 | 133.823 | 153.6371 | 144.3451 | 129.7611 | 115.1681 | 122.2478 | 89.1681 | 86.885 | 83.1947 | 107.4779 | 100.2478 | 99.1328 | 108.292 | 102.6549 | 291.7168 | 263.4159 | 279.5399 | 301.5929 | 191.5929 | 242.5399 | 241.5132 | 300.8141 | 329.6371 | 282.0531 | 246.3984 | 352.823
830 | 11.1244 | 181.5964 | 269.6179 | 255.5108 | 218.5679 | 179.5143 | 204.325 | 181.8928 | 194.5679 | 184.0821 | 144.8107 | 160.2714 | 180.9715 | 248.3786 | 185.9715 | 195.9715 | 282.4001 | 231.9179 | 174.4857 | 194.3786 | 224.5428 | 184.7036 | 326.0536 | 347.7001 | 198.7286 | 293.2929 | 272.3535 | 267.7571 | 305.9499 | 158.6214 | 137.0571 | 151.1607 | 177.325 | 172.4892 | 168.8107 | 176.1357 | 99.4071 | 298.4607 | 310.9715 | 307.5679 | 332.0821 | 158.3786 | 150.0285 | 161.4607 | 158.7036 | 103.7571 | 82.0821 | 87.5679 | 157.1322 | 274.5358 | 219.4892 | 234.4892 | 197.025 | 155.2714 | 184.2143 | 176.8928 | 137.3786 | 196.8107 | 222.2394 | 251.8928 | 202.0285 | 163.7857 | 162.7286 | 185.0536 | 152.4321 | 204.3821 | 224.0786 | 225.7857 | 256.0035 | 192.8358 | 206.35 | 250.7857 | 246.1072 | 259.9464 | 239.7857 | 239.8142 | 359.3215 | 192.8393 | 197.025 | 186.975 | 254.7857 | 169.5929 | 160.5929 | 171.0821 | 180.4607 | 202.025 | 215.5679 | 238.2143 | 158.35 | 152.0786 | 150.1357 | 124.4892 | 174.1858 | 317.8893 | 263.3535 | 291.6465 | 422.4287 | 257.3856 | 283.4321 | 307.5964 | 414.0821 | 165.2964 | 97.4321 | 92.0536 | 165.2964 | 122.5143 | 121.0285 | 132.4321 | 138.9464 | 129.1607 | 122.675 | 107.6785 | 135.9429 | 80.5428 | 92.7036 | 110.1072 | 117.8643 | 120.1893 | 118.8358 | 105 | 111.7822 | 306.1607 | 266.2143 | 296.3786 | 308.2429 | 198.8107 | 252.9429 | 216.3786 | 282.7822 | 320.4356 | 267.5108 | 309.2929 | 350.675
831 | 11.1354 | 194.7827 | 241.8299 | 238.2488 | 218.9843 | 186 | 212.0157 | 191.8142 | 199.2173 | 190.3795 | 139.3638 | 165.5654 | 168.6047 | 248.1858 | 173.8063 | 183.2016 | 242.8457 | 219.7748 | 171.0236 | 190.1543 | 242.1858 | 192.2016 | 331.8772 | 330.0394 | 192.2252 | 258.6205 | 280.9213 | 273.4189 | 344.0315 | 153.1701 | 159.3717 | 137.9921 | 183.4031 | 191.7669 | 171.4268 | 181.4031 | 113.3953 | 302.9528 | 300.6205 | 306.3717 | 336.9685 | 160.6047 | 164.0079 | 168.7748 | 167.8142 | 104.3795 | 91.4031 | 89.1937 | 134.4189 | 234 | 242.3953 | 254.7748 | 184.6362 | 172.4583 | 170.822 | 180.9685 | 139 | 194.1858 | 196.6598 | 257.178 | 206.7354 | 181.411 | 161.4661 | 182.6126 | 150.2016 | 233.3795 | 205.6205 | 243.0079 | 281.3638 | 174.5969 | 198.2252 | 270.4268 | 243.2409 | 256.9449 | 247.5339 | 262.1465 | 341.0472 | 202.9843 | 182.0157 | 206.2173 | 271.2016 | 153.9921 | 153.4583 | 180.8063 | 185.1543 | 186.0079 | 215.5811 | 219.3795 | 154.8614 | 135.2331 | 150.5654 | 153.6441 | 153.2646 | 293.178 | 281.1937 | 279.4504 | 387.9843 | 314.3638 | 281.8063 | 318.3638 | 416.1465 | 154.9764 | 96.0079 | 86.9921 | 154.9764 | 130.2094 | 138.0315 | 135.2409 | 139.3717 | 117.4031 | 118.8142 | 132.4031 | 116.2488 | 98.7906 | 95.3638 | 98.9764 | 112.1937 | 119.589 | 104.6283 | 107.6205 | 96.3874 | 299.8457 | 253.0236 | 299.2094 | 305.2016 | 196.1858 | 235.6677 | 223.4425 | 281.4976 | 341.9528 | 249.4425 | 263.5339 | 344.7984
832 | 11.1464 | 183.2197 | 258.4971 | 257.0118 | 209.0752 | 183.0059 | 215.9365 | 205.416 | 200.9424 | 177.4336 | 124.208 | 143.0811 | 174.6357 | 242.0693 | 176.1445 | 189.3466 | 265.3584 | 210.208 | 181.2891 | 171.2773 | 237.7803 | 190.2197 | 365.4395 | 355.3466 | 203.7168 | 267.9307 | 242.2832 | 279.3643 | 358.0752 | 139.2832 | 144.5723 | 135.7109 | 180.0811 | 176.9248 | 180.7227 | 179.792 | 108.4336 | 281.1387 | 307.5781 | 296.9189 | 319.0752 | 162.1445 | 168.2139 | 157.2832 | 173.5029 | 94.4277 | 90.8613 | 85.3584 | 142.0752 | 236.1387 | 235.9365 | 246.9189 | 202.5723 | 196.4395 | 178.3643 | 171.2021 | 141.7803 | 194.913 | 218.4454 | 248.1387 | 179.9882 | 185.7168 | 190.5088 | 187.5029 | 145.6534 | 212.9483 | 217.0693 | 247.2139 | 264.0693 | 172.1445 | 206.7227 | 276.9424 | 259.295 | 234.7744 | 214 | 240.3466 | 359.0118 | 195.8555 | 188.5029 | 211.792 | 271.7861 | 152.7803 | 179.3643 | 181.8613 | 164.3525 | 197.4853 | 204.7168 | 206.0059 | 180.0118 | 152.2139 | 135.1387 | 172.3643 | 181.295 | 280.0752 | 275.2197 | 302.8613 | 383.6357 | 298.1387 | 280.0811 | 297.4336 | 386.6475 | 142.3584 | 94.4395 | 86.208 | 142.3584 | 133.2891 | 146.7286 | 147.4454 | 128.4219 | 117.5029 | 126.2139 | 132.2891 | 139.7168 | 92.2891 | 81.9189 | 84.0059 | 111.5664 | 115.9248 | 116.8613 | 113.295 | 90.8555 | 320.0059 | 261.3643 | 300.5781 | 310.4912 | 191.7803 | 259.0928 | 243.7168 | 325.4395 | 316.5029 | 262.6592 | 236.8437 | 339.5088
833 | 11.1574 | 170.06 | 258.5722 | 235.6785 | 206.297 | 168.1553 | 202.3569 | 223.5722 | 179.297 | 168.9046 | 140.0708 | 132.0109 | 189.1308 | 245.0954 | 175.4877 | 209.6785 | 259.3815 | 220.4169 | 179.5231 | 188.8093 | 242.5722 | 186.1062 | 345 | 361.6567 | 197.4169 | 270.1062 | 255.594 | 270.0708 | 342.1799 | 137.8938 | 143.6785 | 144.2616 | 163.8093 | 182.6431 | 166.3569 | 167.0954 | 99.4523 | 300.7602 | 295.6785 | 312.2861 | 315.3924 | 160.8093 | 167.4169 | 162.9046 | 164.3815 | 98.7139 | 87.9755 | 82.7738 | 136.1308 | 241.5123 | 231.1771 | 268.6185 | 196.2507 | 177.1308 | 166.1308 | 191.2861 | 150.1907 | 220.1553 | 186.5477 | 253.06 | 201.0954 | 173.763 | 175.8093 | 183.3569 | 135.4986 | 178.2616 | 217.8338 | 255.9155 | 264.1553 | 174.4278 | 188.06 | 259.1416 | 248.1062 | 248.5368 | 218.5231 | 256.1553 | 341.5231 | 202.8447 | 189.7847 | 191.8584 | 276.6539 | 162.0954 | 162.3815 | 174 | 178.6539 | 217.3107 | 190.0491 | 193.1907 | 158.4523 | 150.2861 | 142.0954 | 155.3815 | 158.3461 | 278.2016 | 268.8447 | 290.9292 | 406.4986 | 306.6785 | 271.7247 | 290.7139 | 383.2016 | 141.5831 | 77.6185 | 89.6322 | 141.5831 | 125.4169 | 117.5123 | 120.5231 | 144.6185 | 113.1308 | 124.2861 | 123.5123 | 136.1308 | 88.7139 | 99.0954 | 76.6185 | 122.1308 | 120.06 | 109 | 100.297 | 93.0954 | 304.9292 | 247.5477 | 289.3569 | 326.7738 | 201.3215 | 221.2616 | 238.7738 | 295.2753 | 309.8692 | 240.8201 | 261.0954 | 325.2616
834 | 11.1684 | 165.0074 | 250.2533 | 230.3697 | 221.7153 | 155.7153 | 202.7079 | 199.2772 | 196.8614 | 166.5157 | 155.4307 | 153.3771 | 186.8004 | 238.4233 | 186.3845 | 205.3235 | 245.8078 | 213.5693 | 195.4769 | 192.3771 | 222.0924 | 206.7616 | 338.3235 | 320.769 | 188.6617 | 285.5157 | 285 | 285.9076 | 313.4307 | 126.2995 | 149.5767 | 149.854 | 170.2384 | 181.3383 | 172.4307 | 167.8152 | 102.4307 | 336.7079 | 302.5305 | 304.6691 | 321.5157 | 166.5231 | 166.0536 | 167.6691 | 156.5305 | 105.854 | 102.854 | 82.9538 | 143.5767 | 239.1996 | 260.2772 | 275.2772 | 183.0998 | 174.8004 | 169.0462 | 175.085 | 141.3771 | 207.4769 | 187.146 | 243.9538 | 197.5231 | 151.3845 | 178.1848 | 187.0462 | 165.1386 | 193.8688 | 215.8226 | 276.5619 | 258.1534 | 184.8928 | 178.9538 | 288.6691 | 264.4695 | 227.2921 | 234 | 244.6691 | 349.8466 | 210.5157 | 208.4695 | 168.0536 | 290.0388 | 164.4843 | 154.5305 | 172.3309 | 193.7079 | 187.9002 | 173.8836 | 193.915 | 158.5693 | 142.4307 | 138.7616 | 145.3845 | 145.1996 | 297.4307 | 285.8152 | 280.8152 | 427.0776 | 302.5619 | 304.2847 | 302.8614 | 399.5693 | 148.4307 | 86.9002 | 66.4843 | 148.4307 | 117.1386 | 117.3457 | 132.4233 | 150.8004 | 121.2921 | 118.0998 | 112.6155 | 140.9538 | 95.854 | 94.0924 | 87.5693 | 122.1848 | 108.5693 | 107.0924 | 115 | 86.9002 | 287.4233 | 238.3697 | 282.5545 | 315.5083 | 199 | 237.1072 | 237.7616 | 272.7302 | 305.5231 | 269.8078 | 255.1386 | 329.6617
835 | 11.1794 | 179.2448 | 271.9675 | 221.2951 | 224.2507 | 163.5163 | 187.7463 | 185.2626 | 199.23 | 158.7788 | 160 | 139.9852 | 174.2714 | 235.5251 | 188.7345 | 183.4985 | 250.5251 | 206.9941 | 193.2389 | 195.2833 | 215.7493 | 207.2537 | 319.0059 | 333.9941 | 180.2567 | 266.4956 | 278.9823 | 286.4837 | 324.0177 | 154.7463 | 157.9911 | 138.9882 | 174.7611 | 187.9882 | 174.7433 | 178.2448 | 104.9941 | 312.7197 | 304.4631 | 311.5044 | 308.764 | 163.7463 | 171.9793 | 172.4985 | 169.2448 | 101.0059 | 98.2567 | 87.2537 | 152.9941 | 255.4897 | 246.0118 | 252.236 | 189.7404 | 158.0089 | 168.5074 | 157.2626 | 144.0325 | 205.7404 | 196.5074 | 242.7374 | 198.7581 | 162.7611 | 160.4837 | 188.764 | 154.4956 | 205.9616 | 242.4956 | 252.7404 | 275.7345 | 169.531 | 185.2596 | 293.2478 | 245.236 | 234.9823 | 236.2567 | 252.0059 | 336.7788 | 189.7404 | 197.2596 | 191.531 | 278.2714 | 176.4837 | 166.7433 | 170.7611 | 176.2389 | 183.2655 | 227.9911 | 214.2448 | 163.0266 | 143.2389 | 132.7345 | 147.7345 | 160.7374 | 299.9941 | 296.4956 | 290.9941 | 397.0236 | 287.5163 | 312.7818 | 309.4926 | 386.9764 | 150.4926 | 79.5074 | 86.0059 | 150.4926 | 108.7522 | 137.7315 | 120.2478 | 130.4985 | 133.7463 | 127.2478 | 114.0266 | 142.997 | 90.7552 | 88.5015 | 76.23 | 111.003 | 106.0059 | 99.7463 | 112.2419 | 82.767 | 278.7581 | 222.7759 | 249.5103 | 291.5281 | 192.23 | 249.7108 | 244.0207 | 309.2567 | 307.0089 | 267.2537 | 242.7404 | 323.5133
836 | 11.1904 | 173.4739 | 236.6327 | 247.7988 | 221.3176 | 173.1612 | 190.0497 | 194.0545 | 185.3127 | 184.8436 | 159.737 | 122.1067 | 188.6352 | 262.3152 | 174.8412 | 180.4218 | 271.5285 | 203.1042 | 178.8436 | 221.3721 | 217.8933 | 209.1588 | 332.3648 | 330.8933 | 191.4715 | 265.9455 | 268.6279 | 268.1588 | 345.9455 | 154.4194 | 145.8436 | 127.7891 | 180.7394 | 175.2109 | 165.1067 | 172.2109 | 100.0521 | 288.8388 | 278.5733 | 319.9455 | 323.7891 | 164.7345 | 153.8933 | 170.737 | 163.2109 | 110.9455 | 101.8436 | 93.3673 | 145.4218 | 244.737 | 256.6848 | 246.417 | 187.6279 | 168.8412 | 170.7394 | 172.3673 | 172.7915 | 197.0521 | 207.4194 | 233.4194 | 203.3176 | 168.4764 | 145.0521 | 196.6873 | 147.6327 | 177.2582 | 238.5261 | 247.4715 | 267.6279 | 194.9503 | 194.2109 | 283.1091 | 235.2085 | 217 | 242.2109 | 246.1636 | 360.4764 | 191.8364 | 204.1067 | 215.1091 | 290.5309 | 159.737 | 163.9455 | 181.2109 | 176.3103 | 197.4218 | 212.9503 | 210.3152 | 182.6352 | 139.1018 | 125.417 | 128.0545 | 153.2606 | 289.5285 | 287.0024 | 277.3721 | 412.057 | 302.9479 | 327.64 | 300.9479 | 377.2036 | 142.2109 | 85.9479 | 87.0024 | 142.2109 | 115.2085 | 120.5782 | 119.3152 | 132.6824 | 128.4218 | 130.2606 | 133.3721 | 136.5806 | 91.2655 | 82.1091 | 65.2061 | 117.6824 | 115.4194 | 103.6279 | 104.7891 | 95.0024 | 290.6824 | 254.2606 | 254.4764 | 318.1588 | 181.9951 | 216.5758 | 263.6848 | 309.1612 | 318.8933 | 260.2158 | 240.3648 | 337
837 | 11.2014 | 175.5792 | 250.3738 | 206.1141 | 210.7247 | 165.2882 | 205.8987 | 181.3481 | 199.9715 | 179.3039 | 150.465 | 121.709 | 176.535 | 256.3638 | 183.0285 | 182.0571 | 251.8417 | 206.7974 | 172.4779 | 193.9429 | 227.6519 | 206.2026 | 344.7404 | 343.9558 | 202.0728 | 269.0157 | 289.4494 | 258.8702 | 345.4365 | 162.3481 | 141.4051 | 131.9272 | 166.5506 | 178.7817 | 157.2753 | 171.6519 | 104.9272 | 309.9272 | 306.0442 | 319.7118 | 326 | 173.0442 | 161.7247 | 167.2468 | 156.5949 | 114.8417 | 94.3766 | 98.0728 | 134.9429 | 244.826 | 239.7846 | 268.0728 | 204.8702 | 167.3923 | 159.5792 | 180.3766 | 159.0728 | 206.3324 | 207.6391 | 247.4051 | 199.883 | 154.9272 | 155.709 | 181.2026 | 150.6362 | 202.1455 | 232.0157 | 256.1455 | 272.7561 | 182.1298 | 187.3195 | 266.5064 | 247.826 | 223.883 | 239.7247 | 222.0856 | 347.4779 | 215.9144 | 196.2753 | 194.9272 | 265.1013 | 168.3609 | 168.6676 | 175.1455 | 200.8417 | 191.3481 | 195.7247 | 210.9715 | 161.7247 | 147.7118 | 144.5949 | 121.9558 | 167.7247 | 280.8545 | 276.0285 | 261.5064 | 384.5221 | 297.2468 | 283.1298 | 299.1013 | 402.1327 | 145.7817 | 88.7817 | 78.7817 | 145.7817 | 130.0285 | 131.0571 | 127.4051 | 146.3039 | 126.4779 | 140.8702 | 115.8545 | 126.4494 | 79.6519 | 66.0571 | 89.0728 | 120.5663 | 121.9715 | 119.2183 | 110.0285 | 82.6519 | 290.2625 | 262.9429 | 235.1455 | 319.6077 | 203.9429 | 235.2753 | 252.5663 | 292.2026 | 314.3352 | 231.6519 | 260.4494 | 339.7532
838 | 11.2124 | 182.1239 | 293.0707 | 232.823 | 217.7655 | 193.8496 | 201.5619 | 183.6151 | 192.2876 | 181.3849 | 136.0531 | 140.5619 | 196.9867 | 225.823 | 186.3982 | 194.5487 | 230.4115 | 203.0133 | 181.0133 | 184.6151 | 227.9734 | 208.261 | 323 | 364.4779 | 197.5752 | 248.2876 | 274.7698 | 256.0531 | 317.3009 | 147.646 | 149.8363 | 139.0133 | 172.0266 | 191.1239 | 161.1637 | 174.5619 | 109.7124 | 298.3185 | 289.8363 | 298.0531 | 315.0707 | 162.5885 | 164.4513 | 158.8496 | 142.9867 | 101.7522 | 97.7124 | 94.4381 | 128.7788 | 246.4248 | 209.8894 | 264.7257 | 189.3982 | 161.7921 | 168.1372 | 171.3451 | 150.261 | 222.7124 | 194.5045 | 250.9469 | 211.0973 | 165.7522 | 167.084 | 183.8363 | 173.9867 | 191.7124 | 213.3009 | 249.1637 | 237.4779 | 191.2743 | 183.916 | 285.6018 | 247.6991 | 233.6592 | 253.3806 | 245.7124 | 351.1239 | 195.1637 | 196.4248 | 198.2743 | 271.1637 | 180.0575 | 151.1372 | 164.1372 | 192.354 | 191.3142 | 195.8628 | 202.7124 | 160.7124 | 119.4381 | 132.1372 | 132.4115 | 162.6858 | 290.7124 | 284.8628 | 280.8894 | 377.4248 | 283.3849 | 296.0133 | 309.7655 | 376.5178 | 154.6725 | 101.1239 | 94.8628 | 154.6725 | 135.1239 | 135.4955 | 132.6725 | 159.8894 | 131.8496 | 135.7522 | 130.8894 | 127.0133 | 89.177 | 85.739 | 86.3009 | 99.2743 | 111.4115 | 106.5885 | 117.7124 | 90.1637 | 274.0797 | 246.1372 | 230.7522 | 326.0707 | 196.6284 | 233.9867 | 232.4248 | 288.7964 | 294.5178 | 232.8363 | 258.4248 | 343.5487
839 | 11.2234 | 173.8004 | 254.0039 | 218.5005 | 225.9036 | 187.7011 | 196.8004 | 202.0993 | 186.1022 | 174.297 | 154.8004 | 134.3009 | 187.5015 | 231.2941 | 181.5979 | 185.4002 | 237.4956 | 207.1003 | 176.703 | 192.0029 | 222.1986 | 195.1013 | 327.4985 | 373.703 | 206.7975 | 264.2951 | 256.9952 | 270.0019 | 338.8949 | 132.1967 | 140.8004 | 145.4995 | 175.1022 | 190.8978 | 170.4995 | 173.3992 | 110.7994 | 264.1977 | 284.0993 | 310.8023 | 298.4966 | 171.3992 | 170.6008 | 156.8997 | 138.5998 | 109.4021 | 96.1003 | 100.999 | 154.4995 | 251.5979 | 227.6982 | 266.9007 | 178.2999 | 181.4031 | 168.0993 | 146.298 | 137.4012 | 225.2989 | 211.3038 | 243.8958 | 196.3992 | 170.703 | 153.0974 | 184.0974 | 166.6008 | 187.4012 | 223.4985 | 246.2035 | 257.7994 | 187 | 202.4012 | 288.0029 | 238.7011 | 220.7975 | 270.406 | 236.6027 | 339.8014 | 200.9007 | 189.0019 | 195.1996 | 272.1022 | 150.6991 | 154.0984 | 157.2015 | 204.5044 | 204.6991 | 200.3992 | 204.6991 | 161.7994 | 125.6991 | 127.3009 | 124.8004 | 152 | 275.9046 | 289.0993 | 294.7994 | 384.3973 | 272.3983 | 297.1013 | 329 | 402.2025 | 135.4012 | 91.6008 | 95.2006 | 135.4012 | 129.7994 | 116.297 | 120.5988 | 157.6046 | 129.8997 | 147.6008 | 136.4021 | 129.9007 | 94.5024 | 93.8004 | 91.1003 | 104.5969 | 113.3973 | 120.4976 | 119.0993 | 97.7001 | 288.1051 | 250.298 | 241.4012 | 307.9971 | 206.9036 | 227.8004 | 234.8987 | 267.4995 | 328.5998 | 239.9952 | 262.9981 | 349.9952
840 | 11.2344 | 173.4976 | 241.2975 | 227.8 | 209.0536 | 177.5024 | 185.2586 | 197.9464 | 183.7365 | 188.1951 | 156.0585 | 128.6244 | 185.1171 | 269.9415 | 187.322 | 184 | 269 | 205.6244 | 165.9268 | 174.7463 | 231.0634 | 195.0536 | 334.8781 | 333.8927 | 219.0683 | 294.6927 | 274.5171 | 259.4341 | 358.5171 | 151.0049 | 135.8146 | 147.1268 | 157.7512 | 200.1317 | 180.2439 | 183.6829 | 106.8829 | 286.3073 | 284.3171 | 291.1903 | 315.5073 | 177.6244 | 158.7561 | 156.8146 | 144.9951 | 93.6878 | 93.3756 | 99.2586 | 152.0683 | 254.2 | 238.5024 | 264.8097 | 178.6878 | 165.2439 | 170.8146 | 164.7463 | 125.5659 | 222.0732 | 191.8049 | 262.761 | 206.9951 | 150.2488 | 174.1854 | 208.3073 | 160.6878 | 186.4927 | 228.6927 | 222.561 | 260.1268 | 185.1268 | 181.8244 | 275.7414 | 240.0536 | 236.5024 | 232.9317 | 222.3073 | 343.1122 | 194.439 | 189.6731 | 194.2537 | 249.4439 | 165.4292 | 167.8732 | 150.0585 | 169.5659 | 206.3171 | 202.2537 | 206.9415 | 158.8195 | 134.8097 | 120.3756 | 124.8097 | 151.6878 | 261.4195 | 292.1268 | 296.8146 | 392.4488 | 281.2537 | 285.5024 | 326.5024 | 378.6927 | 131.0585 | 81.6293 | 91.3122 | 131.0585 | 128.3805 | 131.4439 | 124.0049 | 128.9951 | 128.8781 | 137.3171 | 122.561 | 126.2488 | 80.1219 | 88.8146 | 90.8732 | 122.9415 | 122.6976 | 128.322 | 116.1951 | 97.6244 | 265.6731 | 264.6878 | 243.9268 | 328.6878 | 195.9853 | 234.9902 | 242.1268 | 277.5561 | 316.2635 | 265.6342 | 264.5122 | 365.3318
841 | 11.2454 | 177.0531 | 260.0737 | 259.2449 | 225.7551 | 171.6755 | 171.1918 | 191.3039 | 213.0265 | 173.5959 | 147.1533 | 132.2714 | 200.4041 | 249.9675 | 164.0531 | 177.8348 | 278.0856 | 202.4572 | 186.7551 | 182.6755 | 226.0796 | 207.8613 | 332.5428 | 300.8142 | 196.5899 | 268.3982 | 237.9469 | 269.9204 | 320.3245 | 146.1652 | 138.4632 | 149.8142 | 162.7876 | 176.2389 | 187.8348 | 189.1327 | 96.1387 | 298.3245 | 279.8407 | 272.6696 | 306.7876 | 173.4837 | 153.7611 | 152 | 144.9675 | 94.9469 | 92.3245 | 81.2979 | 143.8673 | 227.6224 | 242.7346 | 270.351 | 185.1387 | 177.0531 | 169.2449 | 167.8082 | 126.4366 | 202.8673 | 204.4572 | 238.3245 | 208.9144 | 156.8938 | 176.7286 | 204.4247 | 158.0531 | 200.4041 | 218.9469 | 229.5693 | 260.5428 | 181.649 | 167.006 | 291.1062 | 246.0471 | 230.6755 | 250.8673 | 230.7551 | 355.4512 | 185.1593 | 212.5899 | 189.5693 | 247.9794 | 185.9469 | 172.649 | 162.5693 | 169.1387 | 200.8673 | 194 | 203.5163 | 137.8082 | 141.3245 | 119.649 | 128.0796 | 151.3245 | 293.8879 | 284.4307 | 286.1593 | 372.1918 | 290.1977 | 277.0796 | 316.4572 | 359.5368 | 145.8407 | 72.9735 | 92.9735 | 145.8407 | 116.649 | 116.351 | 113.649 | 134.4837 | 124.5959 | 131.5428 | 127.6224 | 138.7346 | 88.6224 | 85.9469 | 91.4307 | 121.7876 | 100.649 | 101.1593 | 105.4897 | 96.7286 | 298 | 263.6755 | 270.5959 | 329.6224 | 208.5574 | 244.6696 | 240.5959 | 287.7817 | 286.649 | 240.4572 | 240.5693 | 327.1918
842 | 11.2564 | 170.9793 | 232.7045 | 267.3471 | 228.7665 | 176.3884 | 194.1239 | 214.9587 | 214.7045 | 178.3264 | 121.1136 | 134.6426 | 189.2748 | 238.4504 | 180.2066 | 171.7355 | 293.6322 | 200.0723 | 176.969 | 184.0207 | 221.6839 | 194.0103 | 342.3368 | 328.1343 | 188.1136 | 255.5124 | 250.7562 | 275.6632 | 322.6839 | 154.9587 | 158.7149 | 155.2438 | 173.2748 | 171.4711 | 177.3574 | 177.9897 | 104.938 | 294.9587 | 292 | 256.7665 | 317.6116 | 164.0207 | 172.3368 | 158.062 | 129.0723 | 101.0207 | 91.9897 | 81.6426 | 144.5703 | 240.093 | 260.9793 | 271.0207 | 193.938 | 171.6529 | 169.2645 | 161.7149 | 139.9587 | 211.6529 | 193.3161 | 231.9277 | 202.7975 | 173.0826 | 166.9484 | 177.3884 | 155.3471 | 195.3368 | 235.4607 | 239.3574 | 255.1818 | 185.0207 | 191.6632 | 281.3161 | 220.7045 | 231.6839 | 252.9174 | 237.1343 | 336.8285 | 180.4091 | 200.0723 | 198.6839 | 263.5496 | 182.5909 | 164.5703 | 165.2851 | 178.9484 | 209.6529 | 206.1239 | 212.3058 | 132.3884 | 137.6219 | 122.3471 | 129.4091 | 153.0103 | 270.938 | 275.9897 | 277.3678 | 378.9587 | 310.1715 | 281.4401 | 310.3678 | 355.1136 | 156.3161 | 82.0723 | 93.9897 | 156.3161 | 120.6942 | 115.6736 | 123.4194 | 124.6839 | 122.2541 | 137.2955 | 134.7045 | 145.5289 | 93.6839 | 92.3574 | 81.6426 | 130.9277 | 105.3678 | 99.7768 | 115.2955 | 93.3471 | 286.2128 | 255.2541 | 272.6322 | 329.2955 | 175.093 | 223.0413 | 239.938 | 284.0516 | 282.9484 | 241.7768 | 248.6736 | 339.0103
843 | 11.2674 | 180.2646 | 237.5565 | 262.3205 | 243.8112 | 183.1603 | 198.8013 | 202.1131 | 218.8584 | 175.2547 | 145.4435 | 123.7168 | 185.3304 | 264.8397 | 207.0845 | 181.8584 | 270.2448 | 215.0472 | 173.7925 | 187.3019 | 220.8112 | 199.8397 | 335.3205 | 335.4994 | 205.1131 | 275.3392 | 266.6037 | 273.9528 | 328.0944 | 147 | 159.9714 | 148.472 | 160.6037 | 202.1888 | 175.3677 | 176.3491 | 97.1888 | 293.9813 | 297.236 | 282.9813 | 306.6509 | 168.6981 | 170.9056 | 176.6323 | 139.1603 | 104.6509 | 98.0286 | 79.0944 | 133.6795 | 247.528 | 252.4621 | 271.8584 | 184.7925 | 167.2547 | 154.9528 | 173.4435 | 133.7453 | 210.7453 | 205.359 | 217.6509 | 220.3205 | 179.5752 | 162.5851 | 186.2547 | 156.9528 | 202.9813 | 249.877 | 238.0659 | 238.5664 | 193.1888 | 185.4149 | 260.1976 | 238.123 | 233.2547 | 239.7925 | 251.4808 | 357.528 | 197.7925 | 202.8298 | 204.4435 | 252.8683 | 162.764 | 157.5192 | 157.0472 | 166.9056 | 205.6037 | 221.9714 | 212.236 | 147.1888 | 125.8584 | 131.9813 | 137.0187 | 154.6509 | 278.5478 | 275.7453 | 296.5664 | 370.6509 | 289.123 | 287.3392 | 305.8298 | 378.3963 | 162.0758 | 90.0659 | 92.6981 | 162.0758 | 130.8869 | 128.1702 | 132.6696 | 129.0472 | 106.3019 | 134.8869 | 135.0187 | 114.8112 | 88.2733 | 97.6981 | 78.3963 | 114.5565 | 112 | 117.5093 | 111.8397 | 93.2075 | 269.9813 | 248.0286 | 271.9341 | 309.7826 | 195.7925 | 227.1603 | 218.5752 | 296.0944 | 285.9155 | 251.1317 | 247.2075 | 343.0944
844 | 11.2784 | 188.3775 | 226.4955 | 254.5045 | 227.3274 | 181.059 | 181.4293 | 198.4204 | 205.7728 | 172.7227 | 144.2183 | 114.8909 | 194.3023 | 260.7979 | 188.4204 | 187.2021 | 252.3703 | 219.1682 | 181.5295 | 174.0752 | 221.3112 | 201.5796 | 333.6475 | 304.2522 | 201.059 | 246.5546 | 263.6386 | 256.8229 | 323.3274 | 145.9159 | 150.059 | 154.6297 | 162.6977 | 206.7478 | 169.8658 | 171.941 | 107.5295 | 309.5295 | 290.0162 | 287.6888 | 308.1682 | 166.0251 | 172.059 | 182.1342 | 138.5045 | 107.2522 | 100.3525 | 81.5546 | 138.6047 | 226.5546 | 238.5796 | 258.0501 | 194.3363 | 175.9248 | 164.2021 | 180.5295 | 135.5546 | 203.5206 | 211.9571 | 216.2773 | 204.1932 | 171.3953 | 181.3112 | 180.1091 | 153.5546 | 208.7728 | 223.7728 | 239.6475 | 256.941 | 190.1592 | 163.7979 | 238.9248 | 251.5796 | 226.7478 | 257.6475 | 226.7478 | 337.6386 | 205.5295 | 191.3953 | 210.0841 | 269.2183 | 154.4454 | 173.2183 | 164.059 | 169.5045 | 208.059 | 210.6136 | 218.3864 | 157.8909 | 121.0841 | 142.1091 | 132.6726 | 153.2773 | 309.941 | 281.5295 | 299.0412 | 375.7817 | 293.1843 | 258.9159 | 291.8658 | 374.1342 | 172.4955 | 79.3614 | 95.8068 | 172.4955 | 132.7728 | 145.0251 | 120.0841 | 136.4204 | 124.8747 | 140.3864 | 123.4454 | 117.8408 | 79.9499 | 99 | 81.7227 | 107.1091 | 113.4454 | 106.3023 | 120.4454 | 83.6637 | 283 | 270.5885 | 266.3775 | 288.8068 | 201.7227 | 220.7227 | 220.8747 | 294.941 | 310 | 235.5045 | 244.8909 | 339.7728
845 | 11.2894 | 165.7257 | 235.1771 | 255.1642 | 218.352 | 183.2743 | 213.6264 | 202.7689 | 195.2419 | 177.7365 | 140.7257 | 129.7149 | 185.2527 | 248.473 | 191.648 | 193.9223 | 261.8013 | 223 | 187.1209 | 158.2311 | 226.6588 | 192.7473 | 348.9007 | 310.7473 | 214.1101 | 241.7581 | 251.4169 | 247.9352 | 325.1879 | 142.5054 | 147.0432 | 131 | 168.0108 | 200.2527 | 167.8013 | 155.9007 | 113.1209 | 317.3628 | 265.2311 | 278.9784 | 313.1209 | 152.2743 | 173.527 | 171.1209 | 154.1101 | 116.7365 | 78.7473 | 84.9784 | 134.7041 | 228.1101 | 227.8791 | 251.473 | 197.8899 | 182.4277 | 177.648 | 179.3953 | 130.0108 | 194.7149 | 189.4406 | 221.7257 | 206.9784 | 176.4169 | 190.3953 | 181.352 | 152.4946 | 205.3412 | 201.2851 | 260.1318 | 255.473 | 186.6933 | 156.7041 | 252.527 | 243.4946 | 235.1987 | 257.2074 | 215.3953 | 334.3845 | 207.3845 | 207.2527 | 210.1318 | 271.7041 | 158.4946 | 165.9892 | 176.7365 | 179.8791 | 213.2635 | 212.1426 | 201.5378 | 155.1534 | 121.5054 | 132.8899 | 146.5054 | 148.2635 | 291.6588 | 263.581 | 278.0669 | 372.175 | 278.5831 | 271.1987 | 302.8791 | 368.7257 | 159.5054 | 80 | 95.2635 | 159.5054 | 115.1426 | 143.2311 | 121.6264 | 142.2635 | 149.5378 | 126.1534 | 130.8574 | 133.2635 | 94 | 96.3845 | 85.9892 | 111.3412 | 111.8899 | 103.2311 | 117.3953 | 81.6047 | 290.0993 | 276.9007 | 249.7041 | 290.1318 | 195.1534 | 225.3628 | 244.7905 | 286 | 313.7365 | 233.5486 | 258.5939 | 327
846 | 11.3004 | 176.2281 | 256.264 | 238.5438 | 229.7719 | 171.6141 | 197.5624 | 192.0703 | 202.8594 | 190.9469 | 144.6671 | 145.6141 | 186.7016 | 254.439 | 174.825 | 175.4735 | 257.683 | 216.825 | 184.3687 | 167.7719 | 215.6313 | 201.3329 | 341.2453 | 302.3515 | 218.2984 | 249.5782 | 231.6141 | 259.4749 | 341.6499 | 146.9469 | 135.4907 | 135.6313 | 162.2281 | 200.5438 | 165.2267 | 161.7533 | 117.7016 | 313.7374 | 264.7374 | 286.6844 | 311.9125 | 149.4907 | 164.8594 | 161.8077 | 163.7016 | 116.439 | 91.9642 | 93.4562 | 152.1406 | 238.0875 | 227.5438 | 255.8952 | 188.6844 | 164.9469 | 169.3156 | 154.1764 | 131.9469 | 203.2812 | 220.4735 | 234.1578 | 202.7202 | 156.6141 | 176.3687 | 181.5796 | 149.2109 | 220.2281 | 193.1048 | 242.3343 | 270.3156 | 214.5438 | 187.6485 | 249.6485 | 241.3687 | 261.3156 | 215.9125 | 227.6127 | 337.7188 | 204.5438 | 203.6844 | 207.7719 | 285.6671 | 161.7719 | 176.7891 | 172.9655 | 175.2984 | 204.2984 | 207.8594 | 186.3156 | 145.0875 | 123.2453 | 118.6671 | 139.3687 | 149.3329 | 282.561 | 242.5954 | 294.5796 | 360.9642 | 302.2453 | 289.2109 | 302.9297 | 374.9828 | 157.3859 | 88.4907 | 83.9828 | 157.3859 | 111.6313 | 143.8952 | 134.122 | 142.561 | 123.8952 | 116.0875 | 129.3515 | 130.4735 | 94.7719 | 96.2453 | 90.2281 | 115.8077 | 109.2453 | 112.3859 | 109.5438 | 90.6141 | 289.6499 | 253.0358 | 274.8594 | 303.5954 | 186.6313 | 238.3329 | 226.3859 | 275.9655 | 314.9828 | 216.0172 | 282.6485 | 345.1392
847 | 11.3114 | 182.1682 | 240.3805 | 233.4336 | 224.6283 | 175.3805 | 184.1505 | 183.823 | 199.4424 | 190.0531 | 138.177 | 136.6371 | 185.823 | 252.7611 | 168.9823 | 179.8053 | 268.6902 | 219.3452 | 172.6195 | 176.1682 | 232.9558 | 204.6371 | 340.8318 | 298.9469 | 220.7699 | 277.5576 | 233.3894 | 234.3805 | 314.8318 | 155.6106 | 153.7788 | 147.3805 | 153.4336 | 199.4159 | 184.6637 | 175.6637 | 122 | 296.6106 | 272.6992 | 274.6371 | 312.9735 | 161.0088 | 166.2212 | 145.1947 | 170.7876 | 104.8053 | 111.7965 | 85.823 | 144.0088 | 244.1947 | 228.4071 | 256.5399 | 190.5752 | 170.0265 | 165.8318 | 148.9116 | 136.2301 | 196.9735 | 212.4601 | 234.0088 | 195.9293 | 157.5929 | 167.4071 | 164.6018 | 139.6018 | 214.6195 | 207.9912 | 235.3275 | 266.8318 | 198.6813 | 202.2035 | 250.6637 | 229.2212 | 245.4867 | 221.7523 | 246.469 | 340.0442 | 212.9735 | 199.6018 | 193.469 | 267.6283 | 166.9823 | 188.3894 | 150.2301 | 174.1859 | 201.9912 | 208.0265 | 202.7434 | 148.0088 | 133.1859 | 122.531 | 141.1593 | 149.8495 | 294.5929 | 272.1859 | 282.7788 | 372.4336 | 299.4424 | 274.823 | 294.4513 | 352.6637 | 142.469 | 100.4071 | 80.7788 | 142.469 | 122.1859 | 125.823 | 144.2477 | 145.4336 | 116.5752 | 125.3805 | 121.9646 | 150.3363 | 88.8318 | 91.8495 | 94.9735 | 99.1947 | 110.823 | 113 | 113.5929 | 92.3894 | 278.3629 | 262.2744 | 288.5664 | 324.0265 | 200.3717 | 253.9823 | 219.8318 | 266.7699 | 295.8495 | 228.7788 | 298.7965 | 347.7167
848 | 11.3224 | 179.8633 | 245.7684 | 222.821 | 214.5895 | 181.179 | 206.7473 | 189.1367 | 197.7262 | 176.1052 | 141.4843 | 130.9789 | 182.9263 | 253.2738 | 175.821 | 182.9263 | 249.9052 | 221.9896 | 170.5157 | 175.0948 | 239.7895 | 180.9159 | 318.0948 | 318.2316 | 227.3053 | 285.0211 | 254.6525 | 244.2842 | 310.1578 | 159.3894 | 163.8737 | 154.821 | 151.0315 | 189.0737 | 173.6736 | 160.9789 | 124.4631 | 293 | 296.7369 | 277.5999 | 309.2738 | 161.6947 | 173.3683 | 161.3683 | 171.3053 | 105.0526 | 109.7158 | 82.1052 | 138.5369 | 246.821 | 226.4105 | 281.5157 | 202.6947 | 167.1578 | 163.2105 | 175.3053 | 128.4631 | 203.9474 | 191.5895 | 231.821 | 209.7369 | 156.3053 | 160.8948 | 173.0315 | 148.442 | 205.5369 | 213.4631 | 245.2212 | 262.1578 | 185.7788 | 197.3053 | 257.7366 | 219.1263 | 227.9052 | 235.1263 | 217.6632 | 330.7473 | 226.1052 | 198.5895 | 176.5684 | 262.5157 | 159.8633 | 185.4316 | 140.821 | 182.2842 | 206.2316 | 208.442 | 216.2527 | 146.6421 | 143.3368 | 148.4631 | 143.9159 | 145.7262 | 299.4631 | 290.9578 | 279.7369 | 364.6947 | 279.6632 | 262.8948 | 282.7473 | 360.9682 | 139.1156 | 94.7158 | 84.7158 | 139.1156 | 132.7473 | 126.2105 | 124.3053 | 136.8737 | 127.8737 | 124.6106 | 124.558 | 165.8422 | 82.1052 | 75 | 98.6632 | 106.3368 | 108.7473 | 106.4316 | 105.7369 | 96.4105 | 279.9159 | 289.8107 | 278.6528 | 333.063 | 213.2842 | 250.1475 | 219.6736 | 282.7473 | 308.5576 | 239.2842 | 277.0107 | 336.7366
849 | 11.3334 | 169.4563 | 227.9331 | 240.4897 | 208.9262 | 171.9665 | 196.3962 | 201.235 | 212.1947 | 176.0806 | 133.7316 | 133.8122 | 188.3087 | 237.9597 | 181.651 | 176.8928 | 268.6913 | 199 | 181.2281 | 162.2281 | 222.0806 | 175.5703 | 325.5232 | 314.0806 | 215.235 | 272.0403 | 249.5574 | 253.651 | 319.4228 | 155.6243 | 160.3894 | 160.651 | 166.9597 | 180.7316 | 184.7719 | 172.6913 | 119.5437 | 285.8122 | 289.1878 | 303.8053 | 302.416 | 166.5772 | 171.5909 | 179.0403 | 162.6175 | 102.5034 | 99.9262 | 86.7316 | 137.9597 | 256.0335 | 224.8859 | 285.4631 | 197.8525 | 176.8456 | 163.1612 | 157.3156 | 140.4563 | 201 | 183.3894 | 244.8388 | 194.1544 | 166.6441 | 153.7316 | 198.2616 | 160.7316 | 197.7719 | 222.4966 | 242.181 | 284.9529 | 185.4365 | 180.1612 | 286.3156 | 228.1072 | 245 | 233.1141 | 212.9597 | 330.5437 | 210.8593 | 196.3087 | 188.5369 | 264.349 | 143.5369 | 164.584 | 137.349 | 194.6106 | 205.4631 | 219.0403 | 199.5437 | 162.9529 | 137.47 | 137.2015 | 135.1878 | 161.8859 | 296.235 | 289.8593 | 281.9125 | 374.2281 | 261.8525 | 274.3354 | 287.1947 | 390.47 | 154.47 | 91.2684 | 78.3087 | 154.47 | 139.3087 | 142.2281 | 121.5369 | 151.5703 | 132.8456 | 113.5772 | 116.4966 | 146.8525 | 85.8859 | 79.2281 | 91 | 116.5369 | 114.8859 | 104.6106 | 99.4563 | 90.6578 | 273.7247 | 262.9262 | 257.7247 | 334.054 | 212.0806 | 227.5437 | 232.4631 | 286.349 | 340.0068 | 254.5703 | 242.7316 | 361.9331
850 | 11.3444 | 183.6106 | 216.8776 | 256.9085 | 209.8319 | 178.4042 | 167.9543 | 192.4351 | 200.4734 | 160.6032 | 140.267 | 135.7478 | 178.1681 | 258.0989 | 199.7478 | 162.0457 | 266.2138 | 191.1681 | 185.2596 | 181.4882 | 217.9159 | 197.3968 | 368.5723 | 314.267 | 209.0457 | 259.2979 | 243.6786 | 259.5649 | 309.5575 | 139.9159 | 146.9543 | 158.2979 | 180.5723 | 192.9234 | 175.9543 | 173.6947 | 107.5649 | 289.4882 | 282.0309 | 299.4277 | 319.7404 | 173.3968 | 148.3968 | 166.733 | 153.2138 | 101.9617 | 89.5192 | 84.5649 | 146.7862 | 264.8245 | 223.3053 | 264.1607 | 178.3436 | 174.0841 | 159.1829 | 157.4042 | 150.2596 | 202.7404 | 181.267 | 250.9926 | 193.8702 | 176.6873 | 150.2596 | 196.2905 | 156.3894 | 188.9543 | 228.2596 | 256.5575 | 280.0692 | 163.8319 | 166.6106 | 268.5649 | 237.2979 | 245 | 230.3436 | 226.1372 | 309.4277 | 189.8319 | 198.351 | 190.2596 | 259.7404 | 154.3968 | 159.8776 | 148.4042 | 191.5118 | 200.2596 | 203.2522 | 188 | 166.3362 | 125.9617 | 137.4499 | 137.1681 | 160.3053 | 288.3053 | 268.8319 | 290.6415 | 392.618 | 275.8467 | 267.9778 | 285.4808 | 376.7862 | 137.3053 | 90.3894 | 68.1681 | 137.3053 | 131.7787 | 134.0766 | 125.8702 | 151.6415 | 128.7787 | 110.8245 | 117.4734 | 138.2213 | 79.0841 | 88.4808 | 92.3053 | 112.6032 | 111.5649 | 108.9085 | 113.1755 | 74.6032 | 293.5266 | 257.7404 | 260.5575 | 311.7478 | 187.466 | 237.3201 | 238.1372 | 296.0989 | 319.7787 | 252.0309 | 254.9234 | 333.9085
851 | 11.3554 | 187.8948 | 223.8417 | 246.3156 | 231.6323 | 186.6833 | 176.9745 | 196.1318 | 184.3943 | 165.3167 | 154.5792 | 142.4995 | 180.0797 | 260.2094 | 205.6047 | 174.0531 | 267.6057 | 198.8958 | 179.8682 | 194.9734 | 226.5792 | 205.2104 | 367.5781 | 324.5526 | 207.6312 | 260.0797 | 280.6047 | 265.7109 | 317.0276 | 143.6578 | 135.8417 | 152.8161 | 180.0255 | 197.5781 | 178.7109 | 177.3688 | 112.8161 | 297.1573 | 268.5005 | 279.8417 | 322 | 187.0266 | 157.1052 | 163.1583 | 161.7641 | 111.8948 | 91.7109 | 84 | 153.8948 | 266.9214 | 227.6844 | 233.3943 | 179.5005 | 176.1583 | 171.8682 | 183.5792 | 146.6578 | 212.1583 | 195.1318 | 233.0786 | 197.6844 | 162 | 143.9734 | 173.7896 | 162.8427 | 188.1318 | 222.8682 | 233.1573 | 239.6047 | 188.3167 | 187.8958 | 269.3422 | 243.4484 | 235.6047 | 244.9224 | 228.1573 | 321.6068 | 195.526 | 204.4474 | 187.1052 | 261.5526 | 171.6057 | 170.8682 | 172.7896 | 195.9224 | 205.1583 | 191.5792 | 206.3432 | 147.7896 | 133.2104 | 151.9203 | 123.4208 | 162 | 292.6844 | 279.8948 | 275.1318 | 404.526 | 283.8141 | 252.2646 | 293.1318 | 372.2625 | 140.3422 | 88.7896 | 65.6057 | 140.3422 | 127.237 | 130.3167 | 127.4474 | 144.6323 | 124.6844 | 112.9214 | 111.6844 | 140.9734 | 73.1052 | 86.2891 | 89.9734 | 112.3953 | 113.6844 | 105.0266 | 104.3677 | 79.3167 | 294.3943 | 262.237 | 246.5526 | 317.6047 | 184.3432 | 247.1042 | 240.1573 | 290.1562 | 304.9469 | 241.6323 | 245.7089 | 337.1849
852 | 11.3664 | 186.7016 | 217.2984 | 246.2178 | 245.5403 | 178.1937 | 189.7822 | 193.5644 | 177.2984 | 179.3225 | 154.8628 | 135.6775 | 186.2659 | 238.8387 | 185.2659 | 182.1613 | 274.7016 | 209.5078 | 166.3466 | 184.4838 | 230.5403 | 196.5644 | 345.0241 | 320.7822 | 201.9194 | 270.4031 | 273.1372 | 262.0481 | 323.2335 | 154.4115 | 132.9759 | 151.2659 | 161.2419 | 184.2178 | 178.2659 | 164.2335 | 109.4272 | 301.5487 | 278.0806 | 277.4356 | 323.8387 | 194.2984 | 150.0241 | 163.7256 | 162.9916 | 112.5403 | 88.5078 | 86.7581 | 149.4838 | 267.4838 | 234.6775 | 244.2262 | 185.4031 | 177.1853 | 172.5969 | 166.3947 | 154.6534 | 233.4115 | 201.7581 | 221.0241 | 196.8628 | 162.9194 | 155.089 | 170.0241 | 158.911 | 192 | 217.6209 | 225.1372 | 249.1456 | 211.5162 | 199.8869 | 271 | 248.4513 | 228.1372 | 251.3141 | 217.3791 | 332.6932 | 190.4597 | 203.1613 | 185.0806 | 251.8063 | 173.645 | 169.2984 | 163.0481 | 209.2094 | 224.1131 | 194.1613 | 212.4513 | 148.6209 | 130.0806 | 128.7256 | 129.5728 | 162.4597 | 294.1613 | 287.4356 | 299.6859 | 388.8869 | 280.7184 | 286.5162 | 294.7016 | 371.4115 | 144.7581 | 88.2419 | 77.7581 | 144.7581 | 134.5969 | 140.1853 | 130.2984 | 150.2659 | 129.8387 | 123.5969 | 115.4597 | 138.2984 | 71.5403 | 90.8712 | 79.4838 | 117.5644 | 121.1372 | 116.8953 | 100.1131 | 97 | 304.3068 | 252.5885 | 238.645 | 315.6534 | 188.6125 | 233.2744 | 233.0566 | 253.3225 | 302.8147 | 256 | 228.089 | 330.3141
853 | 11.3774 | 166.0648 | 224.7198 | 236.3126 | 217.6253 | 194.6637 | 176.5604 | 187.8879 | 188.0236 | 162.7286 | 151.4159 | 140.472 | 174.5841 | 234.3923 | 191.0473 | 169.6165 | 281.9115 | 203.0797 | 167.7434 | 184.6077 | 230.944 | 197.0236 | 346.0236 | 315.1121 | 195.4484 | 254.1445 | 260.5368 | 262.1593 | 295.2802 | 153.4247 | 134.8082 | 139.5841 | 157.056 | 188 | 164.2242 | 159.4071 | 86.6165 | 303.5368 | 300.5989 | 286.4159 | 307.5929 | 178.5929 | 151.0236 | 166.2714 | 157.1593 | 100.6725 | 84.4396 | 92.3599 | 150.5516 | 261.944 | 226.7286 | 256.2006 | 181.4159 | 175.0797 | 167.6165 | 162.239 | 156.0885 | 227.7049 | 194.1445 | 206.9203 | 187.7522 | 156.4484 | 174.472 | 173.8555 | 148.9676 | 199.5516 | 197.1209 | 233.7758 | 274.944 | 206.2006 | 189.8879 | 254.9527 | 253.0709 | 232.528 | 235.0236 | 219.472 | 325.1033 | 191.944 | 195.3363 | 196.2714 | 255.1593 | 173.0236 | 162.5604 | 158.9115 | 180.7286 | 213.3451 | 198.6077 | 204.7994 | 150.3038 | 142.2154 | 124.6637 | 130.2006 | 145.5368 | 292.472 | 290.7522 | 305.1209 | 358.1209 | 293.9291 | 288.7522 | 311.8231 | 369.0088 | 148.472 | 84.528 | 81 | 148.472 | 144.2478 | 131 | 127.3363 | 140.472 | 133.4159 | 124.7522 | 130.6313 | 140.528 | 81.3835 | 100.3363 | 89.0473 | 111.4159 | 126.472 | 124.528 | 109.6401 | 94.1681 | 299.6253 | 242.7198 | 246.0473 | 320.8643 | 176.9115 | 224.5929 | 229.0885 | 266.9352 | 342.6786 | 247.9764 | 240.8643 | 322.9912
854 | 11.3884 | 157.6225 | 232.4213 | 222.4528 | 197.6539 | 197.1258 | 184.4022 | 193.8742 | 197.4843 | 167.4022 | 136.0506 | 139.5472 | 184.1382 | 245.8865 | 176.554 | 172.0438 | 282.3461 | 208.5787 | 176.673 | 181.3461 | 228.1258 | 198.2517 | 342.409 | 313 | 203.9494 | 259.918 | 258.4337 | 256.6922 | 304.5281 | 141.327 | 142.5595 | 140.4213 | 165.6854 | 187.0315 | 160.9685 | 174.1888 | 82.7483 | 292.7168 | 293.1012 | 287.0315 | 296.6854 | 163.8112 | 157.0944 | 171.2832 | 171.5472 | 92.3584 | 88.7045 | 89.1888 | 159.4843 | 248.4719 | 214.9371 | 252.6854 | 166.0506 | 184.4528 | 169.5595 | 179.2202 | 157.591 | 201.1258 | 189.7483 | 214.9809 | 196.5595 | 145.0944 | 165.3146 | 190.8427 | 159 | 207.0315 | 191.0123 | 236.5472 | 270.1888 | 196.3899 | 179.8933 | 231.1888 | 259.6607 | 235.3899 | 238.673 | 218.0629 | 340.5472 | 188.1573 | 199.1696 | 202.2517 | 263.736 | 170.3775 | 154.9056 | 161.2832 | 165.5472 | 187.0315 | 204.5472 | 208.3146 | 147.2202 | 151.673 | 126.1573 | 138.7921 | 138.591 | 293 | 305.855 | 285.9371 | 361.6977 | 280.0753 | 283.5157 | 318.9877 | 346.1258 | 149.4843 | 77.2202 | 74.7045 | 149.4843 | 137.8618 | 126.1573 | 121.9685 | 132.7674 | 133.5472 | 123.8742 | 136.3461 | 141.4528 | 102.6854 | 91.5787 | 90.5349 | 106.7045 | 114.409 | 111.409 | 109.4213 | 95.8427 | 296.1191 | 256.2326 | 240.2708 | 314.9685 | 196.7168 | 225.7921 | 224.7798 | 308.7921 | 337.7742 | 223.0191 | 248.0438 | 334.736
855 | 11.3994 | 177.945 | 235.4966 | 232.9381 | 214.9587 | 190.5172 | 220.883 | 181.6135 | 200.9794 | 182.531 | 144.8142 | 137.0069 | 191.0344 | 258.055 | 154.9106 | 184.5172 | 280.9312 | 213.945 | 169.5103 | 155.5998 | 243.3658 | 192.9794 | 331.4828 | 325.9106 | 204.6204 | 278.5034 | 263.5929 | 234 | 326.9518 | 143.0275 | 164.945 | 136.5447 | 168.5516 | 188.9794 | 162 | 166.0138 | 107.883 | 296.5378 | 262.9587 | 273.0894 | 308.9862 | 170.4966 | 147.5447 | 156.0413 | 164.0413 | 104.4484 | 87.9587 | 89.4553 | 153.5447 | 240.945 | 239.8418 | 274.9174 | 174.3177 | 184.5103 | 178.0413 | 172.9931 | 152.6273 | 211.4002 | 198.9931 | 232 | 201.0688 | 162.8555 | 164.9312 | 183.586 | 153.0413 | 201.0344 | 212.9587 | 239.9656 | 253.0757 | 194.5378 | 165.0138 | 247.3452 | 231.5447 | 230.0619 | 234.4897 | 229.4071 | 344.9587 | 189.9518 | 206.0275 | 191.5172 | 267.9174 | 165.4484 | 158 | 157.9587 | 179.3934 | 200.4002 | 197.0413 | 191.0482 | 156.3865 | 140.5378 | 121.9931 | 138.6617 | 150.5103 | 295.9794 | 328.4897 | 295.4484 | 361.6823 | 291.5172 | 293.4278 | 301.0138 | 367.3246 | 158.9381 | 77.945 | 75.4484 | 158.9381 | 126 | 131.4278 | 123.4966 | 134.9243 | 120.5791 | 123.531 | 118.0482 | 133.5654 | 101.5791 | 90.4897 | 75.4828 | 108.4415 | 114.9037 | 100.4828 | 109.0206 | 101.9931 | 310.1032 | 262.0206 | 240.8211 | 316 | 201.0344 | 242.0482 | 206.6754 | 324.055 | 301.9587 | 233.8074 | 277.4002 | 333.4553
856 | 11.4104 | 187.5266 | 228.3672 | 243.0177 | 218.9646 | 186.9823 | 217.1371 | 173.1416 | 199.4203 | 176.4734 | 155.2079 | 128.3672 | 190.0708 | 254.5797 | 164.9469 | 168.7699 | 277.7699 | 223.0177 | 153.2433 | 159.8275 | 243.6637 | 187.3495 | 345.7036 | 334.9292 | 194.6328 | 275.4557 | 255.0885 | 239.0885 | 330.438 | 141.0354 | 155.6991 | 137.6328 | 168.6328 | 192 | 179.3009 | 160.438 | 124.4911 | 299.6505 | 284.2655 | 271.7036 | 299.2964 | 176.0885 | 142.0177 | 161.7036 | 166.6505 | 106.4026 | 96.0354 | 92.438 | 147 | 249.5089 | 246.2079 | 257.9955 | 186.2433 | 174.3495 | 172 | 164.8407 | 136.5443 | 225.4911 | 197.9646 | 227.9292 | 181.3318 | 192.6505 | 177.5443 | 194.916 | 155.1416 | 194.9823 | 223.5797 | 238.8938 | 270.4957 | 192.562 | 185.3894 | 280.6859 | 229.0708 | 239.3186 | 224.2964 | 226.2079 | 340.3141 | 205.6505 | 188.7699 | 186.4557 | 279.4911 | 163.3318 | 162.0708 | 162.4734 | 184.3141 | 188.0309 | 202.1947 | 199.7744 | 166.3849 | 137.0354 | 116.8938 | 132.8098 | 146.4557 | 288.823 | 312.6991 | 298.9292 | 344.1593 | 282.3849 | 287.2079 | 278.646 | 365.522 | 175.1239 | 84.9823 | 90.1239 | 175.1239 | 138.2124 | 149.1239 | 128.0708 | 138.3672 | 116.6328 | 114.9292 | 131.8629 | 129.5974 | 90 | 92.5089 | 83.0885 | 120.562 | 121.8761 | 108.0885 | 104.4734 | 98.9292 | 285.8407 | 248.823 | 244.1017 | 298.6991 | 197.0177 | 223.8053 | 205.4249 | 301.7522 | 312.5797 | 254.3672 | 271.1371 | 324.7345
857 | 11.4214 | 188.4789 | 227.2537 | 240.352 | 218.0423 | 191.7326 | 201.1691 | 189.8594 | 208.0286 | 183.3383 | 131.1829 | 126.7326 | 195.5634 | 267.8594 | 176.0708 | 172.6765 | 266.0423 | 215.6617 | 155.1554 | 169.1406 | 221.8731 | 195.6765 | 345.0137 | 330.4789 | 198.8731 | 282.3806 | 270.944 | 262.7611 | 322.8309 | 166.9725 | 141.6057 | 150.2114 | 157.2389 | 184.704 | 186.6194 | 160.1269 | 119.8309 | 299.1406 | 273.464 | 287.1691 | 308.8457 | 185.6903 | 147.6903 | 161.0137 | 155.1966 | 97.352 | 92.7886 | 84.8309 | 150.1269 | 243.2252 | 234.6903 | 252.4091 | 181.9017 | 174.3383 | 165.7463 | 159.1269 | 134.8309 | 219.7886 | 194.9577 | 232.8594 | 193.888 | 193.704 | 167.4926 | 199.76 | 160.3943 | 201.296 | 237.3806 | 226.7463 | 272.464 | 205.3806 | 195.0137 | 259.2103 | 226.2252 | 245.0137 | 212.4789 | 210.5211 | 311.76 | 211.9154 | 182.7748 | 176.1829 | 272.7463 | 164.4228 | 164.9577 | 165.6903 | 169.8309 | 178.6765 | 224.4651 | 211.8731 | 157.3943 | 129.6194 | 122.4651 | 140.6194 | 152.3383 | 295.1554 | 298.6057 | 275.7177 | 373.8457 | 287.4651 | 288.1977 | 281.4091 | 364.536 | 165.8446 | 82.9577 | 91.2674 | 165.8446 | 135.9292 | 147.6617 | 129.9154 | 141.4228 | 120.8731 | 107.352 | 147.3097 | 131.8731 | 91.0423 | 85.1829 | 91.648 | 113.5772 | 112.9154 | 115.0846 | 103 | 102.8171 | 303.0571 | 230.1406 | 243.3668 | 307.0148 | 195.9154 | 207.7886 | 223.7886 | 277.5772 | 315.4366 | 255.8594 | 260.9017 | 337.6205
858 | 11.4324 | 182.6657 | 234.6003 | 231.1322 | 249.9391 | 193.7994 | 188.4636 | 179.3298 | 208.2645 | 195.8009 | 134.1352 | 141.0684 | 201.2675 | 269.5988 | 170.8633 | 180.5319 | 255.2645 | 223.4696 | 165.7325 | 182.3373 | 220.6003 | 204.5988 | 320.1307 | 335.8678 | 210.9361 | 278.731 | 259.6627 | 251.728 | 306.1976 | 163.9286 | 139.7325 | 150.7325 | 151.7355 | 180.6672 | 179.6003 | 170.4681 | 115.4666 | 294.2006 | 257.4681 | 290.4666 | 317.3313 | 170.263 | 167.4696 | 153.2006 | 150.3358 | 95.0669 | 95.4681 | 81 | 144.4651 | 250.8693 | 224.0639 | 266.5988 | 197.4012 | 192.6687 | 161.6003 | 161.4666 | 142.2021 | 210.1991 | 197.2006 | 236.1991 | 203.8648 | 187.5334 | 165.0684 | 177.1322 | 169.2021 | 219.2021 | 208.6597 | 233.8024 | 253.8009 | 198.5304 | 177.5988 | 247.0699 | 218.3997 | 224.3982 | 235.4711 | 222.7355 | 316.3373 | 194.5304 | 198.0684 | 192.4711 | 255.7979 | 175.0669 | 151.1976 | 151.8633 | 176.6687 | 191.3328 | 221.6642 | 208.4666 | 161.9346 | 134.8693 | 140.4681 | 134.6672 | 157.3328 | 296.6642 | 269.5274 | 279.8708 | 371.1292 | 291.0654 | 287.997 | 289.731 | 386.3997 | 153.6672 | 78.266 | 88.1337 | 153.6672 | 137.9361 | 142.6672 | 128.5334 | 144.0654 | 114.266 | 121.6033 | 128.5973 | 127.9331 | 86.1322 | 76.3997 | 88.5988 | 110.4012 | 109.9331 | 106.3313 | 102.4666 | 112.1337 | 305.1292 | 234.269 | 264.0669 | 319.3313 | 195.6003 | 228.0714 | 227.0015 | 282.4027 | 302.7979 | 241.0624 | 256.1307 | 343.7964
859 | 11.4434 | 180.1829 | 234.3628 | 236.9145 | 240.9822 | 200.2773 | 170.7257 | 173.3688 | 199.6372 | 191.2684 | 134.2684 | 135.3569 | 203.3628 | 278.1888 | 163.7316 | 169.6342 | 267.9204 | 226.6312 | 175.0974 | 171.5309 | 214.3599 | 199.5457 | 333.649 | 334.4513 | 210.3569 | 275.6401 | 258.4632 | 225.3628 | 302.0944 | 155.0974 | 140.9115 | 153.003 | 171.2773 | 185.1829 | 183.1829 | 182.4572 | 107.9056 | 303.003 | 274.3688 | 287.8171 | 285.0796 | 165.0059 | 171.7227 | 159.2743 | 163.6372 | 100.3658 | 89.9941 | 89.1859 | 139.1829 | 253.177 | 237.1948 | 269.1859 | 187.7197 | 181.2625 | 169.5487 | 164.2743 | 145.4513 | 220.1888 | 203.2743 | 234.7286 | 203.7316 | 189.0914 | 166.4513 | 170.9086 | 165.3569 | 218.6312 | 194.5546 | 247.7286 | 261.2743 | 189.3658 | 178.5487 | 267.4572 | 232.2803 | 227.4602 | 258.7286 | 232.4543 | 330.0885 | 197.3717 | 205.4543 | 189.5339 | 261.826 | 163.4483 | 154.7346 | 143.6401 | 186.5457 | 202.0974 | 201.8141 | 203.6342 | 169.6372 | 140.4513 | 142.0885 | 143.003 | 151.7257 | 292.0944 | 268.7405 | 280.9911 | 353.6372 | 287.1829 | 280.003 | 272.4513 | 364.2625 | 153.8171 | 80.4572 | 83.4513 | 153.8171 | 134.0825 | 152.6401 | 127.3628 | 140.7286 | 115.9115 | 130.4513 | 119.2743 | 131.9115 | 89.1859 | 83.7316 | 83.5457 | 107.8141 | 120.4602 | 106.823 | 93.2684 | 97.0825 | 300.7346 | 252.0944 | 256.2684 | 313.8201 | 194.8171 | 231.4454 | 221.9941 | 307.5517 | 316.4661 | 234.6431 | 238.5428 | 336.823
860 | 11.4544 | 161.3534 | 235.7901 | 258.2763 | 212.116 | 203.5359 | 175.8122 | 193.7183 | 207.1382 | 195.7183 | 134.8122 | 134.0443 | 202.558 | 279.3977 | 171 | 168.2321 | 270.2595 | 217.442 | 173.1656 | 158.7183 | 220.8344 | 207.2542 | 338.2595 | 328.442 | 198.442 | 269.7237 | 257.9336 | 244.0886 | 306.326 | 154.8397 | 152.8122 | 135.6794 | 174.5359 | 178.6298 | 187.7901 | 192.5802 | 89.1656 | 306.326 | 276.3038 | 284.884 | 283.8565 | 165.5137 | 174.0443 | 177.6244 | 166.116 | 92.8397 | 82.7901 | 88.1878 | 149.9282 | 235.5137 | 238.3534 | 277.674 | 176.116 | 171.8122 | 175.558 | 172.5802 | 132.1878 | 217.4916 | 195.9557 | 231.9779 | 205.9779 | 169.3534 | 171.6023 | 173.9061 | 155.116 | 204.4198 | 224.5084 | 235.4916 | 264.558 | 197.4641 | 179.5359 | 251.3534 | 258.3924 | 244.2542 | 240.9114 | 228.6519 | 324.326 | 203.1878 | 189.3756 | 180.6023 | 272.2099 | 168.0664 | 167.558 | 161.1603 | 190.9061 | 206.8618 | 186.6298 | 198.884 | 164.3038 | 137.7901 | 133.5359 | 127.9114 | 141.7458 | 280.1435 | 285.4198 | 283.2985 | 364.4863 | 292.9061 | 282.2099 | 264.7679 | 340.8618 | 159.8122 | 90.5802 | 89.7183 | 159.8122 | 123.9061 | 144.4916 | 131.0221 | 141.326 | 130.0443 | 116.6298 | 126.4641 | 119.2595 | 97.116 | 94.3481 | 86.2321 | 112.7183 | 117.7237 | 118.3481 | 99.9504 | 83.558 | 329.7405 | 280.3206 | 259.6023 | 304.7237 | 195.2321 | 230.5084 | 223.1603 | 307.9557 | 315.3534 | 244.884 | 237.9061 | 343.884
861 | 11.4654 | 160.9686 | 223.172 | 266.8594 | 224.4061 | 184.0314 | 178.0079 | 191.0236 | 217.7031 | 197.586 | 147.2734 | 137.4375 | 195.586 | 263.586 | 177.2734 | 168.2891 | 264.414 | 219.2812 | 171.5546 | 157.1642 | 232.7109 | 194.7502 | 331.2734 | 321.1563 | 198.5703 | 250.8751 | 252.2734 | 250.3048 | 311.2734 | 142.5782 | 157.2891 | 130.5467 | 173.8515 | 178.8437 | 183.7266 | 185.0236 | 94.3983 | 303.8594 | 275.5625 | 293.6952 | 294.5782 | 171.117 | 171.1642 | 175.172 | 163.0079 | 95.9764 | 97.5467 | 92.8358 | 149.0157 | 222.2969 | 226.5625 | 271.0157 | 185.5546 | 174.5782 | 166.8751 | 180.9921 | 151.6638 | 213.414 | 199.9764 | 248.531 | 198.5782 | 154.7109 | 179.4297 | 178.1406 | 150.2969 | 209.1249 | 208.055 | 225.7109 | 260.4375 | 193.586 | 179.9921 | 236.1406 | 251.8908 | 253.9921 | 245.531 | 223.1485 | 361.7808 | 192.4375 | 188.9764 | 196.9843 | 271.1406 | 176.0079 | 180.5467 | 152.8908 | 194 | 208.8437 | 197.1092 | 211.117 | 155.0079 | 142.8515 | 135.1328 | 122.2655 | 133.7188 | 283.1092 | 287.8437 | 282.172 | 361.7345 | 294.2891 | 266.3126 | 277.828 | 353.6795 | 152.883 | 83.5939 | 105.8437 | 152.883 | 125.8594 | 134.7109 | 130.4375 | 136.5782 | 130.0157 | 117.9843 | 120.8751 | 125.3904 | 100.2812 | 94.1485 | 88 | 103.1799 | 118.8358 | 115.8672 | 103.0157 | 90.8437 | 328.7423 | 279.7502 | 261.1563 | 305.8358 | 194.1485 | 239.1485 | 218.883 | 305.1328 | 328.6559 | 232.0236 | 250.1249 | 325.8908
862 | 11.4764 | 177.6608 | 233.5561 | 253.1829 | 207.3657 | 195.8865 | 197.7213 | 197.7301 | 210.9306 | 173.3569 | 156.0782 | 137.4956 | 186.5044 | 259.7478 | 179.6696 | 168.1652 | 266.5044 | 218.6696 | 183.826 | 161.8171 | 222.9306 | 192.6519 | 322.6003 | 313.087 | 186.1829 | 239.3392 | 241.2699 | 241.2522 | 332.3127 | 171.4603 | 159.4956 | 126.6003 | 174.8348 | 196.2345 | 174.9218 | 177.7478 | 108.913 | 291.9306 | 267.9306 | 265.4616 | 297.2434 | 176.9218 | 172.9042 | 161.087 | 145.4351 | 102.6696 | 90.1047 | 89.9306 | 150.5738 | 227.9042 | 233.4956 | 274.9042 | 190.8348 | 176.0782 | 170.4868 | 186.3304 | 157.0177 | 202.1047 | 204.3392 | 247.6873 | 187.261 | 171.7301 | 154.5309 | 173.174 | 154.1564 | 202.6003 | 210.3039 | 230.4956 | 257.5826 | 188 | 187.6608 | 235.8348 | 244.2434 | 259.3304 | 242.9395 | 228.5738 | 349.053 | 191.913 | 189.261 | 197.7566 | 269.087 | 169.5044 | 161.4527 | 149.3216 | 184.6784 | 204.6784 | 212.3304 | 194.2005 | 162.4868 | 135.0958 | 144.826 | 117.1829 | 146.5649 | 256.966 | 283.6784 | 292.5561 | 355.1652 | 291.8348 | 275.2257 | 275.0177 | 369.3304 | 154.6519 | 78.4956 | 104.0088 | 154.6519 | 137.8171 | 133.6696 | 111.2699 | 125.8436 | 128.6608 | 112.3481 | 117.4956 | 139.4174 | 102 | 100.1564 | 89.7478 | 104.7301 | 109.5221 | 115.4956 | 97 | 95.4174 | 334.8953 | 270.6608 | 249.5914 | 297.1047 | 194.913 | 221.2699 | 224.1475 | 294.4351 | 327.6961 | 223 | 238.9395 | 321.739
863 | 11.4874 | 181.5949 | 243.646 | 239.8358 | 216.3649 | 198.7227 | 203.8869 | 198.8869 | 186.3431 | 188.5546 | 156.0256 | 138.2153 | 192.3284 | 278.2518 | 180.3796 | 174.354 | 286.6058 | 224.7336 | 174.9125 | 175.7591 | 223.9233 | 200.4051 | 321.3284 | 315.7591 | 193.6569 | 267.5291 | 246.6569 | 245.9489 | 327.938 | 163.6607 | 157.2409 | 141.3904 | 176.3796 | 188.3431 | 167.8358 | 164.7227 | 105.646 | 276.8613 | 274.2773 | 249.5182 | 305.7591 | 188.2773 | 177.0256 | 155.4307 | 137.3796 | 86.7227 | 82.7591 | 75.4562 | 154.2153 | 254.6313 | 241.354 | 268.3176 | 185.2409 | 174.8358 | 178 | 154.6862 | 155.3284 | 195.9489 | 206.9489 | 244.3284 | 189.9233 | 183 | 156.6313 | 188.6313 | 160.5949 | 193.5949 | 219.2664 | 223.4818 | 247.8869 | 187.4051 | 178.5073 | 242.1387 | 242.0511 | 264.5693 | 236.5438 | 238.7591 | 339.2262 | 211.2518 | 192.5182 | 191.0256 | 274.1387 | 152.7227 | 164.7955 | 153.6204 | 178.5949 | 218.2262 | 200.3176 | 205.9342 | 152.1533 | 133.354 | 143.0511 | 122.8722 | 157.5949 | 267.6678 | 273.4307 | 284.7993 | 354.8102 | 289.2153 | 282.4562 | 276.8978 | 375.1642 | 169.5438 | 77.4307 | 84.1278 | 169.5438 | 122.0147 | 143.3029 | 105.9489 | 132.4927 | 123.0767 | 118.8722 | 126.5438 | 133.646 | 88.9125 | 108.9744 | 77.3176 | 123.7336 | 93.4307 | 107.292 | 93.4307 | 86.0767 | 325.584 | 259.1278 | 235.4818 | 312.9598 | 189.2664 | 220.1131 | 224.8869 | 259.0147 | 315.1642 | 247.9853 | 219.4562 | 327.4051
864 | 11.4984 | 191.1077 | 233.321 | 237.6072 | 217.3918 | 187.1785 | 199.4287 | 198.0718 | 187.1436 | 195.8564 | 153.3928 | 135.1785 | 180.3918 | 282.7138 | 177.7497 | 185.2862 | 288.0349 | 240.9292 | 172.0718 | 192.9651 | 234.8574 | 216.3938 | 330.6431 | 323.8574 | 208.6431 | 285.9641 | 268.9292 | 244.5354 | 330.3579 | 140.5713 | 144.8923 | 139.7846 | 154.9272 | 176.3928 | 178.3579 | 171.3938 | 104.4287 | 298.7159 | 283.3928 | 248.7138 | 330.8585 | 182.8205 | 166.4995 | 156.6431 | 142.0359 | 84.8933 | 94.5005 | 83.7508 | 153.6072 | 243.7128 | 234.6779 | 272.9651 | 174.1067 | 167.1426 | 167.0708 | 152.6441 | 148.8564 | 220.6441 | 195.2133 | 253.6431 | 217.2513 | 169.6421 | 164.5354 | 180.1415 | 171.3221 | 194.6072 | 232.2154 | 259.5026 | 240.3928 | 190.0359 | 191.8585 | 253.679 | 241.0359 | 250.6062 | 245.2503 | 218.9272 | 346.9282 | 201.1415 | 206.8933 | 193.2503 | 282.0359 | 150.2862 | 183.4287 | 144.1067 | 168.6779 | 206.4985 | 186.7497 | 208.1774 | 144.8574 | 141.8574 | 134.7497 | 118.4277 | 156.1785 | 305.5364 | 274.6431 | 274.4646 | 377.68 | 302.5723 | 259.7846 | 275.2851 | 370.1067 | 156.999 | 84.7149 | 70.9641 | 156.999 | 122.001 | 132.7846 | 116.679 | 137.3569 | 112.1426 | 134.4646 | 123.1067 | 129.3928 | 76.9641 | 97.6421 | 79.5364 | 113.8205 | 100.7149 | 103.0359 | 92.8215 | 88.5005 | 312.8215 | 272.0728 | 248.4297 | 307.1056 | 203.4297 | 231.8574 | 245.9303 | 275.3949 | 319.8215 | 234.641 | 229.5723 | 336.1077
865 | 11.5094 | 205.0531 | 232.7168 | 267.1151 | 213.4336 | 187.8584 | 207.8142 | 213.1505 | 206.1505 | 181.8053 | 149.2832 | 135.8584 | 190.062 | 274.9026 | 165.0885 | 173.2743 | 272.8053 | 221.3628 | 181.5752 | 180.938 | 229.3274 | 216.469 | 314.4159 | 340.6284 | 204.1858 | 268.5132 | 257.4159 | 242.8584 | 334.6637 | 151.3894 | 149.531 | 150.3009 | 146.1947 | 195.2036 | 191.9558 | 173.3274 | 102.9026 | 288.7433 | 274.9469 | 266.9204 | 321.4601 | 176.0974 | 156.6637 | 146.6106 | 147.7168 | 105.4868 | 86.9911 | 96.9558 | 146.5664 | 232.9558 | 231.7168 | 280.7611 | 179.531 | 167.7168 | 149.469 | 161.6637 | 162.062 | 239.5752 | 186.2389 | 238.0354 | 211.7964 | 176.4868 | 158.5221 | 154.8495 | 168.7079 | 193.7611 | 221.6991 | 258.5044 | 251.7699 | 193.2389 | 184.938 | 240.6548 | 224.4159 | 241.8584 | 232.5132 | 207.0974 | 336.8053 | 204.3452 | 203.3274 | 196 | 270.3716 | 151.9026 | 165.1769 | 143.3363 | 166.9558 | 209.8231 | 193.9115 | 198.4779 | 157.9115 | 127.6548 | 122.7079 | 135.6372 | 154.3805 | 284.5044 | 307.3541 | 272.4248 | 399.1947 | 270.4955 | 277.7346 | 252.8938 | 347.6548 | 146 | 88.5221 | 79.531 | 146 | 139.8142 | 116.6637 | 114.1858 | 132.5221 | 118.2921 | 129.9469 | 121.0974 | 133.3363 | 84.2921 | 83.4248 | 88.2389 | 100.2832 | 117.531 | 113.6726 | 97.7168 | 99.5752 | 320.1947 | 269.6548 | 248.6106 | 310.8231 | 207.9469 | 217.6548 | 252.8495 | 275.938 | 313.5664 | 236.062 | 237.7611 | 325.8938
866 | 11.5204 | 187.2566 | 250.1623 | 259.0978 | 209.7871 | 187.1047 | 198.733 | 183.3578 | 193.4154 | 192.5305 | 143.2094 | 132.5777 | 180.2566 | 276.1588 | 171.0035 | 159.8412 | 263.3141 | 205 | 175.5236 | 171.7906 | 225.8953 | 204.3141 | 323.2164 | 357.2129 | 210.1082 | 265.9494 | 237.1553 | 240.8412 | 328.8412 | 165.3176 | 166.7399 | 146.9424 | 164.5305 | 177.3072 | 187.4189 | 166.7365 | 121.8481 | 280.8447 | 273.7906 | 272.7871 | 313.9494 | 183.6859 | 160.3176 | 141.527 | 146.2094 | 100.4695 | 84.0541 | 101.8953 | 150.2129 | 242.3176 | 230.2094 | 255.3613 | 192.9494 | 171.2635 | 143.6318 | 160.2635 | 153.5201 | 229.1012 | 187.6318 | 221.3141 | 190.5236 | 170.206 | 157.6318 | 168.8481 | 156.6824 | 193 | 204.8412 | 268.1658 | 257.7365 | 188.9459 | 186.5305 | 237.5811 | 223.7399 | 243 | 235.6352 | 223.5305 | 325.4189 | 195.4154 | 192.3141 | 196.6318 | 283.4799 | 160.7399 | 175.7434 | 149.7906 | 168.7365 | 209.3648 | 194.9459 | 197.473 | 172.8447 | 123.9494 | 123.3176 | 136.8377 | 150.8412 | 292.9022 | 308.9424 | 288.5846 | 382.7836 | 261.4258 | 285.8918 | 282.1692 | 372.3787 | 146.6318 | 95.2129 | 75.2601 | 146.6318 | 129.4695 | 125.3717 | 120.7399 | 133.527 | 133.4764 | 125 | 112.8918 | 131.5777 | 95.6859 | 87.5811 | 89.6318 | 110.0035 | 103.1519 | 113.9459 | 103.1588 | 89.733 | 308.206 | 248.8918 | 261.2164 | 321.7365 | 211.2129 | 219.0035 | 242.8412 | 281.3211 | 312.7906 | 266.6893 | 231.3141 | 336.2164
867 | 11.5314 | 176.5762 | 257.7119 | 236.2714 | 242.3551 | 202.1013 | 187.7797 | 186.474 | 200.6775 | 192.6273 | 137.7797 | 142.8811 | 197.2705 | 258.6784 | 159.3225 | 173.4573 | 278.0335 | 214.6608 | 165.4916 | 186.237 | 251.474 | 196.4916 | 331.1357 | 345.2546 | 190.2379 | 273.8643 | 238.4405 | 241.5762 | 334.7286 | 147.7463 | 169.1357 | 128.6273 | 165.2714 | 178.6775 | 190.7286 | 175.6608 | 124.3392 | 307.2537 | 255.3903 | 266.0678 | 293.526 | 188.2881 | 160.7797 | 152.0167 | 158.8132 | 97.7965 | 95.3727 | 87.5427 | 149.8476 | 231.8308 | 220.2714 | 260.9656 | 176.3903 | 175.8643 | 162.0335 | 172.5929 | 163.5418 | 233.237 | 176.4071 | 250.2026 | 205.6097 | 168.7286 | 175.3894 | 178.4238 | 145.2714 | 189.1357 | 195.9154 | 236.9842 | 241.5427 | 184.0678 | 189.2035 | 231.6952 | 242.237 | 233.3392 | 257.1692 | 237.1524 | 313.9154 | 193.0167 | 175.4749 | 202.7965 | 278.6106 | 168.9322 | 169.0344 | 144.2714 | 176.3727 | 210.3727 | 194.5762 | 195.3559 | 160.6106 | 134.4405 | 119.2714 | 130.5762 | 133.5427 | 295.1189 | 277.7463 | 280.0344 | 375.5084 | 277.4405 | 274.8476 | 290.1868 | 371.7463 | 162.4573 | 100.6441 | 79.9489 | 162.4573 | 127.4405 | 126.2035 | 119.2714 | 135.6441 | 124.1868 | 121.7797 | 105.0678 | 121.9154 | 89.9833 | 98.4405 | 88.7119 | 108.6273 | 100.0167 | 114.2203 | 94.6952 | 98.1013 | 309.9489 | 268.1181 | 257.5427 | 295.8819 | 202.4749 | 229.2203 | 226.1868 | 291.3559 | 323.3727 | 229.3401 | 230.5762 | 329.9665
868 | 11.5424 | 181.9381 | 256.3437 | 255.6284 | 226.9011 | 184.7462 | 170.9041 | 172.4336 | 208.3746 | 174.2168 | 132.0619 | 138.1858 | 187.7462 | 249.3127 | 165.4086 | 173.4675 | 274.8731 | 227.2198 | 189.2538 | 177.9041 | 247.2787 | 198.5944 | 319.1548 | 326.1548 | 169.4366 | 277.3127 | 230.1858 | 254.8142 | 313.4026 | 158.972 | 147.966 | 133.1888 | 170.8452 | 198.1888 | 178.5914 | 175.7492 | 121.6254 | 284.2138 | 278.1299 | 266.969 | 283.9381 | 182.4366 | 155.7183 | 151.7492 | 157.1548 | 89.8422 | 92.7802 | 81.6254 | 154.2198 | 249.5974 | 225.8452 | 272.6563 | 177.4705 | 167.4985 | 158.2168 | 176.3746 | 161.5604 | 240.6563 | 192.3157 | 243.7153 | 192.0589 | 167.7492 | 173.8422 | 175.031 | 142.969 | 215.2228 | 228.7553 | 232.003 | 233 | 188.2508 | 184.031 | 218.7802 | 253.5944 | 231.9381 | 249.9041 | 231.4675 | 369.0709 | 199.969 | 191.6284 | 181.7153 | 278.1888 | 166.0619 | 176.3776 | 159.034 | 185.5944 | 210.4056 | 207.8142 | 200.9071 | 139.1858 | 129.4675 | 123.5325 | 126.7492 | 140.7523 | 291.2817 | 273.2198 | 275.5634 | 349.1209 | 301.3467 | 257.5604 | 280.031 | 377.0649 | 145.4026 | 95.7492 | 74.8422 | 145.4026 | 125.7492 | 132.1888 | 115 | 146.5015 | 119.9381 | 135.7523 | 107.2817 | 136.3776 | 83.031 | 90.1858 | 89.969 | 94.8112 | 112.8761 | 121.9071 | 100.8142 | 103.0619 | 323.2198 | 255.7462 | 236.5295 | 283.3127 | 214.034 | 238.8761 | 235.7213 | 283.7802 | 320.7802 | 228.5974 | 236.5944 | 321.3127
869 | 11.5534 | 213.4199 | 235.941 | 240.586 | 228.059 | 188.3845 | 188.4081 | 190.7631 | 211.6804 | 167 | 138.6923 | 118.6273 | 200.7454 | 247.9941 | 167.9823 | 179.6982 | 274.6863 | 228.9941 | 185.6155 | 163.3137 | 254.7159 | 209.0236 | 344.7572 | 314.9823 | 197.4258 | 266.6332 | 250.0768 | 264.3432 | 332.7631 | 164.3137 | 156.7218 | 132.6509 | 171.3196 | 178.2605 | 194.0709 | 162.9705 | 114.9764 | 298.0945 | 315.3904 | 276.0236 | 306.059 | 179.6686 | 157.3432 | 144.9882 | 158.6863 | 98.71 | 94.3491 | 94.3668 | 166.0236 | 260.9941 | 228.3255 | 246.9232 | 182.6627 | 180.0531 | 157.0177 | 174.3314 | 147.9823 | 216.2605 | 181.9351 | 214.9528 | 188.3609 | 159.6509 | 145.9351 | 177.3432 | 144.6686 | 232.0059 | 219.9173 | 246.3432 | 257.7395 | 197.0177 | 187.0118 | 247.7631 | 237.2782 | 240.6863 | 235.9823 | 229.0059 | 356.5387 | 211.6982 | 211.355 | 174.3491 | 280.3255 | 157.9823 | 163.9351 | 157.6332 | 190.0059 | 205.3255 | 210.6568 | 181.9351 | 151.0531 | 113.6332 | 125.3255 | 136.7041 | 153.6804 | 283.6391 | 267.6391 | 275.6568 | 333.3314 | 295.2841 | 257.355 | 281.0059 | 345.2133 | 142.6982 | 98.3491 | 81.7041 | 142.6982 | 124.3373 | 134.3255 | 126.3668 | 142.6391 | 105.9528 | 133.9705 | 113.6804 | 138.645 | 90.0236 | 89.3491 | 90.3314 | 105.3786 | 119.0059 | 111.6273 | 99.645 | 94.3137 | 312.29 | 268.0768 | 234.0118 | 296.7041 | 191.9055 | 240.3255 | 259.71 | 284.0118 | 295.6037 | 232.6391 | 225.6273 | 324.0059
870 | 11.5644 | 187.1444 | 234.1711 | 216.7861 | 224.3815 | 181.3387 | 188.0624 | 189.9769 | 205.148 | 184.0231 | 141.6381 | 133.7896 | 193.5722 | 240.8717 | 150.7006 | 187.0428 | 292.2994 | 213.0196 | 183.1283 | 150.467 | 228.2728 | 205.5098 | 351.8289 | 337.1943 | 209.5954 | 246.7006 | 263 | 232.6346 | 355.1283 | 168.4902 | 158.5098 | 128.6381 | 165.5954 | 204.8128 | 193.7433 | 153.2335 | 104.1907 | 307.8717 | 295.7202 | 270.467 | 315.3191 | 188.1711 | 169.8948 | 154.5757 | 154.5098 | 100.5526 | 90.8717 | 91.148 | 160.467 | 252.5098 | 231.7665 | 255.1087 | 206.5133 | 207.385 | 155.9144 | 159.0196 | 128.6578 | 198.5526 | 171 | 219.9376 | 197.0856 | 167.2139 | 148.6185 | 163.3387 | 157.2567 | 226.8717 | 217.5757 | 257.533 | 263.8717 | 202.0428 | 182.8717 | 262.6381 | 238.2139 | 256.9376 | 231.6381 | 223.8717 | 343.852 | 202.0196 | 204.1052 | 183.1283 | 281.7237 | 159.0856 | 164.5757 | 142.467 | 187.5954 | 191.7433 | 210.3619 | 204.3654 | 150.467 | 123.6613 | 123.3191 | 143 | 164.1711 | 304.8752 | 277.2994 | 270.5954 | 343.2139 | 293.8093 | 267.8093 | 273.8289 | 339.9804 | 137.7861 | 86.0196 | 75.0624 | 137.7861 | 127.7237 | 143.2139 | 130.6381 | 127.1052 | 112.9804 | 130.3619 | 118.0428 | 126.8289 | 84.467 | 85.1907 | 97.4902 | 130.7041 | 112.5098 | 114.533 | 106.2139 | 101.8948 | 292.0624 | 253.082 | 238.7237 | 324.1087 | 205.2798 | 238.3191 | 239.082 | 299.6185 | 301.3422 | 228 | 212.8717 | 312.7433
871 | 11.5754 | 178.3983 | 233.1475 | 236.9558 | 225.6254 | 177.4661 | 196.5576 | 163.9764 | 217.6372 | 180.9085 | 162.4897 | 141.2271 | 203.4101 | 255.3304 | 155.0915 | 180.3746 | 271.2713 | 214.3186 | 170.7492 | 177.888 | 225.8643 | 191.6017 | 328.5899 | 319.1917 | 207.3068 | 264.9558 | 258.1475 | 237.7965 | 359.3864 | 152.5899 | 142.5221 | 123.8407 | 182.7169 | 216.761 | 179.6814 | 157.9322 | 94.761 | 294.6017 | 291.705 | 279.1711 | 276.1799 | 185.761 | 169.4543 | 139.2035 | 155.1593 | 98.6932 | 84.5339 | 87.6932 | 153.2271 | 239.9882 | 232.6136 | 260.8407 | 183.3392 | 192.4306 | 160.9322 | 149.2271 | 134.8643 | 218.1829 | 188.3304 | 212.8289 | 202.4661 | 182.3983 | 138.3628 | 178.1829 | 165.7729 | 225.3864 | 218.1475 | 248.8289 | 263.0797 | 186.3628 | 184.1593 | 251.6017 | 252.705 | 265.7729 | 249.7169 | 211.9882 | 340.3746 | 220.649 | 196.9203 | 180.4543 | 273.9882 | 185.9558 | 172.7729 | 156.0236 | 188.0797 | 193.6254 | 193.6696 | 220.6932 | 161.944 | 136.5457 | 145.1829 | 140.9203 | 162.4543 | 283.0324 | 254.1917 | 240.5781 | 354.9322 | 290.0678 | 280.0118 | 282.4779 | 344.9203 | 137.8525 | 88.0118 | 86.3304 | 137.8525 | 133.1593 | 130.6696 | 124.4543 | 122.6932 | 124.1593 | 150.4101 | 127.3186 | 130.6254 | 80 | 85.4661 | 83.6696 | 116.1238 | 105.5339 | 133.5576 | 102.6814 | 92.4424 | 332.4455 | 257.3304 | 244.8525 | 308.351 | 200.2831 | 224.8289 | 205.8967 | 295.6017 | 321.4779 | 227.3068 | 231.4897 | 309.0797
872 | 11.5864 | 195.6991 | 218.3009 | 234.5953 | 227.589 | 181.822 | 188.1843 | 157.411 | 196.7797 | 169.6504 | 154.3623 | 140 | 196.4788 | 258.0614 | 167.7606 | 181.9386 | 275.2267 | 218.7055 | 166.1165 | 177.1843 | 214.3623 | 217.3368 | 331.9936 | 318.6991 | 201.3559 | 283.0551 | 258.1165 | 235.0064 | 307.0869 | 148.233 | 140.5275 | 136.1102 | 167.0127 | 200.5953 | 183.7606 | 160.2945 | 100.4661 | 277.3009 | 259.3136 | 273.0064 | 264.6441 | 168.8898 | 173.3496 | 156.8093 | 178.8707 | 108.8771 | 85.1165 | 87.2945 | 133.6568 | 241.6441 | 244.6991 | 257.589 | 177.8771 | 186.233 | 139.3072 | 143.767 | 131.8284 | 204.7182 | 201.6441 | 215.4661 | 211.0551 | 192.6441 | 132.411 | 186.589 | 169.822 | 218.2394 | 227.2881 | 233.8284 | 271.7606 | 174.0614 | 190.9386 | 255.4661 | 256.2945 | 275.4661 | 248.1229 | 209.411 | 339.1165 | 209.4237 | 180.4788 | 187.8771 | 260.1229 | 175.1293 | 170.4725 | 153.4174 | 190.411 | 199.822 | 189.5275 | 221.7055 | 156.3009 | 139.7055 | 135.2458 | 117.4237 | 167.7606 | 301.8644 | 237.4725 | 254.8093 | 366.4661 | 298.2881 | 295.9936 | 299.2881 | 355.9936 | 131.5339 | 83.5339 | 79.8898 | 131.5339 | 118.7733 | 129.3496 | 121.2945 | 128.6441 | 121.767 | 132.8962 | 121.1229 | 154.4597 | 84.233 | 90.5275 | 86.5826 | 102.4725 | 106.1165 | 107.5466 | 110.9936 | 78.2394 | 317.0191 | 272.7606 | 253.3496 | 285.0064 | 186.3559 | 208.4174 | 206.8771 | 295.233 | 336.8771 | 238.2881 | 224.7733 | 298.0064
873 | 11.5974 | 203.1534 | 234.534 | 251.9696 | 223.411 | 210.2764 | 200.6628 | 179.534 | 187.1534 | 176.3314 | 149.1534 | 126.3618 | 197.8958 | 248.1042 | 171.7178 | 195.4848 | 266.2084 | 226.178 | 169.1534 | 163.822 | 237.1476 | 209.9016 | 346.3314 | 307.4906 | 187.5152 | 285.2822 | 283.4052 | 227.1288 | 300.0738 | 153.589 | 144.1534 | 146.3068 | 177.8162 | 197.8712 | 187.7178 | 161.4356 | 103.2822 | 308.6078 | 299.9883 | 273.0246 | 290.6874 | 167.3068 | 172.411 | 135.699 | 184.411 | 92.9016 | 89.589 | 79.822 | 139.6382 | 266.2518 | 227.7482 | 230.4414 | 182.7178 | 189.4356 | 149.8162 | 147.7424 | 129.4356 | 226.5832 | 211.178 | 232.6382 | 200.3618 | 185.6686 | 151.6628 | 177.3864 | 155.2084 | 225.767 | 209.7482 | 250.8162 | 298.6874 | 214.3501 | 175.055 | 252.5398 | 237.3372 | 256.0304 | 233.9508 | 207.1288 | 333.5398 | 212.2026 | 185.4848 | 189.1288 | 258.1534 | 180.356 | 171.8712 | 146.8466 | 189.5644 | 211.767 | 196.7424 | 202.6194 | 160.3314 | 142.1534 | 122.6932 | 120.9204 | 156.644 | 291.5948 | 260.4052 | 249.4906 | 385.7916 | 299.411 | 280.9016 | 291.0796 | 359.5644 | 161.0188 | 87.8958 | 78.3068 | 161.0188 | 117.7424 | 132.7178 | 118.1288 | 134.589 | 134.356 | 109.7974 | 122.7424 | 130.699 | 83.8466 | 83.3864 | 83.822 | 108.8958 | 104.1288 | 100.4602 | 107.3864 | 80.7424 | 291.9508 | 238.6745 | 259.589 | 311.5832 | 204.8162 | 210.4602 | 223.2026 | 278.3372 | 332.3864 | 241.5644 | 237.3806 | 303.0492
874 | 11.6084 | 190.8584 | 216.7168 | 247.1283 | 227.1106 | 200.5575 | 195.5885 | 179.2389 | 198.2213 | 193.1416 | 125.1769 | 122.4602 | 207.5708 | 257.4115 | 169.8097 | 173.7168 | 260 | 226.0796 | 178.7611 | 171.9514 | 216.1459 | 206.3805 | 341.2389 | 320.7124 | 201.9823 | 266.7477 | 291.5398 | 234.7611 | 316.9514 | 177.6328 | 138.4292 | 127.5575 | 179.4292 | 188.0486 | 185.8097 | 151.7787 | 103 | 309.1283 | 293.5265 | 280.8407 | 316.7921 | 174.1106 | 144.7168 | 155.8541 | 165.4778 | 99.6018 | 102.6814 | 87.9514 | 141.3495 | 253.8274 | 217.5398 | 238.2523 | 176.4292 | 192.9204 | 140.4778 | 147.8097 | 148.2523 | 225.8584 | 189.1769 | 228.5088 | 195.7301 | 160.0973 | 150.2389 | 177.6505 | 161.4115 | 219.7787 | 184.2079 | 228.3362 | 275.1459 | 185.7344 | 184.7921 | 260.2213 | 231.4602 | 253.5708 | 232.1903 | 200.1593 | 328.0796 | 203.1283 | 192.1903 | 191.6505 | 283.8231 | 169.2079 | 170.8097 | 160.6018 | 191.1903 | 212.3495 | 196.0796 | 189.1593 | 151.5885 | 134.2389 | 140.7124 | 135.4912 | 140.0486 | 274.6991 | 274.3805 | 243.7301 | 348.9247 | 298 | 297.093 | 258.2566 | 353.1593 | 152.0973 | 89.5398 | 90.2213 | 152.0973 | 138.9823 | 130.8097 | 108.2389 | 142.5708 | 124.6681 | 122.5222 | 123.5398 | 132.6018 | 77.8894 | 94.6018 | 84.6505 | 113.4602 | 130.7434 | 115.4115 | 119.3319 | 85.1903 | 297.4912 | 257.5841 | 272.6814 | 306.4778 | 196.2079 | 224.6814 | 241.1416 | 273.1903 | 299.7964 | 241.7301 | 240.6195 | 334.7434
875 | 11.6194 | 204.5595 | 214.4238 | 238.2881 | 249.0443 | 177.41 | 202.1357 | 205.6952 | 186.41 | 185.3795 | 137.5291 | 151.9529 | 206.2881 | 253.0914 | 186.0748 | 181.8171 | 271.1357 | 219.8033 | 163.4405 | 173.0305 | 207.7119 | 211.7119 | 334.2881 | 360.6648 | 203.8033 | 275.59 | 258.3352 | 253.59 | 341.7867 | 168.1829 | 138.9695 | 129.651 | 178.4848 | 181.7729 | 193.9086 | 158.3933 | 114.1357 | 317.3629 | 308.241 | 275.5762 | 318.5457 | 181.1967 | 149.8476 | 153.8948 | 177.5595 | 118.3629 | 92.8948 | 92 | 149.651 | 260.8476 | 206.6067 | 281.3767 | 199.241 | 182.8643 | 148.1052 | 159.6205 | 147.5762 | 241.7867 | 198.5595 | 211.6371 | 186.349 | 164.6205 | 159.6205 | 173.061 | 163.7729 | 198.9252 | 195.0748 | 249.6648 | 261.5152 | 196.4681 | 179.8643 | 259.5457 | 238.6814 | 244.1219 | 261.2105 | 212.8476 | 355.9529 | 215.0748 | 202.651 | 191.5152 | 284.0914 | 157.8033 | 186.3324 | 158.5762 | 185.3186 | 213.9695 | 179.41 | 215.9529 | 148.4848 | 135.4543 | 139.8338 | 127.8643 | 161.241 | 300.9529 | 245.5624 | 262.5595 | 353.0443 | 306.1662 | 282.2438 | 263.59 | 353.2271 | 138.8033 | 101.6205 | 84.349 | 138.8033 | 130.41 | 131.4848 | 110.1967 | 136.8338 | 128.651 | 138.651 | 129.6814 | 129.0914 | 79.7119 | 82.1829 | 89.7119 | 119.939 | 112.7895 | 118.5152 | 98.2743 | 84.5152 | 301.7424 | 245.5014 | 248.7895 | 314.8476 | 196.6814 | 238.651 | 231.1524 | 285.1357 | 306.8171 | 233.0914 | 239.7424 | 319.759
876 | 11.6304 | 199.6799 | 235.1121 | 238.5093 | 228.0771 | 180.3014 | 205.2453 | 182.0398 | 189.3014 | 193.0748 | 140.472 | 131.8131 | 213.3014 | 246.9813 | 178.4159 | 170.6425 | 242.5491 | 251.9602 | 185.6775 | 172 | 213.528 | 201.6799 | 349.6028 | 347.8505 | 202 | 299.8668 | 248.5093 | 278.6215 | 339.1893 | 128.0398 | 143.0187 | 119.4159 | 169.9439 | 171.9439 | 193.9813 | 162.7547 | 116.4907 | 314.4533 | 286.0584 | 284.3201 | 306.4346 | 203.3761 | 185.9416 | 145.2266 | 153.0584 | 112.4346 | 73.6612 | 96.528 | 150.736 | 250.1519 | 208.2827 | 254.0024 | 195.9252 | 167.6799 | 168.8481 | 160.9813 | 142.736 | 209.757 | 211.0374 | 217.5654 | 188.2827 | 153.9066 | 144.3785 | 177.0374 | 165.264 | 189.2266 | 211.5654 | 245.9066 | 242.8879 | 187.1332 | 184.3014 | 260.264 | 240.2453 | 237.736 | 252.1332 | 227.8107 | 335.0584 | 209.6799 | 218.0748 | 175.1519 | 287.792 | 137.1332 | 182.9439 | 153.736 | 185.264 | 223.3014 | 186.8294 | 207.8879 | 145.9813 | 155.8668 | 131.7173 | 126.264 | 151.8879 | 297.4159 | 251.6028 | 250.8879 | 348.6799 | 278.8131 | 286.8294 | 282.5841 | 331.3598 | 144.5467 | 95.4346 | 93.0748 | 144.5467 | 128.0187 | 122.9439 | 115.0187 | 139.2827 | 132.7547 | 138.2266 | 128.2266 | 119.9626 | 90.0561 | 88.0748 | 104.5841 | 105.3972 | 113.5654 | 115.736 | 105.6028 | 108.1495 | 329.1682 | 252.0934 | 244.2827 | 304.9066 | 202.0187 | 238.2266 | 235.0561 | 256.5491 | 313 | 227.736 | 246.0374 | 329.1121
877 | 11.6414 | 205.9705 | 252.5044 | 227.4956 | 213.6313 | 196.8053 | 219.5723 | 178.6696 | 194.3009 | 167.8554 | 145.1357 | 144.2419 | 206.7964 | 238.3304 | 180.9027 | 173.4366 | 260.3776 | 223.8849 | 182.7286 | 159.7286 | 231.8732 | 218.708 | 348.8643 | 333.8643 | 186.6607 | 263.0501 | 238.2625 | 255.0884 | 319.8938 | 160.7169 | 134.0295 | 134.1741 | 169.3009 | 179.7375 | 162.3215 | 165.3009 | 125.2036 | 292.8259 | 308.9794 | 278.0295 | 317.5723 | 190.8259 | 198.3009 | 135.5634 | 166.7759 | 101.3304 | 84.3393 | 81.1268 | 150 | 255.7375 | 236.0768 | 270.9116 | 192.767 | 177.0384 | 176.3009 | 153.0973 | 143.5339 | 197.767 | 185.3893 | 253.2125 | 193.8348 | 165.8732 | 164.0768 | 172.0973 | 160.6313 | 194.9027 | 237.2419 | 253.2714 | 238.5339 | 185.6018 | 187.767 | 246.4277 | 243.8348 | 233.1652 | 216.0884 | 221.7964 | 355.6785 | 198.3304 | 198.2241 | 177.6696 | 270.8259 | 160.9116 | 177.6991 | 157.6018 | 184.4661 | 205.292 | 195.6018 | 176.6902 | 147.3009 | 145.1268 | 127.6991 | 125.4661 | 159.8053 | 312.941 | 266.9705 | 291.0178 | 331.1947 | 294.3098 | 268.7581 | 268.5929 | 355.4455 | 142.3982 | 87.3982 | 80.8938 | 142.3982 | 157.3776 | 132.2714 | 109.8643 | 144.0679 | 119.1947 | 146.2036 | 126.233 | 127.2036 | 79.1947 | 92.767 | 83.6902 | 119.1741 | 107.0295 | 111.1652 | 109.4661 | 91.4572 | 307.3215 | 241.6607 | 236.7964 | 314.5723 | 198.3982 | 241.6018 | 223.4277 | 281.2803 | 322.9705 | 266.8821 | 231.1268 | 343.4366
878 | 11.6524 | 185.6224 | 230.2846 | 239.5851 | 214.3378 | 186.9734 | 188.1543 | 182.5585 | 194.2207 | 148.0904 | 146.2207 | 147.1037 | 179.8431 | 260.9361 | 167.4149 | 174.4415 | 262.766 | 226.0266 | 174.3245 | 164.5718 | 227.6489 | 212.5319 | 343.883 | 341.3511 | 204.5984 | 269.4814 | 263.0532 | 237.4282 | 316.5585 | 149.8431 | 158.2739 | 136.1037 | 183.2473 | 192.3511 | 163.1303 | 181.5851 | 116.3112 | 288.5585 | 301.0772 | 243.8298 | 290.8298 | 174.0904 | 176.4016 | 146.2473 | 150.6224 | 102.117 | 87.4415 | 83.7925 | 157.0133 | 251.9867 | 240.883 | 262.0772 | 197.6755 | 184.8963 | 162.1942 | 145.5452 | 137.766 | 185.5319 | 181.6755 | 227.9336 | 202.0133 | 171.7793 | 158.7527 | 178.5718 | 156.6622 | 204.0133 | 199.5825 | 235.9601 | 257.7021 | 188.5585 | 184.1037 | 240.4415 | 253.5718 | 227.3245 | 228.0399 | 208.0904 | 353.3112 | 193.6622 | 187.883 | 176.1037 | 261.1037 | 171.5585 | 156.7394 | 159 | 190.234 | 223.9361 | 189.2075 | 182.2473 | 127.7394 | 135.3245 | 133.234 | 137.4681 | 137.5054 | 294.0639 | 283.2473 | 247.3352 | 344.9228 | 282.5186 | 292.3909 | 308.976 | 378.2473 | 148.0133 | 90.6755 | 72.1037 | 148.0133 | 136.3883 | 143.7925 | 118.1303 | 131.7527 | 124.3511 | 133.4149 | 115.8697 | 124.5452 | 87.4681 | 78.9734 | 84.5718 | 123.4415 | 104 | 103.766 | 110.5585 | 92.5718 | 274.6224 | 260.3776 | 244.1303 | 310.4282 | 205.5718 | 242.2207 | 215.1037 | 272.5186 | 311.9734 | 240.0373 | 222.1037 | 360.8058
879 | 11.6634 | 181.3746 | 227.7493 | 211.5466 | 224.4985 | 198.8309 | 185.1239 | 194.0816 | 202.707 | 179.828 | 140.4592 | 127.7945 | 174.5831 | 279.8732 | 150.3353 | 179.5408 | 257.0423 | 207.0452 | 184.0816 | 181.2478 | 232.1239 | 194.7522 | 312.1297 | 309.172 | 185.6705 | 265.5015 | 270.2085 | 244.4985 | 338.3717 | 159.6225 | 154.1268 | 128.6676 | 171.9606 | 198.1662 | 157.293 | 186.7915 | 101.1268 | 341.242 | 280.9606 | 254.7886 | 304.9548 | 168.6254 | 159.7099 | 145.2507 | 170.9125 | 103.7915 | 89.3746 | 84.4169 | 131.2959 | 271.3717 | 224.9606 | 257.7915 | 187.1268 | 178.8761 | 184.121 | 148.7493 | 147.0816 | 214.4534 | 191.707 | 228.2901 | 188.9606 | 171.2085 | 177.9548 | 181 | 151.2507 | 196.5015 | 202.0394 | 220.5015 | 283.5802 | 184.2507 | 183.7915 | 237.6254 | 237.0029 | 237.8732 | 220.5437 | 219.2478 | 327.8367 | 202.4985 | 225.7857 | 196.3717 | 267.1239 | 161.7099 | 168.414 | 148.7099 | 190.4169 | 208.8367 | 210.7464 | 172.5437 | 150.4956 | 134 | 129.4592 | 127.5437 | 141.0816 | 298.8732 | 288.5831 | 277.7012 | 350.7915 | 273.8338 | 276.4621 | 265.0087 | 394.6647 | 145.2507 | 93.5831 | 86.8309 | 145.2507 | 144.6225 | 116.7128 | 124.9577 | 128 | 113.5437 | 129.7915 | 133.5802 | 126.9577 | 91.7915 | 87.6647 | 89.3746 | 100.0452 | 110.3324 | 119.414 | 105.4592 | 99.7493 | 283.0394 | 268.5831 | 238.293 | 315.9155 | 187.4198 | 234.8761 | 225.0816 | 289.1633 | 329.3717 | 254.3295 | 237.6225 | 355.5015
880 | 11.6744 | 203.7036 | 229.8038 | 231.1343 | 210.1194 | 207.823 | 180.5693 | 197.8038 | 189.7271 | 172.9233 | 171.9573 | 123 | 177.4115 | 247.6312 | 179.1535 | 168.9425 | 248.7655 | 213.0575 | 181.3731 | 169.7271 | 244.4499 | 208.6844 | 328.9189 | 289.5501 | 216.7803 | 272.646 | 265.9808 | 230.1194 | 316.6696 | 161.5885 | 144.5693 | 123.7847 | 179.2537 | 201.4115 | 163.8422 | 158.8657 | 109.2537 | 323.6504 | 307.5458 | 264.823 | 266.789 | 170.4115 | 157 | 161.6844 | 147.4542 | 115.2537 | 90.8038 | 77.5693 | 138.4691 | 237.612 | 243.5074 | 237.1003 | 187.2153 | 159.3156 | 172.5118 | 137.1386 | 141.9616 | 237.4691 | 197.2153 | 205.2772 | 179.3731 | 155.7271 | 157.4734 | 182.6077 | 162.0575 | 184.354 | 209.2153 | 241.3113 | 294.6269 | 160.4926 | 175.1578 | 229.7655 | 215.9233 | 241 | 216.1962 | 241.4882 | 305.9233 | 200.177 | 173.3008 | 179.4926 | 265.7847 | 169.4499 | 160.9425 | 143.5885 | 160.2581 | 178.0811 | 179.2197 | 201.1535 | 151.5693 | 119.5309 | 133.6269 | 127.2153 | 130.3348 | 284.3156 | 273.7271 | 267.4926 | 350.1962 | 283.4499 | 286.0767 | 271.7036 | 375.4926 | 137.5693 | 77.1194 | 74.1194 | 137.5693 | 140.9616 | 104.177 | 129.2153 | 120.7655 | 105.9808 | 132.4115 | 139 | 121.5693 | 83.9616 | 88.5885 | 88.3923 | 107.6652 | 124.8614 | 119.177 | 97.5693 | 84.1194 | 295.0384 | 243.2772 | 231.9808 | 292.2772 | 213.3496 | 211.2964 | 214.3348 | 295 | 314.9041 | 230.4542 | 258.8806 | 332.9041
881 | 11.6854 | 191.0452 | 209.5968 | 231.471 | 213.3451 | 203.2871 | 192.8742 | 175.1484 | 183.5516 | 187.7709 | 160.4129 | 123 | 198.4032 | 260.2193 | 182.1032 | 157.0226 | 270.6677 | 221.7129 | 163.6775 | 159.471 | 250.2645 | 205.6549 | 370.0935 | 309.8516 | 197.4355 | 264.3903 | 272.3451 | 281.4967 | 323.058 | 144.6775 | 147.0806 | 146.6677 | 170.5742 | 194.6549 | 164.3677 | 165.8742 | 121.7935 | 290.7 | 290.5162 | 292.9322 | 297.6225 | 184.8742 | 157.8161 | 151.3097 | 148.9774 | 95.9646 | 89.3677 | 72.7355 | 158.3225 | 268.8064 | 214.7226 | 237.7129 | 168.4129 | 175.4032 | 172.8968 | 159.3 | 148.1613 | 211.6194 | 189.8387 | 247.1515 | 190.2419 | 148.7355 | 159.9774 | 187.8968 | 156.0226 | 211.3806 | 223.058 | 238.8387 | 277.2291 | 170.5064 | 173.8161 | 239.4258 | 225.058 | 210.8033 | 219.2645 | 237.0226 | 330.5645 | 202.2645 | 184.9322 | 191.1387 | 238.8839 | 178.529 | 180.8516 | 163.4032 | 178.3 | 180.9774 | 192.0354 | 186.1484 | 147.5516 | 125.7935 | 146.4258 | 127.1839 | 140.8742 | 299.5871 | 263.471 | 277.5064 | 374.4838 | 295.7935 | 297.3451 | 250.8839 | 361.0226 | 146.6097 | 82.7935 | 83.058 | 146.6097 | 127.5742 | 133.1967 | 123.471 | 145.9322 | 128.6677 | 128.1032 | 106.3549 | 119.1839 | 94.2419 | 90.4484 | 92.8968 | 114.2645 | 129.6323 | 114.7355 | 96.8161 | 106.9322 | 310.8742 | 232.1032 | 258.7483 | 284.3677 | 191.6194 | 250.887 | 236.3 | 272.9646 | 311.6323 | 237.6903 | 267.0806 | 306.7807
882 | 11.6964 | 177.8874 | 221.5684 | 246.5684 | 219.9705 | 213.5979 | 196 | 198.9946 | 175.5442 | 176.26 | 141.9169 | 133.7694 | 189.7453 | 268.3137 | 176.0295 | 175.7105 | 268.5442 | 238.74 | 187.3378 | 181.1957 | 234.4316 | 197.3727 | 345.6917 | 313.3432 | 205.9115 | 266.9705 | 268.2011 | 234.3539 | 338.4263 | 149.2842 | 180.3083 | 133.7748 | 156.4021 | 189.6863 | 150.7453 | 154.0885 | 120.6863 | 297.4263 | 300.74 | 250.1233 | 291.26 | 168.9464 | 160.4853 | 175.3378 | 147.6863 | 105.0831 | 90.6568 | 82.7694 | 144.6032 | 239.8927 | 225.2252 | 261.3673 | 184.7105 | 167.5737 | 172.3432 | 145.9464 | 124.319 | 247.2493 | 202.0831 | 224.8633 | 208.74 | 167.882 | 146.26 | 169.118 | 172.2252 | 206.4021 | 213.5737 | 239.4853 | 271.3432 | 193.0536 | 160.7453 | 270.9946 | 217.2306 | 214.7694 | 220 | 237.4853 | 328.7158 | 182.2895 | 183.5442 | 188.3727 | 223.8874 | 184.9705 | 182.6863 | 168.8284 | 184 | 178.0295 | 206.1126 | 189.2842 | 151.1421 | 137.1126 | 155.6273 | 128.6568 | 174.6515 | 275.8338 | 241.2895 | 279.3432 | 351.0054 | 290.5442 | 279.9464 | 312.1019 | 350.7158 | 136.5737 | 80.0295 | 76.059 | 136.5737 | 120.8579 | 140.8284 | 122 | 157.7989 | 144.7694 | 119.5442 | 115.5684 | 127.2842 | 88.7158 | 74.4316 | 80.7453 | 115 | 107.6327 | 105.7158 | 123.5094 | 85.6622 | 268.437 | 274.9062 | 257.5442 | 253.3485 | 202.3968 | 235.319 | 246.1421 | 290.3673 | 325.2547 | 267.5094 | 239.0054 | 313.5979
883 | 11.7074 | 184.4071 | 218.2743 | 226.4601 | 219.3186 | 204.2301 | 199.3628 | 173.053 | 186.6106 | 209.9912 | 144.0443 | 115.823 | 222.3098 | 286.6549 | 185.0885 | 177.4779 | 278.7699 | 226.8672 | 172.823 | 155.7345 | 238.5664 | 204.4071 | 327.2301 | 330.6549 | 209.8407 | 263.7965 | 262.7965 | 243.0177 | 315.7787 | 148.4779 | 162.6194 | 132.5221 | 166.6106 | 185.6372 | 177.4248 | 151 | 118.3186 | 308.4071 | 271.2123 | 271.1063 | 293.0443 | 180.1328 | 166.885 | 140.6459 | 131.8672 | 92.0265 | 80.0708 | 96.7699 | 135.115 | 241.2478 | 228.1593 | 230.6902 | 182.2743 | 156.5929 | 156.8672 | 153.7699 | 116.4779 | 213.1238 | 207.5221 | 223.885 | 190.9823 | 172 | 144.6814 | 170.0443 | 168.4336 | 217.4514 | 191.6637 | 227.3894 | 271.8407 | 186.0708 | 179.8584 | 254.3008 | 233.4956 | 225.4071 | 226.7257 | 216.9823 | 321.9557 | 214.9912 | 189.5664 | 189.5221 | 270.7612 | 166.6637 | 180.3186 | 178.2478 | 173.9115 | 201.3806 | 203.7965 | 194.3628 | 146.115 | 133.9557 | 145.2301 | 134.885 | 165.8672 | 268.3186 | 258.0177 | 288.2478 | 361.8142 | 289.8407 | 292.8142 | 262.9468 | 361.6106 | 139.885 | 89.0885 | 89.1328 | 139.885 | 151.1063 | 134.2743 | 116.115 | 153.115 | 128.5044 | 144.062 | 138.3363 | 130.6814 | 79.4336 | 88.6549 | 98.177 | 99.8672 | 118.9735 | 116.6106 | 115.5486 | 86.7257 | 295.1505 | 266.3451 | 240.0265 | 305.009 | 195.9115 | 230 | 234.3894 | 289.115 | 296.8937 | 246.0973 | 227.115 | 321.0443
884 | 11.7184 | 178.323 | 218.706 | 232.236 | 219 | 182.381 | 197.441 | 201.12 | 203.501 | 193.116 | 159.501 | 125.648 | 187.203 | 268.675 | 147.762 | 168.47 | 266.499 | 233.383 | 181.795 | 166.207 | 240 | 194.058 | 346.325 | 319.499 | 201.47 | 271.53 | 251.764 | 231.646 | 316.118 | 139.47 | 145.205 | 163.708 | 146.822 | 178.176 | 190.677 | 177.443 | 94.969 | 301.47 | 291.443 | 287.236 | 293.147 | 167.646 | 167.147 | 138.118 | 152.884 | 89.853 | 91.648 | 94.735 | 130.588 | 260.913 | 209.234 | 246.178 | 168.205 | 165.236 | 162.53 | 167.089 | 152.679 | 184.528 | 205.441 | 226.706 | 195.53 | 154.94 | 135.617 | 191.472 | 152.499 | 230.236 | 203.354 | 257.414 | 247.263 | 191.677 | 179.735 | 261.942 | 220.793 | 259.414 | 239.942 | 210.441 | 316.735 | 201.528 | 185.029 | 186.588 | 249.292 | 161.294 | 177.441 | 150.998 | 193.325 | 222.207 | 201.294 | 202.677 | 144.147 | 151.766 | 124.234 | 136 | 141.675 | 264.588 | 276.501 | 278.911 | 346.234 | 272.087 | 285.764 | 270.619 | 354.617 | 137.588 | 91 | 86.029 | 137.588 | 145.058 | 121.058 | 117.559 | 137.499 | 123.294 | 136.205 | 115.557 | 157.443 | 78.853 | 75.793 | 78.116 | 98.706 | 87.027 | 98.528 | 101.911 | 93.971 | 332.708 | 250.205 | 240.412 | 286.998 | 186.323 | 235.971 | 219.205 | 265.822 | 305.501 | 243.559 | 235.383 | 315.176
885 | 11.7294 | 195.1711 | 219 | 248.71 | 204.29 | 181.5959 | 183.1553 | 194.8859 | 173.1189 | 172.5595 | 147.29 | 140.9794 | 171.3471 | 265.8653 | 189.4563 | 172.1918 | 268.3265 | 226.3471 | 178.8082 | 170.7306 | 245.1918 | 204.9794 | 319.7148 | 318.7306 | 207.7876 | 266.0777 | 243.0777 | 243.71 | 343.824 | 152.71 | 145.5959 | 146.5024 | 160.3058 | 188.2488 | 170.3677 | 171.6165 | 115.2282 | 287.0206 | 305.5182 | 290.7306 | 291.2694 | 159.8082 | 184.3058 | 158.0364 | 157.8653 | 93.4612 | 82.7512 | 94.8653 | 136.9223 | 235.4454 | 213.7876 | 232.6942 | 171.1918 | 156.6165 | 167.4612 | 177.6529 | 155.5388 | 197.4405 | 190.29 | 224.4041 | 187.4818 | 185.7464 | 163.4199 | 202.7876 | 151.7306 | 219.8859 | 205.1347 | 243.9636 | 268.0935 | 182.6165 | 166.0206 | 251.0206 | 223.1918 | 249.4247 | 237.6735 | 186.6371 | 355.8034 | 198.8653 | 209.9588 | 189.4612 | 227.5595 | 167.9223 | 172.6735 | 170.2282 | 172.7718 | 206.8289 | 214.8447 | 204 | 131.8859 | 148.0777 | 137.4405 | 122.1553 | 150.1141 | 270.9223 | 267.7512 | 273.5388 | 354.2488 | 284.5753 | 291.7876 | 285.2488 | 376.3629 | 156.9017 | 71.0983 | 88.4612 | 156.9017 | 129.1553 | 124.1918 | 129.2488 | 141.9223 | 129.9223 | 124.4818 | 137.824 | 129.1189 | 70.3471 | 87.71 | 79.1918 | 90.3471 | 95.71 | 99.3265 | 99.1347 | 89.8082 | 315.5024 | 224.6371 | 239.2694 | 301.0364 | 189.3265 | 226.6165 | 216.1347 | 234.3107 | 294.1553 | 238.8082 | 234.4041 | 316.5959
886 | 11.7404 | 176.938 | 235.6741 | 218.5294 | 194.1018 | 189.8982 | 183.6327 | 165.7949 | 187.3068 | 175.2655 | 149.3879 | 124.5708 | 172.6327 | 247.5708 | 184.7138 | 167.7345 | 247.938 | 230.2655 | 177.1224 | 163.9793 | 235.469 | 215.7758 | 329.0413 | 304.9587 | 217.7758 | 262.3673 | 238.4897 | 238.9793 | 319.9174 | 156.7552 | 137.2242 | 141.2448 | 170.0207 | 191.7552 | 173.1431 | 165.6121 | 104.9587 | 320.9809 | 291.2035 | 269.938 | 286.6121 | 152.8569 | 155.407 | 168.0207 | 144.8363 | 85.2242 | 82.7552 | 99.3879 | 136.2448 | 268.7361 | 224.6534 | 256.3275 | 187.7965 | 153.2448 | 159.2242 | 157.938 | 126.0398 | 206.1431 | 182.7345 | 220.4897 | 202.6741 | 158.5294 | 154.8363 | 184.6932 | 147.6121 | 217.1224 | 226.9396 | 216.4277 | 223.733 | 194.1637 | 182.4292 | 237.5914 | 254.7154 | 246.1224 | 249.2862 | 212.8378 | 362.8776 | 179.6932 | 201.7138 | 190.8776 | 238.1637 | 165.4897 | 160.5914 | 136.2639 | 182.1637 | 211.8982 | 192.4277 | 210.1431 | 143.1637 | 150.5103 | 136.4897 | 130.531 | 173.062 | 279.8982 | 272.1431 | 244.0398 | 341.9587 | 281.7345 | 310.5517 | 285.2448 | 345.7743 | 158.2448 | 68.8776 | 85.4897 | 158.2448 | 121.7345 | 131.1431 | 123.1018 | 153.531 | 120.469 | 137.9189 | 139.3673 | 124 | 90.062 | 91.7552 | 69.469 | 99.531 | 94.4897 | 123.6947 | 107.7758 | 88.1224 | 318.1431 | 234.1637 | 281.124 | 302.2448 | 213.6947 | 212.7138 | 205.469 | 268.6136 | 292.8776 | 218.6932 | 230.4897 | 337.1844
887 | 11.7514 | 166.881 | 236.2203 | 258.4937 | 193.8899 | 198.119 | 192.8987 | 185.1367 | 174.8721 | 193.7975 | 128.643 | 126.4494 | 175.6696 | 236.1013 | 185.6696 | 176.7886 | 270.8063 | 203.4139 | 177 | 156.7709 | 244.6785 | 199.2025 | 311.4228 | 310.119 | 178.0658 | 270.8987 | 240.6696 | 232.6608 | 350.7051 | 160.5595 | 138.6696 | 144.5595 | 145.1937 | 185.7709 | 161.5418 | 163.2203 | 119.0177 | 303.7532 | 277.4316 | 303.4848 | 312.6962 | 133.3127 | 174.1367 | 164.5506 | 144.7797 | 103.5772 | 80.3304 | 113.3481 | 119.9823 | 233.9557 | 234.0089 | 242.2025 | 167.7532 | 170.7975 | 149.9911 | 166.5684 | 143.357 | 189.2025 | 200.6873 | 217.3304 | 188.0924 | 178.0266 | 153 | 210.476 | 153.2291 | 211.6608 | 190.8456 | 237.9165 | 255.2646 | 188.881 | 179.6608 | 267.1456 | 231.4139 | 244.2203 | 241.2114 | 200.9823 | 338.0835 | 201.0266 | 196.4405 | 179.4316 | 264.9165 | 154.3215 | 164.3392 | 167.4848 | 170.6519 | 210.3304 | 189 | 200.3215 | 145 | 130.5329 | 133.3304 | 129.3304 | 136.8456 | 297.9076 | 282.7886 | 246.2291 | 364.0266 | 240.9557 | 288.9734 | 274.3215 | 363.2468 | 141.9823 | 86.7975 | 80.5506 | 141.9823 | 123.6696 | 119.5418 | 138.9076 | 132.7532 | 119.8899 | 117.7532 | 117.643 | 132.0089 | 94.7797 | 101.7886 | 92.9165 | 89.4316 | 116.2468 | 116.3215 | 105.4405 | 86.2203 | 282.5152 | 276.0443 | 237.1671 | 303.7797 | 209.881 | 214.5595 | 197.7709 | 270.4405 | 278.762 | 216.8899 | 259.3658 | 311.524
888 | 11.7624 | 189.4563 | 226.9784 | 237.8372 | 198.5108 | 192.6848 | 172.348 | 186.1957 | 228.9342 | 196 | 144.0433 | 142.3368 | 185.0216 | 260.2606 | 208.5541 | 160.8589 | 227.0874 | 224.3585 | 168.8805 | 156 | 222.5437 | 198.8043 | 322.5325 | 313.7065 | 185.6303 | 233.2069 | 245.5108 | 262.6736 | 302.6744 | 161.9022 | 129.0762 | 141.3913 | 158.239 | 172.3697 | 159.0978 | 149.4675 | 119.1957 | 282.0545 | 290.4346 | 289.9567 | 304.2719 | 153.5541 | 185.1195 | 141.4459 | 127.8589 | 90.6632 | 79.0978 | 81.6199 | 136.0433 | 258.7714 | 246.7281 | 276.9888 | 182.1411 | 173 | 163.4346 | 187.8476 | 144.1957 | 195.1195 | 172.3264 | 198.9567 | 203.1411 | 176.4892 | 175.5541 | 188.7394 | 150.3913 | 213.7065 | 201.3368 | 216.6415 | 262.7065 | 182.587 | 175.3913 | 245.7394 | 209.9567 | 239.4892 | 229.174 | 220.652 | 337.7065 | 186.8589 | 188.7827 | 196.9455 | 239.1307 | 162.0216 | 175.826 | 156.6632 | 162.6848 | 184.7394 | 199.826 | 194.4892 | 140.4892 | 145.1411 | 122.174 | 159.6736 | 140.1195 | 293.6848 | 270.4675 | 298.4234 | 382.3368 | 260.3585 | 315.7714 | 286.5325 | 359.6848 | 165.2606 | 79.0762 | 89.9238 | 165.2606 | 120.3913 | 113.4892 | 136.4892 | 131.8043 | 103.761 | 126.7281 | 114.0978 | 117.6632 | 93.1957 | 79.5437 | 84.2719 | 96.1195 | 93.7394 | 118.6087 | 98.6848 | 85.0978 | 278.9022 | 292.7281 | 258.0649 | 306.7065 | 214.413 | 210.4892 | 215.0433 | 271.8043 | 292.3368 | 239.5541 | 226.0112 | 327.8476
889 | 11.7734 | 219.4337 | 256.177 | 217.6253 | 200.8289 | 192 | 180.059 | 216.177 | 167.3303 | 154.8495 | 126.7964 | 126.6253 | 176.8554 | 247.4542 | 183.5663 | 163.5723 | 289.6697 | 192.2507 | 189.9469 | 162.4012 | 211.7699 | 205.4012 | 308.4542 | 312.1711 | 189.7434 | 235.4012 | 245.0855 | 246.7964 | 341.8083 | 145.5398 | 148.118 | 117.2241 | 169.1446 | 165.5133 | 169.9735 | 168.118 | 97.0531 | 296.4602 | 291.0855 | 270.6253 | 335.006 | 171.5458 | 170.4542 | 148.1446 | 152.5192 | 113.6903 | 85.4012 | 89.8879 | 133.4277 | 271.1446 | 226.9675 | 282.8289 | 176.6844 | 154.7109 | 152.1976 | 141.5338 | 133.941 | 214.2891 | 204.6638 | 225.3481 | 186.7109 | 160.4542 | 177.0855 | 179.5988 | 150.9145 | 194.7964 | 191.1121 | 224.059 | 225.5073 | 173.7699 | 189.6313 | 268.6048 | 241.8349 | 230.7699 | 237.1446 | 222.0855 | 334.3422 | 185.9145 | 186.1711 | 179.7964 | 245.1446 | 158.4277 | 168.7699 | 157.7434 | 159.2566 | 193.8879 | 217.4602 | 188.5133 | 152.8024 | 126.882 | 140.2036 | 117.2772 | 138.2566 | 298.4867 | 258.0265 | 251.876 | 327.3037 | 272.1446 | 290.6519 | 294.4012 | 326.0796 | 156.1121 | 79.8289 | 98.3156 | 156.1121 | 126.4012 | 122.1446 | 128.6844 | 158.5192 | 120.2891 | 106.0531 | 118.5723 | 127.8879 | 86.5988 | 87.059 | 92.1446 | 101.5723 | 99.2301 | 122.6578 | 101.6578 | 89.5723 | 279 | 224.5014 | 224.4218 | 317.059 | 203.1121 | 216.4012 | 224.3156 | 238.1651 | 301.3156 | 255.7168 | 232.059 | 309.882
890 | 11.7844 | 193.2708 | 235.8313 | 200.2453 | 252.8978 | 182.7325 | 192.121 | 191.1976 | 164.707 | 173.2419 | 151.8757 | 127.7802 | 183.414 | 237.6593 | 184.707 | 176.0477 | 220.0068 | 205.6814 | 187.3663 | 156.5128 | 218.414 | 201.3663 | 338.5094 | 325.9012 | 208.5349 | 276.7769 | 255.1942 | 229.2453 | 338.586 | 156.0477 | 148.1465 | 135.3884 | 151.4651 | 185.3884 | 176.5605 | 147.7581 | 109.828 | 299.8535 | 297.4872 | 270.8535 | 312.9778 | 168.3663 | 190.3152 | 166.6082 | 139.2453 | 102.0988 | 74.879 | 90.0733 | 133.9267 | 241.4174 | 263.9234 | 232.9557 | 175.0733 | 172.4617 | 161.1942 | 163.8757 | 145.9745 | 163.1754 | 190.3918 | 226.1465 | 207.2419 | 172.9012 | 138.0766 | 171.586 | 159.3407 | 205.9745 | 213.1687 | 209.2453 | 246.0954 | 175.7802 | 166.9046 | 279.3407 | 213.4906 | 242.9745 | 227.8058 | 244.2419 | 308.0511 | 214.7292 | 192.4872 | 198.3884 | 238.586 | 191.3629 | 169.8535 | 159.8535 | 169.1942 | 204.2675 | 212.5128 | 215.8024 | 151.2198 | 119.4395 | 137.3663 | 139.8757 | 134.293 | 280.4651 | 244.9523 | 259.9745 | 350.7292 | 268.3663 | 307.4617 | 270.9046 | 323 | 131.8313 | 69.8058 | 84.172 | 131.8313 | 110.3186 | 125.7802 | 127.0733 | 122.0766 | 115.5128 | 103.0733 | 130.121 | 151.2419 | 86 | 100.9745 | 74.4651 | 109.414 | 103.707 | 130.414 | 114.9745 | 91.8535 | 293.828 | 267.1176 | 221.9267 | 290.1976 | 202 | 232.7547 | 221.293 | 295.2386 | 280.6848 | 238.4651 | 234.8535 | 317.2675
891 | 11.7954 | 207.9027 | 221.7925 | 228.1102 | 234.4631 | 191.3904 | 190.1829 | 163.646 | 153.7315 | 187.2075 | 144.6096 | 155.2321 | 169.9144 | 260.4759 | 195.3294 | 173.2438 | 232.7808 | 225.7808 | 181.3658 | 165.3904 | 219.939 | 196.3048 | 339.1219 | 304.4631 | 184.646 | 278.1219 | 240.0364 | 231.7562 | 298.5604 | 147.6096 | 134.8535 | 123.8535 | 166.9027 | 220.8054 | 181.6952 | 153.5123 | 119.4513 | 294.3658 | 293.3048 | 271 | 264.9871 | 161.4267 | 137.5358 | 164.2438 | 137.061 | 90.6706 | 99.354 | 79.6706 | 143.3904 | 235.2438 | 259.4877 | 244.9636 | 197.5369 | 150.5241 | 160.1219 | 140.646 | 148.8781 | 189.9883 | 193.6952 | 226.939 | 210.8781 | 191.8417 | 162.2321 | 165.3658 | 150.6096 | 213.5733 | 188.7069 | 230.5369 | 273.354 | 191.9636 | 188.4759 | 244.3166 | 241.9883 | 236.4877 | 206.3412 | 246.939 | 309.7562 | 183.1946 | 181.7315 | 180.2802 | 238.939 | 157.3775 | 157.7925 | 144.0364 | 158.7315 | 189.0364 | 206.3658 | 194.5241 | 143.4877 | 129.3294 | 132.3048 | 132.6096 | 144.3294 | 271.4877 | 287.1957 | 250.6706 | 374.5979 | 259.5487 | 286.4631 | 283.0856 | 341.7808 | 132.8171 | 65.2438 | 81.1219 | 132.8171 | 109.939 | 128.8171 | 120.4267 | 135.9027 | 108.4267 | 127.415 | 110.3412 | 116.3775 | 96.3294 | 84.1583 | 88.0246 | 111.8781 | 114.3294 | 111.2802 | 116 | 83.5487 | 255.5604 | 240.0117 | 267.0738 | 319.9273 | 186.0364 | 229.3048 | 219.1219 | 283.0973 | 294.0246 | 221.0364 | 222.7925 | 288.8898
892 | 11.8064 | 207.0973 | 232.4159 | 236.6593 | 235.854 | 194.854 | 194.7566 | 194.3452 | 183.4425 | 164.2168 | 144 | 156.0487 | 187.0753 | 269.6106 | 163.6548 | 183.4646 | 275.8585 | 206.0708 | 168.6327 | 168.854 | 205.7301 | 207.4159 | 380.8585 | 287.7787 | 188.708 | 296.0753 | 239.9513 | 225.3407 | 290.292 | 156.5133 | 144.4646 | 117.292 | 182.2699 | 198.2654 | 169.6327 | 158.7566 | 104.7787 | 279.7301 | 288.2434 | 281.4646 | 311.4691 | 176.2213 | 153.9779 | 153.5354 | 148.4159 | 106.1726 | 84.8274 | 79.9513 | 124.0221 | 239.7566 | 234.2654 | 264.0753 | 178.0708 | 164.2213 | 159.0487 | 153.2699 | 145.1947 | 187.2434 | 205.4159 | 219.3894 | 186.2654 | 173.9734 | 143.0708 | 172.6106 | 137.6327 | 186.4115 | 211.7346 | 237.708 | 256.9247 | 166.3628 | 167.1681 | 223.9247 | 232.5841 | 225.5354 | 224.9779 | 200.3849 | 338.5399 | 182.9027 | 162.9247 | 168.5354 | 297.9824 | 167.3673 | 160.8053 | 149.6593 | 195.1018 | 207.9779 | 192.6814 | 200.6106 | 150.6106 | 112.8761 | 150.0753 | 129.146 | 156.4159 | 266.2434 | 290 | 282.3452 | 379.8053 | 256.146 | 279.292 | 285.9027 | 324.9247 | 143.4646 | 88.7832 | 75.292 | 143.4646 | 118.562 | 131.854 | 110.4867 | 149.3673 | 126.0753 | 134.708 | 112.8053 | 127.3186 | 88.438 | 70.6327 | 89.9513 | 130.0753 | 112.146 | 125.2213 | 115.0487 | 82.0487 | 289.1505 | 233.2434 | 247.1681 | 342.9292 | 211.6372 | 227.0973 | 231.3673 | 267.7301 | 321.6372 | 227.6106 | 206.7787 | 300.3186
893 | 11.8174 | 191.5821 | 263.8358 | 217.7276 | 196.4916 | 196.9272 | 191.1455 | 146.8554 | 195.5998 | 169.7453 | 129.5457 | 132.8731 | 161.0186 | 259.4002 | 173.5634 | 181.1091 | 279.9272 | 191.5093 | 168.9636 | 178.6362 | 207.8909 | 178.1278 | 313.6194 | 269.6548 | 188.0364 | 270.0186 | 223.6185 | 207.6548 | 293.8545 | 158.9272 | 143.0728 | 148.7995 | 173.3638 | 178.6912 | 181.5271 | 174.4179 | 114.5998 | 274.1819 | 248.4916 | 278.1455 | 294.7276 | 173.1455 | 161.7453 | 138.5457 | 160.5634 | 102.1819 | 89.7817 | 103.1269 | 118.1819 | 232.291 | 241.6726 | 264.0364 | 194.3452 | 161.1455 | 157.0728 | 162.6726 | 146.9272 | 174.4729 | 204.0728 | 219.9636 | 185.9636 | 172.0364 | 151.6362 | 157.5821 | 144.709 | 190.7817 | 193.7276 | 246.6726 | 229.9823 | 190.0541 | 169.8545 | 242.2724 | 263.7995 | 245.236 | 223.1091 | 233.6539 | 312.055 | 195.5271 | 183.1996 | 202.6903 | 240.2919 | 167.0364 | 143.6548 | 162.5271 | 180.6185 | 234.0541 | 198.7453 | 180.764 | 153.8909 | 106.2183 | 115.3461 | 116.4729 | 132.9095 | 258.291 | 288.0728 | 243.528 | 308.6921 | 264.6726 | 265.5093 | 245.528 | 323.0364 | 125.6912 | 79.4002 | 102.945 | 125.6912 | 129.5998 | 129.1091 | 119.6362 | 153.8545 | 125.0728 | 99.3461 | 124.5634 | 134.7453 | 91.8545 | 98.9086 | 94.8181 | 101.1278 | 114.8909 | 113.4729 | 116.9272 | 96.4543 | 304.4907 | 245.5271 | 269.1269 | 334.3638 | 205.291 | 234.709 | 206.9459 | 262.1819 | 338.4179 | 259.7995 | 212.7453 | 320.2724
894 | 11.8284 | 192.9518 | 237.6745 | 244.3255 | 203.7832 | 203.8314 | 212.47 | 185.9883 | 206.735 | 170 | 132.9036 | 132 | 201.9642 | 268.7591 | 156.4336 | 179.0482 | 246.819 | 195.8796 | 194.3737 | 141.9154 | 217.7591 | 201.3978 | 325.6387 | 287.5423 | 221.181 | 273.8796 | 263.9883 | 254.8196 | 331.0846 | 155.0964 | 151.7832 | 129.5059 | 173 | 170.1927 | 170.2891 | 147.6745 | 109.1445 | 327.6751 | 260.6628 | 266.2891 | 331.0846 | 161.2891 | 156.1445 | 160.446 | 146.3613 | 91.265 | 91.9518 | 68.8672 | 154.1087 | 246.6387 | 236.1445 | 261.0723 | 175.4818 | 161 | 158.9518 | 158.1204 | 119.6745 | 200.3496 | 195.2168 | 248.3014 | 187.9518 | 137.8431 | 141.265 | 189.2051 | 155.735 | 225.1569 | 215.446 | 223.5781 | 267.0605 | 190.0241 | 165.1204 | 249.8314 | 212.3008 | 250.8796 | 259.1087 | 191.084 | 327.5905 | 176.4818 | 200.5905 | 203.0241 | 233.1204 | 163.0964 | 153.735 | 155.1927 | 171.2168 | 208.6504 | 194.1204 | 205.3737 | 176.446 | 131.3737 | 140.3496 | 134.5423 | 160.3014 | 281.4219 | 241.1563 | 270.3014 | 349.916 | 284.5182 | 263.0482 | 244.9759 | 348.3737 | 142.5664 | 79 | 82.53 | 142.5664 | 126.0964 | 121.1927 | 102.4336 | 139.3613 | 122.0723 | 131.181 | 116.2168 | 115.4818 | 71.5059 | 107.8073 | 82.3132 | 101.9518 | 103.2891 | 94.4577 | 98.4577 | 78.4577 | 282.554 | 241.1204 | 281.7109 | 282.2767 | 208.9036 | 226.2168 | 228.446 | 287.3737 | 309.7227 | 221.9636 | 219.8314 | 334.6628
895 | 11.8394 | 193.9882 | 260.7169 | 227.2124 | 231.6697 | 194.118 | 185.3303 | 171.1888 | 194.1534 | 159.1298 | 136.9528 | 132 | 177.3067 | 250.2242 | 209.3629 | 163.1888 | 263.7876 | 192.0472 | 179.1888 | 156.8112 | 203.177 | 185.2006 | 332.9174 | 295.9056 | 195.3185 | 286.8466 | 215.5899 | 273.7876 | 308.2831 | 129.3067 | 152 | 136.9056 | 174.9764 | 175.9292 | 175.9292 | 153.9174 | 108.0118 | 282.5545 | 287.6815 | 284.7758 | 316.1888 | 152.1062 | 174.7758 | 148.1534 | 137.1062 | 103.8466 | 94.9646 | 81.8348 | 136.2242 | 221.3067 | 219.2006 | 270.882 | 190.8112 | 171.8702 | 141.2124 | 199.5045 | 156.5517 | 194.0826 | 186.1062 | 232.2006 | 184.0472 | 182.4573 | 141 | 190.9882 | 143.1534 | 224.0236 | 227.8584 | 226.9528 | 217.6017 | 176.1652 | 154.1298 | 233.2006 | 221.8702 | 236.177 | 223.4365 | 210.7522 | 326.0236 | 206.6343 | 199.0236 | 169.4011 | 266.5989 | 145.2124 | 154 | 147.0944 | 142.3421 | 202.0708 | 169.2949 | 152.6371 | 145.3775 | 120.1416 | 130.1298 | 119.1888 | 132.3421 | 265.2006 | 260.7522 | 241.3539 | 355.941 | 266.2242 | 248.177 | 294.4101 | 329.236 | 156.8348 | 88.882 | 81.0118 | 156.8348 | 107.2242 | 141.7522 | 80.2596 | 139.9882 | 127.9292 | 100.3775 | 96.236 | 156.5045 | 71 | 94.1652 | 79.0354 | 99.0354 | 117.823 | 120.6815 | 96.0236 | 95.7876 | 279.0354 | 220.2478 | 246.4247 | 317.5635 | 174.4129 | 213.1534 | 231.9646 | 285.0354 | 304.059 | 222.9764 | 238.7758 | 324.1298
896 | 11.8504 | 173.0064 | 205.0183 | 235.9995 | 234.9867 | 170.0079 | 182.0025 | 175.9995 | 186.9975 | 165.0074 | 151.9872 | 137.9906 | 155.0035 | 223.0118 | 196.9882 | 156.0128 | 305.9862 | 207.0163 | 187.9946 | 149.0005 | 225.9946 | 160.998 | 328.0118 | 297.0202 | 224.9827 | 209.0291 | 291.9783 | 222.0118 | 334.9808 | 157.9862 | 149.9995 | 126.0064 | 148.002 | 192.9975 | 158 | 172.9911 | 117.9911 | 288.9975 | 283.0044 | 275.9941 | 309.9882 | 180.0035 | 161.0039 | 138.0084 | 151.998 | 96.9921 | 95.9921 | 81.002 | 141.0039 | 250.0039 | 230.0039 | 238.0104 | 171.9901 | 162.0089 | 156.9995 | 161.9916 | 147.0035 | 184.9985 | 181.001 | 213.9921 | 174.0089 | 157.003 | 161.9995 | 196.9995 | 141.0035 | 212.0113 | 195.9975 | 209.0054 | 272.9882 | 175.0015 | 169.9946 | 265.9911 | 238.9813 | 227.9965 | 195.0039 | 206 | 300.0217 | 201.9956 | 190.9936 | 199.9837 | 209.0178 | 151.997 | 162.998 | 150.997 | 182.9857 | 203.002 | 188.0158 | 213.9842 | 156.997 | 134.9891 | 120.0148 | 129.0025 | 126.0044 | 295.001 | 265.002 | 262.002 | 359.999 | 270.0094 | 270.9798 | 261.9882 | 356.9965 | 164.9798 | 70.0079 | 83.9946 | 164.9798 | 120.9995 | 126.002 | 128.9882 | 139.9956 | 130.0015 | 123.9921 | 111.9916 | 128.0015 | 87.9951 | 73.0059 | 84.997 | 99.0054 | 116.0015 | 117.9837 | 113.002 | 82.0059 | 296.9916 | 230.001 | 225.0104 | 266.0197 | 166.0015 | 240.0005 | 235.9985 | 303.9798 | 312.9788 | 227.0049 | 238.0094 | 333.9956
897 | 11.8614 | 185.7443 | 241.7571 | 234.6804 | 208.473 | 185.9489 | 186.8466 | 175.1151 | 182.1406 | 179.5781 | 125.821 | 119.0511 | 162.179 | 246.6676 | 173 | 182.0256 | 277.9361 | 239.6804 | 176.6676 | 150.1278 | 214.8977 | 157.2429 | 351.4631 | 337.5909 | 190.2301 | 268 | 247.8849 | 245.3991 | 296.1918 | 130.0767 | 148.8722 | 139.0128 | 151.8594 | 188 | 157.9105 | 155 | 100.2173 | 284.3963 | 291.7315 | 264.2045 | 286.2429 | 186.5142 | 168.7571 | 154.821 | 147.9616 | 81.2173 | 80.1534 | 85.0511 | 148.706 | 257.821 | 237.8594 | 258.5525 | 152.2429 | 179.7315 | 156.179 | 145.2557 | 153.8594 | 182.1534 | 183.1662 | 198.3835 | 192.0895 | 162.8082 | 160.9233 | 196 | 147.8722 | 234.4886 | 190.9361 | 220.179 | 248.6676 | 178.0511 | 159.3963 | 248.0128 | 201.4602 | 220.8722 | 203.0384 | 206.0639 | 343.6548 | 192.8849 | 178 | 167.3324 | 244.9105 | 146.2173 | 158.9489 | 145.2813 | 154.2685 | 206.821 | 219.5653 | 182.2685 | 150.7315 | 112.9744 | 149.9872 | 133.8082 | 135.1406 | 296.8722 | 269.3707 | 265.6165 | 358.1918 | 288.3096 | 230.4858 | 238.3707 | 349.3991 | 124.4986 | 85.9872 | 73.3068 | 124.4986 | 120.1023 | 130.0128 | 105.2557 | 130.8466 | 132.7443 | 108.2173 | 95.2045 | 131.0511 | 77.8722 | 85.179 | 79.0767 | 109.8338 | 118.6804 | 85.4347 | 116.8338 | 93.8977 | 279.8338 | 232.5625 | 246.0256 | 306.294 | 169.5114 | 240.3991 | 232.8977 | 263.6264 | 270.0511 | 236.8338 | 256.5398 | 324.821
898 | 11.8724 | 166.7272 | 223.0752 | 210.4514 | 243.9468 | 182.2006 | 175.5767 | 184.2508 | 193.2508 | 147.3511 | 113.2538 | 122.6991 | 175.5236 | 221.5015 | 172.8997 | 184.1254 | 271.9218 | 215.0502 | 151.2758 | 160.3761 | 207 | 176.4012 | 310.8776 | 306.5266 | 207.7743 | 267.8997 | 239.8526 | 199.9529 | 310.3982 | 135.7994 | 139.1254 | 139.8997 | 141.326 | 187.5988 | 151.1003 | 155.3009 | 116.8245 | 314.3982 | 271.0752 | 279.2227 | 305.0752 | 149.5266 | 150.1254 | 141.1755 | 144.4734 | 97.6991 | 91.9498 | 88.7492 | 126.4514 | 243.6239 | 227.0251 | 223.8495 | 170.9498 | 159.5015 | 169.674 | 164.9498 | 143.2006 | 193.9498 | 195.8495 | 227.8746 | 199.2006 | 148.3761 | 155.2758 | 195.7994 | 138.2006 | 195.5266 | 186.4263 | 233.7994 | 224.329 | 181.8495 | 189.5486 | 249.2758 | 237.0502 | 211.8024 | 205.8746 | 210.8245 | 317.0251 | 183.5486 | 178.0752 | 192.5236 | 238.9779 | 163.0502 | 155.1254 | 167.0752 | 175.7021 | 192.9749 | 186.9529 | 202.5988 | 129.7492 | 110.7743 | 148.7492 | 119.5517 | 146.1003 | 286.323 | 296.6459 | 237.1785 | 372.7242 | 235.5767 | 267.7492 | 267.4514 | 303.4764 | 162.6239 | 84.8997 | 96.323 | 162.6239 | 127.7242 | 131.1755 | 125.2508 | 119.7523 | 113.1755 | 125.0251 | 110.6991 | 135.0502 | 68.7773 | 98.7242 | 84.7743 | 96.7242 | 94.326 | 118.7242 | 104.326 | 85.7994 | 267.2758 | 275.1474 | 247.5236 | 329 | 208.7994 | 194.5266 | 225.4263 | 310.8716 | 273.7242 | 224.9027 | 220.8746 | 311.4263
899 | 11.8834 | 194.9626 | 226.6352 | 227.1779 | 203.2616 | 190.1495 | 197.7758 | 193.1032 | 201.169 | 161.411 | 161.4021 | 111.7847 | 157.6352 | 240.6637 | 170.0836 | 188.5516 | 231.3078 | 217.5605 | 162.0374 | 174.3274 | 206.589 | 192.71 | 343.9537 | 325.9537 | 198.4395 | 264.1121 | 271.468 | 236.7758 | 286.8132 | 128.1495 | 142.6548 | 135.6637 | 153.9253 | 172.8221 | 154.9626 | 166.9626 | 110.0747 | 290.4021 | 273.6263 | 249.5231 | 308.1868 | 169.1032 | 154.8879 | 148.0747 | 124.7847 | 86.1495 | 89.3648 | 79.1868 | 144.3363 | 229.2989 | 227.8505 | 218.6352 | 169.0747 | 178.8132 | 157.71 | 163.2242 | 150.8505 | 192.3363 | 190.0374 | 222.8879 | 205.7295 | 162.589 | 166.8221 | 187.3274 | 145.3648 | 216.4858 | 203.4484 | 225.7758 | 274.2811 | 176.6726 | 171.7384 | 258.3932 | 239.3363 | 241.6922 | 202.1957 | 203.3274 | 318.411 | 166.1868 | 180.9253 | 174.7847 | 276.2527 | 165.0747 | 159.5516 | 170.4484 | 200.9822 | 191.4769 | 223.5516 | 187.5605 | 121.1584 | 104.5036 | 139 | 140.1779 | 149.1032 | 260.7847 | 245.5694 | 282.2527 | 360.4306 | 259.121 | 257.5142 | 285.1495 | 322.6726 | 148.5979 | 80.7384 | 70.3737 | 148.5979 | 117.0374 | 137.4769 | 134.3648 | 148.9253 | 119.29 | 126.411 | 99.6352 | 136.2153 | 98.3274 | 87.4769 | 76.5605 | 87.8683 | 106.9253 | 108.1121 | 115.9911 | 78.71 | 277.9253 | 242.2989 | 230.0463 | 328.3274 | 201.8221 | 215.5979 | 240.8416 | 266.6263 | 264.4573 | 258.9537 | 217.0089 | 328.5979
900 | 11.8944 | 193.3545 | 241.9572 | 206.7945 | 210.149 | 193.5531 | 191.851 | 170.4966 | 154.7449 | 170.262 | 145.4538 | 132.8442 | 173.8014 | 233.6884 | 196.5103 | 175.9075 | 264.6524 | 231.7517 | 162.851 | 157.7945 | 196.6455 | 209.9075 | 316.6027 | 299.0993 | 183.851 | 265.8579 | 233.3904 | 231.4469 | 283.3904 | 132.6952 | 110.5325 | 128.9863 | 151.851 | 192.6096 | 154.3973 | 164.9075 | 111.4538 | 275.9931 | 262.8082 | 262.851 | 312.1558 | 145.7021 | 152.7945 | 150.6952 | 145.6455 | 90 | 73.6952 | 83.7021 | 152.6524 | 236.9503 | 225.0925 | 235.2483 | 171.8938 | 174.6455 | 175.3048 | 168.9007 | 146.9503 | 200.8014 | 191.3973 | 220.9435 | 174.887 | 152.7945 | 186.262 | 171.4897 | 130.5394 | 227.6592 | 215.6455 | 221.0925 | 231.7877 | 194.2979 | 163.411 | 218.6387 | 246.8579 | 208.0993 | 233.0993 | 188.5822 | 328.0565 | 171.4966 | 179.2483 | 194.1062 | 256.6027 | 166.7021 | 146.9075 | 181.3545 | 151.7312 | 179.4897 | 210.9572 | 201.9503 | 150.3545 | 166.8151 | 138.3048 | 119.2979 | 126.4966 | 279.3116 | 285.6524 | 262.8014 | 322.9298 | 288.5462 | 247.1353 | 289.8938 | 339.4538 | 162.8579 | 75.0428 | 79.9503 | 162.8579 | 118.4966 | 124.8442 | 117.7517 | 146.4041 | 101.4966 | 136.4041 | 114.9075 | 116.5462 | 80.9007 | 74.8442 | 90.9007 | 133.8647 | 105.4469 | 111.2979 | 91.1918 | 96.4041 | 275.3545 | 250.0497 | 255.2123 | 310.5034 | 220.4675 | 229.8082 | 212.1421 | 258.4897 | 302.0993 | 232.4469 | 242.0565 | 343.1062
901 | 11.9054 | 181.6814 | 222.4336 | 222.3097 | 212.0708 | 183.1415 | 189.9292 | 180.3717 | 170.9204 | 139.0443 | 135.6195 | 147.6991 | 171.0531 | 265.5664 | 168.4956 | 155.8053 | 256.823 | 225.1415 | 160.4956 | 173.0619 | 208.6903 | 187.823 | 310.7346 | 302.7965 | 182.7965 | 244.7434 | 258.761 | 241.7346 | 309.3186 | 146.1239 | 157.1415 | 164.7699 | 147.9469 | 166.7434 | 163.1151 | 144.7434 | 100.6283 | 294.7522 | 241.177 | 261.239 | 297.239 | 141.4248 | 167.5044 | 163.1327 | 157.0708 | 90.3717 | 86.6903 | 78.4956 | 145.5044 | 236.1858 | 242.0972 | 240.6195 | 189.1858 | 184.0884 | 162.3097 | 166.0088 | 144.6371 | 198.3009 | 198.0088 | 237.2654 | 208.3362 | 167.5044 | 152.2566 | 197.9557 | 161.239 | 200.6991 | 227.3186 | 241.5044 | 263.7079 | 198.0176 | 136.5311 | 249.3805 | 226.239 | 210.9912 | 234.5044 | 236.2654 | 311.8585 | 181 | 183.8761 | 175.9469 | 248.8142 | 161.2478 | 126.8053 | 168.5664 | 203.8142 | 205.7079 | 191.0619 | 201.7434 | 138.7434 | 125.8673 | 125.6814 | 125.8053 | 136.6814 | 249.2301 | 279 | 259.0619 | 376.8318 | 299.1858 | 286.761 | 304.5221 | 327.6991 | 141.3097 | 94.8142 | 78.3186 | 141.3097 | 126.6991 | 139.9469 | 111.761 | 135.4956 | 111.0619 | 125.1858 | 95.6107 | 127.1858 | 78.9381 | 90.1327 | 88.5664 | 93.9912 | 112.8849 | 117.3717 | 111.9557 | 84.8142 | 264.177 | 251.3717 | 222.5487 | 301.3097 | 173.5487 | 206.7522 | 232.4513 | 287.0619 | 303.0708 | 241.5575 | 225.6107 | 324.4513
902 | 11.9164 | 190.7379 | 226.4016 | 228.2621 | 200.3757 | 155.7424 | 201.4016 | 185.0349 | 198.5152 | 170.297 | 144.9258 | 129.1878 | 186.6288 | 256.4016 | 174.441 | 167.4803 | 239.6682 | 197.8909 | 169.6333 | 173.0349 | 204.6682 | 171.856 | 334.9955 | 327.4016 | 206.6592 | 257.1136 | 238.2925 | 265.6986 | 299.0742 | 147.703 | 128.703 | 132.1878 | 133.3363 | 177.1091 | 177.2531 | 155.4803 | 96.7075 | 288.6986 | 257.9606 | 277.6986 | 315.1834 | 161.2925 | 159.4061 | 149.0349 | 142.703 | 96.6682 | 81.6288 | 84.9606 | 139.0394 | 237.2925 | 187.4544 | 250.5197 | 191.5546 | 140.5939 | 164.7728 | 150.9258 | 126.0045 | 218.5939 | 183.0742 | 214.415 | 170.2272 | 160.2227 | 144.7862 | 166.9258 | 176.0653 | 183.1529 | 216.7773 | 233.0349 | 231.0787 | 169.6333 | 189.1091 | 239.7773 | 242.4759 | 227.1136 | 227.5197 | 211.0394 | 339.0349 | 180.4803 | 182.1485 | 161.6333 | 247.4848 | 165.1485 | 139.5197 | 160.9606 | 200.6288 | 170.4803 | 192.8909 | 212.6288 | 149.5546 | 136.9213 | 135.2576 | 137.703 | 146.7773 | 281.144 | 277.8864 | 260.6682 | 344.0742 | 297.9911 | 268.3712 | 264.4016 | 311.8605 | 145.703 | 91.3318 | 69.4848 | 145.703 | 107.297 | 124 | 94.856 | 143.9651 | 112.2227 | 127.7773 | 119.6288 | 129.6288 | 77.6288 | 77.1485 | 82.5939 | 108.9258 | 96.7424 | 121.9606 | 81.7424 | 81.7773 | 282.0742 | 257.6682 | 247.559 | 303.1047 | 196.7773 | 203 | 208.1834 | 286.8122 | 288.4061 | 249.0349 | 249.4061 | 302.7075
903 | 11.9274 | 177.942 | 194.5192 | 240.1061 | 230 | 166.471 | 170.6441 | 175.8555 | 179.8269 | 174.7788 | 143.2212 | 142.6155 | 182 | 227.6343 | 157.5959 | 162.298 | 249.6441 | 208.8269 | 186.3657 | 162.471 | 212.6539 | 193.3559 | 313.7208 | 293.9616 | 167.1633 | 270.0482 | 221.5861 | 239.8555 | 302.2498 | 143.3943 | 125.3461 | 146.1347 | 148.8751 | 166.4327 | 146.8939 | 152.1151 | 117.9135 | 261.6441 | 244.7404 | 250.9902 | 303.4424 | 142.2498 | 152.1731 | 137.6057 | 140.3845 | 104.7404 | 77.5192 | 72 | 152.6922 | 217.9902 | 240.4906 | 255.8751 | 187.7306 | 148.0865 | 140.7208 | 150.6922 | 149.2694 | 225.3943 | 184.8653 | 256.0482 | 197.5673 | 162.0482 | 190.4424 | 166.6922 | 128.5576 | 208.0098 | 213.8269 | 221.4327 | 256.4808 | 186.9714 | 179.7306 | 239.769 | 212.2114 | 239.702 | 234 | 222.0098 | 326.8269 | 173.6539 | 183.8269 | 181.5673 | 263.5771 | 165.3559 | 145.3078 | 147.6539 | 193.7502 | 167.2016 | 202.4424 | 207.8269 | 144.2596 | 111 | 125.8269 | 132.2694 | 143.4808 | 256.702 | 265.6441 | 267.0098 | 344.7404 | 248.7788 | 274.2114 | 232.8653 | 357.058 | 142.4327 | 83.4327 | 86.1829 | 142.4327 | 110.6539 | 124 | 118.4327 | 154.2694 | 114.0482 | 123.8751 | 116.2114 | 123.0098 | 73.8653 | 78.8269 | 90.9518 | 106.9616 | 105.8269 | 110.298 | 95.0669 | 81.0767 | 282.0482 | 265.3943 | 263.7984 | 269.942 | 191.231 | 203.8653 | 200.6824 | 273.5576 | 280.8269 | 235.4424 | 239.058 | 323.0482
904 | 11.9384 | 205.6401 | 202.4705 | 203.5458 | 228.32 | 179.7271 | 187.6047 | 205.4071 | 178.5929 | 183.2965 | 135.7906 | 127.6047 | 180.6165 | 262.7389 | 182.7271 | 174.913 | 264.233 | 206.4071 | 148.8894 | 174.6401 | 210.087 | 176.4941 | 349.4424 | 283.7906 | 188.913 | 257.9247 | 266.3436 | 264.8613 | 323.5295 | 136.1106 | 128.2094 | 134.3318 | 139.3717 | 170.2094 | 185.9247 | 180.8495 | 116.6047 | 279.8894 | 246.447 | 273.4823 | 289.2729 | 165.3082 | 153.6047 | 144.4941 | 152.0353 | 101.7035 | 82.6047 | 72 | 158.5177 | 239.1976 | 213.2094 | 246.3717 | 205.1106 | 148.8024 | 179.8024 | 155.5295 | 130.2094 | 219.3953 | 192.7153 | 243.0353 | 191.0236 | 153.68 | 172.4189 | 173.2094 | 144.8024 | 188.7788 | 211.8024 | 226.8894 | 250.8024 | 157.2729 | 194.5412 | 266.9247 | 223.6165 | 223.6283 | 234 | 201.1976 | 321.7389 | 173.0635 | 181.0118 | 178.1859 | 235.9247 | 149.087 | 140.5694 | 143.3082 | 170.3953 | 200.2094 | 185.1106 | 206 | 149.4705 | 112.3835 | 125.9764 | 116.2729 | 137.0118 | 243.0988 | 282.5059 | 246.0988 | 341.0118 | 257.3953 | 286.0118 | 245.854 | 329.7552 | 147.6918 | 87.4071 | 67.6918 | 147.6918 | 109.2729 | 124.8894 | 119.9365 | 134.1224 | 107.1623 | 112.4941 | 127.4189 | 106.7434 | 83.0988 | 76.4071 | 100.5059 | 97.1859 | 103.4071 | 120.8377 | 136.1224 | 101.32 | 275.2611 | 260.3835 | 229.9012 | 302.9764 | 162.4941 | 210.8259 | 226.4306 | 287.6283 | 279.1976 | 219.1976 | 207.4071 | 311.5177
905 | 11.9494 | 177 | 227.3334 | 242 | 213.4445 | 159.2222 | 183.3333 | 198.1111 | 185.8889 | 182.1111 | 142.8889 | 124.3333 | 171.5556 | 238.3334 | 162.3333 | 167.2223 | 243.0001 | 203.2223 | 156.1111 | 145.7778 | 219.3333 | 183.2223 | 316.6667 | 293.3334 | 178.8889 | 241.5556 | 233.7778 | 223.0001 | 299.1111 | 129.3334 | 123 | 111.7778 | 156.8888 | 161.5555 | 168.6667 | 146.6667 | 110.2222 | 288.3333 | 282.9999 | 284 | 304.9999 | 160.8889 | 154.0001 | 149.8889 | 128.7778 | 97.8889 | 79.1111 | 71.1111 | 144.3333 | 240.8889 | 205.5555 | 265.4444 | 195.3333 | 146.8889 | 174.6666 | 134.2222 | 123.6667 | 219.7777 | 178.6666 | 217.7778 | 173.2222 | 167.6666 | 157.4444 | 167.1111 | 142.5555 | 203 | 210 | 232 | 250 | 178.4445 | 164.3334 | 247.4444 | 212 | 203.7778 | 231.4444 | 203.5556 | 292.7778 | 196.2222 | 174.4445 | 187.2222 | 220.3334 | 159.2222 | 163.4444 | 141.2223 | 172.4444 | 195.2223 | 179.5556 | 206.5556 | 169.7777 | 126.4445 | 141.4444 | 135.6666 | 131.2223 | 247.3334 | 275.6666 | 248.4445 | 333.5556 | 263.6667 | 274.6666 | 280.4444 | 358.8888 | 154.1111 | 83.2222 | 77.2223 | 154.1111 | 131 | 131.2222 | 95.1112 | 116.3333 | 132.4444 | 115.8889 | 113.6667 | 146.6666 | 84.6667 | 72.3334 | 94.2222 | 109.2222 | 100.7778 | 92.2222 | 117 | 85.8889 | 305.3333 | 268.9999 | 227.8889 | 315.2221 | 169.7778 | 191 | 202.1111 | 266.7778 | 280.8889 | 224 | 203.8889 | 300.1111
906 | 11.9604 | 196.3319 | 244.2723 | 218.7404 | 213.9149 | 160.8766 | 178.4936 | 183.7404 | 196.6809 | 152.8511 | 144.0979 | 126.6298 | 197.5319 | 276.8043 | 162.7787 | 183.2723 | 270.4341 | 216.5575 | 147.2723 | 160.4936 | 212.0255 | 195.0766 | 314.617 | 313.7277 | 180.5915 | 264.0979 | 244.5659 | 258.6809 | 284.4936 | 141.3574 | 138.6298 | 128.9617 | 130.417 | 152.0979 | 182.1234 | 173.2851 | 90.4681 | 290.383 | 253.4681 | 257.7787 | 269.8127 | 174.0255 | 183.5319 | 157.1234 | 155.4213 | 90.234 | 79.2596 | 66.4681 | 138.7532 | 238.6426 | 206.0723 | 246.6042 | 179.9021 | 149.5787 | 152.0723 | 146.4681 | 129 | 191.1617 | 162.5915 | 232.2468 | 177.0979 | 159.5915 | 155.7149 | 174.766 | 140.7277 | 189.0979 | 208.1489 | 210.3447 | 254.2979 | 179.6553 | 175.8638 | 240.9021 | 218.1489 | 217.6298 | 210.1362 | 209.9745 | 301.7149 | 160.8383 | 192.5192 | 173 | 246.2596 | 160.2596 | 160.6042 | 173.1489 | 161.6042 | 209.2979 | 198.6426 | 209.5192 | 150.1489 | 136.6426 | 122.234 | 124.0128 | 154.4085 | 273.7532 | 256.1362 | 258.3958 | 349.4553 | 282.9021 | 258.4681 | 274.766 | 319.234 | 153.5192 | 93 | 99.6681 | 153.5192 | 137.0255 | 116.7532 | 117.9021 | 113.5915 | 106.2213 | 109.7149 | 118.7532 | 120.4936 | 88.6426 | 82.5064 | 83.3319 | 118.0128 | 122.0128 | 94.7404 | 102.7277 | 85.1234 | 303.3107 | 239.7149 | 224.0596 | 272.2085 | 190.3958 | 196.0596 | 198.4553 | 273.4936 | 278.766 | 244.6681 | 224.0596 | 319.3447
907 | 11.9714 | 212.1297 | 234.3068 | 223.7286 | 191.5929 | 161.4926 | 182.4071 | 188.8643 | 166.4071 | 165.7286 | 158.4572 | 122.9144 | 182.7139 | 246.8643 | 149.9853 | 169.1003 | 236.5074 | 187.5782 | 139.885 | 162.7787 | 203.4278 | 158.1564 | 323.885 | 324.6431 | 200.0856 | 273.0294 | 275.3717 | 229.8997 | 289.8997 | 150.1858 | 137.8997 | 149.7286 | 170.8436 | 169.5782 | 180.6932 | 152.6431 | 107.5929 | 283.5575 | 270.7286 | 244.5782 | 299.7935 | 159.3215 | 169.9352 | 155.4218 | 130.6785 | 96.0147 | 74.5428 | 85.764 | 136.0501 | 227.9144 | 237.9853 | 254.4719 | 165.8142 | 174.7286 | 177.879 | 165.3569 | 128.0501 | 178.5989 | 183.5782 | 230.2005 | 188.2005 | 174.2005 | 171.6076 | 167.6283 | 149.2005 | 200.0648 | 196.4926 | 227.4071 | 224.2005 | 164.2213 | 179.1504 | 225.3215 | 206.4926 | 221.242 | 203.3215 | 223.3215 | 318.9646 | 180.1858 | 178.4719 | 174.0856 | 241.5428 | 157.9853 | 169.2861 | 162.8496 | 174.0856 | 184.3569 | 188.4572 | 197.3717 | 133.6076 | 128.8997 | 131.8142 | 118.8997 | 135.0501 | 275.2567 | 251.8997 | 254.4632 | 369.2507 | 265.5575 | 275.7286 | 267.2213 | 327.5929 | 146.6638 | 93 | 76.8142 | 146.6638 | 124.764 | 112.2861 | 107.3422 | 131.7286 | 123.8997 | 128.4572 | 119.0354 | 123.0501 | 80.764 | 78.5929 | 105.5074 | 110.5929 | 115 | 103.5281 | 114.7286 | 84.9144 | 267.4218 | 254.2507 | 255.5221 | 298.2567 | 182.5281 | 227.5221 | 219.4719 | 279.0354 | 270 | 218.4218 | 256.6076 | 328.2654
908 | 11.9824 | 165.9956 | 247.3722 | 223.7758 | 188.704 | 171.444 | 181.5963 | 190.9597 | 168.1121 | 169.1795 | 157.8117 | 121.1795 | 199.704 | 242.1524 | 168.2601 | 159.6637 | 229.0718 | 202.2242 | 172.4799 | 155.5919 | 233.2645 | 189.5201 | 350.4124 | 313.0403 | 206.2601 | 235.296 | 261.1524 | 248.2154 | 301.2601 | 143.2242 | 148.3722 | 146.0762 | 139.9238 | 180.8205 | 166.148 | 146.6637 | 106.0359 | 272.8835 | 270.9238 | 260.556 | 276.8476 | 155.444 | 189.6725 | 143.1436 | 136.0359 | 80.4037 | 80.0718 | 93.1524 | 131.0359 | 222.9238 | 252.1165 | 232.5919 | 167.7443 | 171.2242 | 155.9282 | 169.1165 | 124.8117 | 223.1524 | 196.5963 | 207.3319 | 165.9238 | 151.6278 | 151.1839 | 179.9238 | 131.0674 | 180.7715 | 205.556 | 230.148 | 208.5832 | 171.4081 | 158.4842 | 229.8792 | 219.5517 | 253.0088 | 202.1077 | 218.4081 | 305.2601 | 177.6637 | 161.3275 | 178.3363 | 242.7802 | 173.3004 | 153.3319 | 175.3765 | 182.1839 | 191.8161 | 185.444 | 186.704 | 117.4755 | 139.0762 | 134.9282 | 133.5158 | 128.4081 | 288.7487 | 267.6996 | 298.5963 | 340.9238 | 252.8117 | 272.3722 | 275.148 | 326.4799 | 164.8205 | 92.704 | 81.852 | 164.8205 | 131.7084 | 98.2557 | 129.5245 | 131.1839 | 135.704 | 124.1121 | 129.7802 | 118.4799 | 91.556 | 77.7758 | 95.852 | 105.6322 | 114.556 | 126.296 | 110.9282 | 80.3678 | 256.7715 | 224.148 | 226.852 | 315.8923 | 204.1121 | 204.3275 | 194.3363 | 291.556 | 275.4755 | 201.852 | 231.7443 | 285.6594
909 | 11.9934 | 190.763 | 236.5575 | 233.1986 | 187.8014 | 171.9164 | 164.7247 | 202.0698 | 166.3658 | 192.5889 | 170.9547 | 147.6342 | 192.5506 | 226.2508 | 166.4042 | 170.6725 | 241.1603 | 193.2822 | 179.7178 | 160.2822 | 209.6864 | 182.763 | 327.2055 | 303.8536 | 198.9547 | 240.8466 | 239.4808 | 277.4286 | 296.3589 | 132.8397 | 139.1603 | 138.6864 | 152.4425 | 154.6864 | 168.6028 | 161.5192 | 110.5575 | 297.7945 | 281.1986 | 260.2439 | 303.0069 | 154.6342 | 153.8711 | 163.6342 | 142.1603 | 100.1603 | 90.237 | 74.0453 | 136.1986 | 234.1603 | 226.5644 | 237.6028 | 148.1986 | 158.2753 | 147.0453 | 140.0383 | 137.9547 | 200.5192 | 178.2822 | 217.7247 | 176.8397 | 163.0836 | 160.5644 | 191.4808 | 165.1986 | 209.6725 | 203.1603 | 227.4739 | 256.5575 | 166.7178 | 148.0069 | 278.115 | 237.2753 | 208.9233 | 219.3589 | 214.1986 | 300.1986 | 190.4356 | 189.9164 | 168.77 | 232.4042 | 153.763 | 158.5958 | 147.1672 | 184.352 | 185.9617 | 188.3206 | 187.5644 | 147.3972 | 129.9233 | 122.1986 | 142.23 | 127.4042 | 246.1986 | 288.1986 | 273.3903 | 352.1603 | 264.8328 | 259.9547 | 273.2753 | 329.3903 | 136.9233 | 89.5575 | 80.5192 | 136.9233 | 109.7247 | 121.8083 | 101.5261 | 139.763 | 132.5575 | 120.6028 | 119.4042 | 123.9547 | 87.3972 | 85.9164 | 92.115 | 93.9233 | 113.1219 | 127.1986 | 104.4494 | 93.6794 | 288.5575 | 227.2439 | 228.5644 | 293.352 | 196.5958 | 234.6794 | 183.3275 | 288.0383 | 303.3136 | 204.3658 | 212.8397 | 319.5122
910 | 12.0044 | 198.1018 | 227.4469 | 228.8274 | 194.7611 | 166.6947 | 180.8982 | 173.0708 | 183.8274 | 161.416 | 152.9292 | 130.8628 | 169.5222 | 253.9955 | 180.208 | 149.3805 | 243.0354 | 200.5177 | 170.584 | 163.7212 | 225.8938 | 193.208 | 339.8584 | 345.3451 | 188.5222 | 262.3805 | 238.8938 | 239.7212 | 291.1018 | 132.5177 | 142.5885 | 155.2389 | 161.7257 | 173.4823 | 176.3097 | 161.2434 | 104.0354 | 281 | 277.3451 | 272.8628 | 318.3053 | 137 | 126.6549 | 146.6903 | 144.0354 | 99.1018 | 82.5531 | 87.584 | 131.6549 | 240.3496 | 239.3097 | 245.8274 | 148.1416 | 146.2434 | 164.3805 | 138.2788 | 127.5222 | 197.8274 | 186.0354 | 225.9601 | 178.416 | 167.0973 | 171.2389 | 194 | 158.2389 | 188.7257 | 201.9292 | 210.0354 | 248.4823 | 169.1461 | 170.7257 | 256.9601 | 225.9337 | 217.792 | 214.4469 | 212.5885 | 299.7965 | 182.177 | 179.3451 | 196.0663 | 238.615 | 161.6195 | 149.6239 | 168.8274 | 160.5177 | 190.6549 | 190.6903 | 203.0708 | 136.4115 | 140.3451 | 121.1062 | 112.6903 | 140.5177 | 244.416 | 274.8539 | 248.8319 | 347.9955 | 253.0089 | 247.2788 | 257.1018 | 305.3496 | 146.1372 | 82.1726 | 81.2788 | 146.1372 | 124 | 142.1681 | 125.3496 | 146.7566 | 127.7611 | 125.2035 | 128.2035 | 111.4514 | 81.5885 | 75.8628 | 79.2434 | 107.2788 | 118.1372 | 119.7212 | 129.5486 | 89.9292 | 284.1062 | 239.1726 | 241.6549 | 270.5531 | 185.3805 | 234.3805 | 201.1681 | 286.2788 | 281.0663 | 220.9646 | 215.4514 | 299.1062
911 | 12.0154 | 186.7817 | 220.1849 | 224.118 | 205.5968 | 193.8151 | 189.4454 | 186.882 | 183.5546 | 177.6215 | 142.8151 | 136.1092 | 191.9331 | 230.1849 | 186.1849 | 156 | 250.588 | 203.7394 | 162.3275 | 149.294 | 212.6637 | 198.4454 | 323.6972 | 345.8908 | 212.0423 | 268.0757 | 233.0581 | 227.1426 | 286.0669 | 134.8151 | 151.118 | 137.0088 | 161.8662 | 171 | 167.2694 | 171.2606 | 107.1514 | 276.3785 | 277.3363 | 270.7148 | 290.2606 | 140.882 | 129.4366 | 147.0423 | 146.7817 | 93.882 | 90.5546 | 79.882 | 134.6215 | 265.8151 | 228.6056 | 244.6303 | 168 | 156.2606 | 166.5634 | 153.6303 | 149.5634 | 200.3275 | 193.2183 | 202.1761 | 191.294 | 140.6549 | 161.3275 | 188.4542 | 147.9577 | 199.4032 | 189.0423 | 224.4278 | 245.8151 | 208.0088 | 174.1937 | 233.3609 | 236.3028 | 209.7817 | 210.6972 | 220.1937 | 316.5211 | 202.9331 | 179.1514 | 180.706 | 212.8063 | 154.8151 | 162.5634 | 162.8239 | 165.4032 | 189.3697 | 193.0757 | 207.0845 | 125.4789 | 138.118 | 131.7482 | 119.5123 | 143.5546 | 256 | 241.9067 | 267.9331 | 331.2095 | 301.5546 | 265.4032 | 253.3609 | 328.0423 | 141.6303 | 82.2606 | 94.9666 | 141.6303 | 123.2606 | 118.6303 | 144.5299 | 132.4877 | 137.4877 | 109.9577 | 113.3275 | 129.1092 | 92.3363 | 80.5546 | 92.0334 | 122.0757 | 110.6725 | 105.6637 | 110.7817 | 80 | 289.757 | 247.9489 | 239.2606 | 278 | 190.1514 | 245.6215 | 189.4613 | 294.4208 | 269.0334 | 214.706 | 230.3363 | 314.1849
912 | 12.0264 | 177.3943 | 221.3943 | 209.3658 | 199.7316 | 189.4513 | 187 | 194.7339 | 184.0285 | 186.5653 | 141.2114 | 136.8598 | 177.7459 | 236.9145 | 186.8028 | 159.1544 | 256.8741 | 201.8741 | 174.8313 | 155.7886 | 216.4513 | 195.4085 | 340 | 332.9168 | 197.4228 | 258.0855 | 268.2684 | 236.2684 | 305.5772 | 134.5915 | 137.1544 | 121.9003 | 129.1544 | 168.0427 | 147.3231 | 167.4085 | 98.0142 | 253.8313 | 268.2256 | 251.4775 | 288.3801 | 161.943 | 147.2256 | 137.1544 | 139.7601 | 106.6627 | 91.8171 | 93.057 | 149.4798 | 261.2541 | 202.1117 | 243.7886 | 166.8171 | 154.1829 | 146.4085 | 150.2256 | 129.8028 | 214.6057 | 202.2114 | 222.8313 | 194.8313 | 166.2541 | 169.6912 | 172.0832 | 159.4228 | 209.0142 | 184.0832 | 264.6199 | 242.6342 | 185.2114 | 159.1117 | 252.6627 | 214.2826 | 200.3943 | 226.0142 | 201.9573 | 305.0712 | 181.4513 | 173.3658 | 177.563 | 234.3943 | 156.9573 | 149.9003 | 143.5487 | 179.3516 | 188.0427 | 184.2684 | 183.1687 | 133.5772 | 119.8171 | 116.5487 | 133.0285 | 141.2684 | 259.1544 | 285.6627 | 251.3801 | 356.6912 | 296.5083 | 269.6912 | 271.2826 | 315 | 138.2256 | 80.5772 | 86.1972 | 138.2256 | 121.1829 | 120.943 | 116.1972 | 117.1972 | 125.5487 | 125.1687 | 125.2399 | 134 | 87.7601 | 85.563 | 98.8313 | 114.8456 | 96.9858 | 109.2541 | 102.5772 | 81.7744 | 261.6176 | 285.3658 | 238.9573 | 282.3373 | 181.4085 | 258.3111 | 235.2684 | 256.6627 | 273.6627 | 219.6461 | 218.8598 | 312.6342
913 | 12.0374 | 178.5811 | 225.5133 | 222.3511 | 217.5339 | 178.7699 | 186.5811 | 161.3983 | 179.7699 | 151.3098 | 134.6489 | 124.1888 | 190.9056 | 255.1356 | 179.9262 | 172.2094 | 236.2094 | 189.5871 | 163.9056 | 161.0944 | 206.8171 | 195.3039 | 338.115 | 298.5339 | 190.1622 | 238.8171 | 245.2301 | 234.3836 | 314.3039 | 136.1622 | 143.5073 | 148.115 | 141.7906 | 161.4455 | 166.2773 | 165.6283 | 93.5811 | 260.4986 | 264.1416 | 275.5279 | 290.8584 | 178.8378 | 140.6283 | 148.3245 | 149.3245 | 88.7906 | 86.5811 | 78.885 | 126.5811 | 253.121 | 226.1622 | 245.7434 | 162 | 160.4661 | 153.2154 | 142.5811 | 135.7021 | 210.0678 | 192.9262 | 211.4867 | 188.7227 | 151.6283 | 151.121 | 198.7168 | 154.2566 | 207.0944 | 216.3717 | 253.3451 | 241.7965 | 180.115 | 180.6489 | 240.236 | 197.9116 | 207.8644 | 221.3717 | 216.3039 | 290.236 | 169.0944 | 182.7906 | 185.2773 | 238.3039 | 165.4395 | 167.7374 | 157.7906 | 186.9262 | 180.1888 | 167.9322 | 195.7699 | 142.7227 | 118.3511 | 129.3245 | 127.3039 | 126.3983 | 269.0678 | 258.1563 | 254.0678 | 329.115 | 258.0412 | 244.2094 | 253.6549 | 318.1416 | 133.0944 | 83.8584 | 87 | 133.0944 | 124.5339 | 134.6961 | 115.5339 | 121.1416 | 136.2301 | 135.2566 | 118.6077 | 133.3717 | 92.7168 | 94.7434 | 90.8378 | 100.7434 | 104.7699 | 99.236 | 105.4395 | 86.2773 | 305.6696 | 282.0146 | 255.6077 | 300.4189 | 181.0944 | 224.4867 | 219.5605 | 245.0738 | 266.4721 | 250.0146 | 205.1416 | 303.6283
914 | 12.0484 | 178.9956 | 230.7871 | 234.3348 | 206.6784 | 185.6828 | 183.4479 | 177.3348 | 189.3436 | 188.6784 | 122.2217 | 134.235 | 175.4611 | 234.9956 | 158.3304 | 169.8957 | 235.8913 | 220.4479 | 164.4258 | 173.8781 | 223.0044 | 196.0176 | 328.7827 | 299.4302 | 187.6652 | 257.6652 | 230.5565 | 286.9178 | 323 | 131.8913 | 120.3304 | 138.5609 | 141.4435 | 180.4479 | 156.6652 | 163.3436 | 96.6564 | 312.348 | 273.5609 | 242.2085 | 281.4391 | 175.5698 | 146.326 | 145.326 | 137.4567 | 87.3348 | 84.5565 | 83.1175 | 128.5477 | 271.1307 | 224.9956 | 239.6696 | 167.765 | 160.6784 | 182.9002 | 141.6652 | 154.348 | 198.7783 | 172.4391 | 204.8825 | 193.0132 | 157.1043 | 176.2261 | 175.2217 | 159.4435 | 209.4567 | 206.0176 | 231.7783 | 261.9178 | 170.1175 | 172.6564 | 250.2438 | 231.5389 | 232.6608 | 221.6608 | 227.4391 | 312.2173 | 183.2085 | 180.0044 | 173.6696 | 245.0044 | 152.8869 | 140.0866 | 157.2217 | 162.8913 | 195.3348 | 182.1043 | 206.0088 | 154.1087 | 129.6696 | 123.4435 | 134.8913 | 141.0044 | 267.3128 | 233.6387 | 246.5477 | 322.6652 | 242.1909 | 253.2041 | 273.4479 | 329.1131 | 138.3392 | 76.2129 | 85.4479 | 138.3392 | 120.3304 | 126.8869 | 115.3216 | 132.5565 | 122.8869 | 134.0132 | 129.348 | 133.2173 | 72.9912 | 91.1087 | 88.9002 | 97.9956 | 117.6696 | 115.0088 | 97.0998 | 81.9868 | 279.5521 | 236.5477 | 259.9178 | 303.3304 | 188.7783 | 213.6696 | 228.3436 | 267.7783 | 304.8913 | 205.6564 | 219.2173 | 302.4523
915 | 12.0594 | 185.766 | 216.7021 | 226.8515 | 195.7228 | 158.5959 | 181.2763 | 173.5959 | 176.0423 | 170.468 | 128.1485 | 117.5103 | 149 | 237.7876 | 158.5536 | 161.1062 | 237.6166 | 210.5536 | 182.7669 | 178.5968 | 219.7444 | 173.8506 | 317.9577 | 315.9154 | 184.3834 | 247.3627 | 226.8938 | 242.4041 | 311.5329 | 129.8722 | 121.2556 | 129.2979 | 147.4464 | 175.7021 | 158.0639 | 148.8722 | 110.8938 | 279.171 | 265 | 260.4897 | 285.0855 | 149.0216 | 158 | 156.766 | 123.3825 | 86.1701 | 83.234 | 76.0423 | 140.766 | 240.7021 | 228.2556 | 231.4897 | 186.3618 | 142.234 | 166.4041 | 142.5959 | 135.2124 | 196.3618 | 179.1278 | 211.4897 | 177.3825 | 166.8299 | 167.1917 | 176 | 154.4473 | 191.4041 | 185.2547 | 230.766 | 222.5527 | 165.1485 | 189.7021 | 228.1692 | 250.9793 | 235.2133 | 224.4473 | 228.7453 | 312.7453 | 201.9577 | 177.4464 | 169.234 | 237.2979 | 166.531 | 173.8515 | 155.8938 | 162.2763 | 197.2124 | 193.468 | 195.0423 | 154.2556 | 129.2124 | 124.532 | 127.7237 | 145.4671 | 292.5113 | 281.5103 | 259.234 | 326.8722 | 302.063 | 273.8092 | 269.1701 | 331.8506 | 131.8083 | 89.5959 | 76.2556 | 131.8083 | 126.4041 | 132.1062 | 127.4473 | 131.9361 | 123.6598 | 118.8506 | 111.6166 | 141.234 | 89.8083 | 91.7237 | 72.1701 | 102.1917 | 111.1278 | 97.4897 | 112.6598 | 99.0216 | 281.0216 | 252.7444 | 219.1485 | 303.5536 | 180.5113 | 216.7228 | 215.2763 | 270.5103 | 297.7237 | 218.2556 | 229.1062 | 285.7876
916 | 12.0704 | 179.3348 | 218.2611 | 215.9115 | 218.2905 | 159.6652 | 182.1474 | 172.6947 | 186 | 172.7389 | 145.4926 | 130.7242 | 157.1284 | 226.9558 | 147.4484 | 173.4189 | 242.699 | 201.4189 | 150.9263 | 138.7389 | 224.6106 | 186.3495 | 306.5221 | 301.4189 | 170.4337 | 210.1873 | 226.1579 | 249.7095 | 285.8231 | 135.5074 | 112.4484 | 122.0737 | 156.5811 | 175.7832 | 156.9705 | 155.2463 | 101.7832 | 252.9115 | 266.2316 | 254.4189 | 273.9263 | 137.2316 | 155.2905 | 153.2905 | 148.8821 | 92.4631 | 83.7537 | 88.4631 | 141.9705 | 241.5221 | 218.9558 | 231.0841 | 174.5958 | 149.6505 | 164.5958 | 149.3305 | 145.3053 | 194.2021 | 175.2168 | 197.0442 | 183.1769 | 161.0295 | 159.2316 | 176.4926 | 140.3894 | 194.2758 | 209.2463 | 226.5516 | 245.4779 | 181.0147 | 186.3348 | 258.3937 | 225.2905 | 213.4926 | 206.2021 | 204.7537 | 295.8968 | 187.8127 | 194.7095 | 176.6505 | 236.9705 | 191.8864 | 156.5074 | 150.4779 | 174.9705 | 210.0147 | 189.8274 | 204.2611 | 149.1431 | 139.3053 | 125.7095 | 122.4189 | 171.4232 | 265.0442 | 282.4085 | 260.7389 | 338.6652 | 320.4675 | 249.9558 | 267.5811 | 344.3495 | 138.7832 | 90.1726 | 69.4189 | 138.7832 | 128.2905 | 132.8421 | 107.4779 | 131.5516 | 111.5663 | 132.8274 | 126.5516 | 138.7979 | 95.5516 | 81.2463 | 78.7095 | 98.4189 | 104.0147 | 94.3747 | 104.5074 | 83.7832 | 266.2758 | 263.5221 | 237.7242 | 294.4189 | 155.0147 | 237.0737 | 217.379 | 285.941 | 287.9853 | 209.2021 | 228.8569 | 276.68
917 | 12.0814 | 166.3963 | 224.2754 | 231.4139 | 207.4139 | 178.5516 | 164.9303 | 178.3107 | 183.4139 | 183.0168 | 147.2586 | 132.8967 | 181.7414 | 238.3274 | 162.3963 | 190 | 212.3442 | 215.6726 | 169.8791 | 153.6717 | 205.1033 | 168.5861 | 308.9823 | 307.6558 | 179.8623 | 229.2065 | 245.7758 | 255.9312 | 316.3107 | 133.4828 | 127.6549 | 112.6893 | 140.5172 | 169 | 171.534 | 160.1377 | 99.1377 | 261.9488 | 270.7758 | 267.3795 | 289 | 137.6381 | 149.8447 | 147.5861 | 162.1386 | 97.1553 | 78.8623 | 93.6726 | 136.6214 | 230.0177 | 228.7758 | 259.8279 | 161.7237 | 173.8791 | 150.9479 | 176.5004 | 137.8791 | 199.8623 | 185.3619 | 195.241 | 159.1545 | 157.5861 | 165.5861 | 175.4139 | 162.4828 | 187.4484 | 211.5516 | 215.7414 | 253.8791 | 171.5349 | 170.5516 | 239.0856 | 219.5861 | 212.1553 | 219.1033 | 202.1898 | 325.534 | 180.4307 | 195.5004 | 195.707 | 247.6549 | 163.2242 | 155.7758 | 151.8967 | 175.8279 | 204.4139 | 180.4651 | 199.6726 | 169.4484 | 130.5861 | 134 | 131.759 | 136 | 257.8102 | 258.0168 | 257.5693 | 350.8279 | 272.1209 | 268.0512 | 256.1721 | 325.5516 | 135.8791 | 99.2763 | 86.5172 | 135.8791 | 116.1209 | 119.9479 | 114.5861 | 119.707 | 104.2065 | 120.1033 | 113.6726 | 128.7414 | 82.4139 | 80.4484 | 87.2586 | 113.707 | 101.7758 | 116.3795 | 103.2586 | 80.3619 | 263.0688 | 259.5172 | 242.4828 | 306.3451 | 163.6372 | 215.2763 | 197.7237 | 289.466 | 289.9656 | 213.8279 | 214.241 | 291.7758
918 | 12.0924 | 176.291 | 238.5186 | 222.1037 | 195.9366 | 180.0201 | 192.2477 | 168.5015 | 179.5217 | 192.4552 | 147.4582 | 125.0836 | 179.1037 | 255.2941 | 174.1873 | 180.5186 | 247.147 | 206.8328 | 175.5821 | 185.1239 | 214.2709 | 171.3948 | 320.8731 | 287.2105 | 170.1672 | 239.6858 | 248.2709 | 242.6858 | 305.4179 | 136.3343 | 135.853 | 131.9799 | 147.6888 | 163.3112 | 188.144 | 167.1239 | 112.0836 | 249.003 | 276.2508 | 266.4814 | 289.2709 | 140.7321 | 154.2074 | 147.144 | 127.2074 | 91.9799 | 69.9799 | 85.6455 | 109.4179 | 220.8359 | 230.4582 | 254.2105 | 176.6657 | 180.1239 | 156.6223 | 156.8963 | 143.0403 | 195.856 | 193.1037 | 214.1037 | 195.6254 | 158.2276 | 166.7694 | 173.6888 | 154.7694 | 186.7926 | 207.3313 | 214.4582 | 272.3142 | 153.5418 | 177.709 | 251.4753 | 222.9366 | 210.2306 | 225.5619 | 204.856 | 341.6022 | 190.853 | 179.9597 | 183.8731 | 249.8932 | 157.2074 | 153.9366 | 146.5217 | 172.6485 | 196.7291 | 195.7492 | 197.3343 | 159.853 | 134.4783 | 133.1873 | 111.8127 | 133.291 | 258.1239 | 279.6455 | 252.0232 | 342.5015 | 263.9799 | 290.4783 | 270.1873 | 323.7694 | 148.0836 | 76.7492 | 85.291 | 148.0836 | 109.8761 | 137.8127 | 121.4582 | 122.2306 | 124.709 | 133.3343 | 105.9164 | 129.6254 | 80.9597 | 76.8127 | 85.291 | 110.2709 | 107.5217 | 101.3948 | 107.2508 | 95.4179 | 261.9567 | 255.853 | 242.8963 | 298.9597 | 192.6657 | 199.5217 | 217.8127 | 271.7724 | 292.147 | 210.6485 | 236.3545 | 294.5418
919 | 12.1034 | 170.1328 | 214.6991 | 220.1151 | 195.1948 | 166.6194 | 178.4603 | 176.0884 | 185.3186 | 162.7965 | 148.5487 | 126.8672 | 176.5487 | 273.6548 | 169.1681 | 157.5487 | 268.9293 | 201 | 160.7168 | 173.4159 | 213.3009 | 159.5664 | 298.7168 | 323.3275 | 174.5841 | 230.4779 | 246.1681 | 225.2655 | 323.3983 | 151.6814 | 129.9293 | 140.2832 | 166.9646 | 155.6461 | 165.3009 | 152.8672 | 113.5841 | 283.7168 | 278.2035 | 278.1238 | 291.1328 | 173.5841 | 142.5841 | 137.3275 | 120.1151 | 102.8319 | 77.7168 | 84.2655 | 128.8142 | 243.4867 | 223.6194 | 267.1061 | 169.8142 | 161.0531 | 134.7788 | 159.1681 | 125.3009 | 204.8229 | 189.9823 | 204.1858 | 198.0177 | 149.3452 | 148.0177 | 187.5841 | 139.9823 | 197 | 184 | 212.4336 | 273.0264 | 162.9293 | 179.9026 | 211.531 | 215.3983 | 227 | 214.1858 | 219.2035 | 314.469 | 175.8672 | 190.3362 | 167.3806 | 225.0974 | 150.1151 | 143.2832 | 165.062 | 197.4867 | 199.6814 | 192.3806 | 201.9203 | 140.9026 | 124.1504 | 130.4336 | 123.9116 | 128.5487 | 251.7965 | 275.7168 | 279.5664 | 358.0177 | 275.6814 | 271.938 | 258.938 | 317.4779 | 149.3009 | 71.6814 | 80.5487 | 149.3009 | 128.6638 | 132.3539 | 120.5664 | 130.2212 | 135.3983 | 140.7522 | 100.7345 | 132.0177 | 95.5841 | 79 | 79.4159 | 112.4159 | 119.8319 | 98.062 | 112.8849 | 103.7699 | 251.4956 | 243.6991 | 244.8849 | 305.0884 | 182.1328 | 216.0797 | 223.6814 | 278.3893 | 298.3539 | 236.9026 | 236.3186 | 296.8496
920 | 12.1144 | 180.9779 | 216.6362 | 228.2955 | 210.9537 | 155.0683 | 195.817 | 159.8644 | 178.9548 | 174.9769 | 149.9769 | 128.1367 | 179.7497 | 245.2723 | 171.7508 | 160.4543 | 268.9075 | 192.7266 | 159.7045 | 165.1809 | 208.7045 | 161.2955 | 300.3638 | 343.1356 | 170.1136 | 254.0452 | 239.591 | 236.6141 | 326.6814 | 154.795 | 148.5226 | 136.8633 | 164.2944 | 177.0683 | 160.1136 | 154.7277 | 110.8864 | 288.9096 | 266.0231 | 242.409 | 301.0915 | 165.8633 | 140.4774 | 159.204 | 127.7045 | 101.1356 | 80.8412 | 89.409 | 136.2723 | 226.7045 | 212.6593 | 245.2734 | 172.8402 | 151.3186 | 154.0452 | 151.7045 | 128.6824 | 184.9327 | 196.7729 | 206.7739 | 200.0915 | 170.9317 | 138.2723 | 181.0452 | 146.7729 | 196.409 | 197.2965 | 227.5468 | 230.6141 | 181.8181 | 161.3859 | 228.9096 | 224.8864 | 222.5678 | 206.7277 | 208.796 | 300.9317 | 178.3186 | 180.1367 | 184.8412 | 233.5678 | 154.4543 | 150.796 | 176.4763 | 183.9548 | 207.5226 | 202.4543 | 192.2734 | 133.6141 | 127.6141 | 128.1136 | 140.1356 | 134.409 | 272.5457 | 269.3859 | 283.9548 | 346.4995 | 272.8854 | 256.0915 | 246.0463 | 341.0452 | 145.2955 | 86.9096 | 84.3407 | 145.2955 | 134.7266 | 120.9779 | 114.613 | 113.6141 | 138.9769 | 121.1356 | 105.0463 | 121.0904 | 88.4543 | 79.8864 | 82.1136 | 108.0221 | 119.9085 | 118.9317 | 104.7045 | 86.5226 | 271.8181 | 253.3186 | 230.4995 | 283.5457 | 178.795 | 230.7497 | 232.7045 | 262.2965 | 289.3417 | 226.6593 | 231.4322 | 289.2492
921 | 12.1254 | 195.999 | 217.1554 | 227.7689 | 205.1584 | 166.4632 | 166.3088 | 184.6932 | 184.4612 | 167.3088 | 142.6165 | 136.1534 | 180.6175 | 247.5419 | 191.3068 | 146.4601 | 230.2331 | 184.0797 | 167.3098 | 141.6165 | 211.3855 | 162.6155 | 308.4622 | 321.0767 | 177.8486 | 244.8456 | 239.4612 | 250 | 314 | 139.3845 | 145.9233 | 134.6942 | 137.6932 | 188.4632 | 163.2321 | 162.6145 | 105.6135 | 292.5358 | 278.3078 | 239.1534 | 305.997 | 155.6922 | 145.2311 | 142.0767 | 127.9233 | 84.6165 | 87.8456 | 88.9233 | 129.001 | 224.7689 | 224.5399 | 268.3855 | 157.7709 | 159.3835 | 143.6145 | 151 | 143.6145 | 198.6892 | 190.5358 | 229.997 | 206.5358 | 161.3835 | 134.3865 | 168.1524 | 144.2291 | 199.3088 | 230.8466 | 245.7639 | 234.7669 | 173.4601 | 155.3865 | 241.7689 | 227.9233 | 209.8456 | 215.8456 | 226.2311 | 299.3855 | 178.997 | 190.9233 | 185.998 | 229.1554 | 148.1544 | 172.5399 | 152.3855 | 182.998 | 195.6902 | 199.8476 | 201.2281 | 153.1554 | 139.1534 | 130.001 | 128.2331 | 136.6932 | 284.0777 | 248.6135 | 282.0747 | 330.0767 | 250.4632 | 265.6135 | 259.2281 | 333.6922 | 143.5378 | 100.3845 | 81.3855 | 143.5378 | 119.3088 | 139.6922 | 107.4642 | 128.8466 | 129.1544 | 110.4642 | 128.3845 | 117.4642 | 85.5399 | 86.9243 | 75.0757 | 91.1544 | 106.0777 | 114.6155 | 104 | 80.5378 | 260.998 | 257.6902 | 231.004 | 295.0777 | 169.2321 | 221.4612 | 224.9213 | 282.3048 | 309.0757 | 227.1524 | 233.6902 | 286.9283
922 | 12.1364 | 186.0398 | 229.3997 | 236.5222 | 231.2772 | 181 | 180.4808 | 188.8392 | 184.2006 | 176.6799 | 148.8392 | 132.6401 | 200 | 261.997 | 178.1991 | 148.0827 | 258.0413 | 212.2006 | 187.9204 | 152.6401 | 216.4395 | 163.3599 | 319.121 | 313.879 | 189.438 | 236.7994 | 242.4012 | 256.0812 | 306.6386 | 138.3201 | 148.6401 | 148.4793 | 155.6018 | 194.0383 | 175.9602 | 150.8392 | 88.6799 | 269.6401 | 283.4808 | 242.0413 | 283.8009 | 156.2802 | 150.5605 | 138.0796 | 126.7994 | 92.1195 | 83 | 90.6799 | 139.6003 | 226.1608 | 231.559 | 265.1179 | 179.7213 | 146.1195 | 143.0413 | 154.5207 | 135.6799 | 175.1608 | 176.6018 | 212.2818 | 183.9602 | 159.0015 | 164.4425 | 163.0015 | 135.4012 | 206.7596 | 230.5192 | 207.9219 | 221.6815 | 174.1225 | 175.8407 | 240.9204 | 222.3185 | 217.8024 | 219.6416 | 220.3584 | 304.7596 | 160.9617 | 187.879 | 179.2021 | 247.1608 | 156.2006 | 182.4395 | 155.5192 | 164.0398 | 180.5207 | 207.879 | 190.7626 | 162.8392 | 125.7182 | 134.8392 | 144.1991 | 142.4395 | 285.879 | 252.4839 | 250.7979 | 337.6018 | 262.4395 | 257.9617 | 238.2006 | 341.0015 | 144.4012 | 102.2006 | 89.9602 | 144.4012 | 128.3599 | 131.3185 | 125.2389 | 131.0796 | 134 | 125.0383 | 120.5988 | 133.3185 | 88.3982 | 91.5988 | 69.121 | 97.9204 | 109.7994 | 118.5605 | 99.1991 | 81.7213 | 256.1225 | 233.879 | 260.9188 | 301.3599 | 174.9188 | 222.1608 | 212.5207 | 272.3231 | 297.6799 | 222.6416 | 216.6003 | 336.5605
923 | 12.1474 | 190.647 | 224.0059 | 251.3618 | 207.6441 | 180.0029 | 188.0059 | 173.0206 | 181.0294 | 174.0059 | 137.3412 | 133.0029 | 186.7059 | 229.3294 | 169.3294 | 168.3618 | 263.6942 | 210.6912 | 189.6735 | 158.6529 | 214.9882 | 174.9617 | 316.0412 | 304.3294 | 172.3176 | 239.6382 | 256.3324 | 255.3736 | 291.6647 | 138.6676 | 146.0118 | 133.3559 | 162.0412 | 176.3294 | 173.0147 | 142 | 93.9823 | 274.3235 | 292.3353 | 256.6676 | 278.3795 | 152.0206 | 155.3353 | 132.3382 | 119.0029 | 91.3176 | 84.3294 | 91.3294 | 145.9912 | 226.0265 | 238.2793 | 235.353 | 181.3706 | 152.6294 | 161.3235 | 153.3589 | 136.3294 | 192.6411 | 192.3471 | 228.6765 | 185.3353 | 161.7089 | 189.6765 | 174.6824 | 152.6559 | 196.35 | 215.0177 | 235.3147 | 250.297 | 193.6912 | 183.6676 | 248.6559 | 206.3324 | 240.3471 | 228.3618 | 210.6411 | 304.9853 | 177.0088 | 174.6735 | 188.7089 | 248.0236 | 161.3382 | 172.6794 | 148.9912 | 169.6441 | 187.3353 | 198.3294 | 208.7 | 155.9941 | 113.9764 | 123.6735 | 130.6765 | 133.0118 | 273.0059 | 280.7118 | 230.3235 | 358.3324 | 270.9588 | 270.353 | 250.9823 | 342.3795 | 152.0177 | 102.0206 | 81.3647 | 152.0177 | 131.3206 | 116.6618 | 107.6824 | 132.3176 | 131.6735 | 115.9706 | 103.0118 | 122.3176 | 68.9971 | 80.3265 | 82.3265 | 101.0029 | 113.6353 | 117.0206 | 91.3265 | 96.9971 | 281.3412 | 222.6676 | 248.6647 | 293.353 | 180.944 | 229.0059 | 215.0147 | 291.3795 | 295.3382 | 239.3382 | 223.9882 | 328.0412
924 | 12.1584 | 200.5536 | 214.8304 | 236.1357 | 215.7625 | 176.2768 | 176.4179 | 161.7571 | 168.2375 | 173.1018 | 136.4804 | 130.3107 | 171.7179 | 222.35 | 165.4518 | 147.6214 | 246.3786 | 190.5536 | 180.175 | 166.9661 | 220.5875 | 188.5929 | 295.5821 | 306.6554 | 193.3393 | 255.5536 | 246.6607 | 241.0964 | 290.5875 | 137.6554 | 135.9321 | 125.2375 | 148.1304 | 177.6214 | 163.3446 | 145.7911 | 103.5875 | 284.4464 | 289.6214 | 261.5143 | 261.475 | 151.4411 | 151.9321 | 143.475 | 130.7857 | 99.2429 | 84.2429 | 88.4857 | 139.5929 | 210.7571 | 253.9768 | 220.9661 | 158.0679 | 176.2768 | 161.7964 | 143.9268 | 133.8304 | 201.0393 | 184.3786 | 226.1357 | 178.4857 | 147.8196 | 178.8643 | 170.8929 | 147.2482 | 183.9661 | 192.3161 | 235.2482 | 255.7 | 170.4518 | 182.6554 | 243.9375 | 215.6161 | 228.5875 | 212.1018 | 222.1411 | 311.2089 | 174.4464 | 174.4804 | 167.5821 | 233.7232 | 162.4804 | 161.2429 | 149.4857 | 175.35 | 192.2036 | 208.2375 | 183.9321 | 150.35 | 125.8643 | 121.4125 | 126.7571 | 133.9607 | 265.8982 | 265.8536 | 241.4804 | 356.9321 | 279.3554 | 260.4464 | 246.8018 | 309.9661 | 143.7911 | 82.1411 | 69.9946 | 143.7911 | 134.8304 | 121.0339 | 103.8929 | 147.825 | 123.8982 | 131.5196 | 106.0286 | 125.4179 | 74.1018 | 88.7911 | 86.6554 | 99.3446 | 123.2821 | 106.4464 | 99.7911 | 96.5875 | 269.1071 | 223.3786 | 238.2821 | 279.3107 | 204.525 | 214.3161 | 207.7571 | 278.6107 | 292 | 231.5196 | 232 | 296.2089
925 | 12.1694 | 193.6431 | 221.062 | 235.0767 | 193.7905 | 169.0737 | 164.8555 | 169.4985 | 193.4219 | 170.0767 | 142.1445 | 122.2183 | 190.4307 | 212.9233 | 153 | 156.7788 | 251.8555 | 188.6401 | 167.0737 | 162.5015 | 210.2891 | 178.2066 | 318.9233 | 307.4277 | 202.292 | 252.2124 | 227.7139 | 249.938 | 284.9292 | 129.5044 | 130.2153 | 142.2124 | 163.5781 | 177.4985 | 160.7876 | 150.5015 | 101.8555 | 273.8672 | 289.1416 | 261.6489 | 272.581 | 168.7905 | 150.1416 | 162.4336 | 147.7197 | 95.0708 | 79.3569 | 83.354 | 113.5015 | 215.6431 | 217.2124 | 224.354 | 165.5693 | 162.6489 | 149.2861 | 160.0708 | 128.2832 | 190.7817 | 196.2803 | 232.9292 | 172.6401 | 171.3599 | 168.1445 | 174.3628 | 133.3511 | 182.7139 | 184.8496 | 222.7788 | 220.4277 | 166.9233 | 176.646 | 221 | 224.5044 | 211.1504 | 214.7876 | 210.6431 | 308.9971 | 189.5664 | 189.4248 | 178.0737 | 238.4277 | 162.0767 | 154.5723 | 143.2832 | 160.9263 | 202.5723 | 213.4336 | 182.8555 | 140.2095 | 121.5693 | 125.6431 | 125.2183 | 143.5044 | 256.0737 | 287.4336 | 250 | 344.4336 | 247.354 | 265.2153 | 223.1387 | 310.4985 | 150.6431 | 73.4985 | 81.5044 | 150.6431 | 120.7168 | 131.2095 | 117 | 144.5073 | 122.6401 | 127.9971 | 124.8584 | 122.4219 | 78.5723 | 92.7876 | 85.6431 | 100.3569 | 104.9263 | 114.4277 | 102.3599 | 92.7139 | 269.2066 | 239.2066 | 220.2832 | 288.3511 | 175.2153 | 203.6372 | 209.0737 | 282.5898 | 291.2861 | 214.4336 | 229.5015 | 298.2803
926 | 12.1804 | 189.3078 | 230.8003 | 225.2768 | 186.3845 | 167.1691 | 181.446 | 172.5231 | 210.677 | 155.8461 | 128.8771 | 115.5384 | 181.4154 | 221.9847 | 164.0767 | 172.8309 | 248.8771 | 201.0615 | 164.7999 | 159.8461 | 199.6308 | 185.0157 | 336.4769 | 319.2306 | 184.1077 | 248.4002 | 246.3536 | 219.1539 | 281.8923 | 125.5998 | 132.1691 | 136.1848 | 149.2768 | 180.5231 | 157.9538 | 140.4616 | 101.0924 | 253.7075 | 298.9076 | 255.9229 | 265.5227 | 150.677 | 148.8309 | 148.8461 | 129.1539 | 92.5693 | 79.6308 | 83.0309 | 111.5846 | 219.7999 | 222.2616 | 245.446 | 169.4154 | 156.9229 | 146.8923 | 159.0462 | 136.7384 | 199.677 | 201.554 | 233.9538 | 181.3692 | 156.354 | 158.5693 | 158.8461 | 148.3078 | 186.9538 | 196.554 | 235.1386 | 236.6613 | 180.4154 | 176.9691 | 222.1077 | 211.7384 | 193.9076 | 203.4616 | 205.2001 | 322.7999 | 194.2925 | 193.2463 | 168.0462 | 250.5998 | 150.7999 | 152.7384 | 155.4307 | 163.9385 | 190.7694 | 205.3845 | 192.0615 | 146.8771 | 129.4769 | 126.1077 | 125.5536 | 134.4307 | 241.6155 | 269.0462 | 240.7694 | 338.9691 | 260.6922 | 257.5693 | 256.6766 | 318.323 | 145.2001 | 87.9691 | 70.2153 | 145.2001 | 116.6922 | 134.554 | 121.0615 | 133.8152 | 128.4154 | 133.3078 | 118.0462 | 137.4616 | 76 | 84.4154 | 94.2306 | 96.9385 | 112.3692 | 115.5231 | 95.8461 | 92.1539 | 276.0157 | 263.7384 | 241.6613 | 303.3078 | 180.4922 | 215.4616 | 210.4922 | 250.7381 | 302.9229 | 214.5074 | 227.9538 | 313.5231
927 | 12.1914 | 178.0403 | 203.7227 | 232.5664 | 204.3412 | 167.6066 | 192.9882 | 174.5782 | 195.4621 | 166.6303 | 124.5379 | 126.9076 | 184.6303 | 220.6824 | 167.7394 | 172.3934 | 235.3815 | 209.1445 | 182.9194 | 159.5664 | 204.3412 | 159.4336 | 333.6588 | 315.5427 | 178.3697 | 231.237 | 270.6588 | 219.0521 | 284.9597 | 141.237 | 137.1848 | 118.6706 | 153.1327 | 176.4739 | 164.5261 | 134.0521 | 95.9597 | 263.2251 | 293.3578 | 271.1327 | 289.5024 | 153.0118 | 148.0118 | 141.6991 | 127.9076 | 90.3412 | 85.1043 | 70.1967 | 129.7346 | 233.3412 | 220.6185 | 258.5142 | 166.1445 | 169.4621 | 148.8152 | 154.3815 | 136.0924 | 192.0924 | 178.5664 | 232.1327 | 172.0806 | 150.3412 | 151 | 161.6185 | 140.4739 | 185.1327 | 184.2488 | 216.1848 | 256.1848 | 177.1445 | 186.3697 | 217.1327 | 216.8152 | 207.4336 | 194.3815 | 215.3128 | 320.3176 | 188.2085 | 178.0924 | 170.6303 | 251.7394 | 155.9479 | 159.7227 | 156.5142 | 155.8555 | 187.9715 | 210.3697 | 196.7109 | 139.4858 | 128.9479 | 121.5142 | 132.5024 | 145.4336 | 252.654 | 255.6066 | 234.5379 | 344.173 | 269.6706 | 246.9479 | 282.4621 | 326.2891 | 133.5664 | 97.3697 | 77.8152 | 133.5664 | 127.1967 | 125.3009 | 123.4218 | 144.0806 | 123.237 | 130.4336 | 122.1564 | 127.6185 | 72.9479 | 77.3294 | 105.0403 | 99.5379 | 110.7109 | 107.2773 | 95.5664 | 89.3815 | 255.0118 | 256.6066 | 250.5024 | 286.3176 | 195 | 214.0118 | 216.6066 | 284.2251 | 311.6469 | 241 | 230.7109 | 307.5664
928 | 12.2024 | 165.2743 | 213.7876 | 222.2743 | 208.6681 | 163.0576 | 188.6638 | 177.2743 | 178.031 | 169.5486 | 145.1195 | 130.3938 | 189.5177 | 242.7257 | 152.4514 | 175.761 | 247.4204 | 204.6991 | 187.1858 | 155.969 | 202.761 | 174.8452 | 318.3053 | 296.9956 | 167.969 | 241.8142 | 243.4911 | 224.3938 | 296.1505 | 148.2699 | 138.0885 | 130.4823 | 141.2124 | 171.6372 | 166.2124 | 137.031 | 102.031 | 255.6947 | 269.969 | 254.8805 | 269.1505 | 154.1858 | 156.2743 | 125.7876 | 132.5752 | 94.6681 | 85.5486 | 79.3628 | 142.9424 | 242 | 215.2743 | 248 | 177.8452 | 157.5443 | 151.0619 | 153.031 | 130.6372 | 185.8495 | 185.208 | 222.5752 | 158.7566 | 164.5133 | 154.9381 | 170.0576 | 143.1195 | 178.7257 | 205.4514 | 210.3938 | 240.5486 | 182.5443 | 169.2743 | 205.2124 | 217.4867 | 216.1505 | 201.3009 | 224.9158 | 292.3672 | 196.9734 | 188.3894 | 193.239 | 225.8186 | 146.2743 | 166.6372 | 143.6372 | 162.2699 | 199.5486 | 190.1239 | 190.2434 | 147.6372 | 125.969 | 114.9381 | 120.0266 | 149.031 | 270.9158 | 247.9071 | 256.6947 | 328.9381 | 267.6991 | 253.0266 | 265.031 | 330 | 135.4823 | 89.7257 | 85.969 | 135.4823 | 126.5177 | 129.792 | 127.4204 | 128.7876 | 125.9381 | 130.4867 | 119.2787 | 127 | 75.4823 | 81.6638 | 89.9115 | 106.3362 | 118.4204 | 100.7566 | 92.3628 | 86.8495 | 267.2124 | 253.239 | 227.3938 | 277.6638 | 189.0929 | 225.031 | 207.3319 | 273.5443 | 298.6638 | 241 | 218.7301 | 313.4204
929 | 12.2134 | 158 | 227.589 | 209.7207 | 199.1573 | 168.4719 | 190.2232 | 178.5281 | 184.3403 | 160.6854 | 155.4366 | 128.5634 | 173.939 | 247.9695 | 171.8732 | 175.4317 | 263.3756 | 203.5281 | 174.2793 | 165.0915 | 190.6549 | 187.1573 | 289.7256 | 314.9695 | 174.6549 | 260 | 227.9951 | 224.5939 | 301.8122 | 159.7817 | 149.4061 | 147.2793 | 144.4671 | 158.2536 | 156.878 | 141.7159 | 94.5329 | 264.3756 | 276.2488 | 246.5939 | 271.9695 | 146.5585 | 150.2183 | 133.4976 | 143.5281 | 96.122 | 75.4671 | 83.4061 | 134.6854 | 229.0049 | 206.7817 | 241.5024 | 186.9085 | 154.4719 | 150.6854 | 150 | 133.9037 | 183.8427 | 203.1878 | 216.4719 | 158.939 | 158.3805 | 161 | 177.9085 | 138.8171 | 189.2488 | 213.0963 | 215.8732 | 224.3756 | 170.5378 | 162.8122 | 208.4671 | 207.7817 | 224.6549 | 219.9342 | 202.9342 | 289.8073 | 183.6549 | 204.5329 | 203.3146 | 224.8122 | 145.0915 | 161.7817 | 140.4061 | 182.7159 | 182.1573 | 188.5585 | 186.8122 | 145.2183 | 129.4061 | 117.7512 | 141.8732 | 150.4671 | 238.7817 | 249.2841 | 272.2793 | 334.1878 | 270.1829 | 278.5281 | 248.5988 | 331.6244 | 136.0354 | 89.2841 | 89 | 136.0354 | 113.7817 | 118.7159 | 123.8829 | 134.4671 | 124.2841 | 128.0915 | 100.9342 | 127.8122 | 87 | 95 | 87.122 | 106.4012 | 123.411 | 107.4366 | 98.8427 | 86.8732 | 266 | 258.8475 | 219.878 | 293.4366 | 180.4061 | 211.4415 | 218.061 | 274.5329 | 300.6293 | 231.2536 | 217.7768 | 328.1573
930 | 12.2244 | 165.1126 | 242.2345 | 213.7149 | 196.2552 | 167.2965 | 183.4597 | 194.7655 | 192.9793 | 158.1219 | 144.3564 | 118.3058 | 173.0207 | 246.6736 | 176.9081 | 173.315 | 269.3678 | 213.0713 | 179.9081 | 181.1126 | 202.2758 | 183.8368 | 307.7676 | 296.938 | 179.5816 | 253.3058 | 243.2345 | 244.9195 | 291.7035 | 144.6115 | 145.3977 | 139.938 | 153.0919 | 149.4391 | 157.9701 | 151.7448 | 93.8574 | 266.1839 | 277.2345 | 260.6436 | 260.2138 | 156.1426 | 152.4184 | 146.7655 | 143.4483 | 101.7242 | 78.2345 | 78.9793 | 142.1632 | 222.1333 | 204.1632 | 242.4597 | 176.3264 | 139.4897 | 159 | 146.2345 | 141.0713 | 180.8874 | 198.6529 | 215.2965 | 164.7149 | 139.9287 | 166.4391 | 175.2758 | 131.0413 | 190.6529 | 198.3977 | 222.5816 | 225.7655 | 159.2252 | 169.8574 | 222.9494 | 215.2046 | 231.2552 | 228.8161 | 212.6529 | 309.1632 | 176.8667 | 196.3264 | 187.0506 | 232.6942 | 159.0207 | 158.3264 | 151.8781 | 182.2851 | 183.8574 | 196.8874 | 183.3977 | 149.9287 | 129.5816 | 114.2552 | 137.7035 | 154.0713 | 252.8988 | 237.5816 | 270.7954 | 342.2046 | 264.9886 | 277.6115 | 234.4391 | 328.5609 | 130.4597 | 76.7448 | 85.6529 | 130.4597 | 114.5103 | 127.4897 | 108.4804 | 141.4184 | 122.2046 | 136.5816 | 121.9494 | 135.6942 | 92.8574 | 85.7954 | 94.8161 | 107.5909 | 103.7448 | 107.6529 | 97.9793 | 87.3058 | 263.9081 | 245.2552 | 216.3678 | 287.7954 | 186.8574 | 208.9701 | 224 | 280.5517 | 289.0207 | 220.3471 | 239.1839 | 326.9287
931 | 12.2354 | 169.8318 | 228.3864 | 221.8171 | 204.4602 | 188.1682 | 196.2773 | 200 | 186 | 156.3511 | 131.7522 | 118.9056 | 176.5546 | 247.708 | 159.3569 | 156.2272 | 255.5899 | 191.7728 | 173.9852 | 185.8318 | 214.8761 | 189.3068 | 346.3126 | 293.6579 | 182.0148 | 250.8909 | 242.3068 | 267.9705 | 278.1534 | 128.1534 | 136.4159 | 120.7227 | 153.8466 | 163.0295 | 165.6637 | 147.8466 | 102 | 278.8909 | 262.9557 | 273.2478 | 262.2124 | 152.7375 | 146.9705 | 141.233 | 132.4602 | 96.0148 | 76.0148 | 80.1829 | 137.1239 | 239.8614 | 212.0443 | 258.7227 | 182.6136 | 144.6136 | 164.1682 | 147.4602 | 135.2773 | 164.9705 | 191.8466 | 221.525 | 175.8318 | 153.9056 | 179.1682 | 180.9852 | 142.4159 | 186.8614 | 199.7522 | 224.5841 | 232.7227 | 180.7227 | 175.8466 | 237.0148 | 225.8466 | 225.2478 | 207.9262 | 202.4012 | 302.8318 | 166.8909 | 190.5546 | 178.1534 | 235.5398 | 159.9705 | 155.1386 | 150.6342 | 169.4454 | 185.9705 | 185.708 | 190.7818 | 147.1091 | 131.5841 | 117.292 | 121.1386 | 146.1239 | 271.1239 | 239.5841 | 251.1091 | 350.6932 | 242.9056 | 265.0295 | 241.5693 | 302.4808 | 155.3363 | 75.8614 | 79.2773 | 155.3363 | 125.3216 | 127.4454 | 120.9557 | 140.2773 | 131.9852 | 143.3216 | 125.2478 | 142.8466 | 97.5546 | 68.2625 | 79.9557 | 96.2124 | 111.475 | 111.6136 | 94.4454 | 81.0148 | 253.6784 | 247 | 221.9705 | 290.0738 | 188.9705 | 219.6784 | 239.5045 | 290.2478 | 296.4307 | 231.4602 | 231.6489 | 339.6136
932 | 12.2464 | 160.7851 | 217.443 | 213.6578 | 203.2546 | 188.3554 | 183.8117 | 191.1406 | 196.1883 | 167.3554 | 149.9735 | 138.2016 | 179.0875 | 232.2679 | 160.3767 | 175.8992 | 228.6446 | 191.1486 | 167.7851 | 171.0265 | 223.1008 | 180.8249 | 317.2149 | 304.1538 | 182.8992 | 254.2414 | 240.9125 | 245.382 | 281.8859 | 128.3422 | 130.3422 | 126.9867 | 163.4032 | 175.4297 | 159.0875 | 148.1008 | 91.8117 | 293.3157 | 246.1273 | 250.5968 | 261.5093 | 153.2414 | 159.3767 | 135.8594 | 137.4562 | 92.0265 | 86.6446 | 90.1141 | 130.2281 | 241.443 | 220.8992 | 242.3952 | 179.8249 | 161.3157 | 159.0398 | 155.1141 | 129.4562 | 174.7188 | 199.6313 | 220.5173 | 172.5438 | 164.3422 | 182.0133 | 184.0875 | 141.6578 | 182.6843 | 218.4165 | 224.8992 | 223.9257 | 172.8117 | 176.9867 | 250.7454 | 219.4562 | 227.4032 | 203.061 | 209.618 | 307.5173 | 181.3157 | 185.1141 | 179.6711 | 250.0345 | 156.8727 | 159.7586 | 133.8727 | 180.2016 | 186.4165 | 196.4032 | 203.6843 | 145.9735 | 127.9125 | 127.8594 | 116.0133 | 144.1008 | 263.3422 | 263.8196 | 254.8462 | 329.939 | 252.8992 | 277.4297 | 239.6711 | 299.2626 | 161.0265 | 75.6711 | 83.6446 | 161.0265 | 120.382 | 123.5835 | 109.443 | 144.6446 | 118.2546 | 137.496 | 112.7851 | 131.1406 | 94.3289 | 67.3157 | 75.5305 | 115.443 | 117.3554 | 115.4695 | 103.4297 | 88.9867 | 255.0743 | 254.5305 | 219.7586 | 309.1141 | 186.3157 | 222.4032 | 235.1936 | 284.8727 | 297.443 | 234.6843 | 212.2281 | 335.939
933 | 12.2574 | 160.2763 | 218 | 222.0079 | 199.65 | 178.6342 | 173.2763 | 190.0158 | 200.35 | 168.1052 | 153.1711 | 145.0895 | 186.6342 | 217.2763 | 182.8132 | 176.0974 | 233.3579 | 208.529 | 159.5369 | 163.2763 | 227.9105 | 183.9421 | 317.7237 | 298.0158 | 184.9184 | 236.529 | 239.2842 | 228.3658 | 293.0158 | 131.3737 | 132.4631 | 131.5447 | 167.1632 | 179.179 | 159.8053 | 152.9105 | 92.2026 | 292.7158 | 256.8447 | 250.4789 | 248.8368 | 153.7395 | 174.0737 | 152.0079 | 140.7395 | 84.2763 | 89.9921 | 92.6421 | 135.2842 | 226.0658 | 214.7237 | 239.4868 | 172.471 | 161.1711 | 142.6342 | 161.7395 | 132.7395 | 197 | 205.7158 | 219.9763 | 173.8132 | 156.4474 | 167.8948 | 188.4474 | 149.0974 | 191.9342 | 215.7974 | 215.5369 | 215.0079 | 165.9184 | 177.4474 | 252.529 | 221.8289 | 228.4316 | 220.6342 | 219.8053 | 319.7237 | 186.6342 | 184 | 183.9184 | 251.6843 | 146.1553 | 158.8053 | 139.5447 | 179.35 | 192.9026 | 190.6027 | 194.7237 | 145.0737 | 117.179 | 136.7237 | 130.5763 | 136.1632 | 254.5369 | 285.4474 | 256.8869 | 311.9105 | 258.5605 | 278.9026 | 255.3 | 313.7079 | 147.8132 | 87.6579 | 81.9842 | 147.8132 | 123.8526 | 118.0079 | 115.0079 | 146.1711 | 103.7237 | 118.4553 | 108.179 | 117.2684 | 90.0816 | 81.2842 | 95.9263 | 113.7237 | 103.5369 | 111.4631 | 109 | 97.1868 | 274.0079 | 247.6106 | 228.821 | 307.5447 | 190.7237 | 236.179 | 207.1868 | 275.5211 | 292.0816 | 217.9842 | 212.7316 | 321.5526
934 | 12.2684 | 165.8053 | 225.4809 | 229.8702 | 198.7787 | 183.5457 | 178.3378 | 202 | 194.208 | 172.9218 | 131.441 | 132.3112 | 182.1947 | 223.2729 | 178.5324 | 164.6356 | 236.4809 | 190.3894 | 152.5324 | 159.4543 | 213.1032 | 199.1814 | 323.416 | 309.0649 | 188.7271 | 251.1181 | 246.5973 | 231.4027 | 298.4543 | 142.7404 | 139 | 132.8702 | 165.1564 | 182.6106 | 163.559 | 155.8702 | 107.5324 | 277.4543 | 256.8952 | 257.7787 | 270.9484 | 146.6356 | 153.5191 | 160.3378 | 147.1947 | 89.8053 | 84.9351 | 95.1298 | 146.8053 | 214.9484 | 221.3511 | 264 | 189.6755 | 170.2995 | 150.3511 | 154.6356 | 138.2596 | 190.9218 | 187.6489 | 213.2213 | 176.5457 | 156.6106 | 145.3112 | 185.3378 | 148.8569 | 200.7138 | 199.0649 | 225.8319 | 228.0133 | 168.3245 | 180.416 | 218.025 | 223.9218 | 215.208 | 200.7654 | 210 | 301.1032 | 194.8186 | 191.0133 | 181.6489 | 230.6888 | 143.2995 | 149.4676 | 151.1564 | 170.8702 | 175.8436 | 183.2346 | 197.1431 | 134.3378 | 111.2596 | 138.208 | 140.441 | 122.9351 | 262.0266 | 265.5058 | 247.9484 | 337.7803 | 256.9085 | 268.3894 | 264.3112 | 306.4809 | 142.5973 | 99.3245 | 73.7404 | 142.5973 | 138.7787 | 119.792 | 114.9218 | 138 | 99.1298 | 119.4676 | 113.4809 | 118.208 | 77.8569 | 83.4543 | 102.9218 | 102.584 | 104.9484 | 113.792 | 102.4543 | 96.792 | 278.1298 | 244.8319 | 230.0649 | 309.8053 | 200.6239 | 226.0516 | 218.0133 | 252.3245 | 283.5973 | 213.4809 | 216 | 302.2197
935 | 12.2794 | 165.1612 | 224.4031 | 239.1977 | 189.3589 | 173.7256 | 171.4031 | 179.4473 | 179.3264 | 168.3992 | 130.8752 | 127.6372 | 167.9636 | 231.3186 | 171.2822 | 156.5566 | 239.2419 | 185.9194 | 165.4356 | 169.2744 | 201.238 | 184.1612 | 309.0078 | 304.1612 | 182.6008 | 261.8465 | 246.4395 | 229.6411 | 297.238 | 134.0442 | 130.8426 | 147.476 | 164.6047 | 181.3225 | 185.7178 | 157.5202 | 104.1209 | 276.238 | 240.238 | 257.476 | 279.5202 | 135.6775 | 145.9597 | 155.3225 | 143.0403 | 89.1612 | 84.0806 | 84.8426 | 140.4031 | 224 | 254.0728 | 251.0442 | 198.3589 | 179.4434 | 163.8791 | 140.3186 | 132.5605 | 189.2783 | 177 | 216.4031 | 180.6411 | 171.1574 | 139.6775 | 185.1209 | 160.7542 | 181.8023 | 194.6411 | 245.9597 | 238.8791 | 163.9597 | 180.883 | 218.4318 | 231.3953 | 218.0806 | 181.9194 | 207.1209 | 287.7984 | 199.762 | 200.4395 | 179.1574 | 243.5202 | 150.524 | 151.4395 | 166.4395 | 160.524 | 188.3914 | 182.407 | 217.8752 | 148.5163 | 109.8791 | 135.3225 | 124.5202 | 130.238 | 272.7217 | 252.8388 | 249.8023 | 342.9675 | 243.4395 | 266.7984 | 255.3186 | 320.7581 | 146.7581 | 96.3992 | 74.6411 | 146.7581 | 123.6008 | 115.4798 | 103.6814 | 140.8791 | 113.7946 | 125.2783 | 130.1574 | 122.5202 | 78.238 | 75.0403 | 84.0039 | 101.1574 | 110.1612 | 102.2822 | 99.7984 | 89.6008 | 279.8388 | 253.4434 | 229.4798 | 312.0403 | 193.8868 | 207.4031 | 222.6411 | 258.5163 | 295.4356 | 215.762 | 201.6047 | 296.0728
936 | 12.2904 | 170.8427 | 212.5236 | 242.5708 | 200.8742 | 165.3348 | 161.9843 | 179.1146 | 156.5236 | 179.3663 | 141.5865 | 131.0944 | 165.3506 | 252.2405 | 165.9685 | 178.3348 | 228.5708 | 196.8584 | 177.5393 | 176.1887 | 203.5708 | 181.9685 | 292.8427 | 296.5551 | 178.0315 | 250.827 | 243.0786 | 246.6697 | 299.5708 | 125.4135 | 129.382 | 138.3618 | 147.0315 | 176.4449 | 177.7438 | 158.9685 | 100.5079 | 278.5708 | 257.827 | 265.0944 | 281.9528 | 156.7955 | 150.9371 | 137.6494 | 132.6494 | 88.9371 | 87.4135 | 84.3663 | 129.5079 | 233.8427 | 274.5551 | 239.9528 | 184.7753 | 156.6809 | 170.4449 | 160.7483 | 142.3191 | 194.0157 | 173.0629 | 219.7798 | 175.0315 | 162.7753 | 152.9214 | 178.0629 | 164.2517 | 178.4292 | 188.0472 | 244.0472 | 231.6652 | 172.8742 | 172.4764 | 244.0315 | 243.0629 | 218.4607 | 202.209 | 200.0629 | 291.5236 | 195.9528 | 184.2831 | 174.7123 | 235.618 | 141.4292 | 153.9843 | 153.7281 | 164.7169 | 206.173 | 165.3978 | 210.8696 | 159.0315 | 121.8584 | 125.0315 | 123.5079 | 133.5551 | 264.0472 | 256.5079 | 249.382 | 328.3348 | 253.3663 | 273.4764 | 243.7595 | 341.2719 | 148.5708 | 85.7123 | 68.0472 | 148.5708 | 127.3978 | 117.4764 | 109.3348 | 141.5393 | 132 | 133.9528 | 122.7595 | 128.3978 | 83.0315 | 78.9214 | 75.4292 | 103.6022 | 106.9843 | 105.827 | 112.8742 | 83.0629 | 265.791 | 238.0629 | 227.5393 | 295.2517 | 172.4764 | 201.4292 | 215.0629 | 273.9528 | 323.2877 | 228.1933 | 201.2562 | 315.0786
937 | 12.3014 | 172.9292 | 221.593 | 230.4336 | 198.9115 | 178.0177 | 178.1328 | 201.4779 | 160.0443 | 173.8761 | 138.0177 | 126.0089 | 177.0177 | 256.3893 | 172.0354 | 172.8584 | 234.593 | 216.0885 | 169.9557 | 149.8761 | 215.6461 | 182.4867 | 314.0974 | 317.7169 | 178.5221 | 235.2654 | 249.6018 | 239.2477 | 306.6018 | 144.6195 | 141.0354 | 125.0885 | 155.0885 | 178.4867 | 168.0531 | 151.4159 | 104.0354 | 255.3362 | 256.8938 | 260.5133 | 275.9203 | 159.9115 | 142.8938 | 138.0885 | 133.593 | 85.4336 | 91.9911 | 83.4159 | 128.4956 | 234.9203 | 251.3274 | 250.062 | 175.5753 | 148.531 | 161.3893 | 163.3805 | 138.8672 | 192.4956 | 176.062 | 227.9469 | 183.593 | 147.5221 | 152.9557 | 187.1151 | 141.9469 | 172.407 | 189.0354 | 232.9292 | 224.531 | 171.4159 | 166.938 | 237.9646 | 227.9026 | 208.3893 | 209.9026 | 201.0443 | 304.6284 | 186.3893 | 170.5398 | 174.1151 | 243.6372 | 151.5487 | 152.9823 | 137.9911 | 186.531 | 181.3805 | 190.1594 | 200.1151 | 148.9292 | 137.5575 | 120.4779 | 120.9823 | 136.0531 | 266.5487 | 252.9734 | 274.1505 | 340.5133 | 263.5133 | 280.5487 | 252.7258 | 327.2477 | 149.0443 | 78.5575 | 76.0974 | 149.0443 | 126.938 | 130.6018 | 128.0708 | 138.4956 | 134.0177 | 129.938 | 110.0354 | 132.4779 | 84.0266 | 84.5044 | 76.9734 | 93.469 | 108 | 107.4159 | 111.9203 | 79 | 263.1416 | 239.0443 | 225.5044 | 277.4867 | 180.5575 | 220.1239 | 208.9823 | 261.3628 | 322.8761 | 246.9911 | 219.0177 | 297.8938
938 | 12.3124 | 170.1672 | 232.5167 | 225.0669 | 194.1672 | 174.3161 | 190.4164 | 184.5319 | 177.4012 | 154.8328 | 136.8997 | 134.234 | 174.3495 | 245.0334 | 176 | 173.535 | 244.4833 | 208.4316 | 171.7173 | 149.4347 | 219.5988 | 177.8997 | 305.3647 | 330.6322 | 171.1824 | 211.5836 | 250.1489 | 225.9848 | 304.5653 | 141.465 | 141.8997 | 129.8328 | 158.7994 | 178.0334 | 157.9818 | 143.5502 | 104.8997 | 241.1337 | 266.1854 | 273.3678 | 250.465 | 143.7994 | 134.1003 | 152.1337 | 140.8997 | 84.2006 | 86.8663 | 80.7173 | 133.6839 | 231.6839 | 227.3495 | 246.6656 | 170.0486 | 158.7173 | 161.4012 | 156.7173 | 134.8511 | 195.1003 | 173.1824 | 225.1003 | 177.465 | 149.4833 | 157.3009 | 188.6322 | 155.1185 | 165.617 | 209.535 | 218.7994 | 222.3161 | 170.7842 | 174.4681 | 211.7811 | 211.8328 | 203.234 | 211.4012 | 200.3161 | 308.6656 | 177.617 | 179.6505 | 177.6991 | 248.1489 | 163.7173 | 153.0669 | 155.0851 | 172.4316 | 190.2189 | 201.7994 | 211.9666 | 149.7842 | 132.1155 | 113.3495 | 117.9666 | 140.9666 | 267.3495 | 262.9179 | 275.4985 | 337.8663 | 276.8845 | 271.5319 | 271.2827 | 300.0334 | 136.9483 | 90.6839 | 79.766 | 136.9483 | 124.6505 | 150.2675 | 126.6991 | 140.5836 | 117.9149 | 118.3495 | 116.5836 | 129.4833 | 87.5167 | 77.766 | 67.2827 | 92.0669 | 112.1337 | 96.4498 | 104.5502 | 80.0334 | 289.3344 | 237.2827 | 223.9331 | 296.6687 | 187 | 228.3161 | 214.234 | 252.7173 | 307.9666 | 233.5988 | 236.5015 | 288.0669
939 | 12.3234 | 172.884 | 245.6962 | 215.3618 | 198.471 | 174.2901 | 203.3412 | 177.3481 | 183.7099 | 146.8259 | 140.3481 | 137.2969 | 176.3481 | 225.8977 | 183.4061 | 177.8908 | 258.2832 | 193.058 | 163.7168 | 170.4641 | 204.8259 | 183.9931 | 300.2252 | 311.0069 | 170.4641 | 222.9931 | 249.5222 | 240 | 290.942 | 125.471 | 140.058 | 131.8771 | 154.058 | 183.2321 | 143.529 | 150.8191 | 94.5939 | 261.3993 | 270.1297 | 271.3038 | 263.0375 | 141.7031 | 144.4061 | 145.9488 | 129.0069 | 81.5359 | 86.1741 | 78.0069 | 145.3481 | 230.1229 | 218.2389 | 228.5359 | 164.4061 | 169.2321 | 155.5427 | 167.2321 | 123.8397 | 200.645 | 177.2252 | 222.7099 | 175.2252 | 161.1672 | 164.413 | 180.645 | 148.1366 | 167.942 | 214.4198 | 215.645 | 225.9931 | 183.2321 | 179.0069 | 208.9863 | 221.8122 | 223.2252 | 209.7748 | 210.8703 | 298.471 | 187.3481 | 179.2389 | 172.1741 | 239.587 | 170.529 | 148.1229 | 172.529 | 162.3481 | 210.471 | 199.1741 | 193.5427 | 156.413 | 134.7542 | 114.8191 | 126.5222 | 151.1092 | 253.4778 | 258.0717 | 234.5496 | 330.8259 | 264.1946 | 262.2321 | 252.3038 | 301.529 | 134.7542 | 100.2321 | 79.3481 | 134.7542 | 125.2969 | 134.1946 | 116.942 | 149.8771 | 107.3481 | 127.7542 | 114.7679 | 133.7611 | 98.0512 | 83.1672 | 73.6962 | 110.9283 | 117.058 | 103.4641 | 93.3038 | 76.7679 | 279.4267 | 229.413 | 227.2901 | 316.529 | 185.413 | 219.2969 | 231.0512 | 276.9863 | 301.7099 | 231.5222 | 243.0649 | 313.2764
940 | 12.3344 | 162.8805 | 226.6872 | 209.3304 | 179.5958 | 173.587 | 198.5958 | 185.7065 | 170.3392 | 152.6608 | 138.9631 | 121.0088 | 179.8348 | 216.1195 | 181.3392 | 159.8805 | 265.587 | 188.0457 | 168.8629 | 156.348 | 198.2109 | 174.1371 | 327.1564 | 301.3761 | 168.2566 | 258.065 | 253.6696 | 233.5044 | 291.6239 | 131.4956 | 137.2109 | 128.2566 | 154.4587 | 176.7153 | 159.1564 | 158.7065 | 91.7891 | 267.3392 | 239.761 | 244.8436 | 264.1828 | 148.7891 | 148.8348 | 139.7065 | 126.9543 | 85.3673 | 85.2109 | 73.2478 | 133.6784 | 228.3304 | 204.7979 | 249.6889 | 160.174 | 173 | 140.5413 | 171.5413 | 122.3216 | 197.0457 | 189 | 212.587 | 188.0826 | 170.9174 | 155.4218 | 174.3392 | 144.4042 | 171.3304 | 202.8348 | 214.2109 | 235.6239 | 188.0826 | 170.4587 | 245.0281 | 223.6327 | 231.7522 | 212.6151 | 222.2935 | 307.2021 | 190.2935 | 174.4587 | 186.9086 | 230.7153 | 164.5044 | 145.2478 | 152.9719 | 165.2935 | 201.8805 | 205.2478 | 188.826 | 143.0914 | 131.3023 | 132.4499 | 113.8893 | 149.6327 | 246.0826 | 250.5782 | 226.1564 | 318.2566 | 265.1933 | 265.4587 | 261.5694 | 328.5237 | 155.1195 | 97.5044 | 75.7979 | 155.1195 | 107.3849 | 118.413 | 128.9912 | 147.8805 | 109.2109 | 138.9174 | 118.0369 | 130.9631 | 100.5044 | 79.3849 | 87.7065 | 109.2197 | 107.7153 | 111.5413 | 90.6608 | 93.5694 | 254.4218 | 237.2021 | 229.8348 | 315.9174 | 190.4956 | 225.7434 | 242.1652 | 292.4587 | 297.5413 | 229.174 | 219.2197 | 320.4675
941 | 12.3454 | 156 | 214.2862 | 221.8575 | 178.3934 | 161.2497 | 185.768 | 195.1967 | 184.59 | 164.9823 | 136.8751 | 119.411 | 182.4287 | 236.8398 | 161.2674 | 145.8033 | 249.3746 | 187.4287 | 179.0177 | 157.3757 | 201.0895 | 176.1613 | 335.5713 | 300 | 174.3934 | 271.589 | 245.0177 | 218.0354 | 276.9459 | 144.7503 | 134 | 136.6077 | 162.3039 | 162.4641 | 172.5536 | 149.9282 | 93.8928 | 284.9116 | 245.9293 | 235.3215 | 257.8398 | 161.9646 | 162.5005 | 143.6608 | 138.0895 | 91.1249 | 88.0895 | 76 | 116.232 | 234.2144 | 225.7867 | 264.0354 | 168.1613 | 176.3215 | 139.8928 | 162.589 | 139.768 | 204.179 | 193.4287 | 217.411 | 186.232 | 168.3392 | 148.4464 | 173.6431 | 148.9282 | 172.232 | 208.1967 | 222.6254 | 231.4641 | 188.4464 | 179.9646 | 260.8928 | 214 | 232.8751 | 220.5713 | 225.5359 | 300.0531 | 198.5182 | 177.8751 | 183.7138 | 215.3569 | 166.7503 | 143.5713 | 142.0895 | 166.8751 | 197.2144 | 204.1249 | 194.1249 | 135.2144 | 127.9646 | 130.821 | 112.6077 | 141.6608 | 237.589 | 276.3757 | 248.411 | 308.8928 | 259.0884 | 254.4818 | 255.7492 | 333.8387 | 153.6961 | 84.8033 | 79.0718 | 153.6961 | 103.8575 | 129.0895 | 130.0354 | 121.0707 | 115.411 | 131.9105 | 114.0354 | 127.2144 | 82.2674 | 75.3039 | 87.7856 | 101.6431 | 100.6608 | 110.8928 | 100.2144 | 99.3746 | 284.537 | 245.5536 | 228.5536 | 308.9105 | 198.2144 | 221.2674 | 223.5171 | 276.4995 | 307.411 | 222.2144 | 209.9823 | 324.5005
942 | 12.3564 | 167.3176 | 236.8835 | 219.38 | 205.8447 | 170.2788 | 176.6824 | 175.5353 | 196.3412 | 168.4341 | 132.0777 | 128.6976 | 189.9612 | 234.9847 | 165.7129 | 158.6741 | 231.5118 | 198.9223 | 164.2482 | 160.5506 | 207.6976 | 183.7753 | 331.4341 | 302.8294 | 184.8682 | 249.2871 | 260.24 | 225.8447 | 307.007 | 148.1706 | 135.6976 | 144.9071 | 153.7212 | 161.3953 | 160.5506 | 135.9071 | 82.2482 | 291.4188 | 274.4494 | 259.5035 | 253.7212 | 163.7753 | 158.5894 | 141.6047 | 136.7753 | 85.7365 | 95.8294 | 74.8682 | 115.1318 | 236 | 241.6435 | 249.7753 | 176.9071 | 180.6976 | 137.8682 | 147.0777 | 135.2094 | 200.9847 | 198.6976 | 213.6824 | 189.6588 | 157.38 | 149.1318 | 179 | 140 | 166.6047 | 203.8141 | 235.3953 | 219.6435 | 197.62 | 181.7753 | 248.6824 | 233.8059 | 212.233 | 218.6976 | 216.9847 | 274.9847 | 180.9694 | 182.1318 | 183.7518 | 229.24 | 173 | 136.6047 | 153.2247 | 172.8294 | 190.5118 | 181.1941 | 186.4729 | 150.0153 | 122.4188 | 122.1318 | 118.0777 | 148.6588 | 237.9223 | 276.1553 | 261.6588 | 301.2094 | 242.3953 | 238.0777 | 238.6976 | 322.2635 | 151.5271 | 78.4341 | 88.5659 | 151.5271 | 120.0541 | 120.4188 | 115.2094 | 120.8447 | 117.3412 | 124.1706 | 124.1082 | 131.8294 | 82.7518 | 89.3565 | 88.8294 | 101.9071 | 100.3024 | 119.62 | 105.9612 | 90 | 271.5589 | 252.2247 | 219.38 | 292.6824 | 186.9847 | 203.0777 | 219.4494 | 264 | 312.1706 | 219.4729 | 206.6435 | 326.8141
943 | 12.3674 | 166.1711 | 227.4985 | 198.1239 | 206.4365 | 191.8289 | 190.5487 | 166.9853 | 191.4218 | 183.6107 | 142.1888 | 118.4365 | 173.9203 | 233.0944 | 185.0944 | 166.6401 | 221.531 | 198.0619 | 161.7817 | 153.3127 | 198.5929 | 183.6254 | 314.8112 | 314.2508 | 192.0944 | 248.4071 | 252.4513 | 229.3274 | 311.6695 | 133.2802 | 136.4218 | 146.7817 | 132.1711 | 165.7345 | 155.6254 | 140.6726 | 84.4071 | 261.9203 | 276.4838 | 249.9822 | 255.298 | 160.1563 | 158.7198 | 140.7345 | 127.3746 | 86.8909 | 87.0147 | 79.2036 | 148.3777 | 236.5782 | 230.2183 | 236.3274 | 171.8437 | 158.8731 | 142.7817 | 145.0472 | 140.4071 | 199.0944 | 191.9056 | 218.8761 | 191.6873 | 156.3599 | 152.3127 | 181.3127 | 136.531 | 178.4543 | 202.9381 | 229.9056 | 217.469 | 186.4985 | 168.3274 | 229.0147 | 238.0147 | 220.0178 | 228.0944 | 197.1711 | 297.9558 | 168.5162 | 180.1091 | 186.7964 | 237.6401 | 167.7964 | 153.6579 | 147.0147 | 172.6873 | 196.1416 | 184.5015 | 185.3127 | 167.5162 | 118.3599 | 119.531 | 130.5015 | 137.3893 | 246.8909 | 265.7345 | 249.8112 | 304.0944 | 242.6873 | 258.5959 | 220.3421 | 318.2183 | 151.531 | 80.8909 | 84.3746 | 151.531 | 122.9528 | 125.6107 | 115.2036 | 125.4838 | 114.7345 | 116.2183 | 127.0147 | 129.9528 | 85.2183 | 92.1091 | 89.2655 | 108.4071 | 103.0472 | 113.702 | 108.4218 | 84.2183 | 243.6254 | 241.3893 | 226.4543 | 304.236 | 172.9528 | 207.4071 | 216.8584 | 250.702 | 311.7345 | 217.1563 | 207.3599 | 322.469
944 | 12.3784 | 170.2188 | 223.4474 | 206.3044 | 196.8191 | 185.4193 | 191.0575 | 186.2188 | 191.5905 | 182.6003 | 146.324 | 128.3044 | 177.1235 | 232.5049 | 173.2861 | 155.5049 | 234.9425 | 203.0379 | 166 | 153.8191 | 198.0856 | 181.6858 | 311.857 | 305.0575 | 190.9144 | 258.3619 | 231.8668 | 213.5526 | 290.4095 | 129.9046 | 134.8191 | 138.0098 | 155.1137 | 164.6381 | 160.7714 | 162.3521 | 93.1809 | 268.0758 | 256.2384 | 232.1809 | 248.467 | 165.1332 | 167.4095 | 134.9144 | 124.5905 | 81.324 | 78.4095 | 83 | 150.7433 | 225.1907 | 227.1809 | 253.6186 | 144.4291 | 154.9902 | 148.7714 | 143.2763 | 150.3619 | 213.2665 | 191.3142 | 216.6003 | 187.6381 | 156.2763 | 153.4095 | 175.324 | 142.2665 | 176.6003 | 225.1235 | 217.5049 | 221.1809 | 188.9425 | 172.0379 | 234.5807 | 213.467 | 235.4572 | 226.324 | 191.7714 | 300.7433 | 171.6381 | 211.6564 | 172.9621 | 237.7237 | 153.3717 | 170.3619 | 156.1235 | 159.1907 | 191.4193 | 211.5807 | 187.5905 | 158.8289 | 118.2763 | 121.1332 | 129.2384 | 130.7237 | 275.5709 | 257.5526 | 258.0758 | 315.3142 | 259.3044 | 278 | 252.0562 | 309.2763 | 143.6858 | 84.1809 | 83.9523 | 143.6858 | 120 | 131.0954 | 125.4951 | 137.7139 | 125.4474 | 113.7714 | 123.7237 | 116.9621 | 88.0856 | 89.4095 | 82.0954 | 115.4095 | 112.4951 | 108.7237 | 93.8289 | 90.6283 | 252.9046 | 240.0379 | 242.9046 | 321.9523 | 170.5905 | 229.1711 | 210.9523 | 265.209 | 287.6101 | 254.6088 | 212.5905 | 320.8668
945 | 12.3894 | 177.3972 | 224.5752 | 219 | 194.7945 | 176 | 177.589 | 181.3421 | 164.2733 | 161.3421 | 155.4661 | 124.1229 | 179.3559 | 247.2882 | 153.7669 | 169.0826 | 238.1642 | 193.1229 | 181.6716 | 148.178 | 190.3421 | 166.5477 | 311.9725 | 294.6028 | 183.589 | 238.3008 | 238.0413 | 210.0138 | 316.5212 | 149.0689 | 135.2055 | 135.6303 | 154.9036 | 166.0138 | 166.2193 | 157.9311 | 98.2193 | 288.6303 | 247.8083 | 249.2744 | 255.0826 | 156.5477 | 156.1504 | 122.7669 | 139.4661 | 86.8496 | 72.5752 | 71.5477 | 142.0275 | 221.2193 | 214.7394 | 239.8623 | 152.5212 | 164 | 145.7807 | 129.1504 | 128.4926 | 215.3835 | 191.3835 | 217.6441 | 178.1642 | 163.2468 | 153.6028 | 183.8634 | 139.5614 | 183.0689 | 211.0815 | 229.8771 | 235.2606 | 174.6843 | 154.8898 | 242.1917 | 205.0138 | 239.6303 | 215.5752 | 201.4386 | 311.9185 | 189.8909 | 202.4512 | 180.4661 | 243.411 | 159.8634 | 161.1504 | 154.7394 | 164.2606 | 178.9862 | 207.1366 | 205.4799 | 148.3972 | 118.6165 | 128.822 | 119 | 159.3157 | 282.5477 | 263.0551 | 265.3697 | 343.713 | 276.822 | 278.6028 | 280.9862 | 310.822 | 149.0413 | 82.589 | 85.3972 | 149.0413 | 128.4386 | 125.1917 | 123.3697 | 154.2193 | 119.9449 | 116.2055 | 116.1504 | 103.3697 | 87.5752 | 85.3835 | 76.1917 | 105.3559 | 100.1229 | 104.7669 | 96.6579 | 100.411 | 270.8634 | 246.3972 | 257.8496 | 298.6843 | 179.4386 | 213.6843 | 207.5752 | 270.5477 | 295.9185 | 267.3421 | 206.3697 | 321.0138
946 | 12.4004 | 171.4646 | 221.615 | 223.9204 | 194.615 | 171.0796 | 159.3938 | 171.1549 | 161.3761 | 164.6859 | 148.3938 | 120.3806 | 177.9204 | 252.6194 | 169.9116 | 166.2389 | 234.2301 | 194.916 | 175.3938 | 146.8451 | 182.615 | 163.9204 | 303.0796 | 315.2965 | 182.615 | 230.7655 | 254.4558 | 222.4558 | 306.938 | 147.3141 | 132.9248 | 132.0044 | 153.916 | 182.7611 | 169.615 | 159.6859 | 91.7743 | 286.2345 | 237.9292 | 252.0044 | 245.4734 | 147.1549 | 148.385 | 136.4514 | 143.4646 | 87.2345 | 76.3806 | 80.6062 | 142.3097 | 236.916 | 206 | 236.6062 | 171.8451 | 165.2301 | 141.1549 | 144.1417 | 133.9116 | 193.9336 | 193.9204 | 231.3806 | 184.0708 | 151.5486 | 143.5442 | 191.7699 | 144.4558 | 167.1681 | 203.8407 | 233.0044 | 243.385 | 165.3806 | 161.7567 | 235.4646 | 209.4602 | 232.9292 | 212 | 203.3097 | 306.8583 | 191.9292 | 193.3009 | 185.6947 | 237.0044 | 173.3053 | 157.0752 | 163.8363 | 167.4646 | 182.5354 | 201.6903 | 197.9336 | 145.5398 | 124.6903 | 136.9204 | 117.7699 | 160.6239 | 255.9336 | 275.3053 | 264.0752 | 328.6327 | 283.0752 | 272.8495 | 277.1549 | 334.9116 | 150.6947 | 79.7699 | 72.084 | 150.6947 | 134 | 134.4558 | 111.0044 | 155.7699 | 116.3053 | 121.9204 | 109.615 | 121.1417 | 84 | 70.6991 | 67.0044 | 100.2301 | 107.4514 | 102.4602 | 106.2301 | 87.8539 | 258.4734 | 236.7743 | 241.6283 | 271.5442 | 184.385 | 213.6062 | 211.3806 | 286.3717 | 295.7787 | 247.3141 | 217.9912 | 300.2433
947 | 12.4114 | 161.7266 | 214.472 | 223.236 | 208.1741 | 166.7453 | 167.8201 | 175.0187 | 182.9194 | 160.7079 | 137.3727 | 133.7827 | 182.882 | 244.8633 | 171.0806 | 151.9813 | 226.2173 | 201.7453 | 158.0993 | 165.5654 | 177.354 | 174.528 | 311.2921 | 302.2792 | 181.118 | 229.9813 | 240.9252 | 252.2115 | 285.6086 | 143.1367 | 139.1554 | 129.5093 | 143.1799 | 180.0806 | 166.236 | 173.9007 | 96.6647 | 271.3353 | 236.646 | 248.882 | 250.8201 | 164.1928 | 154.2734 | 139.8446 | 140 | 86.1367 | 90.4101 | 83.472 | 126.1986 | 246.882 | 188.4346 | 245.7453 | 156.6892 | 162.8633 | 158.1928 | 145.4533 | 133.8259 | 184.5093 | 193.236 | 229.7266 | 169.6705 | 140.6273 | 140.1367 | 172.8072 | 150.3727 | 185.5035 | 220.0374 | 220.472 | 240.4907 | 177.528 | 164.7079 | 240.7827 | 221.0374 | 226 | 198.1986 | 202.882 | 305.5467 | 201.3108 | 193.472 | 199.9381 | 232.6273 | 167.2173 | 144.5712 | 160.5712 | 152.7079 | 197.9194 | 205.2547 | 204.1928 | 150.9007 | 124.4907 | 136.236 | 120.764 | 140.9626 | 240.8633 | 272.9813 | 249.6892 | 342.2489 | 247.3598 | 275.9007 | 264.0806 | 326.0432 | 139.2173 | 75.236 | 77.1741 | 139.2173 | 115.8072 | 131.5899 | 111.646 | 145.5899 | 126.528 | 118.7266 | 103.0993 | 120.5712 | 75.8446 | 85.5841 | 73.2921 | 104.764 | 97.0432 | 114.6647 | 101.354 | 80.2547 | 266.9568 | 236.646 | 221.354 | 295.1122 | 170.1986 | 210.8259 | 200.9439 | 267.7511 | 283.7453 | 237.4907 | 206.6705 | 309.8388
948 | 12.4224 | 170.7926 | 217.0359 | 227.3963 | 206.4056 | 164.0811 | 183.4774 | 172.2433 | 184.2433 | 163.5944 | 129.3245 | 134.5226 | 175.0452 | 235.9641 | 183.1888 | 164.3511 | 212.0452 | 200.3604 | 166.7926 | 160.0093 | 201.504 | 170.6848 | 314.1622 | 284.4415 | 176.8019 | 234.0359 | 240.5133 | 249.3697 | 286.8378 | 137.9641 | 142.0811 | 132.9189 | 133.2793 | 183.8737 | 158.883 | 144.0997 | 104.9189 | 242.613 | 255.3511 | 242.3245 | 266.4774 | 150.0545 | 153.5226 | 145.8737 | 142.5585 | 95.0359 | 80.0093 | 78.3604 | 126.9548 | 229.4508 | 201.0266 | 220.0545 | 161.0718 | 167.3963 | 150.4508 | 139.2793 | 120.883 | 190.4774 | 181.4056 | 222.1622 | 176.1888 | 149.3152 | 149.6755 | 172.8737 | 139.766 | 205.3604 | 203.613 | 217.2793 | 249.8737 | 171.1263 | 155.4415 | 230.0093 | 225.7207 | 234.3152 | 199.5944 | 206.5585 | 310.4415 | 195.0093 | 174.9282 | 204.883 | 230.4415 | 162 | 154.5492 | 148.8019 | 154.3152 | 194.1263 | 186.8112 | 209.883 | 170.9907 | 124.9189 | 128.883 | 120.4415 | 136.2793 | 259.4681 | 265.5226 | 246.3963 | 341.5319 | 246.7474 | 269.766 | 251.2433 | 321.9548 | 141.0359 | 71.0811 | 70.9282 | 141.0359 | 128.6662 | 124.7207 | 122.0359 | 145.117 | 132.2793 | 125.234 | 101.5585 | 131.1888 | 75.4774 | 91.3245 | 80.6396 | 102.5226 | 88.4774 | 119.0811 | 92.883 | 84.1981 | 260.1715 | 245.117 | 223.7567 | 305.3245 | 184.387 | 217.0718 | 208.6303 | 265.27 | 288.7567 | 222.5678 | 213.1888 | 297.7752
949 | 12.4334 | 174.7404 | 186.4483 | 217.3097 | 198.3923 | 170.1711 | 177.5251 | 175.5192 | 170.5693 | 165.0885 | 150.4248 | 125.4808 | 165.4366 | 232 | 180.9616 | 170.3923 | 208.3038 | 191.5251 | 162.9174 | 161.4307 | 220.9115 | 167.3038 | 295.702 | 296.6903 | 173.0442 | 255.6018 | 231.3539 | 238.6962 | 289.6519 | 152.2537 | 139.6962 | 120.3097 | 139.2154 | 166.5309 | 167.7345 | 153.0325 | 95.5693 | 272.3747 | 245.7463 | 250.3865 | 246.1828 | 160.121 | 152.9558 | 140.2655 | 128.3539 | 95.7404 | 84.0384 | 75.3923 | 128.0885 | 220.3038 | 229.6461 | 220.3422 | 187.9057 | 157.3097 | 144.5634 | 149.1269 | 125.823 | 180.6135 | 191.6461 | 214.1327 | 193.5192 | 155.9558 | 150.7404 | 175.0885 | 148.3422 | 182.8347 | 200.7788 | 202.3539 | 235.7905 | 194.3364 | 161.1711 | 225.5634 | 216.5251 | 235.7404 | 207.6077 | 217.1269 | 286.1828 | 176.2212 | 167 | 187.0059 | 227.0442 | 169.823 | 159.1327 | 160.6903 | 166.1711 | 183.4366 | 197.5133 | 199.8289 | 149.3598 | 126 | 133.823 | 117.6962 | 135.6962 | 252.0943 | 238.879 | 255.8673 | 313.702 | 240.7463 | 249.6578 | 270.1653 | 310.702 | 148.9115 | 77.823 | 74.0826 | 148.9115 | 127.6578 | 114.2212 | 130.5634 | 138.2212 | 131.6962 | 117.3097 | 125.1653 | 137.4366 | 88.4307 | 76.5693 | 78.3923 | 104.5634 | 108.6018 | 121.9115 | 97.823 | 82.0885 | 257.4749 | 230.3982 | 234.823 | 304.2596 | 176.1386 | 213.2655 | 216.6962 | 287.3865 | 280.9174 | 209.1327 | 205.0943 | 301.9057
950 | 12.4444 | 163.0368 | 213.1663 | 215.9779 | 191.0221 | 169.3505 | 163.0368 | 174.3505 | 171.6421 | 163 | 156.6937 | 132.6274 | 173.6274 | 230.0074 | 158.059 | 170.3284 | 205.6789 | 201.6126 | 169.62 | 154.38 | 207.059 | 170.6568 | 295.6347 | 300.6789 | 179.3063 | 262.3432 | 238.941 | 225.38 | 280.7011 | 146.059 | 137.0074 | 115.6568 | 162.5905 | 171.9337 | 152.0811 | 151.3947 | 90.6642 | 269.0958 | 263.2326 | 256.3358 | 249.2768 | 143.7748 | 150.0147 | 139.3137 | 147.2326 | 97.3211 | 92.9926 | 80.6421 | 121.3505 | 222.9926 | 222.059 | 237.2989 | 197.3358 | 149.3358 | 155.6347 | 151.3432 | 120.7011 | 190.6053 | 179.4095 | 216.3137 | 180.3947 | 158.3284 | 155.6421 | 166.3579 | 153.3432 | 179.6347 | 203.3432 | 195.3284 | 231.6421 | 187.4169 | 161.6789 | 232.6495 | 204.6937 | 216.7305 | 218.2989 | 219.3432 | 308.5389 | 198.2326 | 178.9558 | 182.9853 | 250.5758 | 160.0516 | 145.3726 | 141.4316 | 179.9632 | 198.2695 | 211.6568 | 182.0516 | 127.6863 | 117.3653 | 134.6789 | 113.6789 | 130.3505 | 253.2916 | 251.24 | 253.0221 | 320.941 | 249.5979 | 247.3137 | 272.0368 | 299.3432 | 148.3432 | 87.3137 | 84.6495 | 148.3432 | 115.3505 | 114.3137 | 127.6863 | 134.3284 | 130.3358 | 118.6347 | 123.7158 | 132.3432 | 86.6937 | 92.2547 | 98.2547 | 108.3284 | 102.7232 | 112.7084 | 98.0147 | 88.6347 | 260.6716 | 248.8968 | 223.059 | 287.4021 | 173.6347 | 216.9632 | 218.6568 | 269.4242 | 320.8305 | 216.6274 | 221.5684 | 304.6937
951 | 12.4554 | 181.001 | 254.2655 | 210.882 | 198.1475 | 173.765 | 178.295 | 176.7355 | 164.8525 | 163 | 131.352 | 129.2055 | 183.7355 | 231.706 | 160.824 | 170.3235 | 222.942 | 211.706 | 148.2634 | 148.6765 | 197.647 | 180.118 | 328.4131 | 311.177 | 197.883 | 237.999 | 240.9115 | 219 | 295.6185 | 138.6765 | 133.294 | 126.471 | 180.4415 | 191.824 | 149.7945 | 159.9125 | 83.5585 | 251.941 | 248.5869 | 258.0295 | 263.0885 | 155.3835 | 147.3235 | 148.765 | 138.6755 | 100.353 | 87.235 | 92.118 | 118.3235 | 241.589 | 249.8546 | 244.0295 | 168.5869 | 159.1475 | 145.7935 | 158.118 | 129.53 | 192.9115 | 181.177 | 227.118 | 176.059 | 163.7355 | 152.235 | 180.589 | 141.8525 | 175.2055 | 197.2645 | 236.5901 | 224.8525 | 183.4415 | 168.118 | 220.117 | 205.059 | 218.824 | 220.294 | 205.823 | 301.646 | 188.3225 | 165.3815 | 184.3235 | 229.8514 | 148.2645 | 159.6185 | 129.353 | 168.0875 | 198.882 | 210.9705 | 180.412 | 137.177 | 125.8535 | 132.3235 | 110.647 | 122.588 | 256.9705 | 251.1465 | 242.5585 | 323.588 | 277.2655 | 248.647 | 274.4415 | 310.8535 | 157.1475 | 78.4995 | 84.294 | 157.1475 | 113 | 124.4415 | 129.059 | 128.235 | 114.4405 | 110.823 | 127.824 | 128.294 | 83.6765 | 84.7345 | 75.8514 | 90.7345 | 107.8535 | 106.3235 | 94.647 | 92.3235 | 270.824 | 261.647 | 238.6775 | 301.6775 | 186.7945 | 230.7945 | 214.588 | 267.1475 | 310.1749 | 222.6765 | 226.882 | 294.9115
952 | 12.4664 | 170.6474 | 229.938 | 200.1121 | 183.7138 | 154.2699 | 162.5147 | 169.3569 | 173.7758 | 155.4233 | 126.5103 | 130.5103 | 174.9793 | 222.6681 | 166 | 156.2242 | 216.8466 | 204.0457 | 152.9086 | 167.5974 | 179.7802 | 176.4233 | 303.4276 | 296.3362 | 196.7345 | 218.7345 | 246.2655 | 212.8009 | 271.938 | 132.1121 | 131.3112 | 122.7345 | 159.8923 | 161.1828 | 148.4897 | 151.469 | 73.1121 | 261.7095 | 234.3112 | 256.9336 | 264.6224 | 158.3569 | 148.3776 | 145.1121 | 142.4646 | 101.6888 | 86.0664 | 84.2905 | 128.3319 | 240.3569 | 249.7802 | 247.0664 | 166.0207 | 179.1534 | 150.7095 | 143.4026 | 140.8215 | 192.7552 | 182.1785 | 224.8009 | 171.1121 | 161.1785 | 145.556 | 200.7095 | 156.975 | 185.4646 | 207.3983 | 226.3819 | 220 | 162.8923 | 174.0664 | 246.062 | 224.9086 | 226.0664 | 252.062 | 206.8879 | 294.1328 | 179.0664 | 167.7095 | 184 | 218.1328 | 141.1785 | 160.7345 | 142.3983 | 165.5103 | 183.9793 | 189.3362 | 198.1534 | 146.444 | 126.4897 | 116.1578 | 105.1785 | 135.8422 | 260.8215 | 269.1077 | 263.7965 | 313.4233 | 295.6431 | 253.5103 | 268.3569 | 318.3776 | 152.3569 | 76.444 | 64.4026 | 152.3569 | 117.8215 | 125.2448 | 130.3112 | 134.6431 | 120.7758 | 120.8422 | 115.0914 | 104.9586 | 82.6224 | 73.2448 | 83.4189 | 97.1534 | 120.1991 | 108.7552 | 118.7965 | 94.7552 | 270.4897 | 236.2035 | 248.6224 | 313 | 176.5354 | 232.9336 | 194.7802 | 285.7758 | 288.4897 | 211.9793 | 219.556 | 329.8836
953 | 12.4774 | 161 | 237.4347 | 229.9509 | 179.2065 | 151.7119 | 171.3314 | 173.413 | 160.3697 | 154.8043 | 133.9076 | 125.9892 | 177.0108 | 230.6195 | 158.9892 | 173.8368 | 213.5054 | 208.413 | 161.7011 | 159.174 | 183.2173 | 189.826 | 290.9076 | 295.413 | 179.6794 | 239.5379 | 240.1848 | 225.4238 | 276.6303 | 128.2989 | 143.6195 | 133.0217 | 153.9076 | 171.6412 | 154.413 | 143.2989 | 97.3423 | 276.1141 | 243.1141 | 258.8043 | 283.6303 | 161.0108 | 142.6902 | 150.413 | 153.9076 | 100.5978 | 80.6902 | 71.2881 | 123.8859 | 227.587 | 230.7827 | 234.6794 | 168.8967 | 177.587 | 167.9184 | 151.826 | 123.3697 | 212.9293 | 176.4946 | 216.3913 | 168.7011 | 147.0816 | 146.8043 | 196.1848 | 149.8751 | 184.9892 | 184.9567 | 238.239 | 240.3314 | 157.6087 | 167.2881 | 258.8967 | 230.8967 | 203.8642 | 209.5112 | 197.3805 | 324.7444 | 190.5163 | 169.4946 | 184 | 220.7011 | 143.9076 | 159.1033 | 139.587 | 153.9783 | 190.9184 | 188.413 | 198.6902 | 138.1848 | 119.7935 | 130.7336 | 124.7336 | 144.4022 | 288.239 | 296.826 | 254.6686 | 333.8368 | 285.2772 | 246.8859 | 249.2772 | 333.0217 | 133.2772 | 92.0217 | 77.7336 | 133.2772 | 131.2173 | 104.3697 | 124.3913 | 115.8642 | 129.1033 | 120.9892 | 133.6412 | 104.8152 | 89.7119 | 88.826 | 81.3805 | 91.3805 | 112.4837 | 102.9892 | 116.6794 | 95.2989 | 265.8967 | 235.5163 | 253.6087 | 302.4837 | 189.6303 | 236.2065 | 191.2065 | 286.3913 | 297.2173 | 217.5163 | 220.8043 | 330.174
954 | 12.4884 | 170.9872 | 212.0447 | 227.5924 | 203.828 | 177.1146 | 155.0319 | 170.5796 | 173.4012 | 179.5414 | 123.8726 | 133.7006 | 177.8599 | 231.0064 | 151.7198 | 177.5796 | 216.4267 | 196.3058 | 174.8407 | 155.5669 | 189.4267 | 179.879 | 303.7006 | 317.5477 | 177.5669 | 230.7389 | 253.8344 | 215.5924 | 286.4267 | 115.1593 | 138.2994 | 131.8663 | 157.4267 | 157.3185 | 172.2675 | 143.7134 | 99.7261 | 245.758 | 258.414 | 248.586 | 287.0064 | 162.5733 | 151.414 | 134.7389 | 140.3058 | 97.1465 | 78.7134 | 86.5477 | 125.707 | 231.1337 | 230.2802 | 255.3949 | 163.7198 | 161.1593 | 149.4586 | 146.8726 | 124.9872 | 201.0255 | 190.6942 | 209.0064 | 158.2994 | 153.414 | 143.0064 | 159.1848 | 136.5796 | 177.0064 | 189.4076 | 225.4586 | 229.0255 | 171.414 | 164.7134 | 250.8663 | 208.5988 | 203.9872 | 209.2547 | 203.414 | 304.8982 | 194.2866 | 192.2547 | 196.8407 | 226.707 | 170.2547 | 154.293 | 133.1465 | 155.8471 | 183.8726 | 194.1401 | 193.1465 | 137.5669 | 114.4331 | 134.293 | 117.5924 | 129.3058 | 274.7453 | 234.0893 | 260.9809 | 341.1465 | 262.5924 | 247.9936 | 257.9809 | 336.1465 | 135.5605 | 73.0319 | 82.0064 | 135.5605 | 116.7389 | 111.6942 | 124.8535 | 131.6815 | 120.7261 | 124.4204 | 118.6051 | 113.9936 | 72.3121 | 83.8726 | 76.7134 | 103.8344 | 110.8535 | 100 | 102.4395 | 89.293 | 274.2739 | 213.6051 | 256.7134 | 291.5796 | 202.9936 | 238.7134 | 225.1018 | 276.8663 | 276.3185 | 226.2802 | 203.4523 | 298.7453
955 | 12.4994 | 163.3894 | 200.177 | 223.9027 | 216.6283 | 167.8584 | 163.1416 | 181.062 | 163.8584 | 169.407 | 121.177 | 130.0177 | 164.6637 | 233.354 | 156.531 | 177.177 | 221.354 | 200.885 | 174.9203 | 162.0797 | 184.9203 | 183.1593 | 319.6106 | 314.115 | 173.1947 | 255.6549 | 241.407 | 230.77 | 275.3894 | 135.3983 | 125.292 | 128.2743 | 146.3894 | 173.8496 | 160.584 | 145.4513 | 101.0797 | 263.9293 | 270.9823 | 267.947 | 259.6017 | 169.2567 | 163.2567 | 138.6106 | 129.646 | 94.5487 | 73.9203 | 84.5663 | 112.0354 | 234 | 237.0797 | 243.5044 | 171.4337 | 160.354 | 147.2567 | 152.1593 | 128.2743 | 200.2567 | 181.761 | 235.301 | 146.7433 | 156.5487 | 146.8053 | 178.6549 | 129.646 | 175 | 178.584 | 205.115 | 223.177 | 176 | 157.7433 | 226.23 | 223.2213 | 206.5487 | 223.354 | 216.708 | 313.77 | 184.5663 | 179.5044 | 174.4247 | 236.2567 | 166.938 | 156.354 | 129.0973 | 151.7433 | 197.1416 | 185.5663 | 170.9557 | 133.1947 | 110.0973 | 126.9203 | 108.823 | 123 | 245.407 | 262.6019 | 233.6194 | 340.7257 | 274.1416 | 268.1416 | 261.0973 | 327.7433 | 130.292 | 72.4337 | 79.2743 | 130.292 | 103.9203 | 121.6283 | 126 | 124.584 | 119.9027 | 105.9557 | 102.1947 | 126.885 | 74.885 | 83.354 | 83.531 | 92.8584 | 117.8053 | 96.3717 | 94.646 | 90.6283 | 256.9557 | 226.2213 | 257 | 292.6283 | 198.469 | 215.053 | 228.5663 | 262.3894 | 279.7876 | 208.407 | 214.1416 | 284.646
956 | 12.5104 | 174.4971 | 206.1657 | 226.4759 | 205.4547 | 161.738 | 184.7591 | 189.6898 | 171.0212 | 172.3314 | 133.0693 | 144.711 | 175.4971 | 236.4759 | 169.3314 | 173.5723 | 233.3314 | 188.027 | 161.9307 | 152.1927 | 205.8767 | 162.6474 | 293.7438 | 287.8613 | 174.6898 | 263.5723 | 239.9037 | 239 | 264.3584 | 136.8825 | 132.3314 | 136.1175 | 147.1657 | 169.2409 | 145.1445 | 146 | 86.027 | 267.8825 | 237.1022 | 262.2409 | 280.2562 | 157.9788 | 141.6474 | 142.262 | 130.2139 | 100.6416 | 74.9518 | 78.0482 | 114.8555 | 234 | 227.9307 | 237.9518 | 152.8613 | 159.0482 | 144.0963 | 152.3102 | 132.4277 | 202.262 | 174.5241 | 194.343 | 178.684 | 154.5241 | 145.3102 | 166.6474 | 138.3314 | 185.3314 | 180.1175 | 222.4007 | 219.5723 | 166.4066 | 175.6628 | 226.1175 | 226.0963 | 222.235 | 222.7861 | 225.9037 | 314.6204 | 173.6204 | 167.3102 | 161.7861 | 245.6416 | 173.0693 | 156.5241 | 143.4971 | 182.946 | 199.5723 | 180.5241 | 177.7591 | 143.5453 | 112.6898 | 131.6416 | 112.4277 | 143.6628 | 266.7803 | 243.4124 | 228.3796 | 326.2409 | 271.4066 | 245.4817 | 261.4759 | 316.1445 | 145.4489 | 77.4759 | 84.1657 | 145.4489 | 103.4759 | 128.1657 | 128.9518 | 108.4066 | 111.4066 | 120.8767 | 103.6898 | 121.4066 | 87.1175 | 80.5723 | 73.4547 | 99.2832 | 107.4547 | 101.6416 | 109.235 | 89.0482 | 258.5934 | 220.2409 | 230.4336 | 294 | 176.0752 | 203.7861 | 236.0693 | 274.235 | 270.2409 | 214.2832 | 213.6204 | 308.0905
957 | 12.5214 | 185.2517 | 223.005 | 222.4985 | 191.2468 | 166.5015 | 170.4936 | 205.2547 | 166.2468 | 174.4995 | 130.998 | 139.7463 | 162.2438 | 233.2488 | 162.4965 | 147.2418 | 230.998 | 190.2527 | 173.7552 | 141.2478 | 203.4965 | 178.0079 | 296.5045 | 272.4975 | 163.996 | 252.2468 | 229.2458 | 236.7493 | 270.2527 | 124.2468 | 124.7463 | 129.997 | 137.7463 | 157.998 | 148.7522 | 146.7502 | 82.2507 | 260.4985 | 253.2597 | 238.2438 | 265.4916 | 147.7483 | 151.0059 | 158.5054 | 118.996 | 91.7463 | 77.5005 | 81.5015 | 123.2517 | 239.2517 | 220.2488 | 232.2478 | 154.7532 | 164.7522 | 139.7493 | 143.4975 | 146.7542 | 202.7502 | 160.4955 | 215.0129 | 220.2597 | 154 | 155.2537 | 173.005 | 143.5005 | 183.7483 | 206.2577 | 225.4985 | 236.0059 | 175.004 | 170.996 | 216.996 | 213.4965 | 211.4946 | 211.4965 | 205.4926 | 315.7512 | 195.0079 | 186.2567 | 168.5025 | 241.2478 | 170.2478 | 159.7512 | 161.7542 | 180.7453 | 216.7562 | 210.0099 | 200.2557 | 149.5005 | 103.4965 | 127.998 | 130.5054 | 134.4946 | 279.2507 | 252.2577 | 239.2527 | 302.9941 | 267.2498 | 236.5005 | 237.9921 | 325.7542 | 133.9941 | 75.7493 | 77.7473 | 133.9941 | 111.5025 | 107.4926 | 125.4985 | 112.5025 | 104.999 | 122.2478 | 107.2507 | 128.5035 | 74.995 | 86.5025 | 84.7552 | 104.7502 | 97.7483 | 120.5054 | 111.2488 | 88.7502 | 262 | 214.2498 | 228.5025 | 318.7582 | 170.5005 | 212.7532 | 218.2428 | 275.4985 | 280.7552 | 231.7542 | 207.999 | 299.7443
958 | 12.5324 | 179.3746 | 237.1505 | 219.4749 | 170.4616 | 161.1372 | 161.7124 | 168.0604 | 175.9632 | 168.6622 | 135.8628 | 129.1372 | 146.087 | 250.3009 | 143.7492 | 147.3879 | 223.6622 | 201.3879 | 191.2006 | 164.9263 | 190.087 | 188.2876 | 317.0133 | 283.7257 | 187.4514 | 232.2242 | 233.3879 | 239.0502 | 276.8127 | 137.0133 | 133.2006 | 127.7625 | 127.8997 | 174.3009 | 162.4381 | 143.9498 | 93.6755 | 246.7994 | 275.2006 | 250.3009 | 266.913 | 139.8997 | 157.7625 | 146.4616 | 121.8628 | 101.7257 | 82.5752 | 70.7994 | 134.1505 | 260.826 | 212.1372 | 249.826 | 175.5384 | 152.5118 | 158.0635 | 137.9498 | 136.5118 | 201.4749 | 158.2876 | 215.7994 | 185.0102 | 147.8997 | 155.9498 | 181.8127 | 135.6121 | 192.6755 | 206.3746 | 214.8495 | 243.5251 | 160.6991 | 155.5619 | 251.8895 | 200.8495 | 236.5015 | 200.3746 | 231.5517 | 325.3879 | 197.6622 | 183.087 | 174.0502 | 225.2743 | 174.1003 | 153.3746 | 140.0737 | 172.1873 | 203.9365 | 208.5619 | 196.087 | 155.3378 | 108.3879 | 119.8997 | 123.7994 | 121.3746 | 292.9632 | 242.4616 | 236.6622 | 316.0635 | 274.6254 | 282.7523 | 235.3378 | 317.0368 | 153.1638 | 67.3746 | 84.1505 | 153.1638 | 110.9498 | 127.4514 | 117.1372 | 121.8628 | 124.5885 | 100.9365 | 109.5251 | 119.0368 | 90.5885 | 101.9632 | 80.8495 | 93.5619 | 105.1505 | 118.3746 | 110 | 82.8997 | 252.087 | 234.5885 | 221.087 | 304.1238 | 183.2006 | 208.3746 | 198.0368 | 291.5384 | 299.7257 | 219.2242 | 204.7124 | 304.6755
959 | 12.5434 | 184.7483 | 212.881 | 226.7483 | 188.2449 | 166.7483 | 176.4966 | 160.4238 | 173.8014 | 164.6755 | 126.3776 | 131.649 | 160.8211 | 231.2055 | 171.5428 | 173.2449 | 233.6224 | 205.5497 | 171.7551 | 142.7816 | 214.8939 | 182.8014 | 306.5034 | 301.1721 | 165.7816 | 251.0197 | 216.6293 | 228.3776 | 250.1061 | 138.9007 | 138.5497 | 134.9735 | 142.2714 | 184.649 | 153.6027 | 153.8476 | 93.1259 | 253.8476 | 248.0068 | 259.0993 | 236.6823 | 165.8939 | 163.4238 | 162.6952 | 137.9469 | 98.2517 | 77.8014 | 81.7218 | 141.649 | 225.1592 | 216.9735 | 238.1789 | 170.9272 | 163.4966 | 144.6293 | 140.8741 | 135.0993 | 193.2517 | 166.7483 | 202.702 | 157.0531 | 159.1721 | 151.1259 | 184.5497 | 140.7483 | 178.9537 | 211.7483 | 216.649 | 255.6224 | 168.9469 | 166.7218 | 233.0068 | 227.4435 | 217.3313 | 201.8741 | 206.3578 | 306.3048 | 200.649 | 166.8279 | 169.5762 | 238.0463 | 165.9272 | 146.351 | 142.8476 | 161.702 | 203.4238 | 197.2517 | 179.8279 | 139.9537 | 124.1721 | 132.7218 | 118.4503 | 130.6224 | 259.8082 | 257.1456 | 238.8741 | 329.7483 | 270.8014 | 272.2055 | 234.6755 | 309.9007 | 140.8544 | 77.3973 | 83.9007 | 140.8544 | 127.8211 | 130.5762 | 109.5762 | 120.9007 | 127.1259 | 125.9932 | 96.0531 | 125.0728 | 72.2055 | 90.5034 | 70.2517 | 114.7945 | 108 | 127.6224 | 107.6755 | 78.6755 | 278.4435 | 239.4503 | 224.1986 | 286.9272 | 185.4503 | 222.2714 | 205.6224 | 269.1061 | 300.9007 | 214 | 220.2714 | 311.8741
960 | 12.5544 | 165.7477 | 186.8962 | 226.6386 | 202.0841 | 180.8068 | 180.2129 | 175.3811 | 192.4651 | 169.5098 | 137.1682 | 122.7674 | 177.0197 | 221.6386 | 170.7674 | 170.5546 | 205.5152 | 194.9803 | 162.6386 | 137.9356 | 200.1735 | 199.8909 | 305.1485 | 291.6189 | 178.2523 | 243.8318 | 223.5939 | 221.8515 | 275.0591 | 121.4705 | 124.8318 | 130.703 | 135.9803 | 156.8765 | 150.7871 | 154.6386 | 118.7621 | 277.4651 | 230.3417 | 244.2576 | 267.1432 | 174 | 142.9606 | 155.6189 | 130.1932 | 87.3417 | 88.5939 | 82.8515 | 134.3417 | 209.8515 | 202.4705 | 245.5939 | 176.6583 | 175.0841 | 136.6386 | 140.4258 | 118.6833 | 181.5546 | 180.0197 | 211.0197 | 158.5098 | 150.4061 | 139.7674 | 181.0644 | 135.1288 | 199.1879 | 205.3417 | 219.5742 | 234.5993 | 191.1038 | 175.7227 | 238.9553 | 221.8318 | 198.7674 | 202.2129 | 220.4007 | 281.8962 | 175.2379 | 181.1038 | 173.5098 | 228.0447 | 151.9803 | 154.4454 | 139.703 | 153.4902 | 186.1091 | 199.7227 | 194.1038 | 146.0197 | 130.3614 | 132.2773 | 117.2129 | 140.297 | 286.7818 | 274.8068 | 226.6189 | 324.1288 | 239.8765 | 258.703 | 237.1485 | 323.9553 | 153.1038 | 84.1485 | 85.3614 | 153.1038 | 121.1932 | 125.0644 | 105.6386 | 126.297 | 120.4902 | 126.3417 | 114.8265 | 131.9356 | 85.4651 | 86 | 79.0197 | 109.4061 | 91.4705 | 128.6386 | 95.1932 | 85.0841 | 287 | 255.5295 | 219.703 | 282.4258 | 158.2379 | 197.8765 | 216.0841 | 252.3417 | 316.5295 | 222.6583 | 242.3167 | 308.2773
961 | 12.5654 | 159.2006 | 201.9853 | 230.7965 | 191.2094 | 180.8024 | 173.6047 | 175.0029 | 148.4365 | 182.9912 | 138.6018 | 116.0029 | 168.0088 | 240.9853 | 160.8053 | 156.8082 | 220.9824 | 193.5988 | 187.5812 | 150.9912 | 201.1947 | 185.0147 | 321.9882 | 310.3835 | 164.8141 | 229.6076 | 264.5724 | 250.5782 | 264.0147 | 134.5871 | 151.3776 | 117.0088 | 131.4012 | 181.7759 | 157.3953 | 150.8024 | 99.6194 | 276.6047 | 236.7935 | 258.3865 | 284.3953 | 150.0176 | 128.2065 | 157.5959 | 142.1888 | 89.7965 | 90.4012 | 88.3953 | 147.1888 | 215.3953 | 212.3894 | 218.6224 | 156.6165 | 171.4041 | 154.3865 | 132.8053 | 137.1829 | 180.5988 | 185.3982 | 217.9971 | 159.2006 | 158.1918 | 139.3982 | 187.1947 | 136.9971 | 161.2329 | 179.8171 | 219.2006 | 231.9971 | 201.5959 | 178.5988 | 259.7876 | 202.8112 | 215.9853 | 221.9853 | 200.6194 | 296.1859 | 176.7935 | 190.7965 | 176.5988 | 232.7935 | 146.6018 | 157.7994 | 145.9941 | 165.59 | 200.9853 | 177.8171 | 194.2035 | 161.7906 | 130.2006 | 111.8141 | 128.9912 | 124.4129 | 277.0147 | 229.2359 | 244.5841 | 336.3894 | 265.5753 | 271.3894 | 253.1888 | 301.6194 | 148.4071 | 76.2065 | 78.0059 | 148.4071 | 114.8024 | 115.2065 | 104.2006 | 133.5959 | 115.0029 | 124 | 120.2006 | 121.0088 | 75.8112 | 85.2006 | 86.7965 | 94.8082 | 108.5841 | 111.2124 | 93.5988 | 80.6047 | 275.0088 | 260 | 226.7935 | 277.2035 | 161.3924 | 197.1947 | 222.7965 | 250 | 310.6076 | 192.2241 | 239.8053 | 310.1976
962 | 12.5764 | 164.6819 | 213.4936 | 224.6947 | 191.2468 | 177.5649 | 168.7532 | 178.8702 | 137.6234 | 155.1553 | 140.4351 | 130.4224 | 170.6819 | 246 | 160.6234 | 166.1755 | 221.3181 | 196.4351 | 171.2723 | 153.1883 | 212.7404 | 175.1298 | 334.117 | 297.3308 | 183.5394 | 253.7862 | 243.1553 | 234.4606 | 273.6107 | 165.7862 | 141.1425 | 135.1043 | 133.4351 | 173.7659 | 163.0585 | 159.7404 | 111.6692 | 287.1755 | 239.8117 | 248.3893 | 280.3181 | 154.5521 | 155.2213 | 143.5776 | 121.6489 | 89.3766 | 68.084 | 77.8245 | 136.3893 | 231.6107 | 209.5064 | 248.3383 | 163.9872 | 165.1298 | 147.6362 | 155.3511 | 152.7404 | 190.7404 | 177.883 | 214.1298 | 165.4936 | 167.4936 | 152.9872 | 180.883 | 126.6362 | 211.6894 | 165.883 | 225.4936 | 257.3511 | 181.8957 | 157.084 | 220.9798 | 213.6107 | 210.4479 | 193.7202 | 205.3638 | 296.1298 | 183.0585 | 194.4351 | 172.9415 | 231.7532 | 167.1043 | 153.9415 | 152.0585 | 173.8702 | 195.4479 | 194.7277 | 180.8245 | 136.5904 | 123.5064 | 132.9745 | 101.1553 | 137.0458 | 273.6234 | 279.5011 | 267.0458 | 331.0713 | 263.4479 | 247.4021 | 261.0585 | 299.0585 | 136.2596 | 69.9415 | 81.6819 | 136.2596 | 130.2341 | 123.5521 | 100.7532 | 125.2596 | 122.117 | 126.4351 | 115.1298 | 129.3638 | 93.1043 | 72.8245 | 80.6947 | 109.8575 | 114.8117 | 128.1043 | 106.1755 | 92.7989 | 284.1755 | 229.1553 | 255.7862 | 286.5521 | 171.3053 | 198.1883 | 211.8245 | 273.5394 | 291.7659 | 206.7277 | 223.3893 | 289.084
963 | 12.5874 | 171.768 | 210.056 | 244.4238 | 196.944 | 166.2881 | 169.648 | 187.4159 | 178.3757 | 161.1839 | 138.528 | 141.4159 | 165.408 | 235.2881 | 156.88 | 177.2399 | 222.472 | 188.7601 | 168.472 | 170.3039 | 194.4002 | 183.0639 | 288.2084 | 312.7759 | 172.5201 | 261.648 | 222.8161 | 224.88 | 283.592 | 150.5762 | 127.9361 | 122.6961 | 148.8319 | 160.9361 | 146.6961 | 157.056 | 101.1681 | 290.824 | 241.648 | 258.1839 | 272.408 | 157.824 | 140.5762 | 146.592 | 134.1279 | 100.5359 | 72.0639 | 72.528 | 124.7601 | 230.88 | 250.8477 | 247.1121 | 171.944 | 173.8879 | 145.824 | 147.8161 | 134.4002 | 204.5359 | 159.5201 | 219.592 | 169.472 | 163.232 | 142.8161 | 195.4799 | 139.6559 | 188.0963 | 200.2557 | 207.2241 | 244.8721 | 177 | 159.4159 | 231.5998 | 206.2881 | 210.472 | 206.5998 | 180.8082 | 297.472 | 163.4002 | 188.408 | 179.4159 | 221.1121 | 158.8161 | 147.232 | 171.1279 | 167.5841 | 183.9361 | 195.056 | 199.4238 | 138.2399 | 137.6559 | 143.944 | 132.7277 | 141 | 281.4159 | 264.3362 | 210.0245 | 328.176 | 261.824 | 262.4238 | 243.0481 | 314.8319 | 141.4159 | 72.296 | 80.528 | 141.4159 | 119.1681 | 127.648 | 107.4159 | 115.5841 | 106.6961 | 119.5841 | 100.8161 | 133.648 | 87.2881 | 70.824 | 79.824 | 118.12 | 109.232 | 123.9361 | 93.3441 | 93.528 | 316.6638 | 227.768 | 239.7522 | 298.0639 | 187.0079 | 209.5359 | 231.2478 | 255.1042 | 284.704 | 210.352 | 220 | 295.5359
964 | 12.5984 | 155.438 | 219.8717 | 218.8937 | 166.2212 | 181.562 | 205.9603 | 174.7832 | 180.3097 | 173.1991 | 129.6372 | 143.8363 | 162.3274 | 222.9646 | 177.7434 | 170.6372 | 253.1063 | 209.5797 | 179.8717 | 155.6017 | 205.0531 | 195.8717 | 298.0708 | 293.584 | 178.1991 | 245.2743 | 251.7788 | 236.5443 | 272.4557 | 141.8186 | 134.3628 | 129.0354 | 132.7654 | 165.6903 | 178.9603 | 148.4734 | 87.9646 | 307.7257 | 273.7788 | 264.3451 | 270.1637 | 172.2168 | 129.3097 | 137.9646 | 139.6726 | 89.6194 | 86.5443 | 72.8363 | 131.3628 | 232.5089 | 233.2389 | 246.6726 | 150.4203 | 190.2168 | 151.0177 | 141.6549 | 129.1637 | 203.6549 | 192.7965 | 224.3451 | 184.2168 | 152.8009 | 158.3983 | 174.7477 | 124.6017 | 204.2523 | 191.2743 | 209.6903 | 220.0929 | 190.3806 | 149.292 | 255.2346 | 200.6549 | 196.7832 | 229.3983 | 188.3806 | 264.5486 | 182.416 | 187 | 176.8186 | 238.2346 | 149.3097 | 146 | 150.7477 | 158.4734 | 166.9469 | 199.0177 | 190.6194 | 131.6372 | 116.7477 | 113.2212 | 116.7477 | 131.8009 | 278.8186 | 241.2743 | 226.27 | 358.9425 | 270.3628 | 251.9469 | 239.8363 | 328.8717 | 143.8363 | 78.0177 | 85.854 | 143.8363 | 122.6903 | 122.146 | 110.6726 | 156.6505 | 126.416 | 126.3628 | 101.3451 | 110.584 | 79.9823 | 65.9823 | 81.6726 | 98.9292 | 108 | 105.2743 | 80.8009 | 90.4911 | 249.4069 | 217.292 | 235 | 332.6151 | 189.1637 | 216.1814 | 209.2389 | 260.0354 | 302.3983 | 231.7434 | 213.3097 | 275.4203
965 | 12.6094 | 185.943 | 213.5142 | 213.8486 | 197.3373 | 189.2429 | 185.8456 | 166.9085 | 141.6627 | 178.3943 | 140.7286 | 123.6342 | 176.4258 | 235.4258 | 153.1484 | 162.2114 | 247.9685 | 207.2114 | 191.3343 | 168.1229 | 196.1199 | 190.3628 | 293.5142 | 276.2714 | 168.9685 | 245.3943 | 259.6972 | 230.5142 | 252.1799 | 142.6972 | 136 | 125.06 | 141.7286 | 173.7886 | 174.9685 | 136.8171 | 93.6372 | 310.1514 | 251.1484 | 245.4828 | 267.4543 | 149.5427 | 153.4573 | 135.1514 | 140 | 89.5457 | 78.8171 | 83.1829 | 115.1799 | 238.94 | 222.9085 | 219.8456 | 150.2429 | 161.6027 | 158.7886 | 157.9715 | 123.9085 | 211.4858 | 199.1514 | 219.9085 | 161.5427 | 145.9085 | 151.8171 | 178.4858 | 137.9715 | 190.3313 | 180.3628 | 194.0285 | 236.3658 | 187.06 | 160.5772 | 235.2399 | 209.3343 | 224.5487 | 221.9685 | 213.063 | 314.8546 | 178.5142 | 171.7256 | 166.6657 | 237.7571 | 161.5772 | 144.3028 | 154.4858 | 129.8456 | 174.1829 | 177.937 | 200.7286 | 130.8486 | 111.1514 | 129.063 | 113.6972 | 113.0285 | 262.7256 | 266.003 | 231.1514 | 336.997 | 262.6657 | 281.2459 | 242.5457 | 306.3913 | 138.9085 | 76.4543 | 81.9085 | 138.9085 | 119.7571 | 125.2429 | 119.4858 | 136.1484 | 108.0884 | 114.4228 | 104.5457 | 114.4858 | 73.06 | 71.7886 | 73.5142 | 103.4858 | 108 | 102.8486 | 107.003 | 84.9085 | 275.7316 | 230.2744 | 228.2114 | 320.3313 | 194.94 | 219.5457 | 202.3028 | 267.0915 | 276.2998 | 232.6057 | 215.3943 | 306.6402
966 | 12.6204 | 164.4523 | 224.0521 | 231.2173 | 218.6347 | 167.6175 | 169.8087 | 177.1913 | 184.9303 | 175.5565 | 143.8609 | 120.8609 | 167.8087 | 249.1913 | 132.5044 | 162.7217 | 235.6696 | 187.0609 | 170.6175 | 157.2261 | 190.5565 | 184.6957 | 292 | 293.7999 | 168.7217 | 227.9218 | 246.2261 | 219.5305 | 296.3477 | 130.9479 | 147.1913 | 127.4435 | 162.9391 | 156.0609 | 166.9739 | 156.5217 | 95.8609 | 304.8348 | 283.0173 | 246.3043 | 273.0261 | 133.8087 | 160.5826 | 130.6957 | 146.8869 | 90.8609 | 74.4174 | 83.2783 | 132.6608 | 247.7478 | 215.9739 | 262.3477 | 172.5217 | 168.0521 | 149.6696 | 144.6435 | 131.6087 | 194.0609 | 203.3043 | 212.9739 | 171.6347 | 152.7478 | 153.4435 | 180.8609 | 141 | 180.1131 | 187.6087 | 209.9391 | 225.3653 | 187.7217 | 175.0521 | 253.3825 | 182.5914 | 247.2173 | 214.8348 | 209.2522 | 312.087 | 181.3043 | 140.5914 | 183.9391 | 237 | 154.5305 | 150.8869 | 154.2783 | 152.5477 | 200.1043 | 189.4956 | 187.5044 | 126.6957 | 119.6087 | 119.2261 | 110.5565 | 130.6608 | 277.2173 | 282.1913 | 237.8869 | 322.5305 | 286.826 | 285.2783 | 244.7217 | 284.7827 | 133.6957 | 76.8609 | 81 | 133.6957 | 127.6087 | 138.0521 | 117.5565 | 110.3392 | 116.913 | 123.1913 | 134.2695 | 107.3913 | 65.9739 | 73 | 87.4956 | 115.3304 | 102.8348 | 112.4695 | 90.4783 | 90.8869 | 271.8087 | 255.3825 | 264.0173 | 295.4783 | 171.0349 | 201.0609 | 188.2261 | 268.8609 | 315.7651 | 228.5565 | 216 | 300.9479
967 | 12.6314 | 160.8732 | 207.6637 | 211.2979 | 213.1416 | 183.2095 | 179.3511 | 174.6342 | 186.8879 | 182.8584 | 121.2979 | 99.171 | 160.7611 | 257.9853 | 164.0532 | 174.3511 | 240.1121 | 217.18 | 167.8732 | 168.0974 | 186.5074 | 186.6195 | 315.5753 | 310.9705 | 187.3363 | 268.6579 | 228.2832 | 237.2095 | 281.2979 | 145.59 | 168.2095 | 127.1268 | 158.1416 | 180.9411 | 158.1416 | 142.5368 | 85.5221 | 319.7168 | 260.1858 | 228.6637 | 283.6047 | 159.0679 | 176.7168 | 124.7611 | 109.581 | 84.0147 | 94.0826 | 83.8732 | 109.8053 | 220.1858 | 230.7168 | 229.8347 | 158.5368 | 161.2684 | 149.7463 | 137.6342 | 127.7611 | 195.3658 | 193.5221 | 217.2389 | 183.6047 | 139.1563 | 152.2537 | 186.2389 | 164.5753 | 170.2684 | 195.9853 | 198.1563 | 247.4484 | 196.7316 | 177.8732 | 248.2684 | 207.6874 | 215.0737 | 189.5516 | 210.6195 | 300.3953 | 184.6195 | 166.5605 | 183.5074 | 235.2537 | 147.7611 | 143.2684 | 151.3805 | 152.6342 | 191.7758 | 212.9558 | 215.5605 | 135.6047 | 141.9558 | 126.6047 | 122.2242 | 141.8584 | 262.5368 | 236.8495 | 208.4395 | 341.0826 | 290.1268 | 268.41 | 254.6047 | 293.3511 | 126.0147 | 84.8584 | 81 | 126.0147 | 100.1858 | 123.41 | 122.2389 | 114.8584 | 111.1416 | 116.2684 | 122.41 | 133.9411 | 73.7316 | 78.2389 | 78.6489 | 94.2979 | 102 | 109.6342 | 108.829 | 79.7758 | 263.0147 | 259 | 230.7079 | 312.9558 | 180.9705 | 222.4484 | 213.0679 | 234.9468 | 334.3511 | 217.5221 | 221.2389 | 289.3953
968 | 12.6424 | 180.4799 | 189.0619 | 206.2291 | 200.4892 | 179.8019 | 155.322 | 168.6873 | 159.4365 | 150.3529 | 139.2508 | 106.6254 | 175.9381 | 235.0929 | 181.3963 | 173.3437 | 213.5511 | 198.9783 | 167.1146 | 150.5201 | 186.8854 | 172.8328 | 301.291 | 297.9474 | 183.8019 | 243.1238 | 245.4799 | 216.0929 | 304.5635 | 140.031 | 135.582 | 141.1672 | 145.4892 | 159.322 | 151.6873 | 140.8854 | 92.8545 | 294.5511 | 262.1981 | 233.0836 | 273.4892 | 163 | 142.6966 | 154.9907 | 123.4799 | 91.8545 | 80.1765 | 97.2817 | 125.4799 | 214.2291 | 208.2074 | 223.1146 | 184.3344 | 173.2817 | 146.4582 | 137.8854 | 130.5418 | 196.8854 | 189.3437 | 202.9474 | 187.6563 | 154.709 | 152 | 175.4892 | 140.5511 | 182.2817 | 196.1146 | 211.9381 | 223.5511 | 201.5418 | 184.1981 | 225.7492 | 200.4892 | 203.8019 | 211.678 | 211 | 303.4272 | 181.4582 | 196.678 | 153.7802 | 236.7709 | 146.1146 | 148.1981 | 148.3437 | 163.5108 | 211.2508 | 185.0093 | 190.7802 | 140.5418 | 116.6656 | 127.1146 | 125.7709 | 143 | 324.6378 | 259.2198 | 255.3561 | 334.2601 | 284.6873 | 275.7399 | 223.2384 | 305.6254 | 136.5108 | 71.8328 | 88.0836 | 136.5108 | 119.9071 | 131.6254 | 117.6873 | 131.0526 | 118.8545 | 102.6037 | 103.1765 | 114.9783 | 73.2291 | 92.2817 | 80.5418 | 111.3653 | 105.5418 | 117.8545 | 115.5418 | 94.8235 | 271.7399 | 267.8545 | 247.1362 | 288.5511 | 174.1455 | 210.0619 | 206.3746 | 262.7616 | 302.3529 | 209.8019 | 213.1455 | 302.1672
969 | 12.6534 | 172.2271 | 204.0571 | 243.7051 | 203.4887 | 177.2045 | 175.3412 | 191.3412 | 153.3068 | 174.7277 | 143.7955 | 125.0571 | 157.352 | 206.8633 | 165.0452 | 156.8407 | 225.2616 | 218.4435 | 174.1819 | 178.5232 | 193.2842 | 188.9548 | 319.648 | 260.0904 | 195.5684 | 265.9322 | 240.8181 | 236.3412 | 314.2842 | 128.2271 | 137.2842 | 137.6136 | 156.5684 | 147.0226 | 173.4435 | 163.4435 | 112.8525 | 289.2045 | 269.2842 | 275.296 | 274.6932 | 173.7729 | 143.3864 | 151.8181 | 156.5232 | 89.409 | 78 | 84.6362 | 123.5113 | 240.9322 | 205 | 231.0797 | 170.9429 | 157.0452 | 178.3186 | 171.2164 | 123.8181 | 227.5232 | 198.8751 | 211.7729 | 170.0452 | 169.5684 | 158.2842 | 183.8751 | 137.8977 | 197.4661 | 203.1819 | 217.591 | 234.3638 | 204.6932 | 159.8633 | 223 | 212.4661 | 216.4661 | 180.8859 | 197.5339 | 336.3186 | 190.8751 | 176.6588 | 179.6254 | 223.5339 | 155.8751 | 147.2045 | 136.3294 | 178.4661 | 186.1701 | 171.1249 | 181.6136 | 150.8751 | 123.7729 | 118.9203 | 117.0226 | 142.1023 | 314.1475 | 237.8633 | 224.2949 | 311.4542 | 299.2616 | 240.1927 | 245.9322 | 314.1819 | 120.9429 | 79.8751 | 71.9429 | 120.9429 | 117.6136 | 118.6362 | 106.2271 | 120.4316 | 121.7955 | 119.8525 | 118.0571 | 122.7729 | 62.2271 | 94 | 77.409 | 104.1249 | 118.5684 | 103.7384 | 89.9655 | 83.5339 | 282.8751 | 273.4887 | 216.7836 | 290.3864 | 175.6932 | 194.5339 | 204.1023 | 266.1023 | 321.3412 | 215.2842 | 209.3068 | 296.8181
970 | 12.6644 | 165.5398 | 211.4602 | 223.4292 | 195.8097 | 181.5502 | 167.9897 | 180.3495 | 167.5605 | 159.7994 | 137.6298 | 115.1696 | 178.6608 | 213.1003 | 155.7197 | 148.6298 | 222.4498 | 222.8201 | 154.0693 | 152.879 | 184.8997 | 191.91 | 301.0693 | 286.031 | 161.5088 | 262.6298 | 225.4395 | 218.9793 | 265.8583 | 135.1903 | 152.5605 | 134.2699 | 161.6401 | 165.1107 | 166.8997 | 160.5398 | 98.6194 | 303.5605 | 246.3392 | 239.9586 | 273.1799 | 175.91 | 144 | 121.879 | 146.3495 | 105.3806 | 69.8097 | 71.1696 | 112.0796 | 262.2006 | 222.2906 | 248.3806 | 159.8997 | 172.2906 | 155.6091 | 153.1592 | 137.5605 | 214.6194 | 216.3806 | 212.09 | 168.91 | 150.9793 | 143.5295 | 180.4498 | 141.6401 | 171.6991 | 200.3599 | 221.6401 | 228.7197 | 159.4984 | 166.1003 | 205.7094 | 206.7197 | 188.879 | 238.8821 | 211.4705 | 305.4188 | 182.8997 | 173.09 | 187.5502 | 231.1003 | 157 | 144.2699 | 172.3113 | 180.91 | 193.9204 | 198.211 | 179.1799 | 143.8097 | 114.9897 | 121.6401 | 114.1799 | 146.5502 | 280.1489 | 262.3009 | 263.6815 | 339.031 | 250.9483 | 268.7611 | 260.8304 | 330.4705 | 139.9307 | 76.4498 | 70.91 | 139.9307 | 124.2803 | 118.8201 | 125.0207 | 122.6401 | 104.7094 | 142.0207 | 126.3702 | 107.6194 | 70.1003 | 84.8997 | 73.3599 | 110.2803 | 91.789 | 90.1696 | 100.6505 | 85.6401 | 277.6298 | 241.2389 | 232.1107 | 294.6401 | 190.5605 | 212.1107 | 217.6505 | 257.8097 | 312.1696 | 220.5502 | 219.0103 | 308.7404
971 | 12.6754 | 170.5339 | 198.1652 | 220.0777 | 230.0482 | 171.8545 | 158.6991 | 179.9223 | 169.9223 | 166.3009 | 126.8545 | 121.3786 | 177.3107 | 226.9125 | 144.8545 | 151.6893 | 222 | 209.1652 | 150.1554 | 156.4563 | 195.9902 | 191.0777 | 296.233 | 270.5536 | 185.6697 | 269.3786 | 237.8348 | 197.6312 | 305.2715 | 154.4464 | 143.8545 | 130.3107 | 154.6214 | 177.1455 | 156.7768 | 156.3107 | 96.0777 | 318.8348 | 248.6116 | 245.2232 | 288.6795 | 144.641 | 155.0679 | 142.9804 | 134.8545 | 70.1071 | 85.6018 | 93.0581 | 121.1455 | 222.4955 | 208.3205 | 235.2428 | 159.9223 | 161.0875 | 139.1652 | 139.9321 | 135.3107 | 214.8446 | 184.7964 | 195.3982 | 163.4661 | 146.233 | 130.9321 | 190.1455 | 125.3982 | 204.0482 | 212.9125 | 218.3107 | 227.0777 | 179.9027 | 166.0777 | 220.6018 | 207.8446 | 215.5143 | 227.4759 | 227.7572 | 293.6991 | 201.3688 | 183.1455 | 165.8643 | 204.3303 | 153.3107 | 152.3009 | 152.9419 | 167.1652 | 195.9223 | 199.1554 | 154.0973 | 131.9321 | 132.4464 | 132.1455 | 119.5339 | 157.1455 | 264.0875 | 245.6312 | 257.8545 | 339.233 | 259.8348 | 272.9223 | 277.6795 | 308.9419 | 135.5437 | 78.767 | 79.3009 | 135.5437 | 121.3107 | 118.0777 | 113.1652 | 118.3884 | 113.1455 | 108.9518 | 107.6312 | 125.3688 | 81.1455 | 71.0875 | 79.4563 | 108.233 | 101.9125 | 100.9902 | 95.5437 | 94.3009 | 276.0777 | 246.3009 | 237.6893 | 278.3982 | 179.0875 | 191.8643 | 226.3786 | 237.6312 | 299.9321 | 209.9321 | 208.0098 | 304.4661
972 | 12.6864 | 178.4769 | 216.6269 | 221.8692 | 200.2885 | 161.6539 | 189.7769 | 214.5807 | 170 | 188.4961 | 130.6731 | 133.2154 | 161.1116 | 230.8038 | 155.2154 | 150.1308 | 227.6077 | 189.3077 | 161.2154 | 127.0924 | 171.7654 | 178.85 | 308.15 | 286.7577 | 194.5423 | 261.5885 | 267.973 | 240.8614 | 290.3077 | 147.5885 | 143 | 134.6731 | 136.2423 | 166.7846 | 149.4577 | 150.3923 | 89.4577 | 288.2231 | 280.7769 | 272.1692 | 247.0078 | 157.8884 | 132.6347 | 133.7846 | 148.0192 | 72.6077 | 92.6077 | 71.6347 | 122 | 235.823 | 231.2999 | 218.1116 | 183.3653 | 172.15 | 145.4769 | 138.0654 | 148.0846 | 221.5423 | 192.2807 | 210.823 | 178.8884 | 146.9346 | 146.823 | 183.5231 | 150.1692 | 201.3923 | 205.5885 | 207.7193 | 235.4115 | 162.3731 | 157.5885 | 218.2615 | 220.15 | 211.4577 | 203.5693 | 187.877 | 289.2615 | 181.5039 | 165.3077 | 158.3923 | 217.8884 | 173.5615 | 147.3923 | 158.4769 | 160.3923 | 196.9346 | 196.1962 | 201.5345 | 135.6731 | 131.1962 | 113.3731 | 125.6077 | 145.85 | 243.3731 | 272.0384 | 266.3461 | 338.0654 | 291.8422 | 270.1962 | 238.8116 | 290.177 | 152.7577 | 79.9346 | 59.4385 | 152.7577 | 129.4115 | 125.4769 | 123.2154 | 128.2807 | 125.2154 | 122.823 | 118.15 | 137.2807 | 71.7193 | 72.8038 | 74.3923 | 112.6731 | 99.2615 | 125.3653 | 104.3461 | 91.2615 | 307.7769 | 216.1578 | 232.3923 | 305.0384 | 163.9808 | 217.1038 | 208.3077 | 285.5345 | 297.1308 | 223.0192 | 221.9538 | 331.1038
973 | 12.6974 | 161.9557 | 236.9381 | 229.5752 | 186.6371 | 207.3983 | 180.6371 | 181.9645 | 152.0088 | 175.7964 | 144.2567 | 123.5841 | 163.7876 | 245.2036 | 141.7964 | 150.9469 | 216.6371 | 195.5752 | 157.2655 | 160.9824 | 188.9381 | 208.3009 | 322.2567 | 313.5664 | 180.7964 | 245.8495 | 242.5398 | 241.1593 | 298.469 | 156.469 | 135.4248 | 111.3274 | 150.1505 | 160.3186 | 166.9912 | 148.1062 | 112.6726 | 304.9381 | 246.0707 | 248.4336 | 267.6726 | 157.1062 | 148.0443 | 154.7788 | 145.2124 | 89.0974 | 70.2743 | 71.8938 | 127.6814 | 236.0531 | 213.115 | 254.8762 | 161.3274 | 206.1417 | 150.7345 | 151.2567 | 123.4336 | 189.8052 | 203.4159 | 202.531 | 161.0619 | 146.0531 | 134.7433 | 168.7964 | 141.5841 | 216.1505 | 210.6814 | 183.3274 | 241.6814 | 165.7345 | 169.3098 | 237.885 | 204.9026 | 214.7876 | 187.7964 | 206.7788 | 306.0443 | 173.3717 | 189.5664 | 163.6814 | 218.0531 | 157.0088 | 149.8407 | 155.2124 | 170.4159 | 159.1238 | 178.9557 | 204.0531 | 137.8938 | 135.7345 | 130.9381 | 123.1593 | 127.9557 | 285.5576 | 258.8495 | 229.1238 | 327.5841 | 261.8052 | 269.0531 | 263.4602 | 315.5133 | 127.4867 | 85.6814 | 75.9912 | 127.4867 | 101.5929 | 117.4779 | 113.5841 | 128.0531 | 118.4248 | 102.2212 | 135.0974 | 127.5841 | 87.0974 | 75.8407 | 73.0531 | 92.1681 | 95.2124 | 117.531 | 105 | 106.1505 | 269.2831 | 225.3629 | 201.6991 | 287.115 | 205.6107 | 218.0531 | 238.2478 | 292.7876 | 272.3805 | 220.2124 | 238.1505 | 315.0088
974 | 12.7084 | 174.4287 | 219.7753 | 225.204 | 202.3063 | 160.1218 | 175.204 | 180.9592 | 152.9184 | 173.0816 | 136.3672 | 120.1224 | 161.1224 | 230.6529 | 138.1224 | 160.592 | 236.1431 | 205.592 | 172.3471 | 137.1017 | 210.1431 | 184.1017 | 307.6529 | 291.0201 | 194.3879 | 241.1632 | 249.6328 | 228.5305 | 284.6121 | 149.3264 | 146.5103 | 135.8983 | 171.1431 | 158.0816 | 159.3672 | 147.0408 | 112.0816 | 282.0201 | 263.1839 | 255.6328 | 271.8776 | 138.7753 | 182.5718 | 128.1833 | 150.7552 | 80.408 | 83.3879 | 88.3063 | 142.3879 | 231.204 | 225.4287 | 256.0408 | 176.3063 | 170.5914 | 144.2856 | 124.1833 | 111.4488 | 188 | 153.1626 | 209.6736 | 174.3879 | 152.7144 | 150.3063 | 168 | 151.5512 | 218.9184 | 175.5098 | 206.9391 | 240.0816 | 186.1431 | 167.1224 | 238.0408 | 209.7552 | 248.5718 | 194.6736 | 196.4897 | 316.592 | 167.2448 | 185.2448 | 167.8368 | 254.4494 | 187.6535 | 144.2448 | 125.2649 | 156.6121 | 180.0207 | 223.0822 | 174.2649 | 139.9184 | 144.6328 | 123.3672 | 127.796 | 128.9184 | 274.5713 | 269.5103 | 260.5718 | 295.3465 | 281.1023 | 231.5914 | 273.6328 | 305.4897 | 108.7345 | 85.0408 | 84.6736 | 108.7345 | 129.7351 | 114.1224 | 113.9592 | 113.6121 | 130.4695 | 123.0615 | 112.0201 | 140.4287 | 78.408 | 89.4287 | 59.5713 | 96.7552 | 103.6328 | 89.1833 | 102.1224 | 77.2649 | 308.2454 | 256.6943 | 229.7351 | 280.2448 | 201.2856 | 181.5506 | 212.1833 | 246.9586 | 300.7351 | 225.7552 | 211.1833 | 290.9793
975 | 12.7194 | 175 | 235.5153 | 222.0855 | 198.1426 | 170.6293 | 193.4582 | 156.7128 | 180.2016 | 167.1711 | 123.3707 | 124.8574 | 167.8004 | 241.6578 | 153.5438 | 145.4562 | 213.6843 | 239.0305 | 171.057 | 175.8309 | 165.3401 | 194.6578 | 318.6578 | 313.3157 | 177.5132 | 284.7169 | 232.5132 | 241.6008 | 294.6864 | 138.3136 | 152.8289 | 129.2281 | 147.7128 | 142.4562 | 146.3707 | 156.7149 | 112.9715 | 298.4868 | 268.8574 | 243.3707 | 269.0855 | 175.888 | 157.7699 | 173.6314 | 118.9409 | 85.8289 | 64.5703 | 68.5988 | 129.3992 | 246.5438 | 221.1426 | 235.5988 | 187.6864 | 173.8574 | 138.1711 | 131.7434 | 116.8289 | 208.4012 | 198.6028 | 216.8004 | 156.5418 | 153 | 146.1426 | 208.8024 | 158.8004 | 196.6558 | 191.4868 | 242.9735 | 236.1141 | 168.5418 | 179.6293 | 237.0285 | 208.057 | 239.3136 | 237.7454 | 198.9145 | 286.8839 | 200.0305 | 145.1691 | 172.8574 | 234.6273 | 155.9695 | 159.5438 | 151.2016 | 187.0876 | 170.3136 | 198.7699 | 182.7149 | 158.4582 | 109.055 | 107.4562 | 114.3992 | 113.4562 | 272.057 | 285.5438 | 238.6843 | 301.7719 | 280.057 | 268.8595 | 269.1426 | 299.1711 | 133.2586 | 73.3422 | 70.4277 | 133.2586 | 130.0285 | 110.1141 | 114 | 122.7149 | 122.2566 | 132.7434 | 121.6864 | 122.5418 | 94.5153 | 81.2566 | 95.9165 | 90.1996 | 116.6293 | 107.4297 | 114.6293 | 80.8574 | 278.9124 | 275.4868 | 265.9735 | 291.6578 | 209.7434 | 218.8595 | 230.4297 | 257.6293 | 297.1426 | 204.6273 | 219.7149 | 332.7454
976 | 12.7304 | 176.9676 | 206.4867 | 236.7567 | 188.1622 | 163.1298 | 177.2758 | 192.3998 | 148.5353 | 169.9513 | 130.8702 | 138.7729 | 162.0973 | 261.6756 | 156.9513 | 182.3835 | 224.8053 | 239.0162 | 173.9513 | 146.5029 | 210.2375 | 198.9351 | 345.562 | 261.8598 | 206.5133 | 261.4056 | 237.9027 | 244.9513 | 295.9838 | 140.9513 | 115.6165 | 116.2109 | 153.8864 | 176.4322 | 147.9676 | 153.0649 | 111.0324 | 275.3893 | 241.4542 | 243 | 294.5782 | 196.6756 | 147.1622 | 121.876 | 140.6269 | 97.8053 | 86.6269 | 77.8378 | 133.9189 | 210.6002 | 235.7567 | 249.7567 | 183.0811 | 195.6431 | 146.854 | 139.8702 | 119.9513 | 205.0649 | 173.438 | 191.4218 | 173.708 | 161.854 | 165.6756 | 178.5191 | 136.3731 | 193.0487 | 190.0324 | 220.3893 | 212.3893 | 176.854 | 166.2271 | 240.9351 | 186.3569 | 215.3893 | 220.3082 | 236.3835 | 299.7729 | 179.3569 | 172.5295 | 182.8378 | 268.4322 | 130.4056 | 149.1784 | 151.0162 | 159.4705 | 166.0649 | 178.3244 | 218.416 | 120.6327 | 135.5458 | 121.7567 | 127.7729 | 119.8864 | 276.9189 | 291.9027 | 241.9351 | 329.5458 | 290.8216 | 258.1947 | 222.7625 | 283.2596 | 129.0811 | 73 | 84.7567 | 129.0811 | 114.2596 | 109.0162 | 138.5944 | 127.9189 | 113.146 | 122.1784 | 129.8702 | 115.1136 | 80.2433 | 81.9838 | 88.146 | 107.708 | 110.1136 | 107.0162 | 95.3244 | 92.8053 | 292.7567 | 240.584 | 259.1298 | 284.1298 | 194.2596 | 233.7729 | 222.146 | 278.6593 | 312.7404 | 205.9676 | 219.0162 | 333.0162
977 | 12.7414 | 191.941 | 228.9096 | 244.9685 | 182.0236 | 170.9685 | 179.9882 | 183.0393 | 171.9056 | 170 | 125.0236 | 124.059 | 173.9528 | 259.0118 | 141.0629 | 169.055 | 248.9056 | 213.1022 | 168.0236 | 140.0236 | 228.9292 | 191.0315 | 332.055 | 294.8663 | 184.0904 | 228.1297 | 222.0629 | 223.0865 | 293.0118 | 160.9214 | 145.8781 | 131.9371 | 175.9135 | 156.0826 | 172.9017 | 157.9803 | 78.1297 | 299.9017 | 256.9371 | 226.0668 | 226.2713 | 170.1061 | 167.9174 | 156.8585 | 126.059 | 81.0668 | 77.0393 | 69.0354 | 150.9332 | 239.8821 | 222.055 | 241.0354 | 159.0944 | 180.0629 | 151.9803 | 122.0708 | 131.9528 | 226.9135 | 198.8978 | 204.945 | 174.9961 | 171.9607 | 148.0708 | 215.8506 | 131.0197 | 198.9764 | 195.9764 | 197.0904 | 240.886 | 189.9489 | 153.0511 | 228.0511 | 214.886 | 224.9607 | 249.8821 | 227.0393 | 311.9528 | 184.9764 | 182.9607 | 183.9961 | 234.1376 | 176.8152 | 157.9646 | 135.0629 | 173.941 | 214.8074 | 167.0432 | 192.1061 | 142.9096 | 133.0118 | 126.9803 | 105.0904 | 134.941 | 288.9528 | 244.1887 | 255.945 | 333.9843 | 291.9961 | 260.9882 | 255.8663 | 310.8899 | 122.0275 | 78.9764 | 63.0865 | 122.0275 | 140.8939 | 116.9685 | 113.1022 | 133.9764 | 128.9371 | 120.0079 | 100.1179 | 123.9646 | 74.0236 | 81.0039 | 96.9646 | 109.9921 | 110.9961 | 115.9646 | 90.0197 | 86.0275 | 271.0865 | 265.8978 | 217.1651 | 289.9764 | 192.0079 | 242.9646 | 201.0826 | 256.0904 | 300.0511 | 227.9135 | 222.9843 | 316.0668
978 | 12.7524 | 175.1923 | 221.6823 | 272.5737 | 204.8495 | 201.7576 | 178.9833 | 189.9331 | 161.9248 | 167.8913 | 114.2006 | 124.9498 | 174.9666 | 229.0585 | 151.2424 | 129.2424 | 204.0585 | 181.3929 | 185.9164 | 147.0418 | 205.0752 | 210.8328 | 312.3344 | 286.0585 | 189.8412 | 272.9415 | 216.1003 | 260.7074 | 291.9164 | 134.2174 | 140.1003 | 110.0836 | 142.8997 | 172.791 | 142.0167 | 165.7325 | 83.1087 | 328.7492 | 278.9415 | 238.1923 | 268.7994 | 170.0167 | 150.9498 | 137.9749 | 139.9582 | 87.0836 | 81.0418 | 69.2341 | 148.8662 | 238.908 | 233.6991 | 209.2424 | 162.7994 | 194.8244 | 149.908 | 153.8746 | 134.8997 | 216.0752 | 178.8913 | 195.3427 | 167.1756 | 157.0167 | 158.9164 | 198.8161 | 166.7408 | 217.1087 | 182.1254 | 204.326 | 257.7826 | 186.1505 | 180.8077 | 234.9916 | 215.9164 | 230.0334 | 221 | 215.791 | 297.2257 | 170.2341 | 181.0167 | 170.2257 | 233.0585 | 174.7994 | 154.908 | 159.8997 | 175.9666 | 211.9415 | 186.2508 | 187.1421 | 138 | 132.1923 | 137.8412 | 104.1839 | 141.0167 | 261.2508 | 292.6907 | 224.1923 | 348.6405 | 296.9749 | 273.883 | 241.4431 | 309.1338 | 139.0418 | 75 | 77.0334 | 139.0418 | 113.1756 | 132.9833 | 107.1003 | 104.3845 | 106 | 124.9164 | 112.8662 | 120.1338 | 77.1756 | 92.8412 | 84.8746 | 98.1003 | 114.9498 | 108.9415 | 91.0502 | 84.9331 | 280.1254 | 279.9248 | 258.7158 | 300.1421 | 199.0502 | 201.1588 | 211.092 | 273.1338 | 306.1505 | 192.1672 | 223.9666 | 338.7576
979 | 12.7634 | 197.5457 | 184.1239 | 222.7434 | 187.3511 | 173.2478 | 177.0413 | 181.8761 | 153.5782 | 155.413 | 138.1446 | 119.0207 | 171.3924 | 235.9174 | 178.8849 | 158.4956 | 211.4956 | 228.3304 | 175.5044 | 152.0413 | 213.5457 | 191.1859 | 351.174 | 293.1033 | 171.1033 | 266.5989 | 228.7228 | 226.3511 | 281.2979 | 159.6076 | 151.7315 | 120.4337 | 131.4956 | 148.6195 | 144.1446 | 134.4956 | 96.3717 | 298.7522 | 272.6402 | 260.5044 | 245.1033 | 172.2272 | 145.2272 | 135.0826 | 134.9587 | 96.6076 | 85.8967 | 96.4424 | 132.2979 | 227.9587 | 198.5782 | 237.6489 | 140.0945 | 174.3098 | 139 | 139.0826 | 122.8967 | 225.4543 | 166 | 235.4631 | 187.7935 | 158.7315 | 148.8967 | 176.938 | 136.5989 | 229.5044 | 196.8554 | 242.112 | 231.8761 | 203.7109 | 158.5369 | 234.5782 | 206.1239 | 234.2478 | 221.1239 | 191.5782 | 323.2979 | 197.7522 | 182.8554 | 196.5663 | 239.7728 | 151.1446 | 144.062 | 147.8141 | 172.2065 | 204.4837 | 215.6076 | 204.1652 | 138.1652 | 154.5457 | 119.2891 | 125.7935 | 142.8348 | 290.5457 | 257.0119 | 246.5663 | 307.8999 | 294.1033 | 260.3098 | 292.8436 | 324.7315 | 144.0413 | 75.0207 | 81.0826 | 144.0413 | 133.6489 | 130.6283 | 119.2065 | 149.6696 | 106.3304 | 115.0207 | 97.5369 | 135.3598 | 97.9174 | 74.1446 | 70.3717 | 109.8554 | 109.1033 | 102.1239 | 96.4631 | 76.8141 | 295.1239 | 270.4218 | 225.8673 | 316.3598 | 204.8554 | 220.1446 | 221.3598 | 288.7315 | 323.112 | 212.3924 | 219.7935 | 309.6076
980 | 12.7744 | 175.6047 | 191.68 | 257.1106 | 203.3082 | 185.3624 | 179.4282 | 175.8682 | 181.0329 | 174.1435 | 144.1435 | 119.9341 | 189.7694 | 232.9553 | 127.6141 | 182.0329 | 234.8353 | 243.1106 | 153.5483 | 153.8023 | 192.4612 | 199.5388 | 312.1977 | 297.5718 | 176.5271 | 293.7812 | 262.8023 | 242.9341 | 248.4612 | 140.5059 | 139.1647 | 140.3082 | 154.6706 | 177.9671 | 151.8235 | 157.1765 | 113.2094 | 287.6259 | 302.3082 | 237.7577 | 249.44 | 181.9129 | 156.0329 | 139.2965 | 133.0659 | 78.593 | 80.9341 | 70.7577 | 99.8565 | 225.8682 | 227.4165 | 222.5483 | 192.1977 | 188.407 | 139.9883 | 143.1647 | 118.2965 | 246.0776 | 166.7906 | 210.7577 | 178.1318 | 146.9224 | 144.2635 | 174.6259 | 164.0117 | 205.2753 | 190.3624 | 200.56 | 226.3953 | 189.7365 | 183.8353 | 261.0117 | 211.3412 | 244.5176 | 227 | 218.9671 | 291.3506 | 186.2635 | 175.9671 | 176.1318 | 230.68 | 157.7365 | 147.3624 | 140.4494 | 181.5059 | 180.8235 | 196.3741 | 211.4729 | 145.8353 | 132.8023 | 133 | 116.0988 | 135.3624 | 269.9224 | 303.5176 | 227.153 | 396.1553 | 298.6706 | 274.9012 | 238.0329 | 312.593 | 145.6376 | 76.6588 | 84.7365 | 145.6376 | 117 | 113.3624 | 128.4729 | 133.8023 | 121.8023 | 115.7694 | 122.407 | 105.4941 | 93.2753 | 81.4282 | 87.0447 | 103.1647 | 113.3412 | 108.1977 | 71.1318 | 68.6588 | 300.5388 | 243.0659 | 266.8023 | 286.0659 | 197.3741 | 226.3741 | 191.8235 | 275.3741 | 283.5365 | 230.2094 | 210.56 | 291.0988
981 | 12.7854 | 163.5929 | 239.5525 | 230.3667 | 182.5477 | 195.2763 | 191.8191 | 171.9095 | 181.1406 | 150.2665 | 119.8142 | 118.0452 | 182.9095 | 261.6333 | 173.643 | 182.231 | 229.8643 | 216.0049 | 197.9144 | 149.7641 | 205.7738 | 187.0404 | 321.0758 | 285.769 | 190.8692 | 257.5477 | 257.5428 | 241.7237 | 262.7237 | 127.0404 | 144.769 | 121.0856 | 145.3166 | 177.4523 | 176.1357 | 134.2213 | 90.9499 | 307.9902 | 281.6381 | 258.9596 | 233.7237 | 151.3118 | 157.6785 | 148.769 | 135.2714 | 95.3215 | 78.5929 | 92.5477 | 125.588 | 222.588 | 267.4621 | 267.5929 | 197.0501 | 170.231 | 168.4572 | 147.0501 | 127.6333 | 219.8594 | 189.231 | 232.6834 | 181.1406 | 174.5428 | 151.1406 | 192.6834 | 134.8191 | 184.4071 | 201.1357 | 217.588 | 236.9144 | 181.7286 | 178.6834 | 232.0905 | 193.3118 | 202.4022 | 227.2714 | 217.5929 | 330.2763 | 193.8643 | 176.1761 | 180.3166 | 278.1455 | 150.3166 | 157.7738 | 181.5978 | 167 | 204.1406 | 178.4071 | 195.9095 | 141.3166 | 127.3166 | 132.0953 | 118.6834 | 145.6834 | 296.3215 | 260.6785 | 260.0953 | 342.6333 | 288.2763 | 272.3619 | 239.4976 | 328.5073 | 135.4523 | 94.8692 | 77 | 135.4523 | 116.8191 | 124.4523 | 113.3166 | 128.2714 | 115.7738 | 109.2714 | 105.1357 | 119.6381 | 72.9951 | 92.643 | 60.357 | 107.5477 | 94.2714 | 114.1809 | 76.2665 | 87.2763 | 287.6333 | 244.4572 | 261.0905 | 289.1308 | 179.6333 | 207.7738 | 215.9095 | 258.0404 | 356.643 | 207.2262 | 225.9144 | 293.8191
982 | 12.7964 | 154.8274 | 209.7478 | 216.6549 | 172.8274 | 178.9646 | 189.093 | 171.1505 | 163.7478 | 176.3097 | 137 | 119.2876 | 181.2876 | 273.1592 | 146.0354 | 166 | 227.9779 | 195.6328 | 173.907 | 184.8716 | 201.9779 | 208.9425 | 383.9867 | 303.0354 | 166.9425 | 249.9558 | 269.8053 | 257.8629 | 277.8274 | 147.907 | 160.1947 | 142.4469 | 152 | 185.6194 | 178.9425 | 158.4469 | 110.7699 | 347.584 | 274.4027 | 236.9425 | 248.7124 | 178.5398 | 171.0221 | 163.907 | 142.5531 | 82.6682 | 70.6903 | 82.3672 | 137.3097 | 234.5398 | 234.4823 | 259.5752 | 177.2876 | 154.8629 | 156.1371 | 127.1726 | 140.9425 | 236.7345 | 173.6903 | 226.8629 | 161.8495 | 184.0442 | 134.7257 | 185.885 | 130.2522 | 193.5177 | 202.0442 | 230.2876 | 214.9779 | 176.115 | 169.9867 | 236.0133 | 221.2876 | 232.5177 | 232.0221 | 208.7124 | 316.4956 | 190.3097 | 200.2522 | 186.8274 | 238.885 | 157.4027 | 152.0796 | 150.8495 | 167.5177 | 186.4027 | 189.1859 | 194.7478 | 147.0796 | 133.0796 | 113.8629 | 112.3451 | 138.7124 | 282.2876 | 275.0575 | 241.9779 | 353.0088 | 272.3097 | 280.5752 | 248.5619 | 298.8408 | 145.1726 | 72.0354 | 77 | 145.1726 | 112.5398 | 133.0221 | 119.1371 | 134.4602 | 111.115 | 116.323 | 108.6903 | 112.2301 | 92.5619 | 65.1505 | 88.6549 | 98.0575 | 99.885 | 117.1947 | 103.115 | 73.323 | 301.9779 | 235.8187 | 260.2964 | 310.4114 | 193.4602 | 205.4735 | 214.5177 | 279.4823 | 327.1371 | 214.0709 | 204.3806 | 290.6903
983 | 12.8074 | 152.5585 | 220.7433 | 211.5585 | 171.0472 | 163.117 | 203.5791 | 190.3491 | 175.6509 | 163.9528 | 136.1622 | 124.9774 | 183.6961 | 246.0719 | 163.0698 | 165.0226 | 243.0924 | 205.9302 | 157.9548 | 150.768 | 215.4867 | 207.7906 | 347.6489 | 318.2546 | 167.1869 | 280.8131 | 279.6489 | 268.5791 | 272.9056 | 130.9076 | 146.1622 | 117.9076 | 152.2793 | 163.3491 | 176.7433 | 133.2793 | 107.4189 | 307.7906 | 282.2567 | 235.7906 | 243.232 | 171.6283 | 156.885 | 145.8604 | 164.9281 | 108.5339 | 82.1396 | 71.6509 | 126.3491 | 229.0944 | 225.4415 | 267.0452 | 182.5585 | 170.6057 | 142.9076 | 130.9774 | 139.232 | 216.6981 | 183.883 | 237.5791 | 146.1417 | 152.1396 | 163 | 184 | 120.2341 | 201.6509 | 172.2341 | 234.6509 | 232.9548 | 179.6756 | 139.5133 | 265.6489 | 225.7906 | 242.3963 | 215.8604 | 205.1869 | 337.998 | 179.1396 | 187.9302 | 184.1396 | 235.883 | 162.3943 | 138.6756 | 133.4661 | 175.3019 | 193.0698 | 224.1622 | 206.8604 | 133.6057 | 118.4887 | 127.9302 | 119.1869 | 135.3265 | 287.4189 | 275.5113 | 257.8604 | 308.887 | 262.8152 | 288.7433 | 227.7926 | 328.232 | 147.0226 | 89.6283 | 77.768 | 147.0226 | 106.1869 | 118.2567 | 106.3963 | 140.883 | 112.7207 | 135.9754 | 119.4415 | 115.5811 | 68.9302 | 84.9774 | 82.9302 | 97.4641 | 98.6981 | 104.3491 | 105.9774 | 94.3717 | 317.0226 | 280.3244 | 216.2094 | 270.2341 | 200.3717 | 242.7187 | 222.6509 | 270.3019 | 315.9322 | 245.2074 | 227.4887 | 301.7906
984 | 12.8184 | 162.8737 | 205.7242 | 219.3284 | 187.1348 | 177.5074 | 158.9853 | 192.5368 | 172.2316 | 151.3137 | 124.3432 | 140.299 | 154.7242 | 284.7831 | 165.0674 | 153.1495 | 229.2758 | 203.5221 | 182.3579 | 161.0821 | 185.5306 | 206.3137 | 304.4779 | 295.6568 | 183.3284 | 262.9473 | 237.0527 | 223.9032 | 250.9116 | 143.4926 | 134.0968 | 128.9326 | 156.4105 | 168.9853 | 160.8211 | 137 | 112.6716 | 305.4105 | 296.7831 | 231.7684 | 236.6127 | 180.2463 | 180.8505 | 143.6716 | 124.8211 | 88.0147 | 84.0821 | 68.3137 | 131.8211 | 255.7684 | 217.5074 | 239.2758 | 188.9326 | 190.8505 | 154.5895 | 145.6421 | 131.1348 | 223.5368 | 169.5747 | 194.3811 | 187.5895 | 155.4779 | 162.0147 | 185.0674 | 148.9326 | 196.4253 | 202.1642 | 230.3284 | 257.7684 | 200.1116 | 170.3726 | 224.2842 | 221.0295 | 262.9705 | 216.0527 | 217.7979 | 300.0527 | 183.1348 | 188.6421 | 184.1937 | 221.1642 | 140.7537 | 159.9326 | 153.3284 | 166.3284 | 193.3432 | 211.5221 | 202.3726 | 154.0147 | 125.6568 | 127 | 132.6189 | 149.88 | 292.4253 | 269.2463 | 255.4253 | 370.7979 | 287.9032 | 272.2463 | 262.9473 | 318 | 134.8211 | 95.2905 | 86.8505 | 134.8211 | 122.739 | 134.4253 | 124.5895 | 127.2316 | 111.0527 | 110.1495 | 112 | 111.1495 | 68.2463 | 97.0147 | 80.9326 | 79.0527 | 107.8358 | 111.5453 | 119.2463 | 86.0821 | 301.701 | 257.9179 | 221.4632 | 298.1937 | 193.1495 | 200.7242 | 217.5074 | 262.8884 | 356.2316 | 209.2316 | 232.1116 | 300.2316
985 | 12.8294 | 193.6784 | 222.6401 | 223.1888 | 209.7728 | 174.6431 | 172.1711 | 165.2832 | 185.2448 | 168.0767 | 120.0207 | 163.6401 | 173.1504 | 260.1888 | 177.5664 | 166.1888 | 215.003 | 190.3039 | 165.6048 | 160.8672 | 222.7345 | 217.8849 | 320.1504 | 301.3952 | 187.0944 | 244.2479 | 256.8849 | 233.056 | 314.1681 | 146.6401 | 125.2272 | 118.8879 | 160.8112 | 179.528 | 173.9647 | 140.3039 | 108.4336 | 310.7552 | 272.1888 | 220.8319 | 275.3952 | 181.9616 | 169.5104 | 140.944 | 135.8879 | 77.3776 | 83.0177 | 81.8672 | 138.9233 | 243.5104 | 212.3776 | 225.5694 | 177.0944 | 177.1504 | 149.4336 | 161.3569 | 153.0177 | 197.2272 | 177.3216 | 221.1888 | 182.8112 | 169.4513 | 150.7168 | 195.8672 | 138.1328 | 190.472 | 201.8289 | 235.5104 | 242.4896 | 179.0944 | 143.2655 | 262.7935 | 202.0207 | 279.997 | 235.3185 | 187.0974 | 318.4689 | 206.6224 | 203.3185 | 164.7552 | 222.8112 | 140.3599 | 148.3776 | 155.3952 | 169.2448 | 186.8496 | 191.2241 | 174.1328 | 144.0384 | 133.3776 | 126.7168 | 108.4543 | 115.3776 | 287.7935 | 271.3392 | 249.6608 | 334.9056 | 294.8849 | 272.4513 | 241.6048 | 315.9233 | 145.5104 | 65.5664 | 73.9056 | 145.5104 | 129.1121 | 127.7168 | 118.6784 | 139.0177 | 134.7552 | 122.3392 | 110.3009 | 122.3952 | 71.6608 | 86.5664 | 70.0384 | 102.0944 | 104.7728 | 136.7905 | 120.7728 | 87.0944 | 278.4543 | 252.9409 | 246.7345 | 281.6815 | 204.0177 | 219.1504 | 211.8112 | 283.9056 | 364.3362 | 220.1681 | 212.8879 | 313.4336
986 | 12.8404 | 179.6131 | 200.5334 | 227.3471 | 196.5064 | 198.5064 | 193.3201 | 170.0271 | 179.0669 | 183.9459 | 146.4395 | 139.9331 | 164.7866 | 261.8933 | 181.0796 | 166.7197 | 260.4395 | 220.5064 | 179.6131 | 152.5605 | 201.4268 | 191.5207 | 313.4936 | 290.4809 | 187.3599 | 284.2532 | 226.7866 | 229.6545 | 286.4666 | 126.9873 | 136.6799 | 135.0796 | 156.3328 | 175.7468 | 209.6529 | 167.3057 | 102.1465 | 317.6799 | 273.0398 | 250.3471 | 258.9331 | 173.8137 | 182.0796 | 150.6401 | 155.2803 | 82.1736 | 64.1067 | 70.8933 | 119.7468 | 258 | 216.6401 | 275.0525 | 176.8264 | 171.5605 | 144.8535 | 137.4936 | 133.3599 | 210.4936 | 189.5732 | 220.0127 | 181.5334 | 146.8137 | 149.2803 | 186.1736 | 149.8535 | 194.0398 | 179.7197 | 247.0127 | 228.5334 | 180.1067 | 165.8535 | 275.9459 | 228.9729 | 230.9331 | 203.3073 | 233.0127 | 280.4411 | 202.6799 | 171.8408 | 171.6799 | 222.1736 | 161.9331 | 155.3073 | 140.8535 | 162.7468 | 193.9331 | 159.3742 | 188.414 | 152.293 | 137.8535 | 125.9204 | 131.0398 | 121.0271 | 303.1863 | 268.3073 | 254.6131 | 333.2532 | 268.7341 | 249.0669 | 259.1338 | 296.6401 | 158.5064 | 70.8933 | 72.0398 | 158.5064 | 111.6401 | 124.1465 | 105.2532 | 119.8933 | 140.9331 | 116 | 97.414 | 108.2803 | 79.6003 | 89.7595 | 81.2803 | 101.3997 | 94.2803 | 107.1736 | 107.6131 | 87.7866 | 302.2134 | 217.7341 | 226.8137 | 316.6799 | 189.0541 | 211 | 212.7739 | 284.4936 | 309.5732 | 194.9204 | 228.7595 | 307.8933
987 | 12.8514 | 169.5467 | 209.4022 | 244.5014 | 187.9263 | 184.4533 | 190.7649 | 186.762 | 184.6686 | 150.4759 | 123.9292 | 128.9774 | 164.9037 | 258.7394 | 165.5949 | 157.3088 | 236.8584 | 208.3569 | 170.9745 | 172.5722 | 205.4759 | 220.1671 | 328.1416 | 326.2153 | 183.9037 | 275.7394 | 225.7139 | 272.3598 | 279.0255 | 153.4278 | 136.0226 | 123.0453 | 137.4504 | 180.6912 | 183.5751 | 129.3088 | 95.9518 | 316.0708 | 264.6431 | 268.5722 | 250.7139 | 188.5241 | 163.9774 | 153.0255 | 162.6686 | 90.6912 | 68.0935 | 68.9292 | 126.238 | 258 | 219.9774 | 242.4759 | 167.3569 | 193 | 153.9037 | 150.5949 | 129.6657 | 224.5467 | 184.6912 | 195.7139 | 183.5014 | 160.2153 | 159.1671 | 195.0482 | 154.9774 | 187.7847 | 170.7847 | 221.5241 | 230.2635 | 182.1898 | 172.762 | 238.3116 | 213.4278 | 222.119 | 227.7876 | 205.9774 | 323.9774 | 200.4759 | 199.5978 | 166.0255 | 232.3569 | 151.9292 | 177.8839 | 150.3796 | 169.4759 | 182.8584 | 195.4051 | 214.6204 | 139.0708 | 144.5241 | 138.6686 | 123.4759 | 135.5014 | 292.6402 | 291.2408 | 242.1671 | 323.7876 | 297.9774 | 260.2606 | 276.4051 | 305.4504 | 145.762 | 69.762 | 66.7365 | 145.762 | 116.4051 | 119.2606 | 100.9037 | 119.3569 | 131.2861 | 114.8102 | 123.5014 | 119.238 | 92.881 | 71.9518 | 90.3343 | 90.8555 | 105 | 115.5722 | 95.881 | 86.238 | 303.762 | 246.0255 | 241.1671 | 310.7876 | 217.289 | 210.0482 | 231.8357 | 290.2182 | 305.881 | 213.4986 | 214.1643 | 306.6431
988 | 12.8624 | 181.9189 | 237.6681 | 223.8997 | 213.0619 | 174.444 | 151.7375 | 184.3436 | 161.444 | 158.0694 | 119.6756 | 116.6254 | 177.6873 | 242 | 166.4557 | 166.3436 | 220.0811 | 213.1003 | 188.6681 | 160.556 | 205.3244 | 212.4248 | 345.3127 | 312.0811 | 197.6062 | 259.7876 | 229.9498 | 241.1505 | 262.5133 | 157.5678 | 148.2433 | 150.7684 | 159.4557 | 172.9691 | 137.3127 | 136.7684 | 101.6873 | 325.9691 | 262.9189 | 255.3746 | 254.9498 | 187.2316 | 150.5752 | 132.8378 | 139.8378 | 82.5752 | 87.1932 | 64.2818 | 126.1622 | 254.7182 | 205 | 241.7994 | 170.3938 | 194.1814 | 167.3436 | 152.2433 | 142.3938 | 232.587 | 173.556 | 204.413 | 167.3629 | 145.9498 | 153.5251 | 186.2935 | 136.8495 | 203.1003 | 185.444 | 218.6564 | 209.0811 | 189.2935 | 170.0811 | 216.9072 | 223.3436 | 221.4248 | 205.5752 | 192.8378 | 308.7375 | 203.2124 | 169.4941 | 149.3821 | 232.2433 | 145.1814 | 137.1121 | 163.9306 | 173.7876 | 167.5251 | 188.5059 | 195.0309 | 148.3127 | 138.6371 | 116.7567 | 128.444 | 119.7567 | 327.1313 | 257.8378 | 236.5251 | 305.3821 | 288.1195 | 281.2006 | 272.2625 | 327.8495 | 145.9691 | 68.7876 | 85.9498 | 145.9691 | 113.3127 | 134.6873 | 115.1313 | 120.6873 | 129.1505 | 107.1814 | 105.9189 | 119.0309 | 91.6062 | 81.2316 | 78.5251 | 110.4248 | 106.444 | 105.5251 | 95.2625 | 89.444 | 299.8997 | 225.7065 | 237.8879 | 288.0502 | 176.1814 | 205.8879 | 201.9189 | 243.0192 | 306.8378 | 233.1814 | 244.3436 | 306.4941
989 | 12.8734 | 184.411 | 211.7277 | 212.4405 | 187.8614 | 185.5792 | 154.1633 | 179.5693 | 177.3167 | 181.8417 | 143.5693 | 134.7178 | 169.1436 | 242.7178 | 146.0148 | 163.7227 | 213.8564 | 223.2674 | 157.8466 | 151 | 182.7227 | 206.594 | 351.2723 | 305.5693 | 192.703 | 267.5841 | 228.3118 | 258.411 | 296.5595 | 123.4701 | 134.0197 | 133.1288 | 141.0246 | 184.7079 | 146.7178 | 125.1583 | 95.0099 | 335.2674 | 271.1534 | 241.1583 | 252.3068 | 199.9901 | 160.4258 | 145 | 151.4258 | 90.2723 | 65.3019 | 85.2822 | 116.297 | 237.0197 | 209.3068 | 217.7326 | 169.5545 | 200.8515 | 161.9951 | 134.2871 | 144.2822 | 210.7178 | 163.7129 | 224.703 | 182.1386 | 140.0049 | 158.4356 | 176.0049 | 118.7277 | 219.297 | 196.7227 | 221.7079 | 215.6334 | 181.7326 | 167.3019 | 249.9704 | 217.8515 | 214.1583 | 214.2773 | 205.1436 | 305.9951 | 195.9901 | 184.4208 | 184.1386 | 218.8811 | 151.4208 | 125.4652 | 141.594 | 176.9901 | 173.4405 | 175.0148 | 185.4455 | 152.406 | 124.8663 | 132.4159 | 135.5595 | 135.1288 | 316.8022 | 271.7227 | 238.2773 | 341.4307 | 322.4307 | 307.4208 | 274 | 306.1386 | 156.9901 | 76.297 | 80.7227 | 156.9901 | 118.8417 | 123.8466 | 114.8515 | 114.297 | 145.406 | 116.8663 | 110.8515 | 105.2822 | 88.1386 | 96.1436 | 82.4307 | 101.0049 | 114.4208 | 107.9951 | 97.2871 | 94.6932 | 287.1485 | 236.1386 | 251.9754 | 298.1583 | 183.8564 | 223.4208 | 207.4258 | 263.1436 | 319 | 237.2674 | 237.5595 | 324.0049
990 | 12.8844 | 167.5197 | 222.0776 | 222.792 | 196.4288 | 194.2207 | 178.2207 | 177.4676 | 208.1425 | 161.4415 | 140.3766 | 136.9739 | 168.2856 | 246.3766 | 168.208 | 176.3378 | 213.4676 | 204.117 | 143.909 | 151.7793 | 191.5973 | 231.8181 | 333.7532 | 302.6883 | 180.2468 | 274.4281 | 255.3244 | 240.117 | 284.1819 | 171.3505 | 155.5064 | 119.117 | 168.1947 | 172.3244 | 150.8441 | 144.4676 | 104.9739 | 315.9612 | 281.6622 | 262.1819 | 276.4415 | 189.4027 | 154.8181 | 146.4027 | 148.7793 | 76.0261 | 83.7532 | 82.7144 | 126.5712 | 263.1819 | 234.0649 | 238.1298 | 152.117 | 193.2207 | 158.6495 | 136.9351 | 140.3117 | 215.4676 | 167.6107 | 212.2468 | 179.6495 | 152.5458 | 169.6495 | 178.5712 | 132.3505 | 229.1037 | 202.9478 | 215.8702 | 282.8957 | 202.5973 | 184.9739 | 218.0522 | 216.1431 | 237.5197 | 208.8314 | 207.7144 | 299.5324 | 181.6622 | 178.8963 | 176.3505 | 250.6495 | 144.3378 | 165.6361 | 166.6622 | 165.9351 | 188.9351 | 189.2595 | 204.4415 | 126.5585 | 137.8702 | 119.4936 | 121.7793 | 121.117 | 257.6368 | 275.1425 | 230.9612 | 340.2595 | 321.5712 | 297.2207 | 268.5452 | 305.208 | 142.6622 | 86.5712 | 90.6883 | 142.6622 | 102.0261 | 113.6495 | 108 | 129.4027 | 118 | 130.3378 | 109.299 | 102.7144 | 86.5846 | 94.5064 | 85.7793 | 104.9739 | 106.8702 | 103.2468 | 98.2207 | 73.3639 | 294.9351 | 229.7532 | 218.1037 | 309.6749 | 179.5712 | 210.1037 | 205.7144 | 268.9873 | 320.7144 | 218.7405 | 223.3117 | 323.6101
991 | 12.8954 | 190.3186 | 194.5399 | 205.3806 | 213.3186 | 189.1593 | 175.3451 | 176.6371 | 168.8141 | 156.1947 | 134.646 | 127.0089 | 157.3186 | 245.6903 | 187.6371 | 184.1238 | 220.708 | 218.8141 | 144.5576 | 157.0089 | 182.4867 | 218.6549 | 323.1327 | 303.354 | 182.6283 | 234.5576 | 236.2213 | 253.469 | 296.8407 | 157.3363 | 143.177 | 136 | 153 | 151.0354 | 154.708 | 153.2213 | 92.9911 | 332.8407 | 268.6637 | 264.4158 | 266.823 | 198.0089 | 143.5044 | 152.3186 | 150.4778 | 84.1416 | 75.9911 | 89.646 | 111.3629 | 268.4424 | 227.4867 | 230.0177 | 169.3363 | 190.0089 | 164.0884 | 143.5133 | 140.1504 | 216.823 | 196.4867 | 216.9823 | 191.3097 | 184.1327 | 177.9468 | 163.531 | 119.177 | 209.177 | 181 | 225.1593 | 238.3274 | 192.8141 | 173.1593 | 237.4778 | 239.6283 | 255.7787 | 232.1416 | 216.1593 | 299.1859 | 172.3629 | 199.1593 | 166.0354 | 256.2565 | 157.1681 | 142.1681 | 152.9911 | 173.5222 | 194 | 172.3629 | 193.9823 | 137.5222 | 148.3363 | 130.646 | 128.5222 | 138 | 303.8762 | 242.2036 | 241.9557 | 325.5487 | 305.0177 | 288.292 | 242.1947 | 329.5133 | 135.2124 | 68.1681 | 85.8053 | 135.2124 | 115.354 | 127.8319 | 112.2036 | 133.469 | 116.9911 | 139.3008 | 134.9734 | 115.0266 | 101.7964 | 83.6903 | 86.8053 | 93.3274 | 114.4778 | 106.8053 | 95.3451 | 99.1859 | 302.5222 | 226.1947 | 204.2302 | 275.885 | 165.708 | 196.062 | 219.3717 | 291.4601 | 326.8053 | 237.3186 | 224.4956 | 291.6903
992 | 12.9064 | 177.3087 | 216.2783 | 232.9022 | 205.1804 | 187.887 | 176.4065 | 164.6913 | 162.6913 | 170.737 | 127.5109 | 129.113 | 153.5109 | 257.0152 | 168.2935 | 165.7739 | 241.6544 | 215.7587 | 171.2261 | 162 | 177.5261 | 216.6913 | 302.9174 | 317.7065 | 167.6761 | 268.6239 | 259.0609 | 236.9022 | 301.7217 | 156.4739 | 150.6239 | 136.9022 | 153.7217 | 170.4587 | 175.113 | 173.7152 | 88.5413 | 339.3456 | 266.2783 | 221.6913 | 258.8348 | 203 | 149.4283 | 142.5565 | 138.7217 | 73.0674 | 71.9022 | 80.5261 | 130.5261 | 240.15 | 218.3761 | 245.7739 | 173.1652 | 191.2109 | 128.6609 | 149.737 | 131.3609 | 209.5565 | 183.4065 | 206.4587 | 176.3761 | 164.6391 | 153.413 | 178.8348 | 129.3304 | 216.263 | 185.6913 | 223.3456 | 234.8196 | 184.5261 | 167.3761 | 224.2783 | 216.9174 | 229.3304 | 221.2478 | 210.737 | 309.0978 | 192.0674 | 193.9174 | 180.9478 | 213.1804 | 158.3609 | 147.15 | 146.1957 | 182.5717 | 194.9022 | 185.2109 | 186.3456 | 150.1804 | 149.0978 | 119.5413 | 139.0152 | 139.4435 | 324.2109 | 265.1652 | 226.7435 | 346.6848 | 315.7217 | 268.7891 | 261.2478 | 329.7826 | 157.3913 | 72.6087 | 71.0826 | 157.3913 | 123.9326 | 127.9022 | 130.6544 | 119.9696 | 115.4283 | 121.3609 | 120.9022 | 128.3761 | 84.2478 | 93.9326 | 74.2478 | 94.6913 | 102 | 92.6239 | 100.9174 | 105.4891 | 313.0152 | 239.4739 | 237.6544 | 301.0304 | 189.7217 | 226.3087 | 237.9326 | 264.4891 | 322.5478 | 231.3456 | 219.2935 | 308.9696
993 | 12.9174 | 163.1308 | 213.9636 | 234.6873 | 205.4218 | 200.6145 | 152.7345 | 184.0727 | 181.3018 | 166.1564 | 143.76 | 117.3491 | 172.0727 | 248.1927 | 162.3599 | 189.3018 | 228.3018 | 240.2074 | 149.24 | 160.2655 | 193.4327 | 229.1346 | 298.5782 | 329.8073 | 195.4946 | 279.2763 | 230.709 | 237.7237 | 295.5565 | 142.88 | 155.4946 | 134.8328 | 156.0364 | 168.9636 | 162.5782 | 134.9017 | 90.8073 | 342.6764 | 255.6764 | 246.2763 | 254.9745 | 197.411 | 164.2291 | 138.12 | 142.5782 | 85.4582 | 74.2655 | 91.4218 | 133.1346 | 256.1091 | 208.1091 | 271.4218 | 180.3018 | 178.8181 | 166.8436 | 144.9636 | 131.2655 | 205.5054 | 160.6982 | 206.8909 | 170.9273 | 162.2291 | 173.5673 | 166.1091 | 149 | 220.4582 | 199.8692 | 227.2546 | 240.3599 | 192.5309 | 161.1564 | 220.6145 | 211.4218 | 246.3018 | 235.2291 | 200.9528 | 306.5418 | 205.2291 | 189.7709 | 164.1672 | 217.2763 | 161.9273 | 161.3746 | 139.1564 | 170.4691 | 199.5782 | 176.0945 | 192.1819 | 146.9528 | 144.0364 | 124.3127 | 124.411 | 148.12 | 303.5309 | 277.6982 | 276.3963 | 314.3382 | 318.4218 | 282.9164 | 270.6037 | 299.8909 | 144.8545 | 87.4582 | 77.9273 | 144.8545 | 110.7709 | 131.2291 | 118.6509 | 135.1091 | 131.5782 | 119.7237 | 124.2291 | 119.2655 | 95.3382 | 78.4582 | 87.4582 | 109.0619 | 102.5782 | 101.3491 | 99.8545 | 82.1091 | 301.3018 | 236.0945 | 228.3491 | 288.9745 | 192.0364 | 209.8181 | 221.1091 | 248.24 | 361.8947 | 240.6509 | 210.6618 | 316.5928
994 | 12.9284 | 192.1799 | 220.6652 | 225.205 | 200.9498 | 197.1549 | 161.0251 | 179.0752 | 170.8201 | 168.1298 | 126.205 | 120.9498 | 166.6652 | 242.8495 | 181.8244 | 171.6445 | 222.7405 | 196.6401 | 168.4351 | 154.0251 | 207.4896 | 196.1003 | 307.1505 | 323.8746 | 183.1549 | 293.41 | 260.5398 | 225.4351 | 257.6401 | 141.3806 | 143.7699 | 115.9706 | 150.1549 | 165.0044 | 168.0752 | 164.3097 | 90 | 310.3599 | 237.301 | 257.3348 | 272.2094 | 176.8702 | 158.7448 | 146.1799 | 138.4395 | 81.0502 | 70.0752 | 84.4896 | 104.2007 | 244 | 202.1505 | 264.6947 | 168.59 | 194.4896 | 158.7197 | 150.8451 | 123.1799 | 212.5147 | 172.8201 | 219.8201 | 181.8702 | 159.615 | 161.7655 | 157.2803 | 149.615 | 214.615 | 172.4602 | 204.3304 | 260.0294 | 179.4351 | 163.5398 | 219.41 | 212.0752 | 240.9455 | 232 | 193.2257 | 312.795 | 199.1298 | 191.3599 | 186.5147 | 233.4602 | 170 | 145.0752 | 141.9498 | 181.5147 | 193.7993 | 210 | 175.6903 | 130.615 | 141.0251 | 138.1003 | 103.4602 | 150.2345 | 289.41 | 287.9749 | 285.6739 | 304.7405 | 309.8495 | 267.1298 | 252.5104 | 313.2301 | 155.6696 | 78.1298 | 90.5104 | 155.6696 | 114 | 128.41 | 115.4602 | 122.385 | 130.9248 | 110.5104 | 120.795 | 111.59 | 80.615 | 74.8702 | 84.8953 | 86.3555 | 106.6401 | 106.1799 | 113.1298 | 68.9749 | 287.7448 | 268.1549 | 233.795 | 313.59 | 187.1799 | 232.2552 | 210.4351 | 271.5354 | 310.4853 | 233.7699 | 221.4896 | 278.9455
995 | 12.9394 | 190.5212 | 227.9577 | 222.9577 | 192.5654 | 192.1731 | 166.5212 | 184.2635 | 170.7404 | 161.5635 | 129.6077 | 114.7385 | 172.2192 | 228.7788 | 153.3481 | 152.4327 | 242.8269 | 206.0423 | 170.3058 | 153.6538 | 187.175 | 212.8692 | 324.6981 | 305.5212 | 175.5654 | 285.0039 | 247.0885 | 225.1327 | 267.9115 | 163.4808 | 138.7827 | 141.6962 | 143 | 145.5173 | 180 | 147.3923 | 89.7827 | 314.2135 | 292.3885 | 255.0404 | 250.475 | 182.1327 | 153.9115 | 141.4788 | 121.6923 | 89 | 78.7404 | 70.2596 | 134.2615 | 252.0404 | 226.6519 | 243.6538 | 165.2615 | 175.9135 | 150.7808 | 153.0019 | 120.6519 | 197.4788 | 176 | 222.1308 | 188.6538 | 160.6962 | 177.1346 | 169.5654 | 150.9135 | 220.6942 | 179.8269 | 230.2154 | 236.2596 | 183.0442 | 162.2577 | 229.9096 | 223.7827 | 254.0039 | 232.8692 | 222.4808 | 310.2615 | 186.4788 | 181.7404 | 168.8712 | 234.7423 | 172.1731 | 158.3038 | 141.1712 | 163.8712 | 174.6038 | 207.175 | 188.9135 | 129.7404 | 144.7827 | 147.4808 | 112.5654 | 122.6962 | 292.9558 | 278.5673 | 234 | 334.3885 | 289.6942 | 265.3442 | 268.2615 | 317.7827 | 138.9115 | 70.0423 | 98.4385 | 138.9115 | 113.7827 | 122.3942 | 119.7846 | 121.3038 | 121.6077 | 126.0442 | 120.2173 | 113.2596 | 84.3038 | 82.3058 | 92.6115 | 107.7404 | 108.4365 | 106.0423 | 102.2173 | 68.9115 | 284.8673 | 257.9577 | 240.1712 | 307.0019 | 181.8269 | 244.2596 | 214.9135 | 283.4385 | 323.5654 | 229 | 210.9538 | 301.3481
996 | 12.9504 | 191.1784 | 219.7552 | 215.2144 | 195.8216 | 202.9848 | 175.6736 | 167.1176 | 156.9336 | 174.556 | 135.0968 | 122.148 | 163.2656 | 253.8977 | 154.9032 | 172.6888 | 230.8672 | 217.6888 | 159.5256 | 142.3624 | 185.9543 | 220.592 | 300.1632 | 310.5408 | 174 | 252.444 | 236.444 | 213.1839 | 287.74 | 141.3776 | 139.8368 | 138.148 | 147.592 | 190.6736 | 176.3264 | 140.9848 | 90.3776 | 339.2905 | 311.1936 | 270.9129 | 275.9488 | 165.1328 | 164.7856 | 140.3624 | 150.2448 | 87.6224 | 72.5104 | 83.3776 | 131.2144 | 282.8368 | 229.2296 | 230.9848 | 163.592 | 177.74 | 161.5712 | 138.2144 | 124.148 | 195.2144 | 185.8727 | 213.26 | 175.296 | 155.5408 | 143.74 | 169.3776 | 150.2144 | 229.408 | 178.888 | 242.3416 | 244.0968 | 178.0664 | 184.8824 | 248.2241 | 224.8368 | 223.0512 | 227.5048 | 199.9184 | 305.3776 | 189.2656 | 176.888 | 160.6584 | 209.3624 | 176.7856 | 157.7192 | 163.556 | 154.0512 | 220.2296 | 193.3264 | 187.296 | 124.1992 | 143.3112 | 129.2808 | 114.444 | 119.8368 | 296.556 | 264.5104 | 239.0512 | 347.9129 | 300.9336 | 290.964 | 264.5256 | 317.4592 | 148.408 | 83.0664 | 72.2656 | 148.408 | 116.2144 | 98.6736 | 105.1328 | 128.0664 | 124.112 | 117.852 | 117.0968 | 120.8672 | 88.3112 | 71.7552 | 68.5256 | 95.7704 | 94.0664 | 112.8672 | 104.8368 | 86.2144 | 314.7247 | 245.852 | 261.8672 | 301.1687 | 179.5104 | 239.6985 | 213.296 | 263.0055 | 319.2448 | 230.148 | 234.852 | 306.1176
997 | 12.9614 | 182.9824 | 224.4041 | 223.3392 | 205.469 | 204.292 | 185.5811 | 188.5811 | 156.0472 | 165.4189 | 125.6283 | 128.4189 | 183.2595 | 229.6756 | 172.1121 | 173.0649 | 213.6135 | 220.4838 | 171.1446 | 157 | 214.0649 | 233.0973 | 303.4337 | 312.6283 | 178.8378 | 265.9351 | 251.3865 | 247.3865 | 312.5635 | 145.5162 | 142.6135 | 140.3068 | 155.9852 | 193.0825 | 164 | 145.9203 | 99.354 | 304.5959 | 275.5663 | 225.7906 | 265.531 | 174.646 | 156.9027 | 154.0324 | 139.5487 | 86.8702 | 90.7581 | 84.1298 | 142.2419 | 275.4542 | 222.9852 | 239.0649 | 175.1298 | 193.8554 | 147.9824 | 151.9027 | 127.5162 | 211.8054 | 208.8406 | 196.177 | 179.6135 | 157.1446 | 155.2595 | 174 | 158.3392 | 217.8702 | 196.4365 | 218.5487 | 238.4986 | 187.177 | 172.8526 | 217.9647 | 220.5987 | 239.5162 | 196.3392 | 205.6608 | 301.0501 | 212.1622 | 193.7109 | 179.1298 | 224 | 161.646 | 144.1121 | 154.4189 | 171.4838 | 213.5014 | 183.9027 | 195.4838 | 143.3068 | 140.5162 | 143.8554 | 122.1298 | 135.4337 | 283.3068 | 276.7109 | 260.1121 | 317.3039 | 301.2567 | 262.5487 | 271.3068 | 322.6283 | 136.8702 | 89.0324 | 99.0324 | 136.8702 | 128.4513 | 110.0324 | 120.6932 | 130.646 | 110.6756 | 115.6932 | 107.6283 | 116.1918 | 81.646 | 79.3068 | 78.9351 | 94.5162 | 99.7906 | 99.2419 | 106.646 | 92.646 | 336.0176 | 246.354 | 255.7405 | 329.6636 | 201.6283 | 191.0472 | 214.7109 | 303.469 | 314.8378 | 234.2419 | 226.1622 | 323.7109
998 | 12.9724 | 199.8662 | 210.2542 | 214.7458 | 192.3712 | 177.5084 | 178.5084 | 184.0502 | 176.0167 | 174.5251 | 136.2375 | 134.4749 | 167.1338 | 263.3043 | 177.1204 | 167.5084 | 209.6087 | 219.9666 | 172.612 | 153.1873 | 204.6957 | 233.1505 | 341.1338 | 325.2709 | 192.9833 | 276.3211 | 257.1706 | 254.1873 | 318.0535 | 144 | 130.9833 | 138.5585 | 135.2709 | 166.3211 | 161.9666 | 140.4214 | 108.1706 | 323.8294 | 250.5753 | 227.0836 | 278.3746 | 159.4749 | 168.796 | 147.9498 | 130.9331 | 94.6789 | 87.4582 | 73.5251 | 133.8495 | 236.5585 | 208.8796 | 231.4749 | 164.0167 | 188.0669 | 163.087 | 159.7291 | 130.3211 | 227.9666 | 179.0334 | 208.0836 | 169.2542 | 165.7291 | 135.3211 | 177.0502 | 151.5251 | 230.5084 | 181.796 | 211.9666 | 253.3579 | 197.5585 | 199.3244 | 255.0368 | 192.8294 | 236.7291 | 192.8294 | 205.1037 | 291.8094 | 194.9666 | 171.7124 | 167 | 222.4749 | 155.8796 | 164.1171 | 161.4916 | 174.0167 | 192.5418 | 189.6957 | 201.3211 | 141.3043 | 132.1371 | 130.9498 | 119.1505 | 160.4247 | 286.388 | 266.6589 | 263.087 | 355.4415 | 318.9164 | 266.388 | 260.408 | 333.7458 | 148.7458 | 85.2375 | 94.9833 | 148.7458 | 134.5251 | 112.0836 | 123.9666 | 119.0334 | 134.3913 | 124.0502 | 122.8127 | 137.1204 | 73.0836 | 73.7458 | 82.7124 | 93.2542 | 98.5418 | 104.0669 | 91.9833 | 84.592 | 313.0334 | 253.3913 | 271.8662 | 299.1338 | 216.3043 | 217.3712 | 196.0167 | 287.0669 | 322.3746 | 231.1873 | 213.0334 | 302.4749
999 | 12.9834 | 181.8624 | 208.6018 | 216.6647 | 206.0058 | 182.9971 | 176.0029 | 189.2694 | 178.4671 | 175.8023 | 130.0029 | 129.7335 | 189.4012 | 262.8741 | 154 | 164.7365 | 227.6676 | 217.1977 | 155.7277 | 144.6647 | 198.4641 | 246.9429 | 359.6706 | 327.9341 | 190 | 278.6076 | 248.3294 | 241.4012 | 289.1289 | 153.593 | 131.4641 | 139.407 | 139.5329 | 177.6676 | 161.3294 | 168.3353 | 92.1347 | 334.5359 | 266.9312 | 242.7994 | 264.2606 | 161.9283 | 176.4671 | 140.0659 | 119.7994 | 85.9971 | 76.5359 | 76.1347 | 109.3982 | 240.3382 | 233.5988 | 229.1318 | 168.0659 | 194.1376 | 140.3953 | 154.6676 | 138.7365 | 224.3982 | 185.5329 | 220.0688 | 169.7335 | 162.7994 | 148.5329 | 184.4012 | 149.0717 | 226.4041 | 190.2665 | 214.5271 | 230.9312 | 201.3382 | 174.6618 | 222.5959 | 207.5329 | 239.3953 | 209.3982 | 189.9253 | 337.94 | 199.1259 | 168.4641 | 168.8653 | 222.7965 | 175.2694 | 172.4758 | 162.6647 | 178.3324 | 198.1347 | 178.6676 | 211.0688 | 148.6018 | 122.1259 | 122.5329 | 138.8053 | 142.47 | 311.2665 | 285.1405 | 237.7306 | 335.6735 | 293.8741 | 293.1318 | 244.2665 | 328.7365 | 147.2006 | 76.6047 | 86.6706 | 147.2006 | 132.8711 | 120.8711 | 115.8682 | 124.2006 | 113.9341 | 132.2006 | 126.8053 | 116.9312 | 84.8023 | 74.5988 | 77.3324 | 96.3982 | 103.0688 | 112.0029 | 95.3953 | 94.3382 | 319.5329 | 243.3265 | 240.2723 | 320.0688 | 218.0058 | 233.1376 | 205.1289 | 295.8023 | 298.9341 | 221.5988 | 221.3982 | 303.8624
1000 | 12.9944 | 186.5707 | 208.4115 | 226.7876 | 187.6239 | 188.7035 | 168.6239 | 179.5575 | 178.3938 | 170.3452 | 116.2124 | 123.7035 | 183.8849 | 237.5089 | 152.6062 | 158.5753 | 226.5089 | 224.8849 | 173.3761 | 161.1992 | 200.1946 | 215.1814 | 344.1637 | 316.9159 | 188.0487 | 253.9027 | 262.7212 | 236.4425 | 310.739 | 178.0487 | 137.6548 | 111.0487 | 144.9027 | 174 | 176.8363 | 157.3761 | 88.3938 | 341.1682 | 271.0973 | 249.4602 | 279.4734 | 178.2124 | 176.9513 | 138.5398 | 119.7699 | 91.7035 | 69.7876 | 75.739 | 120.739 | 226.739 | 233.8186 | 227.4734 | 166.8186 | 179.1151 | 154.9336 | 149.0487 | 120.5885 | 231.2788 | 183.0973 | 210.8849 | 173.4602 | 162.2124 | 141.0796 | 181.1151 | 141.1505 | 215.4602 | 192.1151 | 235.4911 | 249.8718 | 185.7876 | 188.8186 | 236.8849 | 213.4602 | 249.1946 | 213.4911 | 222.2301 | 309.6548 | 220.3097 | 178.8363 | 172.8849 | 241.854 | 162.4911 | 145.5266 | 160.8008 | 181.1637 | 201.9204 | 174.1637 | 193.8008 | 143.3938 | 150 | 134.593 | 118.1637 | 134.739 | 312 | 265.2478 | 246.3274 | 310.2788 | 273.5266 | 282.553 | 249.1814 | 325.9204 | 156.4293 | 68.1992 | 73.9513 | 156.4293 | 112.0973 | 110.6903 | 113.6239 | 122.2788 | 110.4425 | 127.7699 | 111.4602 | 117.1946 | 77.6726 | 82.0664 | 84.0664 | 105.5089 | 93.8849 | 102.6726 | 114.1151 | 80.1814 | 319.6062 | 258.9645 | 219.8363 | 315.3452 | 194.8849 | 220.6239 | 218.0973 | 294.5266 | 298.2301 | 227.3938 | 228.2788 | 315.261
1001 | 13.0054 | 169.2243 | 225.5821 | 228.761 | 197.8732 | 187.2243 | 176.5447 | 180.1268 | 173.8505 | 179 | 114.8211 | 115.239 | 182.4553 | 240.5074 | 153.0748 | 167.47 | 235.0374 | 217.9252 | 166.9105 | 180.6342 | 184.9479 | 222.2163 | 334.7237 | 299.7837 | 192.0227 | 272.7316 | 249.1936 | 241.1121 | 308.4473 | 164.8505 | 134.239 | 112.6942 | 149.1789 | 178.6569 | 175.7984 | 147.1268 | 94.3284 | 366.701 | 268.8585 | 261.2911 | 266.6942 | 179.4406 | 175.3058 | 139.5153 | 122.4406 | 82.0748 | 74.3805 | 83.3431 | 131.8358 | 246.5674 | 224.7463 | 223.4259 | 159.7837 | 181.4179 | 149.6195 | 149.239 | 114.1869 | 223.2684 | 179.6942 | 217.0227 | 181.2163 | 151.7984 | 122.9479 | 186.3284 | 167.3805 | 222.9252 | 192.3805 | 227.5447 | 267.5221 | 186.0147 | 182.3658 | 234.5821 | 222.3805 | 223.47 | 200.8879 | 215.1936 | 296.6342 | 194.8505 | 185.6569 | 167.6716 | 257.8732 | 155.4179 | 167.6195 | 137 | 173.1789 | 222.0895 | 175.2016 | 171.7463 | 147.5821 | 151.1642 | 156.1789 | 116 | 152.239 | 304.7237 | 272.0668 | 244.3284 | 319.1495 | 289.7984 | 246.9853 | 259.4179 | 335.6115 | 172.3205 | 91.1268 | 79.8732 | 172.3205 | 108.4032 | 125.9626 | 120.3805 | 124.4553 | 107.8358 | 116.761 | 122.7089 | 98.1642 | 82.5821 | 86.2163 | 81.8131 | 108.7984 | 93.6195 | 108.4553 | 114.0895 | 90.4179 | 316.5821 | 232.1268 | 233.3511 | 313.5221 | 188.2163 | 230 | 214.6942 | 303.231 | 301.9626 | 232.7463 | 224.9252 | 303.8732
1002 | 13.0164 | 198.3933 | 217.293 | 217.5167 | 199.09 | 198.496 | 171.4833 | 181.3368 | 160.9769 | 185.9769 | 132.9435 | 124.3599 | 182.6632 | 246.5606 | 160.5732 | 157.0334 | 233.9331 | 206.7866 | 164.293 | 165.1465 | 206.707 | 210.8766 | 313.3599 | 308.09 | 197.316 | 285.203 | 231.9539 | 249.3264 | 280.887 | 150.4602 | 144.2699 | 125.9666 | 134.09 | 188.7866 | 168.91 | 142.8766 | 101.8201 | 331.5502 | 288 | 261.3933 | 277.2364 | 190.293 | 155.9666 | 158.7762 | 141.7866 | 86.5398 | 70.4268 | 72 | 128.6967 | 268.0334 | 226.5732 | 251.8997 | 165.3599 | 185.7636 | 157.8201 | 155.3264 | 131.4164 | 204.1234 | 190.5398 | 235.6632 | 178.0669 | 155.2238 | 138.6401 | 185.3264 | 166.7636 | 211.7866 | 186.6067 | 222.7866 | 239.2699 | 166.2699 | 190.2699 | 236.3033 | 218.4268 | 223.9539 | 209.3576 | 188.8535 | 292.0669 | 187.7405 | 191.8431 | 166.5502 | 255.4498 | 156.1234 | 171.8766 | 141.5502 | 157.7866 | 207.113 | 189.6736 | 177.2134 | 145.3599 | 142.7762 | 147.7636 | 118.7301 | 152.8662 | 284.2699 | 250.5836 | 246.3599 | 332.9331 | 279.8766 | 267.0334 | 264.9769 | 314.2238 | 141.1234 | 83.8431 | 84.1234 | 141.1234 | 125.5167 | 125.8535 | 123.4037 | 126.1799 | 106.2134 | 105.2134 | 124.7301 | 104.0334 | 83.6967 | 75.4833 | 62.1569 | 95.8431 | 97.2699 | 110.4833 | 107 | 85.0565 | 304.6528 | 232.7301 | 237.3391 | 286.7866 | 186.2803 | 226.6632 | 222.8097 | 282.1569 | 300.3368 | 235.4833 | 216.8201 | 316.9204
1003 | 13.0274 | 195.1062 | 214.2538 | 221.6313 | 197.9469 | 169.888 | 172.4189 | 189 | 173.8436 | 187.1651 | 140.4218 | 114.1062 | 176.2625 | 276.2124 | 162.5753 | 164 | 224.1033 | 211.5753 | 148.944 | 161.2596 | 210.0618 | 213.5753 | 310.9971 | 313.5753 | 170.0473 | 262.1593 | 249.1651 | 230.8436 | 298.0531 | 162.3687 | 133.5811 | 124.9971 | 137.6814 | 187.8938 | 168.1593 | 138.6313 | 101.8967 | 333.7935 | 291.7876 | 263.472 | 276.9411 | 169.2625 | 152.1564 | 133.9469 | 147.2065 | 72.7375 | 80.1033 | 77.3658 | 126.1062 | 272.1593 | 226.3658 | 229.4218 | 152.2655 | 197.1062 | 161.0531 | 139.3687 | 124.1004 | 208.6844 | 175.1593 | 227.6844 | 185.3716 | 175.6342 | 148.8938 | 179.1534 | 171.1622 | 209.3156 | 187.0531 | 216.528 | 237.4189 | 173.5723 | 179.5811 | 238.5782 | 227.7876 | 253.7905 | 242.4837 | 205.6814 | 304.1062 | 208.1622 | 177.7905 | 174.7905 | 241.528 | 156.8967 | 159.4247 | 149.7905 | 165.1004 | 195.6284 | 196.4778 | 178.7375 | 136.6844 | 128.3629 | 151.8466 | 122.1593 | 130.6814 | 280.2095 | 269.2625 | 243.9971 | 329.7316 | 280.0502 | 275.8938 | 273.7404 | 350.7316 | 141.2655 | 75.7876 | 81.7404 | 141.2655 | 129.3156 | 135.4218 | 133.3215 | 126.7345 | 114.9971 | 108.6313 | 125.0029 | 112.8938 | 82.3687 | 72.3156 | 78.1033 | 84.6313 | 94.472 | 104.4749 | 104.1593 | 77.2625 | 280.3127 | 250.3629 | 186.472 | 295.6785 | 192.584 | 214.5811 | 201.7905 | 287.6873 | 320.6254 | 229.7905 | 221 | 332.2153
1004 | 13.0384 | 190.3441 | 251.0811 | 219.3928 | 198.0324 | 199.0162 | 183.6396 | 185.7207 | 166.0162 | 158.6072 | 137 | 112.2955 | 176.7045 | 275.8703 | 178.6721 | 167.6072 | 231.0324 | 217.5063 | 170.2306 | 166.1784 | 178.8378 | 231.6396 | 332.1982 | 317.8667 | 188.0973 | 262.8865 | 225.1982 | 228.591 | 294.0324 | 162.6396 | 128.591 | 132.0973 | 156.8865 | 191.0162 | 158.3928 | 143.6072 | 93.3279 | 317.3117 | 290.1622 | 281.8216 | 300.7369 | 173.3117 | 160.3117 | 130.4252 | 152.5388 | 74.2631 | 84.0811 | 89.9838 | 124.2955 | 255.5063 | 234.3928 | 236.4937 | 158.7243 | 185.4577 | 159.9838 | 135.0487 | 140.7532 | 210.2955 | 169.6559 | 222.0811 | 177.8865 | 162.6883 | 151.6883 | 195.3928 | 148.0649 | 214.2631 | 187.6234 | 217.1982 | 255.5261 | 202.6721 | 170.6559 | 246.591 | 224.8505 | 242.4415 | 197.0162 | 209.8018 | 305.9027 | 187.3604 | 176.6072 | 169.0162 | 230.3928 | 150.9513 | 151.1982 | 155.1657 | 176.8343 | 213.2793 | 185.8703 | 173.0487 | 144.1982 | 149.3928 | 127.7369 | 127.1495 | 148.2468 | 293.1982 | 270.6883 | 252.7369 | 342.8343 | 286.409 | 272.1297 | 251.7694 | 356.6199 | 132.9676 | 80.3928 | 71.2468 | 132.9676 | 131.3117 | 126.7532 | 115.8703 | 131.4252 | 125.3766 | 109.6721 | 115.6072 | 120.9351 | 80.0162 | 74.6396 | 84.0487 | 88.6234 | 104.2955 | 103.2631 | 100.9513 | 83.2793 | 306.7892 | 262.2108 | 200.5585 | 309.8505 | 170.6559 | 219.7567 | 197.9838 | 287.7729 | 339.4739 | 225.9838 | 227.5585 | 309.0324
1005 | 13.0494 | 184.5772 | 236.2369 | 216.4228 | 201.825 | 190.1956 | 187.3402 | 182.4022 | 167.4022 | 170.0826 | 130.1956 | 121.4228 | 165.1956 | 283.8348 | 176.2989 | 170.9174 | 229.7217 | 198.3815 | 169.0511 | 147.062 | 194.9174 | 233.2989 | 339.8141 | 291.2989 | 174.6391 | 279.7217 | 243.0207 | 235.2783 | 292.3815 | 164.6391 | 134.5978 | 130.5468 | 166.5772 | 194.4435 | 155.7631 | 142.8348 | 96.0413 | 314.5565 | 275.7837 | 280.8652 | 294.7217 | 172.5978 | 169.4641 | 129.887 | 130.4022 | 88.7837 | 72.2989 | 86.5565 | 132.7424 | 243.5261 | 217.1337 | 262.7631 | 189.5978 | 174.4228 | 162.6804 | 132.7424 | 126.938 | 217.3815 | 168.6185 | 215.4022 | 197.1033 | 171.9076 | 143.2163 | 186.2989 | 143.7424 | 214.1543 | 189.938 | 225.1543 | 264.8348 | 198.5978 | 173.7011 | 253.2783 | 203.3609 | 242.6087 | 201.8044 | 198.4435 | 315.2783 | 185.7011 | 186.0413 | 164.9587 | 222.3196 | 164.825 | 164.5978 | 168.5152 | 158.0413 | 210.4739 | 193.8348 | 174.1446 | 154.8761 | 143.3609 | 124.4435 | 140.4739 | 148.113 | 301.8348 | 273.1033 | 243.6598 | 323.0207 | 275.6391 | 269.6087 | 245.1239 | 330.1859 | 139.3815 | 82.5261 | 77.9174 | 139.3815 | 131.9587 | 121.7837 | 124.175 | 122.0413 | 126.8044 | 111.0413 | 123 | 120.1543 | 79.0207 | 76.6391 | 80.0413 | 99.7837 | 103.8967 | 113.0207 | 109.7217 | 85.2369 | 325.5152 | 235.3196 | 207.5359 | 293.1239 | 181.1859 | 236.3913 | 219.3924 | 309.5359 | 323.7011 | 225.9587 | 230.4532 | 310.7837
1006 | 13.0604 | 180.9351 | 227 | 236.9853 | 203.9572 | 178.7625 | 182.0073 | 179.8348 | 175.41 | 176.59 | 120.8776 | 127.5324 | 161.5177 | 260.9498 | 169.4749 | 153.1297 | 231.1224 | 204.7625 | 159.978 | 157.9351 | 184.5324 | 225.7699 | 319.8849 | 295.4027 | 167.7699 | 283.59 | 241.8274 | 231.41 | 296.295 | 159.5324 | 140.6903 | 140.4247 | 155.8849 | 196.5398 | 176.2803 | 132.6401 | 97.8849 | 315.9853 | 285.59 | 243.0502 | 294.7124 | 170.2301 | 184.8201 | 148.8849 | 132.7699 | 94.7124 | 74.9927 | 73.8849 | 131.1873 | 255.3023 | 191.6548 | 253.317 | 169.6047 | 178.4174 | 155.1873 | 145.6401 | 126.6401 | 221.6475 | 172.1077 | 219.1799 | 202.4174 | 182.3097 | 145.7479 | 183.3525 | 149.9425 | 223.4382 | 192.1077 | 210.1151 | 244.0649 | 190.9425 | 176.6475 | 242.7124 | 209.5251 | 256.3097 | 223.4602 | 206.8849 | 313.8776 | 186.8849 | 178.3672 | 164.5251 | 228.4027 | 166.2521 | 163.2876 | 154.705 | 165.1726 | 196.2301 | 185.0502 | 189.41 | 140.5324 | 146.3525 | 130.4174 | 124.4676 | 125.9425 | 283.1799 | 299.2154 | 241.59 | 339.1004 | 283.9278 | 302.6977 | 266.8703 | 336.3672 | 146.1151 | 93.9498 | 75.9927 | 146.1151 | 132.5826 | 123.8348 | 110.41 | 131.6401 | 133.3023 | 117.8201 | 123.3525 | 102.295 | 76.4174 | 77.8776 | 78.7124 | 111.3525 | 93.295 | 111.4749 | 125.2228 | 74.9425 | 316.6401 | 236.1151 | 201.3452 | 294.7772 | 200.41 | 210.3525 | 240.4321 | 287.8348 | 327.705 | 235.3953 | 213.8776 | 328.3452
1007 | 13.0714 | 188.9056 | 227.3648 | 246.5968 | 190.2144 | 190.7552 | 176.1072 | 165.1888 | 171.4336 | 183.12 | 132.0128 | 115.3392 | 177.9056 | 255.296 | 168.472 | 147.4032 | 226.1888 | 219.3088 | 184.6096 | 161.528 | 184.0128 | 215.0816 | 316.7296 | 309 | 166.5664 | 285.3776 | 238.5536 | 230.7168 | 293.176 | 153.176 | 152.6912 | 131.2144 | 149.4464 | 178.8112 | 190.0688 | 144.824 | 95.4592 | 338.7296 | 276.7984 | 249.4336 | 291.2016 | 176.9056 | 182.704 | 140.9872 | 130.1072 | 95.2144 | 85.2704 | 76.2016 | 122.6608 | 248.8368 | 193.4976 | 233.2272 | 163.9696 | 178.7552 | 145.5664 | 154.9056 | 141.0128 | 220.6352 | 184.0816 | 212.176 | 191.8112 | 160.4592 | 164.7984 | 189.472 | 153.1072 | 254.8544 | 203.7168 | 219.1072 | 236.824 | 183.1632 | 182.9312 | 227.528 | 216.7296 | 252.3344 | 235.7168 | 207.3776 | 310.056 | 190.296 | 174.1456 | 173.1888 | 243.0944 | 152.6736 | 166.0688 | 155.2704 | 169.9056 | 209.1072 | 183.408 | 193.824 | 139.2832 | 150.2832 | 129.6608 | 125.7168 | 139.3216 | 287.4848 | 309.3008 | 233.8928 | 363.5408 | 296.704 | 301.5024 | 282.4464 | 328.4976 | 142.3392 | 76.6224 | 83.7168 | 142.3392 | 131.1504 | 119.4032 | 114.824 | 147.1072 | 119.176 | 117.528 | 125.824 | 113.0384 | 74.9312 | 86.4592 | 74.472 | 116.3776 | 105.4976 | 104.6352 | 127.584 | 78.1072 | 323.352 | 224.3136 | 216.2832 | 280.7424 | 194.9744 | 215.0128 | 216.2016 | 254.2192 | 311.1248 | 241.0304 | 223.9184 | 332.704
1008 | 13.0824 | 189.2625 | 223.852 | 222.7375 | 195.9331 | 203.7375 | 186.8855 | 167.1145 | 166.7876 | 186.556 | 135.2291 | 112.5586 | 185.8044 | 249.8996 | 175.3604 | 170.5084 | 228.8687 | 233.3436 | 152.6538 | 155.7709 | 204.1982 | 208.4916 | 316.8687 | 311.6396 | 181.7709 | 289.6062 | 244.5084 | 228.3938 | 289.6229 | 150 | 125.4582 | 137.3102 | 148.0167 | 179.5251 | 167.8185 | 143.852 | 100.2625 | 343.7709 | 271.6564 | 237.2458 | 295.2124 | 178.7709 | 170.3771 | 140.7876 | 135.6062 | 108.4749 | 79.0978 | 82.852 | 137.5084 | 245.0476 | 217.7542 | 251.9665 | 182.6705 | 190.2291 | 155.4916 | 148.852 | 145.7375 | 205.6705 | 183.5251 | 208.2458 | 188 | 156.2124 | 152.4916 | 197.1145 | 164.6396 | 217.4582 | 199.5084 | 223.0978 | 254.3295 | 181.6731 | 192.3604 | 222.9022 | 212.7207 | 261.4891 | 228.8687 | 203.6873 | 305.3295 | 207.148 | 205.8855 | 168.3269 | 249.5251 | 170.7375 | 156.6396 | 155.1313 | 173.2793 | 218 | 169.6898 | 196.6229 | 150.2793 | 145.4414 | 143.3771 | 112.8353 | 157.704 | 304.3771 | 270.4749 | 232.5418 | 353.4582 | 281.7375 | 270.8353 | 259.8996 | 342.5727 | 138.0502 | 71.2625 | 81.7709 | 138.0502 | 121.9189 | 136.0978 | 119.5084 | 145.8185 | 121.6564 | 111.3938 | 128.2458 | 127.5894 | 86.2458 | 84.4749 | 79.852 | 112.6873 | 124.0978 | 110.0335 | 107.296 | 91.148 | 302.6705 | 212.8378 | 214.8353 | 293.148 | 188.9189 | 220.1145 | 213.8018 | 252.9048 | 296.2149 | 213.2793 | 227.1145 | 321.1313
1009 | 13.0934 | 196.5044 | 236.8585 | 213.938 | 180.6726 | 192.9911 | 186.4956 | 160.8849 | 171.7699 | 171.062 | 122.3805 | 128.8319 | 188.3716 | 235.9647 | 167.8849 | 170.2743 | 233.6195 | 225.5575 | 133.7346 | 155.3805 | 213.9291 | 209.8938 | 318.8938 | 317.1681 | 188.3894 | 277.7699 | 248.1681 | 231.2743 | 289.7788 | 147.6637 | 117.2832 | 121.5044 | 149.4956 | 191.2832 | 153.7257 | 142.8407 | 102.0531 | 348.4425 | 272.823 | 238.3363 | 267.6459 | 183.0531 | 169.8319 | 144.6017 | 126.4956 | 99.8319 | 76.062 | 76.7788 | 133.7699 | 224.7257 | 219.3894 | 242 | 171.4602 | 178.938 | 151.8319 | 144.7257 | 139.2743 | 183.9469 | 188.2743 | 201.938 | 181.7699 | 150.4514 | 148.8319 | 199.062 | 175.2301 | 208.5044 | 200.8319 | 218.8938 | 265.9291 | 186.6548 | 180.9911 | 235.2832 | 214.9027 | 224.5044 | 220.3805 | 188.5044 | 330.9469 | 206.9911 | 200.823 | 156.3363 | 242.9822 | 168.9469 | 170.7346 | 161.2832 | 174.9911 | 211.3805 | 190.5575 | 188.6017 | 151.9911 | 133.8407 | 144 | 113.0709 | 140.4956 | 295.6548 | 264.9469 | 243.8319 | 353.0709 | 291.6284 | 276.9116 | 242.8496 | 312.2035 | 151.0531 | 76.9469 | 87.2212 | 151.0531 | 136.9469 | 127.2212 | 126.2832 | 130.1681 | 128.2743 | 121.2832 | 125.0531 | 122.8407 | 90.8407 | 84.3983 | 76.1151 | 101.3983 | 118.7257 | 119.2212 | 119.4425 | 92.938 | 281.3363 | 231.9911 | 212.3452 | 300.7788 | 202.7788 | 214.6637 | 201.3452 | 290.4336 | 332.7346 | 220.8319 | 218.1593 | 326.8938
1010 | 13.1044 | 203.6067 | 247.113 | 208.41 | 196.4351 | 195.2552 | 177.7866 | 163.2301 | 163.6067 | 177.795 | 121.6402 | 131.8201 | 169.8117 | 248.9498 | 168.2218 | 158.7699 | 235.1632 | 207.9163 | 159.0418 | 144.5983 | 221.5104 | 207.5649 | 324.5983 | 304.1381 | 183.7782 | 276.8368 | 255.6234 | 220.5732 | 288.9916 | 159.6653 | 126.0084 | 127 | 157.6569 | 191.9749 | 168.0418 | 149.9916 | 96.9916 | 343.9916 | 262.0084 | 252.0418 | 258.4686 | 186.4268 | 173.2218 | 135.41 | 131.4435 | 94.7866 | 81.1883 | 78.4017 | 126.41 | 235.8284 | 208.3515 | 235.1716 | 187.3849 | 173.8117 | 153.6151 | 145.7866 | 141.8284 | 194.2301 | 186.0084 | 198.4184 | 177.6234 | 155.7866 | 151.8201 | 212.6234 | 174.5565 | 205.1632 | 194.1799 | 223.795 | 246.5983 | 188.8703 | 178.0335 | 250.4351 | 238.0251 | 235.2218 | 228.477 | 190.3849 | 311.1046 | 201.6234 | 198.4435 | 162.41 | 229.8368 | 174.8117 | 184.3933 | 167.1967 | 161.59 | 200.1967 | 187.3766 | 184.2301 | 132.5649 | 149.0251 | 149.6234 | 119.5481 | 145.0418 | 293.4519 | 272.0167 | 247.6234 | 357.1381 | 304.1632 | 292.9498 | 269.0251 | 296.2301 | 140.3682 | 86.4268 | 79.5732 | 140.3682 | 131.5649 | 126 | 127.3766 | 132.4017 | 124 | 131.2134 | 115.1716 | 118.7448 | 87.9498 | 93.59 | 83.8117 | 110.9916 | 120.1883 | 120.0084 | 114.9916 | 81.7866 | 280.9833 | 237.8703 | 233.0251 | 303.205 | 201.59 | 222.6485 | 207.1632 | 261.7364 | 338.3598 | 222.6151 | 217.8368 | 335.41
1011 | 13.1154 | 208.0698 | 216.414 | 226.9027 | 202.3766 | 202.1671 | 172.1721 | 174.8279 | 176.3491 | 176 | 138.0698 | 138.586 | 172.7581 | 258.3217 | 170.2045 | 152.5536 | 222.414 | 185.8653 | 169.8604 | 158.4888 | 239.8878 | 216.1122 | 328.5536 | 278.6883 | 174.3441 | 285.0324 | 264.414 | 216.7955 | 296.3491 | 167.6833 | 136.8653 | 131.1396 | 163.9252 | 193.3491 | 178.4464 | 148.2419 | 96.8978 | 338.5162 | 267.8978 | 253.7531 | 261.4738 | 182.7531 | 166.5112 | 146.0374 | 145.3441 | 97.0374 | 72.8604 | 81.8978 | 128.7581 | 255.5212 | 204.6608 | 239.4888 | 181.1721 | 175.9302 | 151.5486 | 149.2793 | 140.823 | 218.6608 | 179.0649 | 201.2743 | 192.6234 | 161.7631 | 154.4464 | 222.6559 | 164.7955 | 201.9351 | 189.4838 | 221.586 | 234.409 | 206.1347 | 179.2369 | 251.409 | 235.409 | 237.6185 | 238.1995 | 193.2793 | 299.768 | 207.1022 | 201.5811 | 167.2419 | 238.8604 | 176.1022 | 178.6883 | 155.2369 | 161.3117 | 203.5536 | 193.0748 | 194.586 | 138.6284 | 152.2045 | 153.4464 | 120.4563 | 170.7631 | 303.3067 | 269.7207 | 256.8279 | 344.5212 | 296.3815 | 282.4514 | 251.9202 | 308.2419 | 137.6933 | 84.409 | 65.4464 | 137.6933 | 121.0698 | 129.3117 | 124.3815 | 128.4464 | 130.2095 | 135.1721 | 114.1072 | 118.8703 | 78.2793 | 87.1022 | 86.7581 | 106.7581 | 116.414 | 120.1022 | 115.7257 | 80.3117 | 270.4464 | 250.995 | 231.6509 | 291.3392 | 199.2419 | 237.1022 | 195.6559 | 255.1446 | 332.2793 | 221.3766 | 236.3815 | 324.0973
1012 | 13.1264 | 206.3112 | 225.9514 | 228.3922 | 193.4263 | 192.5251 | 179.0989 | 173.0162 | 193.1313 | 173.0162 | 132.0486 | 128.6224 | 172.2788 | 269.7196 | 159.8687 | 164.5413 | 220.8687 | 202.5413 | 165.2788 | 164.9336 | 210.8525 | 218.7196 | 337.9838 | 287.6402 | 177.9676 | 274.59 | 253.4911 | 226.705 | 298.2123 | 145.1475 | 142.8849 | 135.7212 | 149.1151 | 203.3112 | 176.6888 | 158.0989 | 106.5413 | 322.6386 | 260.4911 | 238.8363 | 242.9352 | 169.1151 | 153.6726 | 149.6062 | 141.295 | 98.0486 | 66.1475 | 78.3274 | 127.4263 | 269 | 233.3776 | 271.0826 | 167.6386 | 176.41 | 149.5251 | 155.4587 | 122.8687 | 226.9174 | 174.8039 | 213.4765 | 202 | 174.0162 | 140.7537 | 222.0162 | 178.9676 | 212.1637 | 192.5737 | 216.3112 | 234.1977 | 192.4425 | 184.4603 | 222.6386 | 225.8201 | 251.3128 | 206.1799 | 195.6224 | 333.705 | 194.9011 | 185.705 | 171.9838 | 246.2139 | 176.6888 | 180.3938 | 140.8525 | 174.0989 | 208.295 | 205.7537 | 190.1637 | 157.7537 | 146.9838 | 144.0162 | 134.3598 | 174.0648 | 295.6888 | 266.9514 | 249.0486 | 349.4749 | 301.4425 | 282.7699 | 243.5591 | 314.2626 | 152.1313 | 68.4263 | 60.705 | 152.1313 | 118.8849 | 128.0324 | 128.59 | 129.2463 | 128.7699 | 142.9514 | 122.7374 | 140.6062 | 82.3274 | 78.7374 | 78.6062 | 107.5575 | 122.9676 | 111.7374 | 119.7212 | 101.624 | 272.5251 | 235.2788 | 224.2463 | 292.4603 | 201 | 248.7715 | 204.8201 | 290.3614 | 346.5575 | 225.2139 | 239.7374 | 324.3128
1013 | 13.1374 | 200.4385 | 239.7542 | 217.3491 | 191.8073 | 208.824 | 190.4385 | 176.0167 | 187.6676 | 176.8938 | 129.648 | 115.5615 | 199.8742 | 231.6844 | 155.2458 | 163.6676 | 231.1564 | 192.4581 | 167 | 145.1229 | 206.7374 | 201.7346 | 341.1229 | 302.4916 | 181.6145 | 273.0698 | 246.7709 | 225.9302 | 287.5782 | 146.1927 | 140.3855 | 131.3687 | 157.7542 | 182.9665 | 175.1062 | 169.8771 | 101.7207 | 307.7542 | 233.5978 | 245.3156 | 266.4553 | 167.2291 | 163.1229 | 139.2458 | 137.6844 | 87.3156 | 72.0167 | 71.0698 | 134.1397 | 266.8073 | 246.4385 | 267.581 | 159.7709 | 175.2989 | 156.176 | 143.6145 | 130.5251 | 211.1229 | 198.9637 | 225.1592 | 198.9302 | 176.5782 | 144.5949 | 225.8938 | 196.2095 | 205.2458 | 192.8771 | 207.8073 | 258.1927 | 174.528 | 196.6509 | 222.2262 | 224.9134 | 255.0028 | 213.6648 | 189.1397 | 322.4051 | 185.3156 | 194.5782 | 168.1062 | 259.176 | 188.7011 | 183.2989 | 150.3324 | 181.4916 | 202.4916 | 202.1397 | 186.3156 | 144.0531 | 154.5084 | 146.1397 | 115.1229 | 160.6313 | 310.3324 | 279.8771 | 235.2458 | 356.8575 | 295.3687 | 279.5251 | 262.2654 | 322.1927 | 162.4553 | 83.0335 | 74.8408 | 162.4553 | 114.6313 | 127.8938 | 124.4385 | 134.6145 | 127.7179 | 147.5447 | 124.0167 | 121.9134 | 81.2095 | 73.4385 | 80.5251 | 115.3855 | 120.9134 | 113.0167 | 107.9134 | 112.1592 | 280.9302 | 226.9134 | 237.0698 | 310.7905 | 206.7011 | 248.0727 | 213.5615 | 301.2989 | 340.352 | 238.6145 | 232.6844 | 338.0894
1014 | 13.1484 | 196.0408 | 241.5492 | 228.8849 | 198.4676 | 201.4749 | 181.9832 | 181.3933 | 168.8849 | 186.5492 | 133.8849 | 119.5084 | 214.9665 | 231.8034 | 159.3117 | 166.0743 | 237.1799 | 187.5827 | 167 | 157.0743 | 200.9017 | 218.8849 | 340 | 291.2374 | 179.1559 | 276.5492 | 245.8274 | 224.705 | 291.0408 | 143.1391 | 139.6883 | 129.8034 | 157.295 | 182.5827 | 167.2542 | 152.5157 | 87.3525 | 314.9592 | 231.6235 | 239.3358 | 273.295 | 164.1151 | 160.1966 | 145.5659 | 130.5899 | 88.5492 | 76.9425 | 79.9592 | 128.9257 | 260.8441 | 233.9257 | 252.5659 | 174.6067 | 179.7218 | 155.8609 | 143.8609 | 134.0816 | 208.1966 | 195.4173 | 212.705 | 195 | 181.8441 | 156.1391 | 236.4508 | 189.5157 | 211.1151 | 199.8609 | 214.0168 | 259.1966 | 174.4843 | 175.1966 | 241.6475 | 212 | 238.7793 | 219.6715 | 206.0168 | 322.1318 | 182.0408 | 189.9257 | 162.9592 | 251.1966 | 189.6883 | 190.4268 | 155.5899 | 165.7291 | 207.0168 | 218.5659 | 188.9017 | 154.3285 | 167.5659 | 152.1966 | 116.705 | 165.3525 | 318.295 | 274.6883 | 237.9592 | 362.9665 | 285.6883 | 288.9425 | 247.3693 | 325 | 165.6883 | 87.024 | 90 | 165.6883 | 122.0575 | 130.7866 | 117.7866 | 133.9592 | 137.6883 | 143.8609 | 133 | 112.6067 | 97.9592 | 81.6642 | 98.5084 | 112.4341 | 118.3693 | 117.9425 | 105.3693 | 100.1559 | 299.091 | 226.6235 | 241.9017 | 302.6067 | 216.2542 | 220.4508 | 226.0743 | 301.2134 | 336.5659 | 238.8609 | 220.1799 | 332.5251
1015 | 13.1594 | 203.4307 | 237.295 | 207.8849 | 195.1799 | 192.4101 | 179.7994 | 182.0944 | 172.4307 | 179.0531 | 120.7581 | 123.1475 | 186.9793 | 236.7788 | 168.8525 | 167.643 | 244.3777 | 200.6637 | 160.0531 | 176.2419 | 188.5693 | 199.2743 | 340 | 289.0207 | 180.6844 | 276 | 231 | 228.4631 | 290.0944 | 148.0413 | 143.1151 | 135.2419 | 147.9793 | 193.8112 | 175.5575 | 127.2212 | 89.4631 | 317.7581 | 272.9352 | 233.2419 | 275.0944 | 177.705 | 159.3894 | 151.3687 | 132.7994 | 86.6106 | 75.2419 | 91.0944 | 124.1151 | 259.3156 | 215.6844 | 253.7374 | 176.2212 | 185.2006 | 156.9793 | 160.2626 | 126.6313 | 214.3363 | 173.5369 | 218.7788 | 199.6313 | 177.5369 | 151.3156 | 247.652 | 170.2419 | 211.7374 | 199.5899 | 221.7581 | 247.7374 | 184.0324 | 174.3894 | 248.7994 | 213.3894 | 249.2006 | 210.2626 | 205.8849 | 327.7168 | 184.7994 | 199.9352 | 162.5162 | 252.2419 | 179.6844 | 197.8112 | 163.8201 | 168.6726 | 209.2006 | 238.7257 | 181.6637 | 162.8112 | 165.4955 | 142.5899 | 123.705 | 164.6844 | 297.4011 | 265.1475 | 246.7788 | 337.295 | 284.4838 | 286.7788 | 262.7788 | 335.652 | 161.2419 | 71.8525 | 86.7581 | 161.2419 | 131.1682 | 120.6106 | 114.5575 | 136.7581 | 140.1888 | 156.0944 | 134.3894 | 119.7788 | 95.6637 | 77.1061 | 93.8112 | 97.3687 | 138.8732 | 129.6726 | 119.3894 | 80.8436 | 307.938 | 231.8112 | 241.6106 | 303.7581 | 213.4425 | 237.6726 | 243.8525 | 296.1061 | 333.5693 | 236.2743 | 217.652 | 336.9587
1016 | 13.1704 | 219 | 227.7704 | 199.2788 | 181.8034 | 202.4263 | 191.9017 | 179.9179 | 174.6882 | 164.3441 | 137.443 | 121.4754 | 179.0492 | 241.4263 | 165.7704 | 155.7542 | 244.7044 | 191.4754 | 162.4592 | 169.0653 | 179.1805 | 202.2626 | 333.3441 | 317.8693 | 171.3441 | 261.262 | 228.6229 | 220.9179 | 294.5737 | 158.6721 | 149.8195 | 129.0162 | 141.4754 | 183.4263 | 174.3933 | 118.295 | 88.0983 | 323.5084 | 290.6391 | 250.3118 | 280.5246 | 188.1805 | 158.1475 | 139.3441 | 128.262 | 84.5246 | 82.8034 | 86.5408 | 136.0492 | 246.7866 | 213.4754 | 257.2626 | 161.5899 | 185.4592 | 160.9347 | 161.967 | 122.967 | 217.3441 | 173 | 221.0492 | 192.1637 | 182.2296 | 160.1805 | 264.2788 | 173.0492 | 208.1313 | 180.5408 | 227.9838 | 253.1643 | 175.6559 | 176.4754 | 238.0815 | 227.8363 | 235.1966 | 216.2458 | 189.1966 | 314.2788 | 195 | 216.6882 | 161.6559 | 246.4916 | 174.1475 | 189.8034 | 177.1966 | 178.4916 | 215.1643 | 238.5899 | 196.2464 | 155.7542 | 160.033 | 136.8525 | 129.9017 | 162.4754 | 291.2626 | 254.4424 | 248.5737 | 347.2626 | 289.1475 | 293.3279 | 261.7212 | 327.0815 | 159.7704 | 80.1805 | 88.2296 | 159.7704 | 133.6229 | 120.9017 | 118.1475 | 139.6559 | 149.1475 | 161.5246 | 130.7704 | 131.0821 | 83.1475 | 77.1643 | 78.6721 | 92.4754 | 145.6229 | 142.8195 | 114.3441 | 70.6397 | 283.1966 | 226.6559 | 237.1475 | 292.8687 | 215.5576 | 257.9508 | 245.5246 | 280.4754 | 318.4754 | 234.5084 | 219.0653 | 316.3441
1017 | 13.1814 | 220.4631 | 226.3894 | 203.5123 | 197.2173 | 189.3687 | 194 | 163.6843 | 166.3402 | 175.0453 | 140.9793 | 121.5123 | 165.8072 | 241.0492 | 168.2911 | 162.4631 | 219.4877 | 200.2911 | 170.586 | 163.3402 | 194.7788 | 209.2212 | 327.4631 | 328.8318 | 168.3894 | 244.0246 | 234.2911 | 215.9017 | 305.1928 | 140.3687 | 134.7089 | 123.8525 | 148.828 | 179.6352 | 175.7542 | 123.2173 | 107.4593 | 334 | 277 | 269.8771 | 275.1229 | 198.9017 | 160.8525 | 139.3157 | 122.7296 | 87.9017 | 81.1475 | 75.4877 | 134.5123 | 233.414 | 237.4101 | 257.3933 | 173.3855 | 185.2949 | 159.8318 | 156.8771 | 121.7788 | 214.3894 | 181.7788 | 220.9754 | 186.7788 | 179.2212 | 158.7089 | 256.8072 | 170.0492 | 212.0983 | 188.0207 | 225.3441 | 245.639 | 169.8318 | 182.8525 | 219.8771 | 231.7581 | 232.9508 | 211 | 185 | 324.3648 | 205.7296 | 201.0246 | 176.8034 | 221.3687 | 183.1928 | 186.1966 | 183.7296 | 169.9508 | 219.3441 | 223.5614 | 215.1966 | 154.6598 | 150.0039 | 143.414 | 130.5369 | 156.6598 | 304.0737 | 257.1928 | 229.9301 | 357.1475 | 284.0492 | 294.0737 | 248.2212 | 326.4347 | 154.4877 | 87.9754 | 86.1966 | 154.4877 | 127.0983 | 115.6843 | 111.0983 | 134.8072 | 151.8525 | 155.6352 | 126.9508 | 134.1475 | 87.3157 | 83.7827 | 79.3894 | 97.3894 | 128.8564 | 133.0737 | 114.8034 | 88.5123 | 278.0246 | 234.4877 | 235.9508 | 283.0492 | 209.3933 | 260.9508 | 233.7827 | 273.6843 | 326.3157 | 250.8525 | 204.5614 | 326.5576
1018 | 13.1924 | 221.5 | 215.4999 | 214 | 201 | 179 | 189 | 166.5 | 168.5 | 187.5 | 135.0001 | 129.5 | 166 | 241.5 | 177 | 163 | 238.5001 | 212.5 | 160 | 167 | 200 | 182.9999 | 336.5 | 316 | 175.5 | 242.5 | 244 | 218.5 | 305 | 132 | 128.5 | 125.5 | 155 | 165.5 | 182.5 | 135 | 117 | 326 | 262.4999 | 272.5 | 282 | 193.5 | 177 | 153.5 | 140 | 93.5 | 68 | 73.5 | 127 | 232 | 239.9999 | 252.5 | 186 | 171.5 | 162.0001 | 159.5 | 127.5 | 216.5 | 177.5 | 231.5 | 196.5 | 189.5001 | 150.5 | 252.5 | 172.5 | 212 | 214 | 234.5001 | 225.5 | 171.5001 | 190.5 | 249.5001 | 216.5 | 235 | 212.5 | 196 | 330.5 | 217 | 202.5 | 183.5 | 225.0001 | 182.5 | 170.5 | 177.4999 | 171.5 | 207.5 | 210.5 | 222.0001 | 152 | 149.0001 | 132.9999 | 125.5 | 158.5 | 300 | 291.5001 | 236.0001 | 334.4999 | 305.0001 | 283.5 | 239.5 | 334.9999 | 149 | 82 | 81 | 149 | 139.0001 | 118.5 | 114.5 | 129.5 | 158.5 | 149.5 | 130 | 137.5 | 94 | 80 | 88 | 103.5 | 116 | 140.5 | 112 | 94 | 278 | 235 | 225.5 | 308.0001 | 195.5 | 237.9999 | 243.0001 | 263 | 349.5001 | 251 | 217.0001 | 321.4999
1019 | 13.2034 | 204.6066 | 210.7582 | 228.0738 | 201.1967 | 180.9262 | 185.0246 | 166.2418 | 174.2459 | 170.7541 | 141.3156 | 123.1435 | 170.8033 | 246.0492 | 171.3648 | 160.9754 | 253.4139 | 208.3156 | 163.7582 | 164.5123 | 191.9017 | 185.4671 | 340.9262 | 319.5615 | 165.1927 | 243.5369 | 245 | 219.0492 | 289.4385 | 140.5369 | 134.0738 | 122.5369 | 158.6106 | 168.2704 | 168.2418 | 136.5369 | 101.4385 | 326.1967 | 264.3934 | 252.582 | 291.4385 | 179.9017 | 192.6352 | 148.7296 | 141.3894 | 92.4385 | 71.7582 | 69.9754 | 119.4877 | 219.3156 | 217.4877 | 261.4385 | 171.9754 | 165.3894 | 164.6804 | 160.0738 | 116.8279 | 210.9262 | 179.3688 | 235.3648 | 188.291 | 189.0206 | 166.3688 | 262.5369 | 179.0492 | 216 | 211.8033 | 236.4837 | 246.3442 | 195.1967 | 183.291 | 255.5575 | 218.1475 | 256.004 | 219.1229 | 207.5123 | 332.5369 | 206.2418 | 195.2664 | 172.2664 | 237.8525 | 174.1229 | 175.8073 | 155.9017 | 163.3156 | 205.3894 | 221.004 | 211.3854 | 159.5861 | 155.6804 | 138.9794 | 122 | 151.6312 | 289.7787 | 279.1887 | 249.2418 | 317.4877 | 326.9754 | 286.7582 | 241.0246 | 317.4385 | 145.5615 | 81.6598 | 70.6804 | 145.5615 | 143.2173 | 125.4139 | 127.6352 | 137 | 151.3156 | 138.7541 | 123.8279 | 139.8279 | 90.9508 | 94.1475 | 88.9262 | 108.0738 | 118 | 137.6804 | 105.0983 | 93.3402 | 282.5861 | 229.8771 | 213.5123 | 301.701 | 199.5369 | 225.295 | 251.1927 | 263.5861 | 326.578 | 252.1721 | 228.4139 | 307.7336
1020 | 13.2144 | 186.3771 | 229.3933 | 234.6721 | 197.9837 | 174.8525 | 194.9179 | 165.9671 | 184 | 159.4262 | 131.3771 | 118.4425 | 170.8688 | 255.8688 | 163.3771 | 165.7704 | 246.3279 | 193.1312 | 174 | 166.5738 | 196.9179 | 211.8688 | 330.1312 | 308.8854 | 172.9342 | 258.1146 | 257.0654 | 224.1475 | 288.0492 | 140.9508 | 146.4425 | 122.9508 | 149.8854 | 184.7704 | 163.7704 | 139.5738 | 100.0492 | 322.9837 | 264.7871 | 242.295 | 274.8362 | 177.0492 | 196.9508 | 141.1475 | 129.6558 | 84.7542 | 75.705 | 59.5575 | 124.2458 | 246.098 | 229.0654 | 232.7708 | 167.8525 | 177.7867 | 158.295 | 157.2296 | 122.59 | 200.1312 | 194 | 232.0983 | 173.705 | 178.3933 | 158.4429 | 268.7213 | 161.0658 | 212.3279 | 192.9837 | 223.3442 | 249.5083 | 182.0166 | 170.2787 | 236.4262 | 230.295 | 261.8362 | 221.3771 | 198.5575 | 326.6558 | 200.1967 | 183.9017 | 172.4425 | 230.4262 | 174.8033 | 178.0329 | 149.9017 | 160.7213 | 182.115 | 233.6558 | 202.7375 | 149.3608 | 140.9017 | 152.2296 | 138.7867 | 139.6229 | 288.3442 | 265.3604 | 241.6229 | 335.885 | 302.3937 | 305.3933 | 235.705 | 315.5246 | 163.2129 | 89.0492 | 74.7867 | 163.2129 | 128.3279 | 123.0492 | 119.3608 | 132.2787 | 139.2787 | 127.4262 | 123.295 | 133 | 90.5738 | 93.705 | 87.5738 | 105.7542 | 111.1804 | 117.1312 | 112.6721 | 99.0654 | 262.9183 | 226.5738 | 227.1146 | 297.8029 | 203.0983 | 235.4262 | 233.2296 | 269.0983 | 321.475 | 257.9508 | 228.1475 | 319.0983
1021 | 13.2254 | 188.8319 | 227.8731 | 223.5044 | 196.59 | 180.0531 | 188.5044 | 164.2625 | 182.3894 | 168.2006 | 132.7581 | 116.2625 | 176.3894 | 256.0206 | 165.8319 | 167.2419 | 241.9262 | 190.8319 | 168.6312 | 168.5369 | 206.0738 | 209.3362 | 351.4544 | 310.9588 | 188.3156 | 263.0206 | 241.6931 | 228.6106 | 300.2744 | 129.2625 | 148.5575 | 118.7787 | 150.885 | 184.0944 | 172.7581 | 145.295 | 104.2213 | 326.9588 | 247.2419 | 237.7994 | 262 | 175.8525 | 176.1356 | 154.7375 | 130.5163 | 86.4425 | 67.3156 | 69.7906 | 133.8319 | 248.5456 | 235.705 | 247.655 | 178.4219 | 173.6725 | 158.6312 | 145.0206 | 129.705 | 187.6312 | 198.295 | 242.59 | 169.0738 | 181.705 | 157.8644 | 268.1681 | 153.7375 | 206.8525 | 184.0738 | 237.5163 | 241.0738 | 176.4219 | 172.4425 | 233.3894 | 217.2094 | 242.5575 | 224.9056 | 201.8113 | 304.9675 | 192.1887 | 187.3688 | 178.3156 | 243.4956 | 164.5575 | 171.6844 | 165.18 | 158.0206 | 187.0119 | 213.5781 | 207.3362 | 134.3156 | 147.0531 | 136.7994 | 150.2419 | 137.705 | 295 | 288.4425 | 239.1681 | 342.2625 | 301.4013 | 310.3156 | 247.7169 | 330.4956 | 163.5781 | 87.3156 | 92.6844 | 163.5781 | 126.6106 | 129.3688 | 106.4631 | 132.8319 | 130.1681 | 128.1475 | 118.7994 | 126.0206 | 88.2419 | 85.3156 | 86.8525 | 110.1269 | 115.7375 | 111.0738 | 116 | 101.9469 | 264.0119 | 224.2419 | 226.115 | 317.6312 | 196.41 | 227.5575 | 218.3362 | 267.2419 | 334.4306 | 244.6519 | 236.3688 | 313.4837
1022 | 13.2364 | 209.475 | 219.4508 | 207.705 | 195.2133 | 183.1558 | 187.5325 | 173.6716 | 177.1558 | 177.5492 | 144.2375 | 115.7867 | 182.6883 | 261.5325 | 170.5492 | 144.7792 | 226.1725 | 206.5325 | 160.705 | 156.9183 | 209.1967 | 187.2708 | 346.345 | 305.9759 | 182.1558 | 257 | 262.3841 | 233.8442 | 299.0167 | 123.8442 | 144.0983 | 126.4342 | 151.0167 | 165.82 | 174.3525 | 147.9017 | 91.6234 | 311.5417 | 269.8108 | 245.475 | 265.295 | 177.295 | 168.885 | 161.2542 | 131.5083 | 80.4675 | 69.9425 | 76.6642 | 131.9592 | 237.7292 | 230.9017 | 261.23 | 182.8608 | 170.7292 | 146.8608 | 132.41 | 130.9425 | 197.2783 | 193.2133 | 243.9592 | 183.7292 | 172.5083 | 171.1558 | 266.7458 | 150.3692 | 228.6141 | 187.1967 | 230.82 | 245.8442 | 188 | 183.4917 | 248.475 | 211.4342 | 228.7625 | 222.3117 | 185.0909 | 301.885 | 197.9258 | 191.4508 | 169.41 | 237.3284 | 155.1558 | 173.4508 | 171.2133 | 171.2208 | 199.9592 | 193.7625 | 193.5083 | 134.1967 | 153.4508 | 123.6067 | 130.525 | 140.0408 | 295.5492 | 279.7217 | 271.7941 | 321.4675 | 325.0408 | 305.2542 | 244.8775 | 322.6809 | 162.4342 | 82.2542 | 85.6642 | 162.4342 | 133.4917 | 130.1558 | 109.8442 | 141.9425 | 123.8033 | 125.0575 | 131.4175 | 124.3933 | 80.6067 | 81.9017 | 80.0575 | 103.7217 | 121.705 | 124.0817 | 115.4508 | 94.4508 | 298.9258 | 223.7458 | 237.5158 | 320.6883 | 193.9425 | 217.6067 | 217.1392 | 279.9258 | 327.0408 | 239.4917 | 241 | 316.885
1023 | 13.2474 | 199.4188 | 207.7709 | 205.5615 | 193.6146 | 182.8073 | 177.349 | 161.4886 | 166.7011 | 161.1563 | 131.9105 | 124.1229 | 168.2261 | 250.2261 | 180.5448 | 141.3521 | 206.7011 | 196.9105 | 159.1229 | 148.6844 | 201.4552 | 182.7375 | 330.9833 | 295.3687 | 164.9636 | 271.0364 | 274.9803 | 250.475 | 280.4552 | 127.5615 | 146.1229 | 136.6844 | 147.0531 | 169.0364 | 170.1927 | 139.1396 | 79.5615 | 292.9302 | 284.824 | 274.7208 | 261.824 | 176.0698 | 171.3854 | 146.0865 | 130.9302 | 75.4917 | 71.1927 | 76.2989 | 130.3687 | 237.7709 | 238.9833 | 241.6313 | 180.9302 | 183.6844 | 147.176 | 146.651 | 126.0167 | 202.824 | 180.9469 | 218.2261 | 193.8771 | 157.3323 | 161.824 | 265.6313 | 139.4385 | 228.349 | 205.2823 | 221.1229 | 241.7011 | 177.8938 | 186.8771 | 250.7709 | 223.3687 | 234.9135 | 219.3687 | 186.0197 | 322.9135 | 193.5948 | 177.525 | 176.9135 | 219.7542 | 164.3521 | 164.5781 | 169.6146 | 170.7177 | 209.0364 | 190.9302 | 205.2823 | 132.0698 | 135.5948 | 130.6677 | 123.176 | 168.5812 | 279.7177 | 268.5615 | 267.7511 | 338.3885 | 322.7011 | 313.6677 | 251.8968 | 319.475 | 165.9469 | 76.0698 | 74.6313 | 165.9469 | 124.525 | 132.0531 | 124.2291 | 143.7542 | 134.9135 | 132.2291 | 128.2261 | 121.824 | 76.4385 | 86.0531 | 81.6146 | 109.4052 | 115.0698 | 116.5948 | 116.6844 | 83.8938 | 286.7344 | 213.6479 | 236.349 | 304.5417 | 196.3156 | 220.7375 | 217.9469 | 301.9833 | 315.7177 | 229.4021 | 235.3854 | 323.3156
1024 | 13.2584 | 183.295 | 227.1137 | 204.8525 | 193.41 | 196.475 | 182.0812 | 146.3112 | 157.5575 | 158.9012 | 121.59 | 121.5575 | 158.3112 | 234.5737 | 176.525 | 167.0649 | 214.77 | 177.9838 | 167.4587 | 155.7375 | 184.82 | 187.4263 | 317.3451 | 302.59 | 163.18 | 272.2463 | 243.8038 | 250.6725 | 281.6062 | 130.8687 | 145.2788 | 133.9838 | 157.3112 | 174.2625 | 171.4425 | 133.5737 | 83.4425 | 296 | 281.7212 | 274.3775 | 283.3924 | 169.5575 | 178.475 | 152.6387 | 125.9675 | 76.705 | 81.0487 | 76.2625 | 138.7375 | 229 | 223.0501 | 258.5074 | 183.4263 | 174.6725 | 173.8024 | 153.3938 | 125.3112 | 187.0988 | 191.3437 | 202.5737 | 170.0501 | 141.5413 | 161.0162 | 268.7375 | 147.6062 | 210.705 | 216.2788 | 213.9675 | 232.5575 | 163.6888 | 167.0663 | 234.5413 | 224.8525 | 223.7713 | 226.0162 | 212.0325 | 319.8038 | 181.1475 | 175.0325 | 166.9188 | 224.885 | 156.2138 | 157.4263 | 179.1637 | 158 | 211.3938 | 216.9499 | 201.9351 | 135.3112 | 119.705 | 145.3112 | 136.0325 | 165.1976 | 267 | 252.3613 | 260.5074 | 333.1813 | 294.6238 | 303.0663 | 268.4263 | 340.3275 | 145.3613 | 83.9012 | 79.4587 | 145.3613 | 120.3112 | 130.8363 | 133 | 142.5737 | 127.7876 | 127.23 | 109.705 | 117.5737 | 81.1637 | 88.2788 | 86 | 104.6401 | 109.1313 | 100.705 | 130.0487 | 77.1475 | 296.8349 | 225.2287 | 224.4425 | 289.295 | 197.295 | 213.9513 | 216.295 | 304.41 | 311.6062 | 244.9662 | 239.6062 | 328.3112
1025 | 13.2694 | 193.7906 | 237.4513 | 201.6558 | 171.8328 | 193.8653 | 192.8977 | 165.651 | 167.3068 | 163.4838 | 131.4464 | 116.6558 | 167.6883 | 237.9302 | 169.1721 | 162.5909 | 209.1769 | 194.3393 | 161.2792 | 161.1721 | 203.1997 | 201.0649 | 316.0649 | 301.8977 | 184.651 | 259.7256 | 225.2094 | 240.6558 | 299.7208 | 137.1023 | 136.3815 | 121.0373 | 151.4513 | 170 | 178.6883 | 140.4464 | 85.414 | 321.1997 | 282.414 | 243.9026 | 297.7256 | 167 | 175.8653 | 152.6932 | 121.1721 | 86.7208 | 80.6234 | 74.3442 | 123.6607 | 220.7955 | 222.4091 | 247.8003 | 182.414 | 167 | 170.0747 | 148.7581 | 121.7955 | 192.4789 | 188.5211 | 211.7906 | 155.3442 | 141.2744 | 161.0698 | 246.6282 | 151.0698 | 209 | 222.6185 | 217.9627 | 228.8279 | 159 | 156.5162 | 228.414 | 216.3815 | 221.4789 | 221.3815 | 214.4838 | 314.1023 | 169.1071 | 166.349 | 162.5487 | 224.1396 | 152.3766 | 151.7256 | 185.3442 | 163.2744 | 202.0698 | 208.2143 | 192.9302 | 152.3068 | 127.9627 | 129.4886 | 132.6234 | 149 | 284.5812 | 263.4416 | 256.2467 | 322.2045 | 286.7906 | 301.3068 | 274.4464 | 355.0325 | 131.8279 | 84.9675 | 80.8977 | 131.8279 | 120.8977 | 124.0698 | 116.0049 | 141.2419 | 126.1347 | 136.9253 | 129.0974 | 116.2419 | 90.8604 | 99.8929 | 84.2419 | 95.1023 | 95.8653 | 105.4464 | 133.1396 | 78 | 296.1445 | 246.2744 | 212.9351 | 280.4513 | 187.8653 | 208.5162 | 229.1347 | 299.2419 | 308.0373 | 248.0049 | 256.5487 | 318.349
1026 | 13.2804 | 207.7783 | 245.5566 | 228.7199 | 170.3683 | 191.9833 | 173.2468 | 176 | 160.4434 | 174.7616 | 127.6234 | 112.6067 | 169.8033 | 238.205 | 174.1883 | 164.7616 | 208.7699 | 211.59 | 164.9666 | 158.41 | 210.8284 | 208.0083 | 314.0334 | 290.6234 | 186.6234 | 261.9833 | 227.3766 | 235.41 | 292.4434 | 124.4434 | 146.5566 | 121.18 | 148.1883 | 178.975 | 168.2384 | 145.5983 | 110.7282 | 314.4685 | 277.2134 | 220.025 | 291.0083 | 155.6317 | 180.5649 | 151.7783 | 132.1716 | 91.41 | 70.41 | 74.205 | 103.4184 | 217.3933 | 219.6401 | 256.9249 | 165.8451 | 179.5649 | 164.3766 | 147.6067 | 130.9583 | 190.6317 | 175.41 | 209.6234 | 168.3683 | 143.8033 | 155.41 | 233.385 | 150.795 | 207.8033 | 196.2885 | 224.4017 | 235.7783 | 182.3349 | 166.18 | 233.9833 | 221.7699 | 228.6067 | 227.9666 | 204.82 | 324.7783 | 181.5399 | 169.7616 | 162.82 | 227.18 | 142.2468 | 165.9499 | 179.2217 | 150.2468 | 193.4184 | 184.615 | 190.8117 | 150.8284 | 134.4017 | 130.5566 | 117.025 | 140.025 | 295.205 | 274.0167 | 235.2217 | 333.9833 | 279.2384 | 317.7783 | 259.8534 | 355.2134 | 148.3516 | 82.9917 | 77.4017 | 148.3516 | 122.1883 | 112.4267 | 123.7449 | 133.4184 | 137.59 | 124.0751 | 132.0334 | 117.9917 | 93.205 | 95.8367 | 76.4184 | 100.9917 | 101.7616 | 103.4184 | 115.8367 | 82.1883 | 274.615 | 227.2634 | 217.3599 | 295.7366 | 181.1967 | 212.795 | 226.2301 | 261.5019 | 324.9332 | 250.3599 | 246.0501 | 308.5983
1027 | 13.2914 | 196.5132 | 234.6283 | 243.7257 | 179.8319 | 175.2389 | 154.6725 | 174.1681 | 173.2036 | 162.6283 | 132.3806 | 124.4337 | 160.4513 | 232.7257 | 165.3982 | 185.6018 | 215.3894 | 195.6814 | 169.3363 | 152.3363 | 196.0619 | 206 | 307.8319 | 301.4868 | 181 | 267.8319 | 228.7257 | 237.2743 | 307.0355 | 139.0355 | 146.0088 | 109.5132 | 141.8407 | 177.0619 | 166.9381 | 153.3275 | 106.0176 | 312.6549 | 277.6637 | 227.4337 | 286.5575 | 173.6461 | 192.6637 | 153.3363 | 131.0619 | 83.5044 | 67.3894 | 72.3894 | 117.9293 | 231.823 | 213.0531 | 238.3627 | 164.1593 | 170.292 | 144.9645 | 142.9469 | 139.7788 | 187.885 | 173 | 198.5044 | 167.4513 | 155.823 | 156.0531 | 233.3363 | 139.177 | 207.6106 | 173.1681 | 213.0088 | 258.0974 | 186.3982 | 166.7257 | 249.6018 | 217.3982 | 213.5663 | 236 | 199.3894 | 320.2301 | 214.3186 | 190.6018 | 167.1593 | 221.0088 | 138.9381 | 168.8407 | 152.0176 | 151.2124 | 196.9381 | 183.4425 | 203.8762 | 137.8938 | 132.1681 | 125.1238 | 138.478 | 137.6637 | 283.6194 | 298.0886 | 230 | 352.6549 | 305.416 | 297.3539 | 246.3894 | 355.6637 | 152.0619 | 68.5132 | 86.1593 | 152.0619 | 129.2743 | 118.823 | 118.0707 | 122.8938 | 125.9557 | 123.0974 | 127.0531 | 123.6637 | 85.2832 | 92.4956 | 77.4956 | 110.3275 | 111 | 100.2212 | 117.3806 | 86.2212 | 274.6637 | 241.9204 | 224.5575 | 292.6814 | 182 | 214 | 206.177 | 259.593 | 301.9202 | 242.9026 | 216.292 | 303.5044
1028 | 13.3024 | 190.9833 | 232.8353 | 220.4439 | 205.9165 | 180.6874 | 178.0311 | 166.7709 | 177.4105 | 167.0645 | 137.7542 | 134.2458 | 155.6229 | 234.3604 | 170.4582 | 162.4773 | 220.6062 | 200.1957 | 156.4105 | 159.9666 | 184.148 | 191.673 | 301.5251 | 310.1313 | 183.4916 | 254.0501 | 224.1313 | 217.5752 | 303.0668 | 143.1647 | 137.7542 | 116.4415 | 142.2291 | 175.7375 | 167.6396 | 151.7709 | 89.5084 | 303.9355 | 279.3771 | 255.9332 | 280.0167 | 185.0167 | 181.2959 | 145.3938 | 126 | 76.2625 | 67.6229 | 78.2291 | 115.0501 | 238.7542 | 215 | 223.7208 | 164.3938 | 167.7208 | 145.8186 | 132.2792 | 132.148 | 204.0812 | 182.9666 | 202.4749 | 170.0979 | 172.0979 | 150.5251 | 224.7709 | 144.0812 | 216.0979 | 183.8353 | 210.9833 | 240.9689 | 176.5084 | 170.2291 | 230.8377 | 217.4749 | 224.3103 | 224.7876 | 189.0334 | 305.2792 | 205.4439 | 196.1313 | 176.7375 | 218.9833 | 145.8687 | 154.2792 | 161.0478 | 158.3771 | 190.1647 | 187.4916 | 210.8854 | 130.8854 | 131 | 126.8353 | 129.8377 | 136.2625 | 265.7876 | 273.019 | 234.3604 | 333.9689 | 308.6897 | 284.7375 | 264.6874 | 346.1647 | 123.9522 | 74.1957 | 79.5418 | 123.9522 | 115.8043 | 136.3437 | 127.8019 | 112.7709 | 108.2792 | 123.4105 | 120.2792 | 127.8687 | 76.6396 | 84.7876 | 84.1146 | 100.0501 | 107.2625 | 109.7208 | 110.9188 | 76.4105 | 270.7709 | 244.8043 | 219.8854 | 307.1623 | 183.2458 | 218.3604 | 201.7375 | 255.9355 | 296.3103 | 217.673 | 196.9021 | 287.5418
1029 | 13.3134 | 183.2015 | 222.2143 | 206.7296 | 196.2271 | 172.3087 | 180.3903 | 154.1072 | 162.7424 | 181.176 | 142.7168 | 124.8367 | 166.7985 | 231.9184 | 174.1888 | 150.352 | 214.4719 | 202.3776 | 160.1633 | 146.3087 | 183.8112 | 199.5153 | 306.5281 | 291.8495 | 187.5408 | 222.1327 | 222.3648 | 200.1888 | 303.7041 | 131.5536 | 133.1888 | 128.9056 | 152.9872 | 165.2959 | 171.352 | 144.824 | 91.176 | 291.7296 | 258.6735 | 271.8112 | 278.2704 | 153.4158 | 159.0255 | 143.9056 | 117.1072 | 88.6097 | 73.7168 | 78.824 | 118.7041 | 224.6607 | 210.5536 | 247.4209 | 167.9872 | 167.2832 | 165.5281 | 129.5408 | 126.7296 | 217.7168 | 171.8495 | 199.3776 | 176.2704 | 168.1072 | 149.8112 | 220.3648 | 152.6352 | 219.0944 | 190.3648 | 207.648 | 232.8928 | 178.176 | 176.5408 | 221.352 | 211.8367 | 228.3776 | 217.3648 | 193.1633 | 324.7729 | 185.3776 | 198.176 | 175.824 | 231.5281 | 147.6352 | 150.2704 | 165.4719 | 149.1072 | 162.6735 | 192.8112 | 208.3648 | 145.5153 | 120.2015 | 118.1199 | 125.4337 | 137.8112 | 278.0561 | 244.4592 | 238.9056 | 325.8928 | 301.4464 | 293.352 | 265.8367 | 349.7985 | 135.4081 | 70.0255 | 74.5408 | 135.4081 | 118.0689 | 138.1888 | 123.2143 | 110.2704 | 113.7985 | 110.648 | 128.3393 | 126.4592 | 80.352 | 81.176 | 79.0128 | 99.8928 | 110.7168 | 108.648 | 105.5408 | 80.1633 | 274.6224 | 228.648 | 204.6607 | 303.9439 | 182.7296 | 213.3776 | 219.2449 | 243.0944 | 328.9617 | 222.3393 | 207.8801 | 282.5408
1030 | 13.3244 | 195.1298 | 210.0575 | 218.9499 | 178.7625 | 184.3098 | 162.9351 | 178.1373 | 166.4174 | 173.2876 | 136.2876 | 127.4174 | 168.4675 | 214.1076 | 159.0501 | 146.2301 | 220.6549 | 185.0501 | 167.9425 | 149.2449 | 196.3599 | 203.1725 | 325.5974 | 271.6401 | 176.0501 | 229.7198 | 227.8275 | 218.0723 | 307.7625 | 134.8275 | 136.18 | 103.4527 | 135.6327 | 145.0501 | 169.82 | 150.1224 | 93.6475 | 279.9926 | 257.3599 | 269.4675 | 276.41 | 146.0649 | 155.5325 | 143.705 | 116.5325 | 93.115 | 78.295 | 77 | 122.7625 | 232.5399 | 232.6048 | 244.4026 | 171.3525 | 158.82 | 183.3024 | 134.885 | 133.1224 | 200.2802 | 169.7699 | 205.3599 | 172.4675 | 154.5826 | 147.4675 | 213.5251 | 145.8776 | 206.9926 | 188.705 | 212.4174 | 231.5251 | 174.1725 | 168.9351 | 226.9425 | 211.18 | 230.4749 | 213.7625 | 208.0649 | 310.5102 | 177.7625 | 193.5251 | 156.5176 | 233.7625 | 154.4749 | 147.7625 | 171.0074 | 161.4824 | 174.9573 | 196.295 | 199.5826 | 166.6549 | 114.6475 | 122.5974 | 145.5974 | 133.5251 | 280.5826 | 270.8423 | 221.2227 | 345.2449 | 294.2876 | 294.41 | 274.2449 | 345.6401 | 142.8776 | 65.885 | 72.7625 | 142.8776 | 108.8126 | 121.1076 | 116.885 | 122.6549 | 121.9425 | 116.0649 | 145.0649 | 117.2301 | 74.9351 | 74.5826 | 85.3599 | 90.7551 | 105.5826 | 110.18 | 101.82 | 80.82 | 281.59 | 246.3673 | 210.5974 | 300.7699 | 183.295 | 216.1224 | 226.0575 | 267.2523 | 331.2876 | 236.4749 | 216.3525 | 298.2449
1031 | 13.3354 | 202.3609 | 228.1338 | 210.8249 | 183.5979 | 179.8456 | 161.9587 | 180.8043 | 169.0206 | 173.9381 | 129.6804 | 133.9794 | 169.2989 | 216.5566 | 147.3815 | 147.9381 | 234.9174 | 183.2783 | 146.5673 | 147.1032 | 194.0826 | 200.6598 | 338.9381 | 281.8349 | 176.9174 | 253.5359 | 243.5359 | 229 | 298.7424 | 127.4434 | 123.8249 | 117.3708 | 131.9174 | 169.6697 | 168.9794 | 149.3815 | 97.2989 | 286.5359 | 272.237 | 250.5054 | 278.9587 | 154.2989 | 146.1239 | 135.7424 | 140.1131 | 85.0619 | 79 | 69.0826 | 117.7011 | 217.2477 | 221.5879 | 253.4946 | 159.1239 | 161.9381 | 180.1032 | 145.5772 | 127.763 | 198.237 | 183.237 | 218.2576 | 173.9587 | 148.0206 | 144.3402 | 226.4946 | 152.5566 | 214.8555 | 185.3609 | 244.0512 | 234.5979 | 174.9587 | 182.474 | 222.7217 | 204.1032 | 237.9587 | 212.6598 | 210.3609 | 282.4641 | 184.5772 | 163.6086 | 164.1544 | 239.2576 | 146.7836 | 140.0619 | 157.2064 | 153.8456 | 188 | 197 | 209.5153 | 147.9281 | 125.5566 | 135.9381 | 135.8456 | 136.5979 | 278.6391 | 252.3509 | 247.2882 | 347.7217 | 298.237 | 295.6391 | 272.763 | 328.7836 | 133.0619 | 65.3609 | 88.1544 | 133.0619 | 110.5566 | 113.6598 | 107.7836 | 120.4228 | 123.6598 | 122.9794 | 121.6292 | 116.2989 | 88.474 | 66.0413 | 78.4641 | 89.5979 | 112.2164 | 104.4228 | 110.2164 | 83.9381 | 253.3097 | 241.5054 | 223.2783 | 301.7011 | 178.7217 | 222.6391 | 220.3815 | 275.0619 | 335.237 | 238.6804 | 204.7836 | 309.9587
1032 | 13.3464 | 192.9351 | 227.9189 | 205.6883 | 199.0974 | 165.6234 | 170.0487 | 179.3766 | 174.0487 | 154.8215 | 129 | 124.9189 | 173.6883 | 210.5747 | 151.0487 | 155.0325 | 212.1169 | 206.1623 | 154.4902 | 166.8668 | 190 | 186.8864 | 302.3475 | 301.4253 | 187.0487 | 257.3117 | 241.9351 | 237.737 | 284.2468 | 139.474 | 131.4578 | 114.526 | 140.0325 | 184.6558 | 172.6883 | 154.3928 | 89.5909 | 297.7045 | 277 | 246.7045 | 254.1169 | 178.1785 | 158.1461 | 137.3766 | 128.8215 | 91.4091 | 76.3117 | 66.3442 | 111.2955 | 247.3408 | 206.3117 | 236.4611 | 158.3766 | 164.3279 | 154.8377 | 149.3279 | 123.6721 | 209.0487 | 179.9351 | 215.2792 | 172.6396 | 169.1785 | 161.474 | 223.5747 | 160.0162 | 213.263 | 170.5585 | 241.1981 | 250.7695 | 178.3442 | 180.5747 | 230.7532 | 205.0487 | 245.3766 | 225.0974 | 202.2792 | 322.3733 | 184.9675 | 190.3246 | 174.3442 | 220.8215 | 149.0649 | 163.211 | 159.7695 | 159.7857 | 188 | 201.0325 | 216.3279 | 142.3928 | 133.6883 | 123.5422 | 132.3766 | 135.9675 | 278.6558 | 261.8831 | 270.0325 | 337.6072 | 298.2955 | 276.8377 | 274.7208 | 329.0487 | 130 | 76.0974 | 88.9351 | 130 | 119.3604 | 114 | 110.737 | 136.8344 | 120.6396 | 129.3766 | 117.0811 | 122.7045 | 93.2955 | 78.1136 | 78.7208 | 93.6721 | 105.9026 | 104.0325 | 125.4091 | 84.3117 | 260.8506 | 222.2468 | 238.0974 | 324.1948 | 191.4578 | 219.9675 | 205.2305 | 280.0649 | 326.5585 | 223.8864 | 219.8344 | 314.6883
1033 | 13.3574 | 189 | 209.997 | 218.6343 | 181.7315 | 184.8998 | 178.4749 | 187.4749 | 183.8437 | 176.8496 | 133.1062 | 127.528 | 178.4218 | 222.7936 | 164.2655 | 166.5811 | 218.1623 | 210.5251 | 155.1032 | 156.4159 | 182.472 | 176.5782 | 309.6903 | 287.8377 | 163.3097 | 234.1002 | 228.4189 | 222.8377 | 290.6343 | 145.2625 | 142.4218 | 110.7906 | 137.8938 | 183.3156 | 154.8377 | 155.9469 | 90.4749 | 312.3186 | 253.0471 | 256.528 | 247.1593 | 168.469 | 167.6844 | 152.3186 | 131.6874 | 87.7876 | 76.3687 | 64.9469 | 128.1623 | 265.2625 | 217.3186 | 247.2685 | 172.6343 | 171.528 | 164.1623 | 142.1563 | 126.7375 | 196.2595 | 164.3657 | 221.5811 | 184.003 | 166.3126 | 144.6784 | 208.4189 | 155.5251 | 216.528 | 178.3717 | 220.6814 | 262.4749 | 164.6283 | 167.472 | 251.7405 | 202.5251 | 232.9439 | 222.1032 | 200.3687 | 326.5221 | 196.6874 | 186.7315 | 176.3687 | 231.8998 | 145.472 | 167.7876 | 159.8407 | 140.3097 | 161.3097 | 196.8407 | 202.997 | 135.0501 | 125.4189 | 126.9499 | 128.1563 | 134 | 305.3747 | 290.0561 | 271.3156 | 330.5782 | 285.0501 | 265.6313 | 239.6754 | 343.6343 | 132.0531 | 75.8407 | 91.1593 | 132.0531 | 105.2595 | 129.7405 | 119.1062 | 138.7876 | 116.2625 | 125.1563 | 121.3156 | 120.2094 | 93.0531 | 81.5782 | 71.0501 | 98.7906 | 105.4749 | 95.7345 | 110.1533 | 87.7906 | 265.8407 | 234.1092 | 213.8878 | 294.2536 | 195.5782 | 210.472 | 210.6343 | 264.8377 | 332.3186 | 231.3717 | 242.3186 | 320.1062
1034 | 13.3684 | 192.4833 | 210.2699 | 224.6967 | 191.2031 | 183.8535 | 189.3599 | 185.82 | 164.7069 | 164.1337 | 118.2802 | 112.8869 | 164.6735 | 228.91 | 150.1902 | 148.0103 | 205.4036 | 214.9666 | 152.3933 | 157.8432 | 190.8432 | 183.3599 | 292.0437 | 296.4165 | 178.1697 | 233.7532 | 233.0566 | 211.2134 | 273.7069 | 139.82 | 132.1568 | 114.3933 | 147.1465 | 176.0334 | 139.7301 | 141.0668 | 95.7866 | 312.4267 | 234.3368 | 250.9434 | 265.3265 | 150.6401 | 154.7635 | 146.1568 | 142.8766 | 79.1234 | 81.8766 | 72.3599 | 132.1234 | 244.4936 | 200.0103 | 248.9434 | 174.5167 | 165.9434 | 182.0566 | 141.7866 | 116.1568 | 187.6067 | 177.8098 | 210.6735 | 175.3702 | 176.0232 | 144.1465 | 210.9666 | 150.2134 | 216.5167 | 201.7198 | 232.4165 | 243.4036 | 167.7532 | 163.3033 | 230.437 | 194.4267 | 218.3368 | 223.5733 | 190.5501 | 301.9768 | 200.2134 | 179.4833 | 169.3368 | 212.0437 | 145.4833 | 155.6401 | 177.2031 | 150.2931 | 171.9897 | 194 | 214.4165 | 133.4833 | 122.3933 | 129.2134 | 129.8766 | 121.4601 | 306.8535 | 277.4936 | 231.2905 | 313.6735 | 264.6735 | 272.6632 | 259.6196 | 321.8303 | 141.3599 | 75.7866 | 93.3033 | 141.3599 | 127.2597 | 126.5501 | 125.8766 | 148.2365 | 116.3933 | 110.8535 | 118.2134 | 107.5501 | 84.2468 | 79.3033 | 73.6632 | 89.8535 | 107.3033 | 92.3933 | 110.0566 | 96.9666 | 268.5733 | 234.3368 | 220.8998 | 289.6298 | 191.91 | 215.3599 | 220.8766 | 272.0232 | 305.9537 | 218.4936 | 231.9768 | 317.1234
1035 | 13.3794 | 189.7463 | 213 | 230.6716 | 224.8132 | 191.7611 | 164.642 | 171.2389 | 167.7611 | 144.2016 | 128.0148 | 107.1268 | 158 | 237.9253 | 139.8732 | 155.7237 | 198.1268 | 193.8957 | 150.1642 | 160.8732 | 180.4031 | 182.0374 | 309.4848 | 278.4778 | 184.3284 | 262.1043 | 235.5821 | 222.0521 | 278.8879 | 141.5447 | 141.8879 | 118.5447 | 156.9626 | 196.3953 | 153.3058 | 135.7089 | 94.8732 | 297.9479 | 262.9027 | 244.1642 | 264.9105 | 161.179 | 143.3284 | 161.5595 | 131.53 | 89.7611 | 76.2016 | 76 | 125.0374 | 243.7984 | 197.0895 | 243.5821 | 162.3658 | 149.2389 | 165.4405 | 141.5821 | 115.2537 | 170.6942 | 171.821 | 225.9774 | 168.2911 | 172.3658 | 150.4179 | 210.4553 | 149 | 215 | 200.4926 | 225.1121 | 251.7237 | 156.4031 | 170.0895 | 213.0374 | 206.179 | 249.0296 | 222.4553 | 197.3432 | 279.4031 | 183.4031 | 195.179 | 177.3432 | 228.0669 | 166.1416 | 152.7089 | 168.2763 | 170.6342 | 177.0374 | 177.6942 | 215.6195 | 125.7837 | 122.2911 | 126.5821 | 134.8358 | 124.5074 | 289.9479 | 262.6195 | 233.8506 | 333.9401 | 297.7012 | 289.47 | 290.8879 | 317.7984 | 124.4405 | 84.0895 | 83.7837 | 124.4405 | 135.0374 | 134.0521 | 114.53 | 133.4405 | 118.4179 | 130.1043 | 123.3805 | 110.0895 | 72.2016 | 95.3058 | 73.4553 | 89.2537 | 109.8358 | 110.7237 | 121.0895 | 93.6195 | 249.0226 | 249.4327 | 214.4031 | 298.8358 | 188.8732 | 210.4926 | 230.0895 | 291.0521 | 298.3805 | 230.5595 | 222.1642 | 317.1268
1036 | 13.3904 | 177.9026 | 230.3098 | 220.7389 | 207.5928 | 178.2477 | 171.7523 | 155.9026 | 165.7876 | 146.7699 | 132.8363 | 121.7036 | 165.9336 | 247.0487 | 165.6859 | 165.1637 | 191.7876 | 188.3274 | 139.6239 | 154.2301 | 188.4248 | 183.6062 | 301.084 | 304.062 | 189.9336 | 243.1504 | 210.4778 | 221.2301 | 257.5928 | 127.8539 | 139.3451 | 132.2611 | 143.8539 | 181.4778 | 165.0487 | 146.0974 | 90.3938 | 272.8053 | 251.4778 | 245.1637 | 259.115 | 163.3938 | 141.7212 | 159.9026 | 136.8186 | 82.0177 | 70.8363 | 73.115 | 130.2124 | 254.9336 | 218.031 | 224.2477 | 152.9513 | 144.7212 | 152.6726 | 164.0797 | 139.3585 | 182.7523 | 181.8673 | 213.3628 | 182.4248 | 161.5088 | 161.8186 | 234.2434 | 142.5088 | 210.6726 | 179.6902 | 248.5709 | 260.4425 | 165.1461 | 176.6062 | 220.3761 | 219.2124 | 242.8053 | 217.3938 | 202.7212 | 290.3098 | 164.7389 | 170.7079 | 170.4601 | 234.2301 | 160.2964 | 145.0664 | 165.3274 | 181.4912 | 196.6372 | 199.1284 | 202.9026 | 134.9823 | 124.1637 | 128.885 | 135.2788 | 117.9026 | 275.6239 | 278.7523 | 255.7036 | 321.1991 | 297.4115 | 294.2788 | 273.1991 | 290.7079 | 134.031 | 74.0177 | 85.0487 | 134.031 | 132.2788 | 128.1814 | 103.9513 | 119.9513 | 123.3274 | 137.8363 | 125.2788 | 123.0974 | 82.7036 | 78.1991 | 84.2611 | 93.1637 | 105.9513 | 115.8363 | 131.2124 | 85.9513 | 255.1461 | 246.1814 | 222.4248 | 287.7389 | 181.5088 | 201.2301 | 202.7079 | 272.9203 | 317.5885 | 219.5265 | 239.7523 | 325.9336
1037 | 13.4014 | 183.5359 | 227.4641 | 238.7395 | 186.7306 | 176.1347 | 183.4012 | 175.473 | 157.1317 | 146.0659 | 129.7994 | 120.3982 | 177.8024 | 224.0599 | 160.5958 | 167.4671 | 211.7395 | 194.4012 | 150.6706 | 154.9341 | 195.4671 | 166.6617 | 306.2036 | 317.2665 | 168.7934 | 242.2694 | 219.3383 | 213.8653 | 283.6766 | 124.2006 | 139.6677 | 130.3982 | 153.4042 | 178.6018 | 158.9311 | 141.1976 | 100.003 | 289.2066 | 223.6617 | 231.3294 | 242.5958 | 173.003 | 129.53 | 138.3952 | 127.0629 | 87.2694 | 69.2665 | 66.8653 | 141.0689 | 227.1917 | 216.9311 | 230.9371 | 157.6018 | 160.4042 | 147.3323 | 152.4611 | 131.8623 | 195.8683 | 189.4671 | 209.4012 | 181.3982 | 164.1347 | 147.6617 | 230.7964 | 153.2036 | 208.2665 | 189.8712 | 238.7276 | 248.53 | 169.5329 | 161.1288 | 239.4042 | 210.997 | 225.4641 | 235.8054 | 202.2665 | 287.4641 | 183.473 | 164.8683 | 183.6048 | 241.5359 | 168.9371 | 146.4012 | 167.7335 | 170.7964 | 199.8653 | 196.0629 | 196.7994 | 133.3982 | 136.7365 | 138.8024 | 135 | 138.2066 | 299.1407 | 281.5988 | 259.5329 | 322.2694 | 293.2006 | 276.3952 | 264.7335 | 314.2095 | 158.6048 | 74.8683 | 89.2006 | 158.6048 | 123.1976 | 119.5988 | 126.2066 | 120.2006 | 111.7964 | 109.8593 | 133.8024 | 124.0659 | 85.0659 | 72.6677 | 82.3982 | 95.4671 | 115.003 | 117.9341 | 111.994 | 81.7994 | 249.997 | 217.7934 | 225.7994 | 285.9341 | 184.1347 | 233.476 | 207.8712 | 264.7335 | 307.8623 | 202.4641 | 233.7964 | 319.4641
1038 | 13.4124 | 192.2208 | 221.7625 | 235.8127 | 194.9332 | 178 | 193.9499 | 199.6623 | 178.1205 | 153.9499 | 127.5084 | 126.2041 | 169.0668 | 218.7291 | 155 | 147.8628 | 213.2876 | 233.2911 | 159.7291 | 158.9833 | 181.8294 | 172.679 | 319.9499 | 298.3544 | 168.2041 | 257.9332 | 249.8663 | 240.3413 | 285.8127 | 127.2375 | 130.5585 | 136.9499 | 151.5418 | 191.4415 | 155.2542 | 154.4081 | 104 | 283.8294 | 249.0704 | 235.6957 | 260.1205 | 168.7959 | 142.8998 | 138.7124 | 118.2709 | 85.7792 | 70.4916 | 86.6289 | 124.6086 | 233.1539 | 223.6957 | 217.3544 | 168.2041 | 168.2375 | 149.7291 | 140.5251 | 136.4415 | 200.9833 | 190.7458 | 221.4415 | 194.6623 | 157.0501 | 141 | 249.8663 | 149.7959 | 192.3377 | 198.9833 | 217.5752 | 235.0501 | 151.1002 | 163.2041 | 227.8461 | 213.7124 | 227.2208 | 225.8461 | 217.6623 | 313.0871 | 169.6253 | 189.3747 | 178.0668 | 218.1504 | 165.7959 | 152.4749 | 181.4248 | 176.4415 | 197.2542 | 213.3747 | 189.2876 | 152.6289 | 135.7792 | 116.642 | 132.7625 | 148.4916 | 300.7959 | 277.0167 | 257.5084 | 330.4749 | 288.7792 | 270 | 249.3377 | 324.7625 | 150.8294 | 68.7959 | 77.321 | 150.8294 | 104.3377 | 109.7959 | 118.5919 | 136.6623 | 115.2041 | 105.9666 | 135.5084 | 118.5251 | 68.3377 | 68.7792 | 82.2375 | 101.2208 | 113.7792 | 124.9666 | 101.0167 | 86.2208 | 284.037 | 236.5954 | 237.679 | 288.4916 | 183.7625 | 203.4881 | 205.0501 | 296.3246 | 340.7828 | 231.0704 | 217.0668 | 345.0871
1039 | 13.4234 | 186.4189 | 224.7906 | 225.9351 | 193.6932 | 152.9822 | 168.9498 | 176.1917 | 173.8702 | 178.2272 | 123.9675 | 125.9675 | 184.711 | 211.6607 | 155.7581 | 157.6785 | 217.823 | 207.3362 | 167.823 | 175.9204 | 192.1623 | 185.3393 | 311.6283 | 292.7581 | 174.0325 | 258.2094 | 263.3068 | 221.9498 | 274.4189 | 127.2419 | 137.6136 | 135.4513 | 158.8555 | 184.3864 | 151.2094 | 147.8702 | 100.9675 | 281.2743 | 257.7257 | 219.289 | 268 | 171.3068 | 160.3717 | 141 | 128.3717 | 94.6136 | 67.9675 | 74.289 | 117.2419 | 246.5811 | 218.6607 | 232.2272 | 165.6932 | 170.5162 | 159.3393 | 130.6607 | 126.3539 | 171.6754 | 188.7257 | 225.7581 | 181.0472 | 167.6461 | 139.4838 | 245.8702 | 153.0649 | 177.1445 | 196.9675 | 222.0974 | 238.0649 | 164.711 | 156.1445 | 243.2272 | 191.7404 | 235.0649 | 224.5487 | 216.9351 | 294.9498 | 172.6136 | 185.6283 | 148.9822 | 230.9853 | 175.8879 | 160.0649 | 189.0325 | 161.8053 | 181.8377 | 211.9026 | 215.8083 | 150.1445 | 139.3068 | 133.7759 | 125.177 | 138.3864 | 288.9026 | 276 | 239.5634 | 348.6785 | 280.177 | 301.0827 | 256.8879 | 324.7581 | 154.3393 | 76.6136 | 74.5162 | 154.3393 | 107.3393 | 125.9528 | 124.3717 | 135.177 | 123.3068 | 114.0649 | 140.3068 | 100.3215 | 78.9204 | 82.1623 | 71.6283 | 99.2094 | 123.3717 | 114.8702 | 120.4691 | 87.2419 | 243.1739 | 233.8702 | 230.6283 | 260.1917 | 189.823 | 209.4691 | 216.4042 | 292.5958 | 324.4336 | 229.1121 | 228.1623 | 335.289
1040 | 13.4344 | 187.0816 | 221.3776 | 224.7704 | 192 | 180.4386 | 156.148 | 172.3928 | 173.0816 | 180.3776 | 132.2449 | 131.1632 | 177.9031 | 225.9489 | 163.7041 | 155.2959 | 203.8215 | 209.6377 | 147.6583 | 145.5614 | 192.3776 | 171.8215 | 349.6019 | 293 | 162.6735 | 239.2807 | 248.0511 | 221.4745 | 275.0816 | 121.6072 | 141 | 138.6224 | 145.8215 | 174.8368 | 153.852 | 151.7041 | 96.9184 | 270.4439 | 261.6224 | 238.4234 | 265.6888 | 156.8215 | 168.6224 | 166.4234 | 133.5408 | 88.7551 | 70.852 | 74.9336 | 106.2143 | 264.4081 | 221.3928 | 247.4745 | 167.852 | 160.2143 | 151.2143 | 146.4897 | 115.0664 | 171.2449 | 188 | 229.0816 | 177.3112 | 172.7704 | 160.5713 | 240.4592 | 155.7704 | 199.4234 | 189.8368 | 232.7041 | 243.0816 | 184.8673 | 159.1632 | 228.4287 | 247.9436 | 205.4134 | 215.9847 | 213.4592 | 292.1632 | 202.1938 | 176.6072 | 167.9642 | 246.4745 | 180.7704 | 168.1632 | 161.495 | 167.5561 | 190.8673 | 190.5103 | 231.9336 | 131.5919 | 134.8368 | 135.8368 | 121.4592 | 130.3776 | 288.3112 | 287.5561 | 240.1632 | 339.3623 | 293.4081 | 301.7551 | 244.8215 | 304.9695 | 140.8215 | 80.7704 | 100.4234 | 140.8215 | 120.0153 | 128.9184 | 113.3623 | 128.3776 | 115.7551 | 122.9336 | 130.4439 | 124.2754 | 92.4745 | 90.852 | 81.8673 | 94.9184 | 134.7041 | 117.9336 | 109.2807 | 89.3112 | 273.7651 | 250.0305 | 242.4081 | 277.1938 | 199.7041 | 225.2449 | 194.0358 | 291.0816 | 304.9847 | 239.6377 | 203.7246 | 322.8368
1041 | 13.4454 | 188.7827 | 187.1268 | 216.3903 | 195.9135 | 183.173 | 162.0443 | 177.9135 | 148.171 | 158.6499 | 128.7384 | 127.5211 | 163.0423 | 238.0443 | 155.0423 | 154.5654 | 214.6539 | 188.171 | 140.2173 | 172.5694 | 205.0885 | 181.0885 | 322.0825 | 264.8229 | 170.7404 | 244.9577 | 234.3903 | 223.2173 | 291.6539 | 126.2616 | 145.6962 | 132.9557 | 132.3903 | 173.7827 | 126.0402 | 162.6097 | 86.6076 | 270.1308 | 241.8672 | 244.4346 | 261.0865 | 149.6519 | 155.1288 | 142.6922 | 128.5211 | 94.6097 | 76.6962 | 91.0885 | 106.9135 | 233.7787 | 229.2616 | 268.002 | 161.9557 | 152.3038 | 149.5654 | 127.7364 | 134.1328 | 170.8692 | 170.7807 | 219.825 | 160.7807 | 173.7827 | 163.0865 | 271.3078 | 161.4789 | 192.1288 | 193.4789 | 233.4346 | 254.175 | 189.7827 | 173.5231 | 243.1328 | 231.7787 | 220.2636 | 227.0885 | 208.3038 | 306.5231 | 188.0845 | 177.3481 | 168.173 | 237.2596 | 174.7384 | 171.5654 | 156.9135 | 158.4769 | 191.8692 | 195.175 | 215.998 | 140.3058 | 147.8712 | 132.4346 | 151.5251 | 150.1328 | 251.4306 | 252.6479 | 274.0905 | 315.4326 | 293.3038 | 304.4789 | 256.4366 | 322.4809 | 140.6962 | 81 | 97.0423 | 140.6962 | 124.5654 | 134.2616 | 130.9155 | 119.9557 | 129.4366 | 138.3058 | 143.4366 | 128.3038 | 102.0443 | 73.998 | 86 | 96.3481 | 118.998 | 137.2193 | 127.4809 | 97.827 | 280.9557 | 235.6499 | 236.825 | 283.4346 | 197.3038 | 222.5211 | 214.9598 | 285.7384 | 309.827 | 210.3441 | 244.3099 | 335.8712
1042 | 13.4564 | 170.7153 | 204.2346 | 217.9749 | 185.8702 | 184.1799 | 167.1799 | 175.0251 | 172.7995 | 159.3599 | 126.205 | 114.8702 | 156.8201 | 233.6401 | 152.795 | 174.0797 | 238.8747 | 192.1298 | 170.2095 | 148.0206 | 193.1003 | 172.2802 | 322.9248 | 293.5694 | 174.795 | 245.615 | 219.2802 | 221.41 | 290.4351 | 150.2596 | 149.385 | 123.8451 | 140.3348 | 177.9749 | 148.2095 | 151.4852 | 69.6902 | 265.4351 | 244.7449 | 254.3348 | 282.2596 | 156.1549 | 137.4852 | 149.8997 | 134.9499 | 95.41 | 83.5649 | 82.2802 | 128.6697 | 266.1343 | 213.5103 | 253.9203 | 161.59 | 172.4647 | 144.4351 | 142.4647 | 133.6401 | 195.4396 | 194.6195 | 205.0752 | 171.1048 | 162.0752 | 146.1003 | 271.2551 | 162.205 | 199.1298 | 204.5398 | 209.9454 | 241.1003 | 174.1003 | 173.8201 | 241.0501 | 189.4055 | 209.5103 | 202.3805 | 200.6401 | 304.4351 | 153.3805 | 169.2551 | 198.5945 | 233.205 | 195.2596 | 187.1048 | 181.8496 | 165.3348 | 207.6947 | 221.0546 | 231.6697 | 169.4396 | 143.2551 | 148.6947 | 148.8702 | 152.0251 | 260.0797 | 269.8245 | 243.2506 | 331.4647 | 284.0501 | 291.6902 | 276.1048 | 328.205 | 161.0797 | 87.3599 | 93.205 | 161.0797 | 103.5353 | 133.615 | 97.2506 | 139.4647 | 130.0251 | 122.9203 | 129.7153 | 130.1799 | 85.7153 | 93.6446 | 72.4852 | 97 | 129.8997 | 133.2551 | 125.2551 | 83.3053 | 278.205 | 235.5649 | 253.0797 | 302.8747 | 211.1048 | 224.9749 | 212.6652 | 279.2301 | 327.8997 | 214.5148 | 210.3008 | 301.0456
1043 | 13.4674 | 191.8329 | 211 | 205.2763 | 210.4475 | 184.1927 | 169.6146 | 165.9272 | 179.3854 | 174.7237 | 121.9636 | 104.7345 | 173.7601 | 206.9744 | 157.0364 | 166.0835 | 239.1563 | 196.6146 | 165.8908 | 164.8329 | 221.2911 | 170.6146 | 275.9488 | 294.1199 | 183.0728 | 261.1456 | 220.8437 | 229.0728 | 267.2035 | 140.6617 | 116.0943 | 132.4946 | 133.3127 | 161.2399 | 151.1563 | 147.1927 | 83.7601 | 267.2291 | 259.1092 | 248.9272 | 282.3491 | 177.3747 | 138.8437 | 148.3491 | 137 | 59.4798 | 74.5054 | 69.3127 | 133.1927 | 285.8801 | 234.8329 | 270.7965 | 167.6509 | 161.0472 | 153.4946 | 127.0108 | 140.8801 | 185.8544 | 181.0108 | 211.6873 | 204.062 | 168.6873 | 151.6873 | 258.5054 | 174.9165 | 176.1671 | 206.1927 | 232.2547 | 246.6873 | 157.8908 | 169.7709 | 256.7601 | 224.593 | 199.3491 | 190.9636 | 182.8544 | 329.6402 | 199.2804 | 174.2655 | 178.7453 | 216.0472 | 196.1563 | 189.3854 | 186.3854 | 167.1927 | 206.3491 | 230.2291 | 232.9636 | 171.1563 | 136.9636 | 142.5054 | 158.1092 | 159.0728 | 291.6402 | 267.3127 | 244.3019 | 325.6981 | 290.0728 | 256.5162 | 253.3598 | 334.4582 | 153.0835 | 83.3491 | 95.4218 | 153.0835 | 128.6766 | 133 | 112.0256 | 127.2399 | 127.3854 | 129.3019 | 137.1092 | 134.2291 | 83.4218 | 83.8544 | 85.1456 | 99.4218 | 121.8908 | 128.5782 | 124.6146 | 87.8801 | 271.5418 | 227.3127 | 209.5633 | 310.4218 | 228.7237 | 263.1348 | 230.9892 | 306.2547 | 315.0472 | 261.593 | 220.6038 | 339.4367
1044 | 13.4784 | 200.4587 | 192.9695 | 211.0153 | 209.6239 | 166.789 | 193.7675 | 149.2478 | 172.4434 | 148.4955 | 121.8196 | 132.5045 | 167.3456 | 220.6697 | 151.4434 | 161.3609 | 215.0521 | 203.5566 | 159.7217 | 175.9174 | 215.8869 | 189.0305 | 290.2262 | 280.526 | 166.1499 | 263.3609 | 249.8654 | 240.8348 | 295.6023 | 127.1652 | 142.4219 | 121.8869 | 131.8196 | 169.2935 | 156.5566 | 147 | 81.4434 | 289.3088 | 221.0215 | 269.1284 | 256.4129 | 170.526 | 138.3609 | 156.0153 | 130.4434 | 75.5871 | 82.6544 | 79.2935 | 133 | 234.728 | 224.6086 | 232.5628 | 181.2935 | 202.0763 | 146.9847 | 154.7828 | 147.9174 | 181.1804 | 166.9847 | 211.5413 | 193.789 | 181.6544 | 142.526 | 292.061 | 163.2478 | 180.6544 | 191.2478 | 238.1804 | 266.211 | 154.1804 | 166.5413 | 235.5934 | 216.1499 | 231.6023 | 199.0153 | 190.474 | 317.9695 | 198.8869 | 185.0153 | 195.7675 | 240.6849 | 185.1652 | 192.2783 | 164.6912 | 171.9174 | 211.5566 | 208.0521 | 233.6391 | 151.9695 | 154.8501 | 129.3456 | 167.5566 | 136.4129 | 283.2478 | 280.5718 | 241.263 | 326.2783 | 273.9695 | 280.9632 | 260.9327 | 348.2935 | 150 | 86.0978 | 92.7217 | 150 | 131.9022 | 127.263 | 122.0978 | 131.1957 | 125.3609 | 125.4434 | 153.9327 | 122.7065 | 84.8196 | 66.8869 | 93.9174 | 95.0826 | 130.474 | 119.8043 | 116.8043 | 102.2935 | 246.2325 | 232.3761 | 214.3914 | 286.4129 | 218.8869 | 207.254 | 254.8501 | 285.1346 | 301.1652 | 259.7065 | 235.8348 | 313.2998
1045 | 13.4894 | 184.3628 | 193.9912 | 209.6726 | 193.0265 | 187.1239 | 184.0265 | 165.1328 | 156.8584 | 177.77 | 126.1593 | 139 | 180.8053 | 217.5133 | 154.9912 | 165.1593 | 224.1416 | 184.2035 | 173.1416 | 152.8761 | 181.3893 | 171.3716 | 320.1239 | 291.3098 | 191.1151 | 261.3363 | 258.4956 | 235.5133 | 273.053 | 127.4956 | 117.5574 | 130.646 | 130.3363 | 161.1858 | 128.053 | 137.8495 | 90.8142 | 289.0088 | 258.5842 | 249.0442 | 259.3186 | 155.5221 | 161.2921 | 154.6726 | 136.4867 | 82.6637 | 65.8672 | 87.823 | 146.3098 | 235.4779 | 207.6902 | 264.593 | 187.3274 | 172.0707 | 146.6637 | 139.5398 | 110.7344 | 171.8495 | 184.1328 | 239.2833 | 180.0177 | 159.8761 | 159.1328 | 303.3274 | 164.1593 | 187.9912 | 210.4602 | 235.5044 | 230.9026 | 175.6284 | 156.8495 | 251.6284 | 219.4867 | 206.5574 | 230.947 | 207.9735 | 309.0088 | 181.8584 | 174.5221 | 156.9114 | 235.354 | 198.8053 | 166.3805 | 163.3274 | 169.6726 | 208.0088 | 206.3274 | 206.5486 | 157.9823 | 143.1947 | 133.6549 | 151.5309 | 169.2656 | 273.3451 | 260.7079 | 248.3186 | 297.8849 | 280.8142 | 273.0265 | 260.6726 | 318.5574 | 150 | 96.1505 | 91.1681 | 150 | 125.177 | 135.1505 | 131.3186 | 137.9912 | 141.6372 | 146.4602 | 134.5398 | 113.3451 | 87.4956 | 70.6549 | 93.3363 | 82.354 | 122.1858 | 113.8407 | 139.1239 | 93.354 | 264.2921 | 218.1947 | 207.1858 | 280.1681 | 192.7079 | 226.2744 | 230.7167 | 263.1947 | 330.6107 | 230.3805 | 220.5309 | 331.6195
1046 | 13.5004 | 190.2857 | 214.4155 | 224.1947 | 180.7143 | 184.4026 | 166.6494 | 159.7143 | 170.0389 | 168.9611 | 128.6883 | 125.4936 | 163.7404 | 249.7661 | 150.9351 | 174.4415 | 221.9351 | 215.4544 | 143.0781 | 151.3766 | 198.6362 | 169.5325 | 304.7404 | 312.5713 | 178.4287 | 257.6234 | 257.3117 | 265.2336 | 263.6234 | 131.3766 | 131.2596 | 119.4936 | 151.1038 | 194.4544 | 147.3245 | 155.4155 | 70.2079 | 266.0521 | 273.9091 | 258.3506 | 265.2208 | 156.3766 | 129.7015 | 159.9091 | 119.4287 | 77.9351 | 72.1298 | 72.9611 | 129.5845 | 247.2857 | 216.8181 | 269.6234 | 166.0521 | 156.4026 | 146.1558 | 140.9091 | 136.7661 | 180.974 | 177.8702 | 228.9611 | 172.9351 | 155.8442 | 152.026 | 333.5453 | 165.8442 | 178.8702 | 183.7664 | 233.3117 | 227.2208 | 186.7532 | 158.3766 | 234.0521 | 208.3377 | 224.4804 | 243.7532 | 227.0389 | 297.026 | 215.2985 | 177.9091 | 170.7921 | 227.0909 | 188.4936 | 173.6623 | 182.5713 | 152.9611 | 178.2989 | 210.3766 | 251.6491 | 152.4026 | 137.4675 | 129.9351 | 147.1558 | 149.9872 | 277.0649 | 249.2468 | 244.0909 | 317.3245 | 318.4544 | 305.4544 | 239.7404 | 356.74 | 198.1166 | 98 | 93.5325 | 198.1166 | 120.6234 | 124.3377 | 127.0909 | 152.5064 | 130.6494 | 167.0389 | 137.5974 | 136.4804 | 79.5585 | 77.9091 | 102.2857 | 87.5974 | 135.1947 | 118.0649 | 133.8702 | 91.8442 | 239.4547 | 258.8959 | 232.0128 | 291.8181 | 182.0909 | 244.8181 | 231.7532 | 287.857 | 327.7143 | 232.5974 | 217.8442 | 334.4675
1047 | 13.5114 | 180.0098 | 221.4258 | 224.4307 | 182.4258 | 171.8662 | 179.416 | 162.2822 | 155.8712 | 160.8613 | 133.2822 | 130.708 | 159.1436 | 250.8613 | 183.4013 | 164.0098 | 208.1534 | 218.5742 | 166.119 | 136.584 | 210.708 | 195.6933 | 300.1436 | 277.4601 | 180.1387 | 259.5693 | 224.4552 | 214.4748 | 303.2528 | 133.7129 | 126.4356 | 121.2822 | 152.4307 | 170.1681 | 152.8564 | 156.4307 | 89.9804 | 273.9902 | 248.4503 | 235.3067 | 282.2724 | 170.7031 | 140.9853 | 155.0049 | 131.416 | 90.7031 | 65.4356 | 79.4209 | 123.4307 | 259.2773 | 231.8466 | 301.5448 | 170.5644 | 137.0147 | 142.5742 | 151.4209 | 120.7325 | 208.6933 | 190.5595 | 214.8662 | 149.7325 | 163.708 | 135.4405 | 353.5595 | 182.2724 | 217.2528 | 189.1338 | 247.5595 | 229.7129 | 170.0147 | 185.5497 | 265.9706 | 211.1387 | 228.1436 | 239.0049 | 214.584 | 299.2822 | 172.3263 | 179 | 171.5742 | 230.2822 | 169.7276 | 186.2773 | 178.292 | 165.416 | 211.5399 | 203.292 | 210.4699 | 148.4307 | 130.1485 | 135.8515 | 143.5742 | 179.2577 | 272.8613 | 260.8466 | 265.2675 | 345.1239 | 293.3116 | 303.4356 | 244.5644 | 321.3214 | 185.5889 | 96.2871 | 94.8564 | 185.5889 | 113.1485 | 111.7227 | 124.2871 | 150.7178 | 142.5595 | 148.5889 | 128.7227 | 133.292 | 93.416 | 88.4209 | 80.8761 | 94.9951 | 120.8712 | 136.9853 | 155.1239 | 86.0049 | 229.7178 | 224.1779 | 226.7227 | 288.0049 | 196.416 | 207.6036 | 234.7129 | 287.0049 | 323.4307 | 235.7129 | 224.8515 | 331.4307
1048 | 13.5224 | 197.1121 | 233.2935 | 201.4129 | 184.7375 | 163.9189 | 181.1313 | 175.1623 | 179.9307 | 169.5561 | 120.969 | 129.3938 | 166.8186 | 220.4631 | 152.5132 | 161.1313 | 204.2625 | 178.0383 | 151.0191 | 167.0119 | 203.3127 | 187.8377 | 297.3938 | 315.3054 | 159.2817 | 262.6062 | 271.124 | 219.7685 | 276.1194 | 132.2625 | 125 | 127.2124 | 130.2817 | 170.2124 | 165.1623 | 156.8687 | 77.4941 | 282.0811 | 261.3747 | 246.6372 | 266.7566 | 163.4439 | 157.8997 | 142.7065 | 116.6253 | 74.7566 | 75.2935 | 87.9499 | 128.2124 | 241.0191 | 227.9189 | 254.0073 | 169.3938 | 155.7183 | 134.1814 | 146.0501 | 123.0811 | 192.1504 | 190.3938 | 201.7065 | 187.6992 | 171.0811 | 154.7183 | 316.907 | 179.7876 | 198.8067 | 164.938 | 230.0191 | 239.5561 | 172.2124 | 161.3318 | 227.6946 | 221.5561 | 233.2124 | 231.9189 | 218.0811 | 319.1121 | 189.1933 | 193.7685 | 201.4057 | 206.6755 | 175.6873 | 170.6253 | 178.7375 | 169.7375 | 201.4941 | 215.8997 | 226.3245 | 167.1121 | 155.9307 | 158.7183 | 140.3938 | 146.7757 | 274.6062 | 303.8305 | 248.1504 | 348.1313 | 308.8496 | 264.6445 | 264.2434 | 332.2434 | 177.6563 | 90.7876 | 92.3938 | 177.6563 | 118.9499 | 128.2434 | 131.8186 | 125.6755 | 159.7685 | 146.7375 | 168.6992 | 133.7375 | 88.1814 | 88.2625 | 84.8186 | 102.9499 | 119.7375 | 140 | 138.1504 | 91.9499 | 268.093 | 253.4868 | 253.6682 | 271.3628 | 185.969 | 220.1121 | 233.2625 | 273.8377 | 311.7065 | 223.8377 | 236.4248 | 310.1504
1049 | 13.5334 | 179.7364 | 236.762 | 193.5949 | 187.6431 | 173.5723 | 166.0226 | 184.9292 | 166.3795 | 153.3795 | 130.4533 | 127.238 | 160.0708 | 238.9066 | 180.4789 | 161 | 211.9292 | 170.238 | 142.7139 | 159.6657 | 180.8554 | 185.119 | 318.1446 | 277.9488 | 156.881 | 285.9066 | 215.5662 | 229.0482 | 288.6205 | 140.8102 | 132.0482 | 142.0964 | 134.9292 | 158.6657 | 159.9518 | 130.5693 | 103.1928 | 277.7139 | 237.5693 | 228.7364 | 264 | 161.119 | 151.1898 | 161.2636 | 121.0482 | 76.4051 | 67.3087 | 75.7846 | 112.2605 | 254.7395 | 229.6431 | 269.7877 | 163.7139 | 143.1416 | 163.8359 | 145.881 | 117.8328 | 186.3569 | 174.1416 | 210.5723 | 179.9036 | 181.6913 | 152.7139 | 299.5467 | 201.9066 | 184.4277 | 181.2636 | 223.4759 | 233.9518 | 200.3118 | 198.4082 | 234.2154 | 227.4051 | 223.4277 | 202.8072 | 206.6657 | 286.759 | 181.5467 | 177.4985 | 196.3087 | 227.6687 | 203.4307 | 201.4789 | 174.5949 | 164.7139 | 191.9518 | 200.3795 | 221.1898 | 158.5467 | 152.0708 | 139.9744 | 165.5497 | 168.3118 | 281.1672 | 280.0452 | 228.2605 | 334.7846 | 290.8554 | 263.4051 | 282.8584 | 327.9518 | 173.4759 | 102.3343 | 105.2154 | 173.4759 | 121.762 | 131.881 | 141.8102 | 130.8102 | 125.878 | 154.0482 | 168.8328 | 141.0482 | 90.5241 | 87.119 | 89.5241 | 96.9518 | 136.7395 | 139.119 | 156.1446 | 87.7139 | 278.4051 | 236.9744 | 226.2831 | 299.8359 | 182.238 | 228.2861 | 238.2861 | 261.4277 | 294.1416 | 252.8359 | 236.238 | 322.8584
1050 | 13.5444 | 169.8535 | 197.6941 | 219.7995 | 176.3869 | 158.9203 | 183.6529 | 177.0668 | 148.8136 | 154.5733 | 131.1067 | 131.4666 | 169.7198 | 216.9872 | 152.8406 | 172.6131 | 222.8265 | 173.5733 | 172.3727 | 162.4666 | 191.3997 | 183.2134 | 288.8406 | 280.9332 | 176.6529 | 245.2275 | 251.6658 | 260.3727 | 307.0797 | 150.0398 | 140.1465 | 129.707 | 153.8663 | 155.2134 | 133.0938 | 137.7198 | 90.9203 | 258.2404 | 242.0398 | 281.3856 | 264 | 153.8535 | 141.9602 | 151.4936 | 125.5733 | 67.1735 | 66 | 78.4666 | 118.0398 | 232.8805 | 221.0668 | 317.6658 | 166.5733 | 179.4126 | 163.5334 | 155.8265 | 133.973 | 206.5463 | 183.6131 | 209.3201 | 174.4267 | 183 | 162.7198 | 326.5861 | 173.734 | 196.3997 | 201.8663 | 207.8136 | 247.293 | 170.054 | 168.2674 | 269.946 | 214.6003 | 244.3329 | 198.1067 | 203.2134 | 310.5861 | 193.3997 | 170.5334 | 198.5733 | 184.5476 | 184.6671 | 195.2802 | 187.3997 | 167.5733 | 203.5064 | 206.0398 | 209.2802 | 141.8136 | 148.3201 | 170.0527 | 175.2532 | 192.5463 | 273.0668 | 275.1067 | 225.1067 | 339.2532 | 289.7866 | 278.293 | 251.9473 | 337.7198 | 170.3201 | 94.1735 | 104.3201 | 170.3201 | 115.7468 | 130.2134 | 123.3471 | 117.707 | 142.4396 | 138.027 | 173.3599 | 136.6401 | 82.0668 | 98.6131 | 98.9332 | 96.8933 | 148.8265 | 156.8663 | 165.2532 | 79.8535 | 250.4139 | 241.1465 | 245.2262 | 288.8136 | 198.973 | 231.6799 | 287.239 | 269.8265 | 333.0925 | 230.2005 | 215.4537 | 341.973
1051 | 13.5554 | 188.9233 | 190.2832 | 232.0561 | 182.2449 | 182.357 | 176.9439 | 188.6785 | 139.7551 | 164.0561 | 111.0767 | 125.6608 | 180.9617 | 252.0355 | 159.8673 | 163.1327 | 215.8495 | 175.8112 | 159.6991 | 148.5103 | 194.8112 | 193.8673 | 325.7523 | 304.6402 | 177.1888 | 211.9262 | 228.9262 | 217.8141 | 280.0206 | 140.1327 | 146.4336 | 128.9056 | 139.6991 | 169.4897 | 138.1505 | 153.4897 | 101.6785 | 290.4131 | 263.829 | 252.6813 | 262.1888 | 153.9056 | 156.3953 | 155.4336 | 138.6785 | 91.357 | 81.3953 | 77.1888 | 134.3953 | 219.1327 | 225.4336 | 289.4926 | 158.8495 | 163.171 | 163.9056 | 175.1121 | 123.3215 | 205.3776 | 160.5486 | 258.8084 | 190.3009 | 179.3776 | 170.3392 | 316.4159 | 190.829 | 187.1327 | 190.4159 | 230.4514 | 258.0561 | 173.2449 | 177.5841 | 235.9645 | 226.5841 | 220.7374 | 208.8673 | 203 | 308.5664 | 178.6991 | 185.3953 | 175.4542 | 224.2803 | 188.3392 | 205.7729 | 201.6785 | 178.8673 | 200.472 | 193.4159 | 208.9056 | 158.1121 | 155.2449 | 161.3215 | 157.8879 | 160.5869 | 265.6608 | 281.3392 | 242.2065 | 348.1505 | 292.7168 | 288.1505 | 284.2243 | 329.0383 | 179.0561 | 94.8112 | 107.6224 | 179.0561 | 127.6785 | 127.2832 | 159.0355 | 122.3392 | 173.0738 | 177.6579 | 152.2654 | 133.2832 | 100.0177 | 96.3776 | 97.2832 | 91.5664 | 159.9617 | 131.8318 | 154.2271 | 82.6224 | 284.1299 | 245.6224 | 227.171 | 269.7935 | 200.0944 | 216.6047 | 263.115 | 272.8112 | 296.3421 | 242.3953 | 231.1121 | 333.1327
1052 | 13.5664 | 191.9179 | 217.5369 | 238.5074 | 189.4253 | 161.1347 | 166.821 | 171.642 | 179.3875 | 160.4105 | 110.8358 | 118.5747 | 184.7537 | 231.2168 | 172.0148 | 192.2906 | 214.0821 | 167.7389 | 146.0673 | 159.8506 | 207.8506 | 181.2315 | 324.4926 | 282.2168 | 173.3284 | 228.2759 | 230.5895 | 233.1937 | 286.179 | 143.5895 | 117.6273 | 114.3136 | 151.7685 | 161.821 | 175.7159 | 155 | 95.6568 | 281.1494 | 275.179 | 262.7685 | 276.6864 | 137.4778 | 134.1347 | 140.3957 | 130.821 | 82.0673 | 71.9852 | 80.6716 | 115.8063 | 247.3727 | 217.7389 | 245.6125 | 168.0969 | 141.7241 | 160.3284 | 164.1494 | 123.8358 | 176.5452 | 165.3432 | 243.8063 | 183.7389 | 161.5599 | 162.7389 | 322.3432 | 200.3432 | 188.7537 | 201.8506 | 209.1347 | 238.8063 | 186.8506 | 144.1199 | 242.0969 | 200.4631 | 205.1494 | 208.1642 | 209.4253 | 315.3432 | 188.0148 | 193.4253 | 184.9327 | 216.1494 | 182.5747 | 181.2989 | 207.5895 | 172.6568 | 209.179 | 216.7832 | 198.9031 | 150.821 | 144.9852 | 139.8063 | 178.9474 | 130.381 | 299.8801 | 245.2841 | 244 | 343.4926 | 296.6716 | 280.7389 | 286.1642 | 333.5074 | 187.3432 | 91.3284 | 113.5074 | 187.3432 | 127.1642 | 133.4253 | 141.8884 | 121.1642 | 153.0526 | 140.6946 | 181.2085 | 127.4926 | 104.7537 | 85.9031 | 97 | 105.6864 | 151.821 | 150.1116 | 173.1937 | 76.5747 | 258.6273 | 203.7767 | 220.4105 | 303.798 | 192.6568 | 237.0295 | 268.2611 | 290.4401 | 288.3284 | 239.4105 | 266.0443 | 325.5747
1053 | 13.5774 | 166.8849 | 210.6981 | 207.3736 | 194.6509 | 162.7208 | 170.6509 | 171.8604 | 151.3736 | 157.2094 | 145.417 | 129.1623 | 179.4189 | 205.7453 | 144.1641 | 150.3509 | 210.2792 | 183.7434 | 143.1396 | 154.4887 | 204.3491 | 169.7679 | 294.234 | 327.4396 | 192.534 | 255.1151 | 225.4189 | 229.4189 | 308.3943 | 146.7906 | 135.4642 | 136.2547 | 153.9302 | 162.8604 | 141.7925 | 135.466 | 108.0226 | 266.9774 | 238.7925 | 227.7226 | 226.8396 | 149.9528 | 132 | 151.0925 | 128.1396 | 76.3491 | 75.6509 | 89.3717 | 118.6509 | 203.4906 | 213.2792 | 228.9774 | 160.6283 | 144.6509 | 135.8151 | 145.3264 | 121.2094 | 189.8132 | 169.7208 | 230.8377 | 177.4189 | 164.6509 | 186.1849 | 345.3245 | 173.0943 | 175.0472 | 210.4415 | 196.7679 | 256.534 | 163.8151 | 172.6264 | 255.0925 | 220.3245 | 204.9302 | 221.9528 | 182.0943 | 307.6283 | 181.5585 | 179.117 | 176.6981 | 208.4887 | 184.7906 | 206.9057 | 204.2792 | 165.4887 | 203.4887 | 179.0019 | 197.0698 | 148.1396 | 148.6509 | 162.1849 | 159.6057 | 149.3943 | 276.0245 | 273.6264 | 257.0226 | 334.6283 | 277.466 | 286.5113 | 288.7906 | 353.534 | 207.534 | 79.8377 | 114 | 207.534 | 135.3717 | 110.7453 | 132.5585 | 130.3019 | 159.3717 | 171.417 | 171.9075 | 133.5113 | 85.466 | 86.8604 | 93.2792 | 92.117 | 161.2321 | 138.9774 | 172.2094 | 83.4415 | 235.5358 | 231.6264 | 230.2321 | 281.8849 | 193.8604 | 222.2566 | 242.0245 | 285.4887 | 316.8359 | 212.0245 | 222.4906 | 317.5585
1054 | 13.5884 | 177.2522 | 221.3098 | 257.7789 | 182.7478 | 179.9646 | 169.115 | 145.6106 | 157.4823 | 140.0354 | 115.0132 | 121.5177 | 177.115 | 228.5045 | 135.4026 | 180.9293 | 213.7699 | 167.0929 | 164.677 | 146.4602 | 206.8274 | 196.3319 | 320.2744 | 277.2787 | 177.9778 | 267.3673 | 230.6549 | 261.9868 | 295.8628 | 123.438 | 135.115 | 116.323 | 132.323 | 143.2079 | 171.0443 | 143.4248 | 100.5177 | 265.0575 | 261.447 | 246.677 | 288.9736 | 138.7478 | 155.562 | 152.9425 | 144.0222 | 77.885 | 60.9203 | 84.3451 | 117.115 | 210.3673 | 222.4248 | 296.801 | 172.2522 | 132.7478 | 148.1372 | 154.3673 | 140.7921 | 218.3319 | 173.7699 | 231.885 | 205.2744 | 172.5398 | 194.5974 | 307.4158 | 188.9071 | 188.1372 | 199.6902 | 235.5842 | 246.6902 | 198.7567 | 157.0929 | 239.9778 | 226.7124 | 210.6549 | 227.7124 | 196.9646 | 306.0575 | 190.4248 | 172.3451 | 172.2301 | 216.4823 | 168.0354 | 181.6681 | 189.8628 | 158.4026 | 229.3894 | 194.8496 | 219.6195 | 158.3673 | 136.7478 | 164 | 176.8496 | 148.1726 | 285.3098 | 238.3008 | 202.3937 | 336.8274 | 278.8274 | 297.3673 | 285.2301 | 310.7035 | 175.0707 | 100.677 | 102.6902 | 175.0707 | 117.1504 | 126.9071 | 132 | 113.0929 | 136.438 | 156.0929 | 174.7699 | 122.6902 | 91.5398 | 92.6549 | 74.1504 | 89.115 | 136.553 | 131.4026 | 176.7124 | 84.9425 | 247.1947 | 241.5398 | 254.562 | 277.1726 | 177.9778 | 231.3673 | 267.3319 | 284.0575 | 287.8982 | 221.3098 | 234.0797 | 326.4248
1055 | 13.5994 | 186.5929 | 202.9046 | 209.4424 | 185.8191 | 174.3166 | 169.9548 | 158.3216 | 149.4071 | 186.7385 | 154.0551 | 150.5979 | 167.4523 | 222.3618 | 169.3718 | 160.0855 | 203.4975 | 178.412 | 155.4975 | 161.2763 | 196.4975 | 178.9046 | 333.4573 | 333.1958 | 187.5025 | 236.4926 | 231.9548 | 206.7137 | 265.4021 | 140.1406 | 102.5378 | 121.6834 | 117.6332 | 153.4573 | 151.995 | 148.7739 | 89.4975 | 242.0855 | 231.4926 | 260.412 | 237.6233 | 146.5929 | 152.2261 | 133.9046 | 123.0403 | 78.9548 | 81.9597 | 66.8141 | 119.8643 | 229.1406 | 238.2763 | 219.8445 | 186.3668 | 176.8742 | 152.8191 | 147.3618 | 114.3116 | 196.1307 | 178.7739 | 228.1809 | 176.4473 | 145.3116 | 140.5781 | 229.5731 | 172.8141 | 176.588 | 204.7286 | 227.4975 | 232.6332 | 153.2615 | 175.0954 | 201.7639 | 182.1258 | 211.9548 | 229.9095 | 199.9095 | 302.1809 | 175.7237 | 179.6382 | 183.4573 | 234.1859 | 152.6784 | 175.2261 | 194.7286 | 168.5025 | 195.6735 | 212.2311 | 228.6382 | 137.995 | 145.5477 | 147.7689 | 156.995 | 143.2261 | 269.7689 | 217.8594 | 274.4269 | 314.0855 | 238.8996 | 289.4071 | 275.4523 | 342.3718 | 160.588 | 85.7689 | 97.2261 | 160.588 | 114.0905 | 152.824 | 134.8643 | 146.3718 | 117.8141 | 162.6382 | 155.9046 | 129.6382 | 83.4071 | 79.6332 | 96.8693 | 88.0452 | 128.3166 | 139.5929 | 148.3569 | 95.5025 | 264.2311 | 220.0403 | 203.4876 | 274.1357 | 176.0452 | 231.0452 | 224.1258 | 289.7286 | 286.9548 | 218.1809 | 270.3265 | 306.9498
1056 | 13.6104 | 187 | 226.1765 | 213.7694 | 169.5599 | 163.3623 | 174.8353 | 182.2095 | 166.4071 | 156.1199 | 126.9881 | 147.1647 | 161.1976 | 237.473 | 181.6377 | 162.8682 | 218.473 | 212.8472 | 143.3953 | 168.7694 | 187.2964 | 164.4611 | 299.1858 | 302.1528 | 163.8235 | 220.4941 | 225.2306 | 211.7365 | 259.1647 | 131.3294 | 122.2754 | 120.0659 | 153.7484 | 150.1318 | 157.7694 | 163.5059 | 90.9341 | 264.2095 | 259.0119 | 245.527 | 268.8472 | 147 | 152 | 146.5389 | 140.3742 | 81.9012 | 77.1976 | 60.1976 | 118.0659 | 216.4611 | 196.4493 | 233.4071 | 163.7905 | 160.6258 | 153.9671 | 147 | 117.8353 | 185.3294 | 205.1107 | 237.6706 | 182.7365 | 172.0448 | 146.7036 | 247.2754 | 179.7365 | 199.2095 | 204.0329 | 219.2635 | 225.2306 | 176.1436 | 170.1976 | 237.7154 | 197.4071 | 223.6047 | 201.9552 | 217.4071 | 311.6706 | 193.3742 | 156.7905 | 185.9341 | 205.9881 | 149.0988 | 158.5599 | 174.6917 | 160.2964 | 189.1647 | 208.1647 | 199.9881 | 151.5059 | 137.2964 | 130.5599 | 130.8564 | 120.7576 | 269 | 240.2095 | 236.4163 | 310.0988 | 289.2213 | 277.3953 | 256.6258 | 342.0659 | 173.5389 | 99.5059 | 117.3083 | 173.5389 | 128.5059 | 130.7905 | 107.9223 | 133.4941 | 131.5059 | 148.4941 | 168.5389 | 142.5718 | 72.3623 | 74.1647 | 82.527 | 100.5718 | 123.1647 | 130.3294 | 147 | 83.4282 | 261.1318 | 233.5059 | 218.4071 | 304.946 | 191.473 | 212.6258 | 207.4941 | 288.0659 | 312.1436 | 227.6706 | 241.054 | 306
1057 | 13.6214 | 180.1445 | 230.9174 | 185.5988 | 197.4012 | 155.1652 | 175 | 183.9794 | 161.1239 | 155.9794 | 135.7935 | 136.2271 | 177.649 | 217.4336 | 153.5988 | 146.351 | 231.7316 | 182.6607 | 138.1032 | 164.1032 | 164.4749 | 194.3599 | 308.7729 | 320.5871 | 198.2567 | 234.6903 | 220.1032 | 223.7522 | 262.9174 | 121.2065 | 135.7316 | 121.9587 | 158.9174 | 151.9587 | 145.2684 | 126.7846 | 86.1032 | 261.0826 | 267.8348 | 231.2891 | 260.2065 | 125.4542 | 149.0619 | 140.1445 | 149.8142 | 70.2478 | 74.0619 | 73.7109 | 138.5664 | 237.5458 | 206.7522 | 222.2478 | 190.4219 | 165.8761 | 145.1858 | 160.7109 | 111.1445 | 194.7935 | 176.6194 | 222.3304 | 184.9587 | 139.702 | 163.649 | 228.4129 | 154.5368 | 181.3923 | 179.5162 | 207.2478 | 215.2065 | 180.9174 | 171.9587 | 205.702 | 202.8968 | 192.6607 | 196.1032 | 194.4955 | 307.1032 | 179.3097 | 155.0206 | 152.702 | 196.1858 | 144.1032 | 158.9794 | 183.7935 | 175.6696 | 182.1445 | 195.2684 | 216.6284 | 123.5988 | 124.2684 | 124.1239 | 153.5045 | 139.5871 | 290.5458 | 272.3393 | 247.7316 | 314.8968 | 304.7109 | 317.1535 | 272.649 | 327.3097 | 154.4129 | 92.1652 | 96.4542 | 154.4129 | 129.9794 | 140.7729 | 113.8555 | 125.1652 | 138.8555 | 144.0826 | 163.1239 | 111.6607 | 83.7522 | 81.8348 | 90.8142 | 91.2065 | 115.1652 | 97.6814 | 130.351 | 81.0413 | 260.0206 | 225.1858 | 223.8968 | 314.8142 | 183.1858 | 203.1858 | 209.9381 | 241.9704 | 308.1032 | 223.1032 | 238.0413 | 296.2065
1058 | 13.6324 | 166.117 | 224.0585 | 217.7243 | 194.0334 | 171.8579 | 162.1086 | 161.1922 | 163.9749 | 194.6741 | 147.8997 | 134.0167 | 162.1337 | 212.0418 | 155.9749 | 166.8245 | 198.2841 | 195.883 | 158.8245 | 160.0334 | 202.6741 | 154.3426 | 274.2925 | 291.2507 | 139.5014 | 209.2173 | 229.9164 | 200.2005 | 274.8997 | 131.9081 | 102.2841 | 130.9248 | 138.1755 | 159.9332 | 130.1253 | 158.7243 | 96.9081 | 269.9248 | 283.8663 | 244.883 | 262.9749 | 153.7577 | 168.8329 | 119.1755 | 130.1671 | 89.8329 | 76.9749 | 69.0418 | 111.234 | 202.3008 | 207 | 232.9081 | 159.2674 | 162.0334 | 145.9916 | 134.2256 | 112.9833 | 187.0668 | 176 | 221.0084 | 151.2841 | 161.8078 | 131.2757 | 223.0418 | 148.0501 | 199.8412 | 199.8245 | 222.8663 | 259.624 | 172.0752 | 163.0752 | 218.883 | 224.8162 | 206.8747 | 195.0084 | 198.9582 | 316.9164 | 168.0919 | 173.8412 | 172.8245 | 228.7243 | 151.9332 | 170.8997 | 172.1003 | 167.0752 | 199.8496 | 184.0919 | 220.9666 | 167.624 | 112.1003 | 134.9081 | 112.351 | 120.1671 | 248.3593 | 280.9332 | 206.351 | 326.8997 | 294.0919 | 263.4596 | 266.0585 | 366.6658 | 151.0251 | 89.0251 | 101.9499 | 151.0251 | 132.9749 | 104.3092 | 116.9749 | 119.0501 | 126.1086 | 142.0167 | 145.1504 | 127.8579 | 97.883 | 90.9248 | 96.9499 | 95.9582 | 106.0752 | 128.7326 | 135.9499 | 97.8579 | 239.1755 | 231.9415 | 208.1337 | 300.1253 | 153.2507 | 203 | 212.9749 | 304.4652 | 310.9749 | 236.883 | 227.0919 | 349.5488
1059 | 13.6434 | 168.9646 | 193 | 209.0433 | 193.9764 | 175.0197 | 175.0354 | 147.0787 | 174.9016 | 174.9174 | 133.9331 | 133.8701 | 180.8151 | 211.063 | 172.9292 | 179.9646 | 208.0472 | 164.0787 | 142.0669 | 145.0472 | 178.0669 | 171.9488 | 309.0394 | 302.8308 | 164.0551 | 232.9843 | 240.9134 | 229.9134 | 285.0275 | 115.0472 | 132.0039 | 122 | 134.0275 | 162.1181 | 157.9331 | 145.059 | 105.0079 | 262.0944 | 276.8701 | 247.9725 | 282.941 | 144.9764 | 143.0275 | 123.1613 | 146.9646 | 81.0472 | 64.0079 | 71.059 | 113.9803 | 236.9331 | 213.0944 | 238.8623 | 178.0118 | 142.1259 | 141.0197 | 147.9685 | 93.1062 | 192.0118 | 199.0354 | 207.0944 | 167.1692 | 180.8347 | 155.0708 | 208.9685 | 128.0551 | 187.9056 | 179.0512 | 193.0433 | 256.8819 | 170.0472 | 151.063 | 216.0826 | 201.1141 | 227.8308 | 207.9449 | 203.9843 | 295.9843 | 195.8938 | 186.9921 | 159.1692 | 223.063 | 166.8426 | 151.9882 | 174.9685 | 157.9882 | 178.0512 | 175.9488 | 184.9882 | 133.1181 | 132.9252 | 116.0472 | 125.9567 | 119.0708 | 269.9174 | 264.9331 | 260.8938 | 306.9567 | 259.0394 | 267.0551 | 248.1259 | 323.059 | 159.9882 | 81.0275 | 76.0708 | 159.9882 | 103.0984 | 124.9567 | 108.063 | 125.9292 | 125.9961 | 132.9961 | 122.9606 | 134.9016 | 82.941 | 84.9331 | 72.0433 | 98.9961 | 124.9567 | 96.0157 | 113.0394 | 82.9646 | 256.9725 | 204.2046 | 221.8898 | 288.9449 | 166.9528 | 203.9488 | 204.0236 | 296.9528 | 301.059 | 216.9843 | 231.1023 | 298.9056
1060 | 13.6544 | 160.3407 | 193.1947 | 219.7728 | 187.7566 | 180.146 | 183.3185 | 167.1298 | 150.1298 | 154.146 | 117.0974 | 101.1947 | 134.2596 | 226.9513 | 154.7566 | 171.146 | 219.9675 | 184.0325 | 158.8053 | 156.7891 | 194.4483 | 158.7728 | 318.4159 | 260.2596 | 178.146 | 228.7891 | 218.9189 | 208.5841 | 291.6106 | 126.7891 | 132.854 | 121.8702 | 140.9189 | 191.1238 | 141.2109 | 159.5781 | 106.9351 | 285.7242 | 244.2596 | 240.5619 | 267.3023 | 139.2434 | 149.7242 | 163.643 | 137.7728 | 92.8864 | 66.2596 | 85.8702 | 109.5192 | 220.0325 | 236.7242 | 204.9249 | 181.0487 | 173.497 | 146.1947 | 140.1298 | 119.7891 | 194.5781 | 207.6755 | 230.351 | 209.3834 | 139.1136 | 172.9838 | 200.8377 | 142.0325 | 164.6653 | 191.8053 | 203.7079 | 227.0162 | 181.6268 | 167.0325 | 236.9675 | 229.5943 | 185.2109 | 194.1623 | 200.2272 | 291.7566 | 168.854 | 184.854 | 201.0589 | 238.8377 | 127.2272 | 149.0162 | 166.7242 | 154.7404 | 190.8702 | 163.1298 | 182.3083 | 162.5132 | 114.2758 | 127.9675 | 115.0162 | 136.9026 | 248.854 | 248.3245 | 234.5192 | 296.3245 | 269.1785 | 281.0974 | 279.854 | 337.497 | 156.5457 | 87.6917 | 93.854 | 156.5457 | 127.7404 | 114.3083 | 123.8215 | 108.4219 | 124.6268 | 131.7566 | 113.1947 | 110.1298 | 68.2758 | 68.0487 | 83.1623 | 97.8215 | 114.1623 | 100.1947 | 122.8215 | 73.9838 | 249.7079 | 254.8317 | 194.4868 | 275.7139 | 155.3083 | 191.3083 | 209.7728 | 284.7404 | 315.8215 | 213.4543 | 256.5781 | 274.8864
1061 | 13.6654 | 181.2852 | 205.2567 | 206.4848 | 172.6863 | 187.4316 | 142.7985 | 174.4582 | 157.6578 | 161.7738 | 123.0285 | 113.2281 | 150.3707 | 223.3441 | 140.8555 | 179.8289 | 217.3441 | 185.7433 | 147.0285 | 144.1426 | 161.0856 | 145.2852 | 283.9696 | 277.2833 | 186.2586 | 215.9715 | 213.6863 | 242.9163 | 267.6578 | 114.5133 | 124 | 114.3422 | 136.4563 | 139.0837 | 154.0285 | 134.884 | 102.8574 | 269.3422 | 259.8859 | 214.3992 | 227.0247 | 153.8289 | 133.6274 | 141.0304 | 123.7433 | 85.7148 | 82.4848 | 78.1996 | 140.4867 | 223.1692 | 219.7148 | 259.7738 | 183.7148 | 144.0837 | 157.5437 | 147.7148 | 107.7985 | 168.8289 | 187.3156 | 192.3403 | 171.5152 | 146.3422 | 171.4867 | 190.6293 | 143.4582 | 204.8859 | 179.8289 | 187.1407 | 227.1445 | 158.2871 | 169 | 234.6863 | 205.1141 | 197.8574 | 203.4011 | 213.1445 | 277.3137 | 159.5722 | 175.4011 | 144.2281 | 229.4278 | 140.8289 | 150.0856 | 150.4563 | 139.2281 | 182.7148 | 171.1711 | 201.0856 | 132.3441 | 130.6293 | 125.8859 | 116.2567 | 131.5703 | 240.7129 | 266.8593 | 265.8289 | 317.0837 | 280.057 | 286.7148 | 271.1711 | 307.3422 | 129.1141 | 69.2281 | 84.8289 | 129.1141 | 112.3137 | 132.3441 | 112.8859 | 134 | 102.5133 | 116.8004 | 124.7719 | 117.3726 | 84.6578 | 71.3707 | 92.6293 | 87.1426 | 123.7719 | 111.9144 | 112 | 73.2567 | 233.2548 | 184.5703 | 224.2281 | 318.7148 | 173.6578 | 209.23 | 196.5133 | 269 | 304.0589 | 241.0856 | 231.2852 | 268.77
1062 | 13.6764 | 189.7758 | 214.7754 | 222.3879 | 162.1632 | 135.4283 | 168.8166 | 155.7143 | 146.1224 | 122.3669 | 124.1224 | 120.5511 | 163 | 202.2651 | 169.5919 | 173.4287 | 195.6937 | 176.9592 | 147.3471 | 148.3063 | 164.6937 | 156.4283 | 315.1228 | 320.5103 | 161.857 | 215.5305 | 203.6529 | 206.3265 | 256.8978 | 132.0816 | 123.9184 | 126.3673 | 151.0206 | 174.6533 | 154.7552 | 164.3879 | 97.7143 | 280.3063 | 256.1632 | 228.8978 | 293.021 | 147.4287 | 154.5103 | 109.0202 | 115.0816 | 76.1632 | 98.6327 | 84.7552 | 122.5511 | 260.878 | 209.796 | 218.1632 | 173.796 | 179.4085 | 142.4897 | 139.3467 | 133.939 | 163.2448 | 165.2242 | 238.2448 | 154.9184 | 157.7552 | 153.7552 | 179.3059 | 124.6327 | 201.4489 | 174.6121 | 226.0408 | 200.3669 | 135.1426 | 167.8982 | 224.8978 | 207.3677 | 193.4081 | 183.2857 | 183.3471 | 288.5305 | 146.918 | 156.3467 | 151.9184 | 243.4287 | 135.0408 | 152.1838 | 165.0614 | 148.061 | 173.0408 | 177.4081 | 203.0206 | 110.5305 | 118.204 | 122 | 124.0614 | 149.4085 | 264.7143 | 229.2651 | 259.1838 | 353.3063 | 281.3471 | 278.5507 | 277.8978 | 319.8162 | 133.8978 | 76.9184 | 78.7552 | 133.8978 | 122.5103 | 110.0408 | 109.1632 | 132.939 | 120.4897 | 109.3063 | 117.9794 | 96.4081 | 73.6121 | 83.9184 | 79.9592 | 91.7552 | 115.1022 | 107.4901 | 111.5919 | 81.9592 | 275.9592 | 205.7548 | 230.9798 | 308.1022 | 162.6121 | 183.2857 | 213.5103 | 270.4283 | 273.4691 | 243.8776 | 239.8982 | 296.0202
1063 | 13.6874 | 161.7434 | 231.5663 | 208.3186 | 166.7965 | 169.0531 | 142.5399 | 150.2744 | 149.2655 | 176.3539 | 127.7965 | 110.3186 | 161.6194 | 231.9469 | 160.531 | 160.6903 | 212.6903 | 176.6903 | 132.9027 | 133.5399 | 181.1062 | 189.7345 | 272.062 | 308.2035 | 181.4159 | 228.3717 | 219 | 215.7523 | 278.1062 | 134.4248 | 122.6372 | 134.6814 | 128.1593 | 144.062 | 148.6814 | 150.3186 | 91.3717 | 262.8849 | 260.8496 | 251.3275 | 225.1771 | 134.3717 | 142.469 | 133.0531 | 117.5841 | 80 | 90.0531 | 78.6283 | 111.8938 | 211.6903 | 203.938 | 222.3717 | 167.8849 | 143.8054 | 153.8407 | 169.3008 | 110.9116 | 168.7345 | 193.3628 | 242.5663 | 153.6903 | 151.5221 | 147.4159 | 208.6725 | 117.6461 | 210.1415 | 189.4248 | 226.7345 | 254.6725 | 162.4248 | 143.2213 | 244.5663 | 170.1151 | 202.7876 | 192.2302 | 170.1771 | 301.4779 | 190.8318 | 186.8849 | 150.3717 | 228.7787 | 136.8496 | 134.1151 | 143.531 | 173.2655 | 174.2124 | 187.2655 | 181.7523 | 123.1593 | 123.2655 | 121.9469 | 103.8054 | 112 | 258.4779 | 258.469 | 239.8938 | 335.8849 | 267.4868 | 311.5485 | 296.7167 | 338.7345 | 154.8938 | 74.6814 | 73.1062 | 154.8938 | 110.5752 | 111.4779 | 113.531 | 107.4159 | 130.8849 | 94.3806 | 140.0973 | 105.1504 | 87.8407 | 82.1062 | 79.9558 | 86.2655 | 94.0531 | 74.3364 | 102.3186 | 81.3717 | 275.2124 | 245.407 | 208.5399 | 285.4601 | 176.8938 | 190.1062 | 201.3628 | 302.1946 | 307.2035 | 239.407 | 215.1151 | 318.7256
1064 | 13.6984 | 174.2153 | 208.2887 | 213.4769 | 181.8501 | 168.4306 | 169.0845 | 172.2153 | 154.7847 | 147.1962 | 141.2807 | 117.3732 | 137.6539 | 228.4498 | 168.2999 | 170.5806 | 222.5152 | 189.7847 | 148.9346 | 159.1691 | 183.9809 | 186.0462 | 289.8884 | 292.7576 | 171.9155 | 234.6077 | 219.3923 | 246.4115 | 281.3078 | 141.3461 | 134.6539 | 128.6077 | 131.4577 | 161.2999 | 142.8038 | 156.0654 | 98.0654 | 244.4386 | 274.2345 | 272.5806 | 263.8038 | 140.7384 | 132.4769 | 134.3923 | 126.8884 | 80.5885 | 91.1962 | 72.2616 | 109.9346 | 223.5423 | 186.7655 | 227.2999 | 147.7384 | 173.5152 | 150.8038 | 139.0654 | 143.7576 | 164.6539 | 180.823 | 215.8884 | 166.1962 | 142.3461 | 137.0654 | 184.8692 | 145.9538 | 176.9346 | 196.0191 | 223.7655 | 230.673 | 169.8692 | 164.5423 | 220.504 | 190.2616 | 199.9809 | 230.823 | 208.4769 | 309.8038 | 171.2424 | 168.1116 | 157 | 208.1116 | 150.8884 | 153.2153 | 152.2807 | 176.0383 | 179.0462 | 191.4769 | 211.2999 | 126.3923 | 127.0845 | 120.5423 | 135.8038 | 113.6348 | 267.1962 | 246.2536 | 238.4577 | 315.8692 | 290.5343 | 251.5694 | 255.4769 | 334.5231 | 152.5423 | 70.0462 | 75.1962 | 152.5423 | 101.7576 | 119.4769 | 122.673 | 99.6156 | 111.327 | 118.9346 | 124.5885 | 91.3732 | 83.9538 | 82.1691 | 97 | 91.5231 | 94.7384 | 115.6077 | 108 | 88.9155 | 279.1308 | 218.7001 | 233.4498 | 259.3078 | 174.7384 | 192.7847 | 190.9809 | 269.6077 | 292.6077 | 211 | 232.3844 | 297.8963
1065 | 13.7094 | 163.6215 | 236.2614 | 207.3206 | 192.6017 | 148.4081 | 157.1652 | 161.3884 | 164.4464 | 150.4661 | 131.3884 | 137.0777 | 148.2429 | 192.5438 | 144.8545 | 137.0099 | 187.0777 | 199.9125 | 148.7768 | 134.0875 | 198.5438 | 197.2713 | 276.4081 | 277.3304 | 182.7472 | 231.6412 | 224.0678 | 235.2035 | 300.5438 | 131.9223 | 142.1356 | 123 | 136.8348 | 138.3206 | 140.5438 | 155.3687 | 98.5339 | 263.1356 | 246.7472 | 237.3785 | 258.7472 | 137.3107 | 126.7867 | 141.1652 | 112.0875 | 88.6116 | 92.291 | 76.4661 | 108.1455 | 216.3009 | 209.8249 | 205.0974 | 144.7768 | 138.7768 | 148.5438 | 141.0099 | 127.0875 | 173.8446 | 163.767 | 200.6215 | 169.6991 | 133.8545 | 136.2133 | 183.4661 | 130.8446 | 176.9322 | 183.0099 | 247.058 | 226.233 | 168.7768 | 158.233 | 240.8348 | 192.2133 | 211.2035 | 215.6313 | 200.4562 | 307.233 | 188.3009 | 182.1356 | 157.4661 | 224.3107 | 135.7768 | 142 | 142.6991 | 149.0875 | 192.5241 | 182.8348 | 188.0099 | 129.5142 | 113.1455 | 114.233 | 131.9125 | 136.2232 | 270 | 210.3403 | 245.1554 | 314.5438 | 242.3983 | 273.3009 | 249.1652 | 340.5241 | 144.7571 | 84.6116 | 79.3206 | 144.7571 | 85.6313 | 112.5438 | 117.6893 | 134.9026 | 117.709 | 115.8249 | 130.7472 | 111.5438 | 69.6991 | 56.6991 | 94.5919 | 98.8446 | 91.5438 | 109.4562 | 107.767 | 101.767 | 278.2035 | 239.8927 | 199.1751 | 278.5438 | 168.3588 | 203.3009 | 204.5339 | 264.3884 | 288.8644 | 210.9223 | 196.2429 | 323.9125
1066 | 13.7204 | 170.4602 | 181.1696 | 220.7507 | 176.4499 | 176.6401 | 171.8098 | 166.9897 | 147.7095 | 156.4499 | 136.3599 | 135.0309 | 160.1209 | 200.7095 | 154.0103 | 150.1696 | 190.0694 | 186.6401 | 155.6607 | 145.7404 | 203.7404 | 155.4293 | 304.5501 | 303.5604 | 157.8894 | 259.761 | 213.5398 | 204.9691 | 300.7919 | 132.1696 | 122.6092 | 124.3496 | 122.91 | 154.09 | 145.6504 | 136.8098 | 92.0103 | 244.5989 | 221.0797 | 231.6195 | 235.6889 | 142.0797 | 147.0103 | 153.3805 | 124.1003 | 84.09 | 72.8997 | 82.7198 | 98.9897 | 207.5501 | 185.3496 | 231.3496 | 152.5604 | 147.4602 | 153.2905 | 150.1209 | 138.5604 | 172.5398 | 160.91 | 206.2905 | 178 | 144.7198 | 117.2493 | 188.2802 | 129 | 188.8098 | 192.5707 | 223.2802 | 229.09 | 174.4011 | 160.3702 | 227.8997 | 172.1696 | 181.329 | 232.3008 | 193.1902 | 309.0103 | 180.9897 | 160.3599 | 162.2802 | 225.4808 | 145.3599 | 140.1106 | 151.7198 | 161.4602 | 173.9203 | 170.3496 | 198.3805 | 102.4293 | 104.4396 | 118.1696 | 120.2596 | 127.1799 | 266.851 | 248.7404 | 243.9409 | 320.1902 | 259.6298 | 264.7301 | 264.3496 | 320.8407 | 131.1696 | 80.7198 | 94.3702 | 131.1696 | 105.2699 | 118.1902 | 114.6298 | 110.2699 | 138.3599 | 92.4293 | 104.8098 | 119.3496 | 75.9306 | 65.4499 | 67.6195 | 98.1696 | 99.3496 | 103.8098 | 106.7994 | 99 | 244.7301 | 206.419 | 226.0797 | 288.9588 | 139.4293 | 196.7095 | 198.1902 | 265.9409 | 311.09 | 210.09 | 211.2699 | 309.8304
1067 | 13.7314 | 169.4996 | 197.3973 | 200.2498 | 182.4317 | 168.7049 | 177.4434 | 174.0344 | 166.6362 | 162.534 | 136.8298 | 106.7385 | 134.2724 | 218.1023 | 144 | 162.4091 | 211.8181 | 182.5909 | 132.3068 | 134.7385 | 190.3864 | 180.1023 | 299.7955 | 286.8525 | 174.9547 | 228.0226 | 220.4317 | 233.5683 | 242.4996 | 143.1819 | 146.8525 | 135.8525 | 123.4317 | 156.534 | 132.3068 | 142.3411 | 84.5566 | 276.2959 | 229.9547 | 243.9095 | 274.8407 | 155.4543 | 136.7955 | 136.5909 | 116.3294 | 85 | 82.3068 | 89.1819 | 109.2045 | 203.8181 | 200.1249 | 242.8525 | 140.1475 | 142.7158 | 137.534 | 122.943 | 123.5909 | 176.466 | 159.6932 | 192.6815 | 177.3864 | 151.2842 | 141.2272 | 178.9547 | 130.943 | 199.1475 | 169.6362 | 215.2842 | 228.6706 | 138.8181 | 153.4887 | 233.6253 | 186.0453 | 213.5683 | 206.3294 | 186.3294 | 299.8181 | 189.8751 | 165.943 | 153.466 | 202.8634 | 148.4887 | 123.3521 | 157.8751 | 154.3638 | 165.6589 | 184 | 182.2045 | 126.5909 | 117.1593 | 127.7502 | 130.2389 | 126.7502 | 235.1023 | 235.6932 | 213.4091 | 313.1249 | 263.0344 | 262.9204 | 277.2842 | 298.0796 | 140.9547 | 86.2615 | 87.7955 | 140.9547 | 108.3068 | 108.7728 | 118.7502 | 114.4317 | 137.8072 | 117 | 113.2045 | 132.2842 | 56.7385 | 70.2045 | 81.5457 | 107.3411 | 113 | 109.8525 | 121.7276 | 95.7276 | 242.1023 | 238.9547 | 233.523 | 323.9891 | 164.3068 | 211.0344 | 192.4543 | 240.42 | 309.9547 | 212.534 | 212.9774 | 298.3068
1068 | 13.7424 | 205.907 | 233.1361 | 220.4582 | 190.9906 | 137.3005 | 155 | 149.9474 | 181 | 174.5108 | 108.7709 | 123.1455 | 158.9906 | 219.5728 | 146.2911 | 164.969 | 216.3652 | 177.7399 | 137.1765 | 151.2601 | 185.4582 | 178.7089 | 300.5202 | 271.4366 | 159.062 | 236.4272 | 234.3747 | 222.4892 | 280.1671 | 134.2817 | 130.5202 | 117.4582 | 134.6253 | 168.2817 | 134.4272 | 123.4676 | 109.4056 | 232.1361 | 213.7183 | 216.3625 | 282.4084 | 174.5944 | 136.1455 | 135.062 | 126.8545 | 84.5418 | 83.7399 | 80.3962 | 111.9164 | 209.6253 | 209.4272 | 223.7709 | 159.2291 | 147.6253 | 151.1146 | 139.4272 | 119.0836 | 163.9164 | 158.4892 | 224.3127 | 169.0216 | 146.6038 | 150.1671 | 161.8019 | 147.7709 | 213.6469 | 185.2601 | 204.876 | 216.8854 | 147.2601 | 145.6779 | 207.5512 | 201.2507 | 202.8329 | 217.5418 | 199.9474 | 310.093 | 176.6779 | 182.1981 | 142.6348 | 228 | 142.2817 | 143.1981 | 148 | 142.5202 | 186.8235 | 181.5944 | 181.4798 | 122.5418 | 102.031 | 107.4272 | 107.9474 | 109.062 | 235.8854 | 233.5728 | 220.093 | 322.5418 | 236.4272 | 273.7493 | 269.3962 | 296.5296 | 124.7183 | 72.3747 | 85.1981 | 124.7183 | 113.1765 | 98.5728 | 100.031 | 128.4892 | 101.2291 | 116.7709 | 114.0836 | 125.5418 | 72.1146 | 70.5108 | 61.8329 | 85.3747 | 110.1361 | 89.3962 | 92.7709 | 68.1455 | 244.9474 | 224.093 | 204.1455 | 282.9784 | 167.3437 | 185 | 213.5418 | 284.3343 | 290.2817 | 224.5108 | 206.1765 | 300.1981
1069 | 13.7534 | 180.8584 | 210.1121 | 223.1121 | 163.298 | 181.6872 | 156.2685 | 165.6342 | 179.6047 | 160.8436 | 107.3805 | 130.7315 | 132.3128 | 223.4926 | 160.5752 | 155.351 | 199.5517 | 169.3953 | 156.4101 | 160.7463 | 189.7611 | 161.6342 | 318.351 | 302.4396 | 175.2537 | 231.4926 | 242.5899 | 233.3658 | 266.2537 | 121.6342 | 148.4779 | 122.0148 | 126.9174 | 154.3658 | 132.1564 | 156.7315 | 126.4779 | 218.8792 | 230.0443 | 273.2624 | 222.3423 | 157.1416 | 146.0148 | 151.5074 | 118.6342 | 80.4926 | 74.1268 | 67.8732 | 118.1121 | 201.7906 | 204.8732 | 223.3953 | 161.2537 | 137.7611 | 151.4926 | 135.1268 | 113.5221 | 172.1416 | 173.5517 | 215.0678 | 148.4101 | 159 | 135.4926 | 170.5369 | 142.1946 | 179.8732 | 191.3215 | 179.0886 | 216 | 158.649 | 126.4691 | 232.9705 | 181.7758 | 213.1946 | 210.8289 | 212.4631 | 330.5752 | 154.8054 | 173.8436 | 164.9027 | 228 | 131.1564 | 138.7758 | 146.8584 | 166.4396 | 172.9174 | 168.4544 | 164.9174 | 123.6933 | 110.7611 | 103.5074 | 123.8879 | 125.7611 | 238.8054 | 239.5221 | 243.2389 | 314.9409 | 235.298 | 290.4336 | 261.1859 | 344.5899 | 136.6047 | 81.4779 | 80.9027 | 136.6047 | 126.4483 | 104.2685 | 109.2685 | 137.8436 | 107.8201 | 116.3953 | 109.5369 | 120.7315 | 72.3658 | 61.1564 | 76.7611 | 96.1268 | 90.4101 | 78.5221 | 91.6342 | 77.8879 | 257.2094 | 244.0678 | 213.5074 | 301.9409 | 168.351 | 183.351 | 209.6195 | 258.0295 | 274.7168 | 213.2537 | 223.7611 | 291.7168
1070 | 13.7644 | 174.9479 | 201.3564 | 216.5826 | 182.939 | 151.8609 | 165.974 | 170 | 174.3132 | 145.3043 | 111.1131 | 122 | 156.9911 | 222.9219 | 137.974 | 146.9219 | 222.1912 | 180.2522 | 171.4695 | 160.5305 | 196.3914 | 166.4174 | 306.5826 | 329.6607 | 177.5565 | 229.8088 | 228.2262 | 228.3043 | 269.5305 | 128.3653 | 133.4085 | 130.2522 | 146.7217 | 151.1131 | 145.1912 | 146.8869 | 114.4695 | 277.2173 | 249.1912 | 222.087 | 260.7128 | 164.026 | 154.2522 | 156.9479 | 122.4435 | 77.6957 | 73.0521 | 66.8609 | 113.6696 | 217.7999 | 205.3914 | 237.8698 | 159.5216 | 143 | 147.8609 | 133.913 | 125.3914 | 179.8609 | 192.9911 | 189.8088 | 165.4174 | 156.3564 | 133.9479 | 184.3824 | 120.3132 | 178.7217 | 168.7827 | 216.965 | 213.3564 | 169.1652 | 170.026 | 217.4695 | 193.5826 | 192.5655 | 193.4524 | 197.087 | 308.8088 | 179.7478 | 160.5305 | 173.2694 | 225.7738 | 145.4435 | 152.8088 | 138.026 | 188.9301 | 190.7738 | 201.6868 | 185.5565 | 151.1302 | 115.5826 | 109.3653 | 129.4435 | 130.1652 | 270.1481 | 251.5305 | 237.8609 | 292.1481 | 258 | 258.7217 | 284.2954 | 327.5826 | 128.1131 | 70.8609 | 90.1042 | 128.1131 | 103.6696 | 112.4435 | 119.8088 | 122.3043 | 136.6868 | 126 | 125.3304 | 110.7478 | 67.5826 | 75.026 | 81.7217 | 96.4435 | 105.7478 | 92.7567 | 96.2783 | 84.5565 | 236.4695 | 214.913 | 217.974 | 276.5045 | 158.3914 | 173.9479 | 204.7738 | 272.974 | 264.8437 | 216.6696 | 227.3304 | 279.3393
1071 | 13.7754 | 183.5172 | 188.9371 | 214.6057 | 164.7571 | 148.2743 | 170.4857 | 169.5457 | 192.2172 | 144.4828 | 116.6372 | 120.3343 | 128.7257 | 239.0314 | 144 | 162.7286 | 212.3628 | 185.7286 | 162.9086 | 163.9429 | 206.9686 | 167.94 | 302.4857 | 304.0029 | 174.3372 | 237.5172 | 241.5457 | 222.4857 | 279.3029 | 138.3057 | 106.06 | 132.3972 | 143.3343 | 159.6657 | 137.4828 | 141.06 | 102.1229 | 259.5457 | 241.3314 | 234.0914 | 227.6971 | 156.1829 | 161.3943 | 163.0943 | 119.6057 | 81.0914 | 63.2714 | 68.2714 | 123.6971 | 201.3314 | 215.3628 | 260.58 | 138.9086 | 147.3914 | 145.7886 | 120.8486 | 134.7571 | 174.3057 | 161.2429 | 201.6057 | 166.7886 | 143.0286 | 143.5772 | 161.0914 | 144.4257 | 177.3029 | 189.3314 | 195.18 | 198.9686 | 164.4543 | 163.0914 | 208.7571 | 189.94 | 223.3657 | 217.6686 | 209.8486 | 310.7629 | 176.0885 | 168.94 | 146.5743 | 211.8486 | 144.7257 | 158.7 | 133.2114 | 143.24 | 175.1829 | 175.7571 | 189.3029 | 121.9086 | 113 | 123.0914 | 125.6971 | 125.9086 | 292.6115 | 261.6057 | 235.6372 | 324.4543 | 258.6057 | 253.3657 | 253.6943 | 324.6971 | 134.3943 | 70.7571 | 66 | 134.3943 | 112.4857 | 109.1514 | 127.9714 | 116.1829 | 110.3029 | 123.1229 | 114.8486 | 104.9686 | 67.4228 | 69 | 81.2114 | 93.6057 | 98.7571 | 114.6372 | 97.5457 | 86.7886 | 226.0914 | 204.5743 | 220.3657 | 289.4257 | 169.12 | 186.3029 | 191.4543 | 278.2429 | 300.0914 | 227.9086 | 215.1829 | 302.12
1072 | 13.7864 | 165.6372 | 210.8363 | 221.6018 | 168.8363 | 137.4204 | 165.1107 | 171.7478 | 143.6549 | 163.0177 | 110.9646 | 111 | 146.1283 | 227.5088 | 146.7832 | 149.8363 | 215.5796 | 176.7655 | 166.6902 | 130.1637 | 212.1061 | 165.9293 | 292.6902 | 274.854 | 140.1107 | 218.4912 | 234.907 | 216.4558 | 280.5088 | 113.4735 | 114.1283 | 107.0576 | 131.0531 | 169.8186 | 154.0531 | 144.7079 | 86.8186 | 259.4558 | 256.0707 | 229.1637 | 233.8585 | 147.146 | 157.1814 | 131.6372 | 124.146 | 76.4912 | 75.8363 | 81 | 122.1637 | 218.1991 | 220.8717 | 221.6195 | 147.7655 | 169.5442 | 136.2168 | 123.9293 | 135.8893 | 153.0753 | 158.146 | 205.3274 | 160.1637 | 158.3628 | 132.2921 | 157.4735 | 130.9823 | 178.3451 | 206.1991 | 211.0354 | 212.9469 | 167.6549 | 158.6549 | 216.4381 | 186.9469 | 202.0177 | 198.6726 | 205.3982 | 263.2567 | 196.0531 | 164.3098 | 161.9823 | 209.3628 | 160.6549 | 129.3274 | 142.7832 | 163.3982 | 166.6372 | 180.3274 | 191.9823 | 129.1283 | 115.2921 | 118.6549 | 121.3805 | 131.1637 | 240.6549 | 259.1061 | 227.8363 | 323.7256 | 257.7433 | 238.7302 | 274.5973 | 322.0177 | 131 | 77.1283 | 69.1107 | 131 | 106.4558 | 112.4558 | 109.5265 | 107.4735 | 111.1814 | 106.0177 | 113.0177 | 114.3628 | 79.854 | 70.8009 | 86.854 | 97.1637 | 101.6902 | 105.5265 | 97.2921 | 78.0354 | 227.7125 | 219.4912 | 202.1283 | 272.2168 | 175.5973 | 192.7478 | 194.9823 | 267.9424 | 293.854 | 233.3274 | 206.4735 | 307.7787
1073 | 13.7974 | 174.352 | 204.7197 | 236.6558 | 172.7523 | 158.5919 | 177.8318 | 189.8397 | 149.2881 | 158.704 | 121.352 | 113.4641 | 157 | 224.648 | 156.0717 | 145.8318 | 197.2242 | 196.352 | 157.0078 | 135.4002 | 183.4081 | 179.6636 | 290.5763 | 273.9283 | 157.0561 | 220.824 | 213.296 | 225.3038 | 274.4798 | 122.0561 | 121.4798 | 132.8957 | 121.6323 | 169.9518 | 140.7601 | 134.0561 | 91.176 | 271.824 | 237.4081 | 229.648 | 267.3116 | 151.0639 | 151.0639 | 139.824 | 128.2399 | 80.9361 | 72.0078 | 83.4641 | 124.9361 | 210.7601 | 205.5998 | 235.704 | 167.528 | 153.3038 | 127.1043 | 142.4159 | 122.4641 | 176.8879 | 164.88 | 205.4159 | 161.352 | 150.528 | 140.4798 | 166.7601 | 137.9361 | 176.0561 | 191.7197 | 202.2321 | 229.4081 | 173.4641 | 159.1838 | 231.7119 | 197.2477 | 195.5919 | 197.88 | 189.1121 | 284.12 | 179.5919 | 169.7679 | 164.1838 | 205.5763 | 158.0157 | 131.352 | 157.352 | 147.4641 | 175.88 | 184.2881 | 193.4798 | 139.472 | 127 | 121.296 | 109.2321 | 127.9518 | 247.8722 | 229.528 | 228.4081 | 306.824 | 238.1121 | 268.88 | 246.472 | 316.824 | 130.648 | 84.8318 | 87.6401 | 130.648 | 116.0078 | 124.824 | 104.8162 | 114.824 | 107.352 | 99.7679 | 106.5919 | 111.9843 | 76.8162 | 79.648 | 80.12 | 99.2321 | 96.5841 | 95.1838 | 107.5919 | 67.12 | 259.3599 | 222.528 | 218.2803 | 263.8083 | 149.7601 | 213.8318 | 197.8879 | 238.4081 | 286.4159 | 227.7836 | 214.704 | 280.4002
1074 | 13.8084 | 174.4936 | 185.8373 | 219.8702 | 194.4351 | 151.0585 | 159.6106 | 164.0128 | 161.1298 | 156.5393 | 122.4351 | 121.4223 | 156.0585 | 223.3766 | 136.9543 | 130.8117 | 210.5978 | 194.6106 | 146.3894 | 156.7532 | 189.8117 | 149.4351 | 305.2212 | 293.9872 | 161.0585 | 227.9086 | 212.0713 | 208.9415 | 258.9415 | 126.6234 | 109.7075 | 113.5192 | 148.3766 | 151.9415 | 148.6234 | 138.6234 | 90.8702 | 265.9159 | 247.3894 | 230.2596 | 234.3053 | 141.4351 | 139.7404 | 135.0457 | 120.7532 | 90.3766 | 62.5192 | 92.3638 | 130.0457 | 213.1627 | 193.0969 | 234.6691 | 168.3053 | 136.5649 | 151.7532 | 134.6234 | 133.2468 | 166.2468 | 165.6106 | 198.3766 | 164.6947 | 149.9872 | 125.3181 | 176.1298 | 146.2468 | 180.6234 | 169.8245 | 208.5649 | 236.7532 | 182.3638 | 149.0128 | 217.6106 | 176.1298 | 188.6234 | 204.2596 | 201.4479 | 284.5192 | 172.2468 | 162.117 | 155.1426 | 223.7989 | 129.8117 | 137.1426 | 156.9287 | 157.4936 | 178.6819 | 195.7532 | 176.8117 | 132.4808 | 125.8702 | 113.1042 | 115.3766 | 108.6234 | 265.0585 | 231.0585 | 235.5649 | 309.2011 | 247.0585 | 271.8702 | 240.6106 | 314.4936 | 132.2011 | 71.1298 | 94.1627 | 132.2011 | 102.5649 | 123.2468 | 116.3053 | 115.6947 | 106.7404 | 98.1426 | 101.1298 | 135.6691 | 87.9287 | 77.6234 | 72.9872 | 102.9287 | 102.8702 | 87.0841 | 100.6234 | 66.0128 | 249.8245 | 223.4936 | 234.7148 | 288.3638 | 155.7404 | 199.9415 | 191.9543 | 248.9543 | 276.7404 | 193.6234 | 220.4479 | 303.6362
1075 | 13.8194 | 168 | 211 | 208.7817 | 192.6018 | 145.5959 | 145.6018 | 176.1976 | 169.2095 | 138.2154 | 116.9911 | 105.7994 | 154.0059 | 220.587 | 155.003 | 132.6077 | 190.8231 | 180 | 155.584 | 158.1947 | 186.1917 | 149.0059 | 282.8201 | 281 | 155.3953 | 256.1828 | 219.3953 | 217.413 | 264.003 | 123.5959 | 129.5988 | 130.5929 | 144.7846 | 151.9882 | 149.2065 | 138.2036 | 91.2154 | 249.6166 | 256.3834 | 240.2006 | 223.9911 | 138.7994 | 134.0118 | 122.2124 | 130.2183 | 88.5899 | 82.2006 | 80.3982 | 113.2006 | 194.0178 | 221.7817 | 216.4012 | 158.1917 | 140.8053 | 154.3982 | 134.6047 | 126.9911 | 167.2124 | 157.4189 | 200.4012 | 172.4012 | 139.0059 | 129.997 | 177.3893 | 144.6047 | 180.6047 | 182.4012 | 212.2036 | 235.9882 | 163.5781 | 161.5988 | 206.8112 | 186.2154 | 193.6195 | 211.7935 | 211.3982 | 299.1858 | 171.4071 | 153.5988 | 173 | 209.1976 | 131.4071 | 152.1917 | 149.8053 | 148.5929 | 173.2006 | 196.5929 | 177.4041 | 117.6136 | 122.6047 | 94.2065 | 112.9882 | 111.8142 | 260.1976 | 233.2183 | 239.2036 | 318.1858 | 244.4041 | 263.3893 | 225.7994 | 313.416 | 140.7846 | 78.4071 | 71.0059 | 140.7846 | 112.8231 | 122.4071 | 110.4041 | 122.3982 | 97.6018 | 111.1858 | 105.7994 | 119.4071 | 77.3953 | 76.2006 | 61.6047 | 95.003 | 100.2065 | 109 | 95.3893 | 78.3982 | 258.9911 | 221.413 | 206.4219 | 282.2154 | 149.2095 | 206.6077 | 207.997 | 262.5899 | 273.8201 | 193.003 | 228.1917 | 314.5988
1076 | 13.8304 | 173.1092 | 210.1485 | 186.9804 | 191.381 | 139.1485 | 146.5098 | 169.1681 | 178.1681 | 156.6583 | 109.6834 | 108.619 | 163.7031 | 208.9607 | 156.2325 | 141.7227 | 214.5491 | 180.2129 | 139.535 | 147.5939 | 176.2773 | 158.7031 | 309.7871 | 277.5939 | 146.4454 | 234.2773 | 218.1092 | 232.4454 | 270.3417 | 126.5153 | 125.0196 | 121.6386 | 127.8319 | 140.6834 | 151.1877 | 142.1485 | 106.465 | 273.4902 | 236.1288 | 238.0196 | 212.4258 | 137.7871 | 153.8319 | 136.2325 | 149.0393 | 75 | 79.5939 | 76.0841 | 112.9356 | 221.1932 | 201.0448 | 217.5742 | 150.4061 | 145.6583 | 152.4258 | 137.5939 | 118.4061 | 184 | 177.2522 | 207.535 | 174.4258 | 144.8712 | 123.0196 | 160.2325 | 144.1877 | 184.0196 | 186.1288 | 215.5098 | 218.0841 | 143.154 | 158.5098 | 226.4705 | 205.2969 | 213.1877 | 196.8264 | 202.6135 | 278.9356 | 179.2969 | 152.4258 | 173.6386 | 203.8712 | 139.0841 | 148.0251 | 156.5742 | 140.4902 | 173.3614 | 187.4258 | 184.0644 | 136 | 124.7423 | 104.7031 | 103.1736 | 130.5742 | 263.3865 | 256.2969 | 248.0448 | 304.7478 | 245.7423 | 252.1932 | 230.1092 | 330.3166 | 121.7031 | 83.9552 | 75.5939 | 121.7031 | 142.5098 | 126.0393 | 118.5546 | 124.6834 | 102.5546 | 97.535 | 107.1288 | 127.0841 | 72.2773 | 75.0841 | 69.0644 | 98.5742 | 103.6779 | 108.3614 | 79.7227 | 76 | 246.1485 | 231.5491 | 237.7031 | 299.381 | 160.2969 | 218.0644 | 202.5098 | 246.0196 | 296.9552 | 202.1092 | 214.2325 | 311.5098
1077 | 13.8414 | 185.0196 | 208.5762 | 205.8811 | 179.351 | 140.0531 | 145.2783 | 158.6028 | 162.8741 | 148.4503 | 125.649 | 119.9734 | 164.821 | 225.5958 | 144.8741 | 148.709 | 192.6294 | 182.351 | 156.172 | 134.7748 | 184.3776 | 164.7748 | 301.7944 | 271.0797 | 141.0531 | 237.4238 | 230.9203 | 221.1986 | 275.172 | 158.2252 | 117.6028 | 124.6755 | 145.828 | 154.8476 | 130.2783 | 134.9469 | 84.649 | 276.7483 | 244.2252 | 228.8014 | 226.8349 | 138.1259 | 163.7217 | 129.1524 | 122.2713 | 74.7748 | 71.9538 | 68.5497 | 108.7748 | 224.6686 | 215.7748 | 221.179 | 158.9469 | 133.3972 | 151.9734 | 131.9804 | 130.7748 | 179.2713 | 157.5762 | 222.1455 | 166.7679 | 148.2587 | 114.7021 | 150.4503 | 122.1524 | 173 | 196.7021 | 205.7217 | 221.1328 | 172.2713 | 156.3776 | 240.7483 | 198.3245 | 189.5762 | 178.2783 | 183.0531 | 282.6559 | 168.9469 | 157.6028 | 175.0993 | 193.2979 | 138.53 | 168.3706 | 148.9203 | 148.2517 | 173.4769 | 180.2182 | 187.3245 | 134.1986 | 109 | 109.8741 | 124.8741 | 127.649 | 278.6686 | 247.0727 | 259.8476 | 317.2182 | 234.5035 | 256.3441 | 254.6294 | 315.0266 | 123.9469 | 69.4503 | 63.4503 | 123.9469 | 132.2713 | 103.0993 | 121.6951 | 137.7217 | 106.3706 | 113.4965 | 115.4503 | 118.4238 | 76.5497 | 67.0993 | 74.351 | 95.8741 | 86.2517 | 106 | 79.7287 | 79.828 | 244.351 | 208.5035 | 234.5427 | 287.1259 | 157.6028 | 221.0993 | 199.9273 | 241.53 | 279.5231 | 222.5762 | 207.3776 | 312.0797
1078 | 13.8524 | 163.8496 | 210.6755 | 227.0634 | 185.1873 | 157.562 | 159.7625 | 168.8628 | 156.1504 | 149.7625 | 123.8496 | 114.8997 | 148.4248 | 207.2375 | 140.5251 | 182.8259 | 211.0501 | 184.6253 | 145.6121 | 140.649 | 191.9631 | 164 | 278.1873 | 302.9234 | 154.5251 | 224.1637 | 214.5118 | 215.4749 | 265.562 | 154.7256 | 134.9868 | 124.6253 | 153.7758 | 141.3879 | 140.2507 | 125.7493 | 80 | 276.124 | 244.2876 | 239.9868 | 268.6253 | 140.8127 | 150.4248 | 136.6755 | 111.6992 | 73.5251 | 84.0132 | 65.8127 | 115.8363 | 197.8997 | 213.3377 | 232.1135 | 151.1741 | 125.2744 | 142.1504 | 149.8259 | 128.1003 | 168.2242 | 169.649 | 197.2507 | 141.4116 | 174.8893 | 123.5251 | 155.3245 | 128.4882 | 170.3879 | 196.9763 | 196.2242 | 249.5884 | 160.7493 | 162.0634 | 228.0132 | 193.3879 | 200.9366 | 198.4617 | 195.8127 | 301.4011 | 163.0738 | 164.7758 | 168.913 | 192.7862 | 152.9262 | 147.5752 | 131.562 | 154.1003 | 182.4749 | 152.6121 | 176.6887 | 132.0369 | 114.2242 | 108.8496 | 119.1003 | 121.1003 | 251.6623 | 240.3245 | 240.2138 | 296.2611 | 253.7994 | 238.0738 | 272.8127 | 313.9262 | 115.9366 | 70.0501 | 68.562 | 115.9366 | 112.438 | 100.2375 | 99.5251 | 117.5383 | 94.5989 | 98 | 114.1504 | 112.7625 | 72.3879 | 62.8127 | 78.2876 | 90.8127 | 95.6623 | 99.5884 | 94.8127 | 95.3747 | 244.4882 | 229.4617 | 209.3614 | 283.1637 | 171.1873 | 211.351 | 210.4749 | 259.0132 | 273.3747 | 225.1504 | 213.5383 | 329.438
1079 | 13.8634 | 168.5044 | 201.998 | 204.5034 | 185.2537 | 165.0039 | 163.4956 | 168.7434 | 149.7473 | 146.5024 | 127.0059 | 117.5034 | 155.4976 | 211.4966 | 139 | 157.5054 | 204.5034 | 182.2449 | 153.001 | 162.999 | 173.5054 | 169.2449 | 290.242 | 331.5063 | 152.7502 | 204.2449 | 232.9961 | 224.9922 | 277 | 135.2556 | 143.5073 | 119.4976 | 139.4976 | 137.7454 | 127.4976 | 132.758 | 84.4956 | 289.759 | 250.242 | 253.0029 | 259.7512 | 139.4976 | 148.5063 | 131.2449 | 125.2546 | 74.9971 | 72.7454 | 68.4937 | 130.76 | 206.4976 | 205.2527 | 212.2478 | 163.7571 | 144.7444 | 136.4966 | 135.4946 | 123.7483 | 181.7532 | 183.0078 | 191.9902 | 159.5024 | 141.9941 | 126.2459 | 158.2576 | 117.4966 | 164.9971 | 171.2468 | 206.7561 | 230.4985 | 172.0039 | 145.9971 | 213.998 | 186.998 | 209.5102 | 204.5112 | 190.5015 | 278.2488 | 185.5034 | 157.241 | 163.9951 | 234.7366 | 137.7395 | 144.4985 | 143 | 148.2498 | 174.759 | 164.9961 | 158.7415 | 141.7532 | 127.2537 | 112.2556 | 112.5005 | 118.2468 | 257.4995 | 246.7541 | 218.9932 | 325.5005 | 272.9932 | 255.0088 | 261.0078 | 293.2449 | 134.001 | 70.9961 | 77.5024 | 134.001 | 103.9971 | 105.9951 | 105.7424 | 104.7454 | 122.7463 | 99.7483 | 108.7463 | 108.5034 | 66.4976 | 65.7434 | 75.001 | 87 | 100.5005 | 86.7424 | 94.7463 | 99.0039 | 236.4937 | 238.758 | 225.7571 | 263.7444 | 174.2507 | 189.5005 | 205.7561 | 247.4956 | 285.7454 | 215.2507 | 196.999 | 313.2546
1080 | 13.8744 | 156.5723 | 204.5934 | 190.3313 | 174.5241 | 154.0482 | 173.0693 | 186.6416 | 159.5723 | 141.0964 | 109.7861 | 113.0271 | 165.8825 | 220.1657 | 139.5241 | 144.4066 | 199.2409 | 191.4488 | 149.4759 | 157.8765 | 163.5512 | 174.5181 | 311.3313 | 316.7168 | 158.8132 | 218.4277 | 235.3043 | 246.3795 | 269.9247 | 119.3102 | 121.7861 | 124.3795 | 145.4277 | 154.8825 | 139.1928 | 117.9578 | 93.5452 | 260.6898 | 260.5874 | 244.5241 | 256.7861 | 141.235 | 135.8132 | 142.0211 | 114.6687 | 89.2409 | 83.5934 | 86.9518 | 105.503 | 215.5723 | 206.6957 | 225.8132 | 146.4548 | 159.3795 | 148.8343 | 146.235 | 128.2139 | 172.5241 | 158.738 | 223.0482 | 150.4277 | 159.4759 | 134.8072 | 141.0271 | 124.3313 | 168.497 | 182.8343 | 190.6205 | 231.3313 | 163.4066 | 152.1175 | 224.7168 | 196.1446 | 186.1235 | 178.9096 | 187.0482 | 286.1928 | 170.8072 | 175.8283 | 177.9518 | 253.988 | 162.6627 | 148.738 | 142.738 | 151.6205 | 165.0813 | 175.9518 | 182.1175 | 131.2139 | 116 | 103.9096 | 112.3102 | 127.738 | 254.2832 | 241.0753 | 235.8072 | 320.0693 | 280.6356 | 238.7439 | 234.6416 | 305.3313 | 127.8554 | 84.3102 | 73.4066 | 127.8554 | 112.2139 | 117.0693 | 126.1175 | 115.3313 | 124.0422 | 108.4066 | 119.738 | 100.3584 | 76.0964 | 87.3102 | 73.3102 | 94.0753 | 97.1657 | 106.0693 | 104.4277 | 84.3795 | 256.7861 | 211.3795 | 208.4548 | 284.1446 | 168.5934 | 185.3795 | 194.0753 | 258.1175 | 294.7591 | 216.1446 | 199.2139 | 293.497
1081 | 13.8854 | 158.531 | 196.646 | 190.6992 | 178.469 | 158.0974 | 163.0974 | 175.6106 | 163.177 | 159.6196 | 109.8053 | 128.0796 | 167.6902 | 216.9204 | 146.4868 | 150.7079 | 212.6283 | 175.469 | 146.354 | 136.646 | 184.1947 | 150.2124 | 295.5132 | 316.0087 | 167.0265 | 219.2124 | 206.9026 | 232.9646 | 258.2301 | 115.5929 | 121.531 | 118.0974 | 137.9823 | 159.7079 | 147.1593 | 145.7257 | 87.0354 | 254.531 | 228.5045 | 243.8053 | 253.2389 | 134.0532 | 146.2212 | 131.2921 | 125 | 98.5132 | 72.9026 | 82.0796 | 120.177 | 218.0796 | 224.6725 | 241.9823 | 152.4159 | 154.469 | 152.9026 | 142.3452 | 124.6283 | 178.6638 | 165.4071 | 222.9823 | 149.5664 | 156.6283 | 127.9381 | 157.4513 | 115.9204 | 158.4868 | 182.2389 | 203.2124 | 226.2124 | 172.4513 | 142.9026 | 238.823 | 208.8407 | 209.2743 | 203.4513 | 193.2921 | 287.8495 | 162.8407 | 156.0354 | 176.3717 | 198.8585 | 139.8053 | 136.7521 | 138.7079 | 154.0619 | 208.5929 | 175.1947 | 171.8053 | 127.6283 | 118.1947 | 121.8672 | 112.9823 | 134.1328 | 249.5045 | 261.5487 | 224.2743 | 316.4071 | 249.8585 | 261.0442 | 232.1151 | 304.6018 | 123.1151 | 83.6106 | 83.5487 | 123.1151 | 105.3362 | 106.2743 | 116.0796 | 108.2921 | 97.0974 | 110.8672 | 117.0796 | 111.3894 | 77.8849 | 87.4336 | 72.3362 | 112.0796 | 94.1947 | 97.7434 | 98.9026 | 73.7079 | 252.9646 | 209.2124 | 219.6283 | 289.9823 | 157.354 | 190.6196 | 197.814 | 247.531 | 283.8407 | 218.6902 | 197.823 | 282.1151
1082 | 13.8964 | 159.1657 | 209.0256 | 210.1145 | 186.7198 | 160.4267 | 166.8599 | 169.446 | 158.4203 | 178.3819 | 106.0064 | 124.1465 | 154.1529 | 216.8407 | 153.8343 | 143.1465 | 193.522 | 181.7134 | 142.2866 | 145.0128 | 190.5733 | 166.8663 | 281 | 280.4331 | 144.0192 | 230.1337 | 212.8855 | 220.7262 | 284.2994 | 109.4716 | 112.1337 | 116.1273 | 128.2802 | 152.7198 | 133.5604 | 150.7326 | 98.1273 | 244.8471 | 248.1337 | 245.1657 | 238 | 153.1273 | 125.5797 | 137.4203 | 126.1465 | 81.4203 | 68.2802 | 73.8471 | 110.8343 | 221.8855 | 195.7006 | 227.9808 | 146.4652 | 158.707 | 145.7006 | 163.8215 | 118.707 | 184.6942 | 182.4203 | 211.2738 | 161.293 | 149.2738 | 143.0064 | 163.1657 | 121 | 175.5797 | 171.2994 | 209.5476 | 234.8407 | 159.5348 | 140.2866 | 237.8471 | 209.8279 | 204.554 | 203.1465 | 201.4267 | 270.4588 | 173 | 169.707 | 175.7006 | 213.2738 | 136.5797 | 113.5925 | 136.3058 | 136.9872 | 186.4203 | 177.5604 | 173.1657 | 118.8407 | 121.4203 | 105.8663 | 113.0256 | 145.8343 | 268.2738 | 264.7198 | 235.6053 | 327.6878 | 262.2674 | 239.4267 | 241.8535 | 324.2674 | 134.8599 | 70.5669 | 85 | 134.8599 | 100.1657 | 107.2866 | 114.1529 | 118.4331 | 106.0192 | 100.0256 | 110.8535 | 114.6878 | 71.2994 | 76.8535 | 65.7326 | 102.9872 | 100.5733 | 98.9808 | 101.446 | 71.8791 | 239.5797 | 219.8471 | 216 | 282.5733 | 164.7518 | 221.4203 | 161.5925 | 248.4524 | 301.1657 | 207.1593 | 202.5797 | 294.1465
1083 | 13.9074 | 177 | 215.8475 | 191.2881 | 187.413 | 156.9076 | 169 | 178.5162 | 152.8968 | 154.5547 | 116.3806 | 124.0108 | 161.4022 | 200.7119 | 142.5762 | 145.1032 | 168.7503 | 181 | 148.3806 | 150.8151 | 183.9292 | 168.3205 | 290.2665 | 292.0708 | 165.5762 | 223.2989 | 233.8043 | 227.0108 | 288.1248 | 133.1356 | 104.7011 | 108.7827 | 128.7827 | 154.9076 | 127.0924 | 152.7443 | 85.4022 | 250.7503 | 235.3205 | 254.0324 | 232.0216 | 137.413 | 129.9076 | 132.7935 | 128.7011 | 79.1849 | 62.8043 | 69.587 | 99.2773 | 235.6302 | 194.2665 | 212.7011 | 176.1957 | 147.3205 | 130.5162 | 132.114 | 114.1957 | 170.7935 | 173.3097 | 210.0708 | 164.0108 | 143.886 | 145.0108 | 180.7011 | 123.0924 | 175.7227 | 180.5054 | 189.5054 | 218.413 | 144.7503 | 145.7827 | 221.0324 | 181.0216 | 195.7827 | 205.1032 | 197.9076 | 287.4238 | 168.5162 | 168.1849 | 159.9184 | 208.7827 | 145.0924 | 134.4838 | 150.8043 | 125.2065 | 173.7227 | 169.8968 | 192.1957 | 112.8752 | 109.9184 | 113 | 131.2065 | 127.4022 | 256.0108 | 259.7335 | 244.6626 | 306.1032 | 246.5054 | 232.9184 | 246.7719 | 308.2065 | 136.1032 | 70.886 | 85 | 136.1032 | 117.4022 | 110.9892 | 119.0108 | 120.2065 | 114.4238 | 113.1248 | 105.9076 | 102.3698 | 79.3097 | 78.1849 | 79.7011 | 89.114 | 102 | 88.1849 | 112.0108 | 91.6903 | 241.5162 | 223.0601 | 209.7227 | 284.5978 | 186.527 | 204.2389 | 179.7935 | 264.3097 | 304.9507 | 213.4238 | 215.5762 | 298.1957
1084 | 13.9184 | 165.4852 | 195.5812 | 206.7803 | 183.0457 | 150.4439 | 173.9794 | 170.8009 | 154.6888 | 170.3067 | 126.8215 | 119.1785 | 156.8879 | 202.9588 | 152.7094 | 153.5812 | 207.2242 | 185.357 | 155.087 | 145.357 | 166.8673 | 155.9336 | 310.7552 | 310.0664 | 176.643 | 237.0709 | 228.1991 | 224.0457 | 278.8467 | 115.913 | 111.1579 | 122.4897 | 139.3776 | 147.8215 | 128.5767 | 127.1121 | 79.0206 | 271.4852 | 227.913 | 233.3112 | 226.0915 | 134.913 | 129.9794 | 134.5103 | 128.6224 | 87.7552 | 64.0457 | 76.4439 | 119.2906 | 212.8879 | 211.643 | 220.4027 | 164.3731 | 135.5561 | 131.5148 | 128.2906 | 118.8673 | 176.8673 | 167.6224 | 218.73 | 155.1327 | 150.8215 | 132.087 | 176.8879 | 125.5103 | 165.0457 | 181.9794 | 202.183 | 215.913 | 178.5561 | 155.4439 | 196.5767 | 179.4485 | 209.4897 | 197.087 | 190.5103 | 278.5354 | 162.6682 | 167.1076 | 151.7345 | 222.1785 | 157.1579 | 147.1785 | 152.0457 | 126.6018 | 163.9794 | 172.6224 | 199.9794 | 137.0915 | 109.5148 | 117.357 | 124.1991 | 126.3112 | 247.4439 | 239.2448 | 223.2951 | 306.8009 | 243.9336 | 234.3821 | 262.5103 | 301.5103 | 129.9543 | 82.8009 | 82.5103 | 129.9543 | 112.8879 | 115.8215 | 115.1121 | 111.3318 | 102.1121 | 94.8215 | 101.3112 | 118.5767 | 73.9336 | 87.3776 | 83.9794 | 79.0664 | 96.7094 | 102.0457 | 100.3318 | 87.1076 | 240.6476 | 237.4852 | 208.0709 | 278.8421 | 172.1121 | 177.1327 | 183.3776 | 264.5354 | 271.3776 | 201.4233 | 225.7094 | 303.1785
1085 | 13.9294 | 142.9115 | 217.706 | 213.6479 | 182.9419 | 149.9115 | 180.1475 | 165.3539 | 154.647 | 144.2055 | 126.8525 | 121.0885 | 157.4406 | 208.1779 | 152.3226 | 171.5005 | 217.4415 | 192.7355 | 141.353 | 159.529 | 182.6461 | 164.1465 | 292.7954 | 303.1475 | 164.7355 | 255.0009 | 221.6765 | 232.353 | 284.6184 | 131.2645 | 127.647 | 123.1475 | 140.2645 | 155.9401 | 128.4406 | 144.0286 | 76.0876 | 246 | 235.5005 | 233.6765 | 239.1475 | 144.765 | 135.5005 | 135.1465 | 121.5889 | 84.0295 | 74.294 | 69.118 | 128.7355 | 205.6765 | 199.412 | 236.7355 | 142.0286 | 136.7355 | 144.7069 | 139.0295 | 123.9705 | 178.0885 | 167.059 | 197.294 | 160.0581 | 143.0885 | 123.2055 | 165.7945 | 123.882 | 178.529 | 190.412 | 220.8544 | 231.588 | 177.471 | 158.7936 | 204.8516 | 200.8535 | 211.4415 | 188.2055 | 181.765 | 296.235 | 171.3825 | 146.5585 | 159.6765 | 220.53 | 166.8535 | 151.353 | 150.3244 | 138.353 | 175.647 | 183.0581 | 203.5594 | 145.2949 | 131.118 | 117.2949 | 123.176 | 125.3825 | 252.7346 | 237.7945 | 254.6774 | 309.4415 | 251.176 | 253.6479 | 258.294 | 283.0599 | 127.1465 | 85.765 | 77.647 | 127.1465 | 109.235 | 107.4415 | 117.471 | 107.47 | 105.4415 | 92.588 | 108.1465 | 102.2654 | 77.294 | 81.1475 | 86.2654 | 80.882 | 89.8525 | 107.4415 | 88.059 | 67.529 | 255.53 | 228.4986 | 223.7364 | 276.2636 | 164.7659 | 189.1751 | 192.3521 | 276.0885 | 286.8221 | 215.7936 | 216.2645 | 318.3521
1086 | 13.9404 | 152.0221 | 216.0074 | 198.0443 | 174.4096 | 145.5904 | 167.6494 | 152.0812 | 157.3432 | 148.2989 | 128.941 | 117.3506 | 160.5756 | 190.1255 | 161.2251 | 160.3358 | 202.6125 | 183.2989 | 135.9705 | 162.6199 | 196.2546 | 179.3506 | 266.3727 | 294.3432 | 165.0369 | 228.3063 | 217.6421 | 228.9852 | 271.7159 | 134.6568 | 118.9262 | 112.3284 | 142.6494 | 175.9483 | 139.6347 | 155.5978 | 86.2841 | 257.7528 | 244.4834 | 230.6494 | 232.3579 | 133.9705 | 128.0295 | 139.6052 | 100.3063 | 80.321 | 78.679 | 64.3579 | 124.3358 | 211.3801 | 198.7011 | 228.6421 | 158.2989 | 136.3653 | 124.3727 | 134.6494 | 131.0369 | 179.3875 | 159.9779 | 206.0443 | 174.6347 | 141.7011 | 130.6568 | 162.7085 | 130.9926 | 183.2989 | 194.0664 | 195.0738 | 240.0221 | 162.6273 | 165.9483 | 225.5756 | 181.6199 | 198.9631 | 196.3284 | 187.4244 | 303 | 172.7011 | 162.059 | 166.0517 | 217.7454 | 152.9926 | 148.6568 | 139.0812 | 144.7233 | 172.2989 | 194.2767 | 181.6199 | 132.3875 | 120.3137 | 107.7454 | 127.9557 | 130.7306 | 274.0221 | 235.3801 | 228.6494 | 311.4022 | 264.0148 | 230.321 | 251.5978 | 272.7897 | 135.2989 | 78.3284 | 83.7011 | 135.2989 | 111.9705 | 102.0148 | 110.0148 | 121.3358 | 103.7085 | 100.6863 | 115.6273 | 99.4465 | 76.9779 | 72.3432 | 68.6716 | 85.9779 | 97.6494 | 95.2989 | 93.0664 | 74.9852 | 246.7011 | 251.2472 | 197.679 | 305.0074 | 144.738 | 210.2398 | 208.9631 | 275.0369 | 302.524 | 214.2177 | 219.321 | 328.2472
1087 | 13.9514 | 155.5634 | 209.9941 | 210.6962 | 183.5132 | 134.0502 | 164.7404 | 167.3864 | 161.0443 | 140.7404 | 113.3481 | 112.9498 | 144.0059 | 215.2979 | 130.2596 | 152.9941 | 209.2332 | 178.8732 | 136.7021 | 144.0826 | 171.8672 | 180.8672 | 289.183 | 291.4306 | 176.0443 | 221.3481 | 228.0561 | 227.0885 | 272.1209 | 138.2655 | 105.6136 | 112.0443 | 145.3098 | 150.1651 | 133.0885 | 131.4306 | 82.3983 | 273.6902 | 275.9498 | 221.823 | 246.6579 | 130.5251 | 133.2153 | 125.7847 | 107.9676 | 80.4808 | 81.3038 | 75.8732 | 126.0443 | 221.9115 | 211.4808 | 226.5251 | 148.6519 | 147.4366 | 130.4749 | 135.2213 | 128.469 | 189.823 | 155.0443 | 206.5132 | 167.3923 | 148.9115 | 135.6579 | 171.9115 | 134.9174 | 173.6519 | 198.4247 | 199.3362 | 236.9498 | 158.6136 | 153.7404 | 209.0059 | 172.1328 | 193.177 | 191.1711 | 206.7787 | 299.1711 | 173.9941 | 171.7345 | 179.6519 | 230.646 | 146.823 | 137.9941 | 150.5575 | 156.8672 | 173.7906 | 177.9557 | 171.4366 | 138.9941 | 119.8732 | 124.4749 | 114.9557 | 141.3864 | 271.9941 | 243.823 | 228.5251 | 328.9557 | 272.177 | 237.7906 | 231.2596 | 299.3864 | 123.2153 | 72.1268 | 84.9941 | 123.2153 | 108.5251 | 108.0885 | 110.5192 | 118.8672 | 110.4749 | 104.6077 | 104.3038 | 124.4306 | 71.3481 | 73.9557 | 70.6962 | 78.9557 | 94.0443 | 90.5251 | 107.1711 | 69.6077 | 263.3098 | 226.9557 | 205.8732 | 305.6077 | 166.0885 | 196.8349 | 208.7464 | 278.7345 | 263.6519 | 193.7906 | 212.5192 | 312.3098
1088 | 13.9624 | 159.649 | 191.0359 | 216.3245 | 166.3245 | 158.9283 | 170.5226 | 160.4056 | 167.3245 | 151.5679 | 117.9641 | 99.6038 | 157.1981 | 209.8113 | 134.9283 | 147.7301 | 234.2246 | 182.5944 | 161.2886 | 136.9641 | 173.5679 | 174.2792 | 304.0717 | 288.9283 | 178.7208 | 216.9547 | 246.6661 | 234.6038 | 254.1264 | 142.4321 | 115.4774 | 122.2886 | 152.8736 | 133.4056 | 137.3604 | 138.298 | 95.3151 | 256.3962 | 265.8471 | 225.1716 | 261.7208 | 135.0359 | 126.1981 | 137.8471 | 132.0359 | 87.3604 | 74.5944 | 75.6302 | 130.5226 | 215.8377 | 205.0265 | 237.883 | 157.7301 | 150.1981 | 125.6038 | 143.5585 | 112.5679 | 175.5132 | 154.8377 | 187.5226 | 164.9547 | 143.1981 | 148.1981 | 169.8019 | 143.4774 | 173 | 173.3604 | 182.6131 | 227.9283 | 173.5226 | 153.7566 | 214.9906 | 183.766 | 203.883 | 183.6396 | 199.5226 | 300.2886 | 159.3604 | 158.1981 | 181.1623 | 210.3151 | 143.3245 | 128.766 | 141.4509 | 152.0811 | 183.1887 | 178.1623 | 172.0359 | 118.234 | 126.117 | 115.6396 | 109.7566 | 131.1623 | 253.0359 | 234.1981 | 240.2434 | 328.4415 | 273.8736 | 255.117 | 239.532 | 282.6755 | 119.8019 | 74.532 | 70.3604 | 119.8019 | 122.4056 | 112.7208 | 105.0811 | 122.3698 | 104.1623 | 101.6396 | 105.5226 | 112.1981 | 72 | 73.0811 | 73.0811 | 80.6038 | 93.8377 | 101.1623 | 98.1981 | 68.4415 | 261.8642 | 215.6302 | 216.0811 | 302.2792 | 175.4056 | 204.585 | 222.6755 | 267.3604 | 267.3245 | 214 | 214.2886 | 319.8736
1089 | 13.9734 | 168.292 | 189.2173 | 219.9007 | 174.7453 | 162.1927 | 170.528 | 177.0807 | 162.702 | 167.236 | 119.7827 | 110.8447 | 152.882 | 224.528 | 152.7267 | 156.8013 | 219.5527 | 168.528 | 167.8013 | 155.18 | 188.8633 | 176.7267 | 295.59 | 304.8633 | 170.6833 | 217.6893 | 225.6087 | 229.82 | 263.882 | 127.5093 | 108.3913 | 129.5467 | 144.6087 | 143 | 140.236 | 154.0933 | 82.9007 | 253.7267 | 267.1927 | 230.8387 | 252.1927 | 130.9813 | 128.2173 | 136.9567 | 117.9193 | 80.9193 | 65.7453 | 68.6893 | 129.7827 | 206.7827 | 187.2173 | 224.3667 | 168.292 | 145.5093 | 137.9627 | 139.882 | 128.6087 | 164.7453 | 156.59 | 195.7267 | 157.4907 | 143.7267 | 141.646 | 170.3913 | 134.1553 | 175.236 | 178.472 | 205.7453 | 238.2733 | 175.3913 | 161.0433 | 213.1613 | 191.0187 | 207.1367 | 190.4533 | 201.3913 | 322.4533 | 163.354 | 163.5713 | 179.41 | 206.472 | 153.9813 | 137.882 | 150.3107 | 147.6647 | 164.7453 | 173.056 | 163.1367 | 112.8447 | 123.9813 | 122.0807 | 103.6273 | 124.938 | 259.7887 | 241.062 | 244.882 | 322.8013 | 253.6833 | 245.1553 | 251.8013 | 280.2173 | 119.646 | 87.9193 | 69.882 | 119.646 | 124.5467 | 114.7453 | 115.5713 | 133.0373 | 106.882 | 100.6273 | 103.292 | 111.236 | 77.59 | 74.6273 | 68.292 | 92.9627 | 94.472 | 105.7453 | 95.3727 | 77.708 | 235 | 203.8447 | 220.236 | 305.472 | 186.8633 | 181.6087 | 212.764 | 278.4347 | 277.9813 | 203.5653 | 217.4473 | 317.1987
1090 | 13.9844 | 154.3053 | 194.1504 | 208.3805 | 175.2301 | 147.4646 | 153.3761 | 175.1415 | 148.1593 | 163.6858 | 119.4691 | 117.8451 | 156.0045 | 216.2301 | 162.0797 | 147.6195 | 224.3716 | 161 | 150.5354 | 164.3761 | 183.146 | 175.3008 | 308.8496 | 307.615 | 160.0045 | 230.9203 | 221.146 | 203.6858 | 260.4602 | 128.4646 | 114.3939 | 120.0797 | 147.0752 | 153.3939 | 139.7655 | 132.4646 | 81.0045 | 259.615 | 254.7744 | 216.239 | 249.7833 | 142.5443 | 133.9203 | 128.4735 | 116.7788 | 75.9292 | 64.6903 | 79.6106 | 119.4602 | 197.615 | 202.5443 | 211.478 | 153.9203 | 150.6992 | 146.0752 | 135.6903 | 122.6061 | 169.0797 | 159.0708 | 196.6106 | 160.7699 | 150.385 | 142.1593 | 171.0045 | 124.0752 | 181.6947 | 178.6858 | 208.5398 | 231.2301 | 168.3053 | 166.4468 | 236.2301 | 191.3894 | 201.6903 | 207.6195 | 196.2301 | 313.4424 | 174.0045 | 177.615 | 174.6195 | 216.6947 | 143.6017 | 138.6947 | 133.615 | 137.385 | 171.0045 | 161.8496 | 173.4735 | 124.0797 | 121.3008 | 124.761 | 106.8496 | 114.9292 | 274.9159 | 250.6858 | 254.1638 | 311.3097 | 245.3142 | 246.6239 | 241.4646 | 295.9292 | 118.6195 | 80.6195 | 68.7699 | 118.6195 | 106.2256 | 107.9159 | 125.3805 | 121.8451 | 108.8496 | 102.3097 | 91.2301 | 121.5443 | 82.3805 | 77.8496 | 70.4735 | 103.385 | 99.3053 | 105.0752 | 107.1638 | 84.3805 | 238.0797 | 213.9248 | 217.4557 | 296.4468 | 178.8362 | 190.6195 | 209 | 286.3761 | 288.0045 | 196.3939 | 182.2167 | 320.9911
1091 | 13.9954 | 159.9312 | 188.3972 | 211.7257 | 162.4798 | 164.9175 | 141.9725 | 155.7807 | 160.3835 | 155.9725 | 125.4523 | 117.1917 | 155.6578 | 213.0138 | 158.2605 | 145.8633 | 203.3835 | 181.6569 | 148.5752 | 150.9862 | 179.3285 | 167.767 | 308.2468 | 292.3018 | 167.2055 | 238.3147 | 213.356 | 205.9037 | 263.9587 | 130.0275 | 125.8358 | 114.2743 | 141.6165 | 157.288 | 137.3697 | 129.6578 | 88.6028 | 274.4927 | 269.5615 | 234.9725 | 271.822 | 143.466 | 129.3972 | 143.2468 | 128.8358 | 83.0275 | 70.1367 | 76.9587 | 115.8083 | 201.3697 | 202.6715 | 233.4523 | 147.8083 | 157.8358 | 144.9862 | 134.3835 | 113.1505 | 177.178 | 160.712 | 182.4385 | 163.9862 | 152.1917 | 142.0688 | 184.5615 | 115.4385 | 169.7128 | 176.534 | 201.8633 | 230.7945 | 165.589 | 150.3285 | 224.6715 | 200.3972 | 205.1505 | 207.4523 | 194.2055 | 291.712 | 176.0413 | 171.8358 | 178.822 | 221.3972 | 120.1917 | 140.2193 | 146.1092 | 141.9725 | 169.0688 | 175.1505 | 190.6303 | 126.2193 | 112.9725 | 117.5477 | 108.6303 | 115.6715 | 264.7807 | 238.6028 | 258.4935 | 326.123 | 257.3972 | 252.466 | 247 | 312.1642 | 127.589 | 76.0825 | 67.6165 | 127.589 | 98.9862 | 106.123 | 119.1917 | 122.3835 | 107.4523 | 102.4248 | 88.0138 | 125.2468 | 69.4385 | 80.822 | 81.2743 | 102.0138 | 94.2055 | 100.411 | 122.2193 | 80.5752 | 244.1917 | 217.3972 | 211.767 | 285.0955 | 161.7945 | 185.6853 | 207.8083 | 271.7945 | 296 | 213 | 181.2322 | 324.0679
1092 | 14.0064 | 166.3997 | 195.9622 | 210.3898 | 156.0576 | 168.6661 | 154.1431 | 157.5428 | 165.2286 | 164.8668 | 120.1431 | 124.7336 | 143.8191 | 209.1809 | 139.2664 | 145.9244 | 199.2187 | 191.7039 | 150.3141 | 149.0855 | 184.352 | 164.7615 | 302.5049 | 284.7434 | 160.676 | 232.8371 | 210.3043 | 232.324 | 273.4095 | 127.6859 | 126.6003 | 114.7911 | 147.0099 | 147.4194 | 153.829 | 114.1332 | 86.3618 | 275.8848 | 270.6283 | 237.3141 | 269.8668 | 135.5526 | 132.8668 | 130.9523 | 122.2286 | 81.0954 | 84.4095 | 79.0378 | 117.6859 | 212.0954 | 185.7714 | 216.676 | 148.8668 | 146.7237 | 146.3618 | 134.7237 | 119.9046 | 175.4474 | 183.6859 | 181.7813 | 169.4572 | 144.171 | 140.9244 | 180.7138 | 109.8668 | 165.0675 | 188.1332 | 196.6003 | 229.5428 | 154.0756 | 157.3997 | 213.5049 | 202.6382 | 218.4572 | 188.6283 | 205.2862 | 321.648 | 177.2385 | 168.5049 | 167.8569 | 224.4572 | 134.2862 | 142.1431 | 148.3898 | 149.2286 | 163.0099 | 184.7714 | 188.5049 | 135.9244 | 116.1332 | 121.6283 | 101.1809 | 102.8668 | 267.7714 | 226.9424 | 231.0378 | 333.171 | 260.8668 | 235.5428 | 251.0954 | 306.8947 | 131.6382 | 64.1431 | 62.3618 | 131.6382 | 102.4095 | 118.4951 | 118.1332 | 114.9424 | 98.4572 | 103.1431 | 93.6003 | 109.676 | 66.7336 | 77.6382 | 61.3141 | 91.6283 | 97.9145 | 95.9523 | 106.5329 | 81.9046 | 252.5526 | 221.2763 | 214.0855 | 309.2286 | 163.8191 | 187.477 | 191.6661 | 251.2944 | 287.3997 | 216.6859 | 221.2763 | 327.6085
1093 | 14.0174 | 160.3275 | 215.2803 | 186.2655 | 163.1416 | 154.3275 | 166.7965 | 146.8289 | 169.1092 | 156.3451 | 128.5782 | 125.8289 | 151.7493 | 207.5782 | 133.1092 | 160.0944 | 178.2979 | 173.2331 | 147.9528 | 138.6254 | 172.7965 | 160.9676 | 301.407 | 300.0944 | 145.8289 | 211.0768 | 208.4514 | 235.7817 | 269.7817 | 124.5634 | 133.5162 | 129.7197 | 149.8761 | 158.8908 | 155.7669 | 123.0768 | 85.2655 | 258.7641 | 245.6577 | 231.5782 | 255.8613 | 142.469 | 127.2979 | 129.6873 | 120.6254 | 89.1092 | 75.2979 | 69.4218 | 117.938 | 219.6873 | 189.9056 | 211.9528 | 155.1092 | 141.5782 | 138.5162 | 135.062 | 118.6401 | 180.3423 | 169.5958 | 198.062 | 165.407 | 135.4838 | 145.3923 | 161.7345 | 126.6549 | 181.7669 | 196.6549 | 204.7817 | 226.8437 | 143.7021 | 144.1563 | 225.4838 | 200.7817 | 213.5634 | 181.3746 | 200.643 | 318.4394 | 185.2507 | 169.938 | 162.3746 | 233.4838 | 141.8761 | 142.1416 | 132.7021 | 148.8908 | 168.8289 | 188.9056 | 188.2507 | 138.7049 | 114.531 | 110.5782 | 102.531 | 108.6873 | 271.062 | 218.593 | 220.1563 | 311.407 | 254.8761 | 237.4838 | 250.2507 | 285.0472 | 131.8908 | 69.6254 | 64.6401 | 131.8908 | 117.7641 | 111.5782 | 113.1563 | 99.4218 | 103.2655 | 105.6725 | 106.4218 | 103.6873 | 81.3275 | 77.0472 | 71.1859 | 87.7493 | 105 | 99.3275 | 93.7965 | 73.8908 | 259.469 | 215.0324 | 219.2331 | 302.1416 | 170.062 | 208.1416 | 182.8113 | 235.6077 | 278.3746 | 216.0944 | 213.7669 | 295.2183
1094 | 14.0284 | 164.6588 | 214.1081 | 201.8919 | 166.1943 | 153.8834 | 172.8682 | 157.4037 | 176.7753 | 137.2635 | 137.5507 | 114.6976 | 161.4341 | 213.0777 | 131.2247 | 144.1858 | 192.7838 | 176.1081 | 155.473 | 139.2095 | 181.0389 | 166.4493 | 281.8834 | 295.473 | 165.0862 | 221.1858 | 222.3176 | 226.0929 | 268.7753 | 121.6824 | 134.2483 | 112.2247 | 150.326 | 164.2483 | 144.6436 | 135.7905 | 92.2095 | 265.1081 | 239.8919 | 252.7061 | 245.7601 | 145.0777 | 119.6436 | 129.3953 | 125.5507 | 83.3176 | 71.1166 | 71.7365 | 116.6436 | 202.8986 | 208.4189 | 221.2095 | 164.5118 | 138.3953 | 136.6436 | 132.8834 | 128.9071 | 184.8057 | 150.8142 | 204.1318 | 153.2635 | 143.8682 | 124 | 167.3801 | 143.5659 | 169.7753 | 190.9915 | 199.4341 | 216.2787 | 163.946 | 141.6976 | 235.6047 | 218.4426 | 198.9611 | 196.4189 | 168.6976 | 293 | 168.6199 | 162.1318 | 161.3564 | 237.0929 | 141.8919 | 130.8682 | 142.527 | 151.6436 | 168.5507 | 188.7517 | 178.5659 | 121.5118 | 121.473 | 116.5118 | 109.0389 | 117.0777 | 278.8682 | 235.5118 | 232.0237 | 306.6436 | 242.7365 | 245.4341 | 243.6047 | 292.2872 | 129.4341 | 69.3412 | 67.0929 | 129.4341 | 122.3716 | 106.6588 | 112 | 107.8142 | 105 | 101.3412 | 100.4882 | 109.9071 | 88.6976 | 81.6824 | 80.2787 | 94.3953 | 94.5811 | 103.6588 | 103.3412 | 77.5118 | 253.3953 | 201.7365 | 223.8446 | 293.9071 | 178.3024 | 192.9611 | 213.326 | 267.853 | 281.2635 | 211.473 | 200.473 | 301.4341
1095 | 14.0394 | 163.4641 | 202.2498 | 221.3392 | 181.5359 | 139.7679 | 169.0895 | 176.411 | 162.9114 | 153.7493 | 141.0717 | 117.4641 | 149.054 | 221 | 129.0177 | 130.4287 | 213.822 | 157.5536 | 160.4464 | 149.5713 | 184.5536 | 163.3747 | 279.8211 | 308.4818 | 170.4119 | 215.9105 | 225.2852 | 216.0895 | 279.9105 | 124.5182 | 125.3578 | 111.357 | 152.0363 | 170.5359 | 148.8751 | 134.8034 | 96.3215 | 237.1789 | 241.9291 | 262.6971 | 267.2321 | 149.4287 | 135.4818 | 130.4641 | 131.7502 | 80.4818 | 69.1612 | 73.5536 | 112.3933 | 187.8211 | 210.8397 | 219.5182 | 175.2321 | 134.5536 | 141.7679 | 120.5536 | 126.411 | 174.3747 | 172.6067 | 194.0717 | 144.3038 | 159.2852 | 123.1072 | 175.2143 | 141.2143 | 173.3215 | 161.8928 | 200.8928 | 205.9105 | 180.589 | 149.3747 | 234.5359 | 217.7511 | 199.9105 | 205.0895 | 174.1426 | 287.1966 | 160.3392 | 159.6608 | 162.4818 | 221.7325 | 147.054 | 138.6962 | 140.6785 | 154.9823 | 171.1789 | 178.4464 | 183.8034 | 130.8928 | 117.5182 | 118.7502 | 105.8574 | 127.6785 | 259.465 | 258.2852 | 233.3755 | 316.2321 | 253.4818 | 245.5536 | 247.8928 | 297.1789 | 138.4818 | 71.4287 | 66.0177 | 138.4818 | 103.3392 | 113.0532 | 108.8751 | 108.1789 | 102.3215 | 99.411 | 108.5173 | 119.0177 | 79.411 | 84.9646 | 65 | 90.1072 | 93.0532 | 92.643 | 101.411 | 81.9823 | 245.5359 | 210.2498 | 209.3392 | 292.7502 | 178.6608 | 189.8928 | 233.7857 | 282.8928 | 299.535 | 203.9468 | 215.2675 | 308.6962
1096 | 14.0504 | 176.3393 | 214.587 | 215.202 | 181.9632 | 148.4676 | 158.8717 | 162.413 | 150.385 | 172.8717 | 124.4956 | 117.9543 | 144.9263 | 206.7802 | 135.3761 | 139.7611 | 186.7433 | 159.2935 | 149.5324 | 151.7065 | 189.9632 | 175.4587 | 315.8526 | 307.0737 | 142.385 | 227.0457 | 224.1931 | 205.8717 | 287.3761 | 122.2567 | 112.4676 | 125.3393 | 132.3761 | 153.5324 | 139.0368 | 138.7891 | 89.7891 | 241.0546 | 242.6417 | 241.8348 | 275.0457 | 141.7891 | 142.3304 | 131.8717 | 128.1283 | 91.4587 | 63.9263 | 69.3304 | 113.1372 | 190.8259 | 204.798 | 218.6328 | 166.0737 | 131.7065 | 135.7891 | 124.587 | 118.3761 | 180.9543 | 180.1194 | 183 | 140.587 | 158.6518 | 136.6785 | 176.6696 | 141.2935 | 177.3393 | 169.8806 | 202 | 210.1652 | 180.3672 | 157.3304 | 223.0368 | 190.0089 | 204.6239 | 199 | 180.7065 | 282.7522 | 159.2478 | 159.8348 | 177.1283 | 215.8806 | 132.8348 | 140.5781 | 139.8717 | 145.8717 | 180.1004 | 178.0826 | 178.615 | 131.0826 | 113.4219 | 116.5044 | 108.9263 | 133.2935 | 248.6785 | 259.9454 | 220.8806 | 314.413 | 263.0826 | 232.615 | 243.4956 | 284.5413 | 133.8628 | 70.2109 | 73.7522 | 133.8628 | 111.8806 | 126.1652 | 99.9543 | 111.1372 | 99.4587 | 90.4587 | 116.615 | 114.8259 | 77.3393 | 82.798 | 61.3304 | 81.202 | 103.413 | 90.587 | 91.0826 | 72.413 | 240.4587 | 219.8348 | 213.7522 | 296.0089 | 175.6239 | 194.2109 | 219.9454 | 277.5781 | 295.6429 | 187.6239 | 220.6607 | 297.2757
1097 | 14.0614 | 163.8052 | 216.232 | 215.4199 | 181.9351 | 150.2251 | 155.413 | 162.6519 | 160.29 | 163.761 | 118 | 113.413 | 159.471 | 200.8329 | 139.826 | 151.29 | 199.1368 | 172.3619 | 141.1809 | 151.355 | 186.1671 | 168.9351 | 318.0953 | 299.5939 | 142.3992 | 232.058 | 214.7169 | 211.8329 | 275.3411 | 134.826 | 125.058 | 121.7541 | 146.7169 | 152.2459 | 134.355 | 133.587 | 84.587 | 260 | 254.7541 | 244.942 | 271.5801 | 146.0069 | 126.6961 | 127 | 129.703 | 87.761 | 78.471 | 65 | 118.2831 | 194.1299 | 203.1671 | 217.7541 | 142.5801 | 133.71 | 135.297 | 132.355 | 114.3481 | 180.1809 | 172.471 | 203.2528 | 135.6381 | 140.297 | 158.239 | 177.232 | 139.761 | 177.5221 | 181.768 | 204.355 | 215.5359 | 166.174 | 153 | 200.9282 | 217.6008 | 192.1671 | 196.645 | 184.5939 | 293.0649 | 167.3039 | 160.116 | 175.4061 | 229.6519 | 146.7749 | 144.3039 | 135.942 | 150.8909 | 188.5152 | 169.9351 | 169.6519 | 122.9351 | 119.587 | 107.4572 | 125.826 | 127.0511 | 274.9489 | 254.6589 | 226.645 | 300.9931 | 265.768 | 218.942 | 239.355 | 297.6589 | 118.355 | 76.826 | 79.355 | 118.355 | 118.116 | 117.4061 | 113.3108 | 127.116 | 98.116 | 103.2459 | 112.3619 | 102.587 | 80.8191 | 84.9351 | 63.5939 | 87.0718 | 106.0069 | 96.471 | 93.297 | 76.8909 | 253.2459 | 221.529 | 209.9351 | 305.232 | 183.4199 | 211.6589 | 187.8121 | 264.3481 | 269.0069 | 196.8329 | 199.2831 | 280.5939
1098 | 14.0724 | 150.149 | 208.1337 | 211.9847 | 176.4164 | 146.6657 | 159.8997 | 169.4833 | 154.5167 | 160.4499 | 126.2159 | 116.4499 | 145.9847 | 219.766 | 145.8997 | 139.234 | 197.4193 | 184.9666 | 142.1671 | 150.617 | 172.0669 | 157.1337 | 285.1337 | 299.7507 | 134.6992 | 220.3677 | 229 | 216.2674 | 273.0487 | 125.9178 | 126.8997 | 117.6657 | 153.7507 | 168.8997 | 137 | 130.5501 | 88.7994 | 259.0334 | 250.6657 | 240.6839 | 259.5501 | 141.9513 | 125.5989 | 130.383 | 129.383 | 78.1671 | 66.9178 | 61.617 | 107.383 | 209.5167 | 206.4652 | 203.0334 | 139.7326 | 142.8663 | 136.1003 | 134.0334 | 103.1671 | 171.5014 | 174.4499 | 228.4164 | 145.2646 | 146.4164 | 155.7507 | 166.234 | 144.6657 | 165.1003 | 186.9666 | 223.4317 | 220.617 | 171.6992 | 168.9484 | 201.3983 | 204.5195 | 190.149 | 200.2827 | 192.9666 | 306.3161 | 172.2674 | 158.9666 | 177.6657 | 226.8176 | 154.2006 | 144.9178 | 149.5654 | 152.8176 | 176.2827 | 173.1156 | 180.8329 | 115.9666 | 118 | 109.8147 | 114.5014 | 120.383 | 279.6839 | 255.9847 | 236.5654 | 303.1156 | 265.6504 | 245.1309 | 249.2493 | 310.5836 | 129.6992 | 84.8329 | 74.7507 | 129.6992 | 116.4833 | 103.7173 | 127.7507 | 119.6839 | 96.9666 | 116.5167 | 114.3343 | 88.9178 | 73.5501 | 68.7841 | 71.4833 | 97.7173 | 109.6839 | 95.0669 | 89.7507 | 78.8176 | 264.1003 | 218.5836 | 216.4986 | 286.0181 | 179.0181 | 220.234 | 174.4499 | 261.383 | 291.532 | 214.3161 | 198.532 | 297.6657
1099 | 14.0834 | 152.5752 | 204.4956 | 199.4779 | 179 | 167.9027 | 168.9469 | 159.0973 | 159.9469 | 160.0177 | 124.5929 | 109.0796 | 145.8673 | 218.5752 | 149 | 151.3275 | 196.7523 | 192.9381 | 150.3806 | 149.9735 | 171.4867 | 165.3894 | 291.4071 | 285.5575 | 142.0177 | 221.3894 | 224.5398 | 207.5044 | 276.0973 | 113.4956 | 130.9912 | 132.4602 | 143.5221 | 165.0619 | 141.9558 | 116.115 | 87.0708 | 253.5398 | 238.699 | 231.0354 | 250.0708 | 140.885 | 133.0442 | 134 | 131.5133 | 84.3983 | 70.354 | 65.4337 | 115.4602 | 204.5398 | 213.0885 | 210.4248 | 146.5133 | 141.5487 | 130.5221 | 138.4513 | 97.0088 | 176.3187 | 173.5221 | 215.1415 | 164.0354 | 142.5575 | 143.0442 | 167.9204 | 141.6194 | 157.5398 | 174.1238 | 223.1592 | 217.4956 | 165.1415 | 172.6194 | 215.5398 | 181.4248 | 205.4602 | 200.5575 | 200.9381 | 313.9823 | 161.5221 | 157.5221 | 178.584 | 220.9646 | 146.0177 | 129.5221 | 155.0708 | 150.9292 | 179.531 | 181.531 | 185.0088 | 125.4248 | 121.9646 | 121.0885 | 111.3894 | 127.469 | 260.6194 | 230.5929 | 249.5044 | 333.3364 | 242.1681 | 254.1592 | 253.0354 | 313.9469 | 134.0442 | 88.0177 | 65.0177 | 134.0442 | 115.0177 | 100.9646 | 120.4956 | 116.4779 | 106.9204 | 110.5487 | 102.0177 | 97.3098 | 73.4867 | 64.4602 | 76.4602 | 106.3629 | 103.5044 | 85.0708 | 88.9381 | 76.4337 | 259.5133 | 217.9823 | 226.5487 | 269.0088 | 159.531 | 215.9735 | 193.3452 | 252.115 | 292.6194 | 220.9912 | 217.9558 | 297.0973
1100 | 14.0944 | 145.0157 | 210.5866 | 201.4921 | 182.0472 | 171.8898 | 185.1574 | 155.618 | 166.5079 | 149.8741 | 123.1416 | 108.6338 | 157.9528 | 218.1259 | 149 | 157.7954 | 215.3505 | 196.9528 | 150.2876 | 139.7954 | 164.8741 | 172.4134 | 303.5236 | 280.0157 | 149.1416 | 223.3348 | 234.7282 | 219.1888 | 277.1888 | 123.1416 | 127.4292 | 126.3348 | 141 | 159.0157 | 153.6023 | 118.236 | 76.4607 | 260.1731 | 214.9843 | 231.0629 | 235.9056 | 145.3662 | 138.6652 | 131.9685 | 131.0157 | 97.0157 | 76.3348 | 78.5866 | 117.9685 | 207.1102 | 224.3305 | 213.9213 | 158.7124 | 150.7282 | 130.0315 | 133.3348 | 100.0629 | 181.2561 | 163.382 | 197.4764 | 162.0315 | 149.2046 | 139.5236 | 169.8898 | 132.0629 | 175.3462 | 168.6338 | 194.3348 | 216.9843 | 162.7124 | 157.9843 | 204.3977 | 193.0472 | 209.4921 | 202.6338 | 182.6066 | 303.3033 | 174.236 | 156.0157 | 167.4764 | 214.3348 | 142.4764 | 132.0787 | 143.4449 | 154.4292 | 164.8269 | 168.8584 | 180.9528 | 139.0787 | 109.7482 | 113.5393 | 116.382 | 127.4449 | 272.3934 | 225.0787 | 239.3505 | 327.6223 | 225.5393 | 234.9843 | 233.764 | 310.3505 | 125.9528 | 89.0472 | 76.2046 | 125.9528 | 117.0629 | 116.681 | 118.9685 | 105.7954 | 111.4292 | 109.5708 | 101.0157 | 106.8112 | 80.0787 | 65.9528 | 77.4449 | 112.3505 | 108.0787 | 91.2203 | 114.7911 | 89.0787 | 260.5393 | 220.5079 | 216.9371 | 297.4564 | 155.4921 | 214.9371 | 205.4134 | 276.5823 | 293.2203 | 214.382 | 224.5236 | 294.6338
1101 | 14.1054 | 150.1613 | 205.5968 | 202.0403 | 168.3549 | 163.9597 | 174.7138 | 176.004 | 167 | 147.2016 | 129.9194 | 118.5605 | 157.6008 | 199.4719 | 146.9194 | 152.2822 | 206 | 186.7178 | 133.3589 | 135.8427 | 172.0846 | 161.2782 | 295.117 | 273.1976 | 150.1976 | 216.6411 | 226.6331 | 217.4759 | 281.7178 | 121.0767 | 123 | 123.2822 | 141 | 165.2016 | 149.0767 | 132.4798 | 90.1249 | 248.113 | 227.5241 | 232.3992 | 233.6411 | 144.2822 | 144.3186 | 127.3992 | 129.3992 | 86.5565 | 74.8427 | 80.879 | 117.0403 | 205.6775 | 217.4316 | 206.9194 | 160.5565 | 151.996 | 126.7984 | 126.121 | 105.0403 | 175.3629 | 155.4798 | 197.5605 | 160.3589 | 148.996 | 132.6775 | 155.7178 | 146.4032 | 196.4798 | 175.4395 | 200.6451 | 223.2822 | 157.7944 | 162.7218 | 209.4435 | 196.5202 | 198.0767 | 196.9557 | 174.1249 | 281.6371 | 180.1573 | 167.4032 | 161.8387 | 206.6008 | 148.2822 | 131.2782 | 138.9597 | 140.6371 | 171.1653 | 165.7218 | 182.6814 | 146.0806 | 98.1613 | 119.121 | 92.8751 | 126.0806 | 279.3146 | 224.2782 | 236.2419 | 308.1613 | 242.5645 | 234.5202 | 239.8064 | 297.879 | 123.5202 | 86.2782 | 85.3589 | 123.5202 | 118.3992 | 128 | 112.3186 | 103.4032 | 109.6008 | 114 | 98.879 | 105.9233 | 87.0806 | 59.3589 | 72.4395 | 90.5162 | 107.2782 | 101.3589 | 102.8308 | 92.9597 | 260.3992 | 223.6008 | 219.7621 | 318.7178 | 169.5645 | 211.5202 | 214.0443 | 289.5928 | 288.2743 | 213.7621 | 223.9194 | 295.1976
1102 | 14.1164 | 161.4543 | 185.8834 | 215.7788 | 170.5709 | 159.2728 | 156.5325 | 167.7669 | 156.8834 | 151.4675 | 132.2596 | 116.8053 | 154.6755 | 186.7139 | 150.3245 | 159 | 203.8702 | 194.9086 | 135.857 | 134.9483 | 185.4675 | 157.5974 | 286.5325 | 287.298 | 145.1298 | 217.8702 | 210.2596 | 219.3762 | 271.8053 | 117.3894 | 113.9483 | 122.0132 | 148.9868 | 167.3377 | 130.4808 | 127.2079 | 94.8834 | 249.8954 | 240 | 220.4159 | 241.2596 | 144.6106 | 137.3377 | 128.7272 | 138.7139 | 83.9868 | 70.2212 | 68.9483 | 119.5974 | 207.0517 | 213.8305 | 233.7523 | 164.3762 | 151.1815 | 143.298 | 134.857 | 119.3113 | 178.143 | 148.0781 | 197.9351 | 160.7272 | 143.3894 | 121.8053 | 168.1682 | 152.8053 | 167.2477 | 167.6106 | 205.8834 | 226.2728 | 162.298 | 161.6106 | 228.5192 | 197.5325 | 196.5192 | 200.506 | 203.3762 | 292.7007 | 165.0781 | 166.8834 | 159.5974 | 205.2728 | 150.2079 | 119.6106 | 156.1033 | 146.3762 | 182.7404 | 180.0517 | 175.8185 | 143.7404 | 112.649 | 107.6238 | 98.2331 | 118.9483 | 268.3245 | 231.2464 | 246.2596 | 309.3377 | 242.6887 | 242.4543 | 242.494 | 309.3762 | 124 | 83.1298 | 78.8053 | 124 | 117.0649 | 115.2212 | 106.4026 | 114.5974 | 104.0132 | 107.6106 | 97.0649 | 127.7139 | 79.9483 | 65.0517 | 74.7272 | 83.7921 | 99.3377 | 98.5325 | 82.9868 | 90.4026 | 244.1563 | 218.5457 | 220.143 | 294.9616 | 180.3377 | 210.4026 | 212.6238 | 243.5073 | 279.5192 | 212.5457 | 222 | 290.1298
1103 | 14.1274 | 159.8289 | 182.9922 | 207.9341 | 171.8367 | 147.2841 | 159.179 | 178.3266 | 146.9105 | 142.8289 | 140.3579 | 117.8132 | 149.4553 | 186.1052 | 149.5526 | 158.4553 | 206.9027 | 189.4788 | 143.7237 | 143.887 | 189.9027 | 163.9027 | 295.1711 | 288.5683 | 156.8053 | 220.9027 | 215.0895 | 229.2763 | 269 | 125.7237 | 110.3579 | 124.8053 | 156.5447 | 157.9184 | 130.6264 | 123 | 89.8132 | 262.9184 | 238.3658 | 220.7159 | 236.2841 | 144.9922 | 137.7237 | 123.2841 | 135.9262 | 90.4553 | 70.4396 | 61 | 115.5526 | 213.3658 | 221.0157 | 241.0235 | 169.3736 | 152.2841 | 142.9341 | 146.5369 | 126.6421 | 177.3579 | 159.9765 | 208.4396 | 161.821 | 143.5526 | 133.1633 | 177.3736 | 135.8367 | 146.3579 | 165.2684 | 193.7316 | 223 | 163.0235 | 152.7316 | 209.3077 | 188.1947 | 214.3501 | 198.113 | 190.9419 | 304.5526 | 168.2606 | 154.7316 | 162.6342 | 207.4474 | 142.7316 | 130.3422 | 144.7629 | 147.0157 | 176.2763 | 182.0078 | 183.9765 | 132.9184 | 118.7316 | 112.9765 | 110.5526 | 122.9843 | 257.7472 | 243.6342 | 242.3736 | 311.3579 | 245.1633 | 242.4631 | 226.9105 | 323.0895 | 131.6264 | 73.0157 | 79.2684 | 131.6264 | 113.6421 | 111.6264 | 112.6264 | 119.8132 | 105.7159 | 98.7316 | 103.9922 | 118.3893 | 73.6342 | 74.0895 | 79.0895 | 95.0816 | 94.821 | 94.0973 | 102.5291 | 81.3736 | 255.9686 | 213.6342 | 205.1947 | 292.5291 | 183.4474 | 190.4788 | 210.8948 | 242.1554 | 288.0895 | 191.292 | 213.8289 | 290.3658
1104 | 14.1384 | 161.9125 | 185.2148 | 205.9258 | 176.1539 | 152.8117 | 158.7719 | 183.6179 | 158.2547 | 152.711 | 140.1008 | 118.7719 | 157.9125 | 200.9524 | 147.2281 | 152.9867 | 219.3555 | 176.7852 | 146 | 142.4031 | 203.4696 | 160.7586 | 309.2414 | 285.6977 | 153.8594 | 221.6578 | 213.7719 | 214.1883 | 257.8594 | 125.2148 | 122.9125 | 116.289 | 138.0609 | 144.7586 | 152.5969 | 135.2547 | 84.0875 | 268.6977 | 254.825 | 225.7719 | 222.8726 | 134.9601 | 127.1883 | 133.825 | 126.6711 | 88.3289 | 74.4297 | 77.711 | 118.7985 | 215.3422 | 211.557 | 219.3156 | 160.7719 | 141.1008 | 127.557 | 153.1141 | 119.6578 | 186.5703 | 186.1406 | 209.0875 | 158.3289 | 143.4563 | 142.2148 | 176.5703 | 119.5437 | 170.6102 | 179.6977 | 202.6977 | 222.443 | 154.1274 | 152.7852 | 202.5969 | 214.536 | 217.4297 | 191.7985 | 176.1274 | 282.732 | 168.2015 | 158.6844 | 154.5304 | 215.3422 | 123.8461 | 133.9734 | 126.4563 | 152.0399 | 174.557 | 172.5437 | 186.1883 | 119.2015 | 113.2148 | 118.5304 | 121.0399 | 138.5703 | 266.1672 | 235.5304 | 233.7719 | 335.6102 | 267.0266 | 239.2281 | 243.268 | 338.4828 | 123.5172 | 71.3555 | 73.6445 | 123.5172 | 103.3156 | 118.557 | 124.0133 | 122.443 | 100.1883 | 104.9125 | 106.7719 | 106.4563 | 80.0133 | 76.6711 | 80 | 90.4164 | 90.7719 | 92.7852 | 107.9867 | 79.4563 | 255.9468 | 213.8859 | 207.5836 | 290.1883 | 168.5039 | 182.7985 | 202.175 | 272.9258 | 278.9734 | 181.7852 | 200.8726 | 277.8594
1105 | 14.1494 | 160.4601 | 185.8466 | 215.292 | 178.1829 | 165.2773 | 167.8171 | 166.1534 | 150.9203 | 159.1681 | 123.9056 | 123.8319 | 150.1976 | 222.8466 | 151.2773 | 152.9853 | 202.0884 | 178.4307 | 159.0944 | 144.8024 | 204.7374 | 157.8466 | 305.3215 | 297.2773 | 181.4367 | 221.8466 | 215.9853 | 211.9705 | 272.3423 | 123.5693 | 128.292 | 111.5399 | 133.2478 | 138.4307 | 146.2123 | 150.1239 | 87.2478 | 283.1239 | 250.2123 | 237.6638 | 232.8024 | 141.7877 | 118.1386 | 147.4307 | 130.2773 | 79.0295 | 68.292 | 75.0589 | 119.8761 | 210.5988 | 221.1092 | 226.5251 | 162.9853 | 149.3865 | 137.1092 | 147.7374 | 107.3215 | 180.7522 | 172.7964 | 192.3215 | 147.8908 | 147.5693 | 133.1681 | 162.2123 | 119.9853 | 179.0295 | 174.1976 | 205.1681 | 213.4601 | 157.292 | 158.9853 | 213.8614 | 228.3362 | 209.0147 | 194.0147 | 189.5399 | 282.0944 | 172.8171 | 154.8908 | 142.4454 | 198.643 | 124.0944 | 130.5546 | 143.6638 | 151.4749 | 175.5693 | 162.7374 | 186.8171 | 135.2036 | 102.4601 | 112.708 | 121.0442 | 143 | 282.4307 | 222.3068 | 234.8466 | 319.5486 | 281.8319 | 232.4601 | 229.1031 | 301.0382 | 121.6785 | 83.6932 | 72.6932 | 121.6785 | 111.0944 | 115.5841 | 122.8761 | 123.708 | 102.5251 | 104.5988 | 104.4307 | 104.8761 | 82.8614 | 80.8466 | 78.8614 | 83.1386 | 94.1239 | 95.5693 | 105.1386 | 85.2773 | 255.3717 | 228.3717 | 231.9411 | 274.8761 | 162.6785 | 182.7374 | 191.9853 | 265.2123 | 280.1092 | 186.8466 | 201.1239 | 274.1239
1106 | 14.1604 | 151.0919 | 187.4184 | 210.5103 | 179.8874 | 147.2251 | 159.7148 | 171.561 | 148.4691 | 141.2045 | 111.6735 | 119.6942 | 144.8161 | 223.8368 | 166.9588 | 150.1839 | 206.1014 | 191.9588 | 159.6942 | 146.1126 | 184.2964 | 154.7655 | 300.3265 | 282.7148 | 178.1744 | 221.6735 | 220.1632 | 218.5816 | 287.0919 | 133.3058 | 127 | 116.2552 | 145.0713 | 149.0507 | 141.8874 | 148.1839 | 94.4184 | 249.1951 | 221.4596 | 227.0619 | 235.2758 | 153.0919 | 131.2139 | 133.6229 | 112.2251 | 73.5816 | 77.4691 | 75.7955 | 118.3265 | 218.9588 | 216.9493 | 233.0919 | 159.6022 | 142.878 | 145.7448 | 135.439 | 111.0507 | 173.5816 | 162.3978 | 198.3771 | 147.3978 | 143.3471 | 126.8368 | 147.4184 | 131.1426 | 174.1632 | 158.9287 | 218.6117 | 211.0713 | 154.8368 | 163.1632 | 193.8067 | 206.9493 | 209.4897 | 178.7861 | 184.9493 | 276.2964 | 168.7861 | 144.5103 | 155.2852 | 212.4991 | 140.3978 | 156.1014 | 145.2758 | 150.1426 | 154.4803 | 162.6529 | 190.9287 | 131.0619 | 114.6117 | 114 | 109.6735 | 130.7861 | 258.1538 | 227.3058 | 239.3265 | 290.3471 | 277.1126 | 221.9287 | 216.7242 | 288.4278 | 133.6529 | 76.3677 | 71.7655 | 133.6529 | 119.8368 | 114.1632 | 124.8161 | 120.9287 | 101.5309 | 111.2139 | 102.8368 | 111.4691 | 69.7861 | 84.1632 | 80.3265 | 92.1426 | 93.9287 | 97.7448 | 97.8574 | 77.6942 | 253.0412 | 216.7354 | 237.0206 | 273.9081 | 171.7448 | 195.4484 | 188.6022 | 263.7955 | 278.2758 | 206.4484 | 195.6942 | 296.0301
1107 | 14.1714 | 142.467 | 188.1878 | 215.7208 | 184.6244 | 139.533 | 154.533 | 171.0609 | 152.4366 | 137.3452 | 113.5634 | 117.1573 | 153.1573 | 214.0914 | 156.8072 | 143.6244 | 212.4975 | 197.2488 | 149.4366 | 131.2792 | 193.7867 | 163.472 | 299.6244 | 281.6903 | 161.9391 | 244.3501 | 215.6548 | 240.9745 | 270.7463 | 138.2183 | 128.7817 | 107.2792 | 142.0609 | 150.467 | 144.2488 | 144.5939 | 106.472 | 248.7258 | 227.0355 | 217.9391 | 240.8781 | 137.3402 | 150.6903 | 131.7208 | 105.7208 | 81.1269 | 74.9036 | 87.3756 | 114.6548 | 205.8377 | 184.056 | 209.6194 | 154.4061 | 141.5025 | 132.1524 | 135.9391 | 114.8427 | 172.8122 | 169.9695 | 213.3452 | 162.6903 | 149.5025 | 140.8476 | 149.9695 | 138.7817 | 175 | 177.3806 | 207.6499 | 227.0659 | 149.2488 | 177.6598 | 206.7258 | 193.6548 | 206.0609 | 181.8781 | 191.2537 | 279.2537 | 170.6903 | 164.5684 | 159.8731 | 214.5889 | 134.9036 | 148.2438 | 141.9695 | 145.3097 | 162.1623 | 171.3452 | 183.8427 | 122.533 | 120.8731 | 121.7208 | 113.3452 | 125.5634 | 257.0355 | 226.2792 | 224.3706 | 298.8781 | 267.6244 | 237.4111 | 244.9745 | 304.8122 | 137.5939 | 72.7512 | 82.2537 | 137.5939 | 110.0914 | 101.9341 | 122.467 | 120.9695 | 100.533 | 108.1219 | 104.3756 | 104.1524 | 67.533 | 83.2183 | 73.0914 | 100.3756 | 102.2842 | 98.4061 | 101.5025 | 77.533 | 235.8731 | 211.2842 | 227.8427 | 303.9136 | 168.8427 | 189.5889 | 196.472 | 265.2792 | 283.8781 | 210.0914 | 197.3147 | 298.3097
1108 | 14.1824 | 142.2434 | 190.8186 | 224.031 | 167.2389 | 154.9115 | 162.031 | 163.969 | 144.5442 | 139.1814 | 120.2434 | 106.6637 | 156.6062 | 208.6062 | 135.7256 | 139.5752 | 215.093 | 188.5442 | 148.8186 | 133.8806 | 184.9645 | 178.062 | 282.2389 | 280.5132 | 159.3318 | 241.6017 | 215.6372 | 242.6637 | 278.5797 | 116.9955 | 121.5132 | 108.062 | 140.4248 | 146 | 141 | 151.062 | 108.938 | 251.2389 | 233.7566 | 235.3363 | 238.0885 | 126.1814 | 144.6637 | 137.6062 | 109.1814 | 90.2434 | 74.6682 | 80.5132 | 98.8761 | 196.6682 | 205.7966 | 211.3673 | 172.1859 | 157.093 | 139.5797 | 136.3628 | 108.3628 | 183.5177 | 163.5132 | 196.3893 | 171.2124 | 146.907 | 144.938 | 170.7921 | 141.8186 | 171.969 | 194.4248 | 204.9425 | 229.938 | 167.2169 | 162.7521 | 192.8716 | 203.9425 | 208.062 | 213.6417 | 204.6372 | 282.3318 | 178.6062 | 155.1459 | 149.5442 | 213.5177 | 132.8496 | 142.3053 | 142.1814 | 139.8186 | 171.938 | 180.4558 | 176.7566 | 124.5752 | 117.2124 | 119.7256 | 113.3628 | 116.4823 | 255.8761 | 226.4558 | 229.9735 | 302.7566 | 267.8186 | 232.4203 | 246.6637 | 310.6682 | 130.1194 | 77.2124 | 85.3318 | 130.1194 | 116.7301 | 107.5487 | 99.2079 | 119.969 | 98.938 | 108.4868 | 103.5752 | 98.2434 | 75.6372 | 84.4248 | 73.6682 | 103.2124 | 107.5752 | 107.093 | 91.6327 | 87.4558 | 240.6991 | 207.4823 | 211.6637 | 321.7876 | 147.2079 | 178.2124 | 210.4558 | 272.7301 | 306.5487 | 194.3009 | 202.3938 | 297.062
1109 | 14.1934 | 161.6992 | 191.3815 | 220.4336 | 155.763 | 156.4336 | 166.474 | 155.8151 | 141 | 147.8958 | 110.0118 | 110.9882 | 150.1966 | 210.237 | 142.1328 | 138 | 217.2891 | 189.9479 | 154.9479 | 142.0924 | 168.1445 | 172.7227 | 288.6992 | 281.8034 | 146.9597 | 227.763 | 227.8958 | 225.3412 | 272.0639 | 116.9882 | 128.9882 | 110.1445 | 149.4219 | 160.8437 | 134.1966 | 143.2487 | 92.0118 | 261.4101 | 239.6588 | 241.526 | 224.8151 | 134.8958 | 138.474 | 143.5664 | 114.1849 | 89.2891 | 63.5378 | 74.3815 | 107.0807 | 197.9076 | 209.3529 | 212.7227 | 170.3933 | 161.1445 | 127.4975 | 136.474 | 110.9479 | 174.9193 | 154.2891 | 181.763 | 163.3412 | 143.474 | 145.3294 | 164.4454 | 137.4336 | 164.4336 | 199.7109 | 202.2487 | 227.237 | 167.3933 | 144.526 | 218.3866 | 221.0403 | 197.1563 | 206.7344 | 210.0924 | 303.9765 | 174.0521 | 155.6992 | 155.2773 | 222.237 | 133.526 | 141.5781 | 130.4857 | 144.0924 | 169.8555 | 176.5781 | 186.3698 | 114.3412 | 113.0521 | 122.4219 | 102.3412 | 133.121 | 274.595 | 225.0521 | 233.5781 | 299.3815 | 267.1445 | 263.6757 | 254.0807 | 306.9597 | 127.474 | 74.2891 | 76.6706 | 127.474 | 129.3294 | 109.5781 | 105.5546 | 121.7109 | 114.7916 | 109.0521 | 94.5781 | 105.3294 | 72.4336 | 78.5781 | 73.0521 | 87.3008 | 90.5378 | 88.2605 | 89.0403 | 81.1042 | 234.6302 | 217.3177 | 200.526 | 293.7865 | 155.4101 | 187.0403 | 204.7227 | 265.5378 | 278.8907 | 189.237 | 203.8555 | 287.3933
1110 | 14.2044 | 170.7384 | 199.2001 | 217.6308 | 151.8461 | 149.8461 | 161.6922 | 158.3078 | 147.9386 | 147.6922 | 121.5541 | 113.9538 | 146.6308 | 221.7232 | 143.323 | 150.6156 | 208.6922 | 199.3078 | 152.9538 | 147.1539 | 168.2616 | 158.7999 | 295.2153 | 278.4307 | 146.4155 | 227.6308 | 226.4307 | 218.7384 | 264.1539 | 118.0614 | 120.5998 | 117.8309 | 148.323 | 155.4921 | 156.4927 | 135.3692 | 88.4155 | 251.292 | 233.0152 | 250.7232 | 243.708 | 145.4155 | 140.6308 | 138.7999 | 126.8309 | 76.9075 | 66.6156 | 65.1691 | 110.0614 | 205.4617 | 212.4002 | 219.6156 | 154.4307 | 153.6922 | 142.0158 | 136.1077 | 132.9234 | 170.0462 | 156.4155 | 193.6156 | 173.1387 | 126.0766 | 148 | 164.354 | 133.3692 | 170.4617 | 195.6922 | 200.0462 | 229.2616 | 147.646 | 148.0462 | 226.1843 | 217.1691 | 196.2001 | 199.3078 | 210.7384 | 294.1381 | 167.8461 | 162.2153 | 165.4155 | 233.0925 | 142.7232 | 136.3692 | 132.2001 | 152.9386 | 177.9386 | 175.7847 | 196.5231 | 134.8619 | 118.2001 | 111.2305 | 104.5693 | 127.1843 | 282.7999 | 237.7695 | 225.8461 | 295.2153 | 269.1539 | 255.9379 | 249.4921 | 305.1539 | 115.1229 | 77.0462 | 72.7384 | 115.1229 | 121.9075 | 108.1539 | 113.8461 | 109.4921 | 115.0152 | 102.8461 | 101.354 | 108.6308 | 72.7847 | 64.5383 | 80.7232 | 90.8771 | 96.7695 | 81.2001 | 83.9075 | 76.8923 | 227.6308 | 194.5687 | 214.1387 | 292.7695 | 173.0462 | 198.3078 | 201.5231 | 273.6618 | 279.9696 | 210.8157 | 206.8923 | 313.6928
1111 | 14.2154 | 159.0678 | 200.1416 | 204.4955 | 169.2921 | 169.8644 | 149.6401 | 175.5045 | 155.8584 | 140.1416 | 129.7847 | 109.7139 | 155.3599 | 213.8525 | 154.7906 | 152.8554 | 221.7198 | 181.1356 | 161.5753 | 154.7876 | 190.2213 | 162.3599 | 306.5045 | 277.8584 | 172.1505 | 234.4307 | 223.9292 | 200.6371 | 281.4337 | 133.2921 | 119.3599 | 113.997 | 146.2861 | 154.7168 | 142.4188 | 149.7906 | 86.4985 | 253.4337 | 238.8614 | 256.3569 | 249.2832 | 145.4277 | 136.4985 | 135.2861 | 147.4337 | 87.7198 | 77.8584 | 75.4337 | 103.4277 | 216.7876 | 222 | 214.1386 | 155.7876 | 148.0708 | 132.6341 | 136.2861 | 138.2124 | 153.7079 | 153.8554 | 197.1416 | 170.4247 | 145.2243 | 139.6401 | 174.8584 | 132.3569 | 176.6431 | 188.1416 | 206.8584 | 230 | 148.7168 | 162.5753 | 209.4985 | 203.6401 | 192.6401 | 192.0678 | 210.6431 | 277.0708 | 174.3599 | 181.2213 | 170.5723 | 212.6341 | 143.2124 | 141.7876 | 142.7876 | 175.006 | 162.7079 | 181.8584 | 179.994 | 137.1386 | 109.4955 | 104.5723 | 113.5015 | 120.1446 | 256.1356 | 225.1356 | 214.997 | 308.4337 | 267.7847 | 230.2124 | 229.4247 | 292.8525 | 131.4367 | 76.1416 | 83.5753 | 131.4367 | 106.354 | 109.3569 | 99.7817 | 113.8614 | 116.3599 | 104.2153 | 104.7847 | 108.3569 | 74.3569 | 72.5045 | 74.7817 | 102.8584 | 104.7139 | 97.5753 | 90.2183 | 70.2832 | 241.5045 | 194.9352 | 208.8525 | 294.2832 | 170.8554 | 181.4218 | 179.8495 | 276.2832 | 271.4926 | 212.0678 | 208.6431 | 298.6312
1112 | 14.2264 | 161.5197 | 190.791 | 179.9606 | 151.8195 | 159.0624 | 158.1072 | 165.9606 | 141.616 | 149.1411 | 128 | 109 | 146.8928 | 197.4803 | 151.8589 | 133.616 | 214.616 | 171.6214 | 145.7177 | 155.3786 | 190.2035 | 147.3392 | 294.3392 | 296.384 | 178.4464 | 218.3392 | 219.1018 | 223.768 | 268.3731 | 121.7177 | 136.4519 | 104.4125 | 138.4803 | 176.6947 | 135.1751 | 137.6838 | 87.9322 | 265.9322 | 219.7856 | 250.7571 | 235.8249 | 151.8643 | 133.3446 | 137.3107 | 128.4748 | 89.1357 | 85.8982 | 74.1696 | 96.7571 | 218.6893 | 214.1357 | 232.5591 | 153.7571 | 147 | 124.5536 | 132.4125 | 128.4464 | 160.0394 | 173.2823 | 206.7965 | 175.1411 | 166.5875 | 134.3446 | 169.791 | 124.1357 | 172.4125 | 190.5875 | 206.3786 | 217.5481 | 145.1357 | 153.9267 | 231.9037 | 191.1357 | 189.9661 | 190.5875 | 190.0285 | 278.6214 | 171.1357 | 166.7856 | 169.3786 | 220.9376 | 130.8249 | 140.4125 | 150.5875 | 162.0624 | 169.0394 | 175.4803 | 173.9322 | 122.1357 | 113.1411 | 95.5142 | 126.4858 | 135.4519 | 260.384 | 222.8304 | 245.3895 | 319.6214 | 259.4464 | 236.8304 | 245.9376 | 291.5536 | 140.4125 | 64.8249 | 84.7571 | 140.4125 | 105.5875 | 110.3107 | 103.4858 | 128.8643 | 123.6214 | 96.1696 | 112.1751 | 117.8304 | 82.5197 | 75.4125 | 70.6214 | 112.8643 | 102.0339 | 84.9945 | 85.2035 | 75.1751 | 249.6554 | 212.5536 | 218.0394 | 300.4858 | 177.8304 | 200.1466 | 180.7626 | 256.9267 | 274.0733 | 207.3107 | 193.2713 | 298.0055
1113 | 14.2374 | 162.6617 | 184.3294 | 189.0266 | 151.6883 | 155.6795 | 166.3353 | 165.6795 | 147.6883 | 152.3294 | 121.9911 | 123.0207 | 136.6617 | 197.6736 | 147.3353 | 154.3767 | 215.3471 | 175.6706 | 142.0118 | 140.647 | 192.6795 | 155.6943 | 294.0148 | 282.9675 | 156.3057 | 241.3826 | 213.997 | 224.3087 | 246.3176 | 118.6795 | 122.3028 | 107.3412 | 154.6972 | 167.3057 | 123.9764 | 143.6913 | 86.6617 | 282.0207 | 246.0591 | 231.9764 | 247.6913 | 153.3294 | 137.0059 | 132.6588 | 128.0207 | 90.3412 | 86.997 | 76.3442 | 107.353 | 221.3383 | 204.6588 | 239.9911 | 146.9941 | 151.0059 | 123.9941 | 124.6588 | 127.6706 | 175.0089 | 178.9911 | 201.6499 | 161.6381 | 139.6233 | 122.6499 | 153.9823 | 137.359 | 169.3323 | 186.3235 | 202.997 | 207.6617 | 173.0473 | 151.3412 | 225.6381 | 207.6972 | 191.6677 | 197.6736 | 187.6795 | 286.6765 | 156.3176 | 161.0118 | 164.6617 | 213.6381 | 148.7002 | 142.0059 | 154.3353 | 155.3412 | 161.9764 | 177.0089 | 188.6854 | 132.6883 | 127.6795 | 100.0148 | 129.3294 | 136.6588 | 283.6883 | 230.003 | 249.3117 | 304.3087 | 256.6677 | 234.6558 | 219.9438 | 291.6617 | 136.6647 | 73.353 | 71.3176 | 136.6647 | 102.6588 | 112.3353 | 117.6795 | 130.997 | 114.9852 | 84.9911 | 105.6499 | 112.9852 | 74.3146 | 69.6617 | 86.6883 | 97.6381 | 103.6706 | 83.6795 | 91.6854 | 84.6736 | 233.3087 | 225.3471 | 196.9557 | 323.359 | 177.6588 | 181.2821 | 193.3412 | 248.3323 | 271.9852 | 214.0089 | 181.6617 | 318.0148
1114 | 14.2484 | 150.1209 | 191.1593 | 198 | 170.559 | 156.6003 | 167.6799 | 165.2404 | 162.4794 | 157.1195 | 116.9602 | 104.1622 | 146.559 | 221.7582 | 135.761 | 162.8805 | 206.4011 | 194.6785 | 143.2802 | 143.5192 | 192.2404 | 170.4395 | 313.2788 | 272 | 166.0384 | 225.3627 | 209.6003 | 213.9602 | 269.5576 | 137.9587 | 114.0398 | 118.1593 | 155.4808 | 135.5619 | 141.1579 | 137.0413 | 97.9189 | 252.963 | 274.1593 | 226.0398 | 243.761 | 149.2802 | 128.1209 | 134.7596 | 136.3599 | 83.4808 | 77.1607 | 67.761 | 106.5206 | 208.0413 | 203.3599 | 249.9189 | 156.559 | 134.6415 | 129.4794 | 134.9189 | 105.8819 | 174.6003 | 170.5605 | 190.5605 | 168.9986 | 128.3997 | 138.7582 | 158.1991 | 150.7596 | 191.4381 | 197.9587 | 218.9986 | 227.0782 | 176.761 | 145.8407 | 204.441 | 217.3201 | 182.4808 | 188.441 | 198.7994 | 298.8393 | 147.6003 | 164.3201 | 160.9602 | 210.7994 | 140.2817 | 136.5206 | 149.5605 | 145.0811 | 177.798 | 162.3215 | 175.2817 | 137.9602 | 120.441 | 100.9203 | 120.5206 | 124.4808 | 249.5235 | 247.9986 | 239.2802 | 312.3186 | 260.3997 | 236.4395 | 224.798 | 291.3599 | 146.1991 | 68.441 | 80.9587 | 146.1991 | 111.559 | 114.3599 | 105.6814 | 120.4808 | 101.8407 | 94.9189 | 95.2404 | 114.7994 | 76.1593 | 76.1593 | 89.2404 | 91.3997 | 107.0398 | 96.8393 | 91.2006 | 77.4808 | 228.3997 | 239.5192 | 207.8378 | 310.2418 | 153.2418 | 210.2361 | 204.8393 | 237.8009 | 290.1181 | 217.6799 | 186.7994 | 314.8407
1115 | 14.2594 | 153.999 | 206.768 | 218.767 | 171.8466 | 170.2291 | 160.3854 | 163.6922 | 159.0777 | 142.0019 | 117.3845 | 87.1544 | 154.0767 | 213.3097 | 137.6146 | 155.1544 | 220.0748 | 177.3873 | 128.8476 | 154.2301 | 177.5398 | 152.233 | 307.5398 | 280.999 | 148.6951 | 241.8428 | 203.8466 | 233.767 | 272.6165 | 111.7728 | 107.3854 | 110.9243 | 144.7699 | 149.2282 | 141.232 | 148.6903 | 101.2311 | 260.2282 | 233.0048 | 253.9971 | 259.4593 | 144.5388 | 123 | 141.8456 | 125.9243 | 81.0767 | 66.7699 | 75.1524 | 112.6913 | 201 | 200.5388 | 221.3883 | 151.6165 | 132.2301 | 130.2311 | 136.8466 | 108.8447 | 164.001 | 176.3068 | 196.999 | 172.8466 | 147.9981 | 146.9233 | 183.0748 | 141.232 | 196.4621 | 181.464 | 214.5398 | 230.7699 | 164.0777 | 171.7661 | 200.3845 | 202.4631 | 200.1515 | 174.6932 | 192.3087 | 311.999 | 152.2301 | 163.3078 | 167.6146 | 198.7709 | 135.1534 | 141.3068 | 153.2301 | 134.1544 | 194.5379 | 169.9214 | 174.999 | 134.9233 | 110.8466 | 108.6913 | 120.4612 | 116.5388 | 236.9223 | 226.9262 | 234.5388 | 311.001 | 262 | 230.001 | 240.8456 | 301.6913 | 155.8456 | 74.0757 | 79.6932 | 155.8456 | 108.6932 | 105.3087 | 110.4602 | 109.7699 | 89.001 | 96.1544 | 93.6922 | 109.001 | 79.3078 | 73.0777 | 75.9243 | 90.9233 | 97.6165 | 89.9243 | 70.6941 | 83.3835 | 238.3068 | 220.464 | 222.0767 | 291.6932 | 162.3826 | 208.464 | 197.232 | 258.6127 | 294.7699 | 208.3087 | 220.4583 | 299.232
1116 | 14.2704 | 139.6823 | 204.2502 | 209.6823 | 165.7728 | 162.9095 | 158.4091 | 145.6823 | 152.4774 | 146.6815 | 119.4091 | 114.5901 | 155.7045 | 206.4091 | 138.1819 | 148.0683 | 206.4552 | 173.7498 | 145.5448 | 156.2955 | 186.795 | 165.5448 | 309.7498 | 291.3407 | 176.4534 | 207.3878 | 204.8181 | 241.5909 | 265.1819 | 125.181 | 133.5901 | 118.6815 | 134.9547 | 160 | 137.4091 | 142.9095 | 90.1366 | 270.2955 | 231.3177 | 247.6823 | 245.7507 | 138.7728 | 133.5679 | 127.7959 | 144.9539 | 88.3407 | 74.5679 | 72.5457 | 111.3638 | 198.8864 | 206.7498 | 232.772 | 159.6815 | 128.6593 | 133.9317 | 144.8634 | 115 | 167.0453 | 174.3638 | 212.9769 | 170.2955 | 145.4321 | 153.7498 | 170.1597 | 137.4091 | 182.7276 | 192.8403 | 190.6823 | 230.1136 | 145.5004 | 163.1597 | 213.6815 | 203.0453 | 200.1366 | 175.2272 | 176.0231 | 299.7959 | 169.0905 | 163 | 151.2733 | 215.9539 | 137 | 152.0453 | 151.7728 | 139.0453 | 160.3647 | 177.7045 | 186.7498 | 124.8412 | 106.8864 | 123.5679 | 118.4774 | 122.0453 | 254.795 | 229.4996 | 220.3185 | 316.1588 | 254.2502 | 219.6593 | 247.9317 | 301.0683 | 146.7276 | 76.1819 | 73.1819 | 146.7276 | 106.4091 | 109.4543 | 124.4543 | 121.0905 | 104.0222 | 96.1136 | 108.795 | 109.2272 | 67.7276 | 74.9317 | 68.8864 | 88.5909 | 99.3407 | 83.5909 | 92.5901 | 78.8412 | 240.5909 | 220.5679 | 221.5909 | 274.614 | 167.4774 | 188.8412 | 189.8864 | 254.5004 | 297.6362 | 212.4543 | 212.8642 | 272.1597
1117 | 14.2814 | 154.2213 | 188.8141 | 207.7345 | 161.1327 | 166.3186 | 169.0354 | 150.9027 | 168.2036 | 154.1504 | 104.9469 | 117.6814 | 154.5664 | 224.9204 | 152.77 | 157.4691 | 210.6195 | 194.9204 | 155 | 146.6814 | 186.5487 | 176.4336 | 310.8496 | 278.23 | 169.3451 | 222.124 | 207.4336 | 224.5132 | 287.6549 | 133.4159 | 125.9291 | 128.3009 | 130.5664 | 134.1946 | 138 | 144.885 | 87.4336 | 266.4159 | 231.8318 | 245.7345 | 243.1682 | 134.1327 | 131.5487 | 121 | 135.646 | 93.8673 | 68.2478 | 70.4336 | 118.3186 | 186.5309 | 203.5487 | 228.23 | 159.2655 | 138.9027 | 139.5841 | 136.8141 | 103.5309 | 182.9027 | 172 | 197.2123 | 171.4336 | 141.7168 | 140.5132 | 159.5664 | 125.8141 | 185.885 | 192.6814 | 180.8496 | 226.6991 | 164.8054 | 169.77 | 226.885 | 210.3009 | 206.0354 | 176.2832 | 182.4691 | 280.0973 | 167.115 | 173.7522 | 158.0531 | 227.4336 | 135.5664 | 147.115 | 136.3805 | 143.4336 | 167.9381 | 181.5841 | 182.115 | 100.9291 | 96.6814 | 117.2478 | 119.8673 | 125 | 249.5309 | 225.9646 | 204.2478 | 330.7522 | 242.3982 | 248.5399 | 229.2123 | 311.1859 | 136.2655 | 71.4159 | 68.4159 | 136.2655 | 93.3805 | 111.5664 | 116.5309 | 126.2832 | 113.4336 | 95.5664 | 113.5664 | 108.1327 | 65.8673 | 71.9823 | 70.1504 | 106.6372 | 94.8318 | 83.7168 | 94.9646 | 87.9027 | 254.3363 | 214.2478 | 208.0973 | 300.5399 | 166.7168 | 190.7345 | 181.8318 | 265.2036 | 304.3009 | 201.6637 | 198.9823 | 287.3718
1118 | 14.2924 | 157.8962 | 182.5418 | 208.5418 | 173.1239 | 164.8962 | 174.4582 | 158.1873 | 152.3977 | 160.8329 | 115.0404 | 114.7291 | 154.7291 | 214.5014 | 140.0432 | 153.2507 | 208.7089 | 187.4179 | 154.2709 | 134.9798 | 192.0202 | 160.2305 | 304.8962 | 279.2911 | 153.438 | 237.4582 | 190.5014 | 220.1873 | 301.3746 | 118.147 | 112.8962 | 106.6686 | 136.562 | 140.7695 | 135.0836 | 143.6254 | 99.6657 | 265 | 217.3343 | 212.2737 | 267.8732 | 163.6225 | 115.147 | 123.9164 | 111.2305 | 86.9798 | 80.7695 | 79.0202 | 112.5216 | 172.5216 | 195.8962 | 237.3113 | 146.7925 | 146.9164 | 138.0836 | 116.6887 | 120.8732 | 163.2104 | 168.3545 | 186.8127 | 158.147 | 145.6455 | 139.1038 | 149.5216 | 128.2911 | 183.8962 | 187.5418 | 184.3746 | 219.1671 | 152.3977 | 143.1902 | 218.3343 | 192.3141 | 191.7723 | 194.2277 | 195.7493 | 280.1038 | 158.1671 | 170.7089 | 159.6254 | 227.2709 | 135 | 135.2507 | 131.7291 | 152.0202 | 174.0836 | 177.8962 | 187.0202 | 111.2277 | 106.1239 | 113.7291 | 109.3343 | 111.8761 | 260.3113 | 224.9164 | 236.4554 | 349.5821 | 259.415 | 261.7291 | 242.1441 | 319.6455 | 139.8329 | 65.6254 | 58.9798 | 139.8329 | 100.3948 | 117.562 | 106.8962 | 120.8962 | 116.1873 | 110.3113 | 129.0404 | 101.8962 | 68.4582 | 78.7493 | 66.6254 | 114.7291 | 103.6657 | 89.8329 | 92.4582 | 81.3343 | 263.6455 | 216.562 | 205.1873 | 299.147 | 151.6887 | 189.3545 | 189.2075 | 274.1873 | 283.3977 | 205.4784 | 210.1239 | 312.3113
1119 | 14.3034 | 168.6038 | 191.6382 | 183.5339 | 189.8624 | 159.293 | 162.3962 | 164.9312 | 141.7758 | 144.465 | 109.879 | 120.1898 | 148.3274 | 208.7414 | 130.7758 | 140.3618 | 222.3108 | 190.1554 | 145.8446 | 136.4484 | 180.172 | 161.414 | 284.465 | 303.5006 | 158.414 | 263.2075 | 205.5006 | 218.0344 | 291.879 | 115.2242 | 101.3618 | 115.7936 | 125.6548 | 168.5006 | 154.7592 | 139.0344 | 96.586 | 250.9134 | 228.5694 | 225.6903 | 266.3618 | 150.5339 | 127.0522 | 132.414 | 125.7592 | 81.0344 | 81.8102 | 81.2586 | 120.121 | 194.2075 | 197.707 | 214.8269 | 157.0866 | 136.879 | 128.8446 | 124.3452 | 119.3618 | 168.828 | 158.8446 | 188.2242 | 155.9656 | 141.8102 | 127.6548 | 151.1898 | 136.1898 | 181.2586 | 198.8624 | 193.414 | 199.2064 | 161.0522 | 153.2419 | 209.5516 | 167.2064 | 191.6726 | 201.7414 | 188.6204 | 276.8102 | 156 | 179.8624 | 156.5172 | 215.1376 | 119.4306 | 131.2586 | 139.414 | 157.9656 | 163.3618 | 172.293 | 175.172 | 123.9312 | 113.9656 | 118.4484 | 116.8624 | 116.6382 | 263.0344 | 231.9312 | 244.8102 | 321.6371 | 260.3274 | 229.3785 | 256.6726 | 298.7581 | 133.8446 | 60.293 | 68.6038 | 133.8446 | 118.3452 | 121.4828 | 105.7414 | 115.293 | 108.1032 | 106.3618 | 120.172 | 119.2764 | 73.4484 | 72.3618 | 51.6548 | 107.586 | 100.586 | 93.4828 | 98.9312 | 88.8624 | 262.0344 | 199.7236 | 212.6726 | 322.9145 | 172.6903 | 182.0688 | 190.7758 | 274.2586 | 297.9833 | 203.0688 | 202.3962 | 303.172
1120 | 14.3144 | 142.0987 | 202.5369 | 208.1623 | 179.6799 | 148.202 | 153.4779 | 160.9705 | 162.1033 | 146.2906 | 119.5664 | 129.5369 | 142.2316 | 217.2906 | 150.3496 | 149.059 | 191.0692 | 190.7389 | 158.0738 | 141.7684 | 184.798 | 167.0148 | 293.8275 | 286.1283 | 165.5221 | 239.5914 | 212.2463 | 213.2316 | 273.6799 | 133.3349 | 107.798 | 116.7242 | 123.2611 | 164.6947 | 146.9262 | 147.798 | 92.4926 | 261.0738 | 237.0148 | 244.0443 | 263.7537 | 146.5221 | 129.2316 | 128.9705 | 126.9705 | 81.5074 | 67.1873 | 84.7684 | 114.9557 | 203 | 195.9852 | 210.2758 | 164.2611 | 142.0443 | 136.5517 | 128.2463 | 114.4926 | 189.0738 | 168.059 | 207.8422 | 175.8422 | 136.2316 | 114.7094 | 150.7389 | 134.2316 | 160.6504 | 168.3303 | 215.5959 | 211.0885 | 163.2316 | 148.1873 | 195.1873 | 193.4086 | 177.4188 | 180.8967 | 180.4779 | 288.5664 | 151.4779 | 157.6209 | 157.5074 | 239.6402 | 123.798 | 156.6254 | 125.4188 | 148.4483 | 170.5517 | 175.5221 | 182.8127 | 110.1725 | 121.7832 | 115.4779 | 121.7537 | 126.7832 | 266.5221 | 250.5812 | 244.5074 | 332.6107 | 269.3053 | 251.1623 | 252.2168 | 296.2758 | 133.2611 | 61.2611 | 77.5221 | 133.2611 | 98.8819 | 109.941 | 115.798 | 111.7389 | 108.0148 | 99.2316 | 104.4483 | 116.202 | 77.2611 | 73.5221 | 70.3644 | 93.6947 | 88.9557 | 103.0443 | 97.2316 | 80.1873 | 275.3201 | 203.5517 | 211.2316 | 317.9262 | 172.9557 | 174.7242 | 187.7389 | 252.143 | 298.4631 | 204.0148 | 213.8275 | 294.2316
1121 | 14.3254 | 154.9794 | 207.2979 | 233.5536 | 176.532 | 143.468 | 148.1701 | 153.6381 | 154.4464 | 158.1918 | 113.2763 | 113.6165 | 131.0423 | 219.234 | 153.1701 | 164.4897 | 203.5114 | 188.468 | 156.1062 | 139.9361 | 191.8299 | 180.2557 | 280.6165 | 312.4691 | 170.0639 | 218.2763 | 212.766 | 218.1278 | 273.5959 | 128.2763 | 113.2979 | 116.532 | 130.8938 | 123.467 | 155.7876 | 131.8505 | 88.9361 | 219.2752 | 216.5525 | 233.2124 | 241.0206 | 131.1484 | 133.3619 | 127.766 | 124.234 | 79.7021 | 68.3619 | 80.6381 | 109.7021 | 224.4475 | 198.0639 | 230.3835 | 158.8722 | 137.3402 | 143.8299 | 146.3835 | 123.9577 | 179.4464 | 172 | 200.2124 | 155.9567 | 131.1701 | 126.5536 | 155.3619 | 139.8938 | 166.2557 | 165.4258 | 207.4464 | 210.8722 | 155.1062 | 153.9577 | 200.9577 | 187.1484 | 195.7454 | 205.4052 | 189.7237 | 302.1918 | 156.1278 | 144.4041 | 161.0639 | 221.4247 | 132.3619 | 155.8082 | 144.5114 | 148.8299 | 178.5959 | 180.8299 | 171.6804 | 123.3835 | 114.0423 | 117.0639 | 105.9144 | 128.234 | 251.1484 | 269.7876 | 249.5959 | 350.7237 | 246.9567 | 239.0206 | 259.9577 | 302.5959 | 137.0639 | 52.8082 | 83.5959 | 137.0639 | 100.9577 | 98.3402 | 115.1701 | 127.0856 | 103.6381 | 95.7021 | 110.1918 | 108.4041 | 73.4041 | 76.532 | 73.4041 | 80.8082 | 86 | 86.0845 | 92.1701 | 72.1701 | 274.6381 | 202.4041 | 228.3835 | 284.6587 | 180.7237 | 201.3413 | 197.7237 | 266.4475 | 272.2546 | 200.4041 | 213.6381 | 306.7876
1122 | 14.3364 | 146.4346 | 211.1131 | 209.204 | 198.7916 | 161.6785 | 160.2306 | 150.4435 | 146.8869 | 146.9911 | 121.6741 | 132.9047 | 159.1309 | 206.5477 | 131.7649 | 137.6474 | 217.2262 | 177.8825 | 154.7783 | 147.561 | 186.7738 | 166.0998 | 282.0044 | 293.4257 | 160.1042 | 220.4479 | 212.2217 | 225.4479 | 261.7694 | 117.9956 | 113.2217 | 94.4301 | 140.0044 | 137.683 | 134.317 | 135.3393 | 86.4435 | 237.6874 | 220.1175 | 261.6874 | 234.7783 | 133.0044 | 127.2173 | 135.0044 | 124 | 71.2173 | 61.439 | 78.2217 | 105.8869 | 204.5388 | 222.3481 | 221.2128 | 152.3304 | 141.2262 | 136.439 | 117.7561 | 109.0998 | 177.3348 | 161.8825 | 203.0044 | 146.4435 | 155.683 | 134.1131 | 139.878 | 122.5432 | 179.3393 | 189.0133 | 197.5521 | 219.8958 | 146.7738 | 164.7827 | 209.4479 | 193.6741 | 199.8869 | 207.9956 | 198.4479 | 292.5477 | 179.7916 | 149.2262 | 138.6519 | 229.3437 | 133.2217 | 141.3259 | 135.6563 | 165.5654 | 165.2128 | 168.7694 | 171.6696 | 133.6696 | 124.2306 | 110.9956 | 118.122 | 106.9867 | 244.4435 | 254.5432 | 236.9911 | 329.0953 | 263.9047 | 253.7916 | 248.9911 | 307.8914 | 131.7738 | 58.561 | 66.3215 | 131.7738 | 96.2173 | 99.8914 | 130.3437 | 120.3259 | 113.6741 | 109.7872 | 98.9911 | 107.7783 | 79.0044 | 81.6696 | 73.5565 | 75.6652 | 83.6652 | 83.8914 | 89.4435 | 72.5565 | 265.9956 | 218.9002 | 223.8825 | 310.2439 | 181.6652 | 188.9867 | 209.561 | 260.5477 | 305.4701 | 197.4435 | 221.3393 | 317.2262
1123 | 14.3474 | 153.8584 | 190.6548 | 194.6755 | 191.5604 | 160.6755 | 146.6076 | 166.6018 | 133.351 | 154.0679 | 118.6755 | 127.351 | 147.4454 | 215.649 | 161.5753 | 140.8584 | 201.6076 | 176.5811 | 159.7434 | 125.4926 | 174.7227 | 160.2094 | 285.5811 | 282.6283 | 145.9321 | 225.9528 | 213.5811 | 227 | 278.5546 | 123.1151 | 128.0207 | 114.0886 | 127.7699 | 161.6018 | 145.9735 | 156.9735 | 95.4868 | 266.7641 | 238.8113 | 253.6076 | 251.6018 | 131.8377 | 141.6018 | 151.2301 | 108.1887 | 76.1151 | 73.2301 | 74.8377 | 114.4868 | 212.8113 | 208.4454 | 213.0472 | 168.3924 | 154.0679 | 131.6283 | 128.5546 | 119.8113 | 170.8849 | 165.3245 | 209.7434 | 145.2094 | 149.5604 | 133.4189 | 149.2301 | 124.9056 | 174.8849 | 186.3038 | 188.0944 | 223.7906 | 155.2773 | 162.2566 | 215.7434 | 202.5339 | 208.4868 | 188.6076 | 192.8849 | 281.0944 | 163.0737 | 177.0886 | 142.2773 | 215.8171 | 130.6283 | 130.0944 | 136.5339 | 144.702 | 175.2301 | 161.0472 | 173.7906 | 150.8113 | 128.7906 | 109.7906 | 114.3038 | 117.6018 | 238.4661 | 224.4926 | 245.649 | 318.0472 | 270.2094 | 259.2094 | 253.6962 | 299.5132 | 122.0944 | 77.6018 | 73.649 | 122.0944 | 113.7641 | 108.1151 | 110.4926 | 120.1623 | 103.5604 | 117.9528 | 97.3717 | 117.4868 | 72.3038 | 76.6755 | 74.7906 | 79.7434 | 83.7906 | 94.4868 | 93.7434 | 83.2773 | 233.1681 | 237.4396 | 212.3038 | 290.7492 | 169.9321 | 187.7434 | 214.3717 | 266.4868 | 286.1681 | 190.6755 | 207.3982 | 305.5604
1124 | 14.3584 | 153.7886 | 184.1971 | 208.2542 | 179.1686 | 142.1401 | 155.6485 | 147.7172 | 163.7198 | 131.3087 | 124.2257 | 112.7601 | 162.0713 | 194.9144 | 134.0687 | 140.7886 | 212.2542 | 193.0571 | 162.6057 | 140.677 | 170.3943 | 170.4371 | 288.4086 | 290.0285 | 132.5629 | 222.9857 | 200.3515 | 235.8314 | 276.7743 | 114.5629 | 114.3372 | 103.3372 | 138.4514 | 149.1401 | 151.8029 | 141.9287 | 99.6057 | 247.1115 | 233.3658 | 236.1543 | 251.1829 | 140.6342 | 129.9429 | 141.3515 | 121.6628 | 83.62 | 72.1829 | 81.2257 | 117.8029 | 196.9287 | 189.3515 | 232.8742 | 154.5344 | 122.4773 | 118.9572 | 118.7458 | 127.2114 | 173.0143 | 180.6485 | 206.9857 | 151.4228 | 151.62 | 138.62 | 167.4514 | 131.0143 | 176.2114 | 159.9144 | 194.8314 | 219.1829 | 154.7886 | 144.9429 | 207.3658 | 187.1401 | 192.5344 | 184 | 178.1543 | 298.2685 | 165.0285 | 173.5629 | 165.0713 | 231.0713 | 138.0285 | 132.8171 | 145.2257 | 135.5914 | 173.38 | 175.2542 | 158.7458 | 140.5486 | 114.5486 | 97.9572 | 124.8457 | 126.8171 | 241.0143 | 223.62 | 228.9287 | 312.9857 | 260.3658 | 229.2944 | 242.3515 | 317.0713 | 127.2257 | 91.6342 | 74.5914 | 127.2257 | 102.9429 | 104.38 | 126.4799 | 134.6485 | 109.6342 | 110.9715 | 98 | 107.9572 | 76.4228 | 74.1971 | 83.8314 | 87.4228 | 94.4371 | 88.1686 | 99.8171 | 83.5914 | 269.1569 | 246.62 | 227.6628 | 291.8314 | 163.7886 | 188.1971 | 190.9144 | 280.2399 | 278.8029 | 195.4228 | 202.1971 | 323.677
1125 | 14.3694 | 166.0423 | 194.5969 | 209.5546 | 161.5123 | 163.2695 | 163.8909 | 164.0089 | 184.2272 | 124.1849 | 113.7728 | 111.6303 | 140.1002 | 217.5301 | 134.7817 | 142.4454 | 201.3274 | 165.2094 | 137.7305 | 159.8575 | 182.2272 | 155.882 | 290.6303 | 292 | 143.0423 | 239.9332 | 219.0089 | 237.1849 | 282.3363 | 130.7483 | 105.1091 | 102.2606 | 146.0757 | 165.3786 | 135.6971 | 132.1091 | 91.8486 | 260.5635 | 227.7394 | 245.2272 | 212.5033 | 135.6637 | 113.7728 | 147.7817 | 118.6637 | 73.588 | 57.1425 | 78.1091 | 122.0757 | 195.2606 | 186 | 234.7394 | 141.0334 | 128.6726 | 149.4209 | 137.0089 | 131.2606 | 177.2606 | 162.8062 | 192.1425 | 153 | 143.2183 | 122.882 | 145.7305 | 131.1849 | 160.6971 | 193.9421 | 201.0757 | 212.294 | 156.4454 | 150.7817 | 216.412 | 194.412 | 183.1091 | 184.8151 | 180.706 | 316.0423 | 175.1514 | 144.9154 | 176.5212 | 211.5457 | 144.0757 | 136.4454 | 140.4788 | 147.2272 | 171.1849 | 174.1091 | 161.5212 | 137.8151 | 102.8486 | 107.2272 | 106.8062 | 123.1091 | 244.4454 | 260.0512 | 213.4031 | 334.824 | 231.9154 | 241.5635 | 212.1002 | 317.1091 | 122.4788 | 83.4031 | 75.6303 | 122.4788 | 113.6726 | 94.8486 | 110.8062 | 116.8062 | 99.7728 | 100.0334 | 98 | 89.5123 | 101.6392 | 69.1091 | 58.2851 | 89.8151 | 105.9666 | 98.2272 | 100.1849 | 72.4031 | 249.8396 | 214.5791 | 228.7394 | 282.588 | 161.3697 | 203.4877 | 185.8151 | 268.3274 | 287.9666 | 197.8151 | 216.6726 | 304.5457
1126 | 14.3804 | 152.4513 | 199.4823 | 209.8274 | 178.6859 | 153.2787 | 157.5531 | 164.8628 | 159.6946 | 151.3054 | 120.9292 | 114.4823 | 149.7213 | 213.2433 | 159.3408 | 147.9646 | 216.2036 | 177.031 | 151.031 | 158.0354 | 207.3408 | 140.4159 | 321.6151 | 287.0354 | 162.5487 | 222.4867 | 216.5531 | 254.3761 | 269.1062 | 128.3805 | 126.3408 | 129.4779 | 139.5531 | 136.0751 | 131.1726 | 143.4115 | 83.3805 | 277.4115 | 236.9292 | 224.0044 | 261.0931 | 153.031 | 122.5841 | 140.8982 | 128.5841 | 74.3097 | 68.8938 | 72.8628 | 127.9646 | 215.8585 | 211.6505 | 210.0044 | 138.1726 | 132.8274 | 142.9336 | 144.4823 | 117.1062 | 185.4469 | 177.8585 | 191.4823 | 172.8585 | 127.761 | 143.8231 | 144.9646 | 132.6549 | 133.0044 | 185.6239 | 198.6903 | 225.8938 | 154.5177 | 155.4823 | 214.0354 | 191.208 | 220.062 | 206.5133 | 196.8938 | 290.0398 | 170.3805 | 144.792 | 146.5574 | 210.1372 | 143.3451 | 148.5841 | 142.3097 | 143.3805 | 169.3451 | 179.6195 | 170.4469 | 140.4823 | 108.4469 | 115.792 | 102.8274 | 115.3805 | 255.7567 | 251.4513 | 193.6239 | 304.4203 | 238.4115 | 246 | 220.0664 | 292.8318 | 130.9292 | 85.9646 | 76 | 130.9292 | 125.2744 | 111.2036 | 102.6903 | 120.2744 | 111.0664 | 108.7567 | 100.4823 | 103.3761 | 82.1769 | 74.6195 | 63.5841 | 93.3097 | 110.4823 | 119.2036 | 92.5531 | 74.1372 | 230.5885 | 232.6505 | 216.4159 | 294.0664 | 188.3054 | 172.2477 | 186.8274 | 279.0664 | 302.4115 | 194.6903 | 195.1769 | 292.3805
1127 | 14.3914 | 173.3521 | 184.0452 | 196.5644 | 169.5644 | 160.0767 | 160.1986 | 171.5575 | 154 | 130.1287 | 106.2055 | 141.0315 | 168.1151 | 225.2754 | 154.763 | 143.1219 | 206.5644 | 163.3657 | 138.2055 | 136.8466 | 181.7698 | 159.8329 | 295.2507 | 268.3657 | 161.8014 | 222.1986 | 210.8014 | 226.0904 | 265.1603 | 122.8014 | 113.3657 | 122.4041 | 136.3205 | 152.5123 | 149.474 | 146 | 88.7178 | 276.6411 | 209.6096 | 216.4808 | 244.4493 | 149.4425 | 144.3137 | 133.1219 | 118.4041 | 78.3589 | 58.5644 | 74.5192 | 102.1287 | 224.1986 | 189.289 | 234.3904 | 151.4356 | 164.0699 | 123.2055 | 129.8849 | 114 | 163.4877 | 155.1287 | 180.2438 | 178.6795 | 151.0315 | 148.1603 | 151.0384 | 140.5575 | 159.9096 | 187.8781 | 189.6027 | 212.2055 | 172.474 | 134.1671 | 213.8397 | 184.1219 | 206.1671 | 198.4041 | 206.7178 | 277.2822 | 165.6411 | 163.6343 | 144.1986 | 209.3205 | 127.0452 | 140.0836 | 128.7246 | 164.6726 | 166.4808 | 167.5644 | 166.9616 | 136.8014 | 109.1603 | 95.1671 | 103 | 134.9932 | 268.9164 | 248 | 203.4356 | 298.6795 | 246.0384 | 231.7246 | 211.2438 | 293.8781 | 151.474 | 87 | 64.2438 | 151.474 | 114.4041 | 123.3973 | 112.9164 | 123.6795 | 88.8082 | 99.2438 | 111.9164 | 109.5192 | 83.7178 | 60.8849 | 61.8014 | 92.3205 | 111 | 102.8466 | 87.6411 | 79.1986 | 255.711 | 221.2055 | 221.5575 | 318.8329 | 162.0904 | 197.7493 | 192.0384 | 270.2438 | 305 | 236.8261 | 205.9548 | 285.1219
1128 | 14.4024 | 160.1076 | 184.4081 | 194 | 156.7758 | 141.5516 | 156.7399 | 169.5919 | 145.4798 | 146.3004 | 120.8924 | 119.5874 | 150.5516 | 193.9193 | 156.4081 | 153.0762 | 213.3722 | 181.3004 | 140.1121 | 170.4888 | 188.7803 | 160.5919 | 292.4081 | 280.3363 | 161.852 | 234.9282 | 219.3722 | 211.4798 | 264.148 | 119.4439 | 124.4843 | 129.3722 | 138.5561 | 167.2242 | 135.1076 | 139.1839 | 94.2601 | 272.5919 | 205.704 | 206.6278 | 237.296 | 153.9641 | 141.1839 | 126.8879 | 116 | 90.9282 | 71.3363 | 67.3318 | 137.0448 | 201.1435 | 182.296 | 219.5516 | 174.4484 | 142.7354 | 142.1525 | 135.5202 | 115.704 | 185.4126 | 163.6323 | 197.5964 | 155.9955 | 156 | 137.7758 | 140.0718 | 144.5561 | 157.4798 | 179.6278 | 189.704 | 209.852 | 164.0718 | 154.7085 | 199.5157 | 186.4081 | 202.852 | 189.1839 | 206.296 | 284.5202 | 162.4439 | 175.5202 | 178.2287 | 206.4439 | 144.4484 | 138.852 | 134.5202 | 161.3318 | 160.8879 | 156.4798 | 163.4439 | 128.3318 | 98.7758 | 127.6368 | 103 | 138.148 | 259.0718 | 249.704 | 231.5605 | 308.3722 | 263.1883 | 229.852 | 218.3722 | 274.5516 | 127.7354 | 85.296 | 83.3004 | 127.7354 | 109.4439 | 114.7758 | 110.7399 | 109.5157 | 96.7803 | 106.3722 | 100.3677 | 111.704 | 85 | 82.7085 | 77.1883 | 94.5561 | 89.6996 | 88.7758 | 92.9641 | 69.7758 | 237.1435 | 208.6278 | 213.6278 | 340.8924 | 177.3004 | 192.0718 | 171.6996 | 252.6637 | 319.4843 | 232.2197 | 209 | 305.3004
1129 | 14.4134 | 172.5576 | 198.8289 | 200.9145 | 174.0148 | 149.236 | 167.236 | 181.9646 | 146.5929 | 143.9498 | 130.0502 | 110.6785 | 173.7936 | 217.3865 | 146.6283 | 169.6933 | 222.7788 | 173.764 | 141.8643 | 138.9704 | 171.9852 | 181.6077 | 299.9145 | 290.7788 | 169.7788 | 230.0855 | 198.528 | 203.0855 | 251.0354 | 125.0502 | 113.1711 | 118.0354 | 127.764 | 143.0707 | 153.6077 | 136.2714 | 85.4926 | 276.3215 | 234.5222 | 228.3363 | 242.1859 | 158.4572 | 133.0855 | 139.8289 | 126.3717 | 72.2566 | 86.1003 | 69.4572 | 105.1061 | 239.3511 | 191.5074 | 228.1003 | 141.699 | 141.4572 | 136.4926 | 149.1003 | 119.4572 | 170.9852 | 163.4071 | 202.7286 | 152 | 139.5781 | 123.8997 | 164.7936 | 130.3067 | 152.5428 | 155.528 | 193.4572 | 238.5222 | 168.9145 | 168.5074 | 233.301 | 187.8643 | 187.4424 | 181.0855 | 191.3067 | 305.8791 | 182.7434 | 168.3569 | 157.2064 | 218.9646 | 139.3569 | 170.9794 | 123.8997 | 147.0354 | 169.5074 | 161.9145 | 166.4572 | 122.6785 | 102.1859 | 109.7994 | 112.5074 | 132.8141 | 260.4572 | 226.6637 | 221.3067 | 333.3363 | 266.8643 | 242.1003 | 260.6225 | 296.9293 | 117.8141 | 78.9498 | 66.2566 | 117.8141 | 110.7286 | 112.1357 | 114.3215 | 111.3215 | 97.2714 | 107.1357 | 107.5074 | 117.1859 | 71.1711 | 59.3421 | 71.3569 | 92.4071 | 88.5929 | 104.2862 | 96.5929 | 79.236 | 206.2064 | 217.3717 | 207.6785 | 321.528 | 161.9852 | 202.1003 | 190.472 | 240.4926 | 276.1916 | 210.1209 | 212.4572 | 288.2566
1130 | 14.4244 | 156.5914 | 180.8382 | 223.915 | 162.0978 | 143.9872 | 149.7148 | 175.234 | 147.8766 | 136.8638 | 129.2468 | 110.8766 | 155.2084 | 214.1106 | 142.3702 | 150.085 | 195.0722 | 168.4936 | 136.7404 | 146.149 | 166.3702 | 174.4808 | 272.0722 | 278.851 | 162.234 | 217.6042 | 220.4214 | 214.2724 | 281.4342 | 120.7404 | 115.383 | 127.3958 | 136.5192 | 163.5448 | 138.5914 | 152.6554 | 74.3574 | 264.7276 | 223.2212 | 229.3702 | 236.8638 | 151.1106 | 123.234 | 127.9744 | 156.9278 | 89.1618 | 71.3446 | 69.1234 | 116.532 | 208.3062 | 183.3574 | 237.8894 | 150.9022 | 163.915 | 111.3318 | 128.2084 | 101.5914 | 185.532 | 180.532 | 194.234 | 180.0512 | 143.1362 | 132.5192 | 137.4424 | 124.4936 | 170.4086 | 175.6682 | 211.532 | 218.4552 | 171.7532 | 146.3318 | 205.6892 | 209.0384 | 183.2468 | 216.8172 | 190.7532 | 301.9872 | 168.468 | 175.766 | 166.149 | 215.7404 | 144.1362 | 146.1956 | 118.4936 | 156.3958 | 174.5064 | 176.149 | 175.766 | 129.0128 | 114.3958 | 98.1106 | 108.7404 | 135.5064 | 284.6682 | 240.532 | 242.6682 | 320.3446 | 238.9488 | 243.1234 | 222.2934 | 286.0978 | 135.4086 | 76.2468 | 68.2596 | 135.4086 | 127.6554 | 116.383 | 99.2212 | 103.234 | 112.7788 | 103.4936 | 104.617 | 113.617 | 88.2852 | 65.5192 | 61.234 | 95.5064 | 87.2468 | 104.3702 | 105.766 | 82.7532 | 223.915 | 204.9744 | 211.383 | 277.6764 | 173.9022 | 205.7532 | 173.8382 | 274.9406 | 267.2468 | 224.532 | 202.4808 | 294.6426
1131 | 14.4354 | 144.2222 | 202.8889 | 208.3333 | 161.7778 | 148.3333 | 163.8889 | 153.5555 | 169.3334 | 124.4444 | 130.7778 | 125.2223 | 153.7778 | 232.5556 | 137.5555 | 144.3333 | 226.5556 | 168 | 151.1111 | 140 | 203.3334 | 154.3333 | 295.5556 | 293 | 137.8888 | 226.6667 | 228.4445 | 198.2222 | 275.3333 | 120.8889 | 107.1111 | 117.4444 | 148.6667 | 153.6666 | 144 | 141.6666 | 85.4445 | 247 | 231.6667 | 219.2222 | 242.2222 | 155.3333 | 146.0001 | 147.3334 | 125.4444 | 75.1111 | 88.5556 | 79.6667 | 113.6667 | 206.5556 | 203.3334 | 227.4444 | 146.7778 | 153.6666 | 132.0001 | 137.4445 | 135.4445 | 183.5555 | 170.5555 | 206.3334 | 179.5555 | 140.4444 | 157.1112 | 145.4445 | 116.8889 | 162.3333 | 187 | 214.8889 | 237.2223 | 176.4445 | 159.8889 | 197.4444 | 183.5555 | 172.3333 | 189.9999 | 195.4445 | 301 | 164.2222 | 164.5555 | 145.7777 | 231.8889 | 148.5556 | 152.6667 | 139.3334 | 144.6666 | 156.3333 | 157.5555 | 177.8889 | 121.1111 | 103.5555 | 104.1111 | 117.7778 | 122.6666 | 261.3333 | 230.5555 | 230 | 328.6667 | 235.8889 | 238.5555 | 238.2223 | 286.6667 | 138.8889 | 76 | 69.8889 | 138.8889 | 119.3333 | 114.3333 | 116.5556 | 119.7778 | 105.2222 | 111.8889 | 108.4445 | 116.5556 | 75.8889 | 55.4444 | 81.3334 | 96.8889 | 124.3334 | 88 | 82.1111 | 82.1111 | 240.3334 | 202.1111 | 218.2222 | 275.5556 | 188.4445 | 187.3333 | 192.3334 | 290.6667 | 287.4445 | 209.2222 | 212.5556 | 300.4445
1132 | 14.4464 | 152.0118 | 197.8894 | 196.087 | 166.5059 | 161.6165 | 154.2846 | 142.8894 | 162.087 | 152.739 | 112.0752 | 117.9882 | 170.2213 | 210.6681 | 139.7035 | 152.1106 | 218.3835 | 178.8142 | 133.174 | 147.1106 | 177.3598 | 160.1106 | 299.9012 | 251.7433 | 155.7272 | 237.913 | 208.2728 | 231.1461 | 261.3835 | 126.4071 | 105.0988 | 128.6165 | 131.9764 | 139.3835 | 146.8024 | 121.9764 | 78.8894 | 245.9012 | 217.6799 | 256.7508 | 249.3083 | 157.8024 | 127.3716 | 135.5811 | 122.8024 | 79.3083 | 85.5929 | 71.9882 | 102.1858 | 205.1976 | 207.8024 | 234.1106 | 164.0236 | 118.6563 | 118.7787 | 144.4071 | 110.261 | 190.2095 | 180.7154 | 180.0634 | 164.5811 | 142.7035 | 161.8024 | 162.3201 | 131.3201 | 165.5059 | 178.9882 | 169.0397 | 234.5929 | 177 | 152.9882 | 224.9366 | 195.3201 | 196.2331 | 186 | 200.5059 | 296.4941 | 180.2213 | 181.0236 | 168.2331 | 234.9012 | 150.8024 | 135.0752 | 133.8894 | 127.6799 | 161.2095 | 170.3201 | 171.6917 | 121.8024 | 106.5059 | 101.3953 | 116.2965 | 130.913 | 248.087 | 264.1461 | 255.9366 | 359.739 | 234.1976 | 213.6681 | 212.1622 | 329.3556 | 133.5929 | 78.7035 | 82.6165 | 133.5929 | 120.7035 | 98.6799 | 106.3835 | 116.5929 | 119.3201 | 107.5929 | 114.4071 | 102.5811 | 83.913 | 66.6165 | 77.6917 | 96.0988 | 92.0515 | 80.5929 | 90.7154 | 83.8024 | 216.7669 | 211.0118 | 217.1976 | 266.9882 | 181.8894 | 213.8378 | 196.8024 | 243.3362 | 278.2846 | 218.7154 | 218.5059 | 296.4941
1133 | 14.4574 | 164.8751 | 209.7886 | 208.7021 | 188.9234 | 147.471 | 148.4326 | 163.9234 | 159.1731 | 154.1731 | 144.712 | 123.3943 | 137.288 | 223.529 | 138.1731 | 184.0581 | 216.0865 | 186.3943 | 141.9617 | 149.8269 | 172.1731 | 160.0865 | 298.1731 | 274.4042 | 148.8653 | 222.5575 | 224.2695 | 201.2015 | 286.4907 | 149.8368 | 136.0581 | 117.2114 | 134.5674 | 158.0964 | 141.5192 | 154.712 | 109.0581 | 241.4326 | 237.0099 | 229.0284 | 235.3844 | 155.2596 | 144.1829 | 131.2596 | 141.2695 | 79.0865 | 70.3844 | 67.3461 | 119.2695 | 218.7021 | 208.9135 | 218.5575 | 146.8171 | 144.2311 | 138.9234 | 129.471 | 101.5192 | 172.7305 | 177.4326 | 208.0581 | 156.6057 | 130.2114 | 133.6823 | 153.9518 | 110.1632 | 177.8751 | 167.9518 | 171.3078 | 201.115 | 175.1731 | 166.6156 | 228 | 214.356 | 196.2596 | 196.9617 | 201.9135 | 295.0865 | 169.2114 | 169.2979 | 175.5674 | 213.9901 | 144.6057 | 142.1347 | 139.3943 | 164.3659 | 196.712 | 164.6922 | 175.5674 | 128.3943 | 111.5674 | 113.7886 | 116.9135 | 117.3844 | 247 | 248.8171 | 223.3746 | 345.644 | 247.7021 | 262.1545 | 245.5389 | 276.4512 | 111.0766 | 77.1731 | 69.3844 | 111.0766 | 104.5575 | 104.3078 | 101.3461 | 102.2979 | 105.471 | 100.6057 | 107.6922 | 122.9234 | 87.7404 | 78.0482 | 68.7787 | 93.2596 | 95.3078 | 93.7021 | 90.1731 | 68.471 | 260.5871 | 220.2213 | 240.7503 | 332.6835 | 178.2596 | 183.2015 | 178.7305 | 303.7701 | 271.5192 | 224.5674 | 192.5093 | 291.4326
1134 | 14.4684 | 153.0393 | 214.7031 | 221.1092 | 168.7817 | 127.4847 | 153.5546 | 167.8515 | 161.7773 | 143.8166 | 128.559 | 107.3362 | 145.1092 | 222.2227 | 161.1441 | 176.8166 | 195.6332 | 169.4105 | 137.4454 | 141.6681 | 169.2227 | 156.2969 | 294.2969 | 259.4105 | 160.0349 | 234.8865 | 215.8166 | 221.1441 | 250.1179 | 119.5983 | 132.5197 | 124.3319 | 146.1092 | 145.1878 | 131.7424 | 141.3362 | 83.3013 | 265.0699 | 240.8515 | 239.8865 | 282.1397 | 151.2969 | 160.8122 | 146.738 | 130.9651 | 65.1135 | 70.8515 | 69.7773 | 101.559 | 211.6681 | 185.8559 | 234.5895 | 157.0349 | 134.0393 | 135.4454 | 138.1834 | 121.3668 | 167.2969 | 170.5197 | 235.0699 | 147.6681 | 137.3319 | 116.1878 | 160.4061 | 110.7773 | 191.0349 | 148.4847 | 202.5502 | 213.738 | 156.4847 | 136.524 | 216.8908 | 185.4498 | 198.7773 | 192.4454 | 189.0393 | 317.2183 | 192.0699 | 155.9651 | 171.3712 | 236.9956 | 136.5939 | 135.5939 | 116.8559 | 131.8952 | 167.5983 | 170.4803 | 177.8515 | 140.1092 | 105.5197 | 121.4803 | 107.7424 | 106.7424 | 256.2576 | 246.0742 | 229.2576 | 319.0044 | 257.3319 | 235.524 | 236.9651 | 295.0699 | 116.4061 | 77.9258 | 74.4803 | 116.4061 | 101.1485 | 102.2227 | 120.441 | 101.9258 | 119.738 | 111.1092 | 103.2969 | 99.0786 | 66.7074 | 66.0393 | 82.8122 | 93 | 71.9301 | 99.6288 | 92.7773 | 79.9607 | 212.2314 | 220.0742 | 201.3406 | 294.5633 | 144.6725 | 209.6245 | 196.441 | 235.9389 | 292.2926 | 205.559 | 197.4061 | 310.441
1135 | 14.4794 | 150.1239 | 199.0531 | 207.9292 | 195.1417 | 134.4425 | 154 | 172.6903 | 171.3806 | 142.0619 | 120.4336 | 109.7522 | 146.9381 | 211.6814 | 132.9822 | 150.6725 | 210.885 | 189.5753 | 159.5133 | 138.1858 | 155.8672 | 161.6283 | 279.9292 | 279.5753 | 146.9292 | 215.3628 | 199.9911 | 218.3097 | 247.9381 | 119.8142 | 129.1858 | 126.8761 | 137.6194 | 167.4514 | 143.1947 | 137.1858 | 105.3894 | 268.8761 | 215.6725 | 226.9292 | 221.2742 | 137.8672 | 146.0531 | 132.9911 | 120.6194 | 73.3806 | 78.5044 | 70 | 113.1328 | 198.8053 | 229.9647 | 205.0442 | 175.823 | 143.3186 | 144.3806 | 127.7433 | 125.8142 | 177.3186 | 164.3717 | 212.6106 | 141.3717 | 139.8761 | 138.4514 | 141.3008 | 141.0178 | 178.8672 | 162.0089 | 214.3806 | 197.177 | 169.0708 | 156.5133 | 256.3364 | 205.5133 | 209.3186 | 191.0619 | 203.9469 | 318.0619 | 172.4247 | 157.8142 | 171.9381 | 230.5575 | 141.6283 | 133.1239 | 149.708 | 149.6372 | 154.6814 | 154.115 | 190.1947 | 121.3008 | 127.5133 | 102.3008 | 110.7522 | 134.1417 | 261.6903 | 223.4867 | 237.5044 | 331.0708 | 265.5044 | 220.8053 | 256.6372 | 292.3097 | 119.8142 | 68.6194 | 78.7522 | 119.8142 | 109.4425 | 138.5842 | 99.4867 | 111.3806 | 105.9911 | 102.6194 | 81.4247 | 105.4425 | 78.1328 | 80.0089 | 61.4867 | 96.7522 | 97.2036 | 94.3717 | 82.6814 | 67.8672 | 253.0266 | 212.4956 | 224.2655 | 290.0619 | 166.3894 | 215.7522 | 208.3186 | 270.3364 | 268.6725 | 206.8142 | 219.5753 | 314.8142
1136 | 14.4904 | 134.7945 | 193.2483 | 190.8441 | 193.1986 | 169.2125 | 142.5958 | 177.7517 | 146.3407 | 159.1062 | 122.851 | 103.3476 | 148.9007 | 202.4469 | 133.851 | 160.4042 | 195.8441 | 169.142 | 136.291 | 146.5531 | 188.2621 | 142.0427 | 299.9076 | 304.7587 | 152.6524 | 222.5531 | 222.7587 | 194.2413 | 267.0069 | 113.3476 | 121.3972 | 107.0427 | 132.2483 | 156.6455 | 146.851 | 148.4042 | 94.6455 | 262.3476 | 255.8152 | 214.5958 | 226.5035 | 139.851 | 144.0497 | 126.2979 | 103.8441 | 80.6524 | 75.1986 | 74.7517 | 121.6028 | 206.5531 | 217.7945 | 209.6524 | 146.5889 | 157.3048 | 125.0427 | 137.4538 | 105.0924 | 153.291 | 147.8441 | 209.0993 | 147.6524 | 121.9434 | 137.149 | 170.4111 | 120.1917 | 177.0497 | 179.1559 | 220.7021 | 194.0993 | 152.8938 | 151.3476 | 220.0358 | 200.3476 | 199.5462 | 202.4042 | 203.0993 | 322.7517 | 169.0993 | 149.4469 | 148.2413 | 194.8372 | 121.0924 | 147.2552 | 147.2483 | 168.1062 | 187.2621 | 167.2552 | 173.8938 | 126.6524 | 128.0497 | 110.5035 | 98.6455 | 122.6952 | 250.5958 | 233.4042 | 218.0427 | 315.8441 | 258.3972 | 228.5531 | 247.5462 | 291.0497 | 120.9503 | 71.8014 | 61.8938 | 120.9503 | 97.6455 | 109.6386 | 96.0993 | 104.3972 | 94.5462 | 96.2979 | 125.6166 | 117.4042 | 74.2483 | 67.6952 | 71.4042 | 89.3972 | 91.3972 | 113.9573 | 93.4042 | 87.9076 | 233.1917 | 216.7517 | 231.7021 | 289.0497 | 174.6524 | 173.2344 | 209.9503 | 247.3407 | 306.9145 | 218.4042 | 217.1986 | 299.7945
1137 | 14.5014 | 147.4769 | 171.822 | 199.6264 | 187.2242 | 119.0176 | 174.7297 | 156.822 | 167.1407 | 164.8132 | 102.7846 | 127.0659 | 164.4022 | 203.9253 | 148.4396 | 144.6352 | 198.8505 | 168 | 149.4396 | 151.8132 | 165.9341 | 149.6637 | 288.4857 | 263.644 | 163.589 | 214.3363 | 206.6725 | 214.178 | 287.2527 | 135.1407 | 122.9253 | 111.7758 | 142.589 | 166.589 | 126.7846 | 135.5231 | 127.6923 | 258.1495 | 225.2703 | 223.6264 | 275.1319 | 146.7385 | 145.9253 | 116.3736 | 122.2527 | 77.1495 | 77.8879 | 63.4484 | 115.2615 | 217.589 | 182.3451 | 210.9626 | 147.8879 | 141.6352 | 139.4022 | 102.3824 | 125.178 | 171.2527 | 170.1033 | 185.8967 | 156.6637 | 133.5143 | 162.0286 | 164.2989 | 112.2615 | 160.6352 | 198.2901 | 217.1495 | 208.4396 | 168.3648 | 146.1868 | 227.6264 | 206.7385 | 210.5516 | 187.5978 | 213.589 | 291.233 | 173.8132 | 145.1495 | 146.0374 | 208.4022 | 135.4022 | 146.0747 | 138.3363 | 172.8505 | 168.7846 | 155.4857 | 179.7385 | 108.7099 | 141.4769 | 111 | 113.4022 | 139.3275 | 242.2989 | 266.7297 | 222.7758 | 313.0747 | 269.5516 | 257.8791 | 242.1868 | 290.0374 | 149.8791 | 74.8879 | 66.7758 | 149.8791 | 103.7385 | 120.5143 | 113.3275 | 128.0659 | 104.589 | 96.9626 | 93.3451 | 109.3363 | 77.8505 | 67.9626 | 72 | 88.0374 | 86.1868 | 87.0835 | 92.0747 | 78.411 | 236.8132 | 202.5604 | 209.8593 | 284.1868 | 179.8132 | 189.2901 | 212.8879 | 260.4396 | 290.7099 | 193.9714 | 212.1868 | 299
1138 | 14.5124 | 132.4011 | 188.5487 | 179.5265 | 192.8496 | 156.9721 | 145.7772 | 168.6741 | 177.7493 | 171.8245 | 127.3481 | 111.4262 | 150.3761 | 220.5738 | 148.0251 | 170.3231 | 222.3983 | 171.8997 | 121.7271 | 116.9026 | 168.8997 | 167.5487 | 292.8746 | 285.3983 | 161.0752 | 218.8746 | 221.5989 | 214.0251 | 258.7522 | 127.2256 | 129.8245 | 129.5487 | 134.2256 | 145.5516 | 131.8496 | 160.3481 | 90.9778 | 237.5265 | 204.5014 | 217.1755 | 229.2285 | 151.8746 | 146 | 122.8245 | 130.7994 | 69.2006 | 64.351 | 79.5738 | 146.1977 | 190.702 | 200.4986 | 227.5738 | 148.9749 | 132.2256 | 152.6741 | 119.5236 | 127.9499 | 138.8524 | 154.4262 | 209.3732 | 168.6991 | 156.4234 | 131.8023 | 163.0251 | 127.5989 | 167.7994 | 183.4011 | 178.9778 | 214.8496 | 169.9749 | 137.2256 | 198.7522 | 180.6769 | 183.702 | 185.0501 | 192.5516 | 317.298 | 160.351 | 184.9721 | 167.4484 | 205.1003 | 125.2758 | 128.4513 | 131.1755 | 163.2507 | 149.4764 | 186.1977 | 139.053 | 155.7715 | 111.7772 | 105.1504 | 134.4735 | 121.4764 | 259.5487 | 246.5516 | 209.351 | 308.1254 | 221.2536 | 205.3789 | 246.8746 | 283.1755 | 125.6519 | 87.6741 | 78.6991 | 125.6519 | 111.7994 | 102.4764 | 116.9248 | 115.351 | 102.0752 | 88.2256 | 96.8746 | 130.4484 | 67.2758 | 65.0752 | 78.8245 | 82.1504 | 105.4986 | 90.8746 | 94.9248 | 91.649 | 217.5014 | 208.8245 | 205.1003 | 275.2256 | 175.1254 | 216.3231 | 192.5265 | 272.6991 | 281.2256 | 195.9248 | 216.8746 | 301.9248
1139 | 14.5234 | 154.706 | 177.1534 | 202.6933 | 177.2045 | 155.0383 | 129.2045 | 167.0256 | 145.4218 | 137.4473 | 130.9617 | 116.9233 | 146.0511 | 187.4345 | 159.8466 | 168.0383 | 179.5623 | 181.8722 | 126.9233 | 141.6677 | 166.0383 | 168 | 275.2301 | 291.9233 | 159.0256 | 196.294 | 224.9617 | 209.0639 | 305.3865 | 123.0511 | 149.7444 | 132.9617 | 139.9233 | 156.8466 | 126.0767 | 121.5112 | 100.8594 | 223.1789 | 213.8722 | 238.7188 | 244.7827 | 139.1661 | 152.9105 | 128.9233 | 131.9872 | 78.8722 | 59.0639 | 96.7827 | 132.1917 | 205.7955 | 210.8722 | 212.2045 | 140.115 | 158.6549 | 142.1406 | 141.7188 | 118.1278 | 137.0128 | 169.7955 | 187.294 | 151.2301 | 148.115 | 146.7955 | 149.1789 | 128 | 180.8339 | 175.1022 | 216.5016 | 196.2428 | 172.9617 | 172.5399 | 201.9489 | 201.7188 | 213.6038 | 174.1406 | 177.1917 | 295.294 | 170.8594 | 175.1406 | 159.115 | 229.6805 | 156.591 | 144.7827 | 162.591 | 163 | 167.7572 | 185.0256 | 159.7188 | 112.5751 | 114.9489 | 98.0895 | 114.2684 | 129.885 | 251.115 | 258.8339 | 231.706 | 354.3993 | 258.5016 | 278.0415 | 256.8722 | 290.8978 | 133.885 | 72.2045 | 62.2173 | 133.885 | 114.9617 | 96.0767 | 94.294 | 115 | 103.9744 | 112.6805 | 104.8978 | 107.3067 | 61.0767 | 63.0256 | 68.1406 | 91.8722 | 100.0767 | 96.9233 | 89.0767 | 83.115 | 257.476 | 209.9872 | 176.3706 | 301.6549 | 168.0895 | 201.2045 | 197.9233 | 301.6294 | 281 | 201.9233 | 216.0128 | 302.9872
1140 | 14.5344 | 157.9985 | 220.9785 | 207.9976 | 193.9917 | 162.9961 | 145.9917 | 152.0073 | 148.998 | 123.0069 | 117.0069 | 118.999 | 134.0059 | 214.9863 | 157.0015 | 152.0078 | 205.9868 | 148.0166 | 145.9907 | 154.9936 | 202.9819 | 162.0029 | 295.9897 | 299.9961 | 173.9927 | 232.9819 | 237.9936 | 202.0034 | 279.0132 | 122.0005 | 130.0098 | 119.0069 | 132.0039 | 152.0024 | 133.9961 | 139.9907 | 97.002 | 233.9946 | 265.9745 | 241.9985 | 240.0024 | 153.9927 | 143.0049 | 119.0049 | 115.0083 | 67.0059 | 64.9971 | 63.0166 | 113.0093 | 213.9961 | 220.9951 | 225.9931 | 147.9961 | 124.0171 | 124.0088 | 124.0088 | 106.0059 | 181.978 | 149.0103 | 187.9995 | 164.9931 | 127.0103 | 134.0064 | 171.9887 | 132.9976 | 195.9927 | 185.9946 | 190.0132 | 272.9623 | 154.0093 | 146.0132 | 223.9892 | 198.002 | 198.0078 | 204.9848 | 171.0029 | 263.0157 | 164.0034 | 181.9966 | 171.9936 | 206.0117 | 160.998 | 145 | 132.0152 | 154.0044 | 180.9936 | 156.0142 | 195.9824 | 131.9902 | 124.9951 | 119.9892 | 120.9966 | 129.0005 | 249.001 | 238.0103 | 247.9922 | 307.0235 | 256.0015 | 238.0201 | 232.0122 | 313.9887 | 122.0059 | 65.0034 | 60.001 | 122.0059 | 108.0034 | 110.9927 | 117.9883 | 122.9961 | 120.9917 | 111.001 | 108.998 | 120.9931 | 80.9902 | 89.9868 | 58.0049 | 97.9971 | 108.9956 | 94.0015 | 101.9936 | 81.001 | 225.0162 | 230.9897 | 222.977 | 286.0078 | 155.0064 | 179.0108 | 203.9971 | 299.0015 | 300.9902 | 185.0083 | 221.9971 | 328.9873
1141 | 14.5454 | 159.9528 | 235.6578 | 188.1534 | 176.8348 | 137.0826 | 129.3068 | 152.2124 | 149.1062 | 132.2242 | 127.8466 | 105.1652 | 158.9292 | 208.0944 | 140.9056 | 150.5279 | 205.118 | 198.6814 | 147.941 | 139.9764 | 188.7522 | 171.9764 | 297.1888 | 263.4249 | 182.7168 | 218.4485 | 223.9056 | 224.8348 | 259.2596 | 137.7522 | 109.0354 | 111.0944 | 139.9764 | 165.6224 | 142.1534 | 123.4957 | 87.7758 | 260.0118 | 227.0826 | 268.646 | 243.236 | 172.6224 | 134 | 130.0236 | 103.177 | 85.8112 | 65.882 | 83.6342 | 111.0826 | 220.7404 | 218.7168 | 222.0708 | 150.2714 | 125.0472 | 131.8466 | 159.9292 | 121.059 | 134.3422 | 141.2478 | 189.8466 | 130.0472 | 146.8348 | 143.059 | 163.9174 | 125.1888 | 144.1652 | 190.5751 | 204.6696 | 209.1062 | 177.7758 | 146.9528 | 218.118 | 195.9528 | 208.0708 | 205.5515 | 180.5429 | 287.5075 | 160.7994 | 164.1298 | 154.882 | 219.9174 | 131.354 | 152.823 | 122.177 | 152.7404 | 143.4485 | 179.5987 | 168.0118 | 141.8466 | 105.3776 | 116.7876 | 110.9292 | 127.0826 | 248.1416 | 242.2478 | 220.0236 | 312.1416 | 235.1062 | 237.0944 | 258.4217 | 313.7876 | 117.4013 | 86.7876 | 71.8702 | 117.4013 | 99.236 | 112.9882 | 82.3186 | 113.0236 | 117.8584 | 91.2242 | 126.7286 | 129.7286 | 59.0826 | 52.3304 | 70.0826 | 100.7404 | 96.0354 | 95.1534 | 109.7994 | 88.0118 | 260.5633 | 225.764 | 228.7758 | 276.295 | 167.1416 | 230.5161 | 165.4367 | 290.6696 | 281.3658 | 189.3068 | 203.8938 | 270.7789
1142 | 14.5564 | 155.6386 | 207.0723 | 201.1446 | 163.9396 | 144.0723 | 155.4337 | 169.47 | 157.8554 | 150.7109 | 115.0723 | 119.1446 | 152.9518 | 215.8795 | 133.0241 | 111.6505 | 214.9036 | 172.0723 | 142.8073 | 137.9759 | 167.8073 | 170.3373 | 312.8314 | 298.6868 | 159.1446 | 255.0845 | 216.4096 | 210.5422 | 281.5541 | 116.9518 | 112.0964 | 119.1446 | 137.6145 | 134.0241 | 154.9759 | 164.6868 | 69.4578 | 261.53 | 234.53 | 239.4337 | 262.5181 | 140.6386 | 133.8073 | 131.4218 | 118.4337 | 70.5782 | 56.1927 | 53.7228 | 117.8314 | 198.9036 | 195.265 | 227.4218 | 172.3495 | 129.8673 | 120.3251 | 154.0241 | 125.9277 | 163.2409 | 161.6868 | 177.6023 | 135.3251 | 133.2409 | 147.735 | 157.2409 | 140.6386 | 158.6264 | 154.8314 | 178.1323 | 217.5181 | 158.7591 | 143.0723 | 227.9277 | 192.3614 | 213.229 | 168.8673 | 225.3736 | 329.2049 | 144.6505 | 174.9036 | 144.8795 | 212.6145 | 160.5663 | 138.3614 | 136.7832 | 131.5782 | 180.5904 | 146.7228 | 168.9036 | 129.2168 | 136.0122 | 99.5541 | 105.1927 | 133.2049 | 260.1686 | 262.6627 | 222.3855 | 324.1205 | 243.6145 | 244.2772 | 210.265 | 295.3977 | 150.8554 | 69 | 60.735 | 150.8554 | 119.3132 | 111.7591 | 109.1205 | 115.1205 | 105.7832 | 109.9036 | 104.2168 | 107.1927 | 66.265 | 79.9277 | 77.0241 | 79.3614 | 98.735 | 108.0241 | 93.265 | 88.8073 | 223.8554 | 205.8314 | 209.9277 | 300.5422 | 178.8314 | 190.53 | 202.6987 | 262.5904 | 311.47 | 214.5181 | 195.1205 | 335.4459
1143 | 14.5674 | 141.6913 | 209.1632 | 207.6549 | 200.9449 | 147.8368 | 173.473 | 148.6549 | 152.1092 | 139.3638 | 118.2183 | 124.0904 | 151.1455 | 210.5634 | 133.7089 | 137.8181 | 211.3275 | 174.8545 | 134.7453 | 137.1819 | 161.2007 | 183.1632 | 305.5634 | 285.8908 | 165 | 218.9824 | 232.2723 | 192.8368 | 302.29 | 115.0728 | 115.0904 | 124.4542 | 122.4002 | 135.7641 | 152.8357 | 151.1268 | 88.1455 | 281.3627 | 256.9096 | 256.3087 | 242.527 | 126.4366 | 126.473 | 108.7277 | 135.4542 | 93.5998 | 64.3638 | 83.1455 | 111.5094 | 195.5094 | 205.7089 | 204.2911 | 147.2007 | 163.5083 | 172.8357 | 154.1996 | 123.1455 | 172.1268 | 148.7453 | 202.1455 | 188.0176 | 143.3638 | 136.4179 | 165.9813 | 126.3275 | 183.5998 | 148.6549 | 223.5634 | 198.946 | 149.7641 | 145.6362 | 224.1268 | 205.763 | 182.2183 | 202.763 | 200.7277 | 296.8181 | 170.7817 | 171.946 | 140.5821 | 198.0551 | 143.5094 | 152.7817 | 128.5821 | 155.2183 | 164.1819 | 175.6725 | 165.1819 | 138.3638 | 96.8732 | 121.0176 | 113.5094 | 101.3638 | 266.8181 | 250.2007 | 237.3451 | 329.8004 | 229.2007 | 216.4917 | 222.5645 | 273.0375 | 143.6538 | 68.6725 | 51.0551 | 143.6538 | 131.0904 | 102.473 | 114.5821 | 120.1092 | 97.8732 | 106.2547 | 111.7266 | 114.4542 | 77.2911 | 77.5094 | 77.6362 | 93.7453 | 88.3638 | 108.3087 | 103.6725 | 81.3275 | 219.5281 | 199.5458 | 207 | 281.4906 | 171.9636 | 211.0176 | 229.7993 | 245.7453 | 289.3087 | 195.9096 | 200.7641 | 312.8181
1144 | 14.5784 | 159.1239 | 186.9513 | 224.3672 | 172.6106 | 169.1239 | 184.3451 | 165.5619 | 154.9027 | 149 | 123.9027 | 102.8231 | 154.6106 | 198.5619 | 125.854 | 134.2169 | 220.0973 | 170.9513 | 127.9513 | 141.1725 | 192.8053 | 160.7566 | 292.7831 | 283.8761 | 164.4646 | 243.3937 | 213.7788 | 215.4867 | 258.5089 | 116.7566 | 92.1195 | 110.8761 | 133.4867 | 155.2212 | 122.292 | 128.8275 | 92.146 | 238.8761 | 277.9822 | 237.1725 | 231.5576 | 138.5841 | 138.6593 | 127.854 | 121.146 | 82.854 | 73.3672 | 86.9027 | 124.5133 | 208.2699 | 198.2434 | 211.9027 | 176.8583 | 125.3142 | 141.6106 | 132.9513 | 126.4646 | 150.5576 | 143.3629 | 204.5398 | 163.0708 | 152.5619 | 122.1682 | 140.5576 | 136.0222 | 172.6593 | 166.8761 | 210.0043 | 224.3407 | 170.5354 | 137.2655 | 202.7036 | 175.3364 | 188.146 | 171.2655 | 219.0265 | 291.3672 | 165.2434 | 195.2964 | 156.5841 | 224.8805 | 156.708 | 146.8053 | 142.9292 | 159.3937 | 169.7788 | 166.4646 | 170.146 | 147.1239 | 120.146 | 96.0222 | 125.7831 | 111.8275 | 262.1947 | 279.8583 | 222.6771 | 348.5176 | 260.3186 | 255.0752 | 261.5176 | 326.8053 | 108.9248 | 61.0708 | 78.8053 | 108.9248 | 106.7566 | 114.708 | 128.1504 | 122.9027 | 119.9292 | 112.9027 | 79.5089 | 100.5841 | 85.146 | 91.1947 | 67.4646 | 87.6328 | 99.4602 | 91.1195 | 95.5354 | 89.6593 | 258.2964 | 213.1239 | 208.2655 | 268.3407 | 171.292 | 186.4602 | 198.7788 | 239.8275 | 271.0487 | 219.7566 | 219.8318 | 308.292
1145 | 14.5894 | 143.2803 | 187.5241 | 213.0364 | 166.1583 | 152 | 152.5487 | 158.5851 | 153.061 | 149.3048 | 121.2684 | 155.1956 | 147.7926 | 192.1948 | 123.7316 | 157.5123 | 221.0245 | 169.8171 | 129.1338 | 126.768 | 188.939 | 157.768 | 268.7561 | 300.4513 | 154 | 213.707 | 228.7561 | 225.6706 | 287.4513 | 111.6342 | 114.3658 | 126.6588 | 142.939 | 141.1583 | 127.8171 | 144.5732 | 95.3048 | 257.0974 | 220.5851 | 223.3167 | 261.8781 | 149.4513 | 132.9755 | 124.939 | 124.061 | 79.939 | 61.4877 | 84.6342 | 115 | 195.768 | 203.3048 | 211.6461 | 136.4632 | 151 | 134.6097 | 132.3658 | 116.061 | 179.8417 | 170 | 178.6215 | 183.2074 | 142.9636 | 145 | 172.5851 | 155.2684 | 166 | 184.1829 | 172.4022 | 230.7561 | 180.5732 | 161.7561 | 234.3539 | 220.6952 | 190.4513 | 195.5732 | 198.9026 | 278.2074 | 169.6952 | 161.8781 | 167.3294 | 203.1219 | 150.1465 | 142.8171 | 122.7316 | 128 | 184.0245 | 157.3412 | 172.5123 | 129.6952 | 122.4513 | 115.6342 | 102.061 | 127.3294 | 265.8781 | 239.2803 | 271.3421 | 300.4268 | 248.0364 | 237.3048 | 215.0119 | 321.4149 | 127.7926 | 81.1465 | 74.4513 | 127.7926 | 101.5732 | 108.3903 | 92.6097 | 120.7561 | 98.4513 | 110.939 | 108.6952 | 112.3658 | 86.232 | 94.3903 | 58.6461 | 99.0855 | 125.6223 | 113.2439 | 105.5123 | 82.5123 | 224.8171 | 197.2803 | 232.2684 | 275.8535 | 176.2684 | 213.7807 | 212.4149 | 255.2074 | 271.3903 | 215 | 197.061 | 315.2803
1146 | 14.6004 | 163.6593 | 208.143 | 226.6558 | 183.7802 | 151.6337 | 159.8279 | 181.7546 | 155.0256 | 152.8279 | 111.6116 | 113.2454 | 159.8535 | 225.4872 | 136.1721 | 150 | 205.5605 | 169.564 | 161.1209 | 153.5605 | 188.0733 | 183.0953 | 268.2966 | 291.1942 | 153.3407 | 238.6081 | 224.6337 | 236.3663 | 279.3663 | 107.3919 | 119.4872 | 107.8314 | 141.1942 | 159.0733 | 126.2454 | 138.2198 | 98.6814 | 272.9011 | 244.4872 | 255.7767 | 260 | 141.0733 | 146.7546 | 124.7326 | 124.9267 | 79 | 69.1465 | 78.4139 | 116.1721 | 222.9267 | 206.7546 | 234.5093 | 156.9488 | 151.293 | 145.0989 | 138.3663 | 117.5861 | 163.5384 | 170.0733 | 216.6558 | 171.3663 | 129.3442 | 143.6814 | 194.0697 | 142.3151 | 166.6593 | 186.7802 | 195.3919 | 226.1209 | 173.3407 | 156.8279 | 209.5861 | 215.707 | 183.3186 | 189.8058 | 182.6593 | 268.857 | 165.7326 | 160.5861 | 156.7802 | 206.8314 | 137.9523 | 140.3663 | 134.8058 | 131.3698 | 167.2418 | 175.9488 | 166.3919 | 125.5861 | 113.707 | 109.1209 | 102.707 | 117.7326 | 264.4395 | 258.0477 | 215.7326 | 306.4139 | 262.6081 | 241.7802 | 244.4616 | 296.8535 | 139.0477 | 68.3663 | 67.3919 | 139.0477 | 96.4651 | 100.4651 | 102.5128 | 115.6081 | 90.5861 | 110.0733 | 103.9267 | 115.8756 | 60.8314 | 84.4872 | 83.5605 | 85.1465 | 89.3663 | 114.8023 | 98.5128 | 74.707 | 224.7838 | 216.2674 | 221.1465 | 289.7326 | 166.7582 | 197.857 | 188.5861 | 246.6628 | 263.3919 | 216.0989 | 198.1465 | 331.9232
1147 | 14.6114 | 170.8023 | 173.1121 | 198.0266 | 180.8289 | 148.4543 | 145.4277 | 167.1977 | 165.5191 | 136.7168 | 132.0855 | 126.348 | 157.6578 | 218.7434 | 150.059 | 147.9469 | 200.0855 | 198.407 | 150.4277 | 147.2301 | 187.6312 | 159.8555 | 311.7965 | 281.941 | 148.7641 | 220.5723 | 219.6578 | 235.8671 | 282.2035 | 125.0855 | 113.3422 | 129.5457 | 131.7699 | 159.7434 | 143.6254 | 139.3746 | 82.7699 | 257.7168 | 237.8289 | 217.8231 | 258.7168 | 143.1977 | 130.8289 | 135.1977 | 124.1711 | 78.4867 | 71.6844 | 71.5133 | 129.0324 | 219.7758 | 192.1121 | 203.941 | 133.1387 | 153.0324 | 157.289 | 143.0855 | 125.3422 | 180.348 | 169.3746 | 189.3953 | 173.7758 | 157.3746 | 127.0855 | 158.0266 | 121.7699 | 174.5723 | 184.4277 | 211.289 | 213.9734 | 165.6844 | 143.1121 | 215.6312 | 210.1179 | 198.8023 | 198.8879 | 191.2566 | 305.8555 | 173.4601 | 169.9676 | 153.1445 | 228.7168 | 150.1711 | 144.4867 | 145 | 175.3688 | 146.8821 | 154.3629 | 184.2566 | 132.4012 | 107.9469 | 99.9676 | 99.5133 | 126.5723 | 267.4336 | 245.5723 | 223.8879 | 301.2242 | 246.711 | 239.5988 | 223.2035 | 297.4809 | 127 | 72.4012 | 83.9734 | 127 | 113.4601 | 117.8023 | 108.8289 | 99.6254 | 97.4867 | 110.1445 | 104.1121 | 91.1387 | 83.1445 | 78.4277 | 77.4867 | 86.9145 | 94.4277 | 87.5988 | 103.289 | 71.4277 | 257.9469 | 210.593 | 223.8555 | 298.059 | 187.4601 | 230.9203 | 196.2566 | 292.2301 | 280.4867 | 228.1179 | 201.1121 | 295.1387
1148 | 14.6224 | 158.2935 | 184.2173 | 209.0978 | 180.2719 | 162.3367 | 150.587 | 179.5108 | 139.9784 | 124.1741 | 130.8475 | 99.0762 | 154.587 | 217.3697 | 141.2719 | 127.4676 | 203.5438 | 159.3151 | 156.2719 | 138.3151 | 175.7395 | 168.2173 | 327.7713 | 293.6633 | 184.8908 | 218.5438 | 216.4892 | 184.62 | 263.9784 | 125.6086 | 117.7827 | 113.2173 | 140.1957 | 154.8475 | 159.7281 | 123.0762 | 91.4892 | 244.1957 | 235.3151 | 243.6951 | 246.2719 | 153.6519 | 128.5108 | 146.9238 | 123.9454 | 73.7827 | 77.1411 | 77.1957 | 107.7065 | 198.1525 | 201.6519 | 213.5108 | 156.8805 | 132.587 | 138.1195 | 143.6086 | 128.9022 | 153.7611 | 152.587 | 212.9454 | 150.1957 | 139.9022 | 130.8373 | 169 | 129.6086 | 169.7065 | 188.2173 | 194.4676 | 202.3151 | 174.2719 | 156.1741 | 201 | 188.4346 | 187.7611 | 188.9568 | 193.0216 | 309.3253 | 155.0432 | 186.9886 | 146.5438 | 214.3151 | 151.6086 | 140.3697 | 143.3367 | 188.1411 | 168.9568 | 198.554 | 186.0216 | 126.587 | 87.3697 | 118.0648 | 107.0546 | 120.6303 | 242.3481 | 242.7611 | 212.6849 | 325.3914 | 265.3914 | 245.6086 | 208.0114 | 319.8908 | 127.587 | 66.9784 | 73.8805 | 127.587 | 97.8805 | 107.0546 | 109.0546 | 118.6633 | 92.9784 | 102.6633 | 112.967 | 114.0978 | 73.3151 | 82.5108 | 71.3151 | 85.3151 | 99.0978 | 94.3914 | 86.4676 | 75.6086 | 235.8043 | 245.8692 | 231.7281 | 291.7179 | 169.1411 | 197.9022 | 197.3367 | 283.6086 | 274.8043 | 206.2389 | 212.7065 | 315.6519
1149 | 14.6334 | 161.1101 | 179.2026 | 210.2203 | 193.1013 | 147.2203 | 156.4405 | 172.9075 | 149.2203 | 133.5683 | 114.5242 | 108.8899 | 160.3304 | 230.7709 | 153 | 143.2026 | 226.0088 | 156.8545 | 165.3569 | 134.8634 | 197.8062 | 165.7356 | 297.5595 | 307.7886 | 147.9912 | 239.6873 | 219.5683 | 218.881 | 273.1101 | 120.4582 | 124.2291 | 106.881 | 140.6784 | 135.9912 | 148.6608 | 135.2026 | 96.6608 | 248.8634 | 229.3304 | 209.7268 | 255.5771 | 133.2114 | 123.6696 | 136.5595 | 106.6519 | 83.7532 | 62.0925 | 78.8899 | 106.1013 | 215.9075 | 181.2114 | 209.2203 | 164.2291 | 139.3216 | 130.881 | 138.0177 | 126.6784 | 167.3569 | 157.0088 | 194.5506 | 151.8899 | 136.7974 | 152.7049 | 166.7974 | 126.6608 | 168.1013 | 180.3392 | 210.7532 | 198.533 | 183.3569 | 167.2203 | 203.7532 | 178.6255 | 203.119 | 204.3569 | 185.3216 | 259.6431 | 141.0748 | 150.5506 | 169.8899 | 210.4229 | 146.6784 | 153.3304 | 128.1101 | 169.6784 | 185.7886 | 176.1101 | 177.4405 | 131.4494 | 102.0925 | 92.3216 | 125.0088 | 109.3216 | 262.1278 | 254.9252 | 219.4405 | 329.7709 | 265.3657 | 238.6961 | 243.1278 | 286.6255 | 133.1101 | 73.6873 | 80.6784 | 133.1101 | 106.1101 | 123.467 | 123.7049 | 131.7974 | 102.5506 | 116.5683 | 87.9735 | 115.881 | 68.3216 | 77.6696 | 67.0925 | 80.8722 | 100.5506 | 97.4494 | 97.1366 | 72.2203 | 235.1013 | 218.6519 | 217.5771 | 323.0265 | 155.1938 | 194.0265 | 186.5153 | 277.3569 | 271.6784 | 192.7532 | 207.7974 | 294.6608
1150 | 14.6444 | 165.795 | 196.0206 | 211.6327 | 197.5928 | 147.6534 | 161.4691 | 154.7758 | 156.1417 | 126.2847 | 139.9395 | 109.4691 | 160.9188 | 233.0825 | 152.3879 | 158.1843 | 215.1843 | 185.674 | 143.2655 | 155.4292 | 174.0812 | 200.0825 | 292.8982 | 296.0413 | 157.2242 | 222.4691 | 209.4691 | 223.4292 | 275.8363 | 108.8569 | 117.0206 | 117.0605 | 129.6327 | 146.326 | 150.4498 | 151.1637 | 99.7964 | 270.7758 | 231.0206 | 253.7964 | 237.1018 | 143.2448 | 120.5103 | 132.1431 | 121.4691 | 103.674 | 77.4085 | 77.8776 | 115 | 199.3673 | 192.2242 | 212.5915 | 155.7964 | 147.7964 | 135.7964 | 122.2448 | 115.8776 | 148.4484 | 150.9587 | 198.2655 | 152.4691 | 124.1417 | 120.8157 | 149.6121 | 133.5915 | 174.4292 | 175.7345 | 233.3673 | 221.0812 | 165.0605 | 168.8776 | 224.4085 | 212.0825 | 195.1431 | 186.795 | 193.7964 | 286.9986 | 172.7758 | 153.8982 | 168.8776 | 225.8377 | 135.3879 | 152.0825 | 129 | 158.0206 | 176.9794 | 175.7758 | 178.4292 | 126.3879 | 119 | 102.6327 | 117 | 119.0206 | 248.3466 | 222.8569 | 223 | 337.8363 | 236.3673 | 216.1623 | 229.5915 | 319.4704 | 132.6534 | 56.7139 | 70.2448 | 132.6534 | 108.1018 | 103.7345 | 92.6726 | 115.8363 | 106.3879 | 104.2655 | 111.6327 | 118.8377 | 79.3673 | 73.8982 | 84.9794 | 95.3879 | 103.7758 | 94.2242 | 77.1829 | 74.1224 | 246.0812 | 233.1018 | 200.2036 | 297.0413 | 178.0619 | 173.1829 | 218.9601 | 260.5296 | 262.9587 | 211.5722 | 192.6933 | 304.5296
1151 | 14.6554 | 188.1504 | 190.4818 | 207.7876 | 166 | 139.7513 | 169.5752 | 147.6165 | 187.6115 | 153.6322 | 116.2487 | 119.4612 | 146.1347 | 208.3678 | 148.1347 | 137.057 | 196.3471 | 171.4248 | 138.8083 | 140.9066 | 190.2124 | 170.9223 | 289.5802 | 283.2124 | 168.4248 | 227.2074 | 217.8446 | 207.829 | 291.829 | 119.8496 | 112.6942 | 139.6736 | 126.8653 | 162.057 | 128.7719 | 137.1917 | 85.3471 | 257.4198 | 224.2694 | 235.8446 | 240.5545 | 143.6529 | 117.5388 | 125.057 | 131.4612 | 84.171 | 67.886 | 77 | 118.3678 | 205.3678 | 198.171 | 225.6165 | 143.9066 | 132.943 | 122.9637 | 124.4041 | 116.886 | 167.7513 | 161.7669 | 195.5595 | 162.1917 | 158.3058 | 141.6736 | 155.2124 | 142.5752 | 159.3678 | 183.1554 | 231.5545 | 237.0777 | 182.4198 | 166.6529 | 212.3264 | 187.7719 | 189.2694 | 209.6892 | 178.8083 | 327.4405 | 163.8653 | 144.6529 | 167.4612 | 199.0984 | 133.4248 | 126.9637 | 130.3471 | 151.8083 | 180.6322 | 166.3264 | 162.6942 | 122.8083 | 117.5182 | 101.886 | 117.5388 | 115.7719 | 255.7099 | 229.6736 | 225.0207 | 347.9016 | 240.4818 | 248.1554 | 245.1761 | 287.2694 | 123.5388 | 70.0777 | 70.3835 | 123.5388 | 116.4041 | 106.4405 | 115.3628 | 127.2487 | 99.9793 | 99.1347 | 106.9793 | 91.2901 | 81.3264 | 65.4612 | 92.1347 | 90.3264 | 94.7306 | 98.9586 | 99.4612 | 73.6529 | 257.3628 | 243.1554 | 215.7306 | 287.3108 | 161.0414 | 200.3108 | 189.7306 | 291.3835 | 280.7719 | 194.948 | 210.6529 | 333.0934
1152 | 14.6664 | 155.383 | 197.2069 | 200.441 | 167.764 | 153.058 | 153.265 | 159.911 | 152.265 | 134.205 | 108.6751 | 117.176 | 149.058 | 223.3849 | 148.706 | 134.2341 | 205.588 | 186.853 | 141.942 | 163.5009 | 198 | 163.559 | 315.4719 | 289.383 | 180.7659 | 197.0561 | 207.8221 | 227.323 | 304.8549 | 147.6189 | 127.942 | 137.9729 | 132.0271 | 154.971 | 147.7079 | 130.677 | 92.824 | 233.1451 | 231.1451 | 222.294 | 215.8221 | 139.5571 | 122.176 | 119.294 | 129.029 | 69.8221 | 80 | 74.5009 | 135.5899 | 224.795 | 181.764 | 233.0599 | 168.412 | 139.588 | 138.3249 | 123.1779 | 127.383 | 177.089 | 144.5281 | 211.559 | 158.176 | 140.265 | 143.648 | 158.5899 | 128.029 | 179.383 | 196.265 | 204.911 | 245.029 | 153.588 | 157.853 | 209.323 | 208.1779 | 193.4991 | 186.1161 | 185.176 | 311 | 166.6751 | 142.1741 | 164.735 | 217.089 | 148.706 | 141.0019 | 138.118 | 155.824 | 159 | 162.882 | 179.853 | 125.2069 | 107.706 | 113.265 | 119.089 | 140.882 | 239.971 | 237.9691 | 238 | 329.764 | 253.6751 | 256.7079 | 265.5319 | 291.4991 | 128.029 | 75.971 | 59.853 | 128.029 | 113.8549 | 94.1161 | 94.205 | 115.412 | 88.617 | 100.882 | 96.205 | 106.294 | 77.735 | 62 | 89.4719 | 86.735 | 89.7659 | 76.3811 | 96.441 | 67.6461 | 240.1741 | 248.0328 | 212.824 | 311.1779 | 184.0599 | 222.1199 | 188.588 | 273.7968 | 301.0309 | 240.5609 | 202.441 | 306.971
1153 | 14.6774 | 144.2478 | 183.708 | 201.8849 | 178 | 159.1061 | 146.292 | 171.4779 | 146.8849 | 146.3628 | 124.584 | 118.9026 | 158.1327 | 189.7788 | 147.1593 | 151.0885 | 209.4779 | 188.5487 | 136.6904 | 150.2744 | 194.9734 | 166.7346 | 295.4779 | 280.7965 | 162.6372 | 236.7256 | 229.0442 | 232.9292 | 276.9115 | 125.1327 | 119.6638 | 107.1416 | 163.4955 | 152.3452 | 119.0708 | 123.4779 | 86.1593 | 260.6106 | 260.5221 | 230.0532 | 241.3452 | 163.292 | 127.8849 | 123.3894 | 133.4071 | 88.3362 | 77.4513 | 59.8407 | 112.5487 | 214.3894 | 192.6372 | 212.8141 | 165.7965 | 141.7256 | 124.1682 | 103.2301 | 119.2744 | 167.7522 | 166.5929 | 207.8849 | 164.5221 | 135.3628 | 129.2035 | 134.9115 | 133.2035 | 168.0885 | 192.6372 | 212.3362 | 244.1504 | 156.6814 | 154.9292 | 217.1593 | 188.0708 | 205.6106 | 205.8584 | 192.6372 | 307.6548 | 189.7522 | 177.8407 | 165.3362 | 206.4779 | 144.9292 | 115.0266 | 132.0442 | 147.8849 | 155.8141 | 168.7965 | 181.2301 | 111.8673 | 105.3628 | 108.3628 | 111.1859 | 139.7875 | 243.8762 | 258.1327 | 238.3186 | 337.7699 | 273.0885 | 227.4336 | 232.4602 | 304.2478 | 132.8849 | 70.4779 | 60.1151 | 132.8849 | 84.7965 | 116.7256 | 106.6814 | 113.7522 | 99.9115 | 104.4071 | 106.6106 | 108.3186 | 77.6991 | 61.6814 | 68.3628 | 92.2744 | 79.6019 | 96.7965 | 97.2478 | 81.5664 | 267.7167 | 199.5045 | 208.3894 | 288.6814 | 171.3009 | 187.8407 | 192.4779 | 237.416 | 277.0442 | 207.4071 | 208.0266 | 303.3894
1154 | 14.6884 | 137.5738 | 197.8284 | 198.059 | 178.3432 | 135.035 | 135.2601 | 168.6809 | 145.118 | 146.7749 | 97.5443 | 96.118 | 146.2601 | 215.5148 | 146.6273 | 143.059 | 214.5738 | 198.7399 | 165.9945 | 159.2306 | 176.941 | 191.6809 | 294.0535 | 285.8579 | 167.3727 | 230.8579 | 220.7399 | 218.5683 | 285.1126 | 110.8579 | 133.3377 | 127.7989 | 144.8284 | 165.5388 | 129.6568 | 126.8579 | 87 | 252.6328 | 253.5388 | 254.965 | 251.6218 | 150.3727 | 124.059 | 133.4262 | 125.1716 | 71.2306 | 66.2306 | 64.6623 | 124.1126 | 223.5683 | 195.3137 | 202.0055 | 164.6809 | 135.3432 | 155.9945 | 118.0885 | 122.7399 | 179.059 | 174.1421 | 200.1126 | 158.2251 | 136.6328 | 129.118 | 146.8874 | 128.6568 | 161.2601 | 197.5443 | 194.0295 | 207.8284 | 156.7159 | 146.2306 | 218.1716 | 196.5978 | 193 | 186.5738 | 190.3377 | 293.2601 | 187.847 | 178.7159 | 148.059 | 212.9465 | 133.4852 | 128.7694 | 130.9465 | 144.2306 | 143.118 | 173.6863 | 193.3432 | 126.1421 | 103.2011 | 106.2011 | 124.3967 | 108.7159 | 292.2251 | 242.3137 | 235.1956 | 326.6863 | 263.1716 | 235.5148 | 253.0831 | 301.0055 | 130.6033 | 71.9705 | 65.6568 | 130.6033 | 92.2601 | 110.6863 | 107.5738 | 122.1421 | 110.8579 | 97.2011 | 97.0885 | 109.1421 | 57.5738 | 61.7159 | 65.1716 | 99.6863 | 117.4503 | 101.1716 | 88.8579 | 71.7694 | 230.0055 | 213.1956 | 220.4852 | 288.8874 | 191.0831 | 182.882 | 196.2011 | 262.9115 | 271.4852 | 197.6273 | 219.7104 | 316.6863
1155 | 14.6994 | 154.413 | 197 | 203.0353 | 176.3225 | 165.8741 | 154.4936 | 143.9194 | 168.4936 | 158.3805 | 108.8614 | 118.3904 | 155.7482 | 219.6549 | 142.3904 | 151.8967 | 225.7129 | 175.8288 | 137.7355 | 166.6902 | 167.9194 | 166.5516 | 278.4936 | 284.2998 | 170.5064 | 234.2645 | 208.1259 | 208.0679 | 273.7935 | 117.7581 | 109.8614 | 117.9547 | 145.6549 | 149.3324 | 130.0226 | 133.7581 | 91.2292 | 275.5516 | 234.942 | 228.355 | 224.4583 | 157.7355 | 133.6322 | 120.8161 | 123.9774 | 79.9774 | 77.1839 | 88.1386 | 119.7808 | 201.2998 | 189.4257 | 229.8967 | 140.8388 | 136.4484 | 135.0906 | 132.4484 | 106.2645 | 193.5969 | 167.7935 | 192.471 | 147.7228 | 155.6902 | 146.4257 | 155.8161 | 127.1839 | 176.4484 | 206.6549 | 196.7935 | 209.0226 | 164.6322 | 156.0806 | 218.0806 | 185.7355 | 194.471 | 197.3451 | 166.3098 | 309.9194 | 141.5389 | 185.7129 | 158.1839 | 231.4484 | 131.5516 | 121.4936 | 147.9774 | 152.058 | 166.8614 | 173.5064 | 194.2645 | 121.6322 | 109 | 111.8161 | 110.1259 | 117.1839 | 268.116 | 237.8967 | 209.7935 | 333.8614 | 265.6549 | 243.7002 | 240.2871 | 325.4031 | 144.8741 | 65.7129 | 65.8388 | 144.8741 | 112.413 | 112.3451 | 115.0353 | 119.8388 | 114.8161 | 103.9194 | 110.7129 | 105.1839 | 69.0806 | 71.1033 | 69.4936 | 99.6902 | 87.2518 | 100.7129 | 94.1033 | 66.3324 | 261.3904 | 193.8614 | 222.9647 | 296.529 | 173.1386 | 170.7228 | 203.1033 | 253.7002 | 269.5516 | 193.942 | 200.2871 | 321.2871
1156 | 14.7104 | 170.469 | 195.6268 | 181.7847 | 162.9425 | 152.7847 | 163.3303 | 152.9042 | 180.6652 | 134.2965 | 123.2537 | 125.8613 | 139.3777 | 222.292 | 153 | 144.0575 | 216.6652 | 151.062 | 141.3923 | 151.6077 | 172.5885 | 167.4307 | 294.7847 | 266.062 | 152.2537 | 229.6268 | 218.0767 | 236.7655 | 268.1195 | 130.7847 | 125.0383 | 101.9617 | 149.469 | 167.115 | 142.1387 | 144.6268 | 107.0575 | 270.5456 | 248 | 256.6843 | 262.3923 | 158.0575 | 134.1578 | 124.5118 | 114.0192 | 70.2153 | 77.4115 | 70.2153 | 136.4882 | 181.885 | 181.7272 | 223.4307 | 151.354 | 135.3732 | 164.3303 | 132.7463 | 104.3732 | 205.761 | 157.2153 | 196.646 | 173.9571 | 143.9425 | 135 | 157.7463 | 132.5118 | 174 | 207.3303 | 189.1578 | 214.2728 | 158.4882 | 148.469 | 212.0383 | 192.9233 | 201 | 186.0767 | 183.9425 | 309.8805 | 175.7847 | 173.7226 | 158.8038 | 229.7847 | 132.6268 | 140.3348 | 140.9617 | 143.708 | 183.2153 | 163.4927 | 190.0192 | 122.9425 | 110.5693 | 107.6652 | 122.2345 | 119.3732 | 251.2199 | 231.823 | 208.8275 | 348.0575 | 270.2537 | 263.8997 | 242.8613 | 300.8805 | 130.6077 | 64.9042 | 75.7655 | 130.6077 | 128.6652 | 104.8038 | 98.6889 | 125.8422 | 110.6652 | 108.5885 | 101.2728 | 106.5885 | 66.9617 | 74.4307 | 81.6652 | 84.2153 | 110.4882 | 91.6652 | 98.6077 | 84.5073 | 266.8997 | 203.938 | 243.823 | 287.646 | 157.3732 | 198.9188 | 206.0383 | 279.1962 | 272.3923 | 205.823 | 210.9042 | 326.6077
1157 | 14.7214 | 153.29 | 190.4169 | 190.8368 | 173.7492 | 153.0816 | 146.8398 | 174.0423 | 165.5408 | 157.8761 | 112.8338 | 115.5015 | 179.9547 | 210.713 | 148.2055 | 135.5438 | 205.0846 | 175.4139 | 139.4562 | 150.4169 | 172.4985 | 160.1662 | 299.8761 | 291.2477 | 143.5015 | 228.1692 | 207.6284 | 242.3323 | 287.293 | 130.4562 | 119.9184 | 109.3746 | 133.3323 | 141.1239 | 153.3323 | 139.8338 | 93.1239 | 250.6314 | 247.5831 | 246.0876 | 257.7462 | 142.0393 | 141.3323 | 143 | 115.2115 | 69.5015 | 76.6677 | 68.4592 | 118.4169 | 209.5438 | 196.1662 | 216.1662 | 163.0846 | 140.7915 | 137.6254 | 139.4139 | 106.6647 | 171.5831 | 155.8761 | 190.3353 | 140.7915 | 152.2477 | 139.3776 | 162.3293 | 149.3323 | 175.2508 | 178.3323 | 196.5408 | 201.0846 | 143.1269 | 136.0846 | 205.8761 | 207.5408 | 203.5015 | 172.71 | 196.6254 | 299.6707 | 175.8731 | 146.5408 | 159.8761 | 232.3746 | 129.9184 | 151.707 | 143.997 | 159.8731 | 177.9154 | 188.9184 | 179.3293 | 131.8731 | 117.6254 | 96.9184 | 111.6647 | 121.0393 | 283.7885 | 229.5015 | 238.7462 | 334.9577 | 254.8308 | 240.9154 | 232.293 | 289.6284 | 137.1299 | 84.3746 | 81.1239 | 137.1299 | 113.5408 | 104.4169 | 122.6254 | 120.9607 | 101.3776 | 108.9154 | 90.7945 | 113.1692 | 75 | 69.4592 | 64.8731 | 82.8761 | 96.1692 | 82.5861 | 96.1662 | 66.1692 | 248.71 | 186.4652 | 236.707 | 286.3383 | 163.4169 | 174.4169 | 196.7492 | 280.2085 | 271.707 | 200.3746 | 221.6194 | 321.2477
1158 | 14.7324 | 146.7527 | 188.2473 | 214.3245 | 168.3378 | 142.6622 | 180.2738 | 170.936 | 157.7925 | 157.4945 | 121.7395 | 119.702 | 155.4548 | 227.4173 | 133.9735 | 149.6622 | 211.0133 | 166.3885 | 131.7395 | 143.6115 | 173.596 | 158.5452 | 302.585 | 277.6622 | 156.532 | 238.0398 | 226.351 | 237.5452 | 283.1965 | 116.3378 | 126.3643 | 108.3245 | 128.9867 | 137.3112 | 147.883 | 141.0133 | 86 | 285.9228 | 242.0265 | 266.5983 | 238.8565 | 132.5187 | 130.585 | 140.1303 | 132.7925 | 77.8963 | 83.6622 | 74.883 | 117.5718 | 217.702 | 190.7925 | 209.2473 | 165.9228 | 139.5585 | 138.6755 | 122.883 | 103.2715 | 171.1037 | 159.6888 | 199.0265 | 141.766 | 143.0772 | 152.9095 | 146.0507 | 138.585 | 179.3378 | 176.1942 | 192.9867 | 210.9867 | 156.5452 | 146.4282 | 209.9095 | 204.4282 | 213 | 186.7527 | 201.2075 | 327.649 | 163.3378 | 145.8565 | 162.3643 | 229.5585 | 137.9095 | 143 | 129.6357 | 143.8058 | 168.4548 | 190.5078 | 157.9735 | 127.2847 | 122.4282 | 109.766 | 101.2075 | 106 | 259.585 | 244.298 | 219.6357 | 330.3378 | 240.3245 | 238.5187 | 244.3112 | 308.1303 | 160.4945 | 80.0133 | 71.5718 | 160.4945 | 106.4548 | 107.1037 | 119.4813 | 136.2207 | 112.2605 | 102.5452 | 105.0265 | 126.5718 | 76.766 | 75.2207 | 54.3245 | 90.4415 | 109.7925 | 95.3643 | 92.5585 | 79.7925 | 250.8323 | 232.649 | 221.1568 | 309.8323 | 162.7925 | 177.766 | 192.6622 | 282.3245 | 263.6622 | 197.5585 | 193.2847 | 313.181
1159 | 14.7434 | 156.9705 | 180.593 | 212.4749 | 168.7965 | 145.233 | 153 | 146.9322 | 148.1357 | 139.1564 | 136.174 | 109.1564 | 151.3982 | 185.5634 | 147.767 | 150.1357 | 208.4278 | 187.3775 | 149.4366 | 120.2921 | 203.3097 | 162.6018 | 285.1357 | 281.1652 | 168.767 | 212.1062 | 217.9322 | 232.0471 | 239.6313 | 112.3687 | 111.8643 | 113.4661 | 132.7374 | 148.0678 | 151 | 131.9026 | 87.3982 | 261.0088 | 231.2626 | 246.7965 | 248.0766 | 146.407 | 118.2626 | 128.3687 | 134.5547 | 74.233 | 80.407 | 81.4956 | 109.9322 | 201.7965 | 180.4366 | 197.3982 | 143.8053 | 139.1652 | 130.7374 | 121.1062 | 136.4661 | 178.7286 | 150.7965 | 185.6991 | 153.1269 | 154.4749 | 137.3392 | 171.2418 | 126.7286 | 171.407 | 189.6018 | 205.593 | 221.9617 | 163.3982 | 154.767 | 199.6991 | 213.233 | 207.8731 | 199.767 | 205.2714 | 326.5132 | 163.3304 | 162.3009 | 145.767 | 230.5634 | 130.4956 | 139.7374 | 142.0383 | 133.8938 | 160.2035 | 161.6608 | 175.4956 | 145.5427 | 126.8053 | 113.2035 | 108.767 | 108.5634 | 248.4278 | 257.407 | 229.9409 | 328.233 | 249.6608 | 254.2714 | 256.233 | 291.4749 | 132.6018 | 76.7286 | 63 | 132.6018 | 102.1652 | 107.0383 | 101.3982 | 131.8731 | 89.2035 | 108 | 91.9322 | 120.0974 | 82.3009 | 74.6018 | 59.233 | 93.1652 | 98.0383 | 82.4956 | 92.3982 | 73.6313 | 225.9617 | 239.9026 | 193.7374 | 279.3687 | 156.1652 | 190.2921 | 193.6018 | 291.8938 | 276.2538 | 200.8938 | 218.5339 | 335.233
1160 | 14.7544 | 153.624 | 199.2453 | 189.7733 | 175.5467 | 147.472 | 157.9066 | 151.9627 | 144.4533 | 161.0934 | 118.9627 | 129.6214 | 154.7733 | 195.2267 | 149.6986 | 145.7173 | 223.2453 | 162.9253 | 141 | 142.472 | 183.3973 | 154.8107 | 286.1146 | 279.848 | 164.8107 | 199.68 | 215.6027 | 199.9253 | 259.2292 | 109.6986 | 114.8107 | 114.5093 | 124.6986 | 147.6986 | 155.6613 | 127.944 | 94.208 | 251.0026 | 246.6613 | 256.4907 | 232.5493 | 132.6613 | 127.528 | 123 | 154.9601 | 74.7547 | 66.416 | 86.376 | 115.6986 | 204.6214 | 180.8319 | 208.624 | 139.8133 | 143.9813 | 120.4907 | 108.9253 | 135.3014 | 185.3574 | 158.0373 | 191.9253 | 168.7733 | 135.208 | 127.6426 | 148.0187 | 139.4907 | 165.0212 | 181.8107 | 218.848 | 226.6587 | 161.6214 | 159.3973 | 202.2453 | 213.7547 | 196.888 | 198.0187 | 193.9813 | 293.9627 | 170.0187 | 166.8693 | 142.792 | 239.4907 | 143.9627 | 133.9066 | 136.8506 | 158.3413 | 157.3786 | 171.848 | 178.4134 | 123.4533 | 113.9813 | 106.944 | 106.528 | 120.4346 | 258.3387 | 240.2267 | 218.8319 | 332.9253 | 260.0934 | 241.5093 | 253.5654 | 283.0026 | 127.7547 | 86.056 | 62.7547 | 127.7547 | 104.2827 | 100.3786 | 107.2267 | 115.9813 | 82.944 | 106.7733 | 97.4533 | 126.5093 | 76.0747 | 69.0187 | 61.472 | 92.0934 | 89.6613 | 94.4907 | 93.32 | 77.0373 | 239.4907 | 228.0934 | 193.0586 | 284.7946 | 167.1146 | 202.1681 | 188.264 | 297.4507 | 301.9041 | 208.9041 | 215.528 | 336.736
1161 | 14.7654 | 164.468 | 201.2881 | 191.6677 | 171.8643 | 154.8338 | 169.6509 | 164.6982 | 158.4406 | 151.1525 | 127.3186 | 108.7119 | 149.9695 | 203.8948 | 158.2576 | 151.5152 | 220.9085 | 175.5152 | 142.5457 | 144.939 | 194.3323 | 151.6982 | 304.9085 | 271.2134 | 159.122 | 212.8034 | 202.6067 | 216.6372 | 295.0137 | 122.8948 | 104.2271 | 114.5457 | 138.4101 | 155.4848 | 158.9222 | 133.1357 | 94.8171 | 279.1494 | 261 | 255.4238 | 253.9222 | 145.7119 | 123.5152 | 124.5457 | 124.3491 | 74.5152 | 79.7424 | 71.6677 | 118.0747 | 189.8948 | 207.7424 | 230.5457 | 171.2881 | 147 | 126.122 | 119.7119 | 124.6814 | 174.1829 | 155.6067 | 204.5152 | 164.2271 | 140.7119 | 146.0305 | 147.3186 | 139.7119 | 195.8171 | 175.6067 | 200.4238 | 213.8811 | 145.8643 | 164.6646 | 208.6677 | 214.8034 | 216.5457 | 193.7119 | 198.8034 | 290.9832 | 172.4101 | 186.4543 | 148.3659 | 233.7866 | 150.7729 | 149 | 155.3323 | 177.0442 | 179.061 | 159.3491 | 156.8948 | 128.4238 | 117.7729 | 115.7424 | 106.122 | 125.8475 | 253.0168 | 244.5152 | 243.4238 | 349.8948 | 245 | 240.5152 | 240.4238 | 322.4848 | 129.3186 | 73.3933 | 65.6067 | 129.3186 | 97.9695 | 117.4238 | 108.4238 | 123.8948 | 97.9253 | 100.939 | 105.7729 | 115.7256 | 71.7287 | 69.6067 | 66.7729 | 81.3796 | 103.7424 | 98.8338 | 86.122 | 72.2881 | 227.8612 | 212.7424 | 222.9695 | 307.6951 | 186.1662 | 181.4406 | 186 | 272.7287 | 279.5762 | 188.3796 | 214.8643 | 344.1525
1162 | 14.7764 | 142.3983 | 203.3805 | 200.7123 | 175.7123 | 160.8407 | 153.5221 | 178.3318 | 169.1725 | 160.0619 | 126.4424 | 108.239 | 149.4292 | 211.792 | 145.2344 | 155.4292 | 210.6504 | 179.969 | 145.1106 | 139.2699 | 176.1106 | 164.792 | 301.9381 | 300.0974 | 175.8584 | 225.5576 | 213.2699 | 230 | 276.5885 | 130.5221 | 108.4779 | 124.9381 | 145.0929 | 150.4911 | 137.2567 | 127.9381 | 81.0487 | 252.5576 | 256.9513 | 242.6017 | 225.5089 | 136.8716 | 128.5089 | 128.7301 | 130.4911 | 77.6195 | 80.3496 | 81.2522 | 109.9071 | 202.3805 | 205.9203 | 235.2699 | 171.2212 | 142.9513 | 133.6814 | 116.5398 | 117.4602 | 190.9203 | 168.6991 | 204.6504 | 158.4911 | 139.6991 | 144.4115 | 160.208 | 135.4602 | 175.8407 | 183.5708 | 200.0178 | 248.8407 | 160.7788 | 140.9071 | 227.6991 | 212.7123 | 214.5221 | 196.4779 | 196.9822 | 276.3364 | 184.2212 | 171.7433 | 177.5221 | 214.6195 | 151.1106 | 146.031 | 141.969 | 160.7788 | 179.8716 | 172.5089 | 169.9203 | 124.239 | 117.3009 | 114.4602 | 120.6991 | 115.5885 | 269.553 | 237.9026 | 236.5398 | 359.4115 | 245.2699 | 248.4779 | 234.0797 | 314.792 | 132.4911 | 66.5089 | 75.4602 | 132.4911 | 95.2699 | 109.7301 | 102.0797 | 130.7123 | 111.2212 | 106.6063 | 110.4292 | 88.1903 | 88.0619 | 82.6991 | 67.9203 | 94.5398 | 103 | 103.761 | 99.3496 | 79.7788 | 211.0533 | 208.761 | 217.3009 | 281.239 | 176.1106 | 201.3496 | 189.7788 | 280.4247 | 284.5708 | 200.1903 | 221.6814 | 345.7742
1163 | 14.7874 | 156.4356 | 203.5153 | 189.7483 | 159.3865 | 150.2025 | 142.7729 | 160.7424 | 147.1534 | 145.6932 | 117.006 | 117.2822 | 157.6932 | 201.7729 | 117.595 | 158.0491 | 207.2822 | 183.2025 | 147.6195 | 142.2576 | 179.1839 | 152.2331 | 308.1228 | 315.2271 | 181.6627 | 239.2271 | 222.1839 | 224.3559 | 284.7669 | 114.411 | 121.7669 | 134.2271 | 124.3865 | 142.411 | 151.7364 | 140.0491 | 92.5644 | 262.5584 | 245.7178 | 223.2025 | 235 | 130.0551 | 135.7424 | 132.2331 | 123.822 | 78.0491 | 82.589 | 64.9754 | 127.5093 | 194.0491 | 209.2085 | 230.3559 | 157.8957 | 132.5644 | 119.4602 | 126.1839 | 118.822 | 180.6627 | 167.2517 | 209.1839 | 154.3619 | 150.6687 | 134.822 | 169.9449 | 137.1043 | 169.1534 | 176.7178 | 214.4847 | 236.7915 | 161.405 | 165.8466 | 229.3559 | 193 | 197.5644 | 206.3805 | 178.8466 | 300.7915 | 175.9754 | 147.1043 | 161.1288 | 217.8712 | 143.4602 | 137.4356 | 155.3619 | 165.6381 | 168.822 | 182.5644 | 170.9509 | 134.6932 | 113.6687 | 111.8712 | 126.3068 | 129.6932 | 254.4416 | 215.4356 | 237.4356 | 337.6872 | 249.6687 | 252.4542 | 237 | 304.4907 | 129.7729 | 74.8712 | 73.7178 | 129.7729 | 99.6687 | 100.8161 | 107.2576 | 108.4602 | 104.6687 | 130.6627 | 111.0737 | 90.7975 | 69.178 | 84.3559 | 65.2822 | 96.5644 | 97.6381 | 98.6687 | 103.5644 | 89.7178 | 256.7118 | 208.4661 | 211.6932 | 296.4907 | 177.7729 | 208.1043 | 203.6687 | 248.3373 | 287.3128 | 206.4661 | 211.411 | 317.0551
1164 | 14.7984 | 158.8284 | 214.7055 | 214.945 | 165.055 | 143.5954 | 159.1165 | 154.411 | 148.8284 | 147.9385 | 136.1651 | 122.4174 | 161.411 | 213.1101 | 140.822 | 149.178 | 213.0064 | 177.7734 | 159.6376 | 143.8771 | 199.1165 | 156.2266 | 287.4174 | 293.5826 | 163.589 | 227.5954 | 236.5211 | 210.2945 | 281.3624 | 122.1229 | 105.6991 | 116.411 | 134.4725 | 141.5826 | 118.589 | 136.1844 | 94.589 | 236.8835 | 245.2945 | 220.1294 | 239.4174 | 139.7541 | 131.4725 | 134.7541 | 125.4046 | 72.1229 | 76.3495 | 64.6991 | 118.3624 | 195.1908 | 214.1037 | 220.4174 | 146.1716 | 139.8899 | 123.8963 | 140.8156 | 118.3431 | 162 | 146.178 | 221.4596 | 172.4174 | 151.1651 | 140.5275 | 156.4174 | 139.4046 | 173.7734 | 180.7119 | 202.1165 | 221 | 147.8963 | 166.5275 | 212.9385 | 189.7605 | 195.1716 | 184.644 | 185.1229 | 281.4661 | 175.4046 | 158.5339 | 165.4725 | 224.7184 | 145.2459 | 142.4789 | 157.2202 | 148.4661 | 176.589 | 173.6927 | 179.233 | 135.7605 | 121.233 | 104.8771 | 119.644 | 137.2395 | 275.4046 | 222.8349 | 230.1101 | 328.7248 | 254.5826 | 226.5404 | 226.3982 | 318.644 | 136.6927 | 74.0615 | 77.4174 | 136.6927 | 110.4725 | 91.4853 | 109.4661 | 108.1844 | 111.3495 | 106.9936 | 96.2945 | 103 | 62.589 | 74.4174 | 67.4725 | 91.5211 | 89.3009 | 101.5211 | 103.822 | 89.589 | 227.7055 | 226.1716 | 215.1165 | 309.1716 | 189.1101 | 213.644 | 207.9936 | 265.233 | 301.5761 | 222.9936 | 212.055 | 326.1651
1165 | 14.8094 | 161.2507 | 208.4778 | 217.0914 | 169.6579 | 150.5575 | 158.8407 | 152.3864 | 159.8407 | 146.0679 | 115.3068 | 125.6371 | 152.0235 | 203.6018 | 145.9765 | 154.4543 | 216.41 | 185.7168 | 138.9321 | 141.0561 | 194.5457 | 140 | 301.9882 | 284.2271 | 166.2271 | 234.2507 | 221 | 203.3303 | 294.3982 | 122.1828 | 105.2715 | 118.9882 | 136.1593 | 137.1357 | 126.1357 | 147.9321 | 93.8525 | 241.7611 | 251.5222 | 241.0796 | 257.6697 | 129.2036 | 133.7729 | 118.3747 | 118.7493 | 78.6253 | 70 | 62.7375 | 136 | 216.7846 | 194.7493 | 223.3303 | 144.6697 | 150.6253 | 141.7168 | 125.0796 | 104.9086 | 170.2832 | 152.3747 | 201 | 178.0914 | 129.0796 | 128.7168 | 165.1593 | 133.9765 | 180.4896 | 187.0914 | 194.7846 | 219.4661 | 169.3982 | 156.2507 | 208.3068 | 185.9882 | 191.5222 | 188.2036 | 203.59 | 274.5339 | 161.0796 | 169.7611 | 160.41 | 242.1711 | 151.8082 | 153.3982 | 135.4425 | 145.5222 | 178.3068 | 151.7611 | 175.9204 | 132.9086 | 112.705 | 95.9204 | 111.2389 | 137.0235 | 254.3303 | 245.8407 | 223.9765 | 358.4778 | 240.3186 | 256.5104 | 220.9411 | 313.6136 | 122.124 | 75.6253 | 83.705 | 122.124 | 110.6253 | 116.5457 | 104.9882 | 117.1711 | 108.9882 | 99.9086 | 99.4543 | 110.9765 | 67.9882 | 80.705 | 72.5339 | 80.6018 | 103.8407 | 84.1593 | 97.0118 | 90.0796 | 227.9204 | 209.9439 | 226.1918 | 296.0118 | 173.7729 | 214.1357 | 191.3982 | 260.3864 | 294.3393 | 212.2271 | 212.3629 | 308.3747
1166 | 14.8204 | 142.4901 | 193.7665 | 209.1908 | 200.7862 | 143.3388 | 160.7665 | 147.1711 | 160.4901 | 157.7862 | 115.0428 | 115.7006 | 135.5954 | 207.6184 | 157.2994 | 164.7862 | 201.4671 | 170.1908 | 139.0197 | 146.7895 | 184.852 | 149.2533 | 303.6612 | 294.6579 | 161.9803 | 225.0625 | 221.3191 | 200.9967 | 285.5954 | 109.2533 | 130.7665 | 120.9803 | 134.8717 | 150.0428 | 140.6382 | 146.4243 | 90.5954 | 257.5723 | 253.4704 | 236.4474 | 268.9375 | 139.4704 | 128.1283 | 126 | 130.8092 | 74.5099 | 70.6382 | 78.3849 | 133.7665 | 204.3618 | 209.3618 | 210.4671 | 160.4046 | 145.8717 | 121.0855 | 122.6809 | 113.7665 | 184.852 | 166.0625 | 206.1052 | 174.3388 | 131.4671 | 123.7203 | 162.2763 | 145.2994 | 164.4243 | 186.5296 | 189.3816 | 225.2533 | 151.3421 | 149.296 | 211.5526 | 193.7237 | 208.4671 | 200.704 | 215.3421 | 282.4671 | 165.0625 | 175.3618 | 147.0625 | 227.7237 | 127.7434 | 147.4671 | 133.0889 | 163.1052 | 173.5757 | 167.8914 | 178.9572 | 135.704 | 110.0197 | 100.5526 | 109.148 | 120.5954 | 248.1677 | 250.3191 | 246.4671 | 344.4046 | 245.7203 | 274.1711 | 264.4046 | 322.0197 | 143.4046 | 68.3191 | 74.6382 | 143.4046 | 102.6809 | 105.2566 | 114 | 103.3618 | 113.852 | 111.6382 | 104.6809 | 117.8323 | 74.1283 | 70.0428 | 78.8717 | 93.5526 | 103.5329 | 88.296 | 91.5099 | 84.1711 | 230 | 204.5296 | 245.7895 | 280.6184 | 181.2105 | 218.1086 | 172.0658 | 265.0625 | 318.4704 | 210.5329 | 214.0033 | 319.5099
1167 | 14.8314 | 132.349 | 194.6333 | 211.0295 | 191.4492 | 159.6627 | 159.6451 | 144.3019 | 145.3785 | 162.0354 | 112.0059 | 121.3844 | 145.9587 | 191.0177 | 147.3078 | 158.1002 | 204.3726 | 182.6333 | 155.3255 | 130.3019 | 182.296 | 162.704 | 290.9823 | 306.3608 | 151.6392 | 227.7276 | 227.9646 | 221.4021 | 284.6922 | 126.0177 | 127.3255 | 109.6451 | 138.349 | 142.6981 | 142 | 147.0648 | 90.3549 | 260.7688 | 236.9646 | 232.9882 | 260.6392 | 124.2901 | 120.6745 | 135.2783 | 129.3019 | 79.349 | 70.3432 | 65.3019 | 117.7335 | 210.6215 | 207.3373 | 223.9764 | 156.0059 | 155.3137 | 124.9529 | 124.651 | 112.9764 | 175.3373 | 172.704 | 210.704 | 175.7511 | 138.349 | 149.0236 | 163.6745 | 135.9705 | 178.651 | 208.6686 | 202.6686 | 237.7099 | 149.5802 | 161.0354 | 215.0118 | 191.9941 | 210.3785 | 186.6451 | 201.651 | 292.9941 | 182.3078 | 177.9764 | 152.7099 | 220.6922 | 140.9941 | 150.0413 | 110.9175 | 166.7099 | 170.6156 | 184.0295 | 186.6333 | 124.6274 | 121.6863 | 110.9705 | 119.6568 | 124 | 272.6627 | 256.9646 | 246.0589 | 328.7394 | 261.0825 | 267.3255 | 251.7217 | 328.7158 | 136.3549 | 72.9764 | 76.3314 | 136.3549 | 98.0236 | 93.3314 | 112.0118 | 106.9705 | 104.3373 | 104.3844 | 100.3549 | 93.6686 | 71.9764 | 66.0118 | 81.0236 | 94.6922 | 95.6568 | 107.9882 | 98.6686 | 70.0354 | 232.9823 | 224.349 | 216.0471 | 275.2842 | 189.7452 | 190.0413 | 175.5331 | 273.3608 | 301.6333 | 212.5979 | 195.9175 | 328.9882
1168 | 14.8424 | 133.8732 | 198.0958 | 195.8451 | 165.0591 | 158.938 | 165.969 | 158.0929 | 138.2478 | 148.2817 | 107.5944 | 105.031 | 159.3127 | 195.3098 | 154.969 | 132.2478 | 195.8112 | 182.6592 | 151.4705 | 141.6873 | 190.5324 | 155.8422 | 287.7212 | 292.1888 | 156.5324 | 217.9971 | 226.5973 | 209.6844 | 282.5295 | 121.3746 | 128.969 | 113.9071 | 133.6873 | 139.5605 | 142 | 127.4056 | 86.7803 | 236.6535 | 243.1578 | 235.969 | 270.6873 | 135.5634 | 121.0929 | 150.5634 | 134.8451 | 72.6254 | 67.6873 | 74.6254 | 99.1239 | 216.0339 | 201.5324 | 227.1888 | 156.7493 | 155.8141 | 130.3465 | 136.5295 | 120.3127 | 177.4366 | 163.7803 | 206.5915 | 151.4676 | 133.3437 | 130.3465 | 161.6873 | 143.5944 | 178.1578 | 197.6902 | 209.5605 | 221.6254 | 178.6563 | 148.3127 | 215.8112 | 194 | 201.5295 | 192.6254 | 201.1578 | 297.7493 | 188.5944 | 179.4705 | 145.5605 | 210.6254 | 141.2817 | 136.3437 | 143.1239 | 158.8732 | 179.0648 | 183.2788 | 188.0339 | 132.5014 | 113.6254 | 121.6873 | 115.0958 | 120.2197 | 272.969 | 262.4705 | 229.0929 | 326.3379 | 228.8761 | 262.0958 | 241.5915 | 311.6563 | 132.4366 | 77.2197 | 72.8761 | 132.4366 | 100.9971 | 103.9661 | 110.7493 | 112.5324 | 104.3746 | 86.6563 | 101.9352 | 97.8112 | 78.969 | 68.8732 | 68.8761 | 86.6873 | 98.6563 | 105.4705 | 95.5944 | 64.8732 | 242.0929 | 226.9042 | 201.031 | 285.8141 | 167.8112 | 188.3718 | 211.7212 | 268.1239 | 310.938 | 225.3465 | 215.0648 | 327.5014
1169 | 14.8534 | 149.1357 | 181.5762 | 196.6785 | 186.0845 | 151.7964 | 163.2881 | 167.5595 | 148.3048 | 158.5261 | 108.3393 | 104.1524 | 157.644 | 213.5762 | 158.6952 | 151.2036 | 205.7119 | 162.0679 | 143.6274 | 147.9833 | 178.0845 | 173.7631 | 268.2202 | 278.3726 | 148.356 | 235.4405 | 206.3393 | 229.5083 | 288.5929 | 126.4917 | 126.2881 | 105.8643 | 130.7286 | 158.9833 | 136.6607 | 135.2036 | 97.9155 | 252.0167 | 232.7119 | 245.0345 | 269.5083 | 129.7119 | 129.4917 | 144.356 | 123.9321 | 70 | 68.644 | 73.7798 | 119.4583 | 198.4238 | 204.3904 | 214.7964 | 165.5595 | 144.356 | 116.4071 | 144.3726 | 119.7119 | 184.356 | 169.9321 | 205.9144 | 158.2202 | 134 | 126.3737 | 167.2714 | 135.4405 | 177.6785 | 179.4238 | 210.4739 | 234.6618 | 167.6774 | 142.3726 | 221.7964 | 202.5428 | 211.5083 | 183.9488 | 191.4238 | 294.1012 | 177.9488 | 166.644 | 160.356 | 220.8142 | 138 | 134.8643 | 142.4572 | 161.6774 | 174.8142 | 191.7453 | 180.0345 | 125.9155 | 118.119 | 129.0512 | 104.6274 | 116.5595 | 263.525 | 248.5762 | 230.7286 | 367.9321 | 241.6452 | 253.0512 | 241.6263 | 315.6274 | 131.881 | 75.3393 | 70.3393 | 131.881 | 117.3726 | 107.3382 | 113.5083 | 103.2881 | 116.6107 | 95.9666 | 114.7286 | 104.8643 | 73.7964 | 83.4238 | 67.7631 | 85.8643 | 100.4917 | 94.7798 | 85.6607 | 77.2881 | 241.5929 | 230.9655 | 209.7631 | 279.6952 | 167.3893 | 200.2536 | 194.7119 | 263.8977 | 307 | 214.6107 | 199.4238 | 313.4405
1170 | 14.8644 | 159.6284 | 183.0507 | 227.7837 | 179.7365 | 152.2601 | 164.9459 | 176.2095 | 140.2095 | 170.848 | 122.7871 | 102.6824 | 157.3682 | 206.9493 | 163.4764 | 157 | 205.8953 | 180.9391 | 147.5811 | 150.7399 | 166.9493 | 176.7939 | 286.7263 | 275.5236 | 152.3142 | 233.7871 | 224.8378 | 225 | 279.5777 | 125.8446 | 126.8412 | 104.9459 | 128.5236 | 166.1588 | 123.3176 | 138.0541 | 107.5777 | 243.7365 | 239.1554 | 256.5304 | 265 | 133.2095 | 134.7365 | 138.7399 | 119.7905 | 75.8919 | 67.2635 | 74.7905 | 135.3716 | 201.7365 | 207.3244 | 217.8378 | 169.0541 | 142.7905 | 135.7331 | 138.5777 | 127.2601 | 183.527 | 163.5811 | 181.9966 | 161.2601 | 133.6318 | 144.9527 | 174.6318 | 134.5236 | 177.8514 | 183.1047 | 191.8919 | 236.4866 | 149 | 137.6824 | 217.473 | 211.7399 | 197.7939 | 183.679 | 197.6824 | 287.9425 | 166.6318 | 173.7331 | 162.1047 | 233.6892 | 134.3176 | 137.9966 | 130.3142 | 144.473 | 188.4257 | 181.6284 | 192.6352 | 136.6824 | 119.2163 | 133.0034 | 115.5777 | 113.7939 | 253.6824 | 251.8919 | 227.4189 | 361.9493 | 273.8446 | 261.7905 | 216.4189 | 316.6352 | 124.5236 | 78.3716 | 73.7399 | 124.5236 | 100.1622 | 92.2567 | 107.8953 | 107.8919 | 121.1115 | 109.9493 | 108.473 | 106.8919 | 65.7905 | 84.8953 | 81.8412 | 89.7331 | 101.3176 | 94.3176 | 81.5236 | 73.0541 | 230.3682 | 212.473 | 210.5845 | 289.6318 | 154.7871 | 184.5236 | 200.0507 | 262.2533 | 291.5338 | 220.2163 | 209.3648 | 325.4121
1171 | 14.8754 | 164.9115 | 185.8141 | 224.0264 | 182.5221 | 152.7256 | 162.0088 | 184.5664 | 148.1859 | 165.2213 | 115.3981 | 106.3363 | 156.0973 | 198.4779 | 156.2832 | 159.6637 | 204.7611 | 200.8672 | 140.7611 | 143.8053 | 167.6107 | 161.3805 | 323.2389 | 275.1061 | 154.1947 | 237.8141 | 242.2832 | 225 | 280.1947 | 114.7168 | 124.6725 | 114.5664 | 142.9469 | 164.1416 | 121.1859 | 137.5664 | 94.1149 | 252.9912 | 244.5752 | 243.1416 | 265.3805 | 139.2832 | 133.7168 | 132.1859 | 125.8939 | 86.3805 | 78.5576 | 70.2389 | 126.2832 | 207.1859 | 181.9469 | 223.4867 | 157.5309 | 140.5221 | 136.4424 | 139.9557 | 127.7256 | 184.0443 | 157.5221 | 198.9469 | 172.7611 | 142.1328 | 130.1859 | 170.9557 | 140.1947 | 155.6549 | 180.0531 | 216.0797 | 208.8851 | 161.5576 | 155.0355 | 208.9645 | 209.3717 | 193.416 | 197.0531 | 197.9115 | 304.6195 | 165.2389 | 174.8229 | 167.8053 | 216.7611 | 133.708 | 139.3451 | 146.6549 | 150.4248 | 176.5664 | 181.9645 | 189.1328 | 144.1416 | 107.7523 | 114.5309 | 115.4336 | 110.9381 | 250.1061 | 260.4779 | 237.2832 | 340.5397 | 268.0443 | 252.292 | 227.0443 | 323.4072 | 140.4691 | 67.7611 | 69.0885 | 140.4691 | 93.7523 | 115 | 115.1328 | 113.4336 | 105.8496 | 99.5752 | 107.5752 | 112.4336 | 62.3805 | 81.4779 | 76.2477 | 97.6725 | 96.5221 | 92.1859 | 97.4691 | 70.6637 | 236.708 | 207.3893 | 194.0443 | 291.6637 | 171.7523 | 194.3805 | 207 | 297.7611 | 276.416 | 206.4691 | 210.9645 | 350.5309
1172 | 14.8864 | 159.7139 | 186.4637 | 200.3201 | 177.5363 | 139.1426 | 159.9991 | 190.0359 | 158.9282 | 181.3579 | 98.2493 | 107.4995 | 150.6431 | 199.1426 | 150.965 | 160.0718 | 213.0708 | 183.2502 | 151.8206 | 149.2144 | 186.678 | 166.7139 | 313.3579 | 271.9632 | 157.035 | 220.7866 | 240.5005 | 228.9282 | 278.7139 | 126.713 | 114.0359 | 129.8565 | 144.322 | 160.5005 | 133.1067 | 147.3569 | 74.035 | 262.8574 | 240.1785 | 239.1076 | 263.6431 | 137.8933 | 136.678 | 140.5713 | 134.1435 | 79.9292 | 89.965 | 73.3211 | 127.2502 | 224.9991 | 202.1775 | 217.4278 | 155.6421 | 148.4995 | 118.7498 | 135.7856 | 111.7856 | 191.3569 | 157.2502 | 196.7866 | 174.7856 | 152.6789 | 140.5354 | 167.1785 | 135.1785 | 174.9641 | 186.1417 | 230.7507 | 229.3579 | 161.5731 | 160.8225 | 199.2493 | 204.5373 | 214.7498 | 195.678 | 191.1426 | 310.678 | 165.1785 | 151.4646 | 179.8924 | 228.6063 | 138.2861 | 129.8565 | 153.7866 | 151.2861 | 183.6072 | 174.6063 | 188.6799 | 145.2144 | 116.1076 | 107.9282 | 113.1067 | 123.7507 | 237.9282 | 251.7148 | 233.1435 | 323.678 | 279.2852 | 237.1076 | 226.8933 | 344.3937 | 140.0009 | 68.6072 | 71.1435 | 140.0009 | 109.5713 | 116.9641 | 126.0718 | 111.8924 | 104.4296 | 101.4637 | 97.2861 | 104.2144 | 71.2493 | 73.8933 | 68.2852 | 91.3569 | 104.4995 | 102.9282 | 108 | 72.6431 | 242.4646 | 202.8565 | 204.1067 | 290.8933 | 176.965 | 188.7148 | 198.7507 | 306.0708 | 292.2502 | 227.2134 | 200.0708 | 325.4655
1173 | 14.8974 | 158.8977 | 199.2153 | 211.4414 | 174.1023 | 158.9893 | 175.3795 | 171.2005 | 151.6312 | 168.4051 | 112.2153 | 117.6205 | 151.4563 | 210.887 | 134.9744 | 147.9233 | 220.354 | 176.5693 | 160.113 | 141.9233 | 192.5437 | 172.3795 | 300 | 281.8721 | 161.6824 | 220.3944 | 229.5693 | 231.354 | 281.1386 | 146.9488 | 113.8358 | 132.6824 | 151.42 | 156.4563 | 139.7847 | 147.7591 | 85.1642 | 250.7079 | 247.6716 | 219.0619 | 258.7847 | 134.8614 | 147 | 140.9744 | 117.9488 | 72.2409 | 82.4818 | 72.7079 | 120.1642 | 218.9084 | 208.8465 | 226.7335 | 165.354 | 154.9744 | 127.8614 | 134.9744 | 113 | 187.3028 | 164.3432 | 200.0404 | 163.0363 | 139.9233 | 146.8102 | 164.1386 | 137.8102 | 173.5437 | 203.5437 | 212.3284 | 213.9744 | 152.6865 | 139.9744 | 202.6824 | 184.9125 | 208.8721 | 208.4307 | 187.8977 | 318.1642 | 165.7847 | 153.2921 | 165.0982 | 235.2772 | 141.1279 | 142.8102 | 138 | 142.3795 | 176.5182 | 191.4051 | 180.3944 | 138.3284 | 108.6205 | 114.4051 | 114.113 | 121.533 | 251.6972 | 232.6824 | 230.3176 | 340.887 | 279.2772 | 221.9233 | 219.8102 | 328 | 126.646 | 71.2409 | 63.4563 | 126.646 | 106.7335 | 121.6205 | 114.3284 | 123.4307 | 91.887 | 109.354 | 98.9125 | 104.2153 | 74.2772 | 72.4818 | 70.3028 | 96.6205 | 102.0619 | 111.8358 | 107.1898 | 74.2665 | 236.5949 | 211.7591 | 218.0767 | 296.7739 | 165.4307 | 188.3176 | 192.0767 | 298.3647 | 295.2921 | 234.8977 | 199.2517 | 326.0916
1174 | 14.9084 | 144.913 | 211.4352 | 194.6696 | 183.4956 | 172.2345 | 175.9218 | 153.0089 | 146.1918 | 164.8259 | 117.7567 | 127.5133 | 163.8437 | 221.8348 | 145.7744 | 142.7567 | 223.7655 | 173.4174 | 145.8259 | 145.5222 | 181.4085 | 175.8437 | 299.1652 | 271.5915 | 147.5826 | 241.7478 | 232.6785 | 217.6519 | 288.6092 | 135.5737 | 120.9218 | 117.3304 | 166.5648 | 155.9041 | 135.0782 | 135.0693 | 87.2611 | 248.3481 | 256.8348 | 220.1137 | 258.2522 | 136.1563 | 138.6519 | 128.3997 | 124.5222 | 78.2522 | 83.2345 | 59.0782 | 123.0959 | 193.3481 | 198.9307 | 224.0089 | 154.9911 | 148.6607 | 133.3304 | 135.3393 | 118.0089 | 167.6519 | 180.1652 | 215.8952 | 160.6963 | 136.8437 | 146.7478 | 162.8437 | 145.6785 | 177.4352 | 198.2522 | 201.0782 | 226.8614 | 172.643 | 144.9307 | 195.2611 | 193.4085 | 203.6003 | 199.8952 | 178.5044 | 311.913 | 168.1741 | 158.9911 | 141.27 | 221.4174 | 154.7478 | 138.9911 | 146.7655 | 147.5133 | 161.5737 | 189.0782 | 203.8348 | 127.913 | 106.4085 | 115.8348 | 106.087 | 125.3126 | 273.0178 | 238.2007 | 231.8952 | 353.9218 | 262.913 | 211.7478 | 225.1741 | 330.5044 | 136.1741 | 73.913 | 56.643 | 136.1741 | 109.0178 | 120.6607 | 109.7567 | 117.3997 | 95.7389 | 108.5915 | 115.3393 | 121.444 | 68.7655 | 88.2611 | 65.2611 | 100.6696 | 91.8437 | 110.1563 | 107.2522 | 81.5826 | 237.2522 | 227.4529 | 227 | 308.6519 | 166.4956 | 202.4174 | 205.1741 | 276.8437 | 303.0782 | 232.1829 | 201.2007 | 352.9041
1175 | 14.9194 | 154.8909 | 219.9842 | 193.1327 | 174.8436 | 166.8752 | 166.2497 | 168.1642 | 173.7346 | 158.1406 | 123.8594 | 127.6873 | 161.2576 | 210.1091 | 154.2576 | 139.8358 | 231.2733 | 178.7267 | 141.7188 | 156.5545 | 177.5782 | 173.2576 | 301.0079 | 274.5624 | 155.5939 | 226.6794 | 218.7963 | 224.0473 | 307.7109 | 121.1248 | 125.4297 | 111.1327 | 148.1327 | 149.8673 | 134.0079 | 126.5782 | 93.8515 | 247.9685 | 246.3982 | 250.8752 | 263.8673 | 137.0236 | 123.5624 | 118.4376 | 136.4139 | 75.2733 | 67.1327 | 63.1642 | 129.9921 | 205.4297 | 201.8358 | 221.9764 | 156.1642 | 133.2576 | 137.0158 | 139.1406 | 122.8515 | 172.7582 | 177.2812 | 193.4139 | 175.8358 | 145 | 139.8436 | 170.5703 | 146.4061 | 193.2891 | 192.6951 | 204.7346 | 252.4376 | 155.7188 | 157.2891 | 212.5939 | 199.0315 | 212.2733 | 203.6255 | 186.297 | 312.0079 | 170.9921 | 156.7267 | 151.8436 | 223.7188 | 155.1485 | 134.1485 | 154.2733 | 164.4455 | 166.0315 | 184.5703 | 192.1091 | 135.7424 | 100.4297 | 111.8436 | 111.5782 | 117.1642 | 277.9764 | 275.3127 | 229.6097 | 351.1248 | 266.0158 | 225.8988 | 227.5624 | 318.5309 | 132.117 | 74.4376 | 49.6018 | 132.117 | 117.4139 | 116.4297 | 111.9921 | 111.3049 | 104.1879 | 114.5703 | 112.6951 | 124.0933 | 85.7267 | 86.6873 | 66.6951 | 98.2733 | 97.4218 | 101.5703 | 102.5545 | 79.7109 | 247.5939 | 237.0933 | 230.0079 | 303.0158 | 165.1485 | 209.8752 | 214.4376 | 278.1248 | 301.5782 | 231.6794 | 218.117 | 332.2576
1176 | 14.9304 | 157.8564 | 206.6961 | 195.0718 | 173.7458 | 173.7901 | 152.4199 | 173.6961 | 183.558 | 152.1381 | 120.1381 | 117.0939 | 144.3481 | 185.0442 | 135.5801 | 140.4917 | 223.232 | 182.3481 | 149.3039 | 150.6519 | 189.3978 | 160.326 | 297.9282 | 266.9061 | 168.326 | 231.0994 | 203.0276 | 237.8343 | 298.9282 | 121.7237 | 118.0442 | 105.116 | 143.884 | 143.5083 | 144.6298 | 129.558 | 90.116 | 259.8453 | 243.8177 | 255.5801 | 268.116 | 148.116 | 123.8619 | 129.6298 | 132.0939 | 80.0497 | 70.3978 | 74 | 122.0221 | 206.884 | 195.4199 | 210.884 | 168.326 | 131.3757 | 147.2099 | 124.7403 | 119.558 | 179.0884 | 170.1381 | 212.0773 | 172.0718 | 150.3039 | 141.8398 | 179.7237 | 135.884 | 194.558 | 175.9282 | 216.326 | 233.5746 | 155.3481 | 148.3923 | 224.442 | 216.326 | 205.116 | 211.9945 | 188.4641 | 300.0884 | 167.442 | 164.768 | 150.7458 | 229.5359 | 150.4862 | 144.0718 | 136.9503 | 168.1381 | 179.3481 | 181.3481 | 177.6519 | 145.1381 | 106.3039 | 104.116 | 120.3039 | 130.2099 | 267.326 | 278.8343 | 237.0663 | 345.9779 | 261.6243 | 247.442 | 226.0939 | 306.3978 | 126.0718 | 76.7901 | 70.3039 | 126.0718 | 104.6961 | 112.1381 | 109.768 | 113.4862 | 123.4641 | 103.3923 | 104.326 | 102.9779 | 84.4862 | 73.884 | 62.7458 | 94.2099 | 96.6519 | 104.5359 | 91.7901 | 81.0939 | 251.4862 | 223.0497 | 219.4144 | 302.1602 | 164.0221 | 205.7403 | 208.3923 | 268.116 | 302.7901 | 225.9337 | 210.7458 | 306.5083
1177 | 14.9414 | 149.5399 | 208.6313 | 208.5428 | 178.2743 | 169.6372 | 153.9115 | 155.1887 | 173.4601 | 154.6313 | 107.6401 | 112.823 | 146.4514 | 186.357 | 131.3628 | 159.0885 | 212.3687 | 181.2714 | 151.4572 | 152.6372 | 192.7315 | 158.82 | 312.1683 | 279.8083 | 176.3599 | 236.7403 | 225.18 | 222.2743 | 280.4601 | 135.8171 | 121.6284 | 111.7228 | 144.0914 | 136.7228 | 147.0944 | 124.4572 | 85.82 | 281.8259 | 237.7403 | 244.5486 | 253.3716 | 145.1829 | 134.5428 | 131.6401 | 132.82 | 77.6459 | 72.823 | 72.6372 | 120.6342 | 213.4514 | 220.534 | 218.3599 | 165.0029 | 152.4514 | 146.3658 | 119.8113 | 121.7257 | 177.6255 | 169.4514 | 218.6518 | 162.8288 | 156.0914 | 140.7345 | 182.4601 | 135.1829 | 176.7374 | 175.177 | 214.3658 | 208.5428 | 164.2655 | 156.8054 | 221.3658 | 212.5486 | 207.18 | 190.7257 | 183.5457 | 289.5399 | 166.6372 | 168.8171 | 147.5516 | 222.1887 | 140.0914 | 146.6401 | 129.8142 | 157.4572 | 179.6342 | 179.3628 | 174.6401 | 137.1829 | 116.18 | 105.7257 | 117.4601 | 140.7228 | 263.0944 | 268.2714 | 226.9971 | 349.6372 | 251.4514 | 243.9115 | 257.711 | 333.3541 | 130.0029 | 77.6342 | 77 | 130.0029 | 107.0856 | 102.82 | 110.6372 | 113.0856 | 111.2772 | 96.3599 | 108.7257 | 117.0856 | 72.7286 | 73.1829 | 70 | 101.5428 | 95.4572 | 108.0914 | 91.2714 | 85 | 248.3599 | 233.3658 | 205.5428 | 292.8171 | 154.0029 | 199.4484 | 196.8171 | 271.5428 | 285.0088 | 228.7345 | 213.0029 | 313.8054
1178 | 14.9524 | 156.3343 | 206.8692 | 212.1337 | 185.7311 | 162.8677 | 155.0654 | 166.6614 | 161.5334 | 152.8692 | 113.064 | 110.4651 | 144.1352 | 201.3997 | 132.5334 | 153.7354 | 224.3954 | 175.6017 | 147.8663 | 143.0683 | 183.0669 | 173.1962 | 331.6017 | 294.4012 | 172.8023 | 223.064 | 223.8677 | 225.0654 | 294.6614 | 133.8009 | 131.468 | 114.468 | 148.5988 | 158.1308 | 136.734 | 134.8634 | 84.3329 | 256.6686 | 237.5945 | 238.9331 | 242.5988 | 154.6643 | 144.6657 | 126.3329 | 128.5334 | 66.3983 | 70.9317 | 70.0669 | 126.3997 | 220 | 228.6032 | 224.1337 | 159.4666 | 150.6017 | 138.2674 | 136.3329 | 122.2006 | 198.6003 | 174.6003 | 191.6657 | 147.532 | 147.0683 | 135.798 | 175.1991 | 141.8648 | 152.734 | 187.7994 | 206.734 | 229.8634 | 173.4012 | 177.4012 | 223.532 | 210.1991 | 205.4012 | 206.5974 | 184.8663 | 303.3329 | 165 | 171 | 147.5974 | 216.3983 | 151.5974 | 137.6003 | 157.596 | 149.2006 | 170.936 | 184.266 | 160.0014 | 139.1991 | 123.266 | 102.468 | 108.7994 | 130.404 | 256.4666 | 276.1323 | 231.1337 | 371.7951 | 261.7326 | 248.3314 | 257.8735 | 353 | 126.3329 | 82.9331 | 67.6686 | 126.3329 | 111.468 | 107.3983 | 119.7311 | 124 | 110.3983 | 101.2006 | 103.6017 | 116.8023 | 76.0654 | 72.3997 | 71.8663 | 106.0669 | 102.1308 | 111.1991 | 102.3983 | 83.6003 | 257.8663 | 222.9346 | 218.4651 | 269.8052 | 150.7994 | 208.2674 | 217.1962 | 286.3314 | 293.1279 | 217.266 | 206.5334 | 343.7326
1179 | 14.9634 | 154.352 | 199.7041 | 210.9577 | 194.1691 | 154.6057 | 154.8171 | 172.5358 | 163.3943 | 152.4081 | 132.7886 | 123.3098 | 137.8309 | 220.239 | 136.7886 | 139.6057 | 245.6902 | 175.0984 | 139.465 | 140.2252 | 177.6902 | 186.8171 | 328.2252 | 285.1268 | 160.9577 | 231.7748 | 228.535 | 216.6764 | 306.761 | 118.9439 | 121.6057 | 113.2252 | 142.0707 | 155.2398 | 142.0984 | 147.7325 | 87 | 257.0138 | 257.465 | 231.3805 | 260.9715 | 150.2821 | 151.4366 | 123.7325 | 129.9154 | 76.4789 | 77.6057 | 77.6195 | 120.8171 | 214.7325 | 201.4228 | 212.9016 | 160.9577 | 138.1268 | 130.1691 | 137.5634 | 122.8309 | 195.0846 | 168.6902 | 203.7041 | 154.0423 | 137.5211 | 145.5634 | 171.8594 | 150.5634 | 168.1545 | 189.0846 | 211.1406 | 251.8309 | 166.0423 | 150.409 | 218.0707 | 210.2114 | 210.9715 | 214.2959 | 190.352 | 320.8447 | 169.7886 | 161.4228 | 165.352 | 224.0846 | 154.5073 | 149.887 | 164.1976 | 149.352 | 168.4504 | 188.9577 | 170.1967 | 139.2114 | 117.4228 | 98.8309 | 104.3382 | 122.3658 | 257 | 269.5496 | 234.7463 | 384.6342 | 268.8732 | 252.7748 | 223.352 | 341.0285 | 148.6333 | 70.5919 | 61.7886 | 148.6333 | 100.1691 | 109.8171 | 122.9016 | 121.6057 | 121.9154 | 102.3098 | 100.2252 | 103.5211 | 73.9016 | 76.8732 | 76.3943 | 109.3098 | 103.0707 | 115.0423 | 107.6902 | 78.1691 | 258.0846 | 210.2537 | 224.6057 | 267.8162 | 163.0984 | 195.3805 | 214.5358 | 288.3805 | 302.8878 | 220.5211 | 214.6195 | 336.0285
1180 | 14.9744 | 155.0531 | 201.5973 | 220.8407 | 180.1769 | 152 | 145.5265 | 144.5442 | 167.4735 | 147.845 | 135.5442 | 111.792 | 149.3673 | 215.4247 | 145.9292 | 154.1903 | 231.6681 | 170.7566 | 134.8761 | 144.0531 | 196.5753 | 173.1062 | 322.7212 | 292.8054 | 160.5265 | 226.4735 | 220.845 | 220.6284 | 274.7389 | 111.4381 | 121.9469 | 128.8407 | 130.4558 | 133.0708 | 141.6858 | 138.0708 | 97.3142 | 262.615 | 243.5442 | 234.77 | 252.3893 | 141.8054 | 138.7566 | 130.77 | 130.0354 | 84.8584 | 73.5265 | 75.7035 | 119.8761 | 211.9469 | 207.6992 | 206.9292 | 162.9823 | 149.7346 | 131.8938 | 136 | 117.1769 | 183.6681 | 164.9823 | 202.6504 | 168.7169 | 138.9646 | 149.4027 | 172.3319 | 149 | 181.6327 | 189.9469 | 213.615 | 236.845 | 158.1239 | 146.6638 | 212.3496 | 200.5796 | 196.4955 | 198.6504 | 197.9292 | 330.1239 | 171.5619 | 149.5265 | 166.5442 | 233.7877 | 137.0885 | 154.1946 | 136.6858 | 153.4912 | 167.2654 | 190 | 190 | 126.1416 | 112.4027 | 101.0354 | 100.9115 | 140.9469 | 244.7212 | 248.1416 | 239.0354 | 358.1946 | 276.9115 | 239.615 | 246.6461 | 329.4735 | 147.8981 | 66.8054 | 68.4735 | 147.8981 | 106.3142 | 104.4558 | 110.5442 | 125.385 | 114.1769 | 112.4027 | 104.0531 | 106.9292 | 72.8407 | 76.5619 | 81.4558 | 97.3008 | 100.7877 | 104.6681 | 99.5619 | 78.9115 | 245.6858 | 209.8938 | 221.5088 | 300.4381 | 162.6858 | 185.5088 | 203.7346 | 285.8761 | 290.3319 | 223.9292 | 228.9115 | 315.1416
1181 | 14.9854 | 156.0275 | 209.593 | 226.4759 | 178.0551 | 158.5448 | 159.1033 | 147.0344 | 158.9311 | 146.124 | 129.9793 | 107.0826 | 159.5172 | 200.062 | 153.0207 | 158.407 | 220.9931 | 161.0344 | 151.5654 | 149.0551 | 207.407 | 167 | 330.1377 | 299.9725 | 161.0138 | 220.4483 | 217.6136 | 223.3726 | 270.062 | 112.9862 | 113.9242 | 122.8967 | 145.0826 | 143.0964 | 127.4759 | 138.0275 | 96.9242 | 269.0895 | 243.0138 | 249.5103 | 244.5999 | 156.5241 | 128.0275 | 140.4759 | 124.9793 | 77.8967 | 69.9862 | 62.4897 | 121.4621 | 223.5585 | 218.9587 | 221.5723 | 161.9862 | 161.9793 | 129.4415 | 129.9587 | 117.0689 | 183.5654 | 174.5585 | 197.5241 | 177.9518 | 152.0758 | 145.4346 | 162.8691 | 147.9931 | 183.5654 | 186.9587 | 216.062 | 241.1653 | 163.0826 | 155.9311 | 219.9931 | 200.0275 | 179.0758 | 189.4966 | 203.5103 | 316.9587 | 165.4828 | 156.0551 | 154.4346 | 239.9587 | 139.0482 | 145.0138 | 131.5379 | 162.5585 | 161.062 | 187.9862 | 183.4552 | 122.0275 | 115.4828 | 103.0275 | 109.5241 | 133.4277 | 237.5379 | 243.0207 | 234.9862 | 342.9725 | 277.469 | 237.531 | 257.938 | 313.8829 | 127.0138 | 76.4897 | 68.9931 | 127.0138 | 115.4897 | 101.9656 | 112.0275 | 123.4415 | 119.6068 | 112.4621 | 95.4621 | 122.5792 | 75.4552 | 74.5103 | 75.4415 | 84.0275 | 107.9656 | 93.4897 | 91.9725 | 78.9656 | 235.5034 | 228.5861 | 232.0758 | 300.4759 | 171.1309 | 180.9725 | 220.5103 | 281.9242 | 297.9862 | 225.9862 | 230.4759 | 325.6274
1182 | 14.9964 | 166.1888 | 211.6538 | 212.1696 | 186 | 168.6101 | 168.8427 | 149.9371 | 154.1573 | 160.3899 | 121.3269 | 112.2954 | 159.9371 | 209 | 154.4528 | 146.5472 | 231.3462 | 168.0629 | 146.0437 | 160.6101 | 186.264 | 170.6101 | 331.4336 | 300.1259 | 157.3269 | 231.0507 | 242.2518 | 219.9563 | 278.4843 | 124.9248 | 111.2518 | 106.9685 | 152.8741 | 149.7797 | 132.2518 | 142 | 93.736 | 256.729 | 244.4843 | 265.9563 | 258.4843 | 146.5909 | 143.8619 | 127.7168 | 127.6731 | 66.0944 | 61.8112 | 66.1573 | 112.9056 | 215.4965 | 228.9878 | 224.264 | 160.5157 | 150.7482 | 128.2203 | 133.799 | 117.201 | 186.8112 | 186.0944 | 207.7046 | 177.7046 | 158.3584 | 140.6416 | 164.6294 | 140.2954 | 193 | 194.9248 | 216.2325 | 251.5909 | 167.7797 | 146.5157 | 233.9563 | 205.5472 | 187.9371 | 199.8304 | 203.4528 | 303.7797 | 178.472 | 168.6416 | 142.9371 | 238.6416 | 142.9056 | 135.1381 | 139.5787 | 176.6731 | 184.4406 | 191.1573 | 188.3462 | 117.7482 | 108.3584 | 109.0629 | 109.3899 | 130.736 | 240.9371 | 242.9056 | 238.1573 | 355.5035 | 280.2203 | 232.201 | 255.1888 | 321.2395 | 127.4528 | 72.9371 | 68.5157 | 127.4528 | 119.1573 | 106.2832 | 116 | 121.1888 | 114.8863 | 104.9371 | 93.0944 | 115.9493 | 73.1259 | 73.9371 | 64.9371 | 95.736 | 98.3584 | 96.6416 | 88 | 81.736 | 237.5472 | 233.264 | 239.9056 | 290.8112 | 181.7482 | 180.6101 | 206.0122 | 275.6416 | 293.9371 | 221.9371 | 206.3706 | 330.0752
1183 | 15.0074 | 155.6312 | 185.6873 | 210.4484 | 186.528 | 153.5191 | 162.9439 | 168.065 | 173.7847 | 164.3924 | 116.528 | 110.7522 | 153.2478 | 192.1032 | 159.8643 | 145.8879 | 230.9115 | 164.1917 | 133.0561 | 155.0236 | 179 | 161.8554 | 309.2478 | 303.472 | 159.9204 | 233.7433 | 238.3834 | 222.9115 | 298.593 | 141.1682 | 121.6401 | 106 | 145.3038 | 147.2242 | 135.2478 | 137.7758 | 103.1121 | 248.8408 | 246.6401 | 261.5191 | 247.4395 | 136.1121 | 143.9115 | 125.8643 | 130.8879 | 70.0561 | 63.3924 | 72.0561 | 112.6401 | 202.6401 | 230.2713 | 217.528 | 180.0089 | 140.8879 | 143.4484 | 146.6962 | 119.0885 | 185.2242 | 192.1682 | 200.2713 | 177.6637 | 146.0796 | 152.3924 | 173.9675 | 149.8408 | 176.1032 | 201.1357 | 222.7847 | 232.1357 | 165.7522 | 157.0325 | 232.6873 | 208.5841 | 199.7287 | 197.3274 | 193.5516 | 298.0561 | 194.5841 | 175.6401 | 166.3452 | 239.8318 | 146.3363 | 132.4484 | 134.0796 | 160.3509 | 183.2153 | 188.6076 | 184.7433 | 128.4809 | 123.537 | 112.0561 | 114.4484 | 133.7758 | 234.2478 | 252.1446 | 240.8879 | 372.5841 | 265.879 | 251.5134 | 245.1592 | 342.4159 | 131.2803 | 73.1121 | 72.6962 | 131.2803 | 110.7994 | 115.528 | 109.6637 | 131.7522 | 110.2567 | 94.0236 | 102.3363 | 97.4159 | 77 | 63.5516 | 67.2242 | 88.7433 | 100.8643 | 92.0236 | 86.4159 | 80.5516 | 234.2478 | 227.0561 | 225.9115 | 304.537 | 170.8318 | 191.3924 | 191 | 280 | 294.6401 | 211.0236 | 201.7847 | 314.3599
1184 | 15.0184 | 153.4273 | 190.9111 | 213.1371 | 188.0806 | 150.5079 | 166.8629 | 180.5968 | 171.8711 | 164.5162 | 127.8063 | 104.734 | 148.4597 | 179.1613 | 160.0565 | 152.6451 | 215.0565 | 171.9676 | 135.621 | 148.0806 | 180.0806 | 161.266 | 309.3225 | 307.8629 | 166.4597 | 215.5968 | 228 | 227.0482 | 287.8229 | 126.7099 | 125.0806 | 109.7822 | 130.113 | 144.5162 | 129.9194 | 149.1289 | 97.4356 | 274.887 | 243.5968 | 235.8146 | 262.3143 | 141.2419 | 135.621 | 130.379 | 133.3225 | 71 | 66.2178 | 72.4597 | 109.5968 | 208.7822 | 222.8629 | 239.6127 | 165.1213 | 142.7822 | 148.2984 | 138.6533 | 117.6533 | 193.3225 | 188.5162 | 180.9759 | 172.5403 | 132.5162 | 156.2984 | 172.0241 | 161.8387 | 156.6775 | 185.2743 | 216.5486 | 249.2336 | 154.3308 | 158.4356 | 223.3225 | 213.7822 | 211.4597 | 189.2419 | 196.8063 | 289.8146 | 178.7099 | 165.0324 | 180.3549 | 239.1613 | 148.2178 | 142.1613 | 132.4032 | 154.5079 | 171.0806 | 190.6451 | 181.1854 | 139.5968 | 128.5727 | 107.0565 | 125.2419 | 129.4597 | 250.532 | 253.8146 | 240.0806 | 382.1048 | 260.5079 | 253.2261 | 226.1371 | 342.621 | 135.4597 | 73.379 | 74.9194 | 135.4597 | 106.5644 | 105.1937 | 110.4838 | 131.1371 | 108.9518 | 96.8063 | 100.4356 | 105.7257 | 86.1854 | 66.266 | 79.1048 | 88.4273 | 109.1613 | 93.7257 | 96.887 | 69.7581 | 239.1854 | 221.5162 | 221.4032 | 316.5968 | 169.621 | 200.1613 | 202.3467 | 268.6533 | 275.9276 | 208.9435 | 228.5079 | 322.266
1185 | 15.0294 | 156.8161 | 210.8683 | 209.5526 | 189 | 160.3422 | 171.5526 | 166.6057 | 165.6313 | 155.6844 | 135.3422 | 108.1574 | 151.3156 | 186.5261 | 161.0787 | 159.4474 | 211.6578 | 198.473 | 150.8152 | 140.1583 | 185.9735 | 161.1583 | 307.4739 | 305.3687 | 170.4209 | 213.763 | 229.1052 | 235.6844 | 288.42 | 116.9735 | 118.2635 | 106.3687 | 129.3943 | 144.5261 | 132.8683 | 149.29 | 94.3156 | 275.6057 | 253.9204 | 246.2361 | 273.0531 | 153.3943 | 137 | 125.1317 | 133.1317 | 73.763 | 69.6313 | 72 | 110.5261 | 219.7365 | 216.5005 | 234.2379 | 150.71 | 131.6322 | 129.4218 | 134.5261 | 105.7896 | 178.2113 | 170.8426 | 204.2882 | 164.1583 | 129.763 | 146.8161 | 168.0531 | 166.6313 | 161.2891 | 175.8948 | 206.6313 | 263.4739 | 147.6313 | 149.7896 | 233.6313 | 218.6578 | 215.9735 | 189.237 | 206 | 289.7365 | 176.71 | 146.2635 | 170.237 | 242.6578 | 137.8161 | 145.6578 | 128.7109 | 167.6578 | 176.4209 | 184.1848 | 192.8683 | 132.237 | 114.1317 | 111.9469 | 109.2113 | 118.5005 | 269.6578 | 251.5261 | 233.2635 | 382.3687 | 254.3165 | 252.8939 | 227.5261 | 347.3156 | 141.0787 | 67.5791 | 61.8426 | 141.0787 | 108.5791 | 97.1052 | 115.4474 | 112.6322 | 107.4995 | 107.1052 | 108.3678 | 116.2104 | 87.3687 | 75.5526 | 78.8161 | 94.5791 | 109.3422 | 104.2104 | 97.0531 | 68.1052 | 245.8948 | 220.9735 | 233.7365 | 320.8417 | 178.1839 | 206.4209 | 195.4218 | 265.0787 | 282.8674 | 226.1574 | 244.4209 | 318.8426
1186 | 15.0404 | 143.6563 | 212.3053 | 220.7331 | 181.0914 | 149.9616 | 179.3437 | 168.0384 | 180.0384 | 165.9927 | 134 | 115.7404 | 157.3894 | 187.0457 | 161.4808 | 170.8629 | 209.6106 | 203.6563 | 164.2596 | 147.6874 | 200.7331 | 164.7331 | 310.9086 | 291.5265 | 167.2212 | 227.298 | 231.1298 | 225.6563 | 293.1828 | 120.4351 | 131.7715 | 108.9086 | 148.298 | 154.649 | 137.6947 | 136.1755 | 95.3053 | 269.1298 | 247.3583 | 245.7477 | 267.2596 | 167.2139 | 130.2212 | 145.7258 | 135.0841 | 77.6947 | 65.9616 | 65.2212 | 108.2212 | 232.2139 | 215.9086 | 203.702 | 167.7788 | 118.3053 | 135.7715 | 140.6947 | 107.9543 | 174.298 | 168.3437 | 209.3583 | 159.2596 | 142.1682 | 132.5265 | 176.947 | 168.0457 | 175.9086 | 177.2596 | 214.8245 | 235.8391 | 154.6947 | 148 | 215.8391 | 213.7861 | 216.6106 | 192.649 | 198.6563 | 327.0694 | 163.8391 | 158.6417 | 155.5722 | 225.9232 | 136.5192 | 159.4278 | 120.3053 | 176.9086 | 179.4351 | 189.947 | 196.5649 | 130 | 121.1682 | 120 | 98.5192 | 121.298 | 251.7934 | 236.2285 | 234.9086 | 391.1225 | 243.3894 | 262.7861 | 238.2139 | 334.7477 | 129.053 | 69.649 | 64.9927 | 129.053 | 114.0384 | 105.3437 | 118.9543 | 112.298 | 108.0914 | 106.3053 | 116.6106 | 117.4351 | 77.4808 | 67.5265 | 70.1755 | 98.3437 | 104.6106 | 100.351 | 98.0384 | 68.4351 | 261.3821 | 244.7715 | 229.2669 | 302.0147 | 177.2212 | 210 | 199.5119 | 262.0914 | 288.8775 | 215.0987 | 237.8318 | 335.5502
1187 | 15.0514 | 151.2754 | 205.8053 | 215.7246 | 170.9597 | 155.121 | 180.3825 | 176.7316 | 162.1808 | 178.3088 | 134.5772 | 115.7316 | 151.9193 | 196.6982 | 159.7316 | 173.0737 | 215.6578 | 181.8387 | 163.6912 | 152.6509 | 200.3422 | 159.1474 | 310.0737 | 286.1544 | 163.7316 | 229.0737 | 227.9597 | 222.8859 | 304.0876 | 107.879 | 132.5703 | 112.1141 | 149.4965 | 154.9597 | 130.3422 | 143.8122 | 93.4228 | 272.8859 | 249.1613 | 240.9263 | 258.0334 | 148.3352 | 135.9666 | 144.9527 | 135.5369 | 72.0737 | 70.5438 | 65.1947 | 113.3894 | 213.9124 | 206.4159 | 231.9402 | 160.879 | 126.235 | 143.4965 | 131.0334 | 120.235 | 173.765 | 168.2281 | 192.6912 | 163.8859 | 145.9597 | 131.7719 | 175.5703 | 145.3755 | 175.6509 | 188.8122 | 211.8053 | 213.3825 | 172.7385 | 151.4631 | 222.242 | 210.1544 | 210.5369 | 201.0403 | 195.8859 | 324.4493 | 155.235 | 158.5703 | 158.121 | 222.9666 | 144.6175 | 155.1474 | 142.0876 | 181.8456 | 186.5035 | 190.879 | 181.4159 | 135.1947 | 113.9931 | 120.5772 | 103.1544 | 132.3088 | 260.6648 | 242.9332 | 230.6106 | 373.7177 | 242.5369 | 259.7316 | 234.9193 | 331.0807 | 135.6245 | 80.3491 | 86.5438 | 135.6245 | 111.765 | 108.6912 | 110.4562 | 122.1544 | 96.2281 | 103.2684 | 100.7246 | 125.0807 | 79.7719 | 58.1141 | 77.235 | 91.879 | 103.1544 | 85.0334 | 111.3491 | 72.0403 | 256.6843 | 249.6106 | 229.0807 | 291.235 | 170.8456 | 214.0403 | 196.8387 | 269.2754 | 272.8053 | 199 | 212.1071 | 345.0334
1188 | 15.0624 | 163.9474 | 201.4105 | 218.3791 | 173.8948 | 149.8525 | 164.0314 | 186.2527 | 146.3579 | 174.6947 | 131.4631 | 130.5581 | 163.6844 | 207.179 | 166.3053 | 160.9262 | 222.5895 | 184.737 | 156.6947 | 149.179 | 202.2527 | 158.4317 | 310.8948 | 287.5895 | 156.7473 | 217.5157 | 225.5895 | 244.4529 | 290.9365 | 108.4317 | 122 | 117.6633 | 145.7684 | 144.9262 | 131.6633 | 141.5683 | 87.6947 | 267.9262 | 242.7261 | 258.9686 | 263.5581 | 133.9474 | 128.6735 | 128.6421 | 134.7684 | 77.0212 | 73.6947 | 62.5157 | 119.2316 | 211.737 | 203.8422 | 240.7261 | 153.179 | 145.3791 | 128.5683 | 138.916 | 116.9788 | 170.5369 | 175.2001 | 211.0424 | 160.1052 | 147.7158 | 135.4105 | 189.7582 | 150.4529 | 184.5581 | 200.7158 | 192.084 | 205.8736 | 166.7261 | 141.6209 | 234.1052 | 206.2842 | 194.4419 | 202.821 | 202.1264 | 296.1578 | 155.5157 | 172.7582 | 178.7896 | 222.7473 | 146.821 | 145 | 141.9051 | 176.2842 | 170.7787 | 183.179 | 185.3265 | 141.9474 | 107.7158 | 118.0526 | 111.0738 | 128.6947 | 271.5157 | 234.8313 | 231.2527 | 365.9686 | 242.3579 | 271.6107 | 235.4843 | 331.6947 | 151.179 | 79.1052 | 87.9262 | 151.179 | 112.0738 | 105.2316 | 105.5369 | 125.9474 | 105.5581 | 103.179 | 88.0526 | 113.3156 | 74.5683 | 63.6633 | 82.2316 | 92.4317 | 99.2842 | 90.5581 | 105.3893 | 76.179 | 247.5369 | 226.6735 | 227.9262 | 299.7684 | 157.6209 | 205.2104 | 193.8422 | 275.4631 | 279.6107 | 200.179 | 217.8634 | 322.263
1189 | 15.0734 | 150.9557 | 192.5752 | 231.2124 | 184.0177 | 147.823 | 167.0354 | 173.3451 | 156.4248 | 167.3894 | 127.1947 | 129.1681 | 172.9911 | 227.8585 | 166.9911 | 151.1858 | 244.062 | 179.9557 | 150.5929 | 141.5752 | 189.3451 | 156.5752 | 316.2035 | 296.4248 | 154.0089 | 224.4337 | 227.2035 | 245.3628 | 282.0354 | 128.239 | 135.8408 | 126.0089 | 143.9911 | 149.6284 | 136.3982 | 147.6372 | 93.0266 | 267.8142 | 231.8053 | 263.1858 | 278.4159 | 136.6018 | 119.8053 | 124.5929 | 134.1947 | 74.3716 | 68.7965 | 70.6372 | 114.3894 | 218.3894 | 202.177 | 228 | 152.1947 | 140.1592 | 140.6549 | 135.5575 | 116.6195 | 178.4159 | 187.8142 | 223.1947 | 166.8319 | 149.1947 | 142.823 | 192.5575 | 154.3716 | 199.4159 | 195.5752 | 185.2124 | 206.6106 | 163.6284 | 158.2744 | 222.7787 | 202.3982 | 191.0177 | 194.177 | 197.177 | 291.2035 | 159.4248 | 169.5397 | 166.5397 | 250.6992 | 155.6284 | 140.7876 | 135.4159 | 155.5486 | 172.2478 | 194.2301 | 197.2035 | 133.1681 | 117.6195 | 108.7787 | 97.9468 | 126.2035 | 253.1592 | 245.6815 | 259.2656 | 367.7787 | 245.2035 | 275.9911 | 236.9911 | 307.5397 | 130.938 | 84.6284 | 72.1681 | 130.938 | 101.3539 | 118.4425 | 117.0266 | 125.5929 | 101.761 | 104 | 92.0177 | 105.8142 | 71.6106 | 76.823 | 81.6018 | 100.8053 | 103.2213 | 92.7787 | 95.5929 | 66.7699 | 263.8408 | 199.7522 | 206.7522 | 303.4071 | 163.4425 | 198.8053 | 204.8142 | 269.9911 | 272.5575 | 200.3982 | 215.9557 | 331.8585
1190 | 15.0844 | 152.6672 | 201.7375 | 232.1578 | 192.9125 | 153.6141 | 190.9656 | 167.7547 | 163.8422 | 168.0703 | 136.4391 | 131.825 | 180.825 | 218.8937 | 155.7891 | 165.1938 | 251.2453 | 176.1406 | 146.5437 | 147.6672 | 189.8953 | 169.4219 | 330.5094 | 297.0875 | 168.2813 | 232 | 228 | 217.964 | 280.7891 | 121.6484 | 127.8062 | 128.6141 | 137.7016 | 157.8422 | 138.4563 | 142.4906 | 99 | 274.0703 | 234.8422 | 241.5781 | 262.964 | 132.0875 | 129.5969 | 123 | 127.4734 | 64.3156 | 67.3859 | 79.6141 | 107.7016 | 241.7907 | 195.7719 | 225.5437 | 141.7891 | 130.3859 | 138.2625 | 138.2813 | 124.6844 | 190.5969 | 187.3156 | 210.3328 | 170.9297 | 141.8594 | 149.2281 | 162.7359 | 145.5437 | 203.6141 | 185.7016 | 197.2109 | 226.8078 | 159.5609 | 161.4906 | 231.5094 | 212.4391 | 195.6141 | 190.8422 | 198.1406 | 325.1594 | 165.6141 | 171.3516 | 176.3344 | 255.8062 | 162.6141 | 149.6672 | 145.7547 | 164.2641 | 179.9297 | 201.6141 | 185.1047 | 136.4391 | 115.8594 | 102.1578 | 104.4219 | 134.9828 | 259.9656 | 262.7719 | 265.4031 | 351.9469 | 275.475 | 282.5969 | 264.475 | 333.3 | 139.7203 | 76.8765 | 66.8422 | 139.7203 | 95.2984 | 115.7187 | 113.175 | 131.3688 | 100.9828 | 113.2109 | 100.2984 | 105.3156 | 72.1578 | 70.3328 | 74.0172 | 92.7891 | 115.9828 | 107.65 | 90.3156 | 67.9828 | 276.0703 | 211.1063 | 213.8078 | 295.7891 | 161.3328 | 206.7547 | 212.9125 | 278.4391 | 301.0719 | 200 | 213.3688 | 345.6141
1191 | 15.0954 | 155.6155 | 224.7798 | 214.7148 | 180.3411 | 160.2635 | 187.2202 | 172.6264 | 163.1209 | 176.2635 | 122.9567 | 115.4512 | 175.0993 | 210.6372 | 156.2635 | 163.4729 | 228.2094 | 181.2527 | 144.3736 | 172.5379 | 211.1534 | 178.4946 | 314.5722 | 292.7473 | 167.231 | 213.8357 | 233.0108 | 217.2744 | 293.7906 | 124.5271 | 90.0668 | 117.7256 | 135.6264 | 159 | 134.3628 | 127.7365 | 92.1101 | 240.2978 | 228.4837 | 250.6697 | 261.0217 | 138.3953 | 130.6155 | 132.3953 | 136.5271 | 65.1318 | 75.1426 | 69.352 | 112.5163 | 241.0884 | 201.2635 | 232.769 | 159.8015 | 137.5163 | 138.3953 | 139.1101 | 115.7256 | 195.3736 | 171.2202 | 209.8899 | 171.5054 | 138 | 133.0884 | 175.0542 | 137.7365 | 197.7365 | 196.1534 | 213.648 | 250.5271 | 147.9892 | 158.6372 | 230.6047 | 208.352 | 193.4946 | 194.5054 | 199.4946 | 319.6931 | 174.769 | 172.231 | 166.5722 | 251.8899 | 159.2419 | 154.4946 | 138.0993 | 153.8249 | 179.2527 | 193.231 | 186.3953 | 152.3953 | 127.6589 | 113.5271 | 105.9783 | 134.3628 | 270.6264 | 239.4512 | 247.4729 | 348.1318 | 285.231 | 283.6155 | 254.1877 | 321.0451 | 147.7365 | 66.7473 | 64.8682 | 147.7365 | 109.9007 | 113.6372 | 102.6155 | 118.4621 | 110.3845 | 120.2527 | 114.2744 | 104.8791 | 77.8899 | 67.3845 | 77.1426 | 88.8791 | 109.0993 | 108.0993 | 99.2744 | 78.0108 | 269.231 | 236.2744 | 237.5271 | 298.769 | 173.417 | 207.8682 | 195.9567 | 283.1209 | 277.3971 | 198.1209 | 202.9675 | 328.4621
1192 | 15.1064 | 165.7729 | 222.143 | 217.7729 | 182.4956 | 167.1932 | 169.4203 | 181.941 | 142.8406 | 174.2522 | 116.7478 | 102.6386 | 174.3865 | 196.1179 | 147.2271 | 160.4705 | 207.4203 | 172.4203 | 141.4454 | 158.0928 | 197.9247 | 161.0339 | 310.2183 | 306.6888 | 160.0841 | 219.6048 | 238.5546 | 224 | 308.8319 | 127.5295 | 102.739 | 116.1092 | 133.4454 | 153.2522 | 139.3024 | 130.3865 | 88 | 228.5797 | 235.4956 | 266.8319 | 250.3952 | 137.6135 | 137.5797 | 139.2773 | 138.2522 | 68.2773 | 72.3363 | 64.9159 | 132.9661 | 229.0841 | 209.4705 | 229.6976 | 161.8657 | 145.1932 | 136.3363 | 127.9749 | 122.4115 | 189.8908 | 152.143 | 210.1681 | 167.2522 | 155.8821 | 138.9661 | 182.3363 | 139.1092 | 179.9498 | 202.7227 | 221.2773 | 248.4203 | 162.8821 | 162 | 217.3363 | 202 | 192.6386 | 196.6386 | 209.4956 | 309.7478 | 170.4203 | 150.3952 | 160.941 | 234.9247 | 132.0928 | 156.1932 | 138.0251 | 152.1343 | 172.3363 | 198.857 | 189.4454 | 149.6976 | 126.7817 | 116.5295 | 108.941 | 125.2522 | 257.5885 | 238.1343 | 255.3275 | 357.0251 | 254.4542 | 266.3112 | 256.1594 | 299 | 134.4203 | 71.7478 | 63 | 134.4203 | 112.5295 | 118.2773 | 108.3024 | 123.3275 | 93.8406 | 107.5885 | 105.0339 | 102.1681 | 82.6386 | 77.6637 | 80.0841 | 91.2773 | 107.7478 | 88.8657 | 97.6976 | 79.0841 | 264 | 234.059 | 226.4793 | 319.3275 | 180.5295 | 212.3865 | 207.6299 | 312.0163 | 259.857 | 206.5797 | 197.1932 | 327.5797
1193 | 15.1174 | 170.3963 | 193.2172 | 210.6794 | 179.8869 | 176.8112 | 155.5378 | 170.0757 | 140.4622 | 157.9813 | 115.4434 | 116.6697 | 198.83 | 206.179 | 147.1603 | 162.3491 | 207.8584 | 181.9715 | 154.9715 | 144.6509 | 197.0659 | 159.1603 | 302.3303 | 295.4719 | 165.5094 | 223.1416 | 250.415 | 216.1888 | 296.0285 | 123.0472 | 126.8112 | 113.7925 | 122.8869 | 161.0659 | 142.6981 | 133.3491 | 86.6981 | 228.1416 | 232.8869 | 262.4906 | 242.9528 | 143.1131 | 151.4622 | 151.0659 | 124.585 | 74.2075 | 65.3963 | 69.2547 | 127.6794 | 227.3491 | 209.3963 | 234.764 | 150.4434 | 150.2547 | 139.1603 | 143.5281 | 126.7453 | 181.1416 | 157.7169 | 208.6509 | 177.6697 | 172.5094 | 166.8772 | 186.4622 | 143.9528 | 192.83 | 202 | 221.3491 | 253.415 | 165.1888 | 162.6509 | 225.3678 | 206.5566 | 195.6037 | 185.2831 | 199.0757 | 305.1888 | 176.7169 | 156.6225 | 157.5378 | 213.236 | 132.9715 | 156.0472 | 155.0187 | 169.415 | 162.7925 | 219.9715 | 176.9341 | 138.4906 | 115.1416 | 115.9528 | 111.3963 | 129.1603 | 247.3963 | 246.3019 | 243.821 | 357.0944 | 236.7453 | 265.1603 | 229.8966 | 308.764 | 130.9528 | 71.7453 | 76.0187 | 130.9528 | 109.3491 | 117.6981 | 110.3963 | 107.9154 | 103.83 | 98.0472 | 97.9528 | 113.6697 | 79.7453 | 85.9056 | 80.9528 | 101.1131 | 96.0285 | 93.0659 | 91.6981 | 77.3491 | 260.7453 | 225.7453 | 231.6225 | 307.821 | 175.3963 | 202.3303 | 187.8022 | 267.1611 | 291.3865 | 215.2547 | 218.5281 | 336.9056
1194 | 15.1284 | 167.6735 | 200.8968 | 198.3574 | 179.6426 | 165.7458 | 146.3574 | 166.3058 | 156.9381 | 156.969 | 120.2955 | 119.3574 | 180.818 | 222.3058 | 149.0103 | 156.031 | 210.3368 | 190.6632 | 154.378 | 156.2749 | 219.5807 | 164.9897 | 319.5394 | 296.2749 | 165.0206 | 239.8968 | 241.409 | 212.6632 | 292.6426 | 118.6839 | 112.4296 | 119.622 | 137.2336 | 163.6839 | 147.969 | 146.2645 | 101.9175 | 256.1717 | 235.2955 | 268.2852 | 230.0722 | 132.7458 | 138.7561 | 142.409 | 142.8762 | 77.6632 | 71.9587 | 81.6116 | 131.2749 | 224.3471 | 207.3368 | 235.3574 | 170.5497 | 155.3161 | 150.2955 | 146.0413 | 121.6839 | 175.3471 | 176.6013 | 212.3161 | 181.6839 | 150.7974 | 165.0619 | 181.7148 | 142.3471 | 190.0722 | 172.818 | 222.9897 | 259 | 164.3161 | 151.0619 | 233.9897 | 216.2955 | 193.0206 | 190.2749 | 184.031 | 311.6116 | 184.3265 | 172.9587 | 154.3265 | 215.9587 | 140.3368 | 160.969 | 131.4915 | 176.9897 | 181.2233 | 194.1754 | 164.3677 | 140.969 | 113.9897 | 114.3471 | 111.9897 | 127.6942 | 275.8452 | 239.0413 | 236.3161 | 364.2852 | 251.5807 | 272.3161 | 231.8865 | 335.8865 | 133.3265 | 70 | 86.3161 | 133.3265 | 109.6632 | 113.0206 | 107.0103 | 108.9278 | 111.0206 | 94.3471 | 104.3058 | 113.7045 | 86.622 | 78.0516 | 72.7148 | 94.0619 | 101.2645 | 101.6529 | 116.2026 | 80.9794 | 264.3058 | 234.6116 | 241.9897 | 304.9587 | 164.0516 | 201.6323 | 203.1614 | 248.591 | 297.0413 | 193.787 | 201.8077 | 316.4503
1195 | 15.1394 | 169.0266 | 204.2448 | 206.1328 | 184.0266 | 151.9203 | 165.6166 | 162.9203 | 160.2979 | 177.5635 | 120.6224 | 133.888 | 180.8614 | 205.41 | 148 | 157.0531 | 213.3776 | 178.1652 | 138.5162 | 153.2183 | 215.1652 | 162.6224 | 310.3834 | 320.9145 | 161.649 | 217.5751 | 227.9203 | 233.2656 | 292.9734 | 110.9203 | 118.5369 | 104.41 | 132.8141 | 152.5427 | 135.4631 | 153 | 96.4896 | 270.649 | 259.2656 | 246.6548 | 245.9676 | 143.239 | 124.5959 | 127.5693 | 126.6282 | 81.3776 | 70.5959 | 80.9203 | 132.9469 | 216.2448 | 194.8407 | 235.0266 | 173.2183 | 145.5162 | 145.8938 | 129.8407 | 118.649 | 180.7552 | 169.4896 | 214.6755 | 168.5162 | 156.9145 | 156.9734 | 180.3776 | 126.1386 | 167.4631 | 180.2921 | 239.5369 | 267.1062 | 170.0531 | 163.9145 | 251.888 | 228.1062 | 191 | 200.0531 | 201.2656 | 313.6224 | 174.1917 | 172.9469 | 166.4838 | 234.8614 | 164.3187 | 172.7817 | 136.2656 | 159.0855 | 193.351 | 183.7552 | 185.9942 | 143.3245 | 121.0797 | 110.2979 | 117.7286 | 121.6224 | 270.0589 | 235.6755 | 239.351 | 369 | 260 | 252.3834 | 260.5635 | 338.2183 | 144.8083 | 70.6755 | 77.1917 | 144.8083 | 113.3776 | 128.5635 | 110.0531 | 112.6224 | 113.7286 | 90.9734 | 104.9734 | 107.2979 | 76.8141 | 79.0797 | 68.6755 | 96.7817 | 103.9469 | 96.2448 | 105.3569 | 91.7817 | 245.3834 | 231.2183 | 218.6813 | 305.6224 | 156.2979 | 221.2124 | 209.8672 | 267.4838 | 276.112 | 203.6166 | 194.0797 | 323.239
1196 | 15.1504 | 155.556 | 220.9465 | 208.9366 | 194.6293 | 161.3806 | 170.4975 | 151.0585 | 161.7512 | 168.117 | 132.0684 | 131.6829 | 168.7413 | 215.9465 | 154.8781 | 150.7463 | 219.1269 | 182.3171 | 153.6343 | 155.1903 | 204.8731 | 168.5659 | 305.5025 | 328.6244 | 165.8147 | 224.0099 | 199.5511 | 253.3171 | 309.1952 | 116.2537 | 115.6829 | 102.5659 | 130.8147 | 149.3756 | 134.8781 | 149.561 | 97.5659 | 254.8682 | 277.9415 | 229.1853 | 259.0634 | 128.6145 | 141.2587 | 122.6244 | 131.5709 | 78.8731 | 85.1952 | 60.8048 | 126.8731 | 207.4975 | 205.5074 | 236.6878 | 153.8682 | 147.5659 | 134.4341 | 139.8196 | 118.6878 | 165.4291 | 174.005 | 209.4975 | 180.883 | 166 | 146.3707 | 179.2488 | 152.7027 | 165.5025 | 186.1853 | 215.3608 | 254.4926 | 181.6293 | 159.2438 | 220.7949 | 223.0585 | 217.1368 | 187.556 | 204.1219 | 327.1318 | 173.8147 | 164.1219 | 176.1269 | 226.8682 | 148.4876 | 164.6194 | 148.7512 | 165.8196 | 197.439 | 190.439 | 196.6244 | 131.3072 | 114.3707 | 104.1853 | 112.4341 | 133.0684 | 274.444 | 260.0733 | 263.3855 | 391.0099 | 251.0585 | 224.8048 | 242.1754 | 349.1318 | 150 | 64.1219 | 78.8781 | 150 | 108.8097 | 114.9901 | 106.4975 | 128.1952 | 102.2438 | 90 | 99.8731 | 101.1853 | 76.1903 | 72.3707 | 69 | 109.2537 | 104.0634 | 98.5025 | 88.4975 | 84.995 | 259.7611 | 229.7512 | 218.005 | 317.0684 | 177.0099 | 222.1219 | 194.6829 | 268.1853 | 293.4489 | 203.6829 | 219.6977 | 354.0099
1197 | 15.1614 | 146.8997 | 217.3982 | 217.8014 | 196.1996 | 167.7001 | 170.8004 | 170.8034 | 180.5025 | 173.3019 | 124.6981 | 124.8997 | 187.3039 | 218.6991 | 161.5005 | 149.1003 | 222.4002 | 192.6008 | 155.999 | 143.998 | 217.7021 | 160.7984 | 315.001 | 328 | 155.3982 | 227.6991 | 242.6078 | 265.7011 | 307.1986 | 108.7984 | 113.1003 | 132.6038 | 123.1986 | 152.1003 | 135.1996 | 145.8997 | 95.3992 | 253.6008 | 237.8937 | 231.2006 | 268.4012 | 128.5015 | 141.5988 | 129.7011 | 132.999 | 81.9007 | 70.5968 | 68.4022 | 117.2989 | 199.3992 | 223.5015 | 227.1986 | 157.5015 | 144.6991 | 126.7994 | 139.999 | 107.7984 | 173.8024 | 176.8997 | 202.7994 | 187.5998 | 179.3019 | 148.3009 | 176.5998 | 151.1976 | 185.5025 | 175.5988 | 201 | 259.6018 | 185.2999 | 171.2026 | 205.1003 | 222.5005 | 193.2949 | 176.7994 | 201.7001 | 334 | 176.7001 | 168.7011 | 164.6981 | 246.6038 | 168.9047 | 166.001 | 149.2999 | 161.7984 | 192.6991 | 192 | 193.8997 | 129.5005 | 116.3009 | 105.5005 | 103.3992 | 124.998 | 258.5968 | 254.1976 | 232.6941 | 406.6008 | 276.4042 | 218.4002 | 230 | 348.2989 | 139.4985 | 63.1003 | 86.2006 | 139.4985 | 108.8004 | 105.7001 | 126.4032 | 123.3982 | 100.9007 | 104.7021 | 105.7011 | 122.8034 | 84.6008 | 67.2999 | 75.3009 | 101.7984 | 103.5998 | 94.6991 | 97.2016 | 85.6008 | 263.2989 | 266.005 | 225.1003 | 314.5988 | 184.1996 | 219.7001 | 189.2999 | 260.3992 | 300.0993 | 224.2036 | 223.6991 | 322.6941
1198 | 15.1724 | 158.823 | 173.5309 | 199.2035 | 180.7522 | 175.1239 | 159.177 | 163.9026 | 169.4778 | 186.8363 | 118.2876 | 122.5752 | 189.0265 | 203.177 | 143.053 | 143.5885 | 218.0133 | 190.0133 | 161.5487 | 155.8098 | 219.7257 | 166.6859 | 344.3585 | 296.6548 | 153.562 | 232.8363 | 220.4069 | 264.0133 | 282.3406 | 132.4957 | 105.4513 | 122.6283 | 133.9602 | 142.3141 | 143.9735 | 149.2743 | 97.9867 | 273.0974 | 229.6859 | 241.5487 | 251.3406 | 149.385 | 118.053 | 143.6859 | 125.0133 | 78.3009 | 81.6594 | 71.438 | 126.823 | 215.5222 | 222.3009 | 227.9867 | 154.8761 | 156.9602 | 124.2876 | 151.2478 | 117.2478 | 183.1372 | 185.2611 | 209.5487 | 169.1902 | 165.7654 | 141.7389 | 186.6859 | 146.8496 | 151.6813 | 177.6991 | 205.2743 | 260.0133 | 172.8893 | 169.0265 | 224.5222 | 206.9026 | 202.2213 | 201.3585 | 172.7919 | 322.6017 | 172.0133 | 172 | 151.8761 | 242.3274 | 170.8893 | 146.2035 | 153.2743 | 174.0974 | 203.5354 | 211.2346 | 183.7389 | 156.6461 | 126.1239 | 121.2478 | 105.9867 | 141.7965 | 248.2876 | 286.1815 | 244.2081 | 375.5176 | 281.1637 | 238.947 | 232.8496 | 339.3009 | 149.9602 | 73.9735 | 75.177 | 149.9602 | 124.9602 | 118.1107 | 124.6017 | 128.6859 | 104.4248 | 115.9867 | 106.8628 | 119.4646 | 74.8893 | 66.2876 | 68.7389 | 107.6859 | 110.1239 | 111.2346 | 109.8363 | 86.5752 | 262.1372 | 240.3937 | 205.3406 | 306.7257 | 173.0265 | 204.3274 | 203.2478 | 264.4115 | 302.9867 | 220.7522 | 216.0133 | 295.0265
1199 | 15.1834 | 168.3481 | 194.9588 | 203.0443 | 193.1171 | 164.9557 | 153.2753 | 172.2183 | 159.826 | 167.5348 | 112.9272 | 116.9272 | 174.8544 | 215.3924 | 145.8702 | 157.6677 | 227.5949 | 169.8829 | 159.2025 | 173.8702 | 210.0285 | 169.5506 | 362.9715 | 293.4209 | 171.6677 | 223.6804 | 260.6993 | 243.1583 | 292.1171 | 135.3038 | 115.769 | 107.4779 | 136.3766 | 142.3481 | 131.3323 | 147.3766 | 84.057 | 258.2595 | 239.0728 | 239.2025 | 265.4652 | 171.2183 | 136.8133 | 154.5221 | 120.1013 | 73.826 | 70.4336 | 74.3481 | 113.1583 | 205.6076 | 220.7247 | 226.3766 | 146.9272 | 147.057 | 139.2183 | 146.8544 | 115.0285 | 197.769 | 167.9842 | 225.3196 | 175.769 | 155.1013 | 150.3196 | 178.6804 | 159.5949 | 158.9145 | 183.6234 | 225.8417 | 238.4336 | 175.9715 | 162.1013 | 234.8987 | 213.769 | 219.9715 | 220.6962 | 189.9873 | 319.4494 | 175.7975 | 180.6962 | 151.8987 | 225.1298 | 169.6234 | 162.3639 | 159.3481 | 172.3038 | 222.769 | 203.7817 | 194.4937 | 140.1867 | 126.826 | 128.4494 | 107.2753 | 148.826 | 257.4209 | 275.1867 | 234.9842 | 383.7405 | 256.9842 | 237.5791 | 260.8133 | 329.0285 | 148.7532 | 75.1013 | 73.6234 | 148.7532 | 120.8544 | 115.7532 | 112.7532 | 120.6804 | 126.0158 | 118.7247 | 103.1013 | 110.3766 | 87.3924 | 73.2468 | 70.0728 | 109.1013 | 118.0728 | 103.057 | 102.8544 | 86.7247 | 257.9272 | 229.7975 | 199.174 | 289.7817 | 170.4494 | 220.4652 | 209 | 278.5949 | 312.2468 | 204.8544 | 222.6962 | 317.0886
1200 | 15.1944 | 161.8933 | 206.7891 | 208.737 | 191.1563 | 170.3176 | 157.4218 | 182.4218 | 158.263 | 173.7395 | 139.005 | 118.6849 | 190.1613 | 202.1017 | 148.5261 | 157.3672 | 240.1067 | 196.9008 | 162.1588 | 175.7891 | 201.1042 | 162.3672 | 336.5211 | 325.005 | 180.9479 | 255.8487 | 246.8883 | 223.4714 | 284.2605 | 136.4218 | 104.0496 | 110.8958 | 129.8412 | 155.7916 | 162.5857 | 151.1588 | 85.3697 | 264.0025 | 288.1663 | 251.7395 | 273.263 | 175.5261 | 144.0521 | 133.4168 | 113.8412 | 61.2084 | 66.6849 | 73.7891 | 112.6328 | 215.2134 | 235.7395 | 248.5832 | 182.5857 | 141 | 149.4218 | 129.7345 | 123.0546 | 200.0521 | 166.6328 | 210.8387 | 189.1067 | 158.4218 | 159.4218 | 192.4243 | 155.8933 | 167.263 | 208.5832 | 232.263 | 253.8462 | 183.4218 | 161 | 236.737 | 235.9504 | 225.2109 | 217.3672 | 198.0521 | 318.5261 | 163.9975 | 162.6278 | 143.4193 | 232.0546 | 177.6328 | 168.3151 | 148.4714 | 174.8958 | 228 | 212.7395 | 187.4714 | 155.0571 | 127.4739 | 120.8933 | 106.263 | 130.3126 | 283.1092 | 299.6378 | 249.8462 | 385.3672 | 256.3697 | 238.4218 | 271 | 342.9504 | 146 | 82.1067 | 60.9975 | 146 | 114.0521 | 117.4218 | 118.8437 | 133.6874 | 114.8387 | 110.1563 | 110.1067 | 125.9504 | 82.2084 | 73.7891 | 83.7916 | 122.7395 | 105.2605 | 108.7916 | 105.6328 | 84.7891 | 265.5807 | 224.6303 | 221.3722 | 300.9504 | 182.0546 | 205.4168 | 204.5782 | 280.0521 | 322.3697 | 204.6849 | 212.7345 | 312.2605
1201 | 15.2054 | 144.0147 | 195.5103 | 209 | 175.2625 | 173.2507 | 159 | 173.5103 | 157.2507 | 176.0029 | 138.5103 | 105.764 | 188.7581 | 215.9794 | 152.4956 | 172.2331 | 228.0147 | 191.0176 | 152.761 | 155.5191 | 214.7346 | 157.0029 | 311.764 | 315.5191 | 172.2596 | 254.0147 | 236.2478 | 249.7199 | 300.7287 | 128.2596 | 108.4897 | 113.7493 | 147.4809 | 161.4985 | 152.0235 | 141.761 | 86.5015 | 261.5074 | 263.0412 | 245.761 | 255.0176 | 168.2566 | 143.7493 | 130.9941 | 126.236 | 70.4868 | 58.2596 | 64.0088 | 132.9824 | 210.5103 | 236.5044 | 256.2507 | 180.2654 | 157.4838 | 141.2596 | 133.2419 | 127 | 184.764 | 197.4721 | 221.2331 | 197.2449 | 157.0029 | 183.4779 | 167.5309 | 154.4985 | 167.7493 | 183.2831 | 240.9912 | 253.7581 | 180.5044 | 167.7434 | 241.4956 | 236.0059 | 208.0176 | 209.0059 | 180.5162 | 341.9765 | 175.4838 | 169.236 | 144.4956 | 262.9735 | 174.0059 | 141.525 | 149.9941 | 182.2449 | 209.2684 | 211.2566 | 189.7434 | 164 | 116.761 | 107.5103 | 105.2507 | 124 | 272.2684 | 257.3037 | 229.5279 | 383.7493 | 266.4926 | 215.2743 | 254.5162 | 334.0147 | 144.5015 | 85.7493 | 85.2213 | 144.5015 | 108.5044 | 115.2537 | 131.7404 | 137.0029 | 116.9912 | 112.9941 | 98.0147 | 122.2596 | 81.7463 | 73.7493 | 82.7551 | 107.7699 | 97.7522 | 115.9971 | 96.761 | 71.2625 | 253.2654 | 248.2243 | 228.2507 | 280.7757 | 189.7463 | 230.7169 | 209.7434 | 301.4779 | 300.2743 | 217.9882 | 219.9882 | 326.4809
1202 | 15.2164 | 139 | 195.8078 | 225.7542 | 178.6156 | 160.8151 | 149.0998 | 157.8151 | 172.2311 | 178.8078 | 136.5231 | 130.7007 | 193.6156 | 211.5767 | 160.0924 | 158.9611 | 221.4769 | 204.0389 | 187.8393 | 143.6691 | 210.0998 | 150.6691 | 330.6082 | 282.3455 | 175.854 | 242.9076 | 262.1313 | 247.8151 | 329.3236 | 135.6618 | 112.7616 | 117.8078 | 141.0536 | 156.6691 | 131.8151 | 146.3771 | 95.9002 | 269.6618 | 267.2773 | 273.9853 | 242.9076 | 154.5767 | 148.5693 | 135.2847 | 132.5231 | 79.5693 | 63.3771 | 75.4696 | 138.2384 | 217.6618 | 210.6302 | 260.5693 | 183.3771 | 160.7153 | 131.9076 | 145.1387 | 131.5693 | 191.4233 | 191.7689 | 229.2847 | 214.2311 | 169.708 | 159.0147 | 179.8467 | 163.3771 | 192.3698 | 184.9464 | 240.1922 | 274.6082 | 173.6691 | 155.5304 | 241.4769 | 250.7542 | 220.2773 | 190.2458 | 195.562 | 337.0536 | 167.292 | 179.3309 | 163.5158 | 277.3309 | 164.3844 | 174.8856 | 164.9464 | 186.2847 | 206.8078 | 208.2384 | 194.2847 | 154.8613 | 121.3771 | 100.9538 | 112.6156 | 141.5158 | 243.1533 | 264.1313 | 224.5693 | 392.3771 | 269 | 224.5158 | 242.146 | 316.0536 | 147.0462 | 70.7689 | 76.7227 | 147.0462 | 132.1313 | 118.5693 | 125.0998 | 133.7153 | 120 | 111.9538 | 111.2773 | 96.9149 | 82.2384 | 76.2847 | 89.3771 | 107.0924 | 107.6618 | 117.7616 | 82.3382 | 82.2311 | 262.4696 | 237.9611 | 234.0924 | 316.1702 | 166.6302 | 212.2993 | 223.4233 | 299.0998 | 285.146 | 200.6764 | 211.0536 | 353.562
1203 | 15.2274 | 165.3108 | 195.4523 | 234.0954 | 210.4062 | 154.6785 | 161.4769 | 205.0739 | 186.2862 | 194.7031 | 150.1554 | 135.3569 | 196 | 225.0246 | 153.4877 | 160.7385 | 208.6185 | 189.1061 | 172.9154 | 152.8339 | 213.1908 | 174.5369 | 335.6785 | 277.0954 | 193.4769 | 247.1908 | 263.0354 | 241.6785 | 297.3077 | 128.94 | 120.7385 | 105.8446 | 137 | 171.2508 | 147.3462 | 145.1308 | 95.1308 | 250.5584 | 259.8446 | 253.82 | 256.4769 | 133.9754 | 142.2615 | 153.7985 | 134.5477 | 82.5477 | 69.0954 | 82.3215 | 145.7385 | 249.6323 | 212.2862 | 248.0708 | 178.2615 | 157.6785 | 144.7031 | 139.4877 | 124.4877 | 210.4769 | 184.6785 | 223.8092 | 195.0107 | 175.5477 | 164.4769 | 183.1308 | 162.9046 | 181.4277 | 200.6077 | 229.7138 | 237.1169 | 168.9046 | 155.6431 | 233.2615 | 248.2615 | 211.2969 | 205.12 | 196.5831 | 348.4769 | 178.4769 | 184.8692 | 167.4523 | 244.9507 | 201.4662 | 160.9154 | 166.6785 | 187 | 209.5477 | 216.5123 | 212.7985 | 171.3462 | 120.9046 | 116.2508 | 104.1661 | 127.6538 | 304.8724 | 268.9046 | 237.6077 | 418.2154 | 272.0954 | 211.4277 | 237.6785 | 318.9646 | 138.7138 | 66.7738 | 72.5477 | 138.7138 | 131.4877 | 128.5123 | 111.94 | 125.2615 | 116.9046 | 120.2862 | 116.2262 | 117.0846 | 89.7385 | 78.5477 | 95.0954 | 107.4523 | 97.0708 | 117.2262 | 84.4169 | 92.4169 | 265.4523 | 260.6323 | 234.4523 | 317.6185 | 206.2047 | 213.06 | 203.7846 | 299.0954 | 294.6077 | 202.5123 | 203.1308 | 353.8092
1204 | 15.2384 | 162.7802 | 186.3525 | 220.0634 | 197.7743 | 175.2257 | 173.0752 | 186.1961 | 175.6357 | 191.1386 | 146.9248 | 114.3466 | 218.0119 | 233.9307 | 141.5663 | 150.4218 | 225.4396 | 204.2257 | 168.1446 | 157.5723 | 218.1446 | 178.8554 | 323.9941 | 329.099 | 187.7802 | 237.2079 | 227.1961 | 256.7228 | 314.301 | 129.9307 | 119.0693 | 112.2198 | 151.1505 | 156.3466 | 160.0752 | 139.2832 | 97.9307 | 277.0178 | 263.8614 | 256.7921 | 249.2079 | 140.7921 | 134.497 | 148.7802 | 135 | 84.5723 | 64.497 | 80.6416 | 118.1268 | 244.6357 | 222.8614 | 243.2139 | 167.3525 | 141.2772 | 134.8495 | 138.5723 | 122 | 200.8495 | 184 | 207.8495 | 181.7109 | 177.5723 | 159.5663 | 224.4514 | 170.6475 | 192.5089 | 202.4277 | 253.7287 | 257.0178 | 205.7346 | 186.8732 | 254.5841 | 264.8673 | 202.2832 | 207.0693 | 182.4218 | 336.4911 | 176.7109 | 177.3525 | 169.3584 | 246.0059 | 168.1901 | 171.0812 | 171.503 | 187.7861 | 227.295 | 211.1386 | 188.9129 | 143.9822 | 131.7921 | 119.4277 | 119.8673 | 135.3643 | 284.907 | 272.7168 | 259.0812 | 409.2772 | 277.7168 | 235.0871 | 244.8614 | 338.295 | 139.9307 | 78.7921 | 87.1505 | 139.9307 | 136.0752 | 110.5604 | 116.0752 | 128.503 | 119.9307 | 120.6416 | 104.2079 | 130.503 | 75.4911 | 90.0059 | 84.9941 | 95.9941 | 104.7921 | 108.3525 | 100.1505 | 97.9307 | 265 | 267.4277 | 243.4337 | 337.5841 | 208.2079 | 213.6416 | 210.3643 | 304.7168 | 315.295 | 197.9248 | 217.7228 | 366.1505
1205 | 15.2494 | 175.1701 | 177.6126 | 228.7749 | 190.4031 | 169.822 | 159.8299 | 164.2251 | 171.2016 | 181.0157 | 165.7592 | 119.3796 | 202.4424 | 223.0236 | 148.5811 | 159.7749 | 231 | 208.4031 | 165.0079 | 162.7906 | 249.3403 | 179.5969 | 316.2094 | 317.4503 | 182.6047 | 237.1937 | 265.1073 | 265.7906 | 321.2016 | 140.5811 | 131.5733 | 103.822 | 146.2173 | 152.5969 | 158.8063 | 148.3796 | 102.1937 | 284.4031 | 260.411 | 268.7827 | 269.1544 | 159.9686 | 141.7827 | 144.4031 | 167.7356 | 96.9764 | 55.8141 | 72.0157 | 121.178 | 230.6204 | 221.8063 | 215.055 | 174.5811 | 160.9529 | 146.1701 | 122.233 | 129.9843 | 210.5733 | 180.0079 | 226.356 | 187.1859 | 193.9686 | 153.0079 | 203.2644 | 176.1937 | 177.8377 | 198.0079 | 243.4346 | 269.1937 | 203.2251 | 177.4346 | 229.0628 | 223.6911 | 190.6204 | 206.7984 | 195.767 | 339.9843 | 177.3953 | 184.5811 | 171.5969 | 276.1466 | 146.4189 | 169.6126 | 176.9921 | 196.7827 | 224.8141 | 212.1937 | 190.5811 | 160.5497 | 134.2016 | 125.3874 | 102.644 | 126.2251 | 309.9293 | 279.589 | 230.466 | 387.4346 | 261.4346 | 239.0079 | 241.411 | 343.7984 | 134.6126 | 66.0314 | 67.8456 | 134.6126 | 132.411 | 116.178 | 121.1937 | 128.4031 | 124.1937 | 114.411 | 117.767 | 120.0236 | 68.6047 | 99.3874 | 74.8141 | 103.3796 | 102.411 | 110.7906 | 102.4031 | 83.8299 | 261.0079 | 239.8534 | 241.2094 | 296.0943 | 185.0393 | 223.3796 | 219.589 | 291.6283 | 315.2094 | 205.5811 | 233.9764 | 357.2251
1206 | 15.2604 | 144.9498 | 217.3467 | 218.2178 | 197.2965 | 163.7571 | 164.1072 | 167.4858 | 174.2429 | 198.4573 | 171.8107 | 121.1893 | 208.3501 | 229.7286 | 154.2429 | 145.9749 | 243.9715 | 199.0821 | 188.3216 | 159.1357 | 224.5712 | 162.9749 | 350.6717 | 328.8358 | 175.5142 | 236.3786 | 253.4891 | 248.3534 | 291.8141 | 136.5142 | 123.6499 | 109.1072 | 143.1893 | 157.8643 | 148.2714 | 156.675 | 109.4858 | 279.1357 | 241.1642 | 271 | 270.9464 | 158.325 | 148.8643 | 149.675 | 149.2462 | 102.4322 | 62.1072 | 79.7286 | 133.7286 | 230.4322 | 238.8358 | 234.7538 | 186.7286 | 155.6499 | 145.9464 | 133.4037 | 119.0285 | 206.7035 | 186.2965 | 235.2429 | 203.593 | 166.3819 | 166.593 | 201.4858 | 172.1357 | 158.407 | 213.2144 | 220.3534 | 269.1893 | 178.9213 | 170.5678 | 238.8358 | 233.8894 | 200.1608 | 219.1608 | 210.5394 | 335.5142 | 195.8358 | 199.9715 | 152.5427 | 261.9213 | 171.5645 | 170.4322 | 156.1106 | 181.1642 | 221.3786 | 224.3501 | 208.4037 | 162.9464 | 128.325 | 132.675 | 109.1608 | 126.2429 | 300.3534 | 273.7035 | 253.6181 | 374.0821 | 246.4606 | 259.8894 | 244.0536 | 295.3568 | 137.0536 | 66.8643 | 74.1608 | 137.0536 | 130.1893 | 123.0536 | 134.1608 | 145.0251 | 143.6466 | 111.3786 | 94.4355 | 139.7002 | 89.8894 | 93.7035 | 79.4858 | 110.0536 | 116.4037 | 107.1357 | 102.8107 | 81.6214 | 269.7286 | 252.4573 | 236.7571 | 323.7253 | 164.5963 | 217.8928 | 231.5394 | 278.2714 | 337.5109 | 218.5394 | 210.2462 | 346.7035
1207 | 15.2714 | 146.8761 | 206.4247 | 223.2301 | 192.4159 | 168.7611 | 173.4071 | 163.2389 | 161.8318 | 211.2301 | 161.3009 | 114.4159 | 232.3983 | 242.6991 | 157.469 | 170.8054 | 242.0619 | 212.6372 | 184.9469 | 160.469 | 231.8142 | 185.3364 | 340.8938 | 320.6548 | 175.646 | 255.7523 | 283.3895 | 234.1239 | 289.1151 | 130.0619 | 123.469 | 117.5841 | 158.6372 | 182.0443 | 141.0619 | 126.7256 | 105.2389 | 250.0176 | 262.5133 | 249.6017 | 274.1151 | 152.0619 | 137.6548 | 160.0531 | 149.5841 | 84.8938 | 65.646 | 59.7787 | 122.0088 | 231.823 | 219.9557 | 236.8849 | 180.7699 | 158.7611 | 163.9293 | 139.469 | 139.0443 | 194.3009 | 185.531 | 218.7168 | 200.4159 | 194.3895 | 176.5841 | 202.177 | 169.354 | 187.0974 | 216.177 | 216.823 | 259.1239 | 200.3364 | 181.5222 | 230.6548 | 237.354 | 217.8142 | 217.0619 | 199.0088 | 321.6548 | 194.2389 | 209.5841 | 161.9912 | 280.0443 | 182.1151 | 184.9912 | 169.9293 | 177 | 236.6372 | 227 | 200.4778 | 177.6372 | 135.2301 | 129.8849 | 110.354 | 131.9381 | 283.6548 | 304.9204 | 238.7787 | 400.8054 | 244.823 | 241.1327 | 248.292 | 303.7523 | 165.1594 | 65.531 | 80.292 | 165.1594 | 127.531 | 137.9912 | 131.2389 | 144.8849 | 137.3009 | 125.8142 | 101.1682 | 134.3009 | 91.708 | 78.0088 | 81 | 121.6991 | 112.5929 | 110.9381 | 109.5841 | 72.1239 | 272 | 277.5753 | 241.7611 | 295.9645 | 197.2125 | 216 | 227.4159 | 288.3452 | 310.0796 | 217.708 | 204 | 353.2301
1208 | 15.2824 | 162.3648 | 217.0354 | 224.1647 | 212.7179 | 154.9646 | 171.6588 | 181.2119 | 162.3412 | 211.3294 | 142.294 | 128.8707 | 206.9291 | 243.3294 | 153.8235 | 166.1411 | 239.3294 | 211.8235 | 183 | 180.2119 | 231.6588 | 163.435 | 297.7408 | 315.4941 | 177.6706 | 260.8353 | 280.1411 | 237.0118 | 320.0709 | 141.5295 | 120.6588 | 120.6706 | 147.7999 | 180.3176 | 148.353 | 137.5413 | 107.3412 | 244.8353 | 265.4941 | 248.3412 | 297.5532 | 146.8235 | 158.3885 | 156.9882 | 137.6352 | 92.6942 | 65.1647 | 74.2119 | 122.3412 | 211.9527 | 238.3885 | 258.5532 | 178.1647 | 151.647 | 147.9527 | 158.3766 | 138.1529 | 201.1883 | 196.6942 | 239.2238 | 159.7408 | 213.6942 | 171.3176 | 201.1647 | 166.4941 | 179.7999 | 223.5177 | 237.0473 | 285.4003 | 190.9646 | 156.435 | 258.9062 | 261.2238 | 226.0118 | 222.6824 | 192.6588 | 364.9416 | 188.8235 | 185.1056 | 175.8589 | 285 | 185.5059 | 171.294 | 157.294 | 205.4003 | 230.8117 | 242.0354 | 226.4003 | 162.6234 | 113.6115 | 113.9646 | 123.3648 | 135.5059 | 312.7415 | 259.376 | 231.4941 | 379.435 | 256.6942 | 257.7179 | 255.6824 | 348.9298 | 181.8589 | 77.5295 | 79.3294 | 181.8589 | 130.3412 | 125.1293 | 135.0118 | 160.706 | 138.3412 | 132.3412 | 112.353 | 138.6824 | 71.7881 | 88.3648 | 78.4941 | 108.1293 | 102.647 | 107.8235 | 109.3294 | 79.1883 | 268.6588 | 269.4705 | 248.0118 | 323.0827 | 215.8589 | 255.2592 | 232.6824 | 285.9882 | 321.3766 | 223.6824 | 212.353 | 349.1529
1209 | 15.2934 | 148.0481 | 207.2861 | 240.9519 | 208.5241 | 172.3422 | 183.7139 | 162.115 | 181.647 | 216.9332 | 144.0856 | 126.9144 | 242.5321 | 246.3904 | 162.3235 | 158.0668 | 249.1711 | 204.2193 | 187.238 | 184.8476 | 229.3048 | 190.2085 | 351.0267 | 287.0294 | 188.1711 | 266.9332 | 287.3235 | 246.4759 | 319.2193 | 138.0668 | 125.9332 | 119.3048 | 165.3422 | 156.115 | 167.7995 | 144.3904 | 86.8102 | 289.0748 | 271.7807 | 222.7246 | 279.9626 | 179.9037 | 144.353 | 146.6765 | 164.6658 | 74.6578 | 69.238 | 79.6952 | 143.3422 | 252.9224 | 241.3048 | 241.8102 | 178 | 164.4091 | 159.2567 | 153.5241 | 130.2193 | 219.9519 | 200.6952 | 233.8289 | 179.123 | 181.2487 | 170.8476 | 205.238 | 192.2754 | 166.8289 | 257.2085 | 265.5802 | 296.0856 | 177.8289 | 178.9706 | 250.5909 | 235.4866 | 242.2567 | 230.7807 | 211.4946 | 339.2487 | 213.4278 | 208.8182 | 175.4572 | 263.8102 | 197.0187 | 179.8663 | 153.1524 | 223.7139 | 224.762 | 247.5428 | 213.353 | 194.5989 | 118.3235 | 106.762 | 109.0481 | 146.1711 | 303.7433 | 282.9037 | 267.4465 | 405.3609 | 247.9813 | 239.115 | 232.4198 | 366.3235 | 139.9252 | 96.9519 | 87.4759 | 139.9252 | 139.4759 | 126.238 | 134.3048 | 158.0668 | 128.8289 | 112.6578 | 120.7807 | 128.1337 | 83.2567 | 68.115 | 74.6096 | 112.6283 | 101.8476 | 114.6283 | 90.353 | 80.1524 | 313.77 | 263.6096 | 260.0187 | 307.115 | 196.8102 | 250.2861 | 217.8957 | 300.2567 | 334.3235 | 225 | 221.6283 | 358.1711
1210 | 15.3044 | 157.0384 | 215.3186 | 232.8215 | 221.618 | 161.382 | 175.6814 | 177.7773 | 174.6814 | 225.739 | 155.3186 | 139.7581 | 213.8849 | 277.0959 | 153.6814 | 143.2419 | 221.7639 | 196.9808 | 179.4012 | 148.8849 | 239.3186 | 177.9424 | 345.6622 | 332.7332 | 175.382 | 259.4012 | 282.1209 | 261.7581 | 330.0384 | 145.5988 | 121.8407 | 124.1593 | 144.9232 | 171.7773 | 146.0634 | 146.7198 | 102.7773 | 301.2994 | 276.4395 | 254.1151 | 283.739 | 178.261 | 120.3628 | 138.9808 | 172.5797 | 76.1593 | 75.1593 | 63.6622 | 136.6814 | 237.7831 | 247.0192 | 233.7006 | 190.8983 | 171.2994 | 161.1401 | 162.3186 | 140.1785 | 208.382 | 220.7773 | 261.2361 | 206.357 | 191.3378 | 166.7006 | 237.8157 | 175.5029 | 179.618 | 227.7447 | 231.3052 | 313.3378 | 184.5988 | 163.3628 | 246.2802 | 236.8791 | 244.1401 | 205.3436 | 209.8407 | 345.0384 | 195.643 | 164.9866 | 181.0192 | 276.3378 | 198.1401 | 195.7581 | 170.1977 | 207.0825 | 217.1209 | 248.8599 | 217.739 | 170.9041 | 133.7581 | 148.1343 | 124.0576 | 147.1401 | 314.7581 | 305.3378 | 261.1017 | 418.739 | 239.9808 | 260.7965 | 240.8983 | 378.3186 | 144.0384 | 73.3436 | 63.2035 | 144.0384 | 120.3628 | 128.7198 | 126.261 | 154.4203 | 133.0192 | 115.8791 | 119.4203 | 135.4587 | 81.7006 | 76.8983 | 78.2994 | 95.9424 | 118.3378 | 111.7006 | 118.8157 | 86.8791 | 303.0825 | 236.3436 | 249.9616 | 317.618 | 212.7773 | 231.6622 | 227.0384 | 278.0634 | 307.6238 | 233.5988 | 207.5221 | 368.5988
1211 | 15.3154 | 158.1278 | 211.7669 | 226.6391 | 199.5791 | 174.6992 | 182.7218 | 181.8722 | 197.4209 | 209.5565 | 174.4435 | 116.7069 | 226.3157 | 266.3008 | 147.6391 | 154.9548 | 260.6088 | 187.2782 | 192.827 | 183.1201 | 219.1956 | 190.6992 | 348.2331 | 329.6617 | 184.3383 | 247.5339 | 267.9174 | 277.9548 | 304.0904 | 147.8722 | 123.6165 | 107.5565 | 154.9548 | 182.8496 | 148.9774 | 149.6165 | 83.3234 | 281.94 | 276.1278 | 238.1956 | 302.4435 | 170.0226 | 151.887 | 150.2105 | 149.4513 | 91.827 | 80.3609 | 71.4661 | 126.2782 | 286.3305 | 249.7444 | 234.7444 | 176.4287 | 168.5113 | 175.827 | 152.6617 | 151.5939 | 204.2556 | 188.2408 | 231.9852 | 201.2782 | 203.5939 | 172.1052 | 229.9174 | 173.7444 | 182 | 221.1278 | 223.2556 | 298.5565 | 185.1278 | 156.5113 | 232.0452 | 242.3609 | 226.5565 | 219.3157 | 229.06 | 350.4887 | 186.7669 | 157.8722 | 170.6617 | 263.3008 | 210.2105 | 184.9174 | 179.9774 | 197.0226 | 255.2479 | 278.654 | 213.7669 | 180.827 | 143.8496 | 128.8347 | 140.9548 | 153.1052 | 307.4061 | 285.3234 | 248.5339 | 406.9174 | 284.3531 | 244.0678 | 276.1427 | 368.6617 | 152.9774 | 94.2931 | 93.887 | 152.9774 | 121.3609 | 122.8948 | 117.1504 | 160.1052 | 125.2782 | 130.9548 | 121.6165 | 133.5113 | 93.2105 | 90.3383 | 109.5261 | 107.827 | 104.4287 | 113.6165 | 103.94 | 88 | 286.0452 | 266.887 | 228.8121 | 304.3008 | 215.1278 | 229.8722 | 226.3835 | 314.1201 | 322.1879 | 222.7895 | 220.6992 | 376.9774
1212 | 15.3264 | 158.8845 | 218.0757 | 274.6456 | 229.6097 | 189.3825 | 176.0398 | 192.6136 | 178.8883 | 259.1835 | 147.8126 | 111.2311 | 229.8845 | 246.3107 | 162.0359 | 167.6136 | 286.4583 | 210.7651 | 196.7689 | 177.5019 | 242.534 | 177.9641 | 354.3068 | 340.3825 | 208.9961 | 272.534 | 286.3427 | 250.8126 | 324.7651 | 146.2311 | 126.6534 | 116.4981 | 160.5379 | 181.3466 | 135.8485 | 146.4621 | 109.1874 | 277.2311 | 296.3427 | 269.4942 | 273.1592 | 161.0398 | 146.3864 | 171.4583 | 165.4583 | 77.1951 | 73.0398 | 71.2311 | 134.7291 | 285.1553 | 256.1913 | 285.4146 | 182.8447 | 158.2709 | 148.8126 | 158.0757 | 112.3146 | 208.4223 | 210.4185 | 227.8845 | 213.267 | 205 | 196.8806 | 237.7291 | 185.4981 | 186.4223 | 209.5019 | 258.3787 | 311.9204 | 200.0359 | 182.534 | 273.3388 | 267.7651 | 224 | 230.8447 | 201.0437 | 381.9563 | 182.4621 | 182.7651 | 164.5777 | 240.6573 | 208.4621 | 183 | 176.5777 | 195.1155 | 297.2631 | 284.7689 | 256.3388 | 188.3068 | 132.6175 | 125.8845 | 117.3505 | 155.7689 | 302.4621 | 268.733 | 244.3466 | 430.6495 | 296.3068 | 275.4942 | 243.8524 | 374.9602 | 147.8087 | 88.2709 | 86.6175 | 147.8087 | 136.1515 | 141.4583 | 116 | 147.733 | 120.4621 | 108.2349 | 110.5019 | 154.2272 | 95 | 84.9243 | 75.0835 | 111.7689 | 113.4981 | 130.8049 | 99.2311 | 96.8447 | 298.1515 | 268.4621 | 221.5777 | 308.1913 | 199.9641 | 283.9524 | 224.2311 | 313.8087 | 311.733 | 213.9243 | 203.5417 | 399.2272
1213 | 15.3374 | 152.7227 | 218.1032 | 237.9556 | 210.6843 | 194.587 | 174.1032 | 186.8259 | 202.006 | 248.0649 | 172.6962 | 139.6962 | 205.7876 | 240.413 | 172.0708 | 160.9292 | 262.0973 | 227.4514 | 190.7227 | 187.6578 | 269.3157 | 177.7935 | 386.3865 | 320.4778 | 186.8908 | 285.8643 | 286.3097 | 265.8319 | 340.5546 | 121.7876 | 127 | 118.8968 | 153.8259 | 158.5811 | 155.5222 | 137.0324 | 106.7227 | 285.9676 | 285.5486 | 247.994 | 316.5281 | 151.0324 | 138.7227 | 149.7876 | 135.7167 | 71.413 | 74.6903 | 59.3422 | 136.8968 | 249.0265 | 248.0324 | 263.3038 | 190.2773 | 157.8968 | 138.7227 | 134.7876 | 133.9027 | 218.8643 | 195.1681 | 260.2833 | 198.8584 | 206.7935 | 196.413 | 266.7995 | 204.9351 | 177.1357 | 234.006 | 245.0649 | 323.8643 | 221.7287 | 178.8259 | 269.1357 | 251.2713 | 236.5546 | 247.2449 | 210.4514 | 344.7492 | 169.4454 | 197.6578 | 179.2449 | 268.4898 | 198.1357 | 212.593 | 187.6578 | 206.6578 | 267.9232 | 274.2389 | 240.4778 | 172.8584 | 117.5486 | 117.0324 | 118.4838 | 158.6903 | 318.1416 | 316.3217 | 246.6903 | 431.3097 | 278.1681 | 234.2654 | 254.2449 | 396.6254 | 181.9735 | 78.9292 | 85 | 181.9735 | 148.7611 | 135.0324 | 134.8319 | 171.1092 | 122.6903 | 127.4189 | 123.3481 | 127.407 | 89.6195 | 74.1357 | 75.3805 | 124.5546 | 110.5162 | 124.9292 | 112.4514 | 111.4514 | 323.3157 | 282.3481 | 268.5281 | 337.6962 | 204.2773 | 266.7876 | 222.2065 | 318.3805 | 329.7287 | 229.1416 | 197.413 | 384.0649
1214 | 15.3484 | 158.3633 | 209.8186 | 208.4557 | 219.8176 | 153.1838 | 192.181 | 188.7271 | 210.4543 | 228.7281 | 168.7276 | 135.7276 | 244.8162 | 272.7257 | 149.3647 | 163.6362 | 301.7253 | 210.819 | 203.6357 | 178.0914 | 264.7276 | 181.6362 | 390 | 320.7271 | 227.6343 | 251.5471 | 301.4538 | 269.8181 | 360.181 | 146.2715 | 132.4543 | 111.7276 | 149.3638 | 180.5443 | 154.3638 | 141.4543 | 96.9095 | 312.4534 | 285.8181 | 246.8181 | 302.91 | 171.8172 | 157.09 | 144.2729 | 133.8181 | 80.0905 | 84.0905 | 69.8176 | 119.7281 | 270.4534 | 257.9086 | 234.5466 | 194.6362 | 155.2729 | 156.181 | 154.7262 | 138.8181 | 239.9991 | 209.3629 | 231.2743 | 227.9076 | 204.2729 | 222.3624 | 271.8181 | 206.091 | 184.1814 | 211.5466 | 260.2719 | 315.0005 | 188.5471 | 155.2738 | 267.091 | 270.8172 | 245.2724 | 230.819 | 212.909 | 340.909 | 195.2715 | 188.0914 | 175.5457 | 295.6353 | 211.5448 | 186.9105 | 172.6371 | 198.0005 | 264.909 | 278.4543 | 265.2715 | 190.9991 | 143.2715 | 124.1814 | 116.2729 | 146.2733 | 326.3633 | 309.2733 | 256.0905 | 433.7271 | 308.7257 | 246.2719 | 234.1828 | 352.6385 | 173.2733 | 89.8176 | 95.9086 | 173.2733 | 135.4552 | 147.6357 | 127.9095 | 153.0919 | 132.9995 | 126.3638 | 99.5466 | 143.9991 | 82.6367 | 80.2724 | 100.5443 | 124.1819 | 113.6362 | 133.0905 | 112.1819 | 91.1828 | 297.8195 | 294.9086 | 280.3633 | 335.5457 | 216.8176 | 296.7257 | 231.0905 | 321.7271 | 330.1819 | 236.4543 | 196.091 | 413.8167
1215 | 15.3594 | 171.8988 | 190.5732 | 257.595 | 208.1012 | 183.0896 | 213.3481 | 165.9665 | 181.5171 | 262.9322 | 168 | 127.6293 | 264.6628 | 269.5506 | 175.5615 | 158.472 | 277.4385 | 253.2243 | 199.472 | 181.6067 | 289.7975 | 167.2586 | 376.1799 | 340.3481 | 177.641 | 259.9774 | 269.8318 | 263.5506 | 357.3933 | 115.8318 | 112.7305 | 128.5054 | 159.1347 | 171.0226 | 165.9774 | 157.6628 | 93.236 | 294.7305 | 267.5732 | 286.6176 | 306.528 | 149.1238 | 163.6067 | 183.6176 | 157.9548 | 88.3707 | 62.8879 | 76.528 | 110.6293 | 261.0226 | 257.1573 | 262.4042 | 196.8427 | 171.5841 | 155.236 | 123.8318 | 132.5506 | 242 | 178.7531 | 243.6628 | 231 | 233.4829 | 202.8879 | 248.9665 | 187.5732 | 193.2921 | 263.359 | 245.4159 | 319.528 | 189.6067 | 197.2243 | 267 | 277.6067 | 241.3933 | 212.4159 | 213 | 372.3255 | 239.4603 | 187 | 199.8762 | 295.236 | 224.9774 | 203.3481 | 159.0226 | 193.3146 | 313.831 | 298.3481 | 224.697 | 194.8427 | 121.1238 | 129.6067 | 111.3933 | 155.1347 | 317.7866 | 328.2695 | 275.4268 | 450.5841 | 266.8544 | 232.3372 | 245.8201 | 337.8653 | 165.5506 | 106.6628 | 85.9439 | 165.5506 | 156.1121 | 129.6519 | 131.6067 | 142.7079 | 151.5054 | 149.0335 | 132.9322 | 111.9104 | 96.7414 | 75.472 | 79.0452 | 127.6854 | 139.7975 | 122.9439 | 119.3707 | 99.1347 | 314.3481 | 281.2586 | 265.3372 | 327.6293 | 206.0226 | 250.2477 | 223.7079 | 321.0787 | 328.1573 | 231.472 | 207.9774 | 381.9891
1216 | 15.3704 | 159.9292 | 200.2036 | 241.4601 | 208.8673 | 188.8009 | 166.4512 | 158.3982 | 203.2744 | 243.5929 | 162.3982 | 123.2655 | 289.3408 | 269.9336 | 158.3938 | 155.1991 | 267.531 | 258.8673 | 209.2699 | 182.9336 | 293.8673 | 218.2833 | 379.6682 | 315.8583 | 223.416 | 301.1461 | 283.8053 | 277.0045 | 330.8583 | 127.9381 | 146.478 | 131.8673 | 165.6018 | 159.7301 | 148.3274 | 147.7964 | 110.739 | 300.469 | 296.8098 | 275.9955 | 282.7256 | 179.6771 | 169.6018 | 146.8539 | 157.1991 | 84.3318 | 85.2744 | 87.2699 | 122.1372 | 264.6682 | 273.8053 | 293.0089 | 211.9381 | 176.7345 | 169.9381 | 148.0753 | 143.2036 | 243.8673 | 210.5443 | 286.0797 | 220.7301 | 212.6592 | 170.1902 | 245.1327 | 230.8143 | 181.8628 | 248.3938 | 267.3408 | 306.9292 | 200.2699 | 183.261 | 280.0708 | 278.9336 | 226.9955 | 212.8673 | 198.0619 | 368.4646 | 187.9158 | 217.8098 | 179.5929 | 302.469 | 255.8762 | 216.2036 | 188.8098 | 242.4824 | 273.1857 | 356.9515 | 231.2699 | 195 | 107.7964 | 136.5354 | 108.1991 | 150.3982 | 317 | 327.1991 | 261.1283 | 453.8673 | 278.8717 | 277.6815 | 280.6107 | 352.8717 | 155.6637 | 82.792 | 71.9292 | 155.6637 | 186.9425 | 168.1461 | 142.2699 | 158.8053 | 143.6637 | 133.261 | 147.2036 | 154.7479 | 92.3982 | 93.6726 | 89.1372 | 142.9381 | 134.531 | 117.3318 | 117.1991 | 101.8673 | 333.739 | 288.4027 | 276.1372 | 332.6018 | 218.0708 | 276.8098 | 236.0708 | 306.9955 | 303.7256 | 239.4027 | 231.4071 | 389.2699
1217 | 15.3814 | 175.0807 | 235.0532 | 218.2438 | 212.7837 | 180.4867 | 182.8644 | 194.8909 | 195.5408 | 259.9725 | 166.7296 | 140.0266 | 251.2713 | 267.1622 | 180.648 | 152.1622 | 304.8369 | 214.5417 | 210 | 185.8378 | 252.3795 | 201.1897 | 400.8103 | 313.0541 | 195.7846 | 278.4601 | 299.1888 | 283.6755 | 398.0523 | 164.945 | 121.5683 | 133.8918 | 158.4326 | 179.8103 | 121.4601 | 130.9193 | 123.3511 | 302.8918 | 249.8121 | 267.4326 | 247.8927 | 182 | 175.6755 | 149.6755 | 147.5408 | 80.2163 | 71.8653 | 67.1897 | 134.3511 | 281.0807 | 303.3776 | 306.3511 | 198.8112 | 159.9734 | 177.6215 | 141.4867 | 135.4867 | 223.1897 | 230.0266 | 289 | 203.9193 | 217.6215 | 239.89 | 291.3502 | 225.4867 | 177.2163 | 240.3785 | 274.6755 | 299.3785 | 197.2163 | 184.8378 | 283.8378 | 283.7296 | 257.2154 | 222.4592 | 214.9725 | 384.0807 | 191.5674 | 210.5408 | 195.0266 | 299.2163 | 190.8396 | 229.297 | 207.0807 | 273.4317 | 296.4858 | 322.2172 | 268.8909 | 205.4051 | 133.4858 | 122.8112 | 100.4326 | 168.9184 | 317.9459 | 304.2979 | 258.1082 | 447.3785 | 277.1622 | 240.3254 | 231.9202 | 374.8103 | 170.1347 | 98.0266 | 83.297 | 170.1347 | 145.4876 | 138.8387 | 167.594 | 169.4592 | 155.297 | 132.9459 | 131.9193 | 157.0541 | 87.2704 | 91.2163 | 77.703 | 125.0816 | 134.9459 | 128.3511 | 128.3511 | 107.6755 | 305.6764 | 297.5133 | 295.9184 | 317.8653 | 212.3785 | 289.4051 | 261.594 | 325.8644 | 357.8094 | 233.3785 | 194.2172 | 390.9459
1218 | 15.3924 | 166.4179 | 207.2954 | 238.0807 | 243.6628 | 197.2478 | 199.3314 | 201.7911 | 203.6239 | 270.5821 | 183.2896 | 140.0418 | 259.5403 | 277.5403 | 200.206 | 189.3703 | 306.0418 | 233.0807 | 180.2954 | 195.5821 | 273.9553 | 213.415 | 408.7075 | 337.9135 | 190.1671 | 283.7075 | 246.3401 | 277.2925 | 366.5461 | 151.6686 | 139.1643 | 156.0389 | 183.8717 | 162.794 | 156.4121 | 135.7493 | 93.3372 | 303.9582 | 282.4539 | 290.9553 | 318.8242 | 177.2089 | 164.5014 | 163.415 | 158.4986 | 88.6239 | 54.7104 | 72.7075 | 141.7075 | 258.0447 | 282.9611 | 264.8386 | 201.8329 | 158.0418 | 153.0865 | 140.0418 | 145.5403 | 248.83 | 208.0029 | 264.0865 | 242.2867 | 240.9971 | 201.8386 | 262.379 | 180.9222 | 175.0836 | 269.7046 | 286.4986 | 343.0778 | 220.9553 | 187.8746 | 263.8775 | 282.0836 | 235.0447 | 269.9524 | 223.6657 | 374.4597 | 259.0749 | 246.4121 | 189.2925 | 278.8775 | 256.9495 | 219.4597 | 205.1254 | 228.0476 | 311.415 | 324.7911 | 271.9582 | 209.8329 | 153.206 | 112.4179 | 108.6239 | 156.585 | 328.5403 | 317.3732 | 238.8357 | 428.794 | 293.2896 | 295.4928 | 245.2896 | 377.9164 | 165.2507 | 79.8357 | 87.8329 | 165.2507 | 165.0389 | 155.206 | 156.5432 | 170 | 187.621 | 146.415 | 151.1225 | 169.4568 | 88.9164 | 83.3343 | 107.6628 | 122.0836 | 113.9193 | 144.3314 | 110.794 | 90.7522 | 372.9913 | 297.0418 | 275.9193 | 334.2478 | 226.3732 | 269.8775 | 263 | 324.1254 | 325.5461 | 216.7104 | 209.2478 | 369.9193
1219 | 15.4034 | 161.1475 | 207.941 | 244.823 | 232.3835 | 195.0885 | 198.059 | 197.1475 | 204.9705 | 316.6137 | 171.3835 | 146.7935 | 287.1741 | 291.5871 | 170.9144 | 179.354 | 319.5871 | 238.8525 | 198.4101 | 215.4101 | 292.4691 | 218.8525 | 356.5928 | 327.354 | 201.646 | 314.0856 | 300.2892 | 247.8849 | 380.5281 | 143.236 | 140.9705 | 115.2683 | 175.295 | 177.5281 | 159.941 | 165.1151 | 92 | 313.705 | 316.9971 | 273.5604 | 285.1209 | 173.118 | 171.764 | 183.4101 | 154.1475 | 103.5576 | 81.1741 | 57.4719 | 145.882 | 290.9676 | 281.0295 | 296.9676 | 195.2065 | 199.7317 | 147.1475 | 144.8525 | 148.9115 | 276.2036 | 230.2921 | 293.0856 | 240.118 | 272.0856 | 236.8791 | 305.6432 | 204.2331 | 186.646 | 242.8554 | 291.8525 | 345.9705 | 222 | 201.5871 | 286.2921 | 245.1209 | 229.1475 | 234.1504 | 230.7935 | 406.0266 | 205.7108 | 227.6194 | 188.0295 | 299.3511 | 233.7964 | 227.7345 | 192.3835 | 258.9971 | 332.3806 | 308.5014 | 268.118 | 202.236 | 142.354 | 137.2331 | 134.2331 | 175.4101 | 334.823 | 350.0266 | 238.9705 | 460.0266 | 287.2065 | 286.354 | 249.882 | 394.4986 | 185.3806 | 101.3216 | 66.6489 | 185.3806 | 173.764 | 149.2065 | 139.5014 | 186.4986 | 179.295 | 157.6755 | 132.5899 | 180.6755 | 103.5576 | 83.9705 | 80.8554 | 125.882 | 125.6165 | 159.5576 | 122.6165 | 101.646 | 360.4719 | 347.4662 | 279.8525 | 337.9115 | 229.9115 | 278.705 | 263.9705 | 339.5281 | 357.9676 | 216 | 223.5871 | 401.0266
1220 | 15.4144 | 176.7247 | 189.3269 | 233.2065 | 235.9312 | 182.2237 | 195.0516 | 178.3269 | 218.7591 | 323.8968 | 178.8624 | 130.2925 | 315.5183 | 328.3634 | 182.7763 | 176.0516 | 292.4817 | 245.8796 | 217.6731 | 193.3957 | 296.9312 | 203.2753 | 386.4494 | 333.8796 | 233.4494 | 292.3957 | 308.8796 | 280.415 | 328.9119 | 140.0516 | 143.9484 | 124.8107 | 182.8624 | 200.6043 | 166.8796 | 160.1032 | 116.5699 | 297.2925 | 329.7935 | 295.6043 | 303.6559 | 182.8279 | 173.9656 | 164.3441 | 168.7419 | 91.2237 | 77.086 | 75.6731 | 117.499 | 296.914 | 281.9828 | 280.3097 | 216.6215 | 188.2237 | 169.6043 | 144.0172 | 156.8624 | 247.5162 | 192.671 | 268.4473 | 241.9656 | 282.8279 | 214.4129 | 304.0516 | 268.8817 | 192.8968 | 299.0021 | 284.1376 | 340.1032 | 219.0516 | 193.1548 | 257.5162 | 300.0193 | 216.2237 | 268.3806 | 194.6366 | 359.8259 | 228.5699 | 206.3613 | 167.3613 | 332.4322 | 241.8452 | 228.9828 | 207.7247 | 266.8796 | 336.9312 | 326.6731 | 258.1721 | 241.3118 | 150.8452 | 135.0516 | 109.4473 | 178.9484 | 366.4494 | 351 | 261.6043 | 428.5678 | 287.9828 | 277.1548 | 250 | 364.5334 | 153.5678 | 85.2925 | 72.8796 | 153.5678 | 175.9656 | 150.9656 | 173.3978 | 167.3441 | 193.7419 | 168.8107 | 138.8796 | 177.0688 | 94.1721 | 98.7419 | 105.5527 | 118.1376 | 150.5699 | 147.2237 | 140.6903 | 109.8624 | 360 | 308.7054 | 311.4494 | 353.7247 | 247.6903 | 264.2581 | 248.2753 | 339.0172 | 354.086 | 216 | 203.3613 | 356.7914
1221 | 15.4254 | 175.0098 | 224.8231 | 284.7445 | 223.0639 | 181.0049 | 204.9509 | 196.9066 | 230.941 | 316.0393 | 206.8624 | 143.9312 | 288.1376 | 284.2211 | 188.9705 | 197.8919 | 312.8968 | 265.9017 | 204.0688 | 199.9656 | 327.8477 | 219.9165 | 404.9115 | 322.059 | 239.9705 | 306.9263 | 322.9312 | 306.8722 | 414.5725 | 163.8821 | 150.9656 | 135.9459 | 179.0197 | 208.9607 | 177.9459 | 162.9853 | 113.0197 | 307.9459 | 302.1376 | 270.1278 | 319.9214 | 199.9165 | 170.0197 | 200.8182 | 144.1229 | 83.0393 | 80.9803 | 76 | 160.7838 | 316.9017 | 300.9066 | 287.9607 | 209.0393 | 199.941 | 166.0197 | 158.9263 | 161.9754 | 258.941 | 220.8575 | 303.8231 | 257.9214 | 282.0049 | 196.0885 | 332.8575 | 259.0541 | 230.8133 | 268.1573 | 305.8919 | 386.769 | 220.9902 | 184.0442 | 312.7248 | 325.8771 | 269.7346 | 294.8722 | 202.9558 | 391.8378 | 235.9656 | 215.9509 | 179.9361 | 299.1671 | 241.0049 | 264.8231 | 211.9803 | 256.0541 | 355.9066 | 337.9459 | 303.7739 | 233.0442 | 153.9853 | 160.8722 | 114.9705 | 174.0246 | 389.887 | 327.1179 | 274.9361 | 449.8919 | 318.8477 | 281.9754 | 231.0934 | 371.9607 | 206.7346 | 92.9607 | 68.0246 | 206.7346 | 176 | 154.9803 | 158.0786 | 173.9656 | 162.1573 | 140.1425 | 129.0491 | 167.0491 | 92.0098 | 93.0295 | 104.0098 | 146.8575 | 174.8821 | 152.9705 | 119.1081 | 106.0197 | 338.1081 | 320.9361 | 305.0344 | 375.8919 | 249.9902 | 286.887 | 260.9361 | 352.9312 | 348.0295 | 225.9509 | 230.8624 | 398.7887
1222 | 15.4364 | 163.1033 | 214.1254 | 244.1992 | 224.1328 | 174.2287 | 207.8598 | 203.9115 | 239.7271 | 289.2803 | 187.2287 | 133.2434 | 325.9557 | 323.1402 | 176.8746 | 211.6164 | 345.0811 | 264.0738 | 236.941 | 216.8967 | 333.4352 | 208.9926 | 400.3172 | 333.8598 | 227.9484 | 312.9705 | 327 | 284.1033 | 389.941 | 156.9484 | 158.9262 | 138.8893 | 197.0148 | 194.1254 | 185.0074 | 155.9779 | 110.9779 | 335.7271 | 303.9484 | 290.2213 | 318.0074 | 189.1033 | 174.9484 | 153.118 | 170.1918 | 93.9926 | 83.9336 | 78.9926 | 128.1328 | 303.4279 | 273.1402 | 299.9041 | 228.3098 | 192.8967 | 174.1033 | 163.9336 | 159.9484 | 280.0664 | 215.0516 | 311.8967 | 233.3541 | 283.2213 | 234.2951 | 350.9852 | 263.118 | 189.0443 | 266.2508 | 290.4426 | 384.9189 | 244.1697 | 192.1844 | 315.9852 | 325.0959 | 270.9779 | 271.059 | 212.2434 | 394.3098 | 258.059 | 213.1475 | 194.8303 | 285.4205 | 263 | 238.3025 | 212.1844 | 282.0811 | 348.2287 | 388.9484 | 307.0369 | 238.7344 | 154.0074 | 137 | 127.0516 | 173.823 | 371.118 | 328.0885 | 265.0738 | 503.9115 | 311.059 | 321.7492 | 219.1402 | 345.1844 | 181.9631 | 82.0074 | 92.9336 | 181.9631 | 183.882 | 172.059 | 170.0295 | 188.0664 | 137.4057 | 168.9262 | 161.8008 | 187.0443 | 89.1107 | 106.1033 | 106.0516 | 135.391 | 156.0443 | 133.1254 | 119.9852 | 110.9705 | 367.9262 | 351.0885 | 315.118 | 352.9041 | 264.0516 | 312.177 | 293.7713 | 392.7492 | 334.3098 | 265.9926 | 199.2213 | 445.0885
1223 | 15.4474 | 177 | 230.5673 | 269.9969 | 242.177 | 204.8033 | 189.0787 | 192.236 | 203.2557 | 326.7443 | 217.823 | 166.1573 | 320.6294 | 341.646 | 160.3737 | 160.3737 | 355.3116 | 274.177 | 229.2163 | 203.2557 | 390.9773 | 208.4524 | 443.118 | 314.9803 | 221.6884 | 308.8427 | 327.6687 | 298.5507 | 382.6097 | 150.0983 | 149.118 | 124.236 | 198.7443 | 210.8033 | 185.2526 | 152.4296 | 107.9213 | 300.1014 | 297.5114 | 319.941 | 318.882 | 202.8033 | 168.1573 | 169.236 | 195.2723 | 92.8623 | 75.059 | 77.8427 | 145.7837 | 360.5673 | 292.649 | 286.8623 | 269.469 | 179.0787 | 187.469 | 155.3934 | 152.9803 | 288.9607 | 223.0817 | 298.4327 | 280.8033 | 312.5673 | 273.5476 | 348.8427 | 278.0363 | 195.1377 | 300.0393 | 349.2723 | 374.9047 | 266.646 | 217.0393 | 314.59 | 338.4917 | 268.0197 | 278.587 | 244.882 | 435.9213 | 265.351 | 233.708 | 172.531 | 342.1967 | 263.295 | 278.764 | 236.705 | 293.118 | 379.649 | 382.6097 | 312.295 | 203.8457 | 155.6294 | 137.5704 | 134.1967 | 150.649 | 385.8593 | 340.5507 | 275.0787 | 493.2784 | 318.587 | 288.6294 | 238.177 | 369.587 | 177.6687 | 82.9607 | 84.1967 | 177.6687 | 168.9637 | 179.8623 | 173.7837 | 196.9017 | 191.646 | 158.9803 | 135.3344 | 193.6687 | 103.8623 | 119.8427 | 112.8427 | 187.1149 | 161.5083 | 150.118 | 118.1573 | 106.9213 | 358.1377 | 362.6853 | 330.6263 | 341.1604 | 271.7474 | 336.2163 | 263.3147 | 359.3934 | 376.649 | 264.882 | 229.354 | 456.528
1224 | 15.4584 | 177.5753 | 210.7258 | 221.7897 | 251.4474 | 194.9041 | 193.2876 | 203.8722 | 216.2876 | 315.1505 | 209.2557 | 173.6165 | 351.5206 | 323.8082 | 179.3196 | 178.7763 | 324.0041 | 283.1278 | 240.1278 | 216.5753 | 340.7351 | 231.3515 | 447.9134 | 313.8402 | 255.5206 | 302.3423 | 360.1371 | 324.8175 | 412.3928 | 155.5753 | 154.6485 | 135.7124 | 186.1918 | 201.8309 | 149.0227 | 124.7031 | 104.032 | 353.6897 | 323.4794 | 317.2876 | 312.7763 | 194.0546 | 176.4794 | 180.7124 | 159.2557 | 86.1598 | 78.4155 | 69.7763 | 134.7443 | 340.2464 | 324.9361 | 281.3103 | 242.8082 | 184.0866 | 161.6072 | 174.8402 | 153.0866 | 287.7031 | 275.4979 | 320.767 | 270.6804 | 292.4062 | 251.5753 | 341.6072 | 231.9495 | 202.8629 | 302.2237 | 313.4474 | 419.8722 | 249.5433 | 219.4474 | 343.2969 | 363.0959 | 269.9907 | 259.4062 | 239.3196 | 431.7763 | 234.0866 | 267.2423 | 199.1598 | 351.6804 | 277.7443 | 267.7351 | 222.3196 | 298.1691 | 411.6165 | 414.1186 | 326.1371 | 246.8309 | 185.5938 | 165.4887 | 143.4247 | 183.2876 | 330.0866 | 367.5526 | 279.4794 | 555.3701 | 299.566 | 318.2423 | 246.5526 | 349.4474 | 210.201 | 81.4794 | 94.2237 | 210.201 | 215.6897 | 172.8722 | 163.8309 | 192.2557 | 174.9588 | 157.8082 | 151.4887 | 225.7856 | 96.9041 | 112.2237 | 104.8402 | 143.3835 | 137.0959 | 156.2237 | 125.7443 | 103.4794 | 366.0227 | 348.3103 | 310.6897 | 397.8495 | 306.8268 | 346.9361 | 279.3835 | 378.5845 | 408.3928 | 258.6165 | 246.1371 | 433.1278
1225 | 15.4694 | 194.3805 | 261.6282 | 275.8672 | 265.5753 | 193.3275 | 203.6373 | 200.0885 | 226.1063 | 350.6638 | 217.885 | 162.0443 | 339.124 | 319.5488 | 189.5753 | 173.6373 | 412.6547 | 288.6373 | 243.2477 | 234.2655 | 363.885 | 242.9735 | 418.009 | 309.354 | 240.2477 | 344.062 | 336.8763 | 291.655 | 395.416 | 172.7787 | 146.2125 | 127.2213 | 192.3098 | 226.646 | 179.646 | 146.7523 | 105.6195 | 314.354 | 336.938 | 325.1592 | 305.1592 | 225.0265 | 189.8495 | 172.1328 | 167.708 | 91.177 | 90.5132 | 63.4868 | 128.7258 | 378.7523 | 323.177 | 320.3362 | 237.8408 | 216.6902 | 180.3983 | 170.2655 | 184.8495 | 309.6195 | 231.7258 | 342.938 | 262.6815 | 333.8495 | 269.1328 | 360.708 | 291.3098 | 228.646 | 308.6902 | 326.4247 | 416.1328 | 266.7523 | 233.1328 | 322.6195 | 366.2213 | 299.6017 | 302.6195 | 247.5397 | 426.5488 | 267.3098 | 216.0355 | 204.354 | 341.469 | 271.8585 | 289.6902 | 232.6638 | 274.1948 | 400.9735 | 447.8672 | 303.4073 | 273.416 | 143.531 | 149.6017 | 125.3805 | 191.0265 | 363.7523 | 355.2833 | 294.2213 | 507.1948 | 344.9645 | 266.1505 | 234.3275 | 365.124 | 186.5753 | 95.1592 | 100.5132 | 186.5753 | 175.823 | 170.3275 | 188.3805 | 201.1505 | 204.0443 | 153.062 | 137.1505 | 189.354 | 94.531 | 119.531 | 100.0885 | 155.531 | 141.947 | 163.6638 | 119.062 | 118.3983 | 398.0178 | 389.4603 | 273.4338 | 363.5753 | 241.7523 | 346.593 | 291.5753 | 369.717 | 390.3098 | 246.5575 | 220.7523 | 437.2655
1226 | 15.4804 | 179.5299 | 232.622 | 271.5865 | 277.7173 | 220.6996 | 238.1519 | 202.622 | 249.4911 | 366.3004 | 236.8304 | 161.8692 | 386.1308 | 354.5266 | 200.4734 | 207.9257 | 362.622 | 324.5089 | 230.449 | 239.7738 | 386.7317 | 263.6042 | 481.9823 | 318.8093 | 222.5865 | 365.6996 | 396.6253 | 345.4945 | 426.1696 | 167.6608 | 193.2084 | 131.7738 | 200.0743 | 217.4734 | 174.6009 | 162.5477 | 118.5477 | 323.4135 | 315.3004 | 306.4911 | 288.5831 | 204.5089 | 167.7705 | 176.0177 | 182.8304 | 95.0565 | 80.0565 | 74.622 | 164.2472 | 393.3603 | 325.9823 | 309.449 | 257.2439 | 210.622 | 189 | 175.265 | 160.3958 | 323.735 | 268.9435 | 321.9224 | 298.8869 | 310.3004 | 273.0177 | 377.6962 | 299.5266 | 222.2439 | 302.3958 | 315.8659 | 419 | 283.9047 | 237.1308 | 335.5477 | 370.4346 | 293.3182 | 314.9823 | 218.4313 | 460.622 | 275.5266 | 260.6785 | 210.3038 | 331.7528 | 313.0177 | 282.265 | 246.7738 | 301.0743 | 425.2794 | 446.6397 | 376.6608 | 302.9823 | 156.357 | 142.1874 | 113.1874 | 171.4135 | 382.2051 | 385.0355 | 298.4346 | 535.8836 | 302.4135 | 291.3958 | 261.1907 | 410.5477 | 199.9612 | 78.7528 | 90.5089 | 199.9612 | 175.5621 | 199.0565 | 175.9047 | 225.5477 | 204.378 | 175.1519 | 162.3392 | 197.1131 | 106.1131 | 130.378 | 102.4523 | 167.3392 | 184.1131 | 178.2827 | 141.6042 | 127.735 | 421.6397 | 421.5089 | 326.3392 | 376.9612 | 256.4734 | 379.9823 | 305.0177 | 447.6253 | 396.8869 | 238.8093 | 235.9257 | 444.6397
1227 | 15.4914 | 157.6845 | 244.9961 | 248 | 274.1701 | 200.239 | 224.6883 | 214.4455 | 243.478 | 390.2026 | 232.761 | 144.7896 | 407.8221 | 380.413 | 176.9961 | 191.0325 | 371.6234 | 334.5831 | 286.0727 | 235.5182 | 445.8337 | 256.174 | 465.6883 | 350.9961 | 200.9961 | 346.4091 | 360.0286 | 289.6156 | 432.0974 | 162.0688 | 180.413 | 128.3442 | 216.5545 | 195.1663 | 214.4207 | 156.3766 | 111.6883 | 345.1415 | 337.5507 | 299.0325 | 314.6558 | 215.2026 | 213.0364 | 193.826 | 181.9273 | 94.761 | 80.2428 | 84.0364 | 135.5507 | 366.4818 | 311.3403 | 363.1454 | 278.0688 | 221.5507 | 188.3117 | 162.3805 | 166.174 | 335.6558 | 268.2065 | 353.6558 | 297.2753 | 335.9923 | 291.1013 | 408.1338 | 326.239 | 243.0688 | 311.6845 | 348.3078 | 423.8182 | 300.0325 | 252.7649 | 329.239 | 361.6883 | 331.2428 | 298 | 257.0727 | 471.3442 | 299.6234 | 271.8623 | 183.1013 | 362.5143 | 331.3078 | 276.7455 | 245.1338 | 320.9961 | 482.7533 | 476.065 | 371.8948 | 287.3078 | 178.4493 | 163.3078 | 131.8987 | 194.0364 | 419.413 | 417.6922 | 288.0364 | 582.6922 | 325.5182 | 296.6922 | 233.6156 | 401.8299 | 214.8987 | 107.9312 | 97.8987 | 214.8987 | 233.761 | 199.8623 | 190.3805 | 219.0325 | 194.2753 | 163.5468 | 167.587 | 199.8948 | 108.2065 | 120.6195 | 110.1377 | 173.9636 | 186.0688 | 183.4818 | 136.3078 | 140.5507 | 446.522 | 428.6922 | 332.1377 | 393.7572 | 237.6079 | 363.4818 | 320.348 | 406.6234 | 398.7858 | 267.5545 | 218.5507 | 474.065
1228 | 15.5024 | 193.9189 | 272.3245 | 250.0281 | 289.0265 | 218.7847 | 235.8658 | 234.4056 | 275.5398 | 420.649 | 217.7847 | 168.9189 | 446.8112 | 389.6505 | 207.1638 | 202.6474 | 359.0031 | 355.351 | 262.9204 | 230.2168 | 405.8393 | 247.1903 | 474.6755 | 375.4852 | 231.975 | 379.7301 | 402.947 | 325.8112 | 471.7286 | 164.6225 | 184.1342 | 134.1357 | 196.2699 | 237.4321 | 169.2449 | 175.4867 | 121.8112 | 325.5975 | 343.4587 | 312.6755 | 309.7566 | 244.8378 | 197.9719 | 202.4041 | 210.0281 | 78.7301 | 70.8112 | 71.4867 | 143.7301 | 377.7051 | 343 | 313.0281 | 280.298 | 228.2699 | 182.6505 | 155.5413 | 155.8924 | 333.8393 | 272.7847 | 351.2714 | 303.2714 | 388.2964 | 306.8924 | 434.9719 | 345.3526 | 241.6474 | 351.8128 | 366.2434 | 487.7832 | 313.1888 | 213.2995 | 347.4602 | 371 | 323.1092 | 298.7301 | 233.5959 | 473.4852 | 282.1357 | 273.1638 | 197.2699 | 382.9189 | 349.3245 | 367.4321 | 275.6225 | 349.7036 | 518.9454 | 504.9735 | 384.649 | 307.4337 | 169.7832 | 178.6474 | 118.5413 | 182.4602 | 426.8658 | 403.6505 | 277.5959 | 563.2964 | 318.8378 | 280.7036 | 268.4321 | 388.0281 | 201.298 | 106.0265 | 82.9189 | 201.298 | 218.4602 | 197.9189 | 184.6771 | 233.5679 | 202.7863 | 200.5959 | 163.4602 | 212.0811 | 113.1903 | 129.4056 | 112 | 185.2964 | 188.8658 | 188.2964 | 153.6755 | 148 | 414.5413 | 416.8409 | 335.2168 | 406.5959 | 326.7566 | 373.6505 | 303.8409 | 391.4883 | 449.9189 | 253.0296 | 224.8643 | 500.2153
1229 | 15.5134 | 189.3568 | 273.0107 | 273.1868 | 277.2527 | 237.6922 | 255.5054 | 236.9449 | 261.3354 | 429.4012 | 239.7749 | 168.1868 | 456.8407 | 431.0934 | 241.3185 | 178.709 | 425.7856 | 349.121 | 298.3354 | 240.7305 | 436.879 | 271.2527 | 472.4946 | 347.8407 | 277.1868 | 387.1593 | 426.4946 | 337.5222 | 449.1761 | 181.6922 | 162.7198 | 143.7305 | 188.8407 | 230.8132 | 203.5712 | 180.0659 | 131 | 388.3461 | 328.709 | 310.6815 | 310.4563 | 245.6156 | 176.588 | 172.6263 | 233.2802 | 88.0276 | 80.0934 | 76.3461 | 153.7749 | 431.6539 | 344.9617 | 336 | 294.3461 | 224.2036 | 221.7581 | 171.2251 | 169.83 | 368.2419 | 292.0659 | 378.7749 | 327.879 | 375.7258 | 318.6815 | 419.2863 | 374.6156 | 213.9724 | 392.4778 | 369.7473 | 479.9785 | 304.4671 | 264.2419 | 363.3461 | 373.5222 | 347.5605 | 309.0551 | 259.7305 | 450.8851 | 294.6263 | 305.17 | 190.3078 | 383.4012 | 355.4288 | 363.8024 | 291.5712 | 371.3354 | 512.5114 | 523.5666 | 394.9893 | 332.6646 | 156.8407 | 154.1485 | 136.2802 | 200.4012 | 448.8407 | 442.7581 | 307.0934 | 554.369 | 320.5497 | 300.3737 | 263.7749 | 412.8683 | 214.412 | 94.2527 | 82.7473 | 214.412 | 238.2695 | 198.9617 | 215.0934 | 243.6432 | 253.4503 | 230.2802 | 179.5329 | 214.4946 | 139.2144 | 134.2802 | 110.5988 | 170.8575 | 208.2251 | 174.5114 | 151.2144 | 149.2144 | 438.1653 | 483 | 349.4671 | 434.9724 | 328.6707 | 416.7749 | 365.9832 | 430.8132 | 448.2527 | 308.3185 | 215.2695 | 485.2312
1230 | 15.5244 | 154.4228 | 248.7567 | 278.8054 | 271.6913 | 225.6342 | 242.9027 | 215.7315 | 285.4799 | 447.9681 | 258.4799 | 172.2198 | 470.391 | 429.7802 | 225.7399 | 204.3658 | 417.6997 | 361.1628 | 321.1057 | 212.594 | 431.2853 | 267.4883 | 491.2282 | 357.9027 | 279.6342 | 381.7483 | 439.7315 | 363.1628 | 482.4312 | 172.2769 | 161.6913 | 116.2282 | 199.3255 | 232.4883 | 183.4883 | 171.9513 | 128.9916 | 379.4631 | 355.8456 | 323.5117 | 345.0571 | 267.9345 | 192.4228 | 171.7483 | 234.0973 | 97.0487 | 80.5772 | 69.8943 | 172.9027 | 440.5369 | 366.4312 | 336.8456 | 289.5856 | 264.8943 | 201.4799 | 158.0168 | 205.5201 | 385.5604 | 282.0487 | 399.2769 | 323.8708 | 442.4799 | 335.9513 | 487.3574 | 403.594 | 207.9111 | 370.2198 | 379.4312 | 539.2685 | 311.2198 | 285.1544 | 356.3658 | 401.1712 | 354.4799 | 332.594 | 228.3171 | 509.0571 | 292.5856 | 275.188 | 205.2198 | 401.4396 | 379.4228 | 391.3171 | 270.7483 | 396.2198 | 563.7802 | 599.854 | 421.7802 | 315.4967 | 167.4312 | 172.8859 | 137.3087 | 212.203 | 459.4312 | 424.594 | 309.1628 | 649.9261 | 359.651 | 301.5688 | 277.5117 | 431.5285 | 207.7735 | 89.0084 | 90.7399 | 207.7735 | 269.5772 | 219.1628 | 216.8456 | 277.3087 | 221.2198 | 232.6829 | 174.8943 | 230.5856 | 149.203 | 135.8372 | 97.8456 | 215.2517 | 193.854 | 227.3658 | 164.2685 | 159.203 | 500.1141 | 482.0487 | 355.1628 | 428.1712 | 374.4228 | 439.2853 | 328.0571 | 429.5285 | 443.854 | 286.8204 | 246.7886 | 533.2517
1231 | 15.5354 | 160.9499 | 287.702 | 309.938 | 288.0059 | 237.6078 | 260.534 | 209.3659 | 298 | 501.0917 | 275.4839 | 191.236 | 516.944 | 416.5458 | 230.2301 | 200.534 | 447.82 | 373.8319 | 322.354 | 242.646 | 483.1978 | 295.584 | 529.5959 | 371.7581 | 285.118 | 412.6698 | 429.708 | 372.41 | 503.944 | 198.466 | 177.416 | 152.472 | 220.7699 | 262.354 | 216.3039 | 180.708 | 111.174 | 364.9618 | 362.646 | 294.354 | 358.4839 | 248.3718 | 195.646 | 195 | 218.708 | 91.4779 | 77.354 | 69 | 187.2301 | 465.4219 | 383.5781 | 345.7699 | 304.6519 | 264.118 | 212.584 | 189.1681 | 193 | 354.8438 | 276.5959 | 429.2419 | 381.174 | 455.944 | 348.2479 | 474.3217 | 434.118 | 240.528 | 389.944 | 398.2301 | 549.882 | 334.8378 | 278.354 | 355.3659 | 430.7139 | 369.3599 | 357.1002 | 231.236 | 517.41 | 308.3599 | 328.5662 | 223.174 | 444.7818 | 384.8879 | 395.6519 | 296.9499 | 415.59 | 557.0356 | 617.3922 | 411.4957 | 363.298 | 183.9881 | 156.8879 | 124.18 | 199.8319 | 485.0738 | 455.118 | 319.236 | 618.3276 | 397.584 | 283.9499 | 250.4779 | 434.1121 | 275.7462 | 107.298 | 96.41 | 275.7462 | 265.174 | 231.4779 | 223.882 | 265.4779 | 244.0119 | 234.1299 | 178.8378 | 244.236 | 135.416 | 127.8879 | 106.8879 | 192.118 | 220.3599 | 221.174 | 176.062 | 149.6638 | 524.3718 | 478.6019 | 371.6078 | 456.534 | 378.236 | 483.41 | 342.3098 | 440.9618 | 468.826 | 249.1358 | 247.2419 | 515.8998
1232 | 15.5464 | 183.7817 | 273.9611 | 301.7394 | 305.3452 | 284.6548 | 272 | 238.8731 | 300.2149 | 580.0914 | 309.912 | 192.088 | 534.9441 | 457.088 | 216.4789 | 213.824 | 444.8629 | 396.824 | 332.9475 | 242.4755 | 535.9577 | 289.1692 | 551.0846 | 358.1269 | 289.3875 | 467.7851 | 434.088 | 367.7394 | 508.3943 | 192.7326 | 187.6091 | 154.0457 | 238.2995 | 271.2538 | 235.2674 | 182.8731 | 109.9509 | 414.7851 | 368.2081 | 302.2115 | 389.1337 | 294.3063 | 194.5634 | 197.736 | 222.4366 | 107.5245 | 81.824 | 70.824 | 175.4332 | 492.9543 | 367.9018 | 362.3875 | 320.3875 | 266.0423 | 205.5178 | 169.9543 | 197.5601 | 401.8308 | 296.1303 | 446 | 379.6903 | 460.6548 | 380.7817 | 547.1269 | 433.9577 | 237.3909 | 402.0812 | 389.4298 | 556.1658 | 370.4789 | 286.0812 | 388.5211 | 455.2995 | 387.1303 | 318.6024 | 232.2183 | 525.5076 | 332.1235 | 277.4298 | 216.4789 | 485.5245 | 403.0389 | 411.648 | 325.1235 | 423.7783 | 653.6447 | 580.8985 | 471.0914 | 380.0355 | 160.9052 | 174.2572 | 130.1337 | 221.912 | 526.2572 | 457.0423 | 316.9611 | 694.9222 | 384.5245 | 307.9543 | 266.1337 | 422.8663 | 232.648 | 117.7817 | 92.2606 | 232.648 | 269.5533 | 253.3875 | 222.3486 | 286.6057 | 275.6937 | 262.7851 | 212.6548 | 248.3452 | 145.2183 | 141.8697 | 121.6514 | 195.8663 | 238.912 | 222.687 | 183.3486 | 186.9188 | 576.1692 | 523.2927 | 415.7885 | 472.0389 | 385.7326 | 480.824 | 378.0914 | 494.5634 | 483.7293 | 290.4399 | 261.9154 | 560.6514
1233 | 15.5574 | 188.1445 | 251.8477 | 306.1309 | 315.8672 | 273.8477 | 272.5703 | 215.1484 | 317.1387 | 557.4219 | 314.1465 | 186.5703 | 610.5664 | 454.8496 | 215.2813 | 223.4336 | 491.9844 | 415.8438 | 389.9922 | 263.5625 | 537.127 | 316.5723 | 572.4141 | 380.5684 | 309.291 | 456.9922 | 437.9805 | 366.5703 | 504.4024 | 230.1445 | 189.9902 | 143.4238 | 262.5801 | 310.4336 | 206.4102 | 164.8516 | 141.5742 | 408.8399 | 417.8672 | 335.7168 | 367.1484 | 289.416 | 213.2695 | 213.7187 | 216.4199 | 89.4316 | 91.5762 | 80.4336 | 186.4258 | 506.8555 | 429.2949 | 381.8633 | 345.1387 | 273.2852 | 231.9941 | 181.2891 | 223.8652 | 383.8379 | 303.5586 | 451.1328 | 410.5723 | 460.2559 | 385.1445 | 569.2832 | 437.5508 | 237.1484 | 438.709 | 423.4199 | 610.7012 | 371.8477 | 319.002 | 398.2735 | 489.8457 | 393.5606 | 357.1406 | 235.4121 | 597.7363 | 370.4336 | 305.4316 | 211.4316 | 474.9883 | 440.1231 | 440.4199 | 362.1504 | 453.5625 | 702.9902 | 670.2793 | 449.2774 | 424.7187 | 207.7129 | 194.5742 | 115.5625 | 228.5703 | 547.1445 | 476.4043 | 292.5664 | 649.4121 | 364.7207 | 319.7168 | 247.2852 | 440.1465 | 258.1406 | 122.5723 | 93.8574 | 258.1406 | 333.0137 | 275.5723 | 219.9961 | 307.2793 | 285.7168 | 250.8516 | 194.4336 | 266.709 | 143.5645 | 141.5703 | 124.0039 | 214.8574 | 247.9941 | 273.2871 | 179.1426 | 158.7031 | 603.4297 | 593.7109 | 389.9824 | 497.8594 | 427.707 | 497.7051 | 356.5625 | 511.8438 | 537.7226 | 264.7051 | 244.7207 | 563.4316
1234 | 15.5684 | 182.3805 | 281.4646 | 335.7211 | 295.9957 | 304.5487 | 277.7036 | 207.2922 | 337.2789 | 592.8675 | 308.2656 | 189.3805 | 631.2656 | 478.4646 | 232.0974 | 216.9824 | 542.4779 | 452.9734 | 410.1283 | 291.6195 | 592.0399 | 314.1549 | 612.5088 | 387.0575 | 300.2789 | 488.5886 | 501.0974 | 374.801 | 577.876 | 219.8893 | 220.7743 | 153.9734 | 242.3496 | 296.0841 | 263.6195 | 183.8584 | 131.4513 | 464.4779 | 390.562 | 335.8983 | 360.9957 | 323.6637 | 256.3406 | 199.3805 | 245.7877 | 81.8584 | 78.5487 | 69.801 | 193.2964 | 523.6152 | 407.8143 | 370.4337 | 361.8717 | 277.7877 | 248.6062 | 169.2964 | 204.885 | 451.363 | 344.3938 | 477.8185 | 416.2081 | 560.0532 | 390.9957 | 582.4204 | 505.3363 | 229.2922 | 459.1815 | 454.4912 | 662.6152 | 403.0133 | 317.1815 | 434.1416 | 522.0575 | 426.5354 | 366.1682 | 283.0575 | 543.6285 | 359.6461 | 305.1373 | 208.1948 | 506.5088 | 509.0532 | 470.8717 | 341.1682 | 485.1815 | 737.6461 | 703.7834 | 487.0798 | 421.5311 | 211.2257 | 189.562 | 142.2257 | 233.8584 | 544.6328 | 544.4955 | 308.1549 | 713.1639 | 352.124 | 315.4071 | 253.1815 | 423.2698 | 272.7435 | 121.8584 | 91.4513 | 272.7435 | 294.8717 | 282.1373 | 238.1948 | 327 | 282.4912 | 268.4646 | 181.1682 | 291.363 | 164.6903 | 144.2257 | 114.323 | 223.6019 | 272.0399 | 271.6328 | 181.239 | 188.3539 | 609.9116 | 604.1549 | 440.9026 | 495.3363 | 448.761 | 530.3363 | 387.8717 | 555.4779 | 518.5045 | 296.2523 | 227.4779 | 562.9824
1235 | 15.5794 | 185.5193 | 290.6864 | 293.8269 | 358.1731 | 310.3254 | 291.0119 | 253.3195 | 375.8299 | 679.182 | 338.3373 | 191.5163 | 684.7398 | 487.6864 | 247.3284 | 264.8299 | 559.5104 | 436.0148 | 396.1879 | 298.1761 | 623.8299 | 320.3491 | 592.179 | 370.4985 | 342.5074 | 527.8269 | 526.5252 | 401 | 549.1968 | 199.6894 | 222.4926 | 130.997 | 279.3402 | 308.0178 | 267.5015 | 193.6657 | 126.8447 | 473.9881 | 417.8565 | 371.6449 | 415.818 | 301.5044 | 223.6775 | 201.8506 | 261.8358 | 92.6686 | 84.8269 | 94.3284 | 186.5074 | 589.3402 | 449.0119 | 413.6746 | 387.8477 | 298.5163 | 239.4956 | 163.6775 | 238.5044 | 499.0059 | 349.6598 | 464.6953 | 467.6864 | 604.3343 | 447.3225 | 625.54 | 530.1701 | 281.6716 | 479.1583 | 477.6627 | 722.9911 | 419.0148 | 340.5015 | 430.3521 | 519.0385 | 429.6835 | 384.1849 | 271.6805 | 617.3314 | 384.1642 | 358.0089 | 238.6568 | 490.0237 | 560.1879 | 505.7072 | 361.5252 | 518.0296 | 769.6864 | 779 | 546.6479 | 481.8417 | 207.3343 | 207.4956 | 135.4926 | 243.4985 | 560.5044 | 512.0089 | 312.1761 | 781.9911 | 397.3521 | 330.9911 | 273.8269 | 402.5311 | 317.0059 | 135.5104 | 85.6746 | 317.0059 | 321.6835 | 322.9733 | 269.4926 | 327 | 308.5044 | 269.8299 | 195.1731 | 334.6598 | 162.6657 | 148.1909 | 127.1672 | 277.5044 | 301.4896 | 288.6746 | 191.5104 | 167.5133 | 666.0059 | 612.0207 | 428.8477 | 517.997 | 437.1523 | 566.3699 | 405.1553 | 572.8447 | 566.3699 | 319.8506 | 238.8269 | 615.5104
1236 | 15.5904 | 210.6425 | 321.6671 | 284.4356 | 366.3082 | 302.0246 | 313.1534 | 240.2287 | 374.8452 | 718.2794 | 352.537 | 221.6918 | 799.5658 | 528.537 | 234.5137 | 260.0767 | 589.5616 | 464.0507 | 431.2822 | 309.6932 | 630.9205 | 350.7685 | 621.4342 | 373.0246 | 356.7945 | 520.7685 | 570.4356 | 407.4603 | 606.6397 | 232.6165 | 206.5904 | 128.5123 | 291.3589 | 337.2822 | 273.589 | 189.1288 | 145.1027 | 456.3329 | 451.9754 | 336.7945 | 389.4616 | 310.7945 | 245.589 | 230.2822 | 269.7685 | 99.2301 | 79.2041 | 84.3849 | 202.7685 | 612.6644 | 474.0246 | 428.2562 | 414.3329 | 325.8205 | 236.4863 | 179.1288 | 252.2808 | 504.8726 | 347.152 | 513.8973 | 500.6411 | 619.4589 | 426.3082 | 702.3315 | 553.0479 | 287.6671 | 471.1014 | 475.1274 | 721.9972 | 454.5876 | 345.6918 | 465.4096 | 586.3329 | 462.1274 | 416.7945 | 295.7178 | 614 | 388.7931 | 378.9219 | 226.8452 | 539.8698 | 602.6397 | 575.4356 | 402.2055 | 572.8452 | 818.2534 | 810.0451 | 527.3822 | 502.1014 | 217.4342 | 204.6657 | 129.6397 | 244.589 | 584.1506 | 535.4342 | 327.6411 | 768.0767 | 432.7685 | 329.1 | 265.6918 | 454.6411 | 329.1534 | 151.5644 | 103.1274 | 329.1534 | 361.126 | 289.3055 | 260.0507 | 341.1767 | 324.0753 | 269.3835 | 201.1548 | 324.974 | 159.5644 | 186.4877 | 129.6151 | 278.3603 | 290.1014 | 301.8205 | 214.9219 | 186.0521 | 680.1274 | 662.0739 | 461.2548 | 517.8452 | 426.8178 | 631.0507 | 391.9219 | 595.0507 | 625.8466 | 346.6671 | 229.6151 | 634.974
1237 | 15.6014 | 181.9853 | 309.8495 | 305.7642 | 364.7376 | 326.7257 | 329.711 | 286.4248 | 407.7376 | 774.5899 | 389.8761 | 233.8083 | 786.5018 | 567.2183 | 224.9293 | 261.3569 | 650.0679 | 491.0385 | 433.1771 | 301.5046 | 693.8349 | 374.3275 | 664.7936 | 392.3569 | 371.711 | 545.6697 | 579.3009 | 438.1505 | 647.3541 | 211.6431 | 196.3688 | 148.2449 | 302.2036 | 340.711 | 288.4661 | 176.3835 | 138.3156 | 472.1771 | 434.1771 | 348.2596 | 395.2449 | 327.4367 | 265.6431 | 231.4101 | 289.8761 | 114.3835 | 97.6312 | 79.4927 | 216.3569 | 665.6165 | 506.9706 | 431.6991 | 429.2183 | 326.9174 | 260.2596 | 175.3954 | 274.8642 | 505.7495 | 385.7404 | 519.3422 | 519.8999 | 690.5871 | 434.5165 | 744.6165 | 637.711 | 280.8349 | 502.6431 | 508.189 | 793.2862 | 502.6165 | 362.6018 | 479.0826 | 618.8587 | 487.9027 | 428.8349 | 294.5899 | 633.5578 | 436.478 | 412.711 | 253.4101 | 593.6284 | 638.944 | 590.2449 | 400.1358 | 609.9559 | 898.3541 | 970.9825 | 583.4248 | 552.6256 | 259.6431 | 228.4927 | 156.5486 | 265.9853 | 673.5312 | 587.422 | 324.2743 | 794.4752 | 447.1239 | 378.8376 | 271.4807 | 450.8908 | 339.5752 | 146.7257 | 118.3569 | 339.5752 | 416.4101 | 340.233 | 279.944 | 410.4101 | 362.3422 | 298.2183 | 193.7642 | 348.9559 | 164.5046 | 177.8761 | 141.9853 | 256.8202 | 318.5752 | 310.2036 | 240.8495 | 172.2183 | 719.8999 | 731.5752 | 497.6578 | 539.6165 | 495.944 | 653.6284 | 416.8908 | 629.5165 | 620.6018 | 343.6697 | 243.9027 | 660.2981
1238 | 15.6124 | 204.8363 | 308.4686 | 361.879 | 414.0807 | 340.7331 | 355.3476 | 278.2871 | 455.8565 | 794.9704 | 408.9371 | 243.4057 | 900.9371 | 586.6323 | 277.7331 | 252.325 | 674.6299 | 555.1815 | 477.4437 | 352.6524 | 732.8518 | 409.5694 | 690.1008 | 382.3049 | 400.204 | 595.955 | 600.834 | 443.8161 | 690.5919 | 238.2219 | 223.2847 | 180.1435 | 340.5919 | 374.3049 | 290.7734 | 177.3879 | 133.7331 | 503.1815 | 465.1815 | 354.0427 | 426.3274 | 369.7331 | 281.204 | 243.0629 | 307.3049 | 112.9395 | 80.3049 | 90.8161 | 204.0605 | 681.3429 | 544.5113 | 425.8766 | 451.8968 | 342.9573 | 266.6548 | 211.9395 | 255.5291 | 602.2622 | 373.6323 | 551.5244 | 612.2847 | 737.121 | 518.3476 | 764.6276 | 673.1412 | 301.4887 | 528.2017 | 567.1839 | 849.0961 | 517.3227 | 386.6928 | 468.3476 | 701.2064 | 522.7129 | 457.4461 | 286.2017 | 721.5089 | 473.5089 | 446.917 | 270.7758 | 637.3452 | 675.6299 | 634.7936 | 432.5291 | 668.1032 | 963.6276 | 1041.2064 | 592.834 | 643.5113 | 269.083 | 234.9193 | 142.4081 | 294.9573 | 741.4484 | 648.6928 | 325.1613 | 888.4686 | 449.6904 | 345.5896 | 248.0807 | 452.4057 | 342.204 | 157.0605 | 109.325 | 342.204 | 444.3856 | 342.917 | 305.8161 | 431.4484 | 374.121 | 326.2017 | 251.1032 | 407.5113 | 210.7556 | 186.3274 | 158.9193 | 302.3654 | 338.3429 | 348.3879 | 233.3476 | 191.0202 | 821.6703 | 751.7509 | 485.2669 | 553.7106 | 527.7331 | 709.5872 | 416.5694 | 722.5291 | 655.7106 | 376.4081 | 266.879 | 736.5492
1239 | 15.6234 | 209.6204 | 323.7306 | 330.8613 | 428.4857 | 367.649 | 348.8938 | 264.2203 | 465.7021 | 963.4407 | 433.1878 | 283.351 | 942.4938 | 607.7101 | 299.4572 | 275.0531 | 720.8653 | 589.4857 | 568.4612 | 371.004 | 841.6856 | 448.4081 | 725.1427 | 411.8693 | 408.7877 | 662.1062 | 656.5183 | 442.6735 | 706.5223 | 282.649 | 244.7591 | 165.1062 | 338.5674 | 404.0856 | 321.5388 | 180.6244 | 157.1878 | 548.302 | 504.6775 | 329.9224 | 417.4081 | 399.8938 | 289.1387 | 219.8938 | 337.9509 | 100.351 | 99.9184 | 90.7551 | 215.5674 | 767.8203 | 537.3836 | 441 | 485.7101 | 377.8938 | 249.4652 | 203.6775 | 281.4326 | 661.7877 | 387.5714 | 644.4001 | 632.0285 | 778.7386 | 519.6815 | 853.9755 | 718.5223 | 315.9715 | 576.004 | 573.302 | 958.4652 | 588.355 | 407.8122 | 466.653 | 681.1713 | 548.4612 | 510.9755 | 319.5103 | 757.9264 | 497.3796 | 474.1102 | 267.7306 | 670.1102 | 753.8814 | 675.8122 | 465.5714 | 674.31 | 1062.9264 | 1030.4732 | 645.7061 | 645.0366 | 253.7877 | 229.3836 | 143.7837 | 330.5428 | 767.2244 | 692.5263 | 354.8368 | 904.1632 | 508.9469 | 379.9715 | 257.9429 | 499.0571 | 349.7061 | 168.351 | 129.4572 | 349.7061 | 473.4286 | 371.8408 | 318.302 | 452.8979 | 406.4367 | 376.1673 | 239.355 | 396.0571 | 206.9755 | 176.1102 | 149.2734 | 342.1387 | 409.0285 | 349.4612 | 226.2449 | 202.5714 | 875.4326 | 829.653 | 469.1673 | 635.3345 | 579.7426 | 789.4612 | 457.5714 | 748.8653 | 739.9304 | 379.5143 | 246.4612 | 773.5714
1240 | 15.6344 | 199.6283 | 338.6874 | 311.1726 | 430.5443 | 373.8865 | 358.1991 | 295.1431 | 469.4588 | 1042.4323 | 425.7434 | 272.5412 | 1032.2644 | 689.577 | 289.4012 | 263.4277 | 742.0606 | 595.4308 | 579.8569 | 401.6003 | 941.3484 | 477.8305 | 788.0016 | 435.5412 | 439.5148 | 651.8865 | 697.6299 | 437.9736 | 787.6034 | 293.0016 | 266.1151 | 152.8289 | 339.5148 | 441.1151 | 315.8585 | 203.2862 | 149.5148 | 581.0855 | 529.2022 | 362.7729 | 449.3453 | 415.6003 | 317.972 | 236.972 | 378.1151 | 97.3142 | 98.2862 | 81.5148 | 212.1711 | 841.2582 | 585.9736 | 447.8585 | 552.0311 | 392 | 295.8009 | 209.4557 | 291.916 | 702.5739 | 430.7745 | 601.6563 | 659.4603 | 864.2613 | 569.0871 | 899.2053 | 798.2317 | 309.3437 | 617.8025 | 605.6283 | 1030.6345 | 619.1151 | 414.1135 | 495.2566 | 777.4059 | 575.1742 | 529.9145 | 297.6003 | 813.3717 | 516.7449 | 497.916 | 276.5148 | 702.1462 | 863.3981 | 707.9471 | 498.0295 | 774.4588 | 1151.521 | 1153.3188 | 673.6579 | 715.6003 | 292.2877 | 277.7449 | 145.972 | 347 | 832.6299 | 788.6579 | 338.7698 | 939.6625 | 520.5723 | 365.7994 | 243.5443 | 523.0575 | 382.6874 | 165.5428 | 115.972 | 382.6874 | 470.9191 | 400.0871 | 349.2566 | 499.6579 | 443.8585 | 412.5428 | 269.6579 | 427.1446 | 223.3997 | 202.6594 | 179.5443 | 365.2566 | 426.1726 | 379.6034 | 237.3142 | 229.0855 | 897.3468 | 874.0311 | 511.2582 | 700.5412 | 676.4292 | 804.972 | 490.7154 | 765.2597 | 811.3406 | 385 | 253.7418 | 805.8009
1241 | 15.6454 | 192.35 | 373.059 | 340.9136 | 466.1181 | 398.3226 | 361.5135 | 299.2409 | 509.0817 | 1138.1134 | 456.4407 | 251.2272 | 1181.8498 | 786.9636 | 313.8909 | 267.7544 | 812.3954 | 651.1044 | 585.8772 | 408.2045 | 1039.1682 | 523.0453 | 827.7681 | 416.1545 | 457.6363 | 689.3316 | 720.7227 | 481.0227 | 847 | 323.6136 | 284.0954 | 158.2773 | 355.4681 | 458.6136 | 334.4228 | 208.8318 | 163.5409 | 585.4681 | 565.0727 | 392.3727 | 492.2045 | 434.9726 | 336.9136 | 255.9136 | 396.0954 | 101.7454 | 110.3363 | 86.3864 | 221.418 | 888.6315 | 619.3864 | 470.7591 | 607.5635 | 401.3954 | 308.6273 | 189.7818 | 331.8772 | 754.1545 | 478.941 | 601.2135 | 722.4681 | 949.009 | 615.4043 | 1003.2863 | 871.5498 | 333.3954 | 678.1044 | 616.6589 | 1144.359 | 645.5272 | 417.8635 | 500.6953 | 890.1181 | 643.2362 | 540.709 | 300.5909 | 842.3316 | 560.709 | 522.4591 | 280.1819 | 771.0817 | 859.9088 | 782.9636 | 513.7955 | 816.4908 | 1306.3042 | 1254.6452 | 698.5272 | 749.6679 | 321.941 | 309.1819 | 153.8318 | 358.0817 | 873.068 | 797.6273 | 322.5635 | 1050.1044 | 542.2636 | 369.6953 | 262.9773 | 521.3274 | 410.3137 | 167.8545 | 123.5909 | 410.3137 | 570.6589 | 426.1682 | 348.4317 | 522.5999 | 481.2135 | 420.3954 | 277.6637 | 465.709 | 220.4091 | 238.6273 | 204.5182 | 376.9589 | 449.4091 | 445.0227 | 242.2272 | 230.0954 | 978.7681 | 942.8134 | 531.1682 | 685.2499 | 706.068 | 843.427 | 515.059 | 819.3954 | 793.6816 | 384.0364 | 246.8408 | 841.5135
1242 | 15.6564 | 196.6599 | 379.3785 | 379.1279 | 504.3299 | 439.3172 | 417.1764 | 313.1663 | 557.9259 | 1236.0384 | 535.6471 | 282.6087 | 1276.4962 | 804.4323 | 323.7596 | 320.8849 | 878.4605 | 714.6215 | 632.4451 | 418.8363 | 1074.6114 | 557.9616 | 887.7494 | 440.9488 | 497.8747 | 786.0026 | 745.2814 | 515.6983 | 854.8491 | 341.5831 | 292.4424 | 160.5576 | 384.5704 | 483.4195 | 334.4323 | 200.2276 | 180.4936 | 636.5986 | 562.2916 | 399.7084 | 494.734 | 480.7724 | 367.532 | 273.6189 | 410.7724 | 121.2276 | 118.4296 | 75.5064 | 267.7212 | 988.0768 | 629.0154 | 489.4962 | 659.1919 | 435.5576 | 302.3044 | 213.3172 | 360.2148 | 791.6087 | 469.8363 | 653.5959 | 772.839 | 1024.4349 | 689.2916 | 1060.2303 | 953.2942 | 361.734 | 721.3657 | 662.4552 | 1208.9029 | 713.105 | 491.2558 | 556.6215 | 929.3427 | 701.1764 | 579.4323 | 301.6087 | 925.8363 | 603.2303 | 528.041 | 271.0768 | 832.5858 | 971.0026 | 822.7137 | 539.8774 | 884.8262 | 1446.4962 | 1348.7878 | 736.9872 | 853.9259 | 327.2686 | 304.1279 | 148.266 | 387.1764 | 932.0256 | 818.6497 | 348.3657 | 1116.1535 | 571.1151 | 418.2788 | 250.0768 | 512.9898 | 426.2686 | 179.4936 | 120.8108 | 426.2686 | 612.6573 | 429.7084 | 381.0768 | 562.0768 | 528.2788 | 454.3044 | 268.5576 | 497.8491 | 235.3427 | 229.7724 | 193.8849 | 437.4296 | 473.4579 | 475.6471 | 272.5959 | 247.0512 | 1045.839 | 1034.1663 | 545.5959 | 723.9616 | 764.0128 | 941.8619 | 541.8875 | 862.4195 | 874.6599 | 395.6855 | 263.7825 | 908.5704
1243 | 15.6674 | 232.1416 | 366.6638 | 402.8938 | 537.8674 | 463.7611 | 409.8231 | 346.9557 | 600.062 | 1266.2391 | 585.3984 | 306.938 | 1349.1329 | 862.6638 | 341.3984 | 299.3098 | 988.9204 | 765.5666 | 684.9027 | 442.9823 | 1164.3098 | 570.7701 | 948.3894 | 430.3274 | 540.2744 | 853.7611 | 815.2478 | 543.6813 | 915.6994 | 366.5753 | 285.9027 | 173.469 | 400.7345 | 511.584 | 381.5133 | 202.9647 | 178.2035 | 707.3804 | 586.8674 | 390.531 | 501.6196 | 508.5666 | 372.0796 | 253.3541 | 436.9736 | 116.7965 | 106.7788 | 77 | 267.3717 | 1016.6196 | 699.2302 | 541.1592 | 707.9823 | 453.5133 | 320.8408 | 217.3185 | 355.1859 | 834.7878 | 510.9736 | 684.4956 | 859.9647 | 1142.9913 | 720.2391 | 1130.1948 | 1062.0888 | 369.1505 | 717.0799 | 661.5576 | 1329.0179 | 802.3717 | 547.0177 | 584.469 | 969.3187 | 694.3541 | 632.885 | 325.4071 | 974.1418 | 651.6902 | 603.1682 | 290.5929 | 910.2035 | 1028.1416 | 942.6815 | 624.2832 | 978.7878 | 1540.372 | 1446.2657 | 772.7614 | 904.8231 | 383.8938 | 337.1859 | 149.8761 | 373.7168 | 962.1595 | 929.5223 | 324.6992 | 1185.1861 | 615.5223 | 421.0177 | 270.3894 | 570.3274 | 513.425 | 187.8231 | 138.4514 | 513.425 | 615.1595 | 440.7081 | 399.2655 | 598.8497 | 543.2302 | 481.0709 | 287.0443 | 541.7081 | 262.8408 | 233.1416 | 169.9204 | 432.416 | 536.6019 | 519.2921 | 292.3451 | 272.2744 | 1143.3187 | 1088.3984 | 579.416 | 728.8054 | 790.894 | 980.7435 | 574 | 891.115 | 949.6994 | 445.1062 | 229.1948 | 946.5046
1244 | 15.6784 | 233.3889 | 408.8889 | 392.9445 | 577.0555 | 483.8333 | 447.2779 | 343.5556 | 616.889 | 1337.2223 | 632.1667 | 319.2779 | 1394.6111 | 924.889 | 383.4444 | 323.7223 | 1030.3336 | 828.8333 | 719.0001 | 473.1668 | 1170.111 | 623.2222 | 968.0557 | 437.6111 | 577.8333 | 884.9445 | 865.3336 | 511.3335 | 1039.7224 | 383.9445 | 314.7223 | 170.6111 | 416.1112 | 507.2224 | 402.889 | 223.8889 | 180.7223 | 716.6114 | 631.8889 | 395.8889 | 530.4444 | 571.5556 | 366.2222 | 282 | 501.9445 | 121.2222 | 111.3888 | 73.3889 | 283.6111 | 1056.8334 | 740.3335 | 533.6112 | 740.3891 | 468.7778 | 339.2222 | 190.5556 | 372.0556 | 940.9449 | 565.9444 | 732.1669 | 907.278 | 1198.6112 | 777.6111 | 1187.7779 | 1165.7777 | 399.7222 | 801.7223 | 701.3334 | 1420.4446 | 812.7777 | 549.889 | 583.8333 | 1030.2779 | 730.5001 | 670.7778 | 323.5 | 1088.3892 | 661.7224 | 633.5556 | 300.2778 | 941.3336 | 1055.2225 | 993.3335 | 614.5556 | 1028 | 1641.2778 | 1575.5559 | 880.5556 | 952.0002 | 367.5556 | 345.7222 | 149.8334 | 401.5001 | 1046.5001 | 986.6112 | 343.2222 | 1306.7225 | 690.6669 | 426.389 | 281.4445 | 567.6111 | 582.6667 | 220.5556 | 149.6667 | 582.6667 | 687.8334 | 518.0002 | 400.2778 | 667.389 | 583.5001 | 521.1111 | 314.889 | 612.0557 | 290.389 | 241.0557 | 184.1112 | 467.0556 | 602.8335 | 548.3334 | 302.2224 | 312.8889 | 1203.1667 | 1154.6113 | 614.3334 | 769.5 | 855.0555 | 1030.4446 | 585.6668 | 851.8333 | 1060.6667 | 481.1668 | 264.5555 | 1024.6113
1245 | 15.6894 | 258.122 | 414.122 | 411.5801 | 591.7248 | 497.9461 | 483.8624 | 371 | 671.8546 | 1350.4355 | 657.7936 | 361.0305 | 1393.5574 | 1004.7475 | 391.5723 | 358.2823 | 1143.1064 | 863.778 | 782.656 | 512.2674 | 1148.1603 | 602.4199 | 1038.549 | 461.7936 | 585.2213 | 934.9078 | 980.9539 | 578.4348 | 1123.488 | 388.4199 | 343.7709 | 162.0688 | 466.122 | 573.939 | 431.8014 | 222.4504 | 204.3128 | 806.5574 | 639.0993 | 407.4504 | 540.8929 | 589.8241 | 388.244 | 291.2823 | 524.061 | 140.3433 | 96.2518 | 70.3816 | 295.9539 | 1094.2823 | 814.3965 | 560.9312 | 807.6028 | 486.4426 | 358.0532 | 208.7709 | 406.0532 | 1094.5114 | 577.4348 | 816.4653 | 979.6716 | 1228.2518 | 794.0149 | 1265.0908 | 1159.183 | 395.5191 | 873.7014 | 746.4426 | 1516.5412 | 802.4121 | 603.1142 | 584.8397 | 1070.9234 | 774.366 | 695.6411 | 332.9007 | 1169.6567 | 721.3589 | 673.183 | 308.1986 | 1021.0383 | 1114.344 | 1067.3965 | 636.0149 | 1049.7553 | 1653.0227 | 1703.8624 | 908.6865 | 995.878 | 376.5418 | 357.7631 | 165.5801 | 445.5496 | 1076.3362 | 1006.6794 | 342.3816 | 1400.8241 | 763.1908 | 459.7177 | 289.2291 | 564.6333 | 595.7021 | 234.1298 | 147.6184 | 595.7021 | 718.9539 | 578.4121 | 414.0532 | 701.5879 | 617.3972 | 529.5114 | 348.0993 | 655.5723 | 312.4582 | 276.4504 | 202.6106 | 495.8851 | 648.5347 | 592.3887 | 337.4277 | 322.1298 | 1244.6943 | 1226.4426 | 623.8702 | 765.5574 | 850.8702 | 1102.2284 | 615.4199 | 819.878 | 1093.9539 | 529.2596 | 247.5418 | 1061.3972
1246 | 15.7004 | 274.0693 | 434.3024 | 426.9087 | 656.3276 | 571.3968 | 514.652 | 357.0913 | 739.1653 | 1272.0425 | 720.9102 | 356.6268 | 1322.041 | 1085.4425 | 411.0693 | 367.3937 | 1218.4898 | 967.8866 | 861.1417 | 495.3732 | 1132.5559 | 607.9543 | 1154.0031 | 497.3953 | 613.1181 | 1055.0283 | 1056.2835 | 620.5575 | 1232.3291 | 389.7213 | 364.4898 | 154.2331 | 493.5591 | 572.7685 | 442.9764 | 230.2331 | 225.0457 | 801.2583 | 658.9102 | 422.1874 | 578.4661 | 654.2819 | 433.1402 | 315.5118 | 558.189 | 172.0929 | 117.6283 | 87.8142 | 335.6047 | 1112.7669 | 890.7921 | 594.7701 | 840.8378 | 519.2567 | 399.6756 | 216.1858 | 431.0457 | 1089.8583 | 642.8394 | 855.6961 | 971.2063 | 1262.6299 | 878.074 | 1390.8394 | 1195.9543 | 412.7008 | 1005.2346 | 781.6756 | 1605.6252 | 824.8142 | 651.652 | 600.7465 | 1110.1402 | 834 | 739.6551 | 343.6283 | 1170.2598 | 723.6063 | 712.6756 | 309.4205 | 1028.189 | 1022.252 | 1138.9543 | 694.0709 | 1100.2787 | 1651.8803 | 1749.7701 | 993.6535 | 976.6047 | 399.2331 | 377.5102 | 167.9071 | 429.9055 | 992.6457 | 1038.7953 | 374.9323 | 1450.1638 | 799.1197 | 460.2819 | 274.8598 | 603.4425 | 617.3024 | 237.4425 | 138.0472 | 617.3024 | 768.2803 | 624.0961 | 455.0709 | 701.2583 | 611.6772 | 558.1181 | 344.9307 | 684.1402 | 305.4898 | 289.3732 | 203.5575 | 537.3244 | 616.2583 | 681.1165 | 330.3276 | 329.3732 | 1255.6473 | 1235.9748 | 641.7228 | 736.0693 | 847.5575 | 1197.7921 | 610.674 | 770.6709 | 1136.326 | 538.6063 | 279.6063 | 1052.3512
1247 | 15.7114 | 268.7758 | 456.7111 | 449.2458 | 710.0473 | 629.2627 | 548.2499 | 346.0818 | 798.0689 | 1153.2072 | 781.7414 | 346.9568 | 1212.2242 | 1143.9394 | 425.599 | 351.1034 | 1282.4437 | 1051.0041 | 920.6206 | 527.3836 | 1066.401 | 609.8406 | 1211.4828 | 497.2586 | 629.4355 | 1143.819 | 1152.2715 | 609.2586 | 1322.3404 | 404.4956 | 402.3275 | 169.3491 | 498.7285 | 578.2113 | 459.2843 | 241.2586 | 217.181 | 861.5301 | 736.1034 | 446.7414 | 598.944 | 704.8535 | 464.1938 | 321.8319 | 599.2113 | 158.6725 | 125.6853 | 86.9655 | 345.8103 | 1075.6252 | 943.1723 | 645.6293 | 846.0216 | 519.9527 | 435.2154 | 228.9913 | 408.3923 | 1041.7069 | 682.7414 | 853.9481 | 911.1681 | 1262.6123 | 992.0386 | 1455.2843 | 1179.2761 | 467.9095 | 1055.0689 | 809.9784 | 1548.6509 | 791.5563 | 703.8144 | 669.1723 | 1090.5347 | 824.8751 | 817.8319 | 363.6422 | 1241.6509 | 734.2329 | 767.7239 | 337.763 | 1053.4786 | 930.7501 | 1170.1034 | 726.0129 | 1051.5218 | 1536.3408 | 1817.6206 | 1053.474 | 956.9182 | 386.6596 | 371.7671 | 171.2586 | 416.7887 | 865.0303 | 1132.2931 | 408.3362 | 1465.8535 | 886.3404 | 518.0905 | 271.4956 | 633.6206 | 625.8663 | 260.0689 | 150.1121 | 625.8663 | 759.9054 | 681.056 | 482.6077 | 728.8062 | 648.8966 | 601.181 | 360.2931 | 702.6077 | 325.0776 | 308.1681 | 201.6982 | 519.4051 | 635.9398 | 691.9784 | 365.1681 | 354.4612 | 1198.6206 | 1212.2154 | 696.8792 | 713.7846 | 835.3147 | 1208.3234 | 615.9008 | 691.6725 | 1115.3666 | 559.3018 | 295.582 | 1091.9655
1248 | 15.7224 | 278.3078 | 495.4994 | 429.3078 | 759.7125 | 707.3855 | 600.4239 | 410.3089 | 829.4814 | 1056.4036 | 821.5028 | 385.0969 | 1130.4375 | 1258.3855 | 484.8467 | 394.4622 | 1361.7114 | 1117.2311 | 965.2119 | 569.9436 | 993.4217 | 549.902 | 1225.9436 | 513.7317 | 590.6347 | 1145.6167 | 1200.9053 | 631.6167 | 1403.0405 | 410.5761 | 407.0564 | 176.7497 | 484.0969 | 526.0564 | 515.5581 | 253.1544 | 242.8467 | 931.868 | 790.2514 | 454.7892 | 631.0789 | 689.1725 | 501.7497 | 332.9042 | 573.8658 | 142.8275 | 122.3844 | 77.3078 | 370.2311 | 997.6911 | 1020.5592 | 665.3089 | 810.1139 | 485 | 462.3653 | 264.6922 | 382.557 | 1010.345 | 707.1364 | 924.8467 | 867.0745 | 1141.7284 | 1098.983 | 1512.2119 | 1057.9775 | 470.1161 | 1124.7339 | 827.5772 | 1500.0553 | 685.1522 | 780.4036 | 735.443 | 1005.3258 | 785.3833 | 809.9414 | 380.8264 | 1211.0372 | 679.8456 | 813.8253 | 326.5772 | 963.787 | 871.0564 | 1185.345 | 776.4825 | 939.7103 | 1355.707 | 1882.4825 | 1113.1172 | 885.1714 | 351.7317 | 419.6156 | 176.6156 | 415.3247 | 791.1331 | 1140.4419 | 404.5006 | 1431.5367 | 938.9233 | 532.2311 | 292.6156 | 680.1736 | 627.9606 | 280.6922 | 157.443 | 627.9606 | 715.8286 | 710.3866 | 509.4239 | 671.3067 | 588.5367 | 644.5017 | 387.7317 | 690.5175 | 297.9031 | 307.5581 | 210.6145 | 519.8456 | 562.9414 | 714.8083 | 359.9808 | 334.2097 | 1194.1703 | 1198.8242 | 728.8456 | 692.5197 | 811.8061 | 1159.2492 | 656.0969 | 632.3439 | 1051.0383 | 557.5783 | 261.3461 | 1039.805
1249 | 15.7334 | 277.6076 | 511.9235 | 431.3215 | 798.8201 | 731.233 | 637.534 | 439.8584 | 827.2889 | 1000.4894 | 923.9264 | 412.7139 | 938.9702 | 1297.8378 | 498.9646 | 409.2153 | 1362.767 | 1155.4781 | 991.5694 | 610.6608 | 924.5751 | 491.6962 | 1327.7231 | 535.2153 | 568.5574 | 1236.2068 | 1252.7847 | 671.6608 | 1489.0149 | 405.6934 | 387.3718 | 180.2125 | 495.4306 | 459.8966 | 543.2833 | 270.6431 | 253.5722 | 1025.4986 | 838.0708 | 464.1091 | 708.2656 | 652.6813 | 517.0857 | 345.3746 | 574.4306 | 166.4278 | 124.8024 | 81.3569 | 380.6608 | 925.0588 | 1101.602 | 707.0708 | 721.6133 | 447.6841 | 466.8024 | 270.8201 | 349.5014 | 953.3066 | 783.6934 | 941 | 754.7372 | 1037.8612 | 1164.8938 | 1531.1062 | 937.699 | 512.3364 | 1252.8527 | 858.0326 | 1394.8258 | 600.4157 | 813.0298 | 749.5014 | 926.8994 | 741.6076 | 813.5666 | 370.1062 | 1157.7521 | 619.6133 | 760.6431 | 342.4455 | 889.8789 | 783.1855 | 1170.3718 | 814.7167 | 868.4511 | 1185.8966 | 1897.9879 | 1181.1771 | 800.1119 | 357.9497 | 426.6962 | 176.233 | 351.0885 | 713.5042 | 1114.7521 | 417.5899 | 1339.3803 | 988.3867 | 557.5871 | 285.7875 | 712.9441 | 585.1268 | 272.5722 | 178.2861 | 585.1268 | 723.5425 | 782.4632 | 543.4809 | 634.9292 | 522.5014 | 681.3952 | 411.2507 | 610.8966 | 268.5899 | 324.767 | 202.0326 | 493.1976 | 498.7195 | 723.466 | 391.1742 | 277.3392 | 1077.5042 | 1101.8229 | 723.9086 | 668.3095 | 729.0942 | 1102.1296 | 686.4278 | 542.3421 | 1026.9851 | 638.6374 | 246.0914 | 945.5014
1250 | 15.7444 | 296.6614 | 572.3756 | 458.1221 | 832.8733 | 743.9038 | 683.2007 | 435.2488 | 730.5134 | 855.7317 | 987.0981 | 438.037 | 822.3321 | 1316.2969 | 511.1702 | 424.1221 | 1378.9214 | 1191.7752 | 1041.2007 | 618.7882 | 818.9982 | 471.0962 | 1446.3515 | 519.5393 | 524.6244 | 1310.7993 | 1307.8316 | 677.679 | 1549.0916 | 410.8474 | 437.4366 | 192.3996 | 451.7752 | 404.321 | 566.5634 | 269.7576 | 250.0305 | 1062.2794 | 860.4126 | 454.1156 | 755.5763 | 611.0241 | 574.1943 | 361.9519 | 531.1267 | 174.1878 | 148.4366 | 86.1092 | 370.5088 | 836.963 | 1175.679 | 711.8363 | 612.5264 | 376.0546 | 481.2969 | 284.8363 | 313.0722 | 888.7447 | 855.6374 | 958.2248 | 676.0481 | 944.2229 | 1227.5236 | 1565.4977 | 834.7493 | 552.4607 | 1318.3821 | 907.0851 | 1255.0398 | 520.5874 | 883.9695 | 766.8557 | 813.9325 | 690.7077 | 871.7271 | 384.8122 | 1096.9565 | 525.9935 | 758 | 358.8668 | 812.6355 | 643.1378 | 1169.1138 | 814.2858 | 754.6226 | 1056.1489 | 1758.0222 | 1194.321 | 706.7142 | 298.4783 | 428.2424 | 170.6244 | 334.9871 | 621.7928 | 1080.321 | 410.4061 | 1222.8955 | 1037.7142 | 582.1332 | 274.6549 | 764.3885 | 521.926 | 290.4736 | 169.4847 | 521.926 | 619.1813 | 821.6068 | 559.6485 | 598.8779 | 481.8538 | 671.3885 | 414.3034 | 537.3932 | 245.2359 | 337.0546 | 240.5393 | 461.8298 | 439.4912 | 699.5939 | 432.5328 | 259.3691 | 949.5856 | 1020.8344 | 743.2664 | 562.5264 | 649.8844 | 1005.4413 | 702.9759 | 489.8298 | 955.6596 | 713.8668 | 243.6068 | 858.1008
1251 | 15.7554 | 332.3578 | 570.3678 | 481.6222 | 877.0974 | 783.6342 | 763.837 | 456.3638 | 636.1849 | 727.3459 | 1046.7256 | 507.3797 | 700.4294 | 1319.002 | 545.002 | 464.7217 | 1307.2325 | 1121.4354 | 1047.2544 | 658.2843 | 686.8966 | 408.4453 | 1450.6382 | 523.7316 | 503.7117 | 1258.5407 | 1394.0954 | 707.0994 | 1573.4473 | 340.5427 | 438.4473 | 180.0954 | 395.6242 | 385.4513 | 571.01 | 276.6382 | 261.1809 | 1120.7396 | 923.0179 | 495.8251 | 787.2724 | 565.4373 | 617.7396 | 370.9026 | 473.8051 | 166.8171 | 155.2684 | 94.002 | 358.0934 | 765.4373 | 1234.2048 | 701.6342 | 558.1809 | 346.7117 | 501.1014 | 267.7217 | 275.6342 | 806.5268 | 914.3638 | 973.2644 | 621.6203 | 814.7912 | 1355.9046 | 1580.0775 | 692.4473 | 592.2903 | 1362.9245 | 918.3618 | 1073.6222 | 462.7117 | 872.0915 | 831.1809 | 735.6222 | 589.7177 | 864.9105 | 386.7256 | 1013.4433 | 490.3578 | 760.1829 | 365.2724 | 739.2604 | 507.5228 | 1073.992 | 859.994 | 629.2584 | 855.3379 | 1568.6918 | 1195.4612 | 604.2505 | 258.1769 | 413.3578 | 188.6402 | 285.5447 | 541.9026 | 1037.4393 | 426.4612 | 1114.5348 | 977.4533 | 607.7436 | 293.2704 | 826.3758 | 441.8151 | 343.3638 | 162.6362 | 441.8151 | 547.4334 | 854.6302 | 559.7276 | 518.9861 | 433.3598 | 701.3598 | 441.002 | 486.1789 | 219.8191 | 334.6342 | 239.2744 | 407.6242 | 430.0775 | 667.5308 | 417.8251 | 223.9085 | 795.99 | 899.996 | 750.1014 | 513.2744 | 577.2505 | 897.6163 | 723.0954 | 434.169 | 840.6998 | 759.5646 | 286.4533 | 716.8111
1252 | 15.7664 | 321.4027 | 603.4293 | 453.8143 | 870.2874 | 763.9645 | 842.7877 | 448.7256 | 619.9202 | 599.5571 | 1059.8054 | 545.9468 | 578.9025 | 1306.8185 | 583.217 | 464.1815 | 1159.23 | 1042.3406 | 971.7167 | 717.8054 | 599.7475 | 369.23 | 1463.3938 | 561.5665 | 422.6724 | 1246.8894 | 1443.4737 | 748.2833 | 1552.5133 | 322.7256 | 411.4956 | 205.4027 | 351.1195 | 353.6901 | 604.4956 | 292.0266 | 254.2389 | 1153.0929 | 985.6239 | 520.7522 | 827.3719 | 490.7079 | 676.0443 | 362.2921 | 421.9823 | 164.1283 | 144.8894 | 95.7345 | 345.9379 | 650.4643 | 1327.4116 | 703.7788 | 514.0042 | 298.77 | 530.4778 | 258.531 | 252.9557 | 715.4246 | 961.6374 | 968.8232 | 533.4069 | 678.252 | 1351.6594 | 1493.8847 | 606.4465 | 699.5091 | 1424.8983 | 921.4027 | 949.5305 | 400.1017 | 898.8232 | 816.7167 | 667.5707 | 530.8007 | 868.4956 | 384.7611 | 919.4643 | 436.5263 | 724.5086 | 363.2478 | 641.8404 | 422.3584 | 987.3271 | 833.3584 | 557.9379 | 692.035 | 1396.9864 | 1225.0177 | 495.9113 | 215.8007 | 381.3406 | 189.9645 | 258.1768 | 493.0485 | 963.314 | 440.5973 | 1042.398 | 927.1148 | 693.5842 | 295.1549 | 864.8628 | 399.783 | 313.5352 | 163.5044 | 399.783 | 458.7079 | 855.1904 | 587.3276 | 461.3495 | 392.168 | 694.0177 | 444.2389 | 447.0396 | 194.7919 | 325.1195 | 243.8717 | 345.0663 | 343.5086 | 591.553 | 437.1106 | 205.4512 | 730.1591 | 825.8226 | 781.3584 | 486.6547 | 487.7167 | 775.4731 | 720.9468 | 395.77 | 712.0308 | 821.3584 | 297.4204 | 653.2608
1253 | 15.7774 | 335.8348 | 649.1476 | 499.71 | 832.2636 | 735.0668 | 886.5272 | 458.6924 | 532.6645 | 470.0264 | 1088.2812 | 542.0228 | 495.2584 | 1229.5573 | 645.4868 | 496 | 1107.2181 | 1002.0264 | 944.0668 | 757.5448 | 505.0844 | 346.348 | 1474.3884 | 591.232 | 367.7364 | 1212.2761 | 1467.1512 | 767.3884 | 1513.7805 | 284.1424 | 410.942 | 228.768 | 325.4728 | 309.1248 | 595.3972 | 309.3304 | 241.7856 | 1101.5976 | 1042.8715 | 528.9912 | 904.5448 | 435.928 | 701 | 377.0668 | 387.5308 | 165.2232 | 153.7768 | 89.3304 | 307.8928 | 508.7313 | 1356.9332 | 741.7504 | 413.3744 | 294.464 | 525.768 | 283.768 | 219.4552 | 634.9368 | 1057.2548 | 991.7452 | 459.2496 | 579.87 | 1404.014 | 1380.4501 | 492.7944 | 756.7452 | 1413.1424 | 964.5764 | 780.9457 | 371.8524 | 925.2408 | 780.522 | 599.3252 | 468.4148 | 834.928 | 395.2724 | 794.044 | 368.5712 | 645.0176 | 370.9332 | 517.4324 | 387.754 | 842.6153 | 815.0668 | 486.8788 | 547.696 | 1220.3465 | 1231.5044 | 416.87 | 177.884 | 360.8348 | 193.304 | 215.8348 | 427.0844 | 880.8385 | 452.5764 | 922.4677 | 830.8032 | 746.4376 | 310.8348 | 870.0404 | 326.638 | 281.6432 | 173.768 | 326.638 | 415.1916 | 890.4464 | 632.942 | 404.6645 | 319.2004 | 666.7136 | 441.4956 | 372.2004 | 153.4376 | 306.464 | 238.8928 | 284.4236 | 275.2812 | 526.9192 | 430.5536 | 175.6696 | 652.2004 | 694.1512 | 776.2724 | 434.5536 | 442.2004 | 642.2269 | 693.7188 | 389.522 | 625.986 | 852.0051 | 308.2408 | 587.058
1254 | 15.7884 | 353.2331 | 685.0207 | 531.829 | 801.9171 | 694.9551 | 914.7876 | 492.3938 | 455.582 | 420.5717 | 1077.9896 | 574.1848 | 425.1572 | 1126.3454 | 684.038 | 512.829 | 985.7012 | 903.2867 | 855.4716 | 755.3731 | 437.3454 | 301.9551 | 1410.4888 | 587.0138 | 339.9827 | 1076.8826 | 1417 | 792.0414 | 1394.9206 | 242.5959 | 397.5717 | 242.2021 | 273.9724 | 274.7703 | 576.9862 | 301.7807 | 229.1952 | 1027.3627 | 1123.582 | 558.0311 | 955.6649 | 369.7703 | 724.2401 | 365.0035 | 330.1675 | 149.9758 | 159.4076 | 101.0138 | 287.7738 | 423.9931 | 1340.5475 | 779.2055 | 340.7634 | 247.962 | 537.5993 | 297.6028 | 200.791 | 548.9551 | 1106.8359 | 977.8256 | 400.9586 | 479.5061 | 1455.1986 | 1306.0035 | 445.5544 | 754.4456 | 1384.0173 | 970.3351 | 658.7116 | 331.171 | 922.8359 | 779.0449 | 512.753 | 416.7669 | 740.722 | 406.8083 | 707.3454 | 332.9724 | 580.9275 | 394.2159 | 457.3903 | 330.3696 | 722.722 | 805.8083 | 408.3627 | 488.5613 | 985.2626 | 1213.9793 | 364.589 | 173.9827 | 321.7358 | 217.5786 | 203.582 | 378.5786 | 772.1434 | 485.582 | 789.9413 | 736.893 | 780.2435 | 309.8014 | 953.5112 | 273.7669 | 290.5648 | 186.4007 | 273.7669 | 351.9448 | 856.7461 | 618.3696 | 326.38 | 262.5751 | 615.2124 | 443.3938 | 297.9448 | 153.1848 | 262.3662 | 242.8152 | 239.7979 | 254.171 | 462.9551 | 447.8256 | 157.1744 | 566.7254 | 597.9206 | 792.2159 | 427.3834 | 367.5441 | 544.1606 | 724.4456 | 362.8014 | 549.9586 | 914.1986 | 290.209 | 560.3593
1255 | 15.7994 | 393.9765 | 707.5428 | 549.9793 | 693.2446 | 623.7874 | 926.301 | 502.2803 | 408.2004 | 375.7639 | 1046.286 | 602.2982 | 389.6519 | 1022.4215 | 720.7168 | 555.7582 | 808.7311 | 774.0264 | 723.7639 | 761.3217 | 385.0442 | 254.7404 | 1301.6106 | 610.9765 | 313.0235 | 947.7432 | 1357.1177 | 839.9765 | 1255.7931 | 228.2625 | 357.7197 | 236.1533 | 238.8495 | 241.1533 | 546.4158 | 292.4337 | 215.5014 | 957.4187 | 1112.1298 | 569.7197 | 1019.8467 | 321.286 | 743.3067 | 379.9765 | 280.9379 | 120.7404 | 174.6726 | 97.7197 | 255.0884 | 397.3509 | 1301.1947 | 787.3039 | 301.5663 | 220.3688 | 535.7611 | 303.3039 | 187.2183 | 457.1384 | 1161.9529 | 983.0678 | 346.4158 | 381.8937 | 1419.3095 | 1287.9379 | 386.1769 | 832.8203 | 1353.985 | 930.9322 | 550.6134 | 296.9144 | 927.5692 | 810.9793 | 450.9379 | 361.9822 | 644.5043 | 431.3452 | 635.2211 | 297.0235 | 503.9586 | 412.9558 | 433.3067 | 298.2183 | 609.159 | 777.8966 | 359.0678 | 425.6783 | 804.4657 | 1187.7404 | 327.7432 | 158.587 | 258.9793 | 202.7168 | 180.5663 | 358.7611 | 681.7018 | 503.8645 | 715.873 | 622.112 | 844.4308 | 297.3716 | 1089.2333 | 238.3923 | 264.9558 | 191.1298 | 238.3923 | 290.3716 | 773.6576 | 571.7639 | 301.5663 | 226.1327 | 550.0114 | 441.3039 | 254.5428 | 132.5663 | 229.2832 | 247.3274 | 224.3716 | 212.8937 | 390.1355 | 462.5663 | 142.0207 | 482.1769 | 504.2211 | 804.7611 | 430.3452 | 316.1091 | 476.9615 | 728.4836 | 371.4337 | 487.5692 | 910.9351 | 310.3894 | 501.0471
1256 | 15.8104 | 418.6549 | 769.2854 | 562.1142 | 601.0244 | 556.3111 | 934.1847 | 507.4144 | 350.6671 | 337.644 | 987.2078 | 671.6318 | 390.2989 | 882.0366 | 792.7678 | 580.5747 | 681.9769 | 664.0827 | 671.61 | 803.4253 | 364.8165 | 240.8736 | 1240.1969 | 633.9538 | 278.9891 | 848.6113 | 1247.3342 | 851.8967 | 1095.6113 | 210.0231 | 333.0231 | 217.3451 | 219.1725 | 226.5978 | 485.1276 | 288.966 | 190.0924 | 851.2662 | 1139.2636 | 563.3098 | 1063.746 | 268.1616 | 724.4022 | 394.0231 | 236.1155 | 111.9769 | 184.4484 | 84.0802 | 221.0122 | 363.322 | 1329.1373 | 815.8153 | 272.7133 | 206.7364 | 516.5638 | 311.4022 | 168.3682 | 377.7473 | 1200.678 | 946.2418 | 304.6902 | 339.0122 | 1329.1738 | 1208.6684 | 330.3573 | 884.1725 | 1264.9089 | 962.2296 | 496.2187 | 268.6902 | 839.3926 | 840.5504 | 397.61 | 325 | 594.644 | 441.1155 | 584.6902 | 270.644 | 448.2771 | 423.4022 | 394.8396 | 279.7935 | 497.6453 | 706.3111 | 329.0462 | 364.8627 | 666.3926 | 1167.4945 | 276.2418 | 151.1956 | 238.7935 | 206.2989 | 172.1264 | 345.0924 | 583.4496 | 524.7825 | 649.4035 | 561.2309 | 922.8262 | 275.0231 | 1152.2296 | 210.288 | 251.5869 | 196.1495 | 210.288 | 283.3329 | 694.3573 | 540.4484 | 270.5869 | 198.1956 | 449.2296 | 440.4715 | 254.1956 | 122.8505 | 209.1155 | 228.6209 | 195.2187 | 168.7364 | 320.3111 | 429.8858 | 146.4715 | 450.1725 | 438.8056 | 825.9647 | 432.0353 | 286.0584 | 421.9198 | 694.1604 | 373.9198 | 437.6671 | 901.6209 | 319.0693 | 475.9307
1257 | 15.8214 | 454.8163 | 890.3818 | 604.6893 | 510.6822 | 508.1848 | 918.1899 | 500.3147 | 311.6842 | 302.2467 | 905.0558 | 710.5036 | 349.0568 | 745.4294 | 923.6934 | 627.2569 | 635.6802 | 546.5584 | 559.9258 | 802.6873 | 342.1868 | 223.1228 | 1121.0538 | 660.8152 | 247.6863 | 720.6182 | 1158.4294 | 884.3807 | 962.3675 | 188.935 | 314.998 | 215.3137 | 210.0619 | 202.7472 | 437.5624 | 290.1279 | 172.2497 | 734.9309 | 1157.936 | 601.065 | 1164.4457 | 222.4975 | 713.6853 | 401.8142 | 214.2497 | 116.8751 | 194.5645 | 97.6254 | 184.3736 | 339.4975 | 1235.9827 | 830.4975 | 256.7513 | 179.3127 | 487.3117 | 324.7513 | 155.7503 | 338.7452 | 1228.8782 | 920.6863 | 274.8112 | 304.5614 | 1216.2416 | 1099.5563 | 257.9949 | 946.8224 | 1174.9918 | 968.0619 | 456.3086 | 229.1848 | 742.5584 | 872.9969 | 360.7492 | 309.2477 | 549.1827 | 448.129 | 539.4964 | 253.6244 | 381.1848 | 429.7503 | 370.8122 | 255.0609 | 381.6172 | 698.8782 | 310.4985 | 331.5614 | 553.8061 | 1127.9959 | 242.3726 | 135.4365 | 223.6873 | 210.1259 | 167.999 | 318.936 | 502.4335 | 500.2518 | 584.3096 | 502.1848 | 1007.7543 | 296.8163 | 1198.3178 | 191.1259 | 208.2477 | 190.1868 | 191.1259 | 266.3076 | 611.9309 | 528.2487 | 221.5594 | 179.8112 | 421.8071 | 424.6863 | 227.935 | 114.8112 | 190.7503 | 214.7492 | 181.5624 | 138.6873 | 265.6213 | 402.6254 | 122.8101 | 436.2487 | 383.998 | 835.9969 | 388.7482 | 280.7543 | 377.3076 | 695.4345 | 356.4365 | 413.5614 | 864.1208 | 308.7503 | 463.4964
1258 | 15.8324 | 491.5988 | 932.9055 | 629.5472 | 463.5501 | 440.8139 | 961.9426 | 542.9426 | 275.3525 | 268.9513 | 812.2609 | 738.4012 | 286.4041 | 635.7623 | 988.5959 | 717.6904 | 554.8052 | 494.2551 | 439.0058 | 778.9571 | 316.6562 | 214.3467 | 994.516 | 680.5015 | 230.702 | 611.9658 | 1009.1751 | 912.0574 | 842.0174 | 153.5044 | 282.6046 | 243.9942 | 199.1519 | 187.447 | 416.3067 | 314.2006 | 160.1032 | 641.9113 | 1095.0661 | 596.0087 | 1214.5073 | 195.0516 | 675.556 | 414.7507 | 199.8539 | 119.3496 | 192.2064 | 105.9484 | 162.1032 | 310.2522 | 1064.556 | 811.5988 | 250.9055 | 181.8997 | 435.0145 | 329.702 | 153.9513 | 295.5501 | 1231.8081 | 873.1119 | 260.3009 | 281.9542 | 1070.9229 | 1022.5073 | 216.5988 | 1096.5298 | 1027.4244 | 931.5102 | 409.4041 | 223.444 | 670.9113 | 843.3009 | 355.7994 | 281.8023 | 481.8081 | 489.298 | 491.8052 | 244.4012 | 334.9571 | 428.851 | 348.6562 | 233.8023 | 318.8481 | 676.1664 | 281.7049 | 287.811 | 450.665 | 1061.8597 | 223.2493 | 132.3983 | 207.6046 | 223.3983 | 167.5472 | 319.9455 | 435.0603 | 542.9913 | 570.2464 | 441.5618 | 1046.4499 | 320.553 | 1246.7994 | 187.3038 | 180.8023 | 188.0487 | 187.3038 | 216.2522 | 536.9055 | 527.7965 | 192.2006 | 168.4499 | 351.0574 | 422.9971 | 194.3038 | 103.4499 | 160.6105 | 212.4985 | 174.2522 | 141.9484 | 221.7536 | 392.9542 | 107.3983 | 386.7623 | 366.4499 | 794.1548 | 359.5044 | 281.6621 | 328.6017 | 664.5015 | 334.5044 | 390.5044 | 813.9542 | 316.8481 | 419.8539
1259 | 15.8434 | 496 | 966.5332 | 638.1514 | 421.4069 | 369.5142 | 935.0046 | 569.6136 | 265.0757 | 242.7413 | 706.2996 | 768.6609 | 257.5978 | 556.6498 | 1023.8723 | 749.3785 | 483.7018 | 439.4148 | 398.5142 | 751.5379 | 297.1593 | 224.2271 | 862.8295 | 718.9039 | 226.235 | 531.2034 | 871.4826 | 894.6294 | 719.6182 | 147.6294 | 243.7413 | 257.2192 | 186.9164 | 191.1435 | 375.7492 | 313.3864 | 138.2949 | 574.0441 | 989.634 | 608.2903 | 1174.0599 | 189 | 628.5063 | 438.8045 | 173.9085 | 116.8407 | 180.4779 | 107.3785 | 143.0678 | 293.6814 | 964.8216 | 828.4779 | 233.235 | 177.9164 | 384.6215 | 331.7051 | 150.3864 | 279.287 | 1177.8927 | 799.347 | 257 | 275.2429 | 915.6024 | 937.2318 | 218.2271 | 1190.5615 | 871.2239 | 897.6215 | 395.3864 | 225.2034 | 578.5504 | 850.1672 | 345.0599 | 259.1356 | 424.1955 | 494.6893 | 453.9763 | 221.9763 | 301.0678 | 418.9921 | 331.0079 | 229.1593 | 296.8848 | 611.2271 | 252.8328 | 269.5694 | 381.0599 | 1021.3028 | 212.5978 | 127.4464 | 172.4385 | 209.2034 | 171.0678 | 319.4385 | 396.9243 | 558.4385 | 552.1119 | 369.2712 | 1077.5016 | 310.8486 | 1251.7729 | 175.3107 | 158.1356 | 177.7492 | 175.3107 | 195.9842 | 475.3312 | 531.7571 | 195.5458 | 162.0678 | 316.53 | 447.8644 | 168.9085 | 96.1435 | 128.3943 | 213.8407 | 146.5899 | 152.1593 | 186.7413 | 377 | 111.2271 | 345.0836 | 339.7334 | 778.7729 | 350.8565 | 233.6893 | 308.1356 | 609.5504 | 325.3943 | 368.4542 | 732.3753 | 330.0079 | 417.4779
1260 | 15.8544 | 534.9038 | 1048.0286 | 653.5542 | 367.7812 | 322.8548 | 853.4458 | 544.7526 | 255.9346 | 229.3865 | 578.8119 | 830.3497 | 252.4376 | 487.6872 | 1038.9346 | 758.7075 | 402.6503 | 377.5992 | 355.6503 | 740.5624 | 292.9857 | 202.9714 | 766.2413 | 734.4172 | 227.1534 | 475.6872 | 769.8119 | 922.6564 | 622.6872 | 161.4233 | 216.6278 | 247 | 186.2188 | 184.8466 | 329.0593 | 324.0143 | 125.1023 | 542.6278 | 901.4684 | 652.9346 | 1141.4458 | 201.8098 | 541.6443 | 444.3804 | 167.1677 | 106.3068 | 203.0654 | 105.3211 | 145.4376 | 281.6278 | 788.178 | 829.2413 | 212.8037 | 169.1534 | 371.0225 | 340.1534 | 143.4601 | 258.6278 | 1095.4909 | 765.0511 | 253.6789 | 267.497 | 802.7894 | 846.2782 | 215.9489 | 1196.2925 | 739.1187 | 902.5256 | 376.1248 | 190.5624 | 489.6872 | 842.0737 | 307.227 | 254.3865 | 358.3661 | 491.0511 | 432.5256 | 204.7955 | 264.0593 | 419.1166 | 341.7955 | 209.8037 | 248.2925 | 573.3661 | 238.6278 | 287.9857 | 357.9346 | 977.6217 | 195.5767 | 123.4887 | 156.2699 | 215.8384 | 151.1391 | 306.591 | 362.8917 | 554.1309 | 504.5624 | 343.6421 | 1142.8895 | 313 | 1209.4541 | 170.409 | 131.088 | 172.3354 | 170.409 | 165.8691 | 400.4827 | 484.4684 | 204.8466 | 161.591 | 271.9939 | 435.5338 | 164.5399 | 87.4744 | 128.9346 | 214.6932 | 121.6278 | 151.7812 | 158.6789 | 378.4233 | 101.8323 | 360.9632 | 311.3211 | 765.3947 | 353.8323 | 221.9857 | 288.2045 | 572.4008 | 320.7158 | 341.1248 | 619.4684 | 365.9407 | 417.7669
1261 | 15.8654 | 594.5488 | 1058.7522 | 656.4512 | 341.0707 | 276.0176 | 771.2121 | 550.1506 | 237.7522 | 249.7876 | 497.6371 | 856.4336 | 246.4778 | 437.9114 | 1069.9117 | 795.9293 | 381.4336 | 356.9469 | 301.3361 | 711.4778 | 261.584 | 200.5222 | 689.6723 | 733.9382 | 205.5043 | 432.2389 | 649.2121 | 902.0352 | 584.5929 | 150.8318 | 187.7167 | 255.2744 | 189.0796 | 179.2389 | 288.7256 | 358.3364 | 119.5929 | 494.2476 | 825.4689 | 659.0886 | 1080.1414 | 193.6105 | 464.8318 | 415.5133 | 179.354 | 99.5133 | 196.9114 | 113.3806 | 146.7787 | 290.1947 | 640.4068 | 815.9469 | 195.2478 | 166.9735 | 321.7698 | 341.8673 | 138.3806 | 256 | 1005.2831 | 745.0176 | 236.3716 | 239.6991 | 691.5306 | 800.0442 | 215.0531 | 1182.4336 | 627.7874 | 855.2741 | 353.9735 | 180.4602 | 424.3361 | 835.0886 | 288.3275 | 266.5133 | 322.4336 | 497.3009 | 432.9735 | 204.1327 | 236.3806 | 438.1947 | 340.646 | 176.6725 | 237.8407 | 522.8318 | 230.646 | 248.7963 | 317.8494 | 869.2652 | 186.2124 | 134.9735 | 143.9645 | 254.8673 | 147.7346 | 299.9114 | 337.1417 | 573.9469 | 491.5398 | 357.8142 | 1194.4957 | 318.354 | 1080.6368 | 163.9469 | 119.9735 | 180.646 | 163.9469 | 156.7611 | 342.7432 | 425.5043 | 198.2654 | 153.938 | 215.9114 | 393.292 | 170.1062 | 88.4867 | 113.1858 | 195.6901 | 122.4071 | 140.1593 | 150.1327 | 350.7963 | 94.5133 | 344.2565 | 316.9469 | 734.8318 | 352.8407 | 218.3275 | 281.5664 | 591.1062 | 295.7787 | 319.4602 | 560.5043 | 389.7256 | 403.4602
1262 | 15.8764 | 595.0334 | 1112.3754 | 683.9117 | 310.0177 | 257.4971 | 641.1256 | 613.4402 | 220.5147 | 243.5284 | 434.5755 | 897.4245 | 229.4971 | 409.0255 | 1132.9588 | 834.9666 | 371.0255 | 337.0432 | 268.501 | 651.5814 | 251.9745 | 203.5304 | 590.0687 | 777.4951 | 209.4539 | 394.5559 | 561.5422 | 861.501 | 544.5716 | 136.5029 | 149.0216 | 263.501 | 181.5088 | 176.4872 | 257.0451 | 360.5461 | 108.0157 | 435.0216 | 769.5363 | 699.4519 | 953.6619 | 178.9823 | 426.55 | 408.5128 | 168.0334 | 89.0196 | 181 | 115.5049 | 132.5069 | 274.0549 | 547.0765 | 808.0196 | 198.5128 | 162.5108 | 288.0078 | 337.5186 | 144.4971 | 249.5128 | 888.1707 | 665.1177 | 221.002 | 220.5226 | 560.603 | 727.102 | 206.0137 | 1190.4892 | 530.5932 | 804.5029 | 362.4853 | 185.9922 | 364.0549 | 812.5755 | 295 | 248.0353 | 295.5579 | 502.0059 | 421.0255 | 201.4951 | 240.002 | 440.0196 | 325.5186 | 180.9784 | 220.0471 | 478.0628 | 212.0255 | 228.9784 | 280.0098 | 802.4971 | 180.4951 | 136.998 | 132.9863 | 240.5383 | 159.4971 | 279.5088 | 338.0255 | 585.4814 | 479.0216 | 327.0746 | 1177.0726 | 328.4912 | 927.1393 | 160.5108 | 123.4951 | 172.5049 | 160.5108 | 169.9922 | 289.0491 | 400.5108 | 183.5088 | 140.9922 | 184.5304 | 369.5245 | 165.002 | 92.4931 | 116.4794 | 178.4814 | 123.5049 | 128.0118 | 136.0177 | 318.0039 | 93.501 | 312.5049 | 308.0216 | 713.5167 | 356.5049 | 233.9823 | 260.5363 | 563.0471 | 289.9902 | 315.5108 | 487.605 | 378.5049 | 386.5363
1263 | 15.8874 | 628.6197 | 1163.2172 | 718.7738 | 301.5113 | 251.3303 | 539.1631 | 603.0952 | 221.6923 | 223.3756 | 352.0273 | 927.819 | 217.1946 | 376.5702 | 1112.0726 | 880.6334 | 352.8869 | 302.7286 | 238.344 | 614.0905 | 256.819 | 175.2172 | 531.991 | 780 | 208.9684 | 355.2625 | 495.516 | 879.9185 | 487.0363 | 142.1583 | 149.2488 | 284.9863 | 173.9321 | 174.819 | 225.3077 | 354.3845 | 101.5339 | 403.0363 | 695.7785 | 713.2625 | 842.8779 | 177.2625 | 362.6517 | 405.0679 | 152.0226 | 80.5339 | 182.0226 | 112.4887 | 129 | 236.2851 | 492.1494 | 781.1268 | 205.2941 | 155.4661 | 253.3214 | 329.5339 | 148.5566 | 234.3077 | 724.8148 | 608.0679 | 212.8417 | 210.5339 | 457.8916 | 675 | 194.3982 | 1112.6565 | 438.516 | 773.8553 | 338.2988 | 188.4661 | 312.991 | 765.819 | 288.353 | 225.3982 | 266 | 508.7149 | 408.5113 | 202.9774 | 234.3982 | 457.0994 | 299.1268 | 190.4661 | 208.7828 | 443.9548 | 201.5566 | 234.8869 | 268.8643 | 774.344 | 190.6697 | 133.9095 | 130.7964 | 239.9185 | 135.4345 | 283.181 | 336.7601 | 612.8959 | 461.8643 | 279.7964 | 1099.0952 | 346.8054 | 818.1631 | 150.3982 | 112.1946 | 173.0679 | 150.3982 | 166.8417 | 238.9458 | 386.819 | 168.2625 | 131.7059 | 166.9548 | 355.4661 | 161.4434 | 82.7059 | 110.6381 | 182.3756 | 121.5113 | 132.2262 | 120.8643 | 316.5113 | 85.8417 | 302.8417 | 301.0905 | 690.6833 | 340.7059 | 229.7059 | 250.6923 | 549.7375 | 285.7964 | 325.8506 | 407.4119 | 397.475 | 380.2714
1264 | 15.8984 | 720.9825 | 1124.2963 | 749.3173 | 283.1504 | 227.7684 | 468.9659 | 603.2701 | 224.764 | 218 | 292.1976 | 924.1888 | 204.5236 | 324.4895 | 1047.3907 | 887.5796 | 353.7596 | 273.1976 | 215.7596 | 608.0516 | 246.382 | 161.4292 | 479.6311 | 790.472 | 191.236 | 318.2963 | 434.674 | 870.2491 | 473.8068 | 159.9956 | 151.0988 | 307.5708 | 160.0044 | 166.4764 | 206.528 | 366.2876 | 127.6093 | 369.3864 | 589.0691 | 723.944 | 772.0175 | 170.6652 | 299.2963 | 417.4248 | 152.4764 | 73.0988 | 161.0088 | 102.0516 | 138.4248 | 233.2832 | 422.5455 | 770.236 | 198.6268 | 149.8112 | 223.4764 | 330.4764 | 139.4808 | 218.6696 | 622.1547 | 583.7727 | 212.2832 | 211.4764 | 397.4336 | 624.2107 | 188.4292 | 993.2063 | 397.0472 | 692.5927 | 295.9572 | 191.7124 | 276.8628 | 764.8068 | 269.4336 | 216.2876 | 250.8156 | 528.472 | 396.4336 | 199.382 | 211.674 | 490.3304 | 281.9528 | 179.5752 | 210.0044 | 419.4852 | 208.1888 | 234.7124 | 246.7684 | 716.6783 | 177.5796 | 130 | 120.4292 | 254.3348 | 137.1801 | 293.0944 | 318.1547 | 645.708 | 453.9056 | 276.236 | 1003.9747 | 373.6136 | 722.8331 | 162.7552 | 103.1888 | 171.2876 | 162.7552 | 159.4764 | 207.6696 | 365.91 | 162.1888 | 114.2876 | 155.0516 | 357.6652 | 158.4764 | 82.0428 | 95.5236 | 162.8628 | 126.7124 | 126.8156 | 124.4248 | 285.5839 | 93.6608 | 302.8068 | 295.0516 | 646.6311 | 328 | 200.2448 | 260.0472 | 537.4852 | 283.8068 | 333.146 | 350.5839 | 423.236 | 365.8199
1265 | 15.9094 | 788.3305 | 1080.172 | 780.6411 | 262.2488 | 211.3923 | 401.5593 | 599.8799 | 217.8564 | 207.2822 | 256.172 | 922.6411 | 221.0768 | 319.1869 | 961.7647 | 875.4307 | 324.167 | 246.818 | 208.6745 | 575.9566 | 250.4307 | 170.7178 | 431.1002 | 824.1151 | 182.0668 | 321.8664 | 387.4541 | 881.9766 | 454.2772 | 151.2438 | 150.5025 | 308.4641 | 148.9282 | 164.9282 | 193.2488 | 368.4307 | 119.99 | 350.5693 | 524.995 | 708.7029 | 650.0085 | 177.2154 | 283.0384 | 404.0284 | 148.7846 | 75.182 | 132.4257 | 103.182 | 137.8898 | 241.6795 | 383.7178 | 767.4975 | 185.823 | 150.2871 | 205.8515 | 324.1052 | 140.7896 | 207.177 | 549.9851 | 541.3156 | 214.7846 | 209.9282 | 335.0902 | 559.2438 | 187.5359 | 888.9417 | 382.4591 | 650.4307 | 290.8948 | 186.8898 | 269.8948 | 744.7413 | 278.3639 | 222.7178 | 240.7512 | 546.0384 | 401.6127 | 215.7562 | 192.3205 | 477.172 | 282.6077 | 171 | 187.1386 | 386.1386 | 195.3873 | 227.7463 | 238.4307 | 677.6745 | 146.1386 | 126.2488 | 123.823 | 249.3923 | 144.9232 | 261.2388 | 295.823 | 640.7079 | 435.3873 | 291.182 | 936.3873 | 381.177 | 630.9517 | 161.9232 | 94.1386 | 157.3539 | 161.9232 | 141.8515 | 203.6795 | 345.4257 | 161.177 | 121.2537 | 143.8564 | 332.0618 | 148.3539 | 87.1052 | 93.8564 | 150.5359 | 136.8948 | 123.7178 | 123.3539 | 252.177 | 91.4591 | 309 | 270.995 | 617.3923 | 338.7178 | 192.5025 | 242.2438 | 517.5359 | 292.6795 | 323.8564 | 304.8515 | 423.177 | 338.2488
1266 | 15.9204 | 776.3681 | 998.3144 | 765.5845 | 239.8136 | 202.3254 | 374.3855 | 610.951 | 199.0063 | 200.1927 | 235.1627 | 879.6272 | 231.1627 | 318.5245 | 855.4708 | 873.3254 | 305.5118 | 223.1991 | 193.0664 | 525.8499 | 234.6209 | 182.7409 | 426.7046 | 865.2828 | 186.4455 | 335.0363 | 368.1264 | 859.7299 | 417.1991 | 122.94 | 157.5482 | 276.7536 | 154.5782 | 161.8073 | 182.3254 | 368.5182 | 109.8973 | 333.4882 | 463.2117 | 744.2338 | 555.9763 | 174.5182 | 252.1754 | 380.0664 | 152.5782 | 81.8073 | 126.0964 | 119.1264 | 121.2291 | 237.4218 | 329.9589 | 762.0964 | 182.8736 | 151.2591 | 193.7409 | 320.0964 | 146.0664 | 209.5782 | 451.248 | 530.9336 | 208.7109 | 202.4218 | 294.8373 | 542.4518 | 181.4218 | 743.9352 | 327.8863 | 646.1327 | 289.97 | 175.1627 | 248.139 | 682.5308 | 251.5972 | 226.5182 | 245.0964 | 535.4581 | 389.6872 | 200.2054 | 179.0363 | 457.8073 | 270.2954 | 174.2891 | 181.4818 | 373.3555 | 193.06 | 217.4518 | 263.1864 | 618.2117 | 149.2528 | 121.9036 | 125.2591 | 264.9937 | 116.7472 | 262.6983 | 298.9036 | 632.2227 | 418.2591 | 300 | 836.4834 | 374.2591 | 552.5008 | 150.1927 | 90.1264 | 153.3855 | 150.1927 | 124.2591 | 187.91 | 333.6145 | 146.5845 | 135.9336 | 135.4218 | 307.5482 | 147.1264 | 91.8673 | 89.8073 | 147.7109 | 124.4581 | 123.6809 | 105.1327 | 246.3555 | 80.1927 | 296.94 | 282.1201 | 574.8863 | 333.7472 | 194.0664 | 234.2227 | 494.4281 | 287.3254 | 329.1264 | 274.1027 | 437.09 | 360.2164
1267 | 15.9314 | 739.7138 | 915.8553 | 752.5162 | 221.1976 | 199.3628 | 342.1296 | 602.0411 | 195.649 | 205.3628 | 224.1534 | 849.6814 | 213.4277 | 300.5605 | 770.9644 | 879.3304 | 290.3391 | 226.1328 | 199.649 | 494.8348 | 215.3186 | 172.6696 | 398.2063 | 912.8261 | 190.4719 | 306.3834 | 347.0972 | 816.1209 | 395.4719 | 128.6933 | 151.2743 | 279.5842 | 162.2419 | 155.5162 | 177.6814 | 360.0767 | 112.351 | 326.8024 | 418.5162 | 760.4482 | 509.5368 | 159.351 | 209.7905 | 371.5162 | 152.9557 | 67.1091 | 116.9115 | 129.4838 | 114.5605 | 239.8466 | 308.2656 | 719.283 | 197.1771 | 156.2862 | 184.7905 | 329.9676 | 135.2743 | 205.4719 | 370.6134 | 508.0972 | 207.6814 | 193.0767 | 287.9115 | 516.2182 | 167.0324 | 624.7992 | 282.23 | 612.7787 | 308.5399 | 179.8466 | 231.0443 | 657.2862 | 252.5518 | 223.1209 | 245.4395 | 528.2862 | 357.351 | 179.6047 | 171.1209 | 463.8466 | 248.9115 | 169.1534 | 182.0767 | 336.6901 | 203.5605 | 209.7138 | 256.581 | 547.7344 | 154.2743 | 126.0443 | 123 | 262.7462 | 125.7375 | 280.7138 | 297.4395 | 653.5723 | 429.4514 | 288.7905 | 723.404 | 366.9557 | 494.8229 | 164.3304 | 88.1534 | 149.1534 | 164.3304 | 132.649 | 184.77 | 295.2506 | 142.4838 | 133.8348 | 141.77 | 296.7905 | 148.5162 | 83.7462 | 81.2743 | 130.9881 | 106.6372 | 117.1209 | 108.77 | 237.1534 | 73.0324 | 289.8024 | 267.8111 | 560.6166 | 322 | 194.4838 | 232.5929 | 466.5929 | 283.2419 | 338.3628 | 273.3304 | 453.1209 | 378.7581
1268 | 15.9424 | 739.7277 | 804.7941 | 799.4033 | 227.5915 | 207.1549 | 283.6434 | 588.9106 | 206.8638 | 198.263 | 208.8358 | 781.1226 | 197.9906 | 279.8077 | 664.1133 | 883.8451 | 275.0094 | 230.5541 | 199.4085 | 473.9626 | 222.5915 | 160.1362 | 413.394 | 897.4792 | 184.7183 | 284.1268 | 339.7464 | 811.8451 | 399.4553 | 146.1549 | 143.1362 | 280.2349 | 166.8638 | 156.5821 | 179 | 377.0468 | 102.2817 | 333.0094 | 401.2536 | 751.6196 | 477.6902 | 147.5634 | 201 | 370.8638 | 146.1362 | 76.474 | 96.3992 | 128.9906 | 124.737 | 243.1362 | 333.1455 | 679.2723 | 184.9345 | 160.2817 | 176.5728 | 318.0987 | 139.737 | 208.8825 | 330.5353 | 451.4886 | 204.9906 | 176.8264 | 260.526 | 490.2723 | 165.1549 | 541.1788 | 282.737 | 568.9439 | 296.0707 | 179.1268 | 197.1975 | 636.08 | 270.5447 | 217.1268 | 238.6996 | 544.3098 | 357.0187 | 174.8358 | 179.4459 | 474.0094 | 259.9013 | 154.9813 | 188.737 | 319.3098 | 202.8545 | 203.4272 | 235.5634 | 495.9345 | 134.6809 | 116.2536 | 118.9906 | 240.8358 | 132.263 | 279.5821 | 287.263 | 699.9387 | 447.4459 | 288.0187 | 612.6767 | 367.5821 | 469.8358 | 171.1362 | 80.8545 | 140.7089 | 171.1362 | 141.5728 | 173.3804 | 263.4179 | 139.1268 | 126.1362 | 147.5634 | 286.8545 | 131.2536 | 71.5728 | 81.7277 | 125.7277 | 106.2911 | 106.5447 | 108.8358 | 226.9906 | 74.5915 | 298.3004 | 249.5915 | 557.0987 | 336.3191 | 179.0987 | 224.8545 | 455.4179 | 303.3285 | 336.4179 | 276.1268 | 476.3378 | 379.2911
1269 | 15.9534 | 710.0667 | 719.1081 | 793.1208 | 215.2782 | 188.7679 | 252.8092 | 614.7758 | 189.6931 | 186.9046 | 196.7345 | 733.0954 | 202.0207 | 267.5103 | 559.8553 | 874.7345 | 273.9046 | 227.9253 | 195.3402 | 447.9586 | 223.7345 | 155.6598 | 408.2162 | 885.5517 | 183.6598 | 297.7218 | 335.9586 | 761.1955 | 381.0874 | 142.3943 | 132.8092 | 282.1367 | 153.2035 | 152.3943 | 165.5437 | 362.7472 | 102.7552 | 312.0127 | 368.1828 | 751.7885 | 440.7472 | 152.0207 | 189.2989 | 363.6391 | 154.531 | 86.3195 | 88.1701 | 128.9253 | 134.9253 | 235.1494 | 307.0874 | 638.7265 | 169.2655 | 148.469 | 186.946 | 306.3402 | 138.8092 | 214.6598 | 308.469 | 415.6391 | 189.1288 | 178.1161 | 233.1288 | 427.4944 | 169.5851 | 435.0381 | 258.992 | 542.2242 | 280.6057 | 176.9253 | 182.4563 | 638.8172 | 253.2242 | 215.5103 | 226.9793 | 533.2782 | 376.4563 | 179.7011 | 183.8299 | 464.7138 | 289.7965 | 156.0207 | 184.8839 | 307.6931 | 187.3736 | 219.3816 | 227.1494 | 453.938 | 136.8712 | 112.4356 | 122.4356 | 249.2115 | 124.4149 | 284.1701 | 289.946 | 710.2035 | 442.469 | 281.7138 | 543.2989 | 390.8919 | 448.9586 | 159.054 | 78.8299 | 137.8299 | 159.054 | 142 | 159.7552 | 251.8092 | 146.2862 | 129.8506 | 129.2035 | 252.6518 | 128.0207 | 86.0414 | 87.7552 | 124.7345 | 102.7345 | 101.5103 | 106.0954 | 214.6391 | 79.2448 | 288.3736 | 236.6931 | 520.7679 | 345.2448 | 170.8506 | 219.9046 | 444.9793 | 311.7345 | 310.1828 | 254.0334 | 509.3816 | 365.7885
1270 | 15.9644 | 654.9379 | 647.6546 | 787.6859 | 199.0133 | 181.3629 | 244.6106 | 615.0265 | 180.1814 | 198.9337 | 182.8451 | 670.8891 | 195.0531 | 252.0796 | 492.1459 | 829.1857 | 257.8584 | 221.0398 | 177.2875 | 440.9867 | 222.9867 | 161.7655 | 370.4424 | 940.4117 | 179.0133 | 288.3008 | 302.1061 | 658.1104 | 348.8584 | 132.8186 | 128.1947 | 274.0928 | 138.4159 | 154.7655 | 150.6239 | 374.2568 | 107.5841 | 289.6239 | 346.0265 | 740.4027 | 412.8318 | 168.9469 | 167.8584 | 360.5841 | 160.8053 | 77.0265 | 86.6106 | 128.6106 | 126.2477 | 217.261 | 278.4424 | 605.8318 | 177.1814 | 160.7125 | 183.2477 | 294.8584 | 138.3761 | 202.8451 | 279.4955 | 409 | 179.4027 | 175.8451 | 222.8053 | 372.8584 | 163.4292 | 382.8053 | 232.2212 | 533.0133 | 294.9602 | 177.8053 | 198.5708 | 609.6016 | 244.0133 | 227.5576 | 225.5841 | 519.4027 | 375.8451 | 182.6239 | 181.208 | 447.0398 | 291.4424 | 158.0133 | 162.6637 | 299.9735 | 196.7125 | 195.1592 | 216.4292 | 433.6239 | 153.1682 | 114.6106 | 122.2212 | 265.3761 | 132.9602 | 270.0663 | 295.208 | 722.8939 | 432.0133 | 264.6371 | 453.7607 | 414.5576 | 446.7655 | 153.1947 | 82.7788 | 129.2345 | 153.1947 | 142.5973 | 152.0398 | 240.6239 | 144.2477 | 121.4557 | 120.3894 | 240.9469 | 130.6106 | 84.6504 | 80.6371 | 120.4027 | 103.1814 | 120.1284 | 104.2212 | 215.7655 | 78.5973 | 276.208 | 237.9337 | 483.4955 | 332.0531 | 188.739 | 219.9867 | 427.4557 | 299.0398 | 310.1284 | 226.469 | 528.1682 | 359.7788
1271 | 15.9754 | 591.3253 | 563.2771 | 819.0217 | 195.7807 | 186.3904 | 232.0265 | 587.8337 | 171.4169 | 184.0048 | 183.3157 | 578.2289 | 177.2458 | 262.5566 | 427.3253 | 773.5663 | 255.0964 | 227.8024 | 176.8024 | 367.4048 | 232.3157 | 171.8771 | 354.8554 | 951.3205 | 183.706 | 279.0964 | 273.1976 | 692.0819 | 346.0964 | 142.1928 | 116.1976 | 261.8289 | 156.1181 | 157.5614 | 136.6361 | 355.1542 | 104.5133 | 276.8554 | 317.6145 | 725.9783 | 397.0265 | 166.0265 | 155.9518 | 358.7325 | 141.1012 | 82.7542 | 91.0964 | 117.8554 | 126.3157 | 244.5783 | 269.5614 | 591.8554 | 166.588 | 164.2458 | 177.8289 | 282.3422 | 137.7325 | 197.2193 | 269.8771 | 408.3904 | 183.2675 | 183.6313 | 215.9036 | 353.6361 | 162.0482 | 338.106 | 225.3422 | 527.9518 | 290.6361 | 165.4169 | 181.053 | 563.759 | 234.6843 | 213.8819 | 245.6795 | 493.3952 | 359.2458 | 178.3904 | 187.9253 | 433.6843 | 280.7325 | 138.9301 | 159.706 | 261.3253 | 182.5663 | 196.6048 | 219.9253 | 422.6843 | 141.5398 | 114.2193 | 113.5133 | 270.2193 | 135.9518 | 282.5566 | 281.1976 | 718.8819 | 438.5349 | 249.8554 | 375.2241 | 435.6313 | 439.1976 | 139.3687 | 82.9518 | 140.4602 | 139.3687 | 135.6843 | 140.5133 | 213.8337 | 136.3904 | 110.5614 | 125.6578 | 245.9518 | 116.1976 | 78.2193 | 81.0964 | 123.0482 | 104.7807 | 124.9036 | 104.6578 | 191.7374 | 84.4867 | 274.3904 | 263.241 | 451.9301 | 308.759 | 178.4916 | 196.3952 | 402.5398 | 309.4602 | 331.3639 | 228.1928 | 502.5181 | 369.706
1272 | 15.9864 | 524.1965 | 492.3307 | 787.5746 | 211.1702 | 186.6219 | 236.1947 | 538.7939 | 172.707 | 168.6219 | 163.7447 | 525.8544 | 157.3176 | 272.0246 | 357.7088 | 731.1228 | 257.1342 | 224.1833 | 178.1588 | 339.5237 | 223.3176 | 160.6956 | 379.4745 | 949.9509 | 190.4877 | 287.3535 | 283.036 | 683.6728 | 361.8167 | 133.2079 | 126.036 | 248.0737 | 142.257 | 155.3781 | 149.2798 | 343.9263 | 107.2193 | 281.5728 | 300.1342 | 717 | 390.6219 | 159.6219 | 175.7675 | 322.4158 | 122.7807 | 75.9395 | 90.6465 | 116.3535 | 135.3535 | 227.5746 | 246.8544 | 587.8658 | 164.2193 | 147.4272 | 184.5974 | 271.293 | 125.6711 | 212.9263 | 244.3912 | 376.9035 | 200.3044 | 175.586 | 196.4518 | 322.7447 | 167.7316 | 300.0491 | 244.7675 | 511.0737 | 272.1833 | 166.0851 | 184.414 | 544.0491 | 234.3535 | 205.3535 | 228.3912 | 485.707 | 345.5368 | 189.1947 | 197.9754 | 432.7316 | 283.5974 | 162.8281 | 158.4026 | 219.6956 | 177.4386 | 191.2325 | 221.8904 | 395.6105 | 143.4386 | 106.293 | 110.6219 | 262.9149 | 127.1588 | 289.5368 | 256.8298 | 699.7807 | 429.0737 | 258.3044 | 341.964 | 455.5728 | 417.9395 | 121.293 | 72.9149 | 127.9886 | 121.293 | 126.6465 | 135.7561 | 200.8658 | 137.2439 | 112.1096 | 123.6465 | 237.7807 | 118.5728 | 67.1833 | 80.6465 | 115.6711 | 99.0246 | 114.1588 | 107 | 179.2193 | 81.1588 | 252.8544 | 248.7447 | 429.8053 | 295.8904 | 179.0605 | 199.9017 | 387.0246 | 301.964 | 334.8904 | 242.8412 | 497.9263 | 361.5614
1273 | 15.9974 | 492.3894 | 416.2624 | 749.6814 | 212.7551 | 174.9498 | 230.3156 | 533.8555 | 172.292 | 165.8289 | 152.8053 | 458.457 | 159.1475 | 270.9027 | 295.8494 | 684.289 | 229.3627 | 220.1711 | 173.3658 | 333.6076 | 213.7316 | 163.416 | 366.9232 | 970.5871 | 183.1209 | 291.2684 | 288.292 | 628.6076 | 346.8023 | 134.416 | 118.6076 | 273.2479 | 152.003 | 160.0738 | 161 | 333.3392 | 108.4631 | 289.9027 | 290.7551 | 685.289 | 368.1681 | 157.4631 | 180.1209 | 317.2213 | 127.2449 | 70.5369 | 93.0738 | 114.5605 | 133.5605 | 223.2951 | 226.3894 | 559.8259 | 162.292 | 147.9764 | 154.3863 | 268.5369 | 130.1475 | 198.5338 | 238.416 | 373.8555 | 198.8525 | 168.1711 | 176.8525 | 302.9262 | 159.8525 | 274.9734 | 245.3156 | 459.5072 | 267.5369 | 162.7551 | 186.1209 | 517.0707 | 245.2449 | 214.3422 | 209.0973 | 470.0707 | 335.292 | 189.6578 | 195.9262 | 414.7049 | 273.6814 | 170.6814 | 152.4631 | 223.0502 | 160.2654 | 200.1977 | 227.6106 | 392.1475 | 153.5369 | 109.2449 | 111.6342 | 252.292 | 133.7818 | 248.2152 | 257.7818 | 704.2449 | 418.7316 | 271.4395 | 316.4129 | 481.6608 | 420.7818 | 146.4426 | 80.0502 | 129.2213 | 146.4426 | 127.1711 | 126.7551 | 186.1445 | 120.2418 | 120.3422 | 143.8319 | 225.1209 | 112.3156 | 72.0502 | 84.3422 | 119.5133 | 104.8791 | 107.4631 | 107.6342 | 171.584 | 80.8053 | 251.416 | 277.7612 | 423.6342 | 306.0502 | 187.6342 | 229.124 | 345.3125 | 294.1711 | 308.8996 | 240.0236 | 553.9323 | 361.6106
1274 | 16.0084 | 416.2372 | 382.4744 | 747.8186 | 210.586 | 180.9302 | 207.4837 | 511.3209 | 174.1721 | 155.5953 | 151.7674 | 386.9163 | 181.1628 | 256.4837 | 252.3721 | 651.7767 | 230.8093 | 207.1302 | 177.5256 | 317.1814 | 212.3535 | 176.4651 | 352.2326 | 935.6837 | 199.3349 | 287.4744 | 277.8884 | 592.786 | 354.6884 | 127.4233 | 115.7581 | 246.4512 | 162.4744 | 163.6465 | 172.6372 | 323.293 | 107 | 274.1907 | 286.6465 | 704.4977 | 355 | 152.7674 | 179.5256 | 329.8791 | 135.8791 | 72.6465 | 90.1814 | 119.1116 | 109.6651 | 218.1907 | 249.1535 | 521.0186 | 175.1628 | 137.7767 | 147.2837 | 256.4233 | 120.4837 | 174.6558 | 241.1209 | 361.6651 | 185.8884 | 163.5349 | 178.1116 | 295.4744 | 143.6558 | 271.7581 | 222.4837 | 431.7674 | 269 | 160.586 | 166.7767 | 482.3721 | 231.2512 | 219.293 | 210.586 | 458 | 344.2837 | 173.0698 | 186.5349 | 393.9488 | 255.0093 | 153.9488 | 154.8791 | 223.5349 | 166.0419 | 202.0093 | 221.0093 | 365.6744 | 142.0698 | 97.1907 | 122.3442 | 238.6558 | 138.0605 | 232.8186 | 270.4651 | 701.8884 | 385.6744 | 264.9488 | 288.6047 | 493.6465 | 399.2 | 149.1302 | 82.4744 | 115.3721 | 149.1302 | 130.9395 | 132.9907 | 172.4744 | 117.1116 | 126.586 | 141.4837 | 204.4837 | 113.4047 | 80.293 | 70.5442 | 106.2512 | 104.1814 | 108.586 | 102.8279 | 159.1814 | 80.414 | 278.0419 | 274.5535 | 397.493 | 322.6977 | 191.8791 | 253.2233 | 345.6279 | 282.4233 | 296.9395 | 226.9488 | 560.3116 | 350.4837
1275 | 16.0194 | 370.412 | 335.8602 | 678.5317 | 182.3539 | 168.8947 | 175.2831 | 478.8475 | 170.412 | 165.4937 | 157.9056 | 338.6479 | 171.5535 | 234.8239 | 254.4937 | 615.0127 | 236.706 | 202.9528 | 159.5771 | 292.9183 | 200.1416 | 168.8004 | 361.9056 | 957.4574 | 188.6007 | 284.3412 | 279.2704 | 545.4011 | 356.7768 | 127.9056 | 127.9056 | 246.6933 | 140.2359 | 148.8475 | 179.6588 | 333.882 | 94.4827 | 276.8584 | 306.1053 | 663.0725 | 357.6352 | 141.118 | 167.5063 | 318.8239 | 134.0472 | 71.6824 | 88.3176 | 125.6352 | 123.0345 | 197.1416 | 230.4011 | 491.2123 | 185.3176 | 143.8349 | 167.1761 | 246.3648 | 129.1289 | 172.9528 | 208.6951 | 315.7423 | 174.7532 | 171.8584 | 174.0944 | 289.0472 | 143.9292 | 264.8004 | 226.5173 | 393.107 | 266.3648 | 169.2468 | 179.4229 | 414.6606 | 230.2232 | 207.4827 | 211.3412 | 434.2831 | 336.5063 | 176.5408 | 197.4937 | 380.4356 | 239.7768 | 139.118 | 164.9056 | 206.8239 | 186.2232 | 186.7768 | 206.4356 | 328.2359 | 138.9528 | 107.7641 | 120.7532 | 248.8113 | 138.3176 | 237.9764 | 268.0708 | 663.7423 | 384.4701 | 300.187 | 303.6933 | 500.588 | 399.8113 | 127.1651 | 78.6824 | 108.2704 | 127.1651 | 123.4356 | 133.7296 | 168.6824 | 126.2704 | 124.706 | 121.8004 | 189.412 | 114.706 | 86.9292 | 70.2232 | 92.0472 | 112.1996 | 94.8475 | 105.2704 | 161.2704 | 84.9292 | 268.5771 | 220.1306 | 346.107 | 314.8239 | 179.5063 | 224.107 | 321.7659 | 284.882 | 313.1525 | 249.0109 | 583.9165 | 354.5173
1276 | 16.0304 | 332.1031 | 309.6445 | 651.7375 | 182.4352 | 169.0664 | 185.1461 | 462.9469 | 166.3289 | 173.6711 | 163.3422 | 305.5383 | 186.1594 | 227.3289 | 221.7344 | 560.4055 | 271.2391 | 176.143 | 165.1063 | 264.5516 | 210.1063 | 172.3953 | 331.1031 | 932.6812 | 194.4484 | 268.5648 | 265.2227 | 519.2758 | 356.6977 | 141.3953 | 132.6711 | 233.1562 | 140.7375 | 150.3953 | 167.9203 | 313.4984 | 96.0531 | 280.3156 | 294.525 | 614.8937 | 337.525 | 132.6445 | 179.4617 | 290.525 | 140.0531 | 73.6844 | 75.5781 | 138.4086 | 121.907 | 223.8836 | 220.382 | 461.8672 | 181.9734 | 147.6445 | 169.9734 | 221.5117 | 125.2492 | 199.1594 | 221.1859 | 293.9734 | 160.9336 | 172.6312 | 179.3953 | 281.6312 | 154.382 | 211.3523 | 241.7242 | 383.3953 | 252.2492 | 173 | 191.3422 | 377.9469 | 222.9203 | 201.6711 | 196.907 | 418.6445 | 339.0664 | 195.4219 | 179.4984 | 336.4055 | 269.8969 | 141.382 | 171.6844 | 189.2625 | 189.6578 | 194.0797 | 210.7242 | 323.3688 | 134.2891 | 124.7109 | 125.0531 | 228.8141 | 126.2492 | 255.7773 | 249.5648 | 629.8937 | 417.1594 | 283.4187 | 292.1695 | 489.2359 | 377.8008 | 141.475 | 58.5383 | 114.3555 | 141.475 | 127.7242 | 134.3555 | 157.9336 | 120.9469 | 122.3289 | 111.9734 | 171.907 | 116.3555 | 89.3289 | 68.2891 | 98.0531 | 103.2359 | 101.093 | 96.5914 | 169.3688 | 92.0266 | 258.6711 | 230.1992 | 325.6578 | 315.382 | 192.8039 | 216.4086 | 311.7242 | 282.618 | 316.9734 | 248.2359 | 612.0664 | 345.5648
1277 | 16.0414 | 290.3981 | 302.5329 | 657 | 184.5329 | 177.4169 | 187.116 | 452.1662 | 178.3009 | 154.1818 | 146.232 | 248.8464 | 185.0157 | 242.7179 | 200.6332 | 499.9969 | 269.8652 | 203.0031 | 171.6332 | 243.0157 | 214.5831 | 155.1317 | 316.0502 | 922.6677 | 203 | 241.1818 | 259.0502 | 493.1317 | 341.232 | 131.6489 | 126.1662 | 213.8495 | 157.6175 | 147.4828 | 145.5988 | 288.0157 | 97.2665 | 246.1975 | 295.6175 | 576.2477 | 342.7179 | 139.884 | 164.3981 | 276.5831 | 129.6489 | 79.1003 | 61.4828 | 124.8652 | 125.9342 | 207.8809 | 250.652 | 417.1474 | 166.3323 | 143.9498 | 160.4828 | 214.7837 | 134.7179 | 194.5988 | 228.1662 | 294.7335 | 170.3511 | 152.9154 | 166.9154 | 241.4138 | 160.0502 | 213.2852 | 223.4483 | 368.1818 | 246.6834 | 155.9154 | 186.5329 | 412.9529 | 228.6175 | 217.0345 | 192.7335 | 360.9624 | 309.8308 | 178.7649 | 184.768 | 318.3668 | 265.8151 | 149.1003 | 148.3981 | 183.3166 | 165.7649 | 192.4828 | 193.1317 | 284.3135 | 115.9655 | 115.0157 | 124.8997 | 212.2665 | 132.3009 | 253.0658 | 231.0658 | 635.6677 | 415.3323 | 261.5831 | 271.116 | 506.6019 | 378.7179 | 139.0157 | 73.6019 | 113.2665 | 139.0157 | 122.4326 | 137.3668 | 150.9498 | 122.4671 | 111.7492 | 110.6834 | 154.0658 | 122.1003 | 80.116 | 77.9843 | 97.8997 | 97.3668 | 107.3166 | 106.0345 | 169.9498 | 76.232 | 239.8652 | 229.2821 | 331.8338 | 315.5831 | 175.6646 | 206.2821 | 278.4138 | 276.9498 | 323.2006 | 225.2821 | 609.4828 | 357.1348
1278 | 16.0524 | 269.2605 | 273.5643 | 620.1292 | 174.3476 | 156.7384 | 188.5654 | 436.4778 | 157.5643 | 152.6524 | 160.9573 | 240.9135 | 192.2178 | 217.9989 | 202.7827 | 478.913 | 238.7384 | 194.8686 | 176.5654 | 246.0438 | 210.2173 | 163.7832 | 312.1303 | 905.347 | 185.6081 | 239.6524 | 285.0443 | 437.0854 | 346.9135 | 124.3043 | 122.3043 | 210.3043 | 141.434 | 141.913 | 153.3049 | 312.6097 | 90.4346 | 275.5233 | 307.2611 | 559.2173 | 336.7822 | 141.2173 | 153 | 249.9557 | 131.3481 | 76.8259 | 67.0438 | 132.6962 | 128.913 | 210.3919 | 258.1303 | 391.2605 | 154.4346 | 131.8692 | 138.2168 | 199.6081 | 128.0865 | 192.5654 | 224.3043 | 289.0432 | 178.3914 | 145.6957 | 168.7395 | 231.6524 | 137.347 | 226.3914 | 237.3487 | 349.2605 | 249.7827 | 172.3487 | 172.1735 | 429.6086 | 251.3919 | 211.4778 | 200.9568 | 371.8708 | 330.1751 | 158.2605 | 170.0427 | 299.5211 | 232.347 | 147.5216 | 144.6524 | 185.087 | 151.5216 | 195.9568 | 197.6092 | 272.6524 | 122.9135 | 129.1746 | 114.6519 | 211 | 126.8692 | 255.6524 | 237.8265 | 580.7805 | 438.2184 | 247.4778 | 272.5654 | 552.7838 | 381.1303 | 146.9135 | 78.0432 | 100.8692 | 146.9135 | 117.3043 | 133.1303 | 134.6951 | 113.1735 | 114.6524 | 115.8697 | 146.2173 | 119.8259 | 66.2605 | 74.2605 | 94.6086 | 87.5649 | 95.1735 | 106.7389 | 144.6513 | 73.5654 | 251.1746 | 212.9562 | 325.9562 | 283.3903 | 181.7832 | 208.7395 | 264.4784 | 271.1303 | 311.6951 | 231.2178 | 599.7389 | 368.087
1279 | 16.0634 | 258.6283 | 276.8673 | 570.0176 | 202.8585 | 149.8319 | 163.3893 | 409.761 | 141.7522 | 147.5044 | 146.9291 | 247 | 171.0973 | 217.4514 | 213.2036 | 441.8937 | 237.4956 | 198.5753 | 143.6017 | 212.4778 | 204.7522 | 177.6637 | 327.6992 | 926.1505 | 178.708 | 245.1239 | 259.1857 | 419.1239 | 328.9291 | 114.7964 | 123.4159 | 217.7876 | 145.4514 | 133.2124 | 150.0885 | 278.5662 | 85.8761 | 278.7168 | 300.0885 | 540.3008 | 323.9203 | 156.9912 | 136.0088 | 231.2124 | 131.4602 | 78.5398 | 78.7876 | 130.0885 | 133.6637 | 227.9115 | 231.9291 | 390.5398 | 159.7876 | 134.0797 | 142.7434 | 194.8319 | 122 | 183.6725 | 192.1415 | 281.7522 | 157.053 | 149.5398 | 171.5841 | 225.7964 | 137.6195 | 232.6637 | 220.3893 | 322.3451 | 264.4514 | 175.9203 | 167 | 387.23 | 242.7168 | 196.7964 | 220.9912 | 344.814 | 338.2124 | 155.4159 | 180.531 | 271.8849 | 264.478 | 144.5841 | 141.6283 | 188.1239 | 162.0354 | 167.8495 | 217.4514 | 285.9115 | 131.1239 | 125.2566 | 109.5841 | 181.9734 | 122.708 | 268.2036 | 262.8231 | 558.2478 | 413.4778 | 237.7522 | 306.8585 | 564.4602 | 374.0442 | 138.8407 | 74.292 | 87.5044 | 138.8407 | 112.0442 | 140.9115 | 143.5753 | 124.2832 | 118 | 110.2124 | 156.3275 | 109.5044 | 67.6637 | 75.6637 | 89.7522 | 90.0797 | 99.2036 | 99.0442 | 136.1239 | 86.6195 | 232.3893 | 204.0442 | 319.1681 | 294.7788 | 176.5486 | 220.7876 | 261.0442 | 279.6195 | 288.0088 | 212.2212 | 604.0797 | 344.9291
1280 | 16.0744 | 224.3086 | 249.8681 | 539.4341 | 202.3151 | 161.8039 | 189.4534 | 431.9711 | 163.7685 | 148.3215 | 155.7267 | 230.4341 | 169.3634 | 219.3987 | 220.6013 | 417.6366 | 235.9582 | 189.1543 | 159.9711 | 204.2026 | 209.4823 | 172.0772 | 347.2444 | 895.7845 | 171.0772 | 235.1961 | 243.2797 | 432.2444 | 355.0129 | 134.0483 | 134.0836 | 199.3923 | 147.3987 | 136.4823 | 135.1961 | 273.4052 | 108.0483 | 277.0418 | 282.3151 | 513.553 | 318.8392 | 138.791 | 161.4116 | 248.8875 | 135.7621 | 92.2444 | 76.9582 | 127.4405 | 114.3923 | 222.6366 | 253.6914 | 344.463 | 167.3215 | 142.7621 | 142.2379 | 190.9582 | 130.6431 | 193.4052 | 209.2508 | 250.4694 | 178.2508 | 146.6785 | 173.1608 | 216.9582 | 142 | 216.9936 | 187.3923 | 328.1254 | 254.1543 | 169.3987 | 176.3634 | 341.9099 | 223.0354 | 205.2444 | 235.9228 | 313.3151 | 314.1125 | 158.881 | 181.0772 | 276.9646 | 269.0354 | 153.3634 | 131.0772 | 176.7556 | 159.7974 | 167.9646 | 170.4212 | 249.6656 | 138.4823 | 112.0772 | 117.6431 | 182.2444 | 133.0836 | 248.2315 | 267.3987 | 513.1832 | 399.4405 | 251.8457 | 283.9871 | 591.0901 | 390.7267 | 130.8392 | 76.881 | 89.7621 | 130.8392 | 102.7974 | 142.119 | 136.3151 | 137.4823 | 114.3987 | 102.6785 | 148.0354 | 113.9228 | 76.4823 | 70.7974 | 80.7974 | 93 | 97.2379 | 105.6431 | 132.6785 | 102.5241 | 233.2444 | 243.775 | 323.0418 | 328.7685 | 177.4823 | 219.6785 | 262.6013 | 251.5884 | 285.3215 | 230.0901 | 621.4052 | 341.4823
1281 | 16.0854 | 209.0698 | 230.4071 | 513.9537 | 180.1514 | 160.8722 | 197.4071 | 397.8494 | 172.2675 | 169.0463 | 159.3373 | 213.7443 | 176.6627 | 209.2094 | 198.5584 | 391.291 | 255.2439 | 191.058 | 172.7325 | 193.8141 | 200.2792 | 158.0118 | 326.3608 | 816.0008 | 174.5929 | 234.6627 | 248.1278 | 411.3608 | 331.64 | 123.2212 | 116.0235 | 180.7443 | 137.942 | 144.1278 | 139.058 | 248.2329 | 96.4886 | 274.6047 | 266.7443 | 475.2329 | 305.5467 | 129.7325 | 149.8259 | 226.2329 | 124.8141 | 88.942 | 81.5929 | 116.2792 | 103.0698 | 186.7678 | 222.7796 | 351.6392 | 161.6745 | 134.7443 | 154.651 | 170.6863 | 137.6627 | 179.9537 | 207.8141 | 235.3373 | 190 | 158.9184 | 161.5467 | 209.1396 | 144.9302 | 199.0118 | 181.9302 | 303.2329 | 248.7325 | 162.1396 | 169.7443 | 325.1396 | 226.058 | 205.6047 | 214.8259 | 302.8722 | 330.1741 | 159.2675 | 166.2792 | 235.1169 | 242.0235 | 159.1976 | 159.4996 | 139.7678 | 173.8486 | 182.5231 | 201.2784 | 229.6047 | 143.9302 | 119.9882 | 125.3953 | 178.942 | 147.2557 | 267.5694 | 258.6745 | 491.3373 | 405.8604 | 222.1051 | 243.6282 | 661.3365 | 390.6745 | 127.0698 | 79.4651 | 84.6745 | 127.0698 | 108.058 | 127.0816 | 131.7325 | 140.7325 | 103.4769 | 109.058 | 137.1396 | 121.3953 | 71.6745 | 82.651 | 81.6627 | 93 | 103.058 | 95.8141 | 122.942 | 82.0933 | 270.9647 | 232.1631 | 311.0816 | 298.4424 | 171.2094 | 224.5929 | 251.5467 | 283.6855 | 289.1976 | 230.0118 | 566.931 | 352.7906
1282 | 16.0964 | 223.6416 | 239.1726 | 483.1651 | 198.5797 | 159.7448 | 192.7655 | 358.6548 | 175.7242 | 164.0826 | 131.1858 | 192.1238 | 194.3865 | 220.152 | 188.5103 | 344.2683 | 254.0619 | 198.469 | 182.6829 | 191.2345 | 204.1932 | 152.5103 | 360.9456 | 761.7167 | 171.0413 | 218.8687 | 237.3377 | 396.0413 | 315.0206 | 139.09 | 125.8762 | 166.5722 | 131.2758 | 137.0619 | 144.2345 | 230.2964 | 86.7655 | 260.3377 | 270.4484 | 468.469 | 300.2552 | 169.4765 | 128.3377 | 204.5722 | 127.4484 | 78.5516 | 74.3171 | 116.469 | 131.0488 | 218.9456 | 210.7242 | 308.1164 | 167.1932 | 146.6416 | 145.848 | 172.9381 | 130.0619 | 178.9587 | 189.5929 | 242.9381 | 187.7655 | 154.3171 | 154.7655 | 186.8893 | 131.848 | 186.0619 | 213.5384 | 294.9794 | 260.1726 | 151.0619 | 176.4278 | 323.7448 | 239.4278 | 201.7655 | 198.5516 | 289.0826 | 350.1726 | 161.2345 | 173.1726 | 253.0075 | 239.9587 | 145.1032 | 158.3377 | 168.9662 | 169.5722 | 177.0619 | 200.6342 | 215.3377 | 147.2345 | 118.7861 | 125.5103 | 150.6754 | 143.5516 | 282.469 | 277.6004 | 440.8406 | 409.4897 | 248.4765 | 245.1726 | 671.0826 | 378.3171 | 130.469 | 74.7861 | 83.4897 | 130.469 | 117.7036 | 144.3452 | 133.4897 | 135.0413 | 127.5591 | 102.3171 | 137.9794 | 122.2552 | 81.6623 | 74.5929 | 88.9587 | 103.4278 | 108.2345 | 89.5103 | 132.6623 | 87.152 | 240.5441 | 212.3171 | 277.6961 | 311.5591 | 160.5516 | 223.2758 | 257.4278 | 300 | 286.2758 | 219.2964 | 526.379 | 354.5103
1283 | 16.1074 | 235.057 | 240.7286 | 496.9568 | 194.1288 | 168.3284 | 181.4002 | 344.943 | 155.0432 | 175.8998 | 157.5852 | 177.6716 | 171.9861 | 230.2999 | 172.8574 | 315.1002 | 247.2143 | 177.2861 | 169.3431 | 189.7286 | 210.0285 | 171.6854 | 350.2575 | 764.6284 | 156.8859 | 207.7001 | 251.1711 | 357.6577 | 311.7286 | 134.8859 | 142.5998 | 170.5713 | 145.8998 | 146.1141 | 124.5575 | 228.7571 | 81.4572 | 274.1711 | 274.5143 | 463.1858 | 282.586 | 143.6292 | 139.1426 | 189.643 | 130 | 72.2143 | 72.5143 | 121.7857 | 122.0717 | 196.9006 | 190.0432 | 323.3138 | 151.0717 | 155.0285 | 148.5713 | 173.7286 | 120.1858 | 181 | 206.1997 | 263.414 | 167.3146 | 146.4572 | 155.5143 | 189.0856 | 130.0285 | 217.828 | 189.172 | 286.1573 | 257.943 | 188.8851 | 166.3716 | 324 | 238.4572 | 186.6145 | 206.5998 | 297.8712 | 344.1573 | 168.057 | 189.8712 | 258.943 | 238.2143 | 132.4287 | 138.1002 | 164.8289 | 159.943 | 174 | 195.5143 | 225.3855 | 138.1573 | 110.1858 | 111.3716 | 157.8998 | 141.7571 | 287.0285 | 273.643 | 446.7139 | 426.6569 | 238.5289 | 286.8566 | 657.1573 | 388.6284 | 142.5998 | 88.8998 | 84 | 142.5998 | 121.5143 | 119.4434 | 115.0717 | 130.7286 | 108.9861 | 119.4425 | 139.7571 | 126.5428 | 79.943 | 70.7571 | 86.4572 | 113.057 | 104.4572 | 116.2567 | 117.3146 | 83.6716 | 239.6284 | 221.1141 | 277.8427 | 310.4002 | 162.5428 | 193.9861 | 234.5004 | 292.4287 | 300.8998 | 213.9715 | 548.0139 | 338.8574
1284 | 16.1184 | 205.4537 | 253.0802 | 469.376 | 196.9248 | 154.8421 | 185.6942 | 345.3083 | 153.386 | 181 | 148.995 | 179.6165 | 170.6942 | 234.3083 | 179.5413 | 309.2306 | 230.6116 | 174.6165 | 167.0777 | 175.9198 | 209.4612 | 175.6917 | 347.386 | 745.1479 | 169.9273 | 258.3207 | 260.8471 | 338.3058 | 321.7719 | 128.6917 | 113.6843 | 149.1479 | 131.7644 | 146.9223 | 134.1579 | 195.9148 | 90.0025 | 253.0702 | 294.2356 | 407.9099 | 277 | 135.6165 | 150.9248 | 201.3884 | 127.6917 | 71 | 57.6116 | 126.0777 | 119.8471 | 195.0025 | 218.3934 | 300.1454 | 161.1579 | 149.0752 | 140.2281 | 184.5413 | 131.8496 | 173.3058 | 189.1479 | 262.8446 | 188.6992 | 142.6917 | 168.3108 | 197.386 | 130.2306 | 195.1454 | 201.0827 | 292.2331 | 230.6091 | 158.1429 | 157.3835 | 322.4612 | 235.4612 | 197.3884 | 193.8421 | 278.9173 | 326.381 | 149.2256 | 198.6165 | 215.4512 | 233.9223 | 129.2306 | 151.4661 | 166.6942 | 156.4612 | 174 | 200.6165 | 221.5364 | 124.2281 | 109.5388 | 113.9248 | 159.9223 | 133.5364 | 265.6868 | 280.0025 | 382.4437 | 432 | 269.4711 | 258.99 | 725.5563 | 389.1529 | 126.7644 | 80.1504 | 88.6165 | 126.7644 | 122.7694 | 124.3884 | 123.619 | 119.2281 | 110.7719 | 124.4612 | 146.9248 | 124.1529 | 74.1529 | 77.9248 | 70.381 | 90.3785 | 95.3058 | 92.6843 | 132.5438 | 87.1554 | 244 | 221.9223 | 290.2331 | 313.1554 | 161.6917 | 202.6967 | 235.2331 | 268.4562 | 308.3083 | 230.6967 | 497.7545 | 342.4636
1285 | 16.1294 | 182.7109 | 231.9852 | 441.0206 | 189.6194 | 162.5074 | 180.9646 | 292.0618 | 173.7611 | 165.3657 | 137.528 | 176.3097 | 162.8377 | 220.1474 | 175.1829 | 272.2595 | 217.4012 | 186.944 | 151.584 | 172.7817 | 212.1269 | 164.056 | 293.4984 | 691.8789 | 167.1829 | 214.5898 | 250.2743 | 324.2743 | 307.8023 | 117.8377 | 112.5988 | 154.5074 | 138.5074 | 152.2537 | 143.6903 | 194.5988 | 84.4012 | 265.3245 | 265.6046 | 379.4926 | 291.8526 | 131.528 | 147.528 | 190.3657 | 127 | 65.528 | 74.888 | 116.056 | 112.4012 | 200.3451 | 236.8171 | 285.3097 | 163.6549 | 153.2537 | 142.472 | 161.4217 | 134.4366 | 165.528 | 186.6903 | 229.7315 | 173.5486 | 156.8526 | 168.0914 | 168.5132 | 137.0354 | 199.0708 | 220.5074 | 269.9852 | 254.2685 | 152.0354 | 172.416 | 300.8937 | 210.7669 | 190.2743 | 215.5783 | 225.8789 | 314.9646 | 173.4868 | 189.056 | 228.0148 | 250.9794 | 143.8526 | 151.528 | 169 | 147.4012 | 195.1063 | 217.6343 | 215.0914 | 118.6549 | 96.7109 | 110.528 | 159 | 136.472 | 242.584 | 284.5634 | 402.4308 | 415.584 | 230.1886 | 254.0354 | 747.7817 | 384.8731 | 127.2537 | 76.7817 | 86.8731 | 127.2537 | 138.6343 | 118.056 | 124.0914 | 121.472 | 96.8023 | 116.1829 | 140.4012 | 114.4012 | 66.7463 | 69.8377 | 81.6343 | 82.2183 | 83.6194 | 107.2331 | 117.112 | 86.6549 | 235.4012 | 224.9086 | 285.1829 | 308.7463 | 182.888 | 225.1977 | 234.0914 | 246.3657 | 301.9646 | 224.2743 | 462.8023 | 338.7463
1286 | 16.1404 | 190.1161 | 238.4981 | 430.442 | 192.558 | 148.5319 | 173.442 | 296.0561 | 158.3559 | 160.206 | 136.794 | 155.9439 | 153.648 | 212.824 | 186.4981 | 249.2959 | 203.8839 | 195.558 | 175.5842 | 170.618 | 181.2398 | 172.1161 | 316.9062 | 662.3259 | 152.2959 | 211.4981 | 277.9662 | 294.7979 | 291.8839 | 122.146 | 116.588 | 150.06 | 145.176 | 163.5281 | 151.352 | 175.5619 | 96.2921 | 259.0899 | 268.7041 | 345.8278 | 301.558 | 131.588 | 143.618 | 180.442 | 143.6741 | 71.146 | 72.266 | 108.236 | 109.206 | 191.4719 | 239 | 310.2021 | 163 | 152.618 | 128.1199 | 180.2021 | 124.4719 | 171.94 | 188.794 | 222.588 | 173.352 | 170.5281 | 159.854 | 195.7302 | 127.0899 | 220.4681 | 205.7379 | 258.236 | 258.236 | 145.266 | 171.442 | 308.4981 | 205.588 | 182.236 | 195.2098 | 238.4081 | 322.5281 | 175.648 | 173.2959 | 225.4719 | 234.5619 | 163.0861 | 150.794 | 148.3559 | 164.0561 | 199.412 | 192.6218 | 215.588 | 141.0261 | 109.6741 | 109.794 | 146.2959 | 149.1161 | 261.0261 | 263.5619 | 387.7979 | 418.94 | 249.3482 | 230.5919 | 764.6741 | 380.03 | 136.146 | 67.4719 | 70.9139 | 136.146 | 116.0039 | 114.206 | 144.4381 | 123 | 107.0861 | 105.266 | 139.588 | 114.382 | 84.8501 | 67.794 | 86.794 | 82 | 88.146 | 86.2098 | 107.824 | 87.588 | 244.1161 | 218.854 | 274.266 | 304.618 | 178.6779 | 222.06 | 226.648 | 284.0822 | 294.442 | 244.0261 | 424.6518 | 337.794
1287 | 16.1514 | 184.1308 | 244.4189 | 436.2567 | 184.3245 | 163.3511 | 176.0315 | 292.1308 | 158.6441 | 164.8378 | 135.3874 | 161.4818 | 152 | 215.2252 | 186.7748 | 261.3196 | 249.3777 | 206.6755 | 174.1308 | 180.4818 | 166.5496 | 151.6174 | 331.8378 | 633.0048 | 151.4189 | 221.4504 | 227.9467 | 307.3511 | 322.8644 | 125.6126 | 122.6441 | 143.1622 | 132.293 | 161.9685 | 151.3874 | 183.707 | 116.1259 | 268.0944 | 292.9637 | 358.9637 | 292.5182 | 147.3196 | 137.3559 | 158.0363 | 143.1622 | 73 | 64.3559 | 117.4818 | 129.9637 | 220.4455 | 240.6126 | 308.1308 | 163.8063 | 151.1937 | 140.1259 | 154.7482 | 128.4504 | 160.293 | 186.5811 | 220.5811 | 172.5811 | 168.1622 | 153.9685 | 180.8111 | 135.2881 | 214.5182 | 214.707 | 266.6755 | 246.5182 | 150.2567 | 172.4189 | 291.0363 | 206 | 193.0944 | 179.1308 | 233.7119 | 332.2567 | 167.5496 | 182.9007 | 215.7433 | 254.8015 | 173.4504 | 154.2252 | 146.2252 | 177.063 | 191.7433 | 180.9685 | 215.1937 | 132.4867 | 108.3559 | 122.9007 | 125.2615 | 145.5496 | 261.3559 | 263.6441 | 352.7797 | 391.9734 | 249.1308 | 277.2155 | 723.0412 | 365.293 | 133.9685 | 68.2252 | 69.4189 | 133.9685 | 129.1574 | 118.8378 | 140.3245 | 102.8426 | 98.9056 | 102.1937 | 140 | 107.7433 | 71.4552 | 87.3511 | 80.5496 | 96.5133 | 95.6441 | 109.6392 | 110.2252 | 78.3245 | 263.1259 | 249.2518 | 264.7433 | 298.3559 | 161.4867 | 196.6174 | 228.2252 | 295.8063 | 320.414 | 213.7167 | 379.7167 | 369.4455
1288 | 16.1624 | 194.2788 | 240.9071 | 407.7123 | 184.4558 | 132.8008 | 178.6372 | 274.4469 | 154.2699 | 166.8186 | 130.9071 | 149.2654 | 161.823 | 211.9071 | 172.084 | 242.0796 | 212.7034 | 196.7212 | 172 | 168.2654 | 183.8275 | 150.0929 | 307.6238 | 586.0707 | 159.3673 | 211.5398 | 242.6505 | 298.084 | 321.177 | 118.6327 | 124.8186 | 141.1814 | 159.2877 | 142.1725 | 128.0796 | 186.1814 | 103.6283 | 240.7123 | 291.4513 | 346.8097 | 299.823 | 153.4558 | 150.7346 | 153.8186 | 133.8141 | 71.3628 | 66.2743 | 104.4469 | 123.5398 | 207.5353 | 222.9911 | 284.7168 | 168.0929 | 141.177 | 144 | 152.7301 | 134.9115 | 153.7257 | 172.084 | 230.6416 | 162.9955 | 142.4424 | 155.4558 | 184.823 | 139.6372 | 209.5442 | 208.177 | 237.0751 | 265.2832 | 129.0796 | 177.0929 | 276.9955 | 189.6283 | 186.177 | 212.1992 | 227.7257 | 306.9866 | 167.6372 | 163.0796 | 209.9071 | 238.8982 | 151.261 | 146.8141 | 147.8186 | 160.9911 | 183.4513 | 187.3673 | 210.9071 | 116.7212 | 120.0974 | 122.7257 | 130.0045 | 121.8982 | 259.1814 | 264.1814 | 344.3628 | 395.6416 | 247.8186 | 271.9911 | 707.0885 | 347.2654 | 146.0974 | 65.7257 | 90.4647 | 146.0974 | 115.1725 | 127.3673 | 150.2788 | 143.0223 | 108.2788 | 111.823 | 126.9026 | 120.7346 | 60.4513 | 69.8982 | 81.4558 | 104.9115 | 110.0974 | 86.7123 | 110.1814 | 76 | 244.8982 | 225.8937 | 230.2565 | 319.1018 | 188.2877 | 195.0929 | 237.1859 | 277.1725 | 305.7168 | 210.7301 | 339.8937 | 350.8052
1289 | 16.1734 | 184.5368 | 235.8456 | 379.3972 | 189.1544 | 133.3088 | 174.0147 | 275.1544 | 167.9558 | 151.2133 | 141.6323 | 145.1691 | 154.8603 | 221.8014 | 163.1839 | 215.3972 | 217.7867 | 190.6765 | 173.6618 | 152.5368 | 181.353 | 152.6618 | 326.0954 | 588.6323 | 172.6323 | 233.0954 | 242.353 | 311.6175 | 309.0295 | 112.8456 | 107.5516 | 156.7867 | 150.2133 | 161.2646 | 132.9705 | 182.6765 | 82.5516 | 238.1544 | 246.7944 | 347.9853 | 295.353 | 132.3972 | 141.5368 | 156.4926 | 128.6765 | 72.6618 | 76.9705 | 100.1691 | 119.3382 | 178.9046 | 230.6323 | 269.1986 | 178.9705 | 139 | 138.1839 | 161.4779 | 128.5221 | 156.3235 | 185.6175 | 252.9411 | 156.8456 | 152.7867 | 149.353 | 208.6028 | 153.294 | 203.1839 | 210.9853 | 233.3235 | 245.9046 | 157.2351 | 163.8751 | 280.8161 | 226.713 | 209.7572 | 189.2575 | 229.4926 | 325.9263 | 153.8751 | 176.2793 | 201.5221 | 244.8014 | 136.8603 | 155.8014 | 162.1249 | 175.2793 | 186.1544 | 172.3825 | 213.3235 | 133.1102 | 111.3677 | 108.706 | 125.353 | 127.8014 | 258.1691 | 271.4779 | 334.0295 | 387.1986 | 247.1691 | 289.6028 | 637.868 | 362.2793 | 136.5368 | 71.647 | 79.2133 | 136.5368 | 125.9558 | 106.5663 | 136.3825 | 109.7944 | 130.9411 | 116.4926 | 116.5221 | 103.2281 | 68.9705 | 77.4632 | 81.1691 | 90.2133 | 96.3825 | 96.6175 | 84.2428 | 66.8603 | 262.4337 | 230.6323 | 285.3158 | 323.1691 | 155.1179 | 200.9853 | 218.2281 | 282.1397 | 288.5368 | 226.1249 | 352.9411 | 301.7944
1290 | 16.1844 | 165.1366 | 222.3525 | 349.7049 | 173.9798 | 157.7655 | 185.6475 | 249.0186 | 165.0978 | 158.9612 | 143.1568 | 134.882 | 142.0388 | 204.6071 | 174.6475 | 233.7655 | 237.8634 | 189.1568 | 154.6071 | 142.4115 | 216.2562 | 163.118 | 329.3137 | 574.9798 | 192.7065 | 232.0978 | 221.6071 | 300.6661 | 344.0994 | 133.0792 | 118.3339 | 152.4115 | 134.3525 | 144.9208 | 135 | 187.9022 | 89.118 | 260.9224 | 254.0202 | 309.3711 | 293.1568 | 149.9224 | 133.0978 | 170.4907 | 130.5295 | 83.118 | 73.0978 | 105.9022 | 113.0978 | 184.118 | 202.6459 | 251.823 | 166.6661 | 138.1568 | 142.059 | 147.823 | 139.6475 | 168.8044 | 178.0388 | 213.1552 | 179.6087 | 145.882 | 151.3727 | 190.2345 | 159.3727 | 169.9596 | 212.8432 | 265.1972 | 265.4519 | 190.1382 | 166.9022 | 288.7453 | 219.823 | 187.1754 | 184.6863 | 212.2935 | 303.1754 | 172.9224 | 147.9596 | 214.3339 | 265.5497 | 123.1956 | 148.7251 | 140.5481 | 153.7049 | 176.0388 | 178.2749 | 187.0186 | 145.4317 | 112.3727 | 125.3929 | 140.0202 | 120.7251 | 237.764 | 274.6863 | 358.9814 | 368.9798 | 225.0776 | 284.7251 | 656.0404 | 392.9814 | 107.8618 | 77.2158 | 84.4317 | 107.8618 | 109.6071 | 105.3727 | 128.7842 | 127.1382 | 106.3323 | 102.6661 | 115.8432 | 111.6475 | 67.6273 | 78.3137 | 95.3339 | 94.5885 | 90.4705 | 91.5683 | 109.3541 | 67.5295 | 239.1754 | 226.2547 | 250.7826 | 306.9798 | 139.4115 | 215.4907 | 228.3339 | 286.5295 | 301.177 | 231.5295 | 323.2733 | 342.7469
1291 | 16.1954 | 204.773 | 218.2891 | 337.3009 | 181.2655 | 187.6638 | 163.1917 | 243.1445 | 164 | 151.5899 | 135.3009 | 150.9647 | 167.3747 | 223.2419 | 158.1799 | 218.3244 | 228.1681 | 183.8672 | 166.3983 | 151.2655 | 200.7581 | 164.1445 | 331.5664 | 500.9969 | 185.7345 | 214.7463 | 238.531 | 298 | 323.6253 | 137 | 133.8319 | 160.4101 | 135.4218 | 141 | 138.4218 | 177.879 | 91 | 230.7816 | 271.5428 | 318.2537 | 281.879 | 134.3244 | 121.7345 | 161.0236 | 131 | 85 | 68.5782 | 105.2891 | 117.1328 | 199.6874 | 227.7965 | 278.9411 | 171.6991 | 131.1563 | 143 | 139.0118 | 135.1563 | 159.0236 | 153.7581 | 219.5428 | 172.0236 | 126.8908 | 132.3244 | 161.1917 | 142.8908 | 182.8201 | 185.6253 | 235.0707 | 259.7345 | 169.3362 | 159.4454 | 260.9144 | 223.8437 | 202.531 | 201.2537 | 206.4336 | 327.0856 | 160.7463 | 177.9293 | 203.3126 | 258.7345 | 144.0974 | 131.6017 | 150.5428 | 155.8437 | 181.6991 | 197.1092 | 197.3983 | 141.0118 | 117.2773 | 122.1563 | 137.0118 | 127.5546 | 258.8083 | 251.9026 | 346.0353 | 376.2655 | 255.2184 | 250.4925 | 570.4658 | 329.5632 | 133.7965 | 82.2773 | 85.1445 | 133.7965 | 112.8437 | 126.531 | 128.8555 | 114.8908 | 101.8555 | 84.6017 | 102.3126 | 105.4454 | 66.1445 | 61.7463 | 86.879 | 83.1681 | 95.1328 | 104.5428 | 106.4454 | 77.4101 | 244.2655 | 207.8908 | 259.9647 | 365.5931 | 157.6756 | 170.9497 | 212.1799 | 300.6874 | 294.5899 | 225.1563 | 288.7698 | 354.5664
1292 | 16.2064 | 167.7447 | 216.2645 | 352.4873 | 189.942 | 148.6125 | 170.2808 | 241.2645 | 165.7355 | 143.9257 | 136.6033 | 146.1902 | 168.529 | 260.8423 | 152.3967 | 193.3062 | 234.6775 | 189.0743 | 157.7192 | 159.942 | 181.3805 | 144.9095 | 335.471 | 508.0905 | 163.174 | 232.826 | 220.3062 | 275.4385 | 281.6867 | 137.8678 | 129.058 | 132.4965 | 149.884 | 148.8098 | 140.7355 | 195.9583 | 80.587 | 270.123 | 254.0417 | 294.0997 | 297.355 | 144.0163 | 135.6195 | 186.768 | 118.8515 | 77.1902 | 76.6775 | 99.7935 | 124.942 | 191.587 | 219.116 | 255.3642 | 165.1902 | 136.942 | 142.1322 | 144.942 | 137.471 | 174.355 | 162.1485 | 221.1322 | 177.8098 | 139.6195 | 148.0905 | 169.1485 | 141.7355 | 199.0163 | 184.471 | 225.529 | 242.3805 | 161.529 | 170.1485 | 272.4873 | 203.3062 | 231.1648 | 208.3388 | 206.8678 | 321.587 | 160.6033 | 170.116 | 212.2808 | 226.6287 | 142.7935 | 158.5035 | 166.0163 | 162.2065 | 160.4385 | 200.8678 | 190.4547 | 133.9257 | 131.884 | 122.7355 | 126.4547 | 128.1322 | 233.4965 | 273.1648 | 325.645 | 367.587 | 280.0905 | 254.5453 | 508.1415 | 344.9003 | 128.587 | 69.116 | 93.6775 | 128.587 | 120.942 | 121.3225 | 112.5127 | 117.2065 | 114.1485 | 106.297 | 116.4873 | 115.2808 | 79.0163 | 75.4873 | 72.8515 | 84.471 | 101.2065 | 102.6612 | 91.9837 | 72.058 | 259.0163 | 200.6612 | 251.7192 | 345.6287 | 163.6033 | 167.3388 | 211.6033 | 281.3062 | 317.297 | 217.058 | 270.116 | 345.4547
1293 | 16.2174 | 177.7208 | 203.6794 | 332.1189 | 168.999 | 164.8811 | 198.4013 | 204.9183 | 181.8408 | 155.3206 | 123.7994 | 144.12 | 171.5202 | 238.7187 | 174.8811 | 189.12 | 227.1996 | 171.5191 | 148.0796 | 146.9193 | 179 | 157.8408 | 332.4798 | 514.5202 | 160 | 212.2389 | 208.1996 | 301.9214 | 338.483 | 113.3588 | 118.3195 | 129.7601 | 142.3195 | 132.3992 | 153.3206 | 157.638 | 93.9607 | 243.5584 | 277.4013 | 295.2803 | 317.6008 | 121.3588 | 162.6412 | 162.8387 | 126.6805 | 87.4405 | 77.12 | 91.0796 | 106.6391 | 183.8397 | 209.0796 | 257.1603 | 164 | 148.5605 | 148.1603 | 135.4395 | 106.3185 | 177.88 | 190.4013 | 222.7601 | 159.6391 | 138.4798 | 149.2399 | 192.121 | 147.2803 | 181.6391 | 192.0403 | 225 | 261.121 | 163.6401 | 142.9586 | 230.1179 | 216.7208 | 212.999 | 191.3992 | 219.3206 | 317.3599 | 153.0796 | 180.3206 | 196.3992 | 240.4809 | 147.2803 | 147.1592 | 161.8397 | 157.7197 | 176.3609 | 185.1592 | 180.1996 | 123.3195 | 106.7187 | 115.9597 | 126.7601 | 130.6401 | 265.0817 | 264.6794 | 308.9193 | 357.1996 | 259.2389 | 256 | 466.6784 | 380.6815 | 133.1603 | 71.4002 | 89.7197 | 133.1603 | 110.5595 | 119.12 | 98.5595 | 118.88 | 127.4405 | 106.4798 | 118.12 | 102.9193 | 71.3195 | 78 | 73.6401 | 112.2813 | 105.5202 | 94.9597 | 115.5212 | 63.9597 | 244.2792 | 230.8015 | 234.1592 | 289.9576 | 177.2006 | 175.0403 | 196.1592 | 286.8004 | 276.9979 | 205.4395 | 286.4809 | 326.3992
1294 | 16.2284 | 167.5073 | 193.969 | 318.292 | 157.0767 | 147.4763 | 173.4454 | 216.062 | 179.5383 | 144.5073 | 124.677 | 133.292 | 159.5073 | 191.2757 | 186.9233 | 197.9233 | 201.0147 | 169.8923 | 141.646 | 161.9543 | 201.3083 | 171.6003 | 352.5237 | 432.0131 | 150.1843 | 212.5693 | 243.5856 | 289.0457 | 285.429 | 106.4307 | 123.2463 | 147.8467 | 141 | 157.6623 | 147.8613 | 172.5237 | 81.7227 | 247.031 | 250.6607 | 303.1387 | 297.6917 | 143.8776 | 155.292 | 158.1077 | 114.615 | 82.7537 | 82.354 | 90 | 97.7537 | 193.708 | 227.6313 | 241.0457 | 157.7537 | 135.7227 | 149.8923 | 150.9543 | 115.385 | 180.677 | 193.1077 | 243.5237 | 161.4617 | 129.969 | 154.354 | 186.0767 | 145.323 | 179 | 185.8613 | 239.2773 | 256.8613 | 178.2773 | 163.9853 | 239.1697 | 213.646 | 200.1843 | 212.2007 | 193.3377 | 315.2153 | 161.8157 | 168.615 | 183.292 | 234.0767 | 140.8613 | 126.261 | 139.584 | 157 | 186.1387 | 176.7537 | 195.062 | 150.5546 | 119.062 | 114.1077 | 127.8923 | 127.4307 | 253.0457 | 290.6623 | 320.385 | 399.7243 | 250.646 | 246.1843 | 462.8923 | 364.4763 | 117.0457 | 65.7537 | 89 | 117.0457 | 89.3687 | 110.0767 | 124.6623 | 102.0457 | 104.0147 | 108.677 | 124.2463 | 104.5693 | 83.385 | 73.5383 | 74 | 91.0147 | 101.5383 | 99.354 | 100.261 | 81.739 | 242 | 233.2153 | 258.77 | 314.2316 | 175.4307 | 209.9086 | 210.062 | 294.2463 | 277.2463 | 206.677 | 269.3687 | 348.093
1295 | 16.2394 | 177.7601 | 209.2832 | 278.1012 | 156.9046 | 154.0462 | 162.763 | 225.237 | 160.0029 | 164.711 | 142.1879 | 141.0462 | 171.5694 | 203.1879 | 176.2399 | 187.2399 | 209.7601 | 170 | 151.8555 | 167.6185 | 202.1908 | 154.0029 | 324.2428 | 471.7543 | 165.2832 | 216.6185 | 228.0983 | 262.5752 | 276.1908 | 107.8092 | 103.1937 | 143.6676 | 133.763 | 169.1416 | 140.6676 | 155.0983 | 81.8092 | 246.1908 | 216.2428 | 299.4769 | 272.3844 | 131.6214 | 133.1937 | 175.1879 | 110.2861 | 66.6214 | 69.4306 | 86.3815 | 108.7601 | 213.0925 | 242.6647 | 243.5231 | 168.7601 | 144.8555 | 140.9538 | 146.6676 | 119.7139 | 181 | 184.8584 | 216.1474 | 173.7601 | 127.1908 | 145.0491 | 190.4277 | 135.9538 | 176.2861 | 207.6156 | 211.1474 | 232.4798 | 136.5781 | 173.3324 | 270.8526 | 230.1879 | 201.7139 | 196.0029 | 194.5231 | 334.9017 | 158.4769 | 149.8121 | 201.9017 | 255.6156 | 136.3815 | 132.1416 | 139.7139 | 151.5723 | 161.6705 | 188.6647 | 183.4306 | 133.1937 | 109.2399 | 100.4306 | 104.4798 | 140.5694 | 248.2861 | 271.3844 | 316.5723 | 367.9104 | 235.526 | 255.8555 | 445.8121 | 388.2341 | 147.5665 | 70.4277 | 85.3815 | 147.5665 | 129.5173 | 121.6647 | 111.7168 | 127.1387 | 95.5723 | 98.1445 | 124.0954 | 126.711 | 70.526 | 73 | 81.237 | 97.9509 | 86.526 | 98.1908 | 104.3324 | 74.9538 | 240.1908 | 228.4769 | 249.3353 | 327.9509 | 148.7659 | 208.5723 | 208.3815 | 273.289 | 271.6676 | 233.2341 | 259.763 | 352.8092
1296 | 16.2504 | 160.6617 | 231.1721 | 303.3413 | 190.0089 | 147.6647 | 147.3294 | 214.997 | 148.8309 | 143.1602 | 124.7448 | 142.9169 | 174.8338 | 206.8338 | 170.4154 | 161.2433 | 214.6677 | 181.003 | 125.4926 | 144.1602 | 205.6677 | 151.0831 | 314.5816 | 421.0682 | 175.2522 | 208.7478 | 230.5846 | 287.5074 | 302.5905 | 119.9199 | 113.8368 | 132.9139 | 144.9199 | 130.5727 | 135.4154 | 164.003 | 85.6677 | 230.4124 | 252.4273 | 298.0831 | 275.5015 | 118.0801 | 137.4184 | 156.8279 | 117.3353 | 70.5015 | 60.6647 | 80.4985 | 117.3353 | 219.5846 | 224.7448 | 268.7567 | 160.8309 | 123.0771 | 159.2552 | 133.1632 | 108.997 | 167.2463 | 196.8368 | 202.9139 | 160.3294 | 138.003 | 158.6706 | 192.8338 | 157.9229 | 185.1691 | 206.3323 | 228.1721 | 241.003 | 151.2552 | 149.2433 | 259.3294 | 220.0801 | 194.6647 | 201.3353 | 188.5816 | 336.0831 | 152.4985 | 167.2552 | 176.4926 | 213.9881 | 119.4955 | 152.2552 | 166.5905 | 157.4184 | 173.6706 | 194.5846 | 183.8338 | 140.1691 | 99.7478 | 135.6766 | 119.4214 | 144.7507 | 291.095 | 261.6647 | 295.8279 | 360.3323 | 244.0861 | 245.997 | 386.2344 | 347.905 | 136.3294 | 66.4154 | 64.8279 | 136.3294 | 106.4926 | 109.2463 | 137.5074 | 118.0801 | 100.5015 | 92.4154 | 109.3294 | 117.0801 | 89.1721 | 71.1662 | 74.6647 | 87.0801 | 86.8338 | 82.4124 | 102.2493 | 89.5876 | 225.3294 | 216.0801 | 248 | 323.4985 | 157.9199 | 216.2522 | 219.003 | 264.5816 | 302.1751 | 198.4065 | 261.7507 | 340.1632
1297 | 16.2614 | 151.5664 | 220.9203 | 297.6371 | 194.8584 | 150.7168 | 167.3717 | 214 | 152.646 | 157.7257 | 108.1327 | 136.4956 | 170.354 | 212.5752 | 172.7876 | 202.6727 | 233.5841 | 198.7257 | 154.593 | 158.7257 | 213.4336 | 169.5841 | 314.9292 | 381.6194 | 174.1416 | 199.6371 | 212.4159 | 275.1327 | 276.1946 | 108.9203 | 127.0797 | 118.9911 | 144.1416 | 129.7876 | 166.593 | 144.5575 | 102.7257 | 268.0266 | 284.8054 | 292.4248 | 285.2921 | 113.2832 | 128.7079 | 153.1416 | 118.9292 | 81.2213 | 78.5841 | 80.9292 | 143.0886 | 212.5664 | 205.3451 | 260.7787 | 160.9292 | 149.8054 | 135.9114 | 118.0619 | 114.5044 | 192.0178 | 179.4159 | 209.4336 | 162.7168 | 142.7168 | 154.4248 | 166.053 | 139.5575 | 152.5486 | 219.9381 | 221.6371 | 236.4248 | 146.4956 | 152.5752 | 228.2654 | 219.9292 | 195.8584 | 230.8054 | 204.7257 | 328.5664 | 170.5841 | 175.5044 | 194.4425 | 206.2832 | 126.3629 | 166.0797 | 155.0619 | 158.9292 | 169.4248 | 160.6194 | 170.9911 | 139.1416 | 102.7168 | 118.5575 | 137.7257 | 123.6283 | 236.46 | 270.2921 | 296.7876 | 374.8673 | 233.8495 | 229.2035 | 354.053 | 340.2832 | 123.8495 | 79.0089 | 71.3629 | 123.8495 | 113.2921 | 94.0619 | 119.5575 | 126.2921 | 104.7168 | 95.7168 | 120.0797 | 117.8584 | 75.2035 | 66.354 | 83.2921 | 97.1505 | 88.8584 | 90.2921 | 101.0708 | 77.9911 | 225.8584 | 214.0708 | 239.6371 | 313.7079 | 160.8584 | 208.6371 | 203.2743 | 295.593 | 306.8584 | 226.593 | 250.8495 | 343.646
1298 | 16.2724 | 170.7714 | 214.351 | 294.1755 | 217.5959 | 176.4204 | 164.2925 | 198.936 | 158.649 | 149.585 | 128.6544 | 136.9415 | 159.6435 | 196.9946 | 175.8245 | 169.2816 | 214.2871 | 174.5796 | 156.0585 | 169.415 | 175.3986 | 156.8775 | 308.4095 | 390.298 | 162.702 | 264.9048 | 229.8299 | 241.0476 | 296.5959 | 108.9415 | 119.5265 | 128.3565 | 116.6966 | 134.7075 | 158.6435 | 155.2395 | 110.5905 | 232.3986 | 250.2816 | 308.0054 | 263.4041 | 108.2925 | 135.532 | 156.766 | 116.1755 | 72.585 | 80.9415 | 96.064 | 112.0476 | 189.4041 | 218.1225 | 240.2286 | 171.3565 | 134.1116 | 139.649 | 138.6544 | 130.064 | 194 | 159.1701 | 192.1116 | 158.2925 | 164.6544 | 139.8775 | 171.532 | 160.5959 | 158.4735 | 174.8666 | 220.0585 | 268.9524 | 162.0054 | 158.649 | 212.819 | 194.5796 | 205.415 | 205.6966 | 244.6014 | 318.585 | 143.7551 | 159.0531 | 221.4204 | 219.181 | 130.766 | 134.9891 | 141.7605 | 147.702 | 172.766 | 198.4844 | 174.7075 | 130.5265 | 115.2395 | 113.234 | 119.2286 | 132.3565 | 264.9524 | 251.2286 | 274.4041 | 372.234 | 249.0054 | 264.7184 | 329.4041 | 325.8775 | 137.1225 | 66.819 | 63.5265 | 137.1225 | 106.468 | 105.2395 | 115.1755 | 90.2816 | 102.1755 | 116.7129 | 107.819 | 118.9415 | 62.702 | 63.1755 | 78.351 | 97.0585 | 88.0585 | 99.4735 | 98.1755 | 82.649 | 258.0109 | 206.468 | 255.0054 | 273.4571 | 165.7075 | 227.7714 | 219.8884 | 279.1701 | 289.1116 | 233.7075 | 222.6966 | 349.649
1299 | 16.2834 | 160.5545 | 197.7856 | 280.647 | 163.6803 | 159.878 | 166.8614 | 198 | 167.5841 | 152.8152 | 120.4621 | 131.2773 | 174.2606 | 219.8447 | 144.525 | 177.4917 | 220.6303 | 201.6136 | 153.1386 | 160.4621 | 184.4455 | 145.5083 | 308 | 401.4917 | 184.8909 | 273.7689 | 206.2015 | 279.0591 | 297.0462 | 113.7689 | 118.0462 | 138.5379 | 115.9538 | 131.1848 | 141.7856 | 155.0462 | 83.3401 | 239.5379 | 228.9242 | 308.0462 | 278.2144 | 123.2606 | 156.9833 | 121.7098 | 136.0295 | 72 | 70.5083 | 89.3697 | 117.6303 | 200.3992 | 214.2311 | 235.1848 | 157.6932 | 149.2144 | 146.6765 | 141.9076 | 126.2311 | 186.3697 | 172.3068 | 214.8447 | 130.3401 | 146.9242 | 161.8909 | 169.1386 | 148.647 | 148.5083 | 192.0295 | 234.3068 | 225.2182 | 154.4159 | 162.8152 | 238.7061 | 202.5379 | 165.9409 | 182.0629 | 204.0795 | 253.1424 | 167.7523 | 166.5841 | 172.4492 | 229.5379 | 122.4159 | 136.8152 | 124.7856 | 159.3992 | 160.6008 | 180.0167 | 184.5379 | 136.6765 | 111.2311 | 101.5545 | 116.0924 | 127.2773 | 246.9705 | 279.5674 | 276.8152 | 380.5841 | 247.1386 | 247.9242 | 332.7689 | 365.0591 | 136.0924 | 80.3068 | 80.1682 | 136.0924 | 127.9371 | 105.0462 | 115 | 115.6599 | 103.9076 | 86.525 | 110.8152 | 116.1386 | 70.5841 | 67.7689 | 74.1848 | 96.0462 | 100.3992 | 112.3992 | 99.9076 | 68.6932 | 235.2015 | 216.4917 | 228.3401 | 311.0591 | 151.6932 | 208.0167 | 207.647 | 280.8614 | 308.0295 | 207.2939 | 229.5841 | 323.2939
1300 | 16.2944 | 157.1018 | 201.8304 | 283.8643 | 204.4736 | 170.5929 | 155.4071 | 184.4749 | 154.4749 | 153.9661 | 125.7965 | 138.7286 | 149.8819 | 218.1018 | 182.6092 | 170.2714 | 225.8304 | 173.0515 | 151.0678 | 147.441 | 210.1181 | 181.711 | 327.3216 | 384.6106 | 168.6106 | 228.5943 | 212.7286 | 258.7802 | 297 | 109.1696 | 121.8643 | 109.0515 | 135.3216 | 138.7286 | 137.1357 | 155.9661 | 86.8304 | 240 | 246.3555 | 279.0176 | 268.3731 | 141.3894 | 132.8819 | 132.559 | 121.5427 | 79.7286 | 73.8643 | 80.3053 | 105.441 | 190.3731 | 219.7965 | 268.8128 | 190.8128 | 123.9159 | 143.1357 | 154.559 | 119.2374 | 160.8819 | 187.4912 | 227.5929 | 166.6771 | 160.4912 | 148.5088 | 173.8304 | 143.1696 | 180.8467 | 175.6106 | 215.6784 | 238.4573 | 178.152 | 160.1018 | 225.5088 | 203.9661 | 220.0326 | 194.5251 | 218.4233 | 306.0326 | 190.2537 | 167.9661 | 161.3053 | 228.0678 | 143.2537 | 137 | 153.9485 | 162.8982 | 161.9322 | 176.1018 | 197.559 | 124.441 | 130.3216 | 120.3216 | 118.8982 | 145.3555 | 270.152 | 258.7802 | 271.2035 | 355.8819 | 246.0339 | 241.2035 | 292.4247 | 292.6119 | 128.2714 | 67.4749 | 75.2035 | 128.2714 | 114.5088 | 116.5929 | 119.8304 | 129.559 | 89.5088 | 134.27 | 105.2035 | 107.3053 | 65.2035 | 68 | 76.8982 | 98.8982 | 76.848 | 89.8141 | 91.3053 | 70.8982 | 227.2374 | 231.4912 | 241.4912 | 296.5767 | 151 | 197.3392 | 196.3731 | 277.1357 | 287.7463 | 216.6269 | 245.4573 | 353.8806
1301 | 16.3054 | 139.3893 | 202 | 235.0815 | 190.3461 | 180.7837 | 179.4593 | 189.8702 | 152.0433 | 162.8053 | 133.827 | 123.3461 | 153.8919 | 198.4326 | 157.584 | 161.1947 | 217.1947 | 169.0649 | 156.8702 | 162.6539 | 176.757 | 149.7354 | 310.3893 | 379.1081 | 151.3677 | 218.1947 | 232.5674 | 275.6107 | 311.6756 | 124.6539 | 114.173 | 133.4376 | 141.8702 | 142.9135 | 139.9351 | 149.1514 | 81.1298 | 245.8702 | 266.5674 | 248.6489 | 257.2379 | 140.0433 | 130.0433 | 170.1781 | 132.7404 | 66.3028 | 65.1947 | 91.7404 | 105 | 199.7837 | 212.173 | 262.173 | 162.6489 | 136.6972 | 124.411 | 117.8219 | 127.8053 | 190.3295 | 182.1298 | 219.1947 | 173.8702 | 136.5407 | 140.173 | 156.3893 | 138.1081 | 166.3461 | 196.5241 | 215.9784 | 226.2812 | 147.6922 | 148.2596 | 265.1132 | 225.5241 | 192.6489 | 196.9567 | 210.1947 | 317.7837 | 179.2596 | 193.4376 | 156.1081 | 240.7188 | 136.173 | 154.6107 | 147.173 | 156.1514 | 162.9784 | 178.9351 | 165.7138 | 109.3244 | 113.3893 | 108.2812 | 125.8486 | 125.4542 | 248.4975 | 237.4542 | 301.3295 | 338.3677 | 258.7188 | 249.8053 | 306.6539 | 354.5724 | 144.6323 | 73.8486 | 72.0649 | 144.6323 | 138.4593 | 107.2163 | 106.3028 | 109.4542 | 100.7404 | 114.4759 | 113.8053 | 121.6756 | 73.8053 | 67.0216 | 78.9567 | 90.1947 | 90.6756 | 101.7188 | 105.6756 | 80.7837 | 227 | 227.1081 | 235.1514 | 333.1781 | 169.589 | 189.173 | 228.2862 | 260.3677 | 291.8919 | 176.8868 | 263.6107 | 363.8053
1302 | 16.3164 | 174.6638 | 207.944 | 237.9626 | 201.8879 | 122.551 | 203.7759 | 175.1401 | 170.8226 | 149.1307 | 123.1027 | 131.916 | 159.944 | 179.1774 | 177.8039 | 154.0654 | 219.972 | 175.9346 | 144.1214 | 145.1681 | 189.8693 | 156.9253 | 365.477 | 356.2148 | 167.8412 | 217.0093 | 208.2335 | 262.1307 | 274.3549 | 117.0747 | 114.9907 | 132.0187 | 118.2241 | 136.0654 | 141.9813 | 157.916 | 80.0093 | 246 | 276.9066 | 284.6545 | 276.8132 | 128.1121 | 131.9813 | 133.3549 | 119.1307 | 73.9253 | 58.0654 | 92.9907 | 95.0934 | 196.0374 | 193.1774 | 232.2802 | 162 | 137 | 144.8039 | 144.7385 | 107.1961 | 190.0093 | 182 | 218.0093 | 158.1494 | 127.084 | 150.8973 | 168.8786 | 146.916 | 173.9253 | 211.8599 | 210.056 | 258.6918 | 152.944 | 182.6731 | 223.4016 | 212.1307 | 190.0187 | 204.9253 | 215.944 | 318 | 148.2895 | 154.3736 | 176.8039 | 263.7852 | 145.9066 | 165.8973 | 136.1027 | 150.056 | 174.8879 | 200.7945 | 165.9907 | 142.6825 | 93.1868 | 110.972 | 112.1307 | 144.8132 | 278.7105 | 303.3743 | 244.5417 | 379.6078 | 235.2241 | 256.9346 | 295.1121 | 348.0747 | 116.2708 | 71.028 | 66.056 | 116.2708 | 146.9253 | 109.972 | 122.8412 | 102.0654 | 102.9813 | 119.944 | 111.028 | 124.972 | 69.0467 | 63.0374 | 78.0093 | 69.1961 | 93.972 | 106.9533 | 111.944 | 87.9346 | 237.8973 | 226.0093 | 265.7105 | 306.2615 | 182.8786 | 204.8506 | 211.1681 | 283.7759 | 327.6638 | 220.5798 | 259.0467 | 357.0654
1303 | 16.3274 | 146.0384 | 203.059 | 270.9262 | 182.9351 | 154.9439 | 166 | 185.0177 | 157.0384 | 149.0679 | 123.0354 | 134.9646 | 149.003 | 225.9055 | 156.9911 | 165.9852 | 227.0443 | 181.0059 | 122.1122 | 141.9793 | 184.0856 | 153.9528 | 330.8996 | 355.9498 | 175.0118 | 238.9705 | 223.0561 | 240.9911 | 304.0295 | 128.9587 | 99.0561 | 112.0915 | 118.0531 | 146.941 | 130.062 | 149.0177 | 102.0148 | 236.9557 | 287.9528 | 271.9587 | 272.9882 | 143.9469 | 136.0325 | 139.0325 | 116.9911 | 73.0325 | 77.9852 | 95.9852 | 101.0325 | 185 | 223.8494 | 220.1506 | 129.0738 | 137.062 | 129.0384 | 128.0413 | 120.9911 | 190.9173 | 168.0443 | 190.0856 | 164.9675 | 159.0059 | 150.9882 | 172.9764 | 137.9587 | 169.0089 | 171.0354 | 235.9941 | 223.0531 | 138.0502 | 169.9557 | 231.9528 | 232.9114 | 185.9852 | 195.0207 | 221.8996 | 335.8849 | 147.0236 | 167.9114 | 175.0856 | 215.0177 | 160.8583 | 136.0531 | 135.1092 | 141.0266 | 168.0207 | 162.0974 | 184.0089 | 119.0797 | 113.9616 | 112.9882 | 129.9646 | 132.9882 | 247.0679 | 267.0413 | 269 | 363.0059 | 266.9498 | 258.0118 | 327.8937 | 358.9705 | 136.003 | 61.0266 | 83.9764 | 136.003 | 109.0177 | 104.9823 | 111.0177 | 89.1004 | 97.0236 | 117.9469 | 97.0443 | 104.003 | 73.0413 | 87.9498 | 74.9528 | 93.9675 | 113.9734 | 74.0945 | 104.941 | 67.059 | 241.0177 | 240.9646 | 233.9764 | 290.0413 | 153.9675 | 172.0561 | 202.0059 | 226.1712 | 284.0354 | 232.9291 | 242.9705 | 334.9823
1304 | 16.3384 | 158.9848 | 222.7256 | 245.9238 | 161.1524 | 136.061 | 165.7561 | 190.3598 | 169.9238 | 171.5579 | 135.1219 | 123.1524 | 150.1372 | 194.3049 | 153.939 | 161.625 | 241.8628 | 182.9238 | 160 | 135.1982 | 212.5579 | 138.3811 | 297.5335 | 338.4512 | 178.6037 | 229.0762 | 241.8628 | 237.8323 | 314.061 | 115.1677 | 118.2286 | 142.6646 | 136.2286 | 127.1219 | 150.6951 | 154.9238 | 106.6951 | 222.1982 | 272.2591 | 258.0915 | 268.8781 | 125.9543 | 146.7714 | 150.1067 | 114.1219 | 83.9085 | 73.0457 | 90.8171 | 111.8018 | 185.3506 | 173.4725 | 270.1921 | 153.7561 | 157.9543 | 142.1372 | 142.0152 | 117.9085 | 163.2744 | 182.8781 | 218.3903 | 154.6555 | 160.6799 | 146.9695 | 165.1829 | 124.3354 | 172.1982 | 183.1067 | 233.7256 | 240.9238 | 155.0457 | 155.061 | 216.3354 | 203 | 181.2134 | 201.9238 | 187.8323 | 297.0457 | 154.7866 | 138.5792 | 203.3293 | 221.7622 | 113.4573 | 153.7256 | 171.5427 | 150.2744 | 175 | 194.7104 | 186.7256 | 145.8171 | 101.2744 | 108.9238 | 117.9848 | 128.9695 | 270 | 280.9848 | 268.6951 | 365.1219 | 250.2744 | 261.5732 | 292.1067 | 349.0305 | 136.6189 | 69.9085 | 76.1829 | 136.6189 | 115.1219 | 99.2286 | 116.9543 | 122.7256 | 104.8933 | 99.9695 | 112.0457 | 105.2439 | 86.7104 | 71.0762 | 59.2439 | 82.939 | 105.0152 | 105.8933 | 84.9238 | 86.939 | 246.497 | 228.7256 | 225.8476 | 304.3811 | 143.0762 | 191.0915 | 204.1829 | 283.8018 | 295.6189 | 208.9543 | 232.8628 | 329.1982
1305 | 16.3494 | 157.3392 | 204.6971 | 241.2203 | 170.8899 | 140.3855 | 149.9174 | 149.9086 | 165.0826 | 143.3855 | 142.5319 | 132.7247 | 159.1377 | 213.2015 | 150.2203 | 201.5319 | 232.8899 | 177.6421 | 160.0826 | 148.3029 | 183.8899 | 162.8348 | 331.7247 | 304.4869 | 153.4405 | 233.0638 | 232.3116 | 227.5232 | 316.6784 | 125.8899 | 132.8073 | 121.2203 | 150.6971 | 135.1927 | 130.9725 | 150.0275 | 87.2203 | 234.8899 | 288.4768 | 264 | 260.8623 | 123.7985 | 132.2478 | 156.2566 | 122.0826 | 77.8348 | 75.7247 | 79.2203 | 99.5782 | 207.174 | 204.2203 | 217.9449 | 139.1564 | 154.7247 | 150.4218 | 143.3855 | 111.9725 | 180.5044 | 175.6058 | 179.9637 | 196.1189 | 140.3029 | 144.8348 | 176.2015 | 146.1101 | 183.8711 | 190.9362 | 216 | 235.8623 | 157.7797 | 158.7797 | 238.0826 | 203 | 194.8899 | 197.2203 | 178.0463 | 300.6608 | 141.1652 | 175.8899 | 159.6696 | 269.4305 | 142.8623 | 135.8623 | 142.5507 | 167.6696 | 175.413 | 176.2478 | 169.6333 | 134.5232 | 118.5044 | 104.3029 | 116.7247 | 127.1927 | 269.0914 | 279.8348 | 250.2666 | 372.7797 | 268.3029 | 233.9174 | 298.174 | 351.0551 | 112.0551 | 64.4956 | 87.3942 | 112.0551 | 122.7522 | 114.0275 | 113.8348 | 105.771 | 98.413 | 98.6608 | 114.9449 | 120.3392 | 68.0275 | 75.7247 | 75.3029 | 78.7247 | 105.6145 | 98.6696 | 80.413 | 82.587 | 214.9362 | 211.0275 | 215.9174 | 328.4493 | 147.7522 | 195.8711 | 215.8073 | 271.413 | 273.01 | 205.8623 | 224.9362 | 342.1652
1306 | 16.3604 | 134.7965 | 193.8805 | 248.2433 | 166.8009 | 153.4823 | 147.1195 | 181.6814 | 167.2433 | 155.8451 | 126.1593 | 122.6416 | 164.239 | 185.3983 | 157.4424 | 184.8009 | 230.2346 | 164.323 | 162.323 | 158.7212 | 178.9247 | 157.4381 | 323.6727 | 356.9203 | 168.7212 | 199.8009 | 208.0398 | 245.5221 | 270.9558 | 121.6416 | 125.4026 | 129.4381 | 139.0442 | 142.3584 | 131.0753 | 150.2831 | 94.7212 | 231 | 268.8451 | 262.9247 | 257.1151 | 151.8009 | 140.6017 | 130.2788 | 124.761 | 73.5134 | 66.1991 | 86.6017 | 119.9203 | 179.4337 | 211.761 | 216.4381 | 178.2079 | 143.6858 | 129.9203 | 155.8451 | 111.5177 | 163.4779 | 196.0442 | 213.8407 | 165.1991 | 151.5177 | 139.7169 | 149.9116 | 149.084 | 145.0753 | 223.323 | 215.6814 | 230.6017 | 151.1948 | 152.1151 | 240.761 | 203.2788 | 190.8805 | 204.8805 | 214.7212 | 324.0398 | 148.1948 | 171.8805 | 148.8762 | 214.6372 | 137.761 | 131.7567 | 161.7212 | 156.3584 | 190.1195 | 184.761 | 191.3628 | 152.9602 | 100.8805 | 114.4026 | 107.0797 | 133.8805 | 237.9558 | 273.084 | 293.4866 | 364.761 | 279.0797 | 231.5576 | 270.2346 | 351.8052 | 114.7965 | 81.4823 | 66.0797 | 114.7965 | 113.761 | 115.8363 | 108.4779 | 132.6814 | 113.0398 | 122.1991 | 112.9203 | 98.1948 | 69.5177 | 65.9602 | 85.4424 | 69.9558 | 91.9602 | 88.0355 | 95 | 67.9602 | 246.2875 | 212.6372 | 213.239 | 308.8805 | 140.1549 | 157.6727 | 208.084 | 286 | 341.73 | 201.9558 | 256.3672 | 348.5177
1307 | 16.3714 | 154 | 191.3127 | 230.5212 | 163.2508 | 141.1563 | 150.9902 | 174.6775 | 149.4169 | 129.5114 | 130.1563 | 114.886 | 169.8958 | 196.0423 | 144.7817 | 180 | 258.3844 | 149.1466 | 146 | 151.8437 | 153.3127 | 167.3746 | 361.759 | 354.114 | 162.9902 | 195.9902 | 210.5635 | 234 | 293.8437 | 112.5831 | 111.469 | 139.5831 | 118.9706 | 151.9381 | 155.4886 | 133.7296 | 87.7394 | 233.0325 | 242.4593 | 240.7003 | 283.4788 | 146.0098 | 131.6775 | 137.0521 | 119.5212 | 107.0294 | 70.4267 | 76.9479 | 118.3648 | 211.2378 | 206.469 | 228.5635 | 134.9479 | 112.9902 | 127.8437 | 128.9902 | 124.7296 | 174.531 | 173.8339 | 209.4267 | 167.9153 | 162.8013 | 156.114 | 193.7069 | 128.0944 | 170.8958 | 205.645 | 208.3648 | 219.8534 | 178.8534 | 177.5929 | 235.4169 | 209.6873 | 188.7296 | 202.1042 | 207.3746 | 324.0619 | 176.7394 | 169.0521 | 169.4267 | 230.7296 | 131.6352 | 148.7492 | 153.8534 | 164.2704 | 191.5929 | 179.3127 | 176 | 152.1563 | 98.5733 | 100.3648 | 108.2704 | 130.6873 | 260.9479 | 252.6156 | 257.4169 | 361.3452 | 279.9056 | 245.9381 | 300.0521 | 325.5016 | 133.3746 | 69.9381 | 68.3127 | 133.3746 | 108.6254 | 134.3323 | 119.9479 | 124.3225 | 113.4267 | 125.5929 | 111.0521 | 126.6352 | 81.5831 | 65.3648 | 71.9479 | 92.6352 | 92.1466 | 113.3127 | 95.1042 | 68.0423 | 206.355 | 227.4267 | 219.1042 | 305.3225 | 168.0521 | 198 | 187.4071 | 284.5407 | 287.9902 | 224.531 | 216.1661 | 358.1857
1308 | 16.3824 | 155.1593 | 196.9356 | 241.2237 | 183.7457 | 145.1593 | 168.0983 | 186.5491 | 155.9051 | 156.1627 | 133.2576 | 130.5491 | 167.3559 | 214.678 | 159.1288 | 178.4542 | 228.2915 | 170.1932 | 146.3864 | 149.4509 | 162.156 | 154.3254 | 319.3254 | 338.9017 | 180.3559 | 214.1932 | 241.3187 | 233.9356 | 289.9051 | 107.4475 | 120.7729 | 131.4203 | 170.3593 | 167.9695 | 127.9356 | 147.1932 | 83.1932 | 269.6136 | 269.0644 | 307.2915 | 273.3559 | 128.1932 | 143.7424 | 137.3559 | 128.9356 | 54.1593 | 60.322 | 76.5797 | 125.1932 | 163.6407 | 213.261 | 254.8746 | 135.2881 | 129.9695 | 126.2881 | 145.583 | 137.9356 | 167.417 | 187.0034 | 200.8034 | 131.8034 | 140.0339 | 140.1932 | 153.417 | 148.0644 | 169.322 | 181.3864 | 215.8373 | 200.9322 | 158.5153 | 152.2576 | 242.8712 | 204.5153 | 204.061 | 204.4509 | 197.2237 | 306.2271 | 172 | 172.4475 | 161.061 | 242.7119 | 124.9356 | 127.0949 | 133.5797 | 150.9356 | 167.3526 | 184.5491 | 176.322 | 153.0678 | 107.9051 | 106.8068 | 96.9966 | 125.8373 | 259.9356 | 281.2271 | 265.322 | 402.1966 | 258.6474 | 262.0983 | 300.3559 | 368.5525 | 122.0644 | 86.8712 | 73.6136 | 122.0644 | 118.0678 | 104.322 | 118.678 | 111.6136 | 103.4509 | 101.0305 | 113.1593 | 119.8068 | 75.417 | 72.1288 | 72.739 | 86.6441 | 112.8712 | 117.5186 | 97.6441 | 86.678 | 231.4509 | 216.3559 | 221.4509 | 293.7085 | 168.2915 | 199.0949 | 212.6136 | 258.0644 | 303.7457 | 215.2271 | 202.8034 | 310.6068
1309 | 16.3934 | 171.7699 | 196.5369 | 255.3952 | 152.6107 | 160.466 | 153.6194 | 178.5427 | 140.9204 | 146.0709 | 134.8524 | 125.9175 | 158 | 208.9262 | 159.5427 | 154.466 | 218.1534 | 172.0796 | 150.8495 | 156.767 | 206.0058 | 133.1446 | 296.9175 | 351.3097 | 170.7699 | 218.1505 | 271.4719 | 234.9942 | 275.2272 | 141.2359 | 131.233 | 122.9233 | 133.3039 | 152.1563 | 127.3068 | 150.8437 | 86.2301 | 262.0825 | 268.1592 | 295.2359 | 263.5398 | 130.7699 | 139.0796 | 131.2213 | 128.1534 | 71.4602 | 65.3835 | 84.6165 | 126.8495 | 201.8466 | 188.9971 | 224.767 | 153.5398 | 115.6136 | 144.4631 | 125.6903 | 134.0088 | 185.8524 | 159.534 | 226.0029 | 158.767 | 126.0767 | 142.0029 | 174.8437 | 148.4631 | 174.1534 | 186.7699 | 226.7728 | 229.3068 | 164.6961 | 157.1505 | 243.9912 | 212.0767 | 229.3214 | 209.3126 | 213.0825 | 300.1388 | 171.4631 | 205.9291 | 187.3952 | 223.7699 | 126.0709 | 143.0767 | 141.7699 | 151.3806 | 185.466 | 175.6223 | 180.233 | 126.0738 | 94.301 | 104.767 | 124.466 | 136.3893 | 255.2418 | 269.3864 | 271.3039 | 376.6136 | 239.2301 | 248.6932 | 290.3097 | 334.3806 | 124.2272 | 84.6165 | 67.9233 | 124.2272 | 91.1505 | 107.6961 | 113.4631 | 107.7641 | 110.767 | 115.9233 | 127.7757 | 116.3097 | 94.3893 | 73.6932 | 96.9262 | 96.3835 | 110.1563 | 94.9262 | 106.7699 | 81.7699 | 238.0767 | 208.3039 | 227.9233 | 304 | 156.3893 | 215.2301 | 206.0029 | 259.3835 | 272.4573 | 205.6874 | 227.2359 | 361.6961
1310 | 16.4044 | 168.0211 | 204.1569 | 213.6018 | 172.9789 | 142.356 | 140.5596 | 163.9367 | 152.9789 | 168.5083 | 111.4027 | 147.3092 | 156.5762 | 196.267 | 142.267 | 139.1146 | 220.267 | 164.3816 | 135.7541 | 162.5294 | 182.8009 | 170.2881 | 322.9367 | 343.356 | 169.2459 | 231.1101 | 230.4238 | 256.7753 | 289.3771 | 120.5339 | 121.911 | 122 | 147.3092 | 142.9789 | 135.3605 | 158.7753 | 88.822 | 239.089 | 246 | 273.733 | 259.5128 | 129.2459 | 128.8899 | 169.1101 | 131.6018 | 79.1358 | 70.4449 | 80.178 | 114.7798 | 199.911 | 203.0468 | 232.2202 | 149.8688 | 123.6229 | 137.2881 | 133.0211 | 101.6486 | 162.2247 | 178.5339 | 214.178 | 165.5083 | 126.644 | 131.8688 | 182.9532 | 140.9321 | 176.7119 | 185.3348 | 213.2459 | 232.3771 | 149.8899 | 171.8009 | 276.7964 | 210.7753 | 176.2926 | 191.7587 | 190 | 357.1946 | 164.644 | 182.1569 | 146.5807 | 223.2247 | 147.1734 | 145.6018 | 137.7541 | 166.5762 | 168.4238 | 151.5385 | 170.1101 | 137.4872 | 121.5551 | 114.8899 | 106.4449 | 117.6229 | 213.4706 | 261.0211 | 285.0422 | 383.822 | 241.644 | 242.5083 | 282.6229 | 350.822 | 136.4193 | 80.4449 | 69.4027 | 136.4193 | 106.9578 | 93.2459 | 105.9321 | 127.1312 | 116.8854 | 114.0211 | 102.8688 | 107.3771 | 75.356 | 70.4449 | 82.3982 | 100.2881 | 99.644 | 83.5128 | 102.3982 | 79.178 | 238.2881 | 220.6184 | 228.5128 | 305.4238 | 139.3137 | 217.0211 | 194 | 266.4917 | 294.3605 | 226.178 | 207.6018 | 342.7074
1311 | 16.4154 | 155.6705 | 213.8731 | 230.3038 | 184.6205 | 144.5821 | 174.6705 | 192.3795 | 162.3923 | 143.3038 | 133.6705 | 129.3923 | 146.5576 | 198.1898 | 143.6834 | 168.8731 | 223 | 197.9872 | 125.8359 | 130.241 | 188.3667 | 162.5949 | 350.0629 | 343.8603 | 181.8987 | 221.1898 | 241.1513 | 237.3423 | 283.3038 | 127.1141 | 120.3923 | 124.2282 | 130.4051 | 153.9115 | 175.4424 | 135.6962 | 88.8231 | 237.468 | 247.8231 | 275.4563 | 273.6834 | 141.291 | 136.4808 | 160.8102 | 143.1385 | 98.1641 | 73.6834 | 80.9872 | 132.5949 | 198.8987 | 220.8603 | 215.2154 | 166.7718 | 132.8359 | 131.4179 | 123.5064 | 137.7846 | 183.8859 | 183.5949 | 214.8859 | 141.7218 | 124.0128 | 152.2154 | 160.7846 | 130.2026 | 184.8102 | 195.6577 | 224.9872 | 227.0128 | 158.4936 | 177.2654 | 247.1397 | 187.291 | 210 | 216.5565 | 189.5949 | 307.0885 | 161.3038 | 190.6577 | 173.6077 | 243.468 | 112.0512 | 159.0629 | 127.4307 | 152.5064 | 182.4936 | 196.8859 | 164.5448 | 121.1898 | 115.6834 | 121.6705 | 113.2282 | 123.5949 | 246.3667 | 256.2666 | 262.0629 | 382.709 | 237.5949 | 222.9756 | 286.2654 | 351.0256 | 113.5448 | 83.9872 | 91.9744 | 113.5448 | 125.0757 | 106.6077 | 95.5064 | 107.7974 | 89.3551 | 105.8231 | 120.5821 | 100.3923 | 79.5064 | 73.9872 | 68.0256 | 93.1013 | 98.0256 | 97.8859 | 87.8231 | 81.709 | 230.1898 | 190.1525 | 240.9616 | 319.3923 | 160.9744 | 205.5821 | 197.1397 | 292 | 333.7334 | 228 | 225.7218 | 313.1525
1312 | 16.4264 | 138.975 | 195.795 | 231.5236 | 195.7064 | 131.7729 | 154.3407 | 176.8643 | 158.0221 | 145.5457 | 115.0443 | 123.0914 | 182.3436 | 191.6814 | 133.3629 | 152.1579 | 222.205 | 185.7064 | 147.5928 | 159.5679 | 164.7729 | 158.4321 | 324.795 | 317.9307 | 180.205 | 211.615 | 211.6564 | 262.23 | 288.1579 | 137.1136 | 116.2493 | 142.41 | 134.795 | 158.4793 | 136.0221 | 133.4543 | 104.3186 | 216.7036 | 265.7036 | 308.5264 | 265.5207 | 136.8171 | 123.9086 | 162.6621 | 104.0886 | 73 | 61.7729 | 70.8643 | 125.7064 | 199.3629 | 194.3629 | 225.5457 | 147.3407 | 156.6371 | 145.2493 | 130.2714 | 125.5236 | 172.1829 | 178.1829 | 203.7507 | 156.4321 | 134.3629 | 160.9557 | 151.0221 | 128.3657 | 189.615 | 172.8143 | 214.8643 | 236.5679 | 154 | 154.7036 | 268.4128 | 183.1579 | 207.615 | 188.6121 | 187.59 | 301.0443 | 164.6814 | 162.2493 | 176.1607 | 222.7036 | 144.2521 | 133.795 | 150.7036 | 154.1607 | 179.1357 | 186.3186 | 193.0471 | 115.7036 | 104.5679 | 104.18 | 132.7729 | 119.5457 | 222.6593 | 300.8421 | 238.4986 | 388.0914 | 238.77 | 270.7757 | 265.8614 | 367.7507 | 142.1607 | 74.3186 | 75.4764 | 142.1607 | 108.6814 | 110.0693 | 101.0221 | 106.795 | 125.3879 | 122.3407 | 109.1357 | 95.9086 | 82.5236 | 65.1136 | 83.2743 | 93.5457 | 111.23 | 90.9529 | 104.1136 | 87.3186 | 223.795 | 228.0221 | 216.1579 | 317.6343 | 144.1357 | 192.3186 | 223.6371 | 291.8864 | 291.5429 | 225.3879 | 237.82 | 349.4321
1313 | 16.4374 | 168.2379 | 202.7552 | 220 | 170.1259 | 129.4965 | 155.7414 | 169.5104 | 164.8672 | 143.2586 | 130.8741 | 116 | 145.8741 | 196.3707 | 147.5242 | 166.6086 | 216 | 162.6431 | 125.2793 | 163.3707 | 163.881 | 154 | 328.6086 | 328.4689 | 171.8603 | 201.6845 | 243.4689 | 230.2862 | 299.2103 | 136.4896 | 124.6155 | 129.119 | 139.4896 | 132.0069 | 144.881 | 137.6293 | 98.7483 | 228.7414 | 270.8189 | 256.6155 | 291.2379 | 148.7345 | 131.7552 | 126.538 | 131.7207 | 73.3776 | 61.8879 | 62.3776 | 101.3845 | 209.6224 | 201.4758 | 220.6155 | 151.2586 | 144.1121 | 152.231 | 146.1121 | 113.3707 | 163.1604 | 163.6224 | 194.6293 | 153.2586 | 145.1259 | 141.4758 | 157.9931 | 102.3914 | 169.8672 | 221.4689 | 206.881 | 240.3707 | 152.4896 | 167.2448 | 218.6362 | 201.3845 | 189.5035 | 204.5948 | 196.9793 | 316.2448 | 168.7414 | 175.3983 | 154.3776 | 234.3638 | 117.3914 | 139.2448 | 160.6017 | 134.0138 | 186.6155 | 182.3845 | 168.9931 | 127.7414 | 107.3638 | 126.6086 | 128.6086 | 119.9017 | 228.5587 | 283.6224 | 258.3776 | 377.4758 | 271.4689 | 231.7621 | 294.8672 | 354.8534 | 120.3776 | 68.4965 | 84.8603 | 120.3776 | 112.8672 | 93.9931 | 108.119 | 113.5035 | 102.1052 | 124.119 | 118.6293 | 106.5242 | 70.2448 | 67.6362 | 63.6362 | 89.119 | 74.7552 | 112.8672 | 101.7276 | 79.9862 | 229.119 | 234.7414 | 237.6569 | 340.2103 | 151.4896 | 188.3845 | 210.3569 | 291.1259 | 321.5879 | 208.2724 | 212.7483 | 342.9862
1314 | 16.4484 | 158.0723 | 208.9671 | 222.4868 | 171 | 127.2434 | 146.0329 | 177.0988 | 158.6578 | 151.5855 | 126.408 | 117.9342 | 140.5791 | 190.3422 | 167.4409 | 149.4474 | 213.6514 | 179.4803 | 157.7631 | 156.3751 | 172.0723 | 153.7237 | 313.3815 | 300.3157 | 155.6185 | 254.816 | 213.6578 | 261.7502 | 261.3486 | 123.5132 | 116.6578 | 123.9671 | 128.3092 | 139.5526 | 147.6843 | 140.6185 | 98.9342 | 220.5526 | 221.8748 | 244.7895 | 273.612 | 133.2763 | 136.1711 | 161.1317 | 105.4538 | 75.1711 | 74.8618 | 64.8618 | 111.1382 | 201.612 | 180.8683 | 217.2169 | 157.9277 | 133.6249 | 134.3815 | 133 | 110.9342 | 195.5462 | 164.0394 | 202.0394 | 162.2763 | 146 | 121.6972 | 149.3422 | 119.1382 | 160.6185 | 196.4932 | 212.7237 | 234.7566 | 140.204 | 163.1052 | 229.4474 | 211.9671 | 195.3486 | 178.5591 | 178.9012 | 310.0329 | 162.0723 | 194.0265 | 159.2105 | 225.3486 | 132.8948 | 137.3157 | 140.3157 | 146.7566 | 173.408 | 188.8225 | 163.2434 | 122.3157 | 96.2763 | 108.6185 | 111.0329 | 145.3422 | 278.4674 | 277.9606 | 261.4145 | 349.8225 | 243.592 | 242.8948 | 280.3686 | 334.3815 | 125.0723 | 65.4145 | 69.7237 | 125.0723 | 106.2434 | 89.3486 | 104.2105 | 116.0329 | 86.9735 | 116.0658 | 122.4474 | 127.2698 | 65.8289 | 77.6185 | 72.7895 | 85.0723 | 81.1052 | 103.6185 | 80.1052 | 66 | 228.9406 | 224.8948 | 269.4474 | 307.4603 | 142.3815 | 199.1052 | 193.6578 | 293.796 | 284.5526 | 229.6185 | 212.6578 | 331.2105
1315 | 16.4594 | 166.3362 | 212.4424 | 243.416 | 171.3009 | 139.3629 | 139.8496 | 161.2567 | 168.6991 | 147.4956 | 102.7964 | 130.3009 | 119.77 | 183.7611 | 142.7611 | 154.4249 | 202.7611 | 173.9116 | 172.5929 | 144.3629 | 186.2036 | 153.8053 | 325.9116 | 321.3009 | 148.9558 | 212.1062 | 227.4602 | 228.3982 | 285.0267 | 111.0089 | 124.4424 | 129.3982 | 121.7433 | 141.7964 | 127.7522 | 131.8496 | 108.4424 | 227.9116 | 270.8496 | 240.7789 | 246.3187 | 133.0442 | 131.9027 | 140.4602 | 125.0442 | 70.3009 | 73.3982 | 68.6638 | 113.5044 | 172.8142 | 202.8938 | 251.938 | 146.9558 | 148.4956 | 142.5487 | 134.6549 | 124.0531 | 171.8938 | 181.0442 | 219.3451 | 167.1593 | 147.5044 | 151.1504 | 141.708 | 120.7522 | 156.2125 | 229.4778 | 207.3893 | 226.8496 | 131.2567 | 166.9911 | 227.9558 | 214.5398 | 205.0353 | 201.4335 | 195.4956 | 299.7875 | 175 | 162.8496 | 172.8496 | 235.4867 | 128.1062 | 153.0353 | 157.0884 | 140.9558 | 152.354 | 172.4602 | 177.1682 | 138.6371 | 99.354 | 99.5487 | 105.9027 | 137.1062 | 220.5044 | 255.0884 | 261.1415 | 338.124 | 262.8407 | 243.3718 | 239.1504 | 342.5487 | 137.8496 | 67.8496 | 67.8496 | 137.8496 | 115.0531 | 106.708 | 119.5044 | 108.3451 | 114.0884 | 104 | 114.1858 | 107.4602 | 69.9469 | 67.4956 | 63.2125 | 97.3982 | 90.5576 | 92.7433 | 90.9116 | 66.6018 | 262.8407 | 221.9116 | 266.4513 | 346.0264 | 149.0442 | 208.5576 | 202.3451 | 304.6991 | 291.3098 | 220.8496 | 222.3982 | 340.4867
1316 | 16.4704 | 143.9057 | 200.6038 | 268.7925 | 172.6745 | 162.0472 | 137.5354 | 176.7665 | 163.5825 | 135.4198 | 98.1156 | 134.441 | 157.8373 | 204.8373 | 162.5354 | 195.0024 | 220.4198 | 189.816 | 154.9316 | 165.2571 | 196.092 | 161.559 | 349.7901 | 325.9292 | 158.5354 | 224.3255 | 249.1156 | 227.2783 | 324.6085 | 124.4198 | 113.092 | 126.6509 | 111.9292 | 134.5118 | 131.5118 | 132.1392 | 96.9292 | 251.6274 | 271.9528 | 294.6981 | 274.0259 | 124.441 | 133.7453 | 158.3726 | 117.092 | 71.5118 | 71.1392 | 88.9788 | 121.5118 | 200 | 190.5118 | 226.559 | 158.8137 | 140.9764 | 138.3726 | 143.8373 | 131.3019 | 159.1863 | 170.9764 | 209.1863 | 180.6061 | 156.3255 | 151.5118 | 154.559 | 125.1627 | 179.0236 | 191.0684 | 196.4387 | 225.0236 | 147.4175 | 151.4646 | 239 | 194.5118 | 176.5118 | 173.4175 | 192.3703 | 282.3467 | 172.0708 | 165.7429 | 171.6745 | 215.0448 | 138.5354 | 126.1392 | 137.7665 | 152.6509 | 157.3962 | 192.1627 | 208.0943 | 115.1392 | 108.1392 | 98.7901 | 115.3939 | 144.4434 | 230.3019 | 241.625 | 254.7642 | 376.5141 | 243.8608 | 277.8396 | 241.1392 | 342.1156 | 134.3962 | 67.6509 | 68.7901 | 134.3962 | 118.0708 | 120.0472 | 122.467 | 99.6509 | 96.0684 | 102.5354 | 91.8821 | 122.9316 | 66.2783 | 60.4646 | 85.0472 | 94.6509 | 102.2335 | 83.7429 | 110.559 | 69.6745 | 246.9528 | 237.6533 | 261.1887 | 299.3726 | 137.6745 | 216.3278 | 194.3019 | 303 | 308.5354 | 216.4198 | 220.6274 | 315.8137
1317 | 16.4814 | 165.7503 | 209.9744 | 240.5998 | 172.7503 | 153.9253 | 137.5262 | 188.0245 | 147.0502 | 127.3266 | 110.2252 | 139.8239 | 157.7001 | 203.6499 | 153.4248 | 162.6756 | 204.4504 | 159.7503 | 136.8004 | 145.2999 | 207.6243 | 153.3757 | 353.2241 | 340.6499 | 152.7503 | 225.1249 | 257.5496 | 238.2999 | 271.0758 | 109.6756 | 126.3746 | 129.1249 | 127.8751 | 135.8506 | 133.9008 | 139.9253 | 108.6745 | 256.3244 | 278.4995 | 285.5496 | 240.5262 | 138.0502 | 117.3501 | 150.8751 | 130.3746 | 69.175 | 72.9744 | 63.4258 | 121.8004 | 200.3501 | 193.6009 | 208.7492 | 163.175 | 146 | 134.7257 | 143.5251 | 134.6745 | 158.3255 | 176.7001 | 203.7748 | 157.1249 | 155.3746 | 147.7748 | 147.7759 | 128.6254 | 171.1996 | 210.2497 | 242.7246 | 220 | 166.2252 | 161.9755 | 239.175 | 197.6009 | 177.3255 | 188.0245 | 217.1996 | 317.6254 | 157 | 179.7492 | 170.3501 | 239.4248 | 129.7748 | 133.5752 | 154.8004 | 151.6243 | 151.0011 | 188.6243 | 186.7748 | 122.4002 | 112.9498 | 107.6499 | 132.2241 | 118.8997 | 238.0502 | 283.3746 | 296.9488 | 340.9755 | 238.7001 | 241.175 | 252.4258 | 357.3757 | 126.7759 | 69.2497 | 74.3244 | 126.7759 | 103.8751 | 108.4248 | 98.7257 | 103.7001 | 112.0992 | 96.4002 | 115.0747 | 99.8997 | 78.175 | 67.825 | 72.5496 | 94.4995 | 89.5752 | 101.7748 | 94.8495 | 66.9498 | 251.3991 | 210.2508 | 224.3255 | 293.1003 | 135.4749 | 181.4749 | 199.9498 | 301.5998 | 296.9744 | 196.5998 | 223.5742 | 319.4749
1318 | 16.4924 | 174.7493 | 195.121 | 229.879 | 181.9366 | 161.3141 | 172.6254 | 176.2478 | 151.6254 | 171.6917 | 114.8761 | 109.1844 | 162.4985 | 197.6917 | 144.3141 | 173.6312 | 217.4985 | 169.1239 | 149.8127 | 143.6859 | 188.3112 | 168.1931 | 329.9337 | 335.0663 | 163.2478 | 223.8098 | 244.8761 | 233.5015 | 288.8185 | 128.6859 | 112.3141 | 125.9366 | 131.2507 | 155.4985 | 154.5649 | 149.7493 | 95.4351 | 237.3141 | 264.9971 | 276.9971 | 269.2566 | 138.879 | 120.3112 | 154.6254 | 112.7551 | 69.6254 | 60.3688 | 82.0692 | 131.8156 | 202 | 213.8185 | 189.6195 | 160.441 | 145.0634 | 155.5015 | 143.9395 | 119.3746 | 173.2507 | 180 | 198.7493 | 159.9308 | 141.6888 | 142.1873 | 175 | 142.4985 | 164.3688 | 198.0663 | 232 | 223.3717 | 172.562 | 172.2536 | 235.6917 | 216.32 | 187.0058 | 178.4956 | 201.5649 | 332.2478 | 160.9337 | 155.3746 | 169.1902 | 235.1844 | 127.559 | 142.4985 | 161.2566 | 136.4322 | 177.8185 | 175.4927 | 183.9971 | 128.8127 | 108.562 | 107.6859 | 106.3112 | 110.4985 | 249.1815 | 265.1931 | 260.1239 | 353.5015 | 247.2449 | 246.121 | 275.5649 | 375.9395 | 146.6946 | 64.3717 | 62.2449 | 146.6946 | 106.1268 | 103.8098 | 121.562 | 110.1844 | 88.438 | 100.0029 | 105.2507 | 92.6225 | 77.5015 | 69.6225 | 66.4322 | 80.4351 | 98.3112 | 97.6859 | 82.4351 | 67.0576 | 223.1268 | 236.3112 | 239.2507 | 306.7464 | 134.4985 | 187.9913 | 195.7493 | 288.4439 | 275.3775 | 193.1844 | 203.8069 | 321.683
1319 | 16.5034 | 170.6145 | 209.8063 | 217.1898 | 182.0078 | 150.2074 | 170.002 | 185.2016 | 149.002 | 149.0078 | 112.1957 | 120.0059 | 168.8004 | 184.1898 | 137.1996 | 147.9961 | 219.0098 | 174.5988 | 149.7984 | 151.1996 | 194.7984 | 144.7965 | 341.6106 | 315.8043 | 174.5969 | 232.8063 | 240.1996 | 227.998 | 262 | 133.4051 | 108.1937 | 130.3992 | 131.1937 | 159.8043 | 146.9941 | 152.8004 | 103.2055 | 230.7984 | 281.9922 | 283.4129 | 245.9883 | 124.3934 | 120.6106 | 155.5949 | 97.1859 | 69.1977 | 85.6047 | 56.7906 | 113.6086 | 204.1957 | 190.9922 | 214.1996 | 149.7906 | 142.7965 | 150.9961 | 137.1957 | 121.7984 | 168.8023 | 183.1937 | 200.8023 | 183.2133 | 137.1977 | 137.6106 | 168.4129 | 144.8082 | 186.4109 | 180.2016 | 226.4109 | 237.6008 | 169.2113 | 154.0039 | 224.9844 | 182.4012 | 161 | 200.5949 | 201.3797 | 334.2152 | 170.4149 | 154.0059 | 157.3992 | 244.8082 | 140.6047 | 148.2016 | 130.2035 | 163.1937 | 153.7945 | 205.6047 | 194.0059 | 130.1957 | 112.7965 | 119.3914 | 111.002 | 117.998 | 267.0176 | 248.3836 | 258.4129 | 352.3894 | 259.4207 | 263.2016 | 262.0059 | 362.0098 | 117.3953 | 78.6008 | 83.2035 | 117.3953 | 98.998 | 116.1996 | 122.4031 | 119.409 | 94.3836 | 91.1918 | 104.3953 | 102.2035 | 71.5988 | 77.8063 | 86.0078 | 88.6047 | 100.6066 | 101.8063 | 93.1996 | 86.0059 | 224.5812 | 238.8063 | 234.6027 | 315.6145 | 142.9961 | 223.407 | 205.3875 | 262.9941 | 264 | 205.8023 | 226.6047 | 342.1996
1320 | 16.5144 | 145.6618 | 197.8476 | 235.881 | 172.357 | 136.9071 | 168.5428 | 178.1858 | 142.2453 | 134.4833 | 117.1858 | 114.9927 | 162.9144 | 200.5501 | 139.0595 | 156.4238 | 200.4833 | 175.5167 | 148.1263 | 150.9405 | 196.7286 | 153.0595 | 324.6618 | 314.4165 | 177.1858 | 215.1263 | 240.5762 | 229.4572 | 267.2975 | 124.0595 | 116.9739 | 130.3048 | 144.4238 | 155.238 | 152.0073 | 151.3643 | 90.7286 | 228.0668 | 291.2192 | 264.6284 | 269.3643 | 134.3977 | 105.5689 | 166.6357 | 121.5501 | 74.6357 | 75.7881 | 74.7286 | 97.2714 | 213.8476 | 203.3977 | 218.8142 | 168.5762 | 144.9144 | 158.5762 | 144.7286 | 122.2453 | 165.4833 | 193.881 | 197.6952 | 152.6096 | 143.0595 | 121.9332 | 143.4833 | 128.2119 | 168.8737 | 181.238 | 208.0261 | 242.5689 | 150.238 | 152.5689 | 248.9478 | 187.8403 | 169.2641 | 212.2714 | 230.4979 | 312.1117 | 139.1524 | 147.2975 | 160.6952 | 227.1524 | 136.7213 | 144.3643 | 125.7547 | 167.3121 | 172.2046 | 189.2192 | 181.5762 | 137.7286 | 121.4833 | 132.762 | 109.9666 | 123.4833 | 233.3309 | 285.8142 | 238.1451 | 366.1597 | 220.8142 | 257.4572 | 248.7286 | 341.7881 | 124.4572 | 77.8476 | 75.1524 | 124.4572 | 107.0261 | 116.3643 | 114.9405 | 103.119 | 122.7025 | 109.6952 | 112.0929 | 100.5094 | 71.881 | 65.8476 | 65.9739 | 82.6023 | 87.6357 | 90.4833 | 90.8215 | 72.7286 | 260.2453 | 228.9666 | 236.8403 | 289.8142 | 152.0595 | 210.2714 | 221.8883 | 275 | 272.6879 | 201.8476 | 216.1524 | 342.7881
1321 | 16.5254 | 158.7404 | 199.879 | 225.9823 | 192.6371 | 141.7581 | 174.4129 | 168.9322 | 138.5015 | 139.3274 | 107.0354 | 135.8436 | 145.7935 | 184.6726 | 149.7257 | 158 | 209.5871 | 170.8968 | 132.6903 | 142.5162 | 185.7226 | 156.5516 | 350.0709 | 336.171 | 160.5339 | 200.1387 | 238.7758 | 223.1387 | 278.4807 | 124.1887 | 108.5015 | 127.3629 | 140.3952 | 164.4985 | 127.7935 | 151.2419 | 84.6549 | 214.7435 | 261.9646 | 286.4985 | 266.1032 | 119.6549 | 127.7581 | 162.4985 | 106.3451 | 76.1032 | 70.9646 | 72.0177 | 104.9145 | 211.3952 | 196.0532 | 238.3806 | 166.7758 | 128.6903 | 159.9145 | 140 | 117.3806 | 168.3097 | 185.1032 | 200.6371 | 144.0354 | 144.5339 | 139.292 | 150.121 | 128.1032 | 187.4484 | 189.6017 | 224.121 | 268.3451 | 165.1032 | 170.9468 | 222.3629 | 209.8259 | 205.6048 | 214.5339 | 189.6048 | 336.584 | 143.1741 | 174.8467 | 167 | 226.0177 | 151.2743 | 142 | 136.4484 | 139.4838 | 193.171 | 166.0177 | 178.6549 | 130.7581 | 127.2242 | 113.6371 | 123.0177 | 127.2065 | 240.2065 | 291.0324 | 253.3775 | 347.537 | 236.5693 | 252.0354 | 251.1741 | 342.3451 | 110.5162 | 75.6194 | 69.5339 | 110.5162 | 119.3097 | 117.1387 | 107.6371 | 112.121 | 102.7758 | 111.9646 | 104.7758 | 114.9468 | 64.4484 | 72.1387 | 55.0354 | 95.3274 | 90.4484 | 96.2242 | 82.5871 | 68 | 243.2743 | 232.8259 | 261.9646 | 304.6726 | 154.4307 | 207.1533 | 197.8644 | 272.3097 | 300.292 | 204.7758 | 211.879 | 337.292
1322 | 16.5364 | 145.1824 | 196.473 | 219.7094 | 186.2026 | 136.3648 | 165.3108 | 182.9257 | 158.1622 | 138.1824 | 120.2906 | 120.6013 | 158.385 | 190.3108 | 170.1622 | 155.1622 | 218.2163 | 172.8176 | 148.6215 | 134.5674 | 167.7837 | 154.0541 | 366.6489 | 318.8378 | 150.1081 | 212.6554 | 236.1081 | 231.1622 | 259.8782 | 115.7296 | 124.8515 | 138.0541 | 121.7094 | 139.3987 | 137.7837 | 156.1622 | 79.2906 | 244.7635 | 250.1283 | 267.0743 | 265.1283 | 121.6215 | 124.0202 | 141.4189 | 112.1824 | 75.6013 | 60.7837 | 80.8919 | 112.3987 | 199.0945 | 211.5609 | 254.3987 | 163.8717 | 133.5068 | 164.0879 | 139.527 | 128.6691 | 156.3987 | 178.2163 | 194.2026 | 156.8176 | 137.1824 | 126.0743 | 160.1485 | 128.7837 | 188.2906 | 176.0606 | 222.7974 | 266.142 | 159.3987 | 153.0743 | 229.9798 | 194.4593 | 215.3042 | 203.8717 | 208.0541 | 304.7094 | 163.9798 | 193.2502 | 166.0541 | 235.6013 | 131.5472 | 145.5472 | 140.1283 | 150.2704 | 173.7094 | 175.3648 | 174.2365 | 130.3309 | 123.7433 | 102.2365 | 122.6691 | 122.8919 | 236.1283 | 257.1622 | 241.1889 | 382.5811 | 247.4391 | 265.0541 | 271.9798 | 346.0541 | 102.0945 | 63.1485 | 65.0202 | 102.0945 | 107.8717 | 122.5609 | 102.858 | 114.5811 | 106.0202 | 99.1824 | 102.3446 | 101.5674 | 67.8919 | 83.9459 | 68.2906 | 95.3648 | 94.8378 | 101.2567 | 93.5811 | 66.1081 | 223.5472 | 213.4391 | 248.2365 | 301.3244 | 150.1824 | 191.8241 | 225.615 | 264.1824 | 286.6013 | 206.6013 | 208.2365 | 321.9459
1323 | 16.5474 | 154.473 | 204.9657 | 231.4534 | 197.527 | 146.9853 | 173.0147 | 175.7365 | 149.9951 | 141.2537 | 116.7561 | 130.2439 | 171.0393 | 196.0245 | 161.2488 | 155.4534 | 207.9902 | 167.0098 | 160.5466 | 157.5074 | 174.7758 | 156.2242 | 327.761 | 334.7169 | 148.4779 | 209.5417 | 234.2291 | 220.5074 | 291.7659 | 121.5466 | 112.4632 | 133.5074 | 129.2242 | 136.9706 | 146.7659 | 150.7316 | 75.5074 | 247.4828 | 270.9313 | 268.7905 | 274.9853 | 140.761 | 137.4975 | 146.2488 | 121.2242 | 82.2586 | 73.9951 | 84.0147 | 105.2439 | 221.0049 | 199.7218 | 243.5123 | 157.9951 | 118.0049 | 143.2586 | 145.4632 | 105.2488 | 153.473 | 173.2144 | 215.4779 | 153.9951 | 143.9853 | 137.4926 | 174.527 | 140.7512 | 177.5417 | 210.2733 | 206 | 235.5417 | 155.9755 | 164.4926 | 219.9853 | 219.7758 | 188.9755 | 188.2929 | 207.239 | 299.7856 | 157.9657 | 168.7316 | 162.7512 | 242.7561 | 135.5368 | 154.5123 | 144.2635 | 136.9902 | 174.2586 | 180.761 | 171.2684 | 145.0442 | 112.9853 | 106.2341 | 104.473 | 129.7463 | 244.4926 | 259.4436 | 288.7414 | 364.5662 | 265.2488 | 257.5221 | 253.0294 | 341.5074 | 128.7316 | 78.5221 | 73.2733 | 128.7316 | 102.7414 | 104.7512 | 123.2586 | 109.5025 | 115.7659 | 104.9902 | 94.5123 | 124.0098 | 73.2537 | 82.5221 | 71.473 | 97.2782 | 97.2831 | 109.7758 | 91.239 | 64.2291 | 235.9951 | 226.0245 | 247.7561 | 289.4681 | 148.527 | 227.0147 | 197.288 | 267.0049 | 296.4926 | 209.2782 | 203.0294 | 325
1324 | 16.5584 | 170.7389 | 216.6017 | 253.2123 | 180.4779 | 162.7877 | 162.6947 | 184.2611 | 153.7832 | 144.7434 | 114.5664 | 131.9115 | 158.7479 | 183.3496 | 158.084 | 178.2566 | 217.1328 | 165.177 | 135.1328 | 149.6062 | 166.5709 | 169.1283 | 323.8717 | 342.73 | 160.1725 | 196.5311 | 244.6504 | 228.009 | 280.9557 | 99.7877 | 125.3406 | 126.6504 | 139.7787 | 151.3451 | 138.0443 | 161.2168 | 72.8275 | 255.1283 | 303.6017 | 248.6992 | 276.6902 | 138.6992 | 137.9557 | 155.3541 | 135.6947 | 77.7832 | 74.9115 | 74.4779 | 106.6504 | 213.5619 | 217.5221 | 235.7389 | 159.9557 | 120.2213 | 136.1725 | 156.2521 | 110.177 | 171.0443 | 186.3849 | 224.3008 | 159.3496 | 146.2123 | 142.7832 | 155.6504 | 136.8672 | 157.2213 | 194.9557 | 209.1328 | 208.4336 | 164.9955 | 166.9115 | 230.3053 | 201.9115 | 206.0885 | 172.9647 | 209.5619 | 280.6106 | 172.2123 | 176.084 | 159.3894 | 230.3849 | 121.3541 | 140.7345 | 138.6106 | 144.0443 | 172.1328 | 177.6549 | 165.7434 | 119.3053 | 118.8672 | 114.9557 | 119.9557 | 128.6062 | 245.8672 | 287.2123 | 286.6902 | 331.3098 | 256.6017 | 253.6594 | 247.5311 | 342.7434 | 140.7832 | 65 | 61.6106 | 140.7832 | 106.6947 | 110.0045 | 113.3008 | 110.3496 | 105.7389 | 104.2123 | 92.7434 | 115.3894 | 74.916 | 75.0045 | 83.5619 | 85.2213 | 82.4824 | 91.3894 | 97.4779 | 77.5221 | 228.5574 | 217.2655 | 249.7434 | 304.0398 | 142.7036 | 209.6459 | 183.3983 | 280.9692 | 298.3894 | 204.27 | 199.3585 | 326.8275
1325 | 16.5694 | 170.5467 | 199.2941 | 237.5536 | 184.4671 | 174.346 | 159 | 183.0865 | 154.6332 | 154.9862 | 115.7197 | 130.647 | 178.8789 | 191.9135 | 153.2872 | 162.5467 | 224.9066 | 175.6332 | 148.3737 | 146.0138 | 186.0865 | 158.8132 | 325.1661 | 320.301 | 153.6332 | 227.353 | 237.7266 | 255.9862 | 280.7336 | 117.3599 | 104.8339 | 121.6401 | 126.0934 | 140.1799 | 140.7266 | 148.6125 | 79.6401 | 239.0727 | 264.4257 | 264.7336 | 262.2872 | 147.353 | 127.8928 | 173.2595 | 137.4671 | 78.9066 | 72.827 | 71.6332 | 97.5398 | 202.3668 | 215.9931 | 223.5191 | 161.9204 | 133.9066 | 130.4533 | 126.5536 | 125.0069 | 169.0796 | 167.4671 | 212.6401 | 166 | 127.5467 | 139.2595 | 158.0207 | 127.4533 | 167.353 | 188.7197 | 215.5398 | 204.2734 | 149.9135 | 166.1938 | 231.8132 | 204.2007 | 211.7266 | 207.827 | 194.5398 | 287.4533 | 155.7336 | 170.7405 | 156.3737 | 215.8408 | 142.5398 | 127.3737 | 139.1661 | 152.7405 | 174.9862 | 186.3599 | 181.7266 | 117.5329 | 108.3599 | 116.9204 | 123.0138 | 119.5467 | 244.654 | 271.827 | 259.1661 | 357.4671 | 234.647 | 272.526 | 271.7924 | 354.353 | 134.526 | 65.5467 | 68.7266 | 134.526 | 104.6401 | 121.5329 | 106.0138 | 117.2734 | 99.5329 | 91.5605 | 107.6332 | 103.8997 | 79.9862 | 87.6263 | 74.2803 | 93.7128 | 99.0934 | 92.474 | 94.9066 | 79.2734 | 218.6816 | 232.5398 | 259.1661 | 292.7336 | 166.7128 | 196.3875 | 203.4395 | 312.5987 | 289.3599 | 239 | 240.8201 | 338.2872
1326 | 16.5804 | 173.9995 | 223.7144 | 247.2856 | 190.145 | 150.5722 | 164.4273 | 176.0005 | 148.1435 | 137.5718 | 108.9995 | 133.0024 | 147.4278 | 203.2842 | 170.8569 | 168.5703 | 213.145 | 175.4278 | 152.2875 | 163.4282 | 182.9995 | 152.1431 | 313.8555 | 337.8607 | 154.856 | 213.5718 | 237.8569 | 237.7148 | 284.0009 | 114.7139 | 120.4292 | 121.144 | 129.5713 | 141.9991 | 139.7144 | 128.4268 | 85.9995 | 221.1421 | 231.8569 | 276.5703 | 272.4301 | 134.4287 | 112.4278 | 152.5727 | 142.288 | 79.9991 | 80.4292 | 68.2856 | 93.5708 | 206.2856 | 217.4263 | 208.1393 | 172.7148 | 134.9991 | 131.2852 | 133.7148 | 130.7148 | 152.8579 | 175.4301 | 215.0005 | 160.2871 | 129 | 129.9986 | 178.8579 | 128.2852 | 159.5703 | 181.9995 | 208.1431 | 208.9991 | 158.7134 | 179.7144 | 226 | 221.144 | 202.145 | 217.7144 | 191.9991 | 291.1416 | 160.4292 | 181.4297 | 166.8574 | 242.5713 | 132.5722 | 140.4282 | 127.5699 | 166.8574 | 162.9991 | 180.8574 | 178.1435 | 110.7125 | 107.1421 | 122.5732 | 133.8584 | 123.5708 | 267.8565 | 281.7144 | 249.855 | 369.4292 | 244.7148 | 254.2847 | 245.9991 | 344.2852 | 113.9995 | 72.7129 | 69.7139 | 113.9995 | 109.2856 | 107.5713 | 119.7148 | 119.7139 | 93.5694 | 109 | 108.2847 | 98.4273 | 64.2856 | 71.7153 | 81.8574 | 82.9986 | 102.8569 | 110.4297 | 90.5718 | 77.7148 | 264.4301 | 223.1435 | 240.4287 | 308.2837 | 155.4273 | 217.1445 | 181.4297 | 271.5708 | 295.8546 | 245.8555 | 229.8603 | 345.288
1327 | 16.5914 | 172.7433 | 222.7021 | 240.2536 | 171.5753 | 146.1711 | 177.7021 | 175.767 | 143.0856 | 141.7464 | 116.3835 | 110.9794 | 156.2979 | 218.1062 | 167.9351 | 174.3392 | 190.7021 | 178.6371 | 148.1948 | 169.6814 | 182.3392 | 149.4041 | 326.233 | 315.7907 | 170.2773 | 209.7021 | 242.0856 | 235.5753 | 271.0206 | 116.6165 | 112.8082 | 110.7876 | 134.2773 | 155.5753 | 136.6165 | 148.1918 | 86.8289 | 241.8113 | 253.9618 | 282.0412 | 256.5959 | 131.2124 | 123.6814 | 142.2773 | 116.5103 | 87.0206 | 72.8082 | 69.2979 | 100 | 211.767 | 232.0619 | 235.5722 | 165.2124 | 141.4247 | 142.0649 | 133.9588 | 125.2979 | 153.0443 | 160.7876 | 208.2124 | 146 | 129.8938 | 145.2979 | 162.3598 | 131.6165 | 165.0412 | 183.4247 | 210.1711 | 216.6165 | 166.2124 | 176.0206 | 221.5309 | 204.5309 | 180 | 211.6371 | 199.6165 | 308.767 | 157.5753 | 175.0649 | 178.5989 | 238.0412 | 128.1711 | 140.7227 | 150.5546 | 157.1474 | 164.0619 | 177.1062 | 179.3423 | 130 | 115.5959 | 107.8732 | 120.7021 | 127.9144 | 274.1062 | 283.3835 | 268.4247 | 366.8732 | 247.3423 | 266.4897 | 257.7876 | 338.9763 | 127.6402 | 76.8701 | 71.0206 | 127.6402 | 113.5753 | 114.0649 | 111.3186 | 116.8495 | 112.6165 | 108.4041 | 114.5309 | 110.2124 | 69.767 | 63.6814 | 72.4454 | 92.9351 | 104.1062 | 96.6165 | 93.8526 | 76.4691 | 248.5959 | 225.534 | 234.8289 | 321.5516 | 163.6371 | 205.6608 | 179.534 | 272.6371 | 311.0825 | 247.5072 | 207.1299 | 321
1328 | 16.6024 | 157.6898 | 216.1057 | 217.967 | 176.8977 | 161.2739 | 177.9307 | 184.5182 | 148.3102 | 160.3102 | 117.6568 | 120.7921 | 159.4818 | 213.8284 | 160.6534 | 157.2773 | 193.4125 | 165.3795 | 179.3466 | 161.109 | 173.3432 | 147.3795 | 314.3795 | 352.0693 | 175.5182 | 205.2773 | 251.6534 | 245.2409 | 274.5478 | 115.9637 | 113.7227 | 118.4125 | 146.033 | 158.1057 | 144.3398 | 147.2773 | 81.9637 | 256.7988 | 289.2113 | 278.858 | 254.8977 | 127 | 131.8614 | 145.3466 | 114.8614 | 79.3795 | 67.8284 | 69.0693 | 106.2046 | 219.2773 | 208.2079 | 221.858 | 163.1716 | 130.8284 | 153.6898 | 145.2079 | 123.2079 | 172.3102 | 170.5841 | 203.0693 | 149.7227 | 137.2739 | 146 | 146.6534 | 125.3795 | 151.9307 | 177.7921 | 214.4159 | 217.5148 | 159.2079 | 166.8284 | 210.3795 | 202.6864 | 174.7261 | 197.7591 | 197.1023 | 322.1716 | 160.4159 | 187.9307 | 193.4556 | 226.1716 | 136.4489 | 134.2409 | 156.6931 | 147.5148 | 152.3069 | 175 | 193.1057 | 125.0363 | 113.8977 | 113.5875 | 118.7591 | 121.7591 | 267.967 | 284.6568 | 272.7194 | 382.2046 | 258.9341 | 271.2409 | 263.2409 | 317.6534 | 145.5182 | 56.4125 | 69.5841 | 145.5182 | 115.4852 | 121.3466 | 107.1716 | 105.2739 | 106.3795 | 107 | 104.9307 | 112.5148 | 83.1716 | 67.2079 | 64.7557 | 84.4125 | 99.8284 | 92.8614 | 107.5182 | 83.2773 | 249.3795 | 248 | 230.5841 | 298.1057 | 151.6205 | 224.2739 | 199.8284 | 265.2739 | 287.6534 | 217.5148 | 227.8977 | 318.5182
1329 | 16.6134 | 155.3874 | 193.9695 | 217.708 | 170.9676 | 167.8397 | 173.8721 | 177.7099 | 151.9027 | 164.2252 | 108.3225 | 128.6126 | 162.0973 | 213.5153 | 170.355 | 157.3855 | 203.9027 | 162.3874 | 160.292 | 135.0954 | 177.8397 | 154.708 | 316.5477 | 348.7099 | 168.3874 | 202.4828 | 252.9695 | 250.5802 | 289.292 | 128.2901 | 121.355 | 124.3874 | 152.0973 | 143.7099 | 165.292 | 148.6756 | 90.4198 | 220.0973 | 270.0954 | 284.6164 | 247.355 | 127.645 | 135.0324 | 137.9027 | 121.9027 | 77.6775 | 69.7729 | 67 | 119.0324 | 208.7424 | 202.9676 | 242.4523 | 166.3874 | 129.5477 | 151.3874 | 145.1279 | 116.0324 | 171.3874 | 191.063 | 202.9351 | 157.0324 | 144.4847 | 143.0973 | 157.9676 | 126.5802 | 151.0973 | 189.1603 | 208.1279 | 230.0649 | 157.5153 | 167.1603 | 212.2252 | 213.6145 | 172.3531 | 201.7729 | 197.8721 | 319.2595 | 144.1298 | 185.1622 | 152.2576 | 227.4523 | 133 | 129.9046 | 145.9981 | 160.7099 | 164.292 | 175.9676 | 183.8702 | 115.2901 | 109.2576 | 108.2576 | 116.9676 | 126.7405 | 243.5191 | 285.3206 | 295.8397 | 389.2271 | 237.7748 | 265.292 | 267.9351 | 329.6126 | 138.7099 | 60.4523 | 74.9046 | 138.7099 | 103.9027 | 110.355 | 114.9027 | 114.7424 | 102.0973 | 107.9676 | 102.4847 | 120.2271 | 81.2271 | 60.6775 | 82.645 | 92.645 | 95.3225 | 97.3225 | 96.5172 | 75.3225 | 245.4198 | 234.4542 | 242.0324 | 289.5153 | 148.4847 | 234.0649 | 198.5477 | 278.6126 | 293.4847 | 220.7118 | 216.1622 | 321.063
1330 | 16.6244 | 148.9867 | 189.3969 | 233.3215 | 173 | 160.0133 | 164.6829 | 169.643 | 161.3481 | 162.2949 | 118.7095 | 134.3437 | 159.0133 | 216.9734 | 160.9645 | 168.6386 | 212.6785 | 168.0399 | 149.7006 | 153.3481 | 167 | 177.3525 | 338.062 | 342.9867 | 160.9823 | 213.6652 | 238.6873 | 263.3658 | 285.062 | 129.3259 | 119.6652 | 125.6873 | 137.2949 | 138.6563 | 145.3259 | 169 | 91.3614 | 217.0133 | 286.6741 | 254.7051 | 261.0665 | 128.6652 | 128.6475 | 148.6873 | 130.6785 | 73.9867 | 77.9734 | 71.3525 | 113.3171 | 201.3215 | 202.3215 | 238.0133 | 154.6297 | 135.9956 | 149.0044 | 151.9734 | 112.9956 | 172.3525 | 208.3348 | 210.0133 | 162.7006 | 133.6608 | 144.031 | 163.0133 | 134.3437 | 158.0576 | 203.0133 | 225.6873 | 214.6164 | 167 | 169.6297 | 220.3392 | 195.3304 | 187.3127 | 213.6563 | 198.7272 | 303.3348 | 145.3747 | 168.3038 | 170.0576 | 224.6873 | 140.031 | 122.3658 | 159.969 | 156.3259 | 158.0532 | 183.6918 | 182.2905 | 125.3658 | 113.6652 | 115.0089 | 119.6696 | 134.6297 | 226.0665 | 304.9911 | 271.2728 | 370.6475 | 230.6652 | 239.969 | 256.5987 | 339.0266 | 126.9601 | 55.0089 | 57.9911 | 126.9601 | 103.9734 | 116.3658 | 113.6342 | 105.9823 | 99.0133 | 109.6652 | 95.6785 | 107.6741 | 70.0133 | 65.0222 | 87.3481 | 87.3038 | 98.0089 | 100.3437 | 94.7095 | 75.3304 | 232.969 | 225.0842 | 238.9956 | 307.7051 | 148.0399 | 224.9778 | 207.0044 | 278.9867 | 280.3171 | 206.7095 | 211.357 | 347.3747
1331 | 16.6354 | 146.1239 | 214.529 | 228.6942 | 179.5949 | 160.7935 | 172.3471 | 161.0826 | 155.1572 | 148.471 | 128.2812 | 133.7935 | 165.6942 | 206.0413 | 153.6775 | 162.5949 | 211.058 | 179.058 | 165.6942 | 162.7768 | 174.9833 | 181.4877 | 371.5536 | 338.3884 | 153.6942 | 213.3471 | 258.4131 | 268.2812 | 297.0333 | 122.529 | 115.8761 | 133.2645 | 118.6116 | 130.8761 | 139.5703 | 158.2398 | 99.4877 | 230.2891 | 297.2891 | 258.1159 | 269.4797 | 133.5536 | 119.3058 | 155.2232 | 134.6116 | 63.4877 | 70.1652 | 81.7355 | 107.3884 | 197.0413 | 206.7188 | 233.9341 | 139.3884 | 134.3471 | 147.1819 | 148.3304 | 108.5703 | 181.6942 | 194.0746 | 215.6529 | 170.3638 | 138.6362 | 159.7355 | 165.529 | 139 | 185.7355 | 201.0167 | 224.2398 | 204.8428 | 162.8348 | 155.7768 | 238.6196 | 196.7768 | 182.1652 | 204.4877 | 212.1159 | 305.7355 | 159.4877 | 159.1652 | 187.6696 | 217.6862 | 147.058 | 139.7768 | 148.0826 | 150.5703 | 181.6529 | 181.463 | 173.0659 | 144.1652 | 113 | 118.3058 | 113.0167 | 110.7109 | 246.6283 | 299.9587 | 253.7188 | 356.0993 | 227.5703 | 236.7602 | 247.5203 | 349.6116 | 114.5536 | 61.7768 | 60.2478 | 114.5536 | 94.7768 | 129.2645 | 110.8015 | 109.1073 | 100.4877 | 109.3471 | 110.0246 | 114.8181 | 74.6116 | 71.1819 | 90.1819 | 81.2891 | 107.2065 | 106.7355 | 103.9341 | 79.5536 | 217.2645 | 255.0167 | 230.7522 | 321.8761 | 163.5703 | 224.0246 | 213.1652 | 273.3471 | 277.8594 | 221.1406 | 225.8181 | 349.3391
1332 | 16.6464 | 149.1249 | 201.8903 | 236.8285 | 189.1249 | 145.6091 | 170.4843 | 177.9381 | 133.5157 | 142.4213 | 111.4376 | 125.4376 | 166.6406 | 209.4376 | 167.8436 | 175.3594 | 204.8285 | 176.4224 | 168.0782 | 174.1097 | 194.6721 | 160.7806 | 370.14 | 341.3594 | 159.6721 | 215.0782 | 258.2334 | 250 | 265.9218 | 114.5624 | 107.1249 | 131.0782 | 121.7503 | 130.3909 | 136.4376 | 142.3127 | 87.0467 | 230.7025 | 295.5461 | 247.8133 | 232.2345 | 137.5309 | 132.7351 | 156.4691 | 124.0934 | 61.8285 | 76.5624 | 82.8436 | 113.2345 | 201.1564 | 219.6091 | 223.266 | 141.6406 | 137.5157 | 144.3127 | 161.4843 | 105.4376 | 175.4528 | 172.8285 | 215 | 149.5309 | 138.9837 | 159.0467 | 165.1097 | 135.7655 | 183.2497 | 195.3442 | 203.6406 | 228.2812 | 152.4843 | 163.5157 | 256.14 | 208.8285 | 179.5776 | 192.4061 | 189.2345 | 300.0152 | 158.5472 | 161.2497 | 174.5157 | 207.126 | 135.0782 | 141.4061 | 153.4376 | 151.7503 | 179.9218 | 166.063 | 192 | 146.6091 | 118.0315 | 120.2345 | 102.6721 | 103.8285 | 243.7351 | 290.8121 | 274.1564 | 363.329 | 229.1097 | 259.8751 | 282.9685 | 342.6873 | 123.9218 | 64.4843 | 78.1097 | 123.9218 | 106.1097 | 128.8751 | 131.203 | 121.0152 | 98.4691 | 106.4061 | 122.3279 | 117.6873 | 72 | 71.5472 | 81.5624 | 94.6406 | 112.3279 | 102.4528 | 92.1879 | 81.734 | 232.6884 | 246.4691 | 223.3127 | 317.4376 | 152.5309 | 228.7806 | 215.9685 | 272.9218 | 284.1249 | 212.3594 | 231.9218 | 340.6417
1333 | 16.6574 | 143.6283 | 201.4071 | 247.5133 | 191.4337 | 144.177 | 159.3097 | 197.6548 | 138 | 133.7523 | 112.8849 | 126.1416 | 157.4513 | 204.938 | 152.4778 | 169.6814 | 218.4867 | 190.9115 | 157.3627 | 179.7964 | 197.0531 | 142.3452 | 348.2566 | 326.3893 | 161.6814 | 222.9469 | 225.2301 | 259.2921 | 284.0709 | 120.5487 | 100.8849 | 131.8849 | 138.3186 | 135.9115 | 141.2301 | 146.6548 | 85.2036 | 214.062 | 259.823 | 280.0885 | 260.301 | 125.6283 | 141.9025 | 169.1151 | 114.4602 | 74.3717 | 74.7434 | 77.8849 | 112.6814 | 206.1151 | 219.2921 | 230.5663 | 146.5753 | 142 | 153.1151 | 150.6814 | 112.8319 | 155.6814 | 186.8584 | 213.1416 | 147.2921 | 118.0885 | 143.4513 | 164.4778 | 133.3452 | 171.1416 | 203.4513 | 210.062 | 226.2566 | 145.3982 | 171.7168 | 232.0265 | 216.5398 | 175.8674 | 186 | 197.9735 | 289.5753 | 164.5663 | 149.5133 | 179 | 229.9911 | 136.6283 | 142.062 | 160.0885 | 161.6283 | 173.9115 | 178.0531 | 191.2566 | 130.6814 | 122.9115 | 117.4513 | 108.7699 | 118.6018 | 259.9645 | 279.4867 | 270.5663 | 382.9911 | 237.7699 | 265 | 279.2036 | 339.4602 | 123.1151 | 63.7168 | 84.1681 | 123.1151 | 108.823 | 126.1946 | 129.8584 | 107.9735 | 98.1061 | 100.7434 | 114 | 109.6283 | 76.0885 | 81.2832 | 76.0265 | 90.2832 | 98.7964 | 87.1416 | 99.7964 | 71.4602 | 243.699 | 250.1946 | 234.7168 | 314.4247 | 151.0355 | 207.3717 | 205.5133 | 279.9204 | 293.8674 | 212.6283 | 228.8849 | 360.5574
1334 | 16.6684 | 141.0801 | 214.0083 | 245.384 | 195.9365 | 147.6327 | 159.9032 | 188.1603 | 138.384 | 152.7763 | 115.2237 | 120.3122 | 147.6077 | 194.9199 | 144.9032 | 166.2955 | 206.409 | 176.7045 | 145.5994 | 155.641 | 188.9199 | 148.7679 | 353.9115 | 315.743 | 142.5526 | 228.7763 | 230.5359 | 260.409 | 303.7929 | 125.0167 | 108.5994 | 131.5359 | 134.409 | 127.8481 | 142.3122 | 143.6878 | 84.7846 | 233.2955 | 261.743 | 278.1686 | 283.025 | 131.1436 | 122.5994 | 157.1769 | 118.9282 | 76.9199 | 76.7679 | 67.1686 | 111.9833 | 201.8564 | 229.2404 | 212.2404 | 156 | 139.3122 | 158.4558 | 134.6244 | 121.384 | 154.5994 | 193.4558 | 214.2237 | 166.8397 | 133.0635 | 140.1353 | 149.9917 | 134.3923 | 163.7763 | 184.3923 | 231.2154 | 219.6244 | 142.9199 | 179.5359 | 233.1436 | 204.7763 | 187.9449 | 189.8397 | 194.5526 | 303.9917 | 158.9115 | 150.0718 | 180.1519 | 209.1519 | 132.1603 | 151.6962 | 158.5526 | 151.4006 | 168.5359 | 178.7513 | 187.6962 | 120.384 | 111.0083 | 109.1436 | 110.7596 | 121.4641 | 241.0801 | 278.5442 | 267.9833 | 358.6962 | 247.4391 | 264.2321 | 279.1686 | 352.7596 | 129.9917 | 74.2237 | 72.9282 | 129.9917 | 95.384 | 111.0801 | 126.4724 | 116.8397 | 85.2955 | 97.0083 | 113.2321 | 105.9282 | 74.9365 | 78.4558 | 75.2237 | 91.2955 | 96.5276 | 89.3756 | 89.4641 | 76.6962 | 213.6795 | 238.1519 | 237.5442 | 302.6962 | 160.7045 | 201.0885 | 211.4474 | 295 | 309.4006 | 233.1186 | 222.3923 | 326.7679
1335 | 16.6794 | 154.2469 | 212.3403 | 238.0747 | 185.3776 | 141.6992 | 166.3195 | 188.3403 | 147.7179 | 144.0187 | 122 | 124.3216 | 156.1888 | 192.4523 | 162.8112 | 170.4709 | 188.5477 | 167.3776 | 157.8299 | 129.4149 | 192 | 163.0768 | 363.8112 | 344.1701 | 141.6805 | 229.9253 | 231.4149 | 235.0187 | 279.4149 | 108.2075 | 125.1888 | 135.5851 | 110.6037 | 132.7365 | 141.9627 | 157.9066 | 73.9627 | 242.2261 | 285.4149 | 271.9253 | 260.9813 | 135.4523 | 128.0934 | 140.944 | 115.1888 | 72.8672 | 74.4336 | 64.0954 | 123.6411 | 193.1888 | 218.0187 | 216.473 | 158.3776 | 130.2448 | 160.0373 | 124.0187 | 116.8485 | 162.8672 | 183.9419 | 237.2469 | 155.5643 | 141.2448 | 151.6597 | 168.3029 | 128.1888 | 159.7739 | 182.944 | 247.1888 | 223.2842 | 147.5851 | 178.4336 | 253.6992 | 196.0187 | 182.4544 | 198.3776 | 186.1515 | 306.8485 | 167.6224 | 177.9834 | 180.8112 | 216.944 | 137.888 | 150.7739 | 146.9813 | 145.9066 | 176.1515 | 187.6784 | 177.2635 | 132.0954 | 110.1328 | 122.5664 | 121.5851 | 123.3589 | 251.8693 | 275.3776 | 262.2054 | 355.3963 | 252.7531 | 257.8485 | 254.6971 | 365.1701 | 128.4896 | 79.8112 | 64.4336 | 128.4896 | 101.1515 | 99.6784 | 112.4336 | 114.2821 | 97 | 90.1888 | 116.3589 | 116.8506 | 65.5664 | 80.4336 | 79.2261 | 96.6597 | 103.8299 | 99.1888 | 87.7925 | 75.3776 | 230.7552 | 240.3963 | 246.2656 | 300.5851 | 147.4149 | 191.9813 | 226.9813 | 289.056 | 296.7739 | 245.9979 | 218.1701 | 345.832
1336 | 16.6904 | 167.969 | 211.7861 | 226 | 187.3658 | 153.5059 | 154.354 | 184.1091 | 163.0428 | 136.9144 | 119.9572 | 141.9027 | 155.9572 | 193.3968 | 158.5487 | 159 | 206.8054 | 165.2802 | 162.3968 | 135.9883 | 189.5487 | 166.2492 | 351.7861 | 333.0663 | 162.6342 | 225.6577 | 229.8171 | 254.8792 | 276.1829 | 111.0856 | 131.4941 | 130.2374 | 118.9883 | 136.4631 | 141.4631 | 166.4631 | 84.9027 | 236.3658 | 276.4631 | 290.5369 | 265.2257 | 127.8171 | 113.8171 | 153.2685 | 113.323 | 72.2139 | 69 | 71.6032 | 114.9572 | 201.1284 | 218.677 | 237 | 166.0856 | 127.9027 | 151.9572 | 138.568 | 111.0428 | 152.4086 | 165.9027 | 243.2064 | 132.2685 | 134.8171 | 146.0856 | 170.5176 | 133.2685 | 165.6342 | 205.8909 | 238.7861 | 242 | 149.5914 | 171.3658 | 251.2064 | 202.8054 | 201.0973 | 206.0856 | 200.1284 | 307.9883 | 156.646 | 198.7005 | 182.677 | 230.9027 | 150.6888 | 155.8171 | 158.1711 | 145.4748 | 182.3658 | 170.9144 | 167.4086 | 137.969 | 118.5487 | 124.7315 | 120.323 | 123.4631 | 250.1208 | 263.8834 | 243.2995 | 353.5487 | 244.7508 | 243.0545 | 243.9027 | 356.9262 | 127.4824 | 74.7315 | 71.6653 | 127.4824 | 100.4203 | 98.031 | 110.2685 | 101.8171 | 95.3658 | 98.9455 | 119.323 | 125.8834 | 69.7743 | 77.8599 | 78.2685 | 92.7198 | 90.4203 | 98.1401 | 95.1284 | 79.8171 | 236.7743 | 232.0117 | 254.9262 | 315.6653 | 146.6342 | 197.0428 | 219.3775 | 278.7743 | 294.0545 | 226.6653 | 229.1284 | 360.7433
1337 | 16.7014 | 152 | 212.1505 | 232.7336 | 180.7915 | 157.5212 | 166.6834 | 163.1583 | 154.2163 | 130.1583 | 113.6332 | 144.7915 | 157.2085 | 207 | 151.2124 | 160.2625 | 203.3244 | 164.1544 | 161.2703 | 152.7876 | 182.6332 | 150.8378 | 344.9961 | 331.0424 | 165.8417 | 208.3127 | 230.5792 | 265.2742 | 278.3668 | 114.4749 | 129.1622 | 120.6834 | 132.8417 | 132.471 | 124.8456 | 148.583 | 82.7413 | 229.7915 | 271.2085 | 289.058 | 264.8958 | 128.5792 | 112.0541 | 158 | 114.3127 | 80.6873 | 69.4208 | 60.1042 | 109.8958 | 217.5753 | 226.9459 | 232.7915 | 159.3746 | 130.3707 | 139.3205 | 167.7876 | 122.417 | 164.7837 | 181.8378 | 226.1004 | 142.8919 | 143.5753 | 145.6873 | 154.7876 | 131.2664 | 167.5792 | 209.1661 | 226.9459 | 232.7413 | 162.888 | 159.3205 | 232.417 | 207.3205 | 194 | 220.8378 | 204.7915 | 313.8456 | 150.6293 | 189.3629 | 181.6873 | 228.7413 | 134.583 | 153.6332 | 165.7915 | 143.2007 | 170.7413 | 167.9459 | 178.5212 | 127.0502 | 116.2625 | 122.5251 | 105.7413 | 123.2587 | 227.5212 | 264.7258 | 269.8919 | 357.9961 | 246.2742 | 238.8919 | 243.4247 | 343.8919 | 139.0039 | 67.4749 | 77.3746 | 139.0039 | 88 | 112.3166 | 112.8958 | 109.7336 | 101.417 | 107.7375 | 108.1081 | 101.4749 | 73.471 | 73.1622 | 77.529 | 91.7413 | 81.3668 | 90.6332 | 99.7915 | 85.6293 | 240.8919 | 237.5173 | 235.1583 | 330.2124 | 152.3668 | 200 | 204.8417 | 278.6834 | 284 | 235.3205 | 225.7954 | 360.6756
1338 | 16.7124 | 152.4331 | 224.7704 | 233.7704 | 165.4709 | 162.4709 | 170.2994 | 163.7326 | 147.9622 | 138.9622 | 113.7645 | 126.439 | 153.9041 | 210.8983 | 159.4273 | 152.4709 | 185.0639 | 170.968 | 153.561 | 151.9361 | 181.8983 | 155.6366 | 344.3052 | 348.3052 | 156.1715 | 213.1017 | 236.0639 | 237.032 | 266.5407 | 112.7326 | 111.8023 | 126.032 | 129.6686 | 140.8663 | 129.7267 | 134.4012 | 76.9302 | 229.6308 | 277.4331 | 287.8605 | 255.9361 | 127.5669 | 123.1599 | 159.2994 | 123.8663 | 76.3314 | 67.4012 | 67.3314 | 107 | 206.7762 | 225.3692 | 243.4593 | 145.8983 | 133.9622 | 135.9622 | 153.0756 | 121.0058 | 171.9041 | 171.4768 | 234.8023 | 150.5669 | 142.7384 | 140.4651 | 156.5349 | 125.032 | 173.4971 | 197.6948 | 223.6366 | 210.4709 | 172.4709 | 148.5988 | 246.5988 | 199.1977 | 201.3634 | 223.9041 | 199 | 290.5727 | 166.5291 | 192.0058 | 173 | 226.3953 | 129.4971 | 142.5029 | 146.1395 | 170.5669 | 145.8721 | 170.2674 | 187.3692 | 129.2355 | 114.968 | 116.9041 | 92.1337 | 126.0378 | 253.6948 | 310.7267 | 271.9361 | 383.2936 | 233.625 | 252.6308 | 229.968 | 350.2674 | 121.0698 | 65.2994 | 72.1279 | 121.0698 | 97.9622 | 109.5669 | 116.0639 | 120.2994 | 101.7384 | 99.9361 | 98.7006 | 101.032 | 75.8023 | 61.4331 | 70.8663 | 87.2296 | 85.5669 | 93.7965 | 90.968 | 84.6366 | 235.5727 | 249.1395 | 255.2238 | 307.6744 | 155.7006 | 200 | 215.0959 | 271.9041 | 298.7267 | 223.7326 | 229.3575 | 367.6744
1339 | 16.7234 | 154.3363 | 212.2183 | 233.6903 | 165.9175 | 155.3452 | 174.6726 | 175.354 | 159.6637 | 146.2094 | 135.1446 | 105.9911 | 157.5901 | 204.8642 | 156.9556 | 148.0177 | 194.3363 | 170.118 | 162.4277 | 143.5546 | 181.2094 | 148.2272 | 346.3718 | 323.646 | 154.0266 | 211.5634 | 234.646 | 238.7728 | 261.9263 | 114.5546 | 109.0089 | 121.0914 | 130.6815 | 137.9911 | 142.764 | 143.9175 | 86 | 239.2272 | 281.118 | 295.7551 | 259.5812 | 122.1003 | 141.2183 | 168.1269 | 114.7551 | 76.2094 | 75.1358 | 71.6637 | 105.2183 | 185.236 | 218 | 249.8554 | 150.1091 | 142.1003 | 157.0177 | 135.4632 | 109.3452 | 167.1269 | 159.8172 | 223.5457 | 148.2183 | 132.0089 | 144.5546 | 156.0089 | 132.118 | 172.2005 | 215.6726 | 215.3363 | 210.9175 | 164.0089 | 139.9734 | 262.0266 | 205.1091 | 199.8642 | 204.8731 | 203.0089 | 284.1358 | 177.2183 | 180.7906 | 171.2183 | 234.6548 | 142.0089 | 132.6637 | 142.6992 | 161.5368 | 142.472 | 172.4543 | 167.9734 | 127.4543 | 112.7817 | 107.2272 | 107.4809 | 124.5812 | 255.1917 | 313.5191 | 278.2538 | 397.5457 | 251.764 | 261.7817 | 252.2804 | 349.7817 | 116.8997 | 68.3363 | 84.4366 | 116.8997 | 109.6637 | 114.7906 | 123.1091 | 116.2094 | 97.6903 | 95.5634 | 109.0266 | 109.8997 | 78.7995 | 64.2272 | 73.7817 | 95.7728 | 95.2449 | 105.7817 | 89.6726 | 83.9086 | 225.5723 | 232.3363 | 273.2005 | 304.1535 | 154.4454 | 202.2272 | 232.3452 | 272.0266 | 324.236 | 218.4277 | 251.5457 | 344.1091
1340 | 16.7344 | 149.1342 | 204.0496 | 240.5919 | 173.3033 | 157.8805 | 171.5919 | 188.8309 | 171.7316 | 141.2886 | 133.4025 | 112.9008 | 167.3879 | 230.364 | 137.4927 | 164.4081 | 191.4228 | 172.6268 | 158.4927 | 153.9853 | 187.7316 | 143.2187 | 368.9154 | 327.1048 | 165.3382 | 194.0294 | 222.5423 | 254.7666 | 292.1544 | 114.9154 | 108.0496 | 120.0699 | 129.8456 | 130.2537 | 134.5423 | 149.4724 | 88.2886 | 233.3033 | 278.5919 | 270.6415 | 265.761 | 124.6967 | 134.8805 | 157.318 | 112.9853 | 70.3732 | 84.4228 | 75.4927 | 109.4081 | 210.3934 | 208.3879 | 228.7114 | 146.2537 | 132.3382 | 151.8952 | 146.9154 | 114.1691 | 167.3033 | 171.4724 | 217.0846 | 158.8162 | 134.7114 | 149.0349 | 164.204 | 147.443 | 168.239 | 219 | 216.5423 | 226.0846 | 151.6066 | 145.5975 | 257.7757 | 209.4927 | 186 | 197.8309 | 211.6618 | 305.7813 | 173.6268 | 173.3529 | 175.4081 | 214.2684 | 149.7463 | 131.9154 | 156.4228 | 156.9504 | 152.9651 | 165.6415 | 169.478 | 124.761 | 118.6618 | 116.4081 | 127.0846 | 136.7114 | 252.1894 | 295.2886 | 272.6562 | 374.1489 | 249.9504 | 263.5423 | 274.4724 | 362.0699 | 121.6268 | 72.2886 | 79.0846 | 121.6268 | 113.0349 | 117.8805 | 109.6415 | 116.7813 | 110.5423 | 109.6121 | 113.0147 | 107.3033 | 76.8106 | 77.9853 | 77.3732 | 90.7114 | 105.2537 | 98.3879 | 96.6618 | 87.6765 | 228.5919 | 237.204 | 244.522 | 321.1342 | 155.4577 | 209.5772 | 233.5073 | 286.0846 | 315.9798 | 206.7114 | 228.1489 | 341.1691
1341 | 16.7454 | 145.2301 | 220.9706 | 230.6499 | 171.4602 | 152.06 | 172.9302 | 183.4798 | 176.7197 | 147.2901 | 124.9804 | 124.6499 | 148.4798 | 229.9988 | 157.1603 | 179.5202 | 192.5802 | 176.6401 | 157.0098 | 159.0098 | 185.6695 | 141.5202 | 364.3599 | 355.3403 | 158.6499 | 200.9009 | 231.4002 | 258.1995 | 287.4896 | 120.2301 | 103.8201 | 127.2999 | 126.0502 | 126.53 | 132.59 | 146.4002 | 92.88 | 223.47 | 264.4198 | 265.8703 | 252.7099 | 137.8201 | 132.8201 | 139.6401 | 116.1297 | 77.6401 | 73.8299 | 74.9498 | 119.3501 | 211.7295 | 195.12 | 234.4602 | 148.6401 | 125.6499 | 147.1603 | 148.47 | 111.06 | 152.2999 | 167.9302 | 235.2705 | 167.4198 | 130.12 | 152.1799 | 174.1102 | 151.1297 | 176.59 | 208.6597 | 215.53 | 232.0502 | 134.7601 | 158.2497 | 223.2497 | 212.2399 | 191.6401 | 197.6499 | 203.7797 | 307.4798 | 166.3599 | 164.9302 | 183.94 | 207.4002 | 131.3794 | 137.7001 | 148.6499 | 142.8494 | 147 | 156.6401 | 175.8397 | 115.47 | 116.8898 | 130.5802 | 122.2399 | 134 | 271.3501 | 290.9498 | 283.071 | 361.9302 | 248.0698 | 262.06 | 271.8703 | 377.7601 | 125.1701 | 78.7601 | 72.8299 | 125.1701 | 108.6597 | 107.8299 | 106.2803 | 124.12 | 104.8299 | 117.7099 | 111.7503 | 102.1701 | 66.1102 | 86.7699 | 72.8898 | 103.9804 | 103.88 | 90.7601 | 103.8201 | 77.59 | 243.5606 | 243.8201 | 244.7405 | 318.1701 | 150.3599 | 204.1897 | 211.9596 | 291.1102 | 300.2901 | 216.2203 | 215.9302 | 330.5398
1342 | 16.7564 | 147.1062 | 234.9778 | 234.2699 | 173.2831 | 144.7301 | 181.0708 | 174.5177 | 155.8008 | 156.7876 | 120.3317 | 123.9292 | 146.2699 | 185.1992 | 168.6239 | 165.814 | 204.3407 | 173.8362 | 151.3761 | 152.8938 | 180.916 | 150 | 357.5177 | 360.2478 | 150.6947 | 225.4601 | 233.3185 | 248.5753 | 279.1638 | 122.5885 | 119.0576 | 122.4823 | 130.6239 | 119.7301 | 131.9646 | 146.3893 | 94.5177 | 224.4823 | 259.4115 | 282.7876 | 252.3761 | 130.8716 | 127.3185 | 143.5885 | 117.5399 | 75.8008 | 68.4823 | 63.1416 | 110.4247 | 192.4115 | 200.2345 | 237.2478 | 161.7522 | 132.6461 | 149.9424 | 149 | 113.3761 | 164.3629 | 181.8584 | 238.1902 | 155.177 | 129.9292 | 147.5531 | 171.8362 | 153.5045 | 168.7301 | 213.8806 | 209.6947 | 243.8584 | 149.6107 | 158.5045 | 228.8098 | 209.9292 | 188.8362 | 208.9868 | 194.3407 | 309.6107 | 167.2345 | 164.3893 | 173.907 | 204.0132 | 117.1284 | 130.4601 | 147.9292 | 132.0222 | 166.2921 | 167.823 | 182.7036 | 112.1416 | 120.1284 | 120.1548 | 120.4115 | 141.2345 | 251.814 | 291.1992 | 283.075 | 360.4247 | 248.1284 | 245.084 | 277.2124 | 344.3803 | 137.7522 | 76.7301 | 66.5177 | 137.7522 | 110.5045 | 105.3761 | 110.3185 | 112.3539 | 97.0708 | 114 | 116.8008 | 110.4115 | 71.5531 | 78.6239 | 71.7876 | 108.0132 | 102.6239 | 92.1062 | 101.2124 | 81.7876 | 243.9202 | 237.8362 | 256.3893 | 316.7655 | 143.0354 | 200.6815 | 195.4823 | 298.4823 | 316.5399 | 220.3539 | 225.5177 | 321.4823
1343 | 16.7674 | 144.4946 | 211.1081 | 229.1297 | 173.4081 | 137.5054 | 186.9135 | 174.4946 | 150.9568 | 157.5595 | 107.3757 | 129.4622 | 159.4297 | 196.473 | 157.0865 | 158.3216 | 215.4297 | 166.9676 | 148.0757 | 163.4189 | 186.1297 | 148.0216 | 361.4513 | 365.3649 | 151.9243 | 220.4081 | 224.4946 | 242.2054 | 273.173 | 118.0216 | 132.4946 | 112.6135 | 126.0108 | 135.2567 | 148.827 | 145.8811 | 86.5811 | 225.4946 | 280.3 | 265.2595 | 249.5703 | 121.4838 | 128.8811 | 146.4405 | 128.0324 | 65.5162 | 71.473 | 66.9135 | 98.4946 | 218.2459 | 214.4297 | 216.646 | 159.6027 | 133.1081 | 152.3324 | 132.6784 | 118.9784 | 174.0973 | 185.5054 | 223.9351 | 152.473 | 128.5378 | 161.3324 | 169.4189 | 150.6892 | 158.5378 | 218.1514 | 215.8703 | 226.7324 | 154.0757 | 163.1081 | 241.5811 | 200.1297 | 203.2351 | 219.0541 | 200.9784 | 305.173 | 169.0649 | 165.8595 | 161.0108 | 209.773 | 133.9351 | 141.2351 | 157.9135 | 156.3865 | 168.7 | 169.5378 | 195.1189 | 106.0216 | 132.9784 | 123.7513 | 113.1081 | 149 | 236.9027 | 298.0216 | 243.4297 | 351.4622 | 265.4297 | 234.4297 | 276.4405 | 312.9027 | 141.5378 | 72.9676 | 73.4189 | 141.5378 | 114.6135 | 111.9676 | 110.8919 | 105.9568 | 103.9027 | 105.5919 | 109.9784 | 111.0216 | 70 | 76.9784 | 80.9676 | 91.0216 | 90.5919 | 96.4189 | 95.9892 | 84.5378 | 227.3865 | 228 | 244.5054 | 300.5919 | 141.0108 | 211.9784 | 199.9568 | 298.0108 | 308.7324 | 213.4622 | 218.5703 | 334.8595
1344 | 16.7784 | 151.0826 | 209.1102 | 223.0826 | 191.6308 | 139.5344 | 178.7797 | 180.5757 | 151.4518 | 149.9725 | 115.4518 | 137.0551 | 154.3416 | 200.5207 | 151.0275 | 176.5207 | 212.3692 | 178.1102 | 151.1377 | 160.3554 | 170.9587 | 147.5207 | 352.3141 | 365.8347 | 157.9862 | 218.8623 | 248.3167 | 245.8098 | 263.5895 | 106.8761 | 134.5207 | 116.1928 | 124.4931 | 148.3692 | 155.8761 | 148.8898 | 86.0964 | 218.3967 | 287.8485 | 263.303 | 240.9725 | 123 | 130.3692 | 150.4793 | 121.4518 | 73.1239 | 76.5344 | 67.3967 | 118.2616 | 228.3141 | 214.9174 | 234.9338 | 146.4518 | 127.562 | 154.3141 | 124.1102 | 121 | 153.3416 | 195.1377 | 231.0138 | 153.4793 | 130.0689 | 159.259 | 182.0689 | 137.562 | 162.6033 | 218.6997 | 226.9862 | 219.1928 | 151.0551 | 157.0551 | 228.4243 | 202.6997 | 202.1488 | 206.3967 | 216.6859 | 291.5344 | 156.9174 | 168.8623 | 176.2203 | 219.3416 | 133.9174 | 144.7108 | 160.4243 | 142.6695 | 165.7134 | 177.6584 | 173.8623 | 118.1928 | 116.2452 | 137.3692 | 103 | 143.9311 | 243.4656 | 298.5344 | 264.6997 | 362.0964 | 249.1902 | 253.6721 | 257.1626 | 343.7961 | 143.0689 | 74.4793 | 74.4105 | 143.0689 | 100.4518 | 98.2728 | 114.9174 | 117.0964 | 100.8347 | 102.5757 | 113.5207 | 114.5757 | 69.4931 | 69.3692 | 85.5207 | 90.5207 | 84.5344 | 95.3829 | 91.9311 | 88.0964 | 243.5757 | 237.1239 | 254.6446 | 305.179 | 145.5757 | 212.9862 | 206.0275 | 279.7659 | 309.8098 | 212.9449 | 221.1239 | 336.3416
1345 | 16.7894 | 148.6932 | 216.4808 | 239.3835 | 189.5781 | 142 | 171.3068 | 183.4041 | 144.8849 | 158.9027 | 124.9794 | 143.0767 | 151.826 | 192.6548 | 155.0767 | 169.174 | 209.7493 | 183.4041 | 169.826 | 140.1356 | 169.5575 | 145.3658 | 357.6904 | 344.6548 | 161.6726 | 224.0561 | 255.4247 | 263.3274 | 285.5753 | 124.478 | 131.3274 | 121.174 | 123.4808 | 145.2301 | 154.7876 | 146.1918 | 88.8466 | 221.9027 | 254.6754 | 272.5398 | 248.3452 | 133.3835 | 138.652 | 142.7699 | 126.3068 | 77.8466 | 79 | 64.1534 | 123.5014 | 204.6165 | 221.4602 | 255.0973 | 144.0384 | 130.9616 | 135.8082 | 132 | 128.2685 | 154.9794 | 185.2713 | 223.6932 | 161.8644 | 136.5575 | 150.3068 | 182.3274 | 136.8082 | 189.7287 | 239.7493 | 234.3068 | 226.2507 | 147.2124 | 161 | 221.9616 | 208.174 | 188.4219 | 199.5192 | 212.348 | 310.6136 | 157.2301 | 177.6904 | 187.3274 | 227.2095 | 135.7876 | 143.7287 | 163.826 | 134.5753 | 176.8466 | 181.7315 | 174.3835 | 138.7493 | 109.9411 | 127.4808 | 100.4041 | 134.8466 | 225.9439 | 271.9261 | 262.3864 | 379.3835 | 248.2862 | 261.3274 | 249.0945 | 349.9439 | 143.3274 | 67.8082 | 70.5959 | 143.3274 | 105.3068 | 97.0561 | 119.9027 | 126.0767 | 89 | 100.2124 | 112.9233 | 118.9616 | 75.2301 | 69.8644 | 83.3658 | 92 | 101.0178 | 91.6726 | 98.4219 | 89.8082 | 245.3658 | 239.2507 | 271.2301 | 318 | 148.4041 | 213.0384 | 213.7109 | 270.7876 | 303.4453 | 213.6726 | 220.6548 | 339.0178
1346 | 16.8004 | 144.1888 | 215.4685 | 232.5245 | 176.8741 | 137.2168 | 165.7133 | 176.2168 | 153.1608 | 155.1818 | 134.3427 | 147.6573 | 155.7483 | 184 | 155.4056 | 177.4755 | 225.0349 | 190.0349 | 163.6504 | 158.1048 | 189.6014 | 156.8811 | 363.8392 | 353.5384 | 153.7762 | 218.8672 | 231.4336 | 251.5595 | 280.8672 | 135.1818 | 128.0629 | 118.3147 | 128.3147 | 135.5874 | 154.5595 | 154.1888 | 77.028 | 232.5315 | 255.2587 | 265.2517 | 240.5245 | 140.8741 | 147.028 | 137 | 128.6853 | 74 | 82.1888 | 73.8461 | 108.8741 | 198.1888 | 229.2797 | 240.7483 | 151.9091 | 126.8112 | 145.3496 | 149.5384 | 125.4336 | 161.6853 | 182.4126 | 221.8461 | 164.6224 | 137.4685 | 149.4965 | 154.6154 | 134.6573 | 196.8392 | 211.9861 | 233.4965 | 223.7203 | 153.8182 | 177.4755 | 239.6014 | 202.1259 | 210.6923 | 200.5315 | 188.5595 | 314.8392 | 159.8112 | 190.2168 | 181.9371 | 231.7133 | 135.028 | 140.2098 | 146.4896 | 146.8741 | 179.3776 | 173.4056 | 193.5664 | 126.3986 | 117.8112 | 128.0629 | 101.1888 | 130.4685 | 240.6992 | 268.916 | 251.7832 | 351.7973 | 254.7133 | 266.0349 | 254.4336 | 347.1608 | 133.1539 | 66.7203 | 71.4056 | 133.1539 | 114.0629 | 116.8461 | 124.1539 | 130.6573 | 94.8461 | 93.5315 | 100.3706 | 108.9651 | 84.1888 | 67.8392 | 86.3776 | 78.7133 | 104.9651 | 92.2797 | 114.3147 | 82.8741 | 242 | 228.2168 | 269.028 | 323.3147 | 148.1259 | 219.3147 | 218.4685 | 282.2517 | 287.6294 | 227.0349 | 218.9091 | 338.7133
1347 | 16.8114 | 143.7374 | 224.7877 | 225.0689 | 191.3128 | 137.3501 | 162.1564 | 169.2812 | 153.2123 | 145.7188 | 132 | 123.8996 | 153.0875 | 187.2626 | 147.4749 | 174.5997 | 217.231 | 166.462 | 157.5251 | 165.1435 | 190.231 | 161.8436 | 375.7503 | 367.9125 | 169.1004 | 213.4189 | 221.3687 | 255.8752 | 279.2252 | 119.1937 | 132.8063 | 128.9813 | 134.0875 | 135.0502 | 144.7374 | 162.9813 | 78.1564 | 246.0502 | 284.3314 | 278.7877 | 237.4189 | 133.5624 | 133.4749 | 142.4376 | 133.7877 | 74.5438 | 75.2123 | 73.5624 | 119.5064 | 199.9125 | 222.7374 | 233.1937 | 162.8938 | 121.825 | 147.1248 | 155.7561 | 117 | 162.4376 | 190.0187 | 222.6499 | 148.1248 | 133.7374 | 144.7188 | 147.0502 | 130.4749 | 176.1248 | 196.1377 | 228.175 | 223.7374 | 144.2497 | 180.0373 | 251.1062 | 201.2812 | 203.5997 | 199.3687 | 198.8565 | 309.3501 | 147.2123 | 182.7374 | 186.4376 | 218.9125 | 125.825 | 128.0689 | 129.4376 | 137.9311 | 177.9311 | 166.5624 | 182.4247 | 124.6815 | 103.0373 | 124.6499 | 111.6126 | 132.7188 | 268.7188 | 304.6815 | 247.2999 | 340.2066 | 258.769 | 260.9498 | 257.9627 | 339.056 | 123.1062 | 75.4376 | 72.7188 | 123.1062 | 112.2812 | 114.9311 | 120.0875 | 122.6686 | 94.0187 | 95.6313 | 100.2439 | 105.3501 | 77.7561 | 68.3314 | 84.3874 | 79.5064 | 97 | 78.6686 | 107.5811 | 77.7374 | 233.2999 | 244.6628 | 250.0373 | 317.6686 | 152.175 | 224.5438 | 202.231 | 286.5438 | 288.2999 | 226.8436 | 229.3501 | 343.8565
1348 | 16.8224 | 151.0089 | 239.6726 | 223.7714 | 190.5427 | 147.6726 | 170.1121 | 188.1299 | 131.6548 | 146.3319 | 123.1032 | 115.9012 | 160.6726 | 188.8879 | 152.5649 | 154.4395 | 200.1077 | 158.3496 | 154.1032 | 155.4484 | 181.4484 | 151.1077 | 355.6326 | 373.1165 | 180.4351 | 206.3185 | 233.3452 | 282.7936 | 275.3185 | 122.6726 | 130.9956 | 126.2153 | 123.879 | 130.9867 | 128.6548 | 184.6815 | 79.9956 | 226.4173 | 286.3274 | 278.6593 | 241.9956 | 126.7758 | 134.1165 | 159.7891 | 138.6637 | 72.2198 | 64.2198 | 64.5516 | 118.879 | 193.4395 | 225.5605 | 218.879 | 169.2242 | 123.8923 | 136.3319 | 149.6681 | 130.9012 | 155.879 | 188.8923 | 229.0089 | 139.556 | 129.8879 | 138.6593 | 165.2286 | 130.5605 | 167.556 | 213.5605 | 245.5694 | 228.7847 | 146.5738 | 171.6681 | 234.7669 | 213.4573 | 189 | 201.3363 | 196.8746 | 305.2153 | 146.2286 | 190.5649 | 192.1121 | 220.7802 | 139.0133 | 150.1254 | 135.1121 | 149.2375 | 174.2242 | 179.2375 | 176.0089 | 133.6637 | 103.0089 | 127.1165 | 116.3319 | 143.3407 | 268.2198 | 307.5472 | 249.4528 | 339.6504 | 265.3274 | 241.6593 | 253.5427 | 333.4528 | 124.5605 | 71.1032 | 66.1032 | 124.5605 | 103.8835 | 124.0133 | 127.6726 | 116.7802 | 93.4484 | 98.1121 | 106.8879 | 100.6593 | 68.8879 | 70.8835 | 70.7714 | 90.556 | 107.0089 | 72.2242 | 87.4351 | 77.7802 | 220.4395 | 241.4262 | 246.1165 | 326.7936 | 139.5427 | 209.4306 | 204.5694 | 287 | 278.7758 | 241.1299 | 246.9012 | 349.6593
1349 | 16.8334 | 140.2448 | 227.9498 | 206.0648 | 171.0432 | 157.5467 | 176.1151 | 198.7483 | 126.1151 | 143.295 | 122.2517 | 133.8201 | 152.3598 | 200.5035 | 154.1799 | 153.41 | 187.3382 | 178.1151 | 142.4533 | 157.8633 | 179.8849 | 161.0718 | 355.6904 | 351.8563 | 158.9282 | 204.9352 | 240.0216 | 281.0865 | 266.5467 | 120.6116 | 113.9282 | 113.1799 | 117.8417 | 137.1869 | 132.7769 | 164.4679 | 73.1583 | 237.2371 | 275.8849 | 270.8633 | 231.1799 | 128.9784 | 131.0432 | 150.6764 | 125.2015 | 77.3884 | 64.2734 | 73.5035 | 100.9065 | 199.2301 | 233.41 | 227.0502 | 168.7266 | 130.9784 | 138.41 | 144.1799 | 120.4031 | 162.9135 | 189.1583 | 240.9784 | 155.9135 | 124.4533 | 146.2085 | 168.1583 | 135.5683 | 173.1151 | 218 | 229.0146 | 229.3166 | 157.3166 | 173.5683 | 235.7985 | 205.6764 | 194.6834 | 209.6834 | 194.7985 | 308.0935 | 170.1869 | 183.6548 | 185.6116 | 212.7699 | 145.8849 | 147.6548 | 129.7483 | 157.3166 | 172.59 | 181.0648 | 178.8849 | 126.4533 | 125.2085 | 139.9568 | 111.0216 | 129.5397 | 256.3382 | 305.1151 | 275.1869 | 347.7336 | 245.7915 | 237.8417 | 231.2015 | 350.0935 | 125.59 | 76.5035 | 61.2734 | 125.59 | 97.8633 | 116.1081 | 119.9282 | 104.2231 | 110.6402 | 98.4317 | 110.5467 | 101.9568 | 77.0935 | 61.4533 | 72.9568 | 96.6834 | 107.0432 | 84.7985 | 88.0935 | 77.7266 | 230.7769 | 231.2517 | 257.8201 | 343.8417 | 140.5035 | 204.9568 | 205.0648 | 273.9282 | 304.8487 | 239.8131 | 240.3814 | 340.1583
1350 | 16.8444 | 148.6446 | 216.1613 | 220.2253 | 166.7419 | 141.8392 | 182.9032 | 201.5482 | 139.8707 | 152.4514 | 122.3549 | 133.7743 | 158.8383 | 190.8392 | 157.7095 | 161.8063 | 193.3545 | 180.2581 | 150.032 | 145.968 | 180.0644 | 157.0648 | 367.968 | 338.3869 | 149 | 213.4194 | 235.8387 | 266.6775 | 282.7739 | 104.5486 | 120.2577 | 127.1608 | 142.064 | 136.4842 | 139.9356 | 153.5158 | 82.6126 | 239.0328 | 272 | 274.0644 | 231.2257 | 129.0968 | 119.7743 | 148.1933 | 113.5162 | 82.4194 | 61.9356 | 64.4198 | 100.3869 | 188.4198 | 224.968 | 233.6131 | 148.4198 | 129.9356 | 151.4514 | 128.5486 | 115.6126 | 183.7095 | 179.4518 | 252.7095 | 161.6131 | 130.2901 | 151.7743 | 170.0644 | 120.3229 | 166.5486 | 212.1937 | 220.3545 | 235.4514 | 155.9032 | 178.0644 | 236.4518 | 205.5158 | 187.968 | 207.6131 | 204.1613 | 303.968 | 169.4842 | 170.8387 | 186.9676 | 205 | 141.4194 | 131.9356 | 148.2253 | 144.8712 | 174.0644 | 187.0964 | 181.3869 | 133.4514 | 123.2261 | 140.1937 | 117.8707 | 126.7739 | 242.0324 | 305.5162 | 265.7747 | 341.7747 | 255.3221 | 280.0636 | 256.6766 | 343.6455 | 120.0968 | 78.4518 | 63.5806 | 120.0968 | 103.9676 | 107.9676 | 123.9352 | 100.5482 | 101.8392 | 99.7419 | 111.6775 | 95.2261 | 77.6131 | 70.7739 | 72.0324 | 98.0968 | 92.2905 | 91.8387 | 89.7743 | 81.8063 | 223.4198 | 229.6131 | 259.5162 | 315.8071 | 146.5162 | 217.9676 | 220.3865 | 279.6771 | 301.6135 | 223.6775 | 221.7743 | 354.8383
1351 | 16.8554 | 144.8406 | 212.8495 | 239.3717 | 164.4425 | 130.9646 | 185 | 195.1415 | 149.9646 | 157.6283 | 133.2301 | 123.4778 | 149.6548 | 190.0443 | 164 | 156.5132 | 208.9292 | 171.8849 | 154.4425 | 164.0886 | 189.5222 | 147.5929 | 355.2566 | 332.3274 | 155.5222 | 210.0354 | 232.6283 | 246.0265 | 290.2212 | 117.7523 | 130.2212 | 120.2477 | 147.9203 | 132.5222 | 134.0354 | 158.0708 | 84.4778 | 225.1505 | 272 | 281.7434 | 233.2212 | 127 | 108.292 | 143.3628 | 119.3009 | 79.0354 | 77.3806 | 63.4514 | 106.7788 | 203.4603 | 203.9557 | 245.6018 | 151.0089 | 138.2655 | 158.4071 | 138.7877 | 125.1858 | 174.584 | 175.1858 | 221.6459 | 152.8495 | 123.9557 | 158.6726 | 163.5132 | 138.0532 | 167.8938 | 230.5311 | 220.5132 | 245.9646 | 157.4071 | 181 | 252.3452 | 209.4778 | 172.292 | 195.292 | 205 | 306.0797 | 161.9646 | 165.8495 | 175.9911 | 217.4514 | 140.4071 | 144.416 | 167.3717 | 141.9646 | 161.584 | 190.292 | 183.6283 | 141 | 109.8495 | 116.4335 | 113.7345 | 125.3274 | 263.0886 | 307.7434 | 266.7877 | 346.3452 | 261.6991 | 268.6812 | 268.5486 | 322.7345 | 126.8938 | 71.8142 | 65.1858 | 126.8938 | 98.9203 | 99.9557 | 125.1062 | 102.4425 | 102.8231 | 117.0089 | 99.3274 | 99.6372 | 70.0354 | 72.292 | 65.8142 | 91.2566 | 88.3717 | 98.1151 | 94.3009 | 83.6283 | 239.6461 | 233.8938 | 248.1062 | 321.8674 | 138.6637 | 208.1769 | 207.168 | 288.0354 | 276.8849 | 256.3895 | 241.124 | 365.8142
1352 | 16.8664 | 154.6033 | 210 | 232.527 | 172.2886 | 137.2365 | 181.3687 | 190.0261 | 149.5791 | 148.1322 | 136.3426 | 118.3948 | 136.2104 | 206.8678 | 173.6834 | 163.3147 | 211.7896 | 179.1043 | 144.7374 | 176.5531 | 190.3687 | 149.2104 | 369.5512 | 357.0763 | 155.1583 | 221.9199 | 250.9721 | 234.8156 | 247.8455 | 132 | 123.5531 | 135.0503 | 150.6834 | 138.2104 | 128.3687 | 171.5512 | 80.6052 | 231.6313 | 262.9218 | 300.7356 | 260.4451 | 131.8417 | 120.5512 | 140.6574 | 123.1844 | 75.1844 | 74.8697 | 77.4469 | 115.2626 | 214.527 | 214.3668 | 248.527 | 151.1062 | 136.3166 | 162.2365 | 141.9218 | 127.2104 | 174.2886 | 189.9199 | 219.6033 | 145.1583 | 123.473 | 154.5009 | 173.3408 | 141.054 | 166.7374 | 245.3948 | 249.7076 | 239.527 | 157.6052 | 174.9479 | 264.5791 | 195.9218 | 188.1825 | 205.1303 | 201.9739 | 296.6592 | 167.0261 | 168.4469 | 168.0261 | 218.7374 | 136.9739 | 145.3166 | 158.1062 | 141.5791 | 168.5512 | 186.8156 | 194.7095 | 122.2384 | 120.9199 | 106.6313 | 115.0782 | 113.1583 | 251.3445 | 311.6052 | 278.3948 | 362.2104 | 250.5531 | 259.1564 | 273.9199 | 326.4991 | 123.9218 | 70.3948 | 78.7095 | 123.9218 | 95.0261 | 97.6574 | 123.2365 | 112.7095 | 107.5531 | 120.7374 | 95.6313 | 108.3948 | 60.1322 | 63.3687 | 61.3687 | 86.1844 | 97.8678 | 94.527 | 94.5531 | 87.4469 | 252.1322 | 261.7877 | 255.9199 | 329.5009 | 141.6574 | 214.3408 | 209.6033 | 293.2626 | 287.7356 | 257.6071 | 239.4488 | 331.108
1353 | 16.8774 | 165.6775 | 220.4976 | 227 | 174.6775 | 149.2626 | 187.6451 | 195.705 | 163.4376 | 150.4101 | 120.8849 | 137.1427 | 141.3225 | 208.9125 | 178.1475 | 169.53 | 213.47 | 171.5624 | 149.8801 | 161.2674 | 184.9125 | 150.6175 | 383.47 | 359.4749 | 157.5575 | 210.6223 | 239.2998 | 248.9676 | 276.5754 | 114.7099 | 117.53 | 126.8298 | 144.0324 | 136.53 | 137.7326 | 163.8573 | 82.8525 | 258.0827 | 277.9952 | 302.3549 | 283.3225 | 136.8525 | 116.5624 | 139.4425 | 109.0324 | 69.6775 | 74.8801 | 79.765 | 104.5624 | 200.9724 | 228.8525 | 258.4376 | 155.7326 | 135.5575 | 154.5024 | 138.47 | 132.94 | 169.8849 | 205.175 | 222.6499 | 149.4101 | 126.53 | 149.47 | 173.5024 | 155.9352 | 166.94 | 216.5948 | 242.7423 | 221.0324 | 150.5899 | 181.4976 | 246.9449 | 210.9952 | 202.6175 | 220.5575 | 222.2302 | 309.4652 | 162.2074 | 172 | 179.8801 | 227.5851 | 129.4425 | 138.3825 | 136.7974 | 149.8801 | 159.6223 | 191.705 | 183.8573 | 111.8525 | 131.235 | 104.06 | 127.7926 | 117.4101 | 251.7602 | 313.8525 | 256.3873 | 365.47 | 255.0276 | 256.1799 | 288.5575 | 350.0551 | 132.2026 | 68.765 | 74.0324 | 132.2026 | 101.3225 | 97.6775 | 118.5899 | 132.7326 | 116.3501 | 101.7974 | 118.3777 | 106.1475 | 64.8801 | 69.0276 | 62.0875 | 88.0875 | 92.5024 | 82.2074 | 87.9125 | 91.6175 | 255.0276 | 251.9173 | 247.7099 | 298.5348 | 154.0276 | 228.0875 | 200.9173 | 285.6499 | 294.9125 | 231.2674 | 218.7374 | 334.3177
1354 | 16.8884 | 168.0384 | 223.2212 | 234.5576 | 168.2212 | 149.3318 | 166.5486 | 195.4808 | 164.8127 | 143.6637 | 131.6343 | 133.2551 | 147.149 | 203.2212 | 165.6637 | 168 | 198.6253 | 183.4131 | 155.3702 | 154.6298 | 167.2551 | 148.4808 | 383.7404 | 341.0722 | 167.298 | 211.2257 | 240.3747 | 263.298 | 277.6208 | 119.1151 | 121.6682 | 131.4131 | 124.6637 | 140.0384 | 127.9955 | 157.0384 | 82.1106 | 247.8082 | 297.9278 | 303.298 | 275.2935 | 138.6298 | 120.2257 | 146.0767 | 125.562 | 75.1873 | 68.4041 | 75.2212 | 100.6682 | 202.9278 | 231.2596 | 266.1106 | 157.9616 | 133.3318 | 163.7449 | 151.9661 | 130.3318 | 164.2596 | 198.2935 | 233.2641 | 171.0045 | 142.0045 | 148.4808 | 165.7404 | 154.9955 | 170 | 210.9661 | 206.9955 | 201.6637 | 159.2257 | 170.9955 | 255.8939 | 202.5869 | 183.4763 | 203.8465 | 230.9616 | 337.8555 | 180.0429 | 169.4808 | 187.8894 | 231.5914 | 132.298 | 138.6298 | 135.5576 | 149.702 | 158.3363 | 173.2167 | 175.7788 | 139.4514 | 122.5531 | 105.4086 | 119.4041 | 133.9661 | 260.851 | 279.7314 | 269.4514 | 371.4086 | 270.0767 | 277.86 | 285.702 | 351.1829 | 130.2935 | 62.9616 | 74.8171 | 130.2935 | 99.5914 | 98.7788 | 117.7788 | 127.4041 | 112.0339 | 94.2596 | 128.4808 | 108.149 | 71.6298 | 75.2596 | 74.0767 | 85.5914 | 91.6682 | 86.8171 | 95.4469 | 72.4763 | 260 | 233.2212 | 268.4898 | 311.711 | 155.851 | 206.6975 | 210.562 | 291.8555 | 304.9661 | 220.851 | 242.6004 | 335.5147
1355 | 16.8994 | 156.232 | 234.4838 | 233.8633 | 174.3471 | 162.6943 | 170.2626 | 199.1367 | 156.1475 | 134.4317 | 126.5899 | 119.5054 | 150.9263 | 195.2212 | 162.8525 | 164.1475 | 206.9784 | 176.7374 | 160.1367 | 152.4317 | 169.5576 | 145.7158 | 381.0737 | 345.9892 | 157.5162 | 213.8633 | 230.0953 | 247.0953 | 279.1889 | 127.3579 | 128.2104 | 136.937 | 114.1475 | 143 | 123.1367 | 165.1367 | 79.7158 | 233 | 294.295 | 305.0737 | 268.3579 | 133.2212 | 131.8525 | 148.1475 | 128.4424 | 90.2734 | 69.9892 | 78.1367 | 104 | 204.4317 | 215.9478 | 248.3057 | 147.937 | 138.9892 | 152.4532 | 148.7266 | 108.3795 | 169.4946 | 186.2212 | 235.6529 | 168.1583 | 143.6529 | 156.6313 | 175.2734 | 136.0108 | 157.1583 | 219.9263 | 234.8831 | 212.3363 | 163.7896 | 159.7158 | 245.8111 | 209.9047 | 178.4209 | 211.9047 | 213.8741 | 319.0323 | 150.6223 | 168.6421 | 168.4532 | 232.2104 | 119.3057 | 149.2734 | 137.4317 | 138.295 | 168.2104 | 169.9892 | 164.5162 | 144.0845 | 110.5791 | 108.6421 | 108.1475 | 135.2212 | 266.705 | 285.3256 | 300.4101 | 339.3273 | 276 | 258.759 | 280.0737 | 337.6529 | 125.2842 | 62.5683 | 76.1475 | 125.2842 | 105.3471 | 108.063 | 123.2104 | 124.4946 | 97.937 | 107.1997 | 109.6637 | 117.705 | 70.0737 | 61.232 | 71.6529 | 88.1367 | 97.5683 | 86.2212 | 107.4209 | 69.3471 | 225.9694 | 239.9892 | 289.5683 | 331.8525 | 148.8633 | 190.4317 | 223.0737 | 281.3795 | 298.5162 | 227.3471 | 243.3057 | 328.9263
1356 | 16.9104 | 139.4931 | 230.8756 | 242.7788 | 160.023 | 150.4055 | 171.8387 | 164.0461 | 165.1244 | 131.0369 | 129.1244 | 115.6912 | 143.4931 | 193.9631 | 160.4194 | 169.1982 | 213.0737 | 167.9631 | 173.47 | 149.0369 | 166.1843 | 160.0507 | 361.6774 | 341.1843 | 147.3825 | 204.4562 | 227.5069 | 241.8894 | 282.4931 | 124.3825 | 120.8387 | 127.765 | 127.7051 | 143.6544 | 141.7051 | 162.765 | 82.2719 | 234.3088 | 305.7051 | 301.3825 | 271.2719 | 126.0737 | 128.765 | 141.4194 | 114.8018 | 82.9124 | 78.2719 | 84.9263 | 106.6175 | 220.0138 | 228.5945 | 240.9631 | 168.8664 | 140.7281 | 161.2857 | 145.6175 | 119.5945 | 161.53 | 188.235 | 255.212 | 155.765 | 148.8157 | 167.8894 | 168.5668 | 119.8756 | 161.7788 | 214.4562 | 252.7281 | 248.1751 | 182.2857 | 179.9401 | 238.5806 | 210.53 | 190.5069 | 215.8018 | 211.8894 | 298.765 | 144.0138 | 176.1982 | 165.5069 | 233.3456 | 125.7051 | 169.3963 | 153.6682 | 132.6912 | 162.1475 | 188.7419 | 171.3963 | 130.8018 | 111.5806 | 115.5438 | 113.1982 | 126.1106 | 255.2949 | 300.6544 | 290.6774 | 323.2719 | 271.4194 | 260.2489 | 271.8018 | 338.8894 | 122.7281 | 62.6912 | 76.6175 | 122.7281 | 117.1982 | 113.3088 | 113.2212 | 115.8756 | 102.5069 | 104.8387 | 95.4194 | 116.765 | 71.6175 | 70.0138 | 70.9263 | 98.8525 | 101.6175 | 100.0138 | 98.5668 | 72.6912 | 227.2857 | 233.2212 | 289.6912 | 318.9493 | 144.6912 | 213.8664 | 222 | 296.5207 | 301.47 | 234.6175 | 228.1106 | 355.5207
1357 | 16.9214 | 156.3334 | 219.6666 | 245.6667 | 159.3334 | 149 | 174.3334 | 183.3335 | 165.6666 | 139.3334 | 123.6666 | 121 | 137 | 197.6667 | 152 | 173 | 211.6667 | 167.6667 | 158.9999 | 168.0001 | 173 | 167.3333 | 343.3333 | 346.6667 | 162.0001 | 213 | 248.6667 | 266.3334 | 257.3333 | 118.3333 | 126.6667 | 135.6667 | 150.6667 | 148.6667 | 127.9999 | 166.6667 | 87.3333 | 221.6666 | 293.9999 | 297.3333 | 269.6667 | 124 | 128.6667 | 149.6667 | 131.6667 | 86.6667 | 66.6666 | 75 | 98.6666 | 220.3333 | 229.3333 | 229.3333 | 191.3334 | 157.0001 | 162 | 142.3333 | 126.3333 | 170.0001 | 191 | 245.3332 | 154.3333 | 132.6666 | 147.6666 | 173.3334 | 136.0001 | 141.9999 | 221 | 266.3334 | 241.9999 | 181.6666 | 177.6666 | 244.3333 | 205.6667 | 202.3334 | 208.6667 | 212.3333 | 296 | 171.6667 | 184 | 187.3334 | 239.6667 | 130 | 164.3333 | 157 | 150.6667 | 171.3334 | 182.6666 | 181 | 126.3333 | 114.6667 | 117.6667 | 117.6667 | 125 | 277.6668 | 309.6667 | 287 | 350.3334 | 248.3333 | 282 | 271.3333 | 339.3333 | 123.6667 | 64 | 76 | 123.6667 | 110.3333 | 109.3333 | 111 | 120 | 106.3333 | 96.6667 | 105.6667 | 115.3333 | 69.6667 | 75 | 77.6667 | 82.9999 | 111 | 88.9999 | 94.6667 | 71.3333 | 245.3334 | 237 | 254.3332 | 327.0001 | 154 | 212.9999 | 205.3333 | 336.0001 | 311.6667 | 238 | 226.3333 | 350.9999
1358 | 16.9324 | 155.4577 | 221.0738 | 256.2213 | 180.295 | 145.1735 | 184.7896 | 186.026 | 154.8525 | 128.3839 | 126.8265 | 127.3948 | 141.4317 | 187.4577 | 159.1844 | 169.6052 | 198.4208 | 168.3579 | 163.9371 | 165.7787 | 203.4426 | 164.2842 | 351.2213 | 354.7158 | 165.9262 | 222.3948 | 234.5943 | 274.9631 | 280.948 | 117.3579 | 125.8894 | 125.3839 | 141.026 | 145.5683 | 138.7267 | 173.3948 | 82.2104 | 246.2322 | 275.1735 | 305.5054 | 285.653 | 140.295 | 127.2842 | 159.0738 | 113.4837 | 84.5314 | 76.295 | 59.4946 | 104.1844 | 206.4946 | 217.2104 | 216.8894 | 167.4837 | 149.705 | 151.5683 | 150.8634 | 117.2104 | 189.2213 | 188.9631 | 217.705 | 152.2842 | 126.7527 | 129.8894 | 163.3839 | 135.4577 | 162.269 | 221.9262 | 274 | 229.2842 | 176.679 | 167.6052 | 256.8634 | 202.6052 | 191.0629 | 198.1735 | 198.7787 | 294.6421 | 159.2733 | 182.6052 | 169.5206 | 219.9154 | 133.4317 | 162.0738 | 164.8634 | 157.2842 | 188.8634 | 166.4946 | 176.5683 | 121.2473 | 122.4686 | 127.8634 | 141.0846 | 116.4946 | 259.052 | 316.7158 | 294.0738 | 353.4946 | 256.3319 | 286.321 | 285.2213 | 338.679 | 120.2473 | 73.1475 | 75 | 120.2473 | 107.7158 | 111.7527 | 115.7158 | 127.7158 | 110.7527 | 98.3948 | 107.9262 | 105.8156 | 63.9262 | 65.5314 | 70.4946 | 87.2582 | 95.3102 | 97.3319 | 96.679 | 68.2842 | 242.2104 | 245.3948 | 256.0109 | 339.0738 | 152.8894 | 202.6052 | 201.0738 | 313.0151 | 313.321 | 232.8894 | 250.7636 | 339.9631
1359 | 16.9434 | 146.5437 | 207.1012 | 262 | 174.8662 | 151.3688 | 167.9537 | 193.8988 | 136.4837 | 143.12 | 117.175 | 132.4563 | 133.6312 | 185.4563 | 155.0137 | 159.0137 | 230.0188 | 169.6913 | 162.705 | 155.8525 | 212.9262 | 157.47 | 365.295 | 329.7325 | 171.6038 | 208.1012 | 230.9125 | 282.295 | 283.175 | 129.06 | 129.2213 | 126.295 | 153.0463 | 149.9125 | 156.6038 | 166.705 | 86.6038 | 266.53 | 273.765 | 310.6913 | 269.8062 | 142.47 | 123.235 | 166.53 | 111.7513 | 72.0137 | 77.7787 | 64.6775 | 113.1475 | 211.6775 | 217.4563 | 229.2075 | 151.235 | 155.1338 | 151.765 | 147.0137 | 114.6913 | 172.88 | 179.705 | 233.5025 | 149.6175 | 142.1338 | 154.65 | 169.825 | 128.6175 | 161.41 | 212.3962 | 239.4374 | 233.53 | 164.5575 | 151.4837 | 239.1887 | 216.2075 | 184.6913 | 208.5163 | 197.8388 | 295.3825 | 153.8388 | 176.8525 | 162.53 | 209 | 131.1612 | 160.5437 | 153.41 | 145.6312 | 189.1612 | 177.2075 | 171.235 | 139.7375 | 117.705 | 115.0275 | 137.4837 | 125.825 | 277.5626 | 312.47 | 282.175 | 355.9125 | 268.4563 | 278.3962 | 285.47 | 348.6775 | 128.3688 | 64.5575 | 67.3962 | 128.3688 | 101.3962 | 101.5575 | 116.3087 | 118.6312 | 106.0875 | 92.3962 | 101.8525 | 94.7787 | 70.9863 | 67.53 | 76.3688 | 78.1012 | 96.3688 | 106.6913 | 101.1475 | 71.1475 | 244.53 | 262.2075 | 261.5437 | 318.1887 | 179.0326 | 212.06 | 229.9588 | 273.1887 | 315.0738 | 232.765 | 235.0412 | 345.2213
1360 | 16.9544 | 164.6992 | 212.6638 | 245.115 | 179.5531 | 169.3673 | 176.5886 | 196.7787 | 142.6638 | 138.9292 | 131.4027 | 118.5221 | 137.4425 | 191.2213 | 167.1815 | 169.0708 | 204.0839 | 165.7787 | 165.3319 | 154 | 196.5221 | 147.9646 | 385.885 | 337.6992 | 172.8894 | 210.8496 | 238.9248 | 264.8937 | 284.0354 | 129.0752 | 135.5177 | 115.9292 | 121.8981 | 156.5177 | 162.8142 | 162.2965 | 87.8894 | 240.1548 | 283.4425 | 296.9292 | 264.6283 | 138.5929 | 123.4071 | 159.1504 | 115.5929 | 77.146 | 72.8894 | 68.2965 | 125.5531 | 215.2965 | 214.0752 | 233.1858 | 148.5929 | 139.8937 | 150.1858 | 144.4071 | 109.3717 | 166.5177 | 197.8098 | 222.1902 | 170.8098 | 126.8937 | 157.1504 | 149.2654 | 139.2567 | 163.8496 | 213.9248 | 253.5488 | 234.5929 | 150.5575 | 168.2169 | 240.2567 | 207.5221 | 193.4425 | 224.146 | 216.1815 | 317.1063 | 164.4425 | 170.7787 | 188.2169 | 242.0665 | 135.3319 | 147.0398 | 141.7787 | 133.261 | 179.9646 | 167.115 | 174.2213 | 141.9646 | 105.5575 | 136.5488 | 133.1106 | 121.4469 | 259.23 | 311.4071 | 297.1063 | 324.5264 | 292.5134 | 274.2965 | 268.9292 | 344.5575 | 139.3319 | 61.8142 | 64 | 139.3319 | 114.885 | 107.9602 | 119.5177 | 127.3673 | 102.2965 | 102.3673 | 93.6681 | 103.9602 | 76.4071 | 73.5177 | 66.9292 | 92.8098 | 103.8142 | 102.0752 | 84.7079 | 76.5177 | 215.3406 | 250.7079 | 267.0354 | 296.0398 | 161.7477 | 204.3362 | 204.7123 | 263 | 316 | 238.9248 | 230.739 | 352.9248
1361 | 16.9654 | 165.8417 | 225.8742 | 232.2733 | 169.6833 | 147.7983 | 187.5792 | 183.5467 | 160.1692 | 144.4533 | 127.115 | 116.295 | 150.3058 | 188.705 | 130.3342 | 166.41 | 189.5792 | 181.8958 | 166.5683 | 134.6725 | 166.3775 | 147.1475 | 368.6617 | 363.18 | 158.3992 | 220.0108 | 240.2842 | 261.7267 | 255.5033 | 119.8417 | 129.1258 | 129.3275 | 128.6225 | 147.9783 | 158.9892 | 163.4317 | 87.7158 | 256.6333 | 302.0325 | 275.2517 | 264.8525 | 143.0108 | 122.5683 | 140.6833 | 122.8742 | 70.9783 | 67.705 | 69.4317 | 114.2517 | 198.5358 | 216.295 | 252.7592 | 167.3275 | 146.7483 | 156.8742 | 150.0108 | 118.4533 | 164.4208 | 175.5033 | 228.0325 | 169.2625 | 118 | 143.6942 | 147.3058 | 123.9567 | 164.4208 | 208.8417 | 248.82 | 231.8525 | 160.6008 | 183.7267 | 240.705 | 185.9675 | 221.3383 | 223.705 | 231.59 | 303.0933 | 166.5683 | 166.1367 | 184.3992 | 230.23 | 138.7158 | 147.0108 | 132.8417 | 152.1908 | 176.2842 | 162.1475 | 170.9892 | 139.7158 | 119.8958 | 117.5033 | 124.6942 | 114.8525 | 262.8958 | 306.9892 | 293.115 | 327.8958 | 282.6725 | 279.7267 | 266.5792 | 332.41 | 144.1475 | 77.3167 | 68.295 | 144.1475 | 99.8092 | 100.8308 | 115.2733 | 123.6942 | 96.9892 | 110.1475 | 112.475 | 96.115 | 69.1258 | 64.9783 | 76.0325 | 99.1367 | 93.5467 | 102.8633 | 84.8742 | 66.5467 | 219.8958 | 241.5683 | 262.8417 | 326.0758 | 144.705 | 186.115 | 209.0433 | 261.5683 | 318.1475 | 235.2733 | 239.0108 | 373.6117
1362 | 16.9764 | 128.05 | 234.825 | 230.7281 | 180.0187 | 142.825 | 187.5438 | 188.4656 | 151.1656 | 149.7281 | 113.2625 | 123.0969 | 153.2719 | 178.9906 | 155.2312 | 156.4469 | 205.5625 | 167.1563 | 171.0969 | 142.2906 | 171.5625 | 148 | 358.2719 | 386.7469 | 173.3875 | 230.0094 | 235.6313 | 244.3406 | 264.1156 | 115.5438 | 123.0875 | 137 | 147.4656 | 133.0781 | 137.3406 | 163.2719 | 84.3594 | 245.4281 | 289.0688 | 298.125 | 236.3313 | 135.5344 | 138.7469 | 162.6687 | 127.4562 | 80.8344 | 57.9906 | 72.9125 | 102.9031 | 221.125 | 201.9813 | 239.1563 | 160.4375 | 136.2531 | 152.7188 | 136.7094 | 112.8063 | 170.2812 | 192.125 | 233.2719 | 169.3687 | 133.2906 | 137.8156 | 159.7375 | 142.2125 | 165.9125 | 214.7375 | 234.7188 | 234.6406 | 174.0094 | 178.7188 | 255.7469 | 196.0187 | 199.6907 | 211.0781 | 235.7281 | 309.2906 | 158.7188 | 194.125 | 195.7469 | 228.0094 | 137.5438 | 148.2719 | 137.2812 | 141.3406 | 179.6406 | 189.2125 | 195.9406 | 118.8844 | 114.6219 | 112.2812 | 115.175 | 116.1844 | 264.175 | 329.7562 | 290.9125 | 335 | 291.7469 | 247.05 | 283.2906 | 333.3687 | 147.1844 | 66.2531 | 64.9031 | 147.1844 | 107.7469 | 91.6313 | 120.2812 | 117.0875 | 107.3781 | 104.4469 | 108.6219 | 101.9219 | 74.7375 | 71.9219 | 79.0875 | 91.4469 | 105.7469 | 110.5531 | 89.4562 | 69.2812 | 232.825 | 234.4469 | 268.0094 | 340.7281 | 113.875 | 209.3969 | 211.9031 | 269.0094 | 300.7969 | 237.3687 | 223.525 | 362.7969
1363 | 16.9874 | 129.0679 | 251.0679 | 230.2596 | 191.9233 | 155.3658 | 182.5575 | 184.5958 | 153.4042 | 144.8171 | 124.0679 | 122.0383 | 141.1534 | 202.655 | 174.7021 | 168.0679 | 231.7316 | 166.4042 | 164.115 | 162.8083 | 185.885 | 154.6637 | 374.2891 | 401.1446 | 169.1534 | 232.2596 | 256.9529 | 254.77 | 286.8083 | 109.0767 | 121.2596 | 145.885 | 136.9321 | 133.4425 | 144.8083 | 159.2979 | 73.3746 | 243.6637 | 272.3746 | 297.8938 | 242.2891 | 148.2891 | 132.4129 | 147.8259 | 126.5192 | 75.6342 | 72.0296 | 82.1446 | 109.885 | 220.8938 | 209.3275 | 233.9617 | 155.7404 | 141.8466 | 175.9145 | 141.3658 | 122.3275 | 175.2212 | 192.6342 | 231.5192 | 171.7404 | 137.5192 | 156.9912 | 172.6254 | 152 | 161.8171 | 232.8083 | 230.5192 | 221.9321 | 144.4217 | 193.0296 | 262.7404 | 205.7021 | 202.8083 | 215.885 | 221.9321 | 316.4808 | 173.0296 | 178.345 | 204.2212 | 213.23 | 138.4808 | 140.5958 | 140.7404 | 156.9529 | 157.3067 | 193.0767 | 203.0383 | 111.7404 | 118.1446 | 116.4808 | 125.5871 | 120.7021 | 247.9321 | 307.9026 | 300.1446 | 352.0296 | 254.3155 | 261.3954 | 267.528 | 374.242 | 136.1534 | 74.0679 | 63 | 136.1534 | 116.2212 | 104.0679 | 123 | 94.528 | 96.4513 | 96.8171 | 102.5192 | 114.885 | 69.115 | 73.0383 | 72.8171 | 86.0383 | 98.6725 | 95.23 | 90.7404 | 74.2212 | 255.7316 | 233.4808 | 267.2979 | 331.3746 | 152.6078 | 202.9704 | 198.8938 | 280.885 | 291.7788 | 241.9617 | 236.2508 | 338.23
1364 | 16.9984 | 140.0217 | 231.9134 | 236.7744 | 203.7852 | 151.4729 | 171.2148 | 188.7744 | 161.2689 | 149.0217 | 125.2365 | 116.4838 | 164.0974 | 216.5162 | 157.935 | 175.2581 | 208.8917 | 156.204 | 162.5054 | 158.9675 | 185.9892 | 154.7419 | 376.9892 | 385.1823 | 158.9783 | 221.4621 | 258.9783 | 261.7527 | 269.4188 | 113.7744 | 135.3014 | 125.6661 | 125.9783 | 150.0541 | 136.4513 | 162.5162 | 80.5379 | 233.9567 | 289.3231 | 298.5162 | 275.8501 | 144.2148 | 126.4946 | 169.861 | 135.7852 | 80.2798 | 84.0217 | 77.4729 | 102.4621 | 210.2256 | 226.0433 | 256.8285 | 162.7744 | 139.9783 | 158.6553 | 135.9675 | 137.5379 | 166.2148 | 201.796 | 231.7527 | 159.9567 | 144.5271 | 168.5162 | 165.4621 | 147.4838 | 161.5054 | 237.2473 | 240.5379 | 213.2365 | 145.796 | 186.9567 | 232.139 | 200.9783 | 188.4296 | 211.4729 | 229.0433 | 311.7311 | 179.7527 | 186.3122 | 187.6878 | 203.2365 | 136.7419 | 142.5162 | 134.2256 | 145.4296 | 145.9892 | 178.204 | 192.2148 | 125.5487 | 114.9783 | 118.5054 | 137.5271 | 115.9783 | 244.5054 | 310.5487 | 327.0866 | 364.7744 | 266.0974 | 281.5379 | 262.2581 | 363.1607 | 144.0433 | 66.9567 | 67.5162 | 144.0433 | 110.2256 | 93.9459 | 107.9459 | 108.8285 | 107.5595 | 107.0433 | 108.0217 | 123.2689 | 75.7852 | 68.9892 | 77.0217 | 81.9892 | 96.2581 | 89 | 102.2906 | 84.7852 | 250.9567 | 240.0217 | 240.408 | 324.9892 | 140.6444 | 213.5595 | 213.065 | 317.1191 | 321.861 | 242.2473 | 246.7635 | 354.5812
1365 | 17.0094 | 151.18 | 250.0099 | 246.65 | 174.1051 | 172.7149 | 158.82 | 184.115 | 179.8299 | 171.6549 | 127.06 | 117.295 | 166.115 | 221.825 | 169.5949 | 182.12 | 212.005 | 171.8899 | 163 | 130.7551 | 152.8701 | 164.71 | 379.825 | 369.055 | 158.53 | 214.94 | 237.1101 | 269.65 | 270.415 | 125.18 | 141.53 | 138.6549 | 109.465 | 145.82 | 154.9499 | 144.1101 | 101.5949 | 234.59 | 284.525 | 291.585 | 270.465 | 136.41 | 132.885 | 174.645 | 110.6951 | 76.88 | 74.525 | 76.53 | 113.535 | 215.65 | 238.415 | 261.705 | 136.6951 | 120.4051 | 137.76 | 133 | 123.4051 | 189.0099 | 201.41 | 244.24 | 152.94 | 138.585 | 138.6351 | 172.71 | 137.585 | 168.885 | 219.4051 | 237.88 | 238.7749 | 152.295 | 177.645 | 261.0149 | 199 | 199.5949 | 191.4051 | 233 | 319.945 | 158.5801 | 165.9901 | 183.53 | 226.4799 | 132.94 | 149.355 | 145.005 | 152.005 | 180.9549 | 189.3 | 193.59 | 134.59 | 108.41 | 119 | 128.525 | 123.18 | 264.8899 | 317.295 | 315.8751 | 345.5801 | 265.82 | 288.06 | 273.71 | 366.475 | 143.41 | 59.94 | 69 | 143.41 | 101.88 | 103.535 | 119.8299 | 132.0649 | 107.645 | 107.94 | 113.825 | 122.705 | 79 | 80.24 | 76.705 | 87.12 | 100.06 | 82.88 | 110.59 | 72.7 | 239.35 | 244.295 | 258.7749 | 346.9499 | 132.53 | 219 | 211.35 | 305.0501 | 329.705 | 262.6549 | 241.88 | 378.8299
1366 | 17.0204 | 147.7817 | 237.7905 | 283.978 | 184.2095 | 176.1136 | 190.978 | 178.1136 | 167.8997 | 144.5766 | 117.118 | 122.6637 | 160.1136 | 197.3495 | 166.4498 | 174.6725 | 203.5634 | 164.7861 | 139.6813 | 136.2139 | 153.882 | 166.4454 | 373.2271 | 356.6725 | 160.5546 | 204.6725 | 227.1136 | 253.3451 | 300.9824 | 125.6681 | 128.0088 | 130.2315 | 126.7641 | 135.2271 | 153.4498 | 159.7641 | 90.6769 | 253.8776 | 278.6681 | 298.3275 | 236.463 | 133.4454 | 142.7729 | 163.6725 | 121.4322 | 83.5502 | 76.441 | 77.7773 | 108.6725 | 207.118 | 224.6769 | 238.4586 | 145.8776 | 135.9868 | 144.1048 | 142.3275 | 127.3275 | 179.8997 | 200.7773 | 220.7949 | 161.3275 | 142.9956 | 125.1136 | 159.6769 | 145.882 | 176.441 | 224.1048 | 254.6549 | 257.1048 | 146.0044 | 171.3363 | 268.3363 | 201.3319 | 205.7773 | 203.8776 | 232.2227 | 318.3363 | 139.5634 | 158 | 181.6681 | 227.7817 | 135.1092 | 147.8908 | 135.0088 | 145.8952 | 177.2315 | 162.1312 | 206.6593 | 121.2315 | 117.1048 | 119.7773 | 116.4498 | 110.4542 | 271 | 321.8864 | 292.1224 | 342.1092 | 280.1003 | 277.3407 | 268.4498 | 321.0308 | 126.4542 | 54.3363 | 82.9912 | 126.4542 | 98.4454 | 121.9912 | 107.8997 | 120.6769 | 103.6681 | 96.8952 | 100.2315 | 96.3495 | 87.5502 | 72.3407 | 76.7773 | 75.7861 | 92.4498 | 93.4366 | 98.0088 | 60.2271 | 230.7817 | 231.0088 | 272.441 | 352.4454 | 153.2095 | 200.3451 | 206.6681 | 295.6681 | 338.3275 | 272.1092 | 211.2403 | 370.0088
1367 | 17.0314 | 134.1563 | 228.8417 | 263.996 | 159.996 | 151.3126 | 179.6813 | 188.8437 | 153.525 | 133.4208 | 113.2104 | 118.4729 | 155.0521 | 195.5271 | 158.4729 | 189.3708 | 208.6854 | 160.2104 | 140.1062 | 147.8958 | 155.4208 | 146.2605 | 360.7355 | 372.1603 | 160.2104 | 220.9499 | 259.9519 | 245.6313 | 282.1543 | 119.4729 | 118.4729 | 137.8437 | 135.3687 | 147.2124 | 139.9459 | 162.0521 | 82.8417 | 236.8918 | 266.9459 | 284.4188 | 269.0581 | 119.5771 | 137.1042 | 164.9479 | 124.6313 | 92.3166 | 74.8417 | 66.1563 | 127.3187 | 206.3687 | 239.7395 | 227.2625 | 153.3687 | 140.4208 | 156.475 | 146.5792 | 116.5771 | 174.2104 | 202.5792 | 227.2124 | 166.3687 | 147.3687 | 141.3708 | 157.3687 | 133.998 | 167.7355 | 250.6874 | 247.3667 | 235.523 | 162.1603 | 179.475 | 278.8437 | 201.2104 | 197.3146 | 228.7395 | 210.6813 | 309.8938 | 150.2124 | 169.8437 | 176.2625 | 214.1563 | 133.6313 | 139.8938 | 152.3187 | 165.1082 | 160.3667 | 164.0541 | 188.6813 | 117.7896 | 116.0521 | 108.9459 | 103.7355 | 132.0561 | 274.1583 | 321.4208 | 286.2104 | 355.6333 | 256.5751 | 272.4208 | 281.7916 | 311.7375 | 121.2104 | 56.9479 | 82.2625 | 121.2104 | 106.6854 | 118.1042 | 112.475 | 109.6834 | 92.7355 | 109.7916 | 104.6854 | 110.3187 | 80.525 | 78.475 | 85.6854 | 79.1062 | 112.1082 | 79.6292 | 105.8437 | 72.2124 | 236.8958 | 234.1062 | 284.2645 | 332.2605 | 181.1082 | 201.3166 | 222.5812 | 298.1583 | 342.5792 | 240.6272 | 228.2665 | 365.2104
1368 | 17.0424 | 153.4524 | 242.4337 | 247.1794 | 168.8393 | 157.0281 | 182.8206 | 195.2075 | 166.2355 | 154.6505 | 124.2262 | 118.6037 | 146.7832 | 188.1794 | 151.3869 | 200.415 | 227.0374 | 177.6412 | 163.6505 | 167.6412 | 176.6505 | 145.0094 | 374.0374 | 368.9813 | 169.6224 | 216.3776 | 242.5383 | 267.4524 | 277.4056 | 129.2262 | 117 | 132.1794 | 132.7925 | 142.9813 | 162.6599 | 187.4617 | 97.2355 | 262.273 | 300.0842 | 295.2355 | 263.1607 | 132.0374 | 142.2168 | 146.7551 | 134.4243 | 85.9813 | 69.1888 | 76.6318 | 100.1233 | 206.1981 | 236.9813 | 238.8299 | 161.2168 | 152.8299 | 167.8206 | 147.8019 | 141.8673 | 174.8019 | 210.2168 | 265.4804 | 148.557 | 138.3776 | 165.2449 | 160.4056 | 130 | 170.6131 | 226.5289 | 236.7832 | 214.5663 | 162.1888 | 165.9626 | 281.1981 | 198.5944 | 180.5663 | 211.5476 | 197.7832 | 302.3869 | 171.6412 | 168.9906 | 187.0281 | 225.4337 | 143.4243 | 154.0374 | 152.3869 | 157.3682 | 176.2449 | 171.8112 | 189.415 | 154.0842 | 119.8112 | 129.2543 | 112.2262 | 127.7738 | 283.8206 | 308.1701 | 294.8206 | 367.8206 | 282.6786 | 263.1794 | 269.1607 | 333.0468 | 154.6786 | 58.8019 | 56.9439 | 154.6786 | 113.0094 | 124.0187 | 115.8019 | 103.1888 | 102.8299 | 101.9719 | 125.443 | 101.5663 | 74.7925 | 68.9719 | 75.1701 | 91.4243 | 92.3401 | 101.4617 | 77.727 | 92.0374 | 268.6692 | 227.9813 | 272.5663 | 317.3776 | 162.9439 | 214.2262 | 229.4056 | 286.1701 | 302.9065 | 230.3963 | 245.4243 | 381.8393
1369 | 17.0534 | 172.0266 | 244.3717 | 280.8231 | 170.5575 | 145.3451 | 188.2566 | 177.2743 | 170 | 145.3451 | 128.6283 | 125.5133 | 151.5133 | 205.5399 | 141.0442 | 188.9734 | 225.3009 | 191.7699 | 167.3717 | 150.0176 | 186.0708 | 177.7523 | 389.3983 | 334.4335 | 147.5751 | 227.0266 | 247.3983 | 269.7434 | 278 | 117.3451 | 125.1416 | 138.9558 | 149.0974 | 145.0708 | 147.0176 | 161.4335 | 88.7876 | 235.8052 | 312.2566 | 314.4691 | 264.6991 | 121.3451 | 133.4159 | 163.9824 | 115.0176 | 73.4159 | 66.3717 | 62.9026 | 100.9558 | 198.6725 | 265.124 | 260.7257 | 176.8408 | 133.2035 | 145.5751 | 143.115 | 121.3185 | 172.5575 | 211.1858 | 214.5662 | 139.9292 | 135.1858 | 136.6194 | 142.2743 | 137.3275 | 175.2566 | 218.1858 | 244.7699 | 221.5841 | 156.9292 | 190.4957 | 263.0884 | 193.9292 | 190.0266 | 249.1506 | 198.4425 | 302.6283 | 166.2301 | 166.3717 | 185.9292 | 224.115 | 130.5309 | 142.5309 | 162.3983 | 150.7434 | 177.7434 | 208.009 | 191.8142 | 147.5309 | 105.5309 | 139.885 | 125.5841 | 126.6283 | 266.4601 | 305.8142 | 285.6017 | 349.646 | 258.4335 | 268.3275 | 266.6283 | 344.5133 | 141.8318 | 66.3275 | 63.2124 | 141.8318 | 119.6991 | 85.2919 | 111.115 | 106.885 | 107.6283 | 123.4249 | 125.9292 | 113.4691 | 69.115 | 61.9292 | 78.5133 | 79.3451 | 95.7699 | 102.3009 | 100.124 | 76.4601 | 266.2301 | 232.5133 | 269.8142 | 304.4159 | 166.7699 | 215.3717 | 243.8408 | 295.2124 | 303.5841 | 256.0532 | 261.0266 | 348.5486
1370 | 17.0644 | 182.4381 | 203.5038 | 256.7683 | 177.4381 | 151.9174 | 166.6858 | 192.8349 | 148.5122 | 179.1904 | 115.7768 | 128.6529 | 167.0497 | 202.5619 | 134.8677 | 179.3884 | 214.0826 | 181.6032 | 170.3058 | 171.4465 | 172.1238 | 175.0826 | 358.1154 | 384.0252 | 165.9671 | 221.7355 | 267.3555 | 246.6858 | 264.7768 | 119.7852 | 119.5619 | 141 | 138.1238 | 135.2561 | 148.6116 | 170.5291 | 91.7852 | 234.6116 | 293.1651 | 311.3884 | 277.5703 | 136.182 | 133.4794 | 168.1735 | 123.2232 | 81.7439 | 65.1735 | 66.4381 | 116.2232 | 216.8349 | 222.4128 | 245.9916 | 161.818 | 131.3058 | 173.0581 | 149.4381 | 113.3471 | 193.4878 | 196.1238 | 225.7936 | 142.3058 | 141.6116 | 142.2232 | 171.8846 | 131.5619 | 164.4297 | 216.3471 | 243.6942 | 224.8265 | 176.6613 | 181.2974 | 259.8265 | 212.8349 | 197.1323 | 223.4625 | 196.5206 | 298.8677 | 164 | 156.0826 | 185.8265 | 215.5619 | 135.2645 | 134.0413 | 145.9916 | 146.6942 | 181.9587 | 191.2064 | 172.1651 | 120.0329 | 102.8265 | 120.3387 | 128.8265 | 122.0413 | 242.1651 | 292.7768 | 294.5703 | 391.2814 | 270.8349 | 280.7439 | 264.5206 | 368.3142 | 138.6529 | 76.2645 | 68.4794 | 138.6529 | 102.818 | 91.7026 | 109.1735 | 97.2561 | 98.0826 | 99.2477 | 105.1651 | 113.6942 | 77.091 | 77.5291 | 79.1735 | 76.8265 | 102.1323 | 99.3471 | 91.2974 | 85.2232 | 250.7768 | 233.1735 | 270 | 310.2645 | 152.4709 | 217.4794 | 215.5948 | 294.6942 | 329.9671 | 243.818 | 242.5122 | 382.1491
1371 | 17.0754 | 177.29 | 226.8723 | 234.0639 | 173.1288 | 133.0314 | 180.4524 | 186.0963 | 155.7424 | 146.7403 | 115.5162 | 125.645 | 153.2251 | 194.29 | 163.356 | 170.4513 | 206.1288 | 187.3874 | 173.5162 | 156.8701 | 180.7424 | 163.7738 | 374.4848 | 366.644 | 170.1613 | 222.5162 | 235.7728 | 266.3236 | 298.066 | 115.1288 | 114.645 | 124.2251 | 129.9675 | 155.6461 | 136.5801 | 171.4838 | 77.9026 | 260.3236 | 316.6785 | 303.29 | 261.5476 | 129.9351 | 148.2586 | 157.0963 | 108.3864 | 75.6126 | 69.1937 | 65.4838 | 114.8063 | 213.4513 | 217.8712 | 220.1927 | 168.0649 | 137.0325 | 147.289 | 147.645 | 103.7738 | 172.8377 | 183.7738 | 242.7424 | 152.2262 | 137.1288 | 150.8712 | 181.5162 | 146.7749 | 147.7414 | 226.9037 | 234.6126 | 230.8712 | 160.0314 | 191.4199 | 265.8712 | 214.4838 | 217.2911 | 221.8712 | 209.4199 | 332.3885 | 159.8063 | 189.2272 | 210.3236 | 230.7749 | 140.355 | 149.7749 | 140.3225 | 157.7424 | 162.8701 | 171.7414 | 180.5812 | 137.6461 | 112.2262 | 115.1613 | 122.29 | 132.7424 | 237.1613 | 324.3885 | 308.7424 | 388.4188 | 244.8052 | 280.4838 | 248.9026 | 343.644 | 128.0963 | 65.4188 | 72.355 | 128.0963 | 97.3225 | 106.7424 | 109 | 115.9686 | 101.0325 | 107.2586 | 93.4513 | 111.3225 | 77.3225 | 74.29 | 68.0963 | 93.7749 | 110.5487 | 78.0314 | 97.2262 | 88 | 256.3874 | 253.1299 | 263.29 | 295.2251 | 153.1937 | 228.0649 | 219.9037 | 283.0963 | 347.5812 | 220.7089 | 253.0974 | 353.2565
1372 | 17.0864 | 160.6814 | 267.0414 | 236.1062 | 165.1917 | 124.7448 | 176.3407 | 193.3614 | 174.1697 | 135.5959 | 116.851 | 126.7021 | 147.4469 | 204.9145 | 152.8303 | 191.9779 | 202.4469 | 179.6386 | 177.4041 | 157.2552 | 206.829 | 153.4897 | 370.4897 | 373.7655 | 170 | 213.6386 | 242.6166 | 259.9365 | 287.9793 | 118.2552 | 122.5103 | 123.5531 | 143.4676 | 151.4897 | 139.9572 | 157.3834 | 97.9779 | 242.0221 | 301.5752 | 314.7655 | 258.851 | 133.9572 | 120.3627 | 179.68 | 120.3186 | 74.851 | 66.5959 | 78.6166 | 116.5531 | 217.9572 | 236.0207 | 241.531 | 155.5324 | 156.7228 | 160.5738 | 130.8303 | 123.2759 | 176.5103 | 195.6166 | 245 | 171.8717 | 128.3407 | 135.8303 | 157.32 | 143.1917 | 189.2538 | 239.2124 | 245.7655 | 241.3614 | 154.2979 | 174.4262 | 262.7448 | 224.2124 | 190.3627 | 228.9572 | 214.5531 | 310.9158 | 173.4676 | 178.1283 | 197.9793 | 232.2979 | 162.2759 | 135.9793 | 146.8083 | 160 | 174.3186 | 179.2124 | 177.0428 | 123.2772 | 135.2759 | 115.851 | 126.1062 | 117.1283 | 276.9986 | 317.3834 | 293.1283 | 348.5545 | 273.8924 | 268.0855 | 286.8497 | 328.6386 | 131.1062 | 79.1697 | 68.7448 | 131.1062 | 105.5103 | 112.4041 | 106.4469 | 107.2345 | 97.7448 | 119.3614 | 99.6593 | 105.8938 | 82.9572 | 60.2345 | 69.4041 | 72.32 | 94.9793 | 85.9145 | 83.6814 | 80.3407 | 260.5531 | 226.3627 | 259.4469 | 338.8069 | 143.7876 | 222.3407 | 240.7228 | 287.8083 | 321.0648 | 230.6166 | 259.4041 | 351.1062
1373 | 17.0974 | 162.3166 | 260.0501 | 262.8998 | 161.41 | 133.4966 | 158.5968 | 195.8633 | 169.2301 | 153.9932 | 124.7699 | 127.8633 | 161.6765 | 184.5535 | 138.7768 | 174.4169 | 216.6765 | 163.3235 | 150.3736 | 167.4966 | 205.82 | 146.82 | 386.2665 | 400.1731 | 163.0934 | 218.0433 | 232.9134 | 270.9499 | 303.1298 | 121.59 | 119.6834 | 133.4966 | 125.2802 | 156.0433 | 149.6333 | 166.2232 | 92.5034 | 261.3098 | 301.4533 | 298.8702 | 265.9066 | 141.0433 | 135.7198 | 157.2369 | 125.59 | 83.6333 | 59.0934 | 73.2301 | 124.7699 | 201.7335 | 256.2665 | 221.8269 | 161.6333 | 156.5467 | 167.4533 | 142.6765 | 118.3667 | 177.1367 | 221.3098 | 249.3166 | 170.6834 | 136.4966 | 153.7198 | 159.9066 | 146.3166 | 184.9134 | 199.5604 | 257.4966 | 227.4601 | 169.5399 | 199.4898 | 287.0365 | 200.9635 | 189.3166 | 209.2802 | 226.2232 | 301.6834 | 174.2733 | 191.4032 | 185.5034 | 231.1367 | 141.8269 | 140.7699 | 141.0934 | 136.6902 | 179.59 | 168.9134 | 159.5968 | 143.3098 | 112.2369 | 135.8565 | 121.82 | 123.4966 | 277.9567 | 327.0866 | 316.7631 | 360.9932 | 284.3166 | 273.7699 | 275.0068 | 349.4465 | 133.7267 | 63.0068 | 87.8565 | 133.7267 | 86.2802 | 108.6834 | 126.7198 | 105 | 115.1298 | 116.6834 | 111.3599 | 92.9134 | 75.3667 | 73.5399 | 88.9932 | 83.5399 | 89.41 | 90.59 | 116.3964 | 70.3667 | 235.1002 | 228.7699 | 267.6333 | 305.5604 | 164.4465 | 217.5467 | 231.0501 | 306.2665 | 333.2665 | 242.4966 | 257.41 | 358.0433
1374 | 17.1084 | 181.388 | 229.9803 | 260.8707 | 183.766 | 121.8657 | 182.2685 | 196.8756 | 177.6318 | 156.1244 | 124.2488 | 131.5025 | 149.9901 | 189.7562 | 140.5025 | 173.6269 | 211.9951 | 162.7512 | 138.1194 | 152.3633 | 181.3583 | 142.4975 | 370.612 | 364.5972 | 169.0049 | 246.1441 | 253.766 | 267.7463 | 298.1194 | 123.7512 | 127.7562 | 122.7414 | 138.6367 | 146.4926 | 153.6269 | 154.8657 | 91 | 234.3534 | 316.0099 | 286.3682 | 245.9852 | 133.2438 | 127.617 | 153.8756 | 122.4975 | 82.3731 | 62.3781 | 72 | 128.6269 | 226.1441 | 216.9704 | 222.3781 | 178.7611 | 148.1194 | 139.1046 | 149.3781 | 108.2438 | 196.2636 | 197.8559 | 224.6071 | 171.7512 | 157.2636 | 155.2488 | 172.383 | 147.8756 | 183 | 221.8954 | 239.7364 | 239.0099 | 205.2734 | 189.1145 | 267.3583 | 199.6269 | 197.0049 | 201.6219 | 222.7463 | 285.2389 | 175.7512 | 190.4975 | 203.2636 | 223.9951 | 155.5123 | 134.1194 | 160.1392 | 155.766 | 184.3781 | 199.3978 | 175.388 | 118.1046 | 113.2537 | 132.8707 | 116.6219 | 119.7463 | 288.383 | 342.1343 | 320.1244 | 338.6071 | 292.0049 | 286.383 | 271.1244 | 344.2438 | 124.3682 | 67.8806 | 78.7414 | 124.3682 | 112.771 | 101.8707 | 131.7512 | 119.8855 | 118.8756 | 108.1194 | 103.3682 | 96.2537 | 66.9951 | 72.4975 | 72.7364 | 88.6269 | 121.3978 | 104.1343 | 93.9803 | 80.383 | 259.8954 | 259.771 | 258.4926 | 312.1343 | 148.7364 | 208.2438 | 229 | 290.612 | 355.2636 | 212.4778 | 266.6318 | 391.3978
1375 | 17.1194 | 167.1297 | 264.18 | 233.3628 | 173.6814 | 151.0768 | 178.0088 | 204.9912 | 163.9056 | 144.4572 | 113.3451 | 106.2507 | 160.4307 | 190.1121 | 161.4219 | 176.6637 | 215.4395 | 176.3186 | 157.4219 | 156.2154 | 174.4484 | 176.6284 | 391.9735 | 378.534 | 163.7846 | 245.5605 | 258.7758 | 264.3363 | 278.3539 | 113.3451 | 120.1209 | 134.3186 | 138.3363 | 166.3098 | 146.0088 | 136.1297 | 96.3274 | 237.9912 | 312.6726 | 300.0944 | 254.5428 | 117.7935 | 139.3186 | 161.9912 | 126.4395 | 68.6814 | 69.2154 | 74.6637 | 98.8111 | 196.2595 | 217.2154 | 227.4395 | 150.8111 | 144.3363 | 173.18 | 173.9735 | 135.4131 | 163.4837 | 186.8967 | 235.2065 | 156.9056 | 129.8111 | 140.7935 | 163.3451 | 137.3451 | 176.7846 | 203.8023 | 226.3451 | 217.0265 | 170.0441 | 181.6726 | 271.1033 | 205.3274 | 213.0944 | 182.3539 | 224.6637 | 300.7582 | 147.5869 | 189.1121 | 187.3539 | 232.767 | 140.2418 | 145.4307 | 153.233 | 160.7758 | 189.4395 | 185.3539 | 186.8791 | 128.2065 | 133.534 | 116.0177 | 113.3363 | 129.6549 | 281.1209 | 312.9232 | 323.5516 | 353.6461 | 282.3451 | 282.6726 | 280.767 | 350.9912 | 154.9647 | 73.4395 | 64.5693 | 154.9647 | 98.3539 | 108.1033 | 111.5781 | 118.4484 | 93.2507 | 104.3363 | 120.6461 | 94.3363 | 70.4395 | 76.4395 | 66.4484 | 84.5605 | 94.9232 | 87.3539 | 93.5516 | 78.4484 | 238.2507 | 261.3363 | 249.0088 | 324.6549 | 152.2154 | 210.5516 | 238.767 | 280.8967 | 339.3539 | 261.2745 | 244.9144 | 387.1209
1376 | 17.1304 | 155.998 | 224.8902 | 281.3114 | 191.8044 | 162.2016 | 165.2974 | 171.7924 | 158.3992 | 151.1018 | 103.8982 | 121.004 | 146.6966 | 203.503 | 167.6008 | 163.497 | 207.8982 | 199.6048 | 151.8982 | 157 | 192.9042 | 155.7944 | 404.002 | 368.3972 | 146.7964 | 233.2974 | 258.0998 | 307.2096 | 292.2036 | 125.503 | 112.6986 | 126.0978 | 145.2016 | 155.497 | 150.4012 | 159.2056 | 102.4012 | 261.505 | 314.7006 | 310.1018 | 274.004 | 132.2036 | 146.4012 | 132.3932 | 135.1018 | 80.503 | 62.7984 | 80.4012 | 107.6028 | 221.7066 | 222.501 | 268.509 | 154.2016 | 134.998 | 159.996 | 160.7964 | 114.6946 | 173.4032 | 191.4012 | 222.5968 | 182.9062 | 123.2994 | 140.8004 | 137.6946 | 126.998 | 171.499 | 213.6028 | 277.2116 | 239.2056 | 161.3992 | 184.6008 | 225.1896 | 188.8962 | 189.7944 | 189.9022 | 210.5968 | 313.8024 | 167.4052 | 184.499 | 192.2016 | 243.002 | 140.7006 | 139.7984 | 154.7006 | 168.2016 | 162.0938 | 180.2994 | 180.7984 | 130.9002 | 127.8982 | 118.501 | 129.2036 | 130.0998 | 260.1956 | 309.9002 | 314.998 | 369.503 | 267.497 | 264.8962 | 296.4032 | 359.2016 | 134.6946 | 71.2994 | 60.2994 | 134.6946 | 131.1078 | 109.9002 | 118.002 | 116.1996 | 108.004 | 105.8004 | 113.998 | 90.3992 | 74.6008 | 77 | 76.8024 | 80.3992 | 100.9022 | 115.6068 | 92.1996 | 73.499 | 241.3014 | 238.495 | 265.1038 | 321.499 | 150.2994 | 206.499 | 225.5968 | 315.1078 | 345.1018 | 233.7924 | 243.8004 | 358.0938
1377 | 17.1414 | 164.1249 | 246.4622 | 263.2753 | 176.6627 | 160.2625 | 195.9372 | 182.5998 | 174.4248 | 158.3874 | 130.3747 | 129.3874 | 139.5251 | 203.175 | 157.0501 | 153.7876 | 239.8496 | 195.6126 | 171.9873 | 148.7876 | 195.9125 | 173.0748 | 367.5879 | 342.3628 | 156.8624 | 222.8751 | 268.0374 | 281.8878 | 319.5497 | 117.8751 | 132.0748 | 148.7247 | 140.5251 | 159.4749 | 153.7375 | 141.0127 | 78.3628 | 259.4375 | 301.3126 | 302.7876 | 263.2251 | 136.7375 | 123.2753 | 149.0748 | 130.5251 | 83.825 | 77.5123 | 86.4749 | 116.2999 | 225 | 239.4248 | 273 | 163.2124 | 144.9499 | 148.8751 | 155.35 | 119.2999 | 164.9626 | 182.8751 | 256.5871 | 167.7502 | 168.6245 | 137.35 | 162.3747 | 146.0748 | 159.1376 | 212.2625 | 271.1376 | 263.8998 | 169.2124 | 160.3628 | 254.6746 | 212.5497 | 178.7876 | 203.7749 | 219.0374 | 302.2251 | 172.7375 | 184 | 175.6627 | 228.4877 | 151.9499 | 128.9626 | 153.175 | 152.5752 | 188.1997 | 179.0875 | 183.65 | 145.5998 | 119.7001 | 106.2251 | 136.4749 | 138.2124 | 278.9873 | 317.2999 | 307.6126 | 344.5378 | 260.5251 | 266.65 | 283.4002 | 341.7502 | 121.0501 | 80.1249 | 64.5625 | 121.0501 | 99.4129 | 113.65 | 129.0374 | 134.2498 | 121.8624 | 121.5123 | 106.6126 | 83.6126 | 72.2625 | 77.9125 | 64.3126 | 85.4749 | 109.2999 | 95.2753 | 81.0501 | 87.5998 | 249.2999 | 247.8624 | 252.4002 | 325.5625 | 167.3373 | 213.2999 | 211.2251 | 282.5004 | 350.5625 | 264.6746 | 262.2498 | 411.5744
1378 | 17.1524 | 143.73 | 229.5796 | 245.2787 | 186.0974 | 146.1283 | 184.2035 | 200.6461 | 174.1504 | 138.6548 | 121.9026 | 108.73 | 142.6991 | 216.8717 | 166.1726 | 155.7743 | 234.677 | 191.3009 | 171.2257 | 149.8496 | 188.6017 | 152.8052 | 379.7213 | 352.0222 | 151.5265 | 260.8408 | 225.5353 | 284.5487 | 311.8274 | 113.3009 | 122.9026 | 133.4292 | 143.6991 | 136.8805 | 151.2257 | 149.1726 | 92.6461 | 240.5044 | 288.9026 | 329.7434 | 259.2257 | 124.0531 | 150.593 | 148.2257 | 117.0531 | 83.0752 | 70.677 | 86.0752 | 106.8274 | 220.3761 | 225.2787 | 257.2787 | 156.6017 | 132.1283 | 160.9469 | 167.0222 | 120.9248 | 175.0974 | 176.4513 | 244.2787 | 158.6017 | 160.9778 | 142.5487 | 166.8496 | 144.3009 | 164.4735 | 236.0443 | 262.6017 | 235.4822 | 136.7079 | 185.7434 | 276.4956 | 215.9248 | 196.4956 | 217.9469 | 235.7213 | 347.2391 | 173 | 180.3009 | 204.5178 | 228.8496 | 140.0531 | 135.3983 | 131.73 | 158.3983 | 191 | 179.9248 | 195.0974 | 133.1283 | 121.7743 | 117.0222 | 137.9248 | 139 | 283.7743 | 330.0222 | 323.6461 | 356.7965 | 279.4204 | 273.4735 | 297.7213 | 328.9026 | 135.7213 | 73.6017 | 64.0752 | 135.7213 | 95.0752 | 128.7965 | 113.3539 | 133.2257 | 107.2787 | 100.8052 | 112.4735 | 120.916 | 75.6991 | 77.0752 | 77.7965 | 83.2257 | 97.9778 | 95.7743 | 84.6239 | 78.8274 | 268.4956 | 260.0974 | 275.9691 | 320.4513 | 156.9778 | 225.0974 | 226.6461 | 310.4426 | 340.8274 | 252.2787 | 260.3009 | 364.2875
1379 | 17.1634 | 153.2449 | 232.6854 | 269.301 | 189.8112 | 165.6156 | 169.881 | 191.6922 | 176.8112 | 171.6717 | 129.4337 | 130.4268 | 160.8044 | 181.4542 | 167.9371 | 183.1751 | 233.0629 | 166.6361 | 170.0629 | 162.1819 | 184.2517 | 167.8673 | 389.4337 | 346.4405 | 152.8741 | 234.9508 | 217.3146 | 289.6854 | 281.0137 | 116.7483 | 135.119 | 144.1819 | 141.1888 | 133.1259 | 151 | 148.1259 | 93.0629 | 284.9165 | 279.5663 | 316.0698 | 244.9439 | 138.9302 | 125.8249 | 144.2517 | 127.2449 | 73.6293 | 70 | 71.9439 | 111.6224 | 220.9371 | 233.3707 | 277.5527 | 185.9863 | 138.4966 | 164.8112 | 147.3844 | 114.4405 | 176.9371 | 196.6156 | 207.3913 | 159.8741 | 162.8112 | 148.6224 | 177.3078 | 158.9932 | 175.3078 | 218.3215 | 272.3078 | 237.6854 | 172.42 | 169.2586 | 261.1327 | 199.1327 | 207.3707 | 211.5034 | 218.2586 | 314.4542 | 187.9932 | 155.6361 | 196.6922 | 245.8673 | 154.9302 | 132.2517 | 160.9234 | 157.1259 | 176.9439 | 193.119 | 195.0629 | 139.4966 | 114.5034 | 121.7483 | 123.9439 | 140.8741 | 277.4405 | 323.5034 | 319.3776 | 349.5663 | 290.3707 | 307.7346 | 291.5034 | 329.8741 | 133.2517 | 71.1259 | 78.9932 | 133.2517 | 115.6156 | 115.9439 | 126.9932 | 118.0068 | 97.5663 | 99.9371 | 105.5034 | 103.3844 | 60.0698 | 73.2517 | 60.2586 | 105.4897 | 87.6293 | 87.5663 | 110.301 | 72.3776 | 295.301 | 245.0698 | 307.0492 | 320 | 147.5663 | 208.1956 | 213.0068 | 311.1259 | 335.3146 | 230.3844 | 248.7551 | 357.1888
1380 | 17.1744 | 147.3545 | 225.4051 | 263.4051 | 177.6583 | 155.6076 | 169.9494 | 183.4051 | 163.7089 | 160.7089 | 117.6583 | 126.3038 | 168.6455 | 197.0379 | 172.7468 | 168.8608 | 226.3545 | 175.443 | 170 | 138.3166 | 196.3417 | 164.2532 | 371.9621 | 357.3924 | 153.9494 | 240.5949 | 233.1392 | 271.9621 | 283.7468 | 112.2532 | 132.2025 | 142.1519 | 113.4685 | 138.6962 | 164.2911 | 139.4557 | 97.7468 | 236.7344 | 264.7595 | 312.1519 | 257.2911 | 144.7468 | 146.7847 | 163.9366 | 122.3038 | 68.2532 | 65.2532 | 65.3038 | 107.2532 | 211.5064 | 247.2911 | 267.6076 | 155.7216 | 163.6834 | 166.8987 | 146.9494 | 126.3417 | 161.8102 | 172.3672 | 243.9239 | 171.3924 | 152.557 | 128.114 | 163.7595 | 116.3293 | 162.7089 | 207.5064 | 259.7089 | 234.2025 | 179.7468 | 169.8987 | 291.329 | 218.886 | 217.4936 | 193.9115 | 221.7468 | 321.4936 | 155.7723 | 187.2277 | 198.8481 | 214.7216 | 137.9621 | 141.4936 | 137.3672 | 167.443 | 185.4936 | 180.7089 | 205.443 | 130.5064 | 110.2025 | 137.1898 | 135.3417 | 132.4557 | 265.6076 | 312.557 | 330.3924 | 373.6834 | 237.8357 | 270.1267 | 284.3545 | 351.8354 | 125.4051 | 71 | 75.2532 | 125.4051 | 101.8102 | 114.0506 | 93.8229 | 119.8481 | 117.886 | 98.1013 | 102.1519 | 111.4936 | 88.4303 | 60.6583 | 75.1392 | 88.0127 | 118.329 | 106.9366 | 106.3038 | 70.1013 | 278.9621 | 233.557 | 296.6583 | 304.8102 | 139.4051 | 221.2405 | 229.0885 | 314.7975 | 332.1519 | 241.3417 | 234.7089 | 380.7341
1381 | 17.1854 | 139.3068 | 265.3895 | 297.6196 | 172.1917 | 162.6932 | 191.1564 | 184.9233 | 149.5368 | 148.4601 | 117 | 108.6902 | 157.4601 | 203.7699 | 167.2301 | 163.1917 | 204.8436 | 170.2301 | 144.997 | 161.0414 | 192.1917 | 143.8053 | 385.4248 | 373.3865 | 171.3098 | 227.5368 | 221.4985 | 259.4601 | 304.1947 | 114.885 | 129.115 | 138.1534 | 137.9647 | 136.115 | 141.9203 | 168.8113 | 92.2301 | 272.4662 | 305.3511 | 314.885 | 274.3481 | 142.115 | 158.5782 | 146.7286 | 116.2301 | 84.3481 | 82.3098 | 72.6932 | 97.3835 | 223.5015 | 239.3451 | 247.7669 | 187.6579 | 165.9617 | 160.2684 | 140.2684 | 122.1917 | 178.3098 | 184.4632 | 222.9203 | 172 | 159.6932 | 139.5015 | 161.0767 | 129.3865 | 167.7699 | 230.0797 | 248.4218 | 232.0767 | 154.0353 | 167.115 | 240.109 | 207.4985 | 208.3835 | 200.6932 | 246.0414 | 330.6549 | 181.888 | 160.1504 | 177.8436 | 228.3865 | 145.6549 | 157.3865 | 166.7729 | 151.6519 | 165.8053 | 200.1947 | 193.4985 | 122.3068 | 99.4218 | 116.8436 | 122.5368 | 142.5782 | 271.7316 | 340.8496 | 303.112 | 366.3451 | 285.9677 | 300.6963 | 286.885 | 340.4985 | 143.2714 | 82.5399 | 81.7316 | 143.2714 | 99.0767 | 114 | 116.0414 | 109.4218 | 111.3068 | 101.8466 | 101.0383 | 121.6165 | 72.6902 | 76.3481 | 69.2684 | 77.3835 | 94.0353 | 96.4601 | 85.8053 | 86.3481 | 257.8053 | 227.2301 | 282.5368 | 335.7346 | 155.3481 | 195.0737 | 230.9617 | 280.3804 | 343.5399 | 258.3481 | 232.0767 | 385.8466
1382 | 17.1964 | 158.4789 | 244.5993 | 260.0422 | 189.531 | 130.8598 | 191.0261 | 203.505 | 174.3226 | 131.4429 | 117 | 138.1923 | 164.7916 | 210.8176 | 162.1303 | 174.6873 | 210.8176 | 172.9218 | 146.9218 | 160.0521 | 169.5993 | 167.3486 | 391.8437 | 402.2444 | 155.4429 | 214.3387 | 204.4429 | 249.2605 | 273.8337 | 139.3486 | 133.8697 | 116.5732 | 161.4007 | 130.1563 | 135.1563 | 158.3127 | 87.1303 | 242.8337 | 303.1042 | 297.469 | 260.3908 | 136.1563 | 163.8697 | 149.8958 | 127.6873 | 89.8697 | 75.2084 | 66.1824 | 105.7655 | 224.9739 | 229.2605 | 278.1663 | 200.6873 | 163.0782 | 161.9479 | 169.2184 | 104.469 | 164.3908 | 207.4007 | 243.4268 | 173.9479 | 144.4169 | 168.2444 | 171.7134 | 139.7395 | 176.7655 | 211.5211 | 248 | 262.1923 | 144.2345 | 173.8176 | 293.5149 | 205.0521 | 208 | 201.9739 | 228.495 | 280.3548 | 167.4169 | 173.6092 | 201.3486 | 224.1303 | 153.7916 | 138.5211 | 155.3387 | 164.6352 | 198.1141 | 210.7395 | 204.6873 | 137.5831 | 127.2444 | 122.8176 | 131.7395 | 135.2084 | 267.1303 | 351.7395 | 328.2965 | 351.3908 | 277.2866 | 286.4169 | 263.6253 | 344.8697 | 146.9218 | 52.8077 | 78.1042 | 146.9218 | 107.7655 | 120.8176 | 111.1563 | 101.2084 | 111 | 84.469 | 116.5831 | 113.2345 | 72.9739 | 80.8958 | 75.8176 | 98.4268 | 102.7395 | 94.0521 | 111.2965 | 80.1824 | 249.2084 | 228.9479 | 287.8437 | 307.7816 | 167.6873 | 221.2705 | 219.3127 | 289.7134 | 380.036 | 260.9479 | 260.2444 | 352.8859
1383 | 17.2074 | 159 | 212.4404 | 259 | 208.7385 | 145.7798 | 183.1101 | 182.3028 | 174.0138 | 144.8073 | 142.6421 | 130.1239 | 139.3579 | 181.4129 | 176.7935 | 186.8348 | 225.7935 | 173.9862 | 135.1652 | 172.8211 | 179.8486 | 159.1239 | 376.2202 | 426.6697 | 171.766 | 214 | 246.4082 | 269.711 | 304.5596 | 138.0275 | 159.6421 | 133.7523 | 151.1514 | 151.6972 | 148.8073 | 156.0275 | 94.8899 | 239.0413 | 314.8348 | 294.0413 | 273.8073 | 127.1239 | 145.2615 | 138.1652 | 122.0826 | 64.3579 | 68.0963 | 71.9174 | 105.0138 | 213.1652 | 244.7798 | 221.7983 | 177.3303 | 152.1514 | 157.0688 | 156.1927 | 122.7385 | 147.234 | 209.9725 | 248.9312 | 176.9587 | 139.0688 | 159.1376 | 165.0963 | 132.1101 | 191.7935 | 192.2615 | 282.5183 | 243.2753 | 139.0688 | 177.9449 | 242.7295 | 192.1789 | 208 | 257.2292 | 259.5596 | 334.2292 | 183.766 | 185.8348 | 172.4129 | 267.3944 | 135.2615 | 166.6009 | 160.9174 | 151.1927 | 175.3303 | 205.0826 | 180.3441 | 132.0826 | 118.1376 | 129.9037 | 116.2202 | 119.2202 | 279.8211 | 299.7295 | 326.0413 | 403.2705 | 292.7798 | 294.8761 | 289.6284 | 376.5596 | 130.234 | 70.7385 | 49.3991 | 130.234 | 119.8348 | 89.4404 | 108.0413 | 138.477 | 104.0963 | 107.6697 | 98.2615 | 99.1927 | 77.9312 | 89.8761 | 75.0138 | 93.0826 | 80.3166 | 100.9037 | 101.1514 | 61.2615 | 278.5871 | 264.5045 | 284.0551 | 306.0138 | 144.3303 | 246.6559 | 231.8211 | 324.5183 | 352.3991 | 260.0138 | 260.0138 | 347.0688
1384 | 17.2184 | 146.0191 | 228.975 | 267.9867 | 190.028 | 120.0383 | 166.025 | 187.9912 | 165.0133 | 171.9602 | 130.0191 | 131.9971 | 167.9573 | 203.9661 | 146.0456 | 197.9838 | 203.0339 | 182.9867 | 161.9602 | 173.9985 | 192.9809 | 176.9735 | 371.0074 | 428.9971 | 182.9838 | 218.9926 | 236.0162 | 279.9853 | 287.0265 | 130.0118 | 121.0574 | 136.9956 | 141.0147 | 126.0383 | 150.9971 | 155.0015 | 100.9912 | 280.9382 | 305.0147 | 310.975 | 283.9853 | 128.9971 | 140.0074 | 146.9867 | 144.9661 | 75.9823 | 80.9809 | 72 | 105.9985 | 219.9897 | 248.9941 | 215.0088 | 169.0118 | 152.9985 | 140.025 | 158.9956 | 118.0074 | 154.9882 | 208.0029 | 240.0133 | 179.9956 | 162.9647 | 117.0618 | 177.9809 | 160.9573 | 159.0486 | 215.9647 | 264.028 | 243 | 196.9146 | 171.0103 | 270.9573 | 224.9514 | 222.9779 | 244.0206 | 230.0442 | 304.0456 | 174.0147 | 174.0177 | 210.9426 | 240.0412 | 154.9706 | 135.0471 | 141.0294 | 154.9941 | 190.9764 | 196.0133 | 173.0103 | 133.9971 | 114.0059 | 129.0015 | 116.9985 | 126.9882 | 267.0191 | 306.9882 | 319.0103 | 374.0442 | 279.0206 | 291.0059 | 290.9985 | 366.0162 | 123.0103 | 79.9867 | 68.9706 | 123.0103 | 116.0059 | 119.9544 | 112.9926 | 128.0162 | 104 | 102.0088 | 98.9985 | 100.9971 | 70.0118 | 75.0221 | 71.0059 | 85.0118 | 103.9647 | 100.0015 | 87.0206 | 96.947 | 271.0118 | 248.025 | 262.0324 | 330.9632 | 154.9838 | 209.056 | 229.0044 | 309.0236 | 363.9823 | 271.9823 | 278.972 | 347.9985
1385 | 17.2294 | 176.4699 | 243.8918 | 270.9784 | 179.9567 | 161.0757 | 197.6105 | 177.3246 | 167.0649 | 151.9026 | 125.9459 | 113.2488 | 149.9135 | 212.9892 | 151.8918 | 165.4003 | 223.9135 | 172.2921 | 153.0433 | 144.8485 | 181.2921 | 192.6754 | 373.6971 | 402 | 161.1839 | 212.9784 | 271.6754 | 291.8053 | 289.2813 | 144.8594 | 130.9243 | 139.1298 | 140.1731 | 141.9567 | 131.9892 | 161.119 | 81.9675 | 278.0865 | 331.6105 | 316.3246 | 292.881 | 133.2705 | 149.0108 | 160.1082 | 123.1515 | 79.8053 | 90.7079 | 73.0649 | 106.8918 | 215.0757 | 258.5023 | 265.5564 | 163.2164 | 136.0541 | 135.4003 | 148.238 | 146.8053 | 152.1515 | 221.7512 | 260.0325 | 164.0541 | 144.9026 | 145.9784 | 185.6754 | 143.9567 | 163.0541 | 220.0108 | 262.0541 | 249.9243 | 150.0649 | 173.119 | 267.7728 | 208.2056 | 244.4158 | 216.3029 | 216.357 | 310.0541 | 180.8918 | 203.6646 | 189.0325 | 245.7944 | 147.0325 | 135.238 | 153.0216 | 171.6646 | 208.8161 | 178.9135 | 167.2056 | 155.9459 | 127.7404 | 105.0865 | 133.2705 | 124.0541 | 309.0216 | 327.9026 | 328.8918 | 378.9675 | 286.4652 | 288.1082 | 322.5781 | 353.9026 | 133.9135 | 69.1515 | 79.9784 | 133.9135 | 117.8377 | 112.8485 | 116.0325 | 119.8269 | 102.1082 | 118.8161 | 115.8377 | 112.0865 | 80.0216 | 73.0757 | 64.2272 | 104.9135 | 80.2164 | 88.1839 | 94.0108 | 81.1082 | 287.9784 | 237.1082 | 293.9351 | 315.2056 | 167.8485 | 234.9784 | 218.4652 | 346.2211 | 339.5301 | 274.6863 | 264.8161 | 363.8918
1386 | 17.2404 | 128.5315 | 234.3004 | 268.5147 | 176.3929 | 167.5147 | 162.6933 | 207.4391 | 173.1849 | 143.2311 | 121.6471 | 135.4685 | 142.7395 | 212.2311 | 142.8319 | 200.9832 | 216.6471 | 199.0462 | 156.8151 | 131.6702 | 207.7227 | 162.8151 | 346.7857 | 402.3235 | 178 | 210.4454 | 242.7857 | 273.584 | 314.7458 | 131.8382 | 124.1155 | 150.4454 | 155.584 | 138.0924 | 132.2017 | 171.1912 | 78.7227 | 284.4748 | 296.5546 | 344.6596 | 281.6765 | 157.3529 | 149.9769 | 169.6071 | 137 | 62.1387 | 64.0462 | 78.6302 | 97.7857 | 222.4622 | 213.7626 | 226.479 | 182.4685 | 141 | 172.0462 | 169.2605 | 128.8613 | 166.208 | 198.6765 | 262.9076 | 168.4916 | 136.7395 | 144.0231 | 156.8551 | 140.2311 | 168.624 | 221.2311 | 266.8382 | 243.5546 | 156.1618 | 183.9769 | 247.624 | 226.9307 | 190.9076 | 243.8382 | 248.2143 | 315.3698 | 170.6765 | 173.2542 | 192.7164 | 227.3235 | 150.1849 | 156.3529 | 154.9769 | 141.2542 | 191.792 | 171.6702 | 186 | 151.3004 | 104.2311 | 113.2542 | 157.0063 | 129.3004 | 310.2374 | 319.0693 | 319.0693 | 376.4391 | 327.4748 | 297.5147 | 284.3466 | 345.8551 | 126.0693 | 82.6534 | 77.9538 | 126.0693 | 103.3235 | 99.4853 | 119.1155 | 104.3698 | 111.584 | 102.0462 | 101.3004 | 119.9769 | 82.3698 | 79.9307 | 85.3929 | 97.1618 | 99.8613 | 104.792 | 94.9538 | 90.9769 | 284.8676 | 247.8551 | 288.7626 | 333.5378 | 154.4622 | 232.6302 | 259.5903 | 275.2311 | 387.6996 | 246.0231 | 249.0862 | 354.3004
1387 | 17.2514 | 152.3098 | 246.469 | 248.5664 | 193.0354 | 147 | 191.2212 | 224.4424 | 180.9646 | 152.7168 | 147.9026 | 113.0354 | 173.115 | 220.8672 | 177.7876 | 159.1682 | 242.8672 | 200.2566 | 149.354 | 159.9292 | 195.823 | 155.6726 | 379.6814 | 416.2832 | 177.6106 | 189.4427 | 274.7256 | 257.239 | 303.6106 | 125.2832 | 128.9646 | 127.8496 | 138.885 | 142.2832 | 181.584 | 138.0266 | 68.0266 | 222.3364 | 319.115 | 289.8054 | 267.8938 | 130.1416 | 148.7522 | 153.3186 | 137.4248 | 68.0708 | 66.177 | 63.8142 | 130.6814 | 242.9204 | 244.8672 | 288.7522 | 159.7522 | 141.531 | 174.4248 | 138.4956 | 122.8584 | 175.0708 | 185.177 | 258.469 | 147.6372 | 167.5398 | 145.2124 | 192.0796 | 149.292 | 195.0708 | 229.6548 | 260.4602 | 265.2654 | 162.7522 | 183.4602 | 274.7434 | 222.7964 | 187.4602 | 236.0088 | 215.2832 | 330.9646 | 157.885 | 183.8938 | 223 | 241.2124 | 157.6106 | 129.708 | 153.7876 | 152.2832 | 183 | 199.4336 | 185.9646 | 162.9026 | 114.4602 | 123.6106 | 115.9558 | 142.1062 | 278.4956 | 322.885 | 323.593 | 394.469 | 264.5222 | 277.885 | 298.9292 | 382.2832 | 129.0354 | 67.9292 | 75.823 | 129.0354 | 116.3274 | 120.7434 | 123.2566 | 119.8938 | 94.7434 | 104.0354 | 113.8584 | 118.823 | 97.115 | 76.9292 | 101.5044 | 103.6106 | 94.0708 | 96.2832 | 92.7876 | 89.292 | 238.3452 | 282.3008 | 319.6548 | 314.3894 | 174.4602 | 217.708 | 200.8142 | 285.2832 | 374.9646 | 248.4868 | 294.5398 | 366.7876
1388 | 17.2624 | 186.14 | 232.7154 | 262.4262 | 192.4739 | 146.1415 | 170.1908 | 182.7631 | 180.0954 | 146.1446 | 117.8092 | 114.0954 | 150.1923 | 189.6662 | 171.8092 | 190.4754 | 210.9031 | 180.8108 | 158.9046 | 157.7138 | 191.2385 | 174.9538 | 369.8554 | 422.76 | 166.4277 | 254.2354 | 240.0477 | 291.2385 | 293.5723 | 132.2846 | 127.0462 | 150.1415 | 161.3308 | 150.3815 | 144.0015 | 165.8092 | 96 | 285.5231 | 295.7154 | 337.4262 | 264.2846 | 134.5723 | 140.4739 | 162 | 148.3323 | 69.9046 | 71.4769 | 85.8092 | 121.5231 | 266.0446 | 215.2877 | 280.8077 | 153.5246 | 157.0969 | 185.38 | 151.6185 | 118.8569 | 178.4308 | 190.2862 | 244.0954 | 165.1908 | 153.9508 | 151.5723 | 166.9046 | 131.5738 | 194.8539 | 195.623 | 273.3338 | 220.0984 | 156.5723 | 194.3308 | 295.4769 | 189.9523 | 201.0969 | 209 | 222.5708 | 329.7138 | 182.3338 | 182.4784 | 221.8554 | 246.0462 | 147.2385 | 150.1446 | 147.8092 | 160.5723 | 182.7138 | 184 | 186.2877 | 133.0477 | 127.4292 | 114.86 | 140.14 | 144.6662 | 292.8108 | 345.7123 | 365.2354 | 380.62 | 306.62 | 301.2369 | 296.6662 | 388.8077 | 131.0969 | 66.1431 | 70.8569 | 131.0969 | 99.5261 | 140.4277 | 103.62 | 116.9046 | 114.4754 | 105.8585 | 110.2862 | 114 | 73 | 74.4277 | 87.2369 | 92.7615 | 95.8092 | 102.9508 | 88.86 | 70.5246 | 274.1415 | 235.3323 | 282.3323 | 325.2862 | 186.1892 | 237.0477 | 225.0507 | 295.1938 | 371.9493 | 288.7615 | 283.0969 | 359.6646
1389 | 17.2734 | 148.6401 | 246.1602 | 233.2194 | 163.8595 | 130.9794 | 174.12 | 198.2999 | 180.2005 | 168.94 | 114.1799 | 116.2999 | 174.1799 | 185.5793 | 166.6204 | 179.1602 | 190.1997 | 195.5004 | 157.6598 | 152.2999 | 197.1397 | 192.8603 | 349.0394 | 395.0608 | 155.4799 | 221.3993 | 240.0403 | 294.4404 | 307.0994 | 117.8201 | 109.7395 | 132.94 | 128.6598 | 156.4404 | 163.2202 | 161.7001 | 94.9203 | 276.4799 | 308.2005 | 307.4396 | 248.6804 | 145.8201 | 112.0394 | 162.4199 | 133.9803 | 68.06 | 81.06 | 80.5004 | 111.4002 | 227.2399 | 238.9006 | 257.6598 | 164.7198 | 177.3805 | 173.12 | 145.0797 | 116.8398 | 206.5201 | 196.7798 | 246.6598 | 169.06 | 133.1799 | 163.5398 | 165.5398 | 162.88 | 154.0394 | 245.5407 | 278.8603 | 261.1602 | 168.4199 | 163.2793 | 304.88 | 191.2194 | 223.5998 | 211.2194 | 214.5398 | 323.7001 | 188.94 | 210.1405 | 198.3402 | 226.7601 | 150.9203 | 170.0009 | 144.6598 | 170.6204 | 177.06 | 184.5398 | 211.2202 | 134.4799 | 135.2802 | 151.3402 | 105.0994 | 137.5201 | 307.0206 | 321.7995 | 331.3796 | 392.94 | 318.3402 | 286.6598 | 291.3196 | 364.7601 | 152.7001 | 68.5201 | 68.2399 | 152.7001 | 129.4602 | 128.5201 | 115.6401 | 116.3196 | 104.06 | 121.8603 | 116.2399 | 113.3402 | 73.5398 | 61.9203 | 72.1799 | 88.2399 | 92 | 84.6392 | 123.7207 | 81.3599 | 259.5193 | 221.5801 | 269.9597 | 328.0608 | 170 | 237.8201 | 264.6804 | 338.5201 | 334.8389 | 283.5801 | 303.7404 | 337.4387
1390 | 17.2844 | 146.1078 | 233.2891 | 268.1445 | 193.1445 | 160.2654 | 177.0841 | 202.205 | 151.9277 | 167.205 | 116.5664 | 121.5059 | 176.9277 | 225.4941 | 145 | 165.2773 | 212.5414 | 173.0118 | 166.1209 | 156.5664 | 213.1932 | 176.4455 | 380.5664 | 353.6978 | 164.3009 | 256.6872 | 228.3969 | 270.4336 | 338.7713 | 115.5782 | 137.5059 | 133.3732 | 140.3009 | 133.1564 | 151.6505 | 156.7832 | 78.5782 | 283.2773 | 282.3128 | 331.1564 | 231.9028 | 142.7832 | 144.4336 | 169.7227 | 119.0841 | 69.5059 | 81.7832 | 57.9396 | 103.5178 | 231.5782 | 209.8305 | 270.9632 | 174.2654 | 150.9882 | 174.6386 | 161.0604 | 130 | 196.8318 | 207.4822 | 253.9645 | 170.6505 | 136.1445 | 170.5545 | 174.4336 | 159.6268 | 185.7109 | 205.8436 | 260.1327 | 247.7109 | 175.3614 | 198.4941 | 301.9159 | 226.795 | 233 | 243.8318 | 222.3495 | 320.0118 | 187.7832 | 182.1564 | 189.59 | 224.2287 | 134.3614 | 125.891 | 154.4941 | 150.6623 | 178.5059 | 195.2405 | 198.9277 | 142.3614 | 124.8673 | 139.3377 | 138.1445 | 130.2891 | 276.8673 | 347.7595 | 350.109 | 393.0841 | 305.8318 | 297 | 309.9041 | 360.5664 | 148.5782 | 61.5059 | 72.1445 | 148.5782 | 120.4941 | 121.0723 | 109.205 | 135.9882 | 106.3009 | 103.7832 | 118.7713 | 103.795 | 81.9277 | 45.8673 | 74.2773 | 92.795 | 94.0959 | 122.9645 | 89.3732 | 85.5545 | 296.4704 | 217.1682 | 285.3614 | 284.4573 | 168.7713 | 234.1327 | 243.1327 | 330.5664 | 393.6268 | 276.4218 | 284.5059 | 404.109
1391 | 17.2954 | 183.3785 | 236.8309 | 266.8711 | 195.3383 | 137.7463 | 190.4926 | 192.1691 | 150.8309 | 158.5222 | 111.2537 | 126.8161 | 174.4778 | 218.6617 | 147.9598 | 156.1691 | 255.8013 | 186.9302 | 144.2833 | 150.4926 | 190.1691 | 191.702 | 373.3087 | 411.3785 | 179.3933 | 226.2389 | 273.353 | 277.353 | 322.8309 | 123.9302 | 142.3933 | 150.3235 | 157.2537 | 148 | 158.9852 | 153.8309 | 86.8456 | 274.4228 | 312.3383 | 341.6872 | 281.4926 | 141.0993 | 149.0698 | 175.6174 | 131.5624 | 77.4376 | 79.4228 | 69.9154 | 122.5772 | 242.298 | 255.1946 | 308.408 | 154.8752 | 138.3383 | 169.2389 | 151.3678 | 129.408 | 169.9302 | 187.7463 | 218.89 | 178.1544 | 139.4376 | 154.1987 | 179.408 | 143.353 | 181.408 | 194.1289 | 251.3678 | 245.353 | 179.3235 | 190.5624 | 286.1396 | 237.2537 | 229.8711 | 215.2389 | 216.4524 | 330.7167 | 182.5476 | 196.8309 | 207.8013 | 241.6913 | 136.8859 | 161.4778 | 149.2685 | 172.5074 | 187.4524 | 220.702 | 199.353 | 145.9007 | 134.9007 | 118.9302 | 137.9704 | 134.0846 | 291.4672 | 318.1544 | 313.0846 | 403.11 | 280.0296 | 296.2389 | 285.3678 | 355.8456 | 154.8161 | 70.0296 | 74.1691 | 154.8161 | 112.8161 | 119.0402 | 100.7759 | 121.055 | 122.9154 | 99.7315 | 102.408 | 115.353 | 79.647 | 57.3383 | 66.1839 | 102.9154 | 117.9557 | 85.353 | 107.0846 | 68.945 | 267.1289 | 245.0846 | 285.9409 | 316.3383 | 155.5369 | 225.0296 | 233.5222 | 325 | 376.2537 | 268.0698 | 290.1544 | 367.9302
1392 | 17.3064 | 155.3874 | 232.966 | 235.712 | 197.7409 | 135.9686 | 182.8692 | 191.966 | 151.2277 | 172.966 | 114.4843 | 112.7409 | 156.9346 | 216.1623 | 176.9006 | 159.2591 | 232.5811 | 196.8063 | 166.966 | 147.4214 | 192.0969 | 158.2277 | 361.2956 | 387.5522 | 163.5497 | 224.3925 | 282.4792 | 287.8352 | 322.6466 | 133.322 | 127.034 | 141.7409 | 158.644 | 149.0654 | 146.1597 | 154.712 | 93.966 | 282.8742 | 316.8717 | 297.9057 | 277.1623 | 150.288 | 137.7409 | 138.9031 | 118.2277 | 91.1597 | 83.1283 | 70.5497 | 118.1937 | 273.7383 | 220.7799 | 299.7723 | 187.2591 | 141.8063 | 160.7094 | 176.7094 | 122.2252 | 178.5472 | 184.5157 | 261.1937 | 167.8377 | 155.678 | 175.5786 | 173.678 | 155.4817 | 176.1623 | 225.5786 | 277.7748 | 256.3195 | 172.1937 | 173.3534 | 270.068 | 239.5157 | 198.8088 | 208.1623 | 239.7069 | 304.9346 | 158.3245 | 195.0969 | 184.678 | 257.966 | 117.39 | 145.3874 | 160.966 | 176.5472 | 216.034 | 186.0654 | 213.7094 | 133.9031 | 121.644 | 130.7434 | 117.1623 | 134.0314 | 329.5811 | 309.9686 | 315.2591 | 364.7774 | 299.3849 | 288.9686 | 308.4817 | 368.5837 | 139.6755 | 89.4817 | 76.0969 | 139.6755 | 100.5811 | 90.2931 | 120.3874 | 102.034 | 120.3534 | 88.5183 | 98.4214 | 129.322 | 67.1308 | 62.9371 | 78.7094 | 100.5472 | 88.0994 | 98.9346 | 107.6126 | 89.2252 | 298.6755 | 247.7434 | 247.8114 | 324.6491 | 181.5472 | 247.8717 | 246.61 | 325.2906 | 378.7094 | 258.678 | 280.0314 | 379.9371
1393 | 17.3174 | 159.5457 | 216.5104 | 261 | 185.7817 | 146.2006 | 164.8555 | 177.5842 | 171.5634 | 152.2536 | 120.2006 | 103.1829 | 147.8732 | 222.1975 | 150.4012 | 169.2536 | 236.6902 | 193.5811 | 150.0738 | 168.9085 | 192.4543 | 178.3451 | 414.9792 | 425.4158 | 178.1091 | 277.2329 | 230.6726 | 253.8025 | 334.7256 | 125.7994 | 144.9085 | 129.6726 | 143.9262 | 159.6549 | 129.2183 | 177.1622 | 77.6195 | 318.3451 | 324.6726 | 323.1445 | 286.5811 | 128.3836 | 127.2006 | 140.0738 | 137.2536 | 75.4189 | 75.5457 | 83.3628 | 119.3451 | 239.0208 | 279.525 | 277.3628 | 194.6341 | 141.3098 | 157.4543 | 173.7817 | 117.7287 | 167.6195 | 180 | 263.5457 | 155.5811 | 162.764 | 154.0915 | 177.9262 | 133.2006 | 184.8171 | 203.9823 | 283.1445 | 226.9116 | 174.4366 | 158.764 | 308 | 233.9085 | 224.5634 | 220.6372 | 212.0561 | 300.4574 | 179.7817 | 194.6902 | 189.6902 | 238.8909 | 150.3098 | 149.3274 | 142.9823 | 162.6726 | 233.4719 | 198.6195 | 210.5634 | 135.6195 | 109 | 145.4719 | 128.1091 | 124.236 | 331.9439 | 320.0915 | 326.6726 | 401.8555 | 278.9293 | 296.9085 | 285.764 | 397.9616 | 119.5281 | 67.6372 | 75.0354 | 119.5281 | 107.3098 | 121.9823 | 124 | 119.4719 | 104.7817 | 110.4719 | 119.9085 | 121.3628 | 86.7817 | 81.6549 | 76 | 88.2006 | 114.9262 | 91.1829 | 105.3098 | 80.1445 | 277.9823 | 260.9439 | 299.4896 | 363.0884 | 168.764 | 253.3805 | 219.9116 | 323.0884 | 379.4927 | 262.9262 | 273.1829 | 395.053
1394 | 17.3284 | 161.8141 | 235.7856 | 261.9715 | 184.2429 | 154.5427 | 181.943 | 213.5427 | 167.6357 | 139.2144 | 128.057 | 121.3643 | 156.7001 | 186.036 | 167.4498 | 156.5787 | 223.0855 | 181.1499 | 166.6357 | 161.4573 | 187.8786 | 174.943 | 376.1859 | 390.6716 | 178.6357 | 227.5862 | 231.886 | 292.1499 | 307.7571 | 117.2144 | 137.7001 | 126.0285 | 129.6716 | 167.9145 | 132.5787 | 152.2999 | 99.3643 | 308.7497 | 321.8786 | 311.1574 | 276.4573 | 155.7571 | 137.7286 | 156.1499 | 126.886 | 85.5427 | 77.0855 | 69.1499 | 114.9715 | 275.6282 | 243.2429 | 264.9715 | 163.1289 | 151.7571 | 156.5217 | 170.7856 | 141.943 | 188.3928 | 182.0644 | 265.9356 | 146.7931 | 169.4858 | 161.6642 | 163.0644 | 142.6357 | 167 | 211.7571 | 264.8426 | 266.7571 | 178.4858 | 168.036 | 306.4213 | 223.5427 | 219.057 | 233.0285 | 235.5427 | 339.7782 | 179.7001 | 183.0285 | 180.3358 | 238.3643 | 160.2714 | 154.0644 | 154.036 | 161.7001 | 220.2714 | 220.8501 | 203.8426 | 156.7571 | 111.5502 | 132.8786 | 130.3358 | 122.8575 | 307.8501 | 328.3928 | 324 | 415.4213 | 312.2144 | 292.979 | 293.8216 | 366.7001 | 131.7571 | 79.1784 | 59.3643 | 131.7571 | 118 | 129.0285 | 124.1214 | 107.4858 | 105.6716 | 99.2144 | 110.2714 | 108.8501 | 83.4213 | 86.6357 | 75.5142 | 98.1214 | 98.8501 | 107.7856 | 113.5712 | 67.5502 | 288.5502 | 240.8575 | 277.9071 | 325.7931 | 177.9145 | 232.8501 | 256.2354 | 288.9505 | 407.0285 | 251.1004 | 291.7286 | 374.6431
1395 | 17.3394 | 146.6686 | 225.2615 | 273.6805 | 187.3373 | 140.1219 | 192.6568 | 201.2615 | 166.2036 | 148 | 116.1397 | 123.5988 | 168.0639 | 207.1976 | 152.0758 | 171.8781 | 204.942 | 176.8722 | 167.7444 | 171.3314 | 188.6048 | 186.058 | 389.7266 | 410.8024 | 176.1337 | 266.7964 | 257.1917 | 286.6686 | 304.9302 | 126.4012 | 148.3952 | 121.1397 | 148.4651 | 185.2556 | 145.7385 | 146.0178 | 102.4012 | 277.0639 | 320.7325 | 328.5171 | 290.8781 | 152.529 | 147.8143 | 151.8722 | 152.3254 | 76.8722 | 58.5409 | 63.6686 | 120.529 | 248.1456 | 246.6048 | 272.4012 | 191.9302 | 149.5988 | 177.5869 | 165.477 | 153.7266 | 186.8083 | 195.3952 | 252.471 | 166.7325 | 168.4532 | 150.9302 | 179.8722 | 139.3314 | 169.0059 | 210.5349 | 242 | 271.0178 | 180.9302 | 187.9941 | 294.7325 | 213.5349 | 208.6107 | 226.1337 | 226.3373 | 306.1456 | 193.3373 | 177.8722 | 190.8024 | 243.6746 | 155 | 165.3893 | 171.4532 | 174.5349 | 215.1337 | 224.4591 | 184.3432 | 154.3314 | 128.7964 | 131.0639 | 139.3314 | 155.5929 | 317.2095 | 330.1515 | 321.9941 | 403.5988 | 318.058 | 326.1917 | 305.6627 | 375.3893 | 132.1397 | 66.8024 | 64.942 | 132.1397 | 116.9302 | 121.1976 | 123.1278 | 116.6166 | 122.1917 | 106.9302 | 107.0059 | 115.5349 | 73.8722 | 85.3373 | 70.9302 | 96.9941 | 104.0639 | 103.0119 | 95.7325 | 82.3893 | 302.052 | 275.3314 | 271.471 | 346.4012 | 196.9941 | 238.4651 | 229.1337 | 330.1278 | 398.2615 | 288.4012 | 298.2036 | 403.1397
1396 | 17.3504 | 152.0221 | 221.447 | 298.0132 | 192.0575 | 124.9294 | 173.4823 | 192.1903 | 173.708 | 154.1328 | 129.708 | 119.2478 | 164.3363 | 200.8318 | 169.5177 | 195.9779 | 220.3495 | 188.708 | 190.3938 | 163.7876 | 197.3938 | 167.292 | 377.1903 | 415.854 | 178.3142 | 257.2478 | 237.9779 | 289.3938 | 298.146 | 129.5841 | 134.9336 | 136.3142 | 146.4602 | 161 | 157 | 183.1017 | 102.0796 | 271.4381 | 319.5841 | 296.0664 | 313.8097 | 141.2478 | 175.2035 | 164.292 | 136.4602 | 86.3716 | 75 | 70.0443 | 109.3938 | 269.7433 | 258.4602 | 278.2124 | 183.9779 | 149.4823 | 150.4602 | 183.9114 | 134.9114 | 200.9114 | 186.8983 | 259.0796 | 164.708 | 140.3142 | 144 | 189.9557 | 135.146 | 181.2699 | 214.8761 | 246.6726 | 304.7433 | 174.292 | 175.7301 | 292.854 | 218.1682 | 225.7211 | 229.0443 | 236.0221 | 330.8097 | 200.1017 | 196.1328 | 199.6505 | 259.1017 | 157.7744 | 146.6726 | 146.2346 | 175.5177 | 216.8761 | 207.2699 | 202.9336 | 149.1239 | 121 | 123.1017 | 137.6284 | 137.3716 | 334.0575 | 363.8672 | 310.6062 | 398.8097 | 301.6284 | 304.4955 | 295.3938 | 357.8408 | 144.2478 | 73.1682 | 82.1017 | 144.2478 | 112.9204 | 119.6505 | 111 | 149.2035 | 104.292 | 102.0443 | 119.2699 | 118.854 | 85.854 | 91.5177 | 64.5841 | 86.1903 | 99.0221 | 128.2699 | 94 | 62.5045 | 275.6948 | 270.5619 | 282.7522 | 347.3938 | 180.9336 | 231.7876 | 233.7965 | 318.146 | 389.6284 | 287.7876 | 302.7876 | 421.5266
1397 | 17.3614 | 162.1161 | 258.8259 | 260.3087 | 174.3826 | 179.0422 | 189.5831 | 205.3166 | 172.6332 | 189.2084 | 130.8496 | 110.7414 | 176.1003 | 197.0079 | 153.1003 | 189.5251 | 203.2744 | 187.3166 | 178.942 | 151.6491 | 190.058 | 172.1662 | 384.4591 | 413.8918 | 183.6253 | 241.934 | 240.2322 | 281.2665 | 302.0079 | 130.467 | 119.6913 | 143.2084 | 149.7757 | 160.5251 | 156.0501 | 171.2084 | 85.1583 | 277.4828 | 320.467 | 320.9657 | 299.4169 | 137.6491 | 151.4169 | 160.7757 | 143.5752 | 71.9499 | 76.1082 | 82.9499 | 102.058 | 242.7995 | 261.1425 | 296.0501 | 180.6913 | 159.8839 | 158.6834 | 162.8496 | 114.6412 | 180.0079 | 198.3166 | 240.5752 | 163.4749 | 148.1583 | 144.1583 | 177.3668 | 136.1583 | 177.3166 | 221.2665 | 278.9076 | 271.1504 | 177.4248 | 180.4749 | 290.8918 | 226.2665 | 194.6992 | 238.7836 | 241.3668 | 323.8575 | 187.7335 | 225.9842 | 216.5673 | 246.5752 | 166.876 | 171.942 | 168.467 | 164.4828 | 216.1425 | 207.2744 | 189.9076 | 149.2243 | 121.9499 | 112.6332 | 150.7836 | 124.3747 | 311.1583 | 332.9076 | 325.6992 | 384.8918 | 317.4749 | 291.124 | 292.0158 | 385.9842 | 137.6412 | 77.1504 | 74.0079 | 137.6412 | 130.7916 | 137.5172 | 110.5251 | 129.6913 | 108.8496 | 115.4248 | 115.6332 | 116.4169 | 83.8918 | 76.3668 | 67.0501 | 98.6834 | 104.8417 | 123.8417 | 94.9499 | 82.6834 | 306.1504 | 270.058 | 293.1583 | 339.8997 | 164.1082 | 217.7493 | 255.0501 | 321.6913 | 403.4169 | 270.7414 | 285.5831 | 447.0501
1398 | 17.3724 | 180.2467 | 227.706 | 250.7769 | 196.1759 | 158.5118 | 200.559 | 204.6116 | 179.2414 | 181.2467 | 142.4174 | 121.1942 | 169.5826 | 218.2887 | 144.3648 | 183.4174 | 222.7586 | 190.8766 | 167.3176 | 181.8819 | 204.7533 | 183.3701 | 364.8949 | 408.1706 | 174.0708 | 220.1181 | 277.5171 | 287.8294 | 320.3884 | 119.7296 | 131.9528 | 137.6352 | 121.4882 | 158 | 150.147 | 168.706 | 85.4882 | 292.0761 | 309.9002 | 351.5171 | 287.7586 | 167.7113 | 144.0236 | 137.9475 | 133.0998 | 77.3412 | 81.8294 | 87.8294 | 117.6063 | 257.3701 | 237.8477 | 289.294 | 191.588 | 139.8766 | 159.559 | 175.441 | 133.3176 | 192.9292 | 197.2704 | 237.6063 | 170.0945 | 149.6824 | 143.9764 | 176.7296 | 136.8294 | 177.1234 | 226.294 | 301.2467 | 254.6352 | 186.706 | 180.441 | 285.1706 | 234.5354 | 224.6588 | 227.3412 | 237.4882 | 351.588 | 183.0472 | 197.412 | 194.8766 | 239.3412 | 132.7533 | 162.0236 | 156.0236 | 181.4646 | 188.5826 | 220.6169 | 214.6352 | 174.6116 | 127.853 | 116 | 140.0236 | 136.4882 | 306.3701 | 352.0053 | 352.7586 | 392.1364 | 322.2178 | 327.1234 | 328.6063 | 357.2414 | 151.8819 | 64.559 | 84.7113 | 151.8819 | 132.6116 | 119.5118 | 111.2414 | 141.7822 | 121.9528 | 121.1234 | 115.9292 | 109.0236 | 79.8766 | 75.2178 | 65.9239 | 97.6824 | 102.4646 | 120.9528 | 98.9764 | 79.4646 | 292.0236 | 275.1995 | 287.5171 | 351.5354 | 171.0236 | 244.3648 | 251.5354 | 336 | 393.8058 | 277.9528 | 293.8294 | 445.5826
1399 | 17.3834 | 155.9645 | 237.8289 | 276.7434 | 168.3039 | 167.0355 | 207.1474 | 194.8289 | 195 | 164.8289 | 129.2065 | 124.0737 | 185.7197 | 236.5014 | 154.1092 | 163.6224 | 213.9499 | 195.2448 | 162.354 | 157.9145 | 222.6342 | 156.1829 | 407.6961 | 409.5487 | 185.1592 | 247.7434 | 252.2184 | 284.4395 | 334.0118 | 135.0118 | 125.1092 | 132.8289 | 123.2065 | 158.5487 | 150.5723 | 172.9763 | 82.6342 | 272.4513 | 323.4513 | 315.7935 | 274.7434 | 144.5605 | 148.4513 | 163.8289 | 120.8289 | 81.5605 | 78.9882 | 87.3658 | 137.4395 | 240.9736 | 273.2566 | 287.5842 | 176.0974 | 150.4631 | 173.2803 | 165.5605 | 127.6224 | 179.0974 | 187.1092 | 259.8171 | 186.7079 | 152.8171 | 144.4868 | 187.6224 | 140.0237 | 171.6579 | 221.6579 | 280.8053 | 248.1829 | 195.3658 | 175.1328 | 286.3658 | 238.525 | 220.2566 | 227.9026 | 235 | 334.0855 | 192.6342 | 214 | 203.8171 | 245.9381 | 155.7552 | 165.1711 | 161 | 173.4513 | 207.0974 | 178.6579 | 209.8289 | 160.8053 | 131.8171 | 116.3658 | 146.8289 | 134 | 287.7789 | 322.5723 | 340.1092 | 451.2448 | 333.1829 | 315.0737 | 344.9645 | 376.6579 | 130.475 | 74.2566 | 62.475 | 130.475 | 120.2684 | 121.2684 | 125.7197 | 129.354 | 115.475 | 110.1711 | 105.0237 | 116.7434 | 84.4277 | 84.5369 | 85 | 97.3421 | 101.0355 | 114.2921 | 97.2921 | 71.2684 | 294.6224 | 252.8526 | 259.475 | 357.9026 | 179.1092 | 249.1711 | 256.9882 | 337.4631 | 387.2684 | 263.6106 | 291.7197 | 451.4777
1400 | 17.3944 | 135.6363 | 244.2433 | 284.5108 | 194.147 | 191.853 | 182.3663 | 196.7541 | 189.1445 | 170.853 | 143.7059 | 114.5133 | 178.8289 | 207.0989 | 162.9252 | 155.9278 | 240.5108 | 178.2941 | 159.8796 | 164.147 | 217.0963 | 162.0748 | 391.417 | 413.5615 | 171.123 | 254.3397 | 277.2674 | 279.2459 | 338.1204 | 144.9759 | 140.7567 | 141.9759 | 146.4892 | 155.3397 | 168.4626 | 156.5615 | 82.5615 | 270 | 322.1711 | 305.1977 | 285.6337 | 154.4385 | 144.0482 | 169.0722 | 133.2941 | 91.0482 | 70.5856 | 88.2193 | 125.2193 | 286.4867 | 261.7807 | 312.3156 | 183.9278 | 161.4892 | 172.1685 | 169.583 | 117.1952 | 187.9037 | 204.7085 | 255.2915 | 172.1952 | 154.5374 | 169.9759 | 181.4892 | 157.0241 | 183.7059 | 236.6337 | 271.1952 | 256.2218 | 189.583 | 203.4867 | 286.8289 | 223.5856 | 207.6096 | 220.6578 | 227.3878 | 334.6844 | 192.5615 | 210.6819 | 201.4385 | 267.0241 | 168.8023 | 160.5133 | 163.9278 | 171.5856 | 211.4144 | 201.8315 | 214.2915 | 150.6096 | 135.6844 | 120.3422 | 151.2915 | 136.5374 | 326.0241 | 342.2193 | 356.147 | 433.3181 | 330.6819 | 309.2218 | 322.2941 | 394.3663 | 139.7541 | 66.6844 | 77.6096 | 139.7541 | 114.0722 | 115.853 | 118.0482 | 125.9037 | 127.6819 | 103.7567 | 120.6578 | 127.4385 | 65.5615 | 77.6096 | 81.877 | 82.7567 | 122.3397 | 124.5108 | 108.0963 | 71.5133 | 284.7807 | 278.8048 | 271.4867 | 346.9493 | 190.6578 | 237.4867 | 251.3181 | 337.7541 | 387.123 | 252.5881 | 289.123 | 408.9785
1401 | 17.4054 | 165.2822 | 236.3776 | 271.3944 | 202.1701 | 182.1701 | 190.4897 | 180.9588 | 171.4317 | 161.7925 | 128.6972 | 126.3028 | 171.5103 | 222.1327 | 144.9047 | 161.1533 | 222.8299 | 196 | 172.7384 | 172.7925 | 206.6018 | 187.9794 | 417.2281 | 411.701 | 182.7551 | 228.9252 | 258.9458 | 298.7757 | 318 | 142.9832 | 143.8505 | 143.9252 | 163.1327 | 135.1121 | 156.3402 | 166.7346 | 85.4729 | 277.6766 | 324.7178 | 337.3402 | 281.8673 | 148.2075 | 137.6598 | 155.7551 | 148.0953 | 82.7925 | 74.4523 | 85.6224 | 122.415 | 271.3776 | 260.8505 | 300.0542 | 199.3196 | 180.415 | 149.6018 | 143.9794 | 121.3196 | 201.0953 | 220 | 237.3028 | 177.1495 | 164.7925 | 171.7178 | 195.4355 | 150.7178 | 170.1495 | 228.0953 | 272.2075 | 290.357 | 163.1495 | 188.1701 | 279.3028 | 228.4897 | 214.299 | 210.1701 | 208.4485 | 353.7925 | 197.9626 | 191.8131 | 195.2075 | 258.4355 | 142.2616 | 173.5477 | 171.8505 | 177.5271 | 211.6972 | 230.6224 | 200.8673 | 149.2075 | 149.3982 | 133.942 | 135.585 | 145.1327 | 321.3776 | 332.9832 | 366.6598 | 438.286 | 312.4355 | 334.8505 | 335.4355 | 403.7346 | 118.357 | 84.1327 | 75.7925 | 118.357 | 103.6598 | 110.5271 | 113.1121 | 139.7178 | 113.7925 | 108.5103 | 110.585 | 118.0953 | 74.2822 | 76.6224 | 71.6972 | 91.4523 | 99 | 102.6804 | 98.2243 | 81.0206 | 283.2281 | 277.7925 | 261.1495 | 330.299 | 179.3402 | 234.4112 | 261.4729 | 318.6392 | 397.7178 | 279.4935 | 296.8131 | 439.1533
1402 | 17.4164 | 172.2419 | 232.4616 | 286.5058 | 198.3407 | 180.5384 | 201.5384 | 215.0442 | 190.7256 | 162.9779 | 139.2198 | 114.6814 | 162 | 215.2198 | 155.5605 | 140.4942 | 234.7907 | 195.7802 | 146.7581 | 171.5605 | 229.0988 | 198.4291 | 397.4616 | 376.4837 | 182.1756 | 221.8791 | 238.1535 | 282.0325 | 318.8791 | 120.6593 | 129.3186 | 135.3407 | 152.4837 | 147.4395 | 149.1209 | 181.2198 | 72.4395 | 302.3849 | 318.6372 | 331.4395 | 282.8465 | 155.1535 | 143.3407 | 147.9233 | 135.0221 | 79.5826 | 80.8791 | 85.8023 | 121.8023 | 272.9558 | 253.5709 | 328.3186 | 214.4174 | 175.6268 | 169.4616 | 165.6372 | 132.9012 | 193.0767 | 213.6268 | 243.4825 | 191.0221 | 167.0767 | 161.0221 | 178.8791 | 141.7802 | 187.9779 | 226.0104 | 273 | 304.8023 | 184.2744 | 185 | 280.8674 | 234.7035 | 241.1651 | 215.7907 | 252.9233 | 347.9454 | 195.7581 | 181.2302 | 205.0104 | 242.3186 | 166.7035 | 170.6372 | 158.6372 | 180.2314 | 227.0546 | 227.0663 | 213.4942 | 154.3953 | 128.8791 | 148.3407 | 134.8791 | 141.9558 | 326.0325 | 317.4721 | 369.0442 | 415.1093 | 294.5605 | 314.6047 | 323.0546 | 409.4291 | 138.5163 | 75.2419 | 77.4174 | 138.5163 | 105.8244 | 122.6814 | 125.8791 | 130.5605 | 114.9779 | 104.9337 | 108.7802 | 113.3732 | 81.0221 | 79 | 81.5605 | 100.5163 | 106.9116 | 91.3186 | 120.9012 | 66.1756 | 269.8349 | 278.7581 | 286.2302 | 357.3849 | 174.0988 | 277.5605 | 255.9116 | 345.9779 | 389.8791 | 250.2965 | 278.7581 | 412.3407
1403 | 17.4274 | 163.6794 | 231.4091 | 265.3756 | 195.7679 | 182.055 | 202.3588 | 201.4641 | 173.1603 | 173.945 | 138.8397 | 110.9282 | 163.1603 | 225.5144 | 154.4641 | 168.5861 | 256.0215 | 193.3756 | 159.0335 | 169.7679 | 226.0383 | 175.6579 | 393.6244 | 369.1053 | 194.6794 | 238.4593 | 263.945 | 293.5407 | 324.5526 | 129.7297 | 132.1435 | 142.0502 | 142.5526 | 151.0885 | 143.6794 | 175.2703 | 79.8014 | 290.8732 | 326.6077 | 337.4091 | 268.89 | 169.8062 | 143.3206 | 139.0885 | 126.3756 | 73.0885 | 79.5909 | 80.7847 | 114.4641 | 291.4091 | 281.9115 | 332.7679 | 217.4306 | 157.8732 | 166.0885 | 174.3038 | 126.6794 | 206.7847 | 194.945 | 301.7679 | 183.3038 | 169.4139 | 154.4641 | 183.1603 | 147.7297 | 191.9833 | 253.1268 | 277.177 | 296.5359 | 204.1603 | 184.7679 | 320.9115 | 228.8014 | 229.5694 | 240.5024 | 241.5359 | 328.8397 | 205.945 | 206.3421 | 233.2871 | 246.3038 | 161.9617 | 177.9115 | 169.6244 | 152.2105 | 242.2153 | 214.8182 | 224.878 | 179.2823 | 134.5526 | 141.823 | 132.8947 | 153.6794 | 344.2703 | 343.3038 | 361.7297 | 448.4474 | 308.3158 | 307.3541 | 335.6627 | 375.5191 | 148.2153 | 72.0167 | 82.2871 | 148.2153 | 94.4809 | 119.1603 | 128.0718 | 127.6077 | 124.7847 | 121.1268 | 108.4641 | 133.6794 | 73.5359 | 80.6244 | 78.8397 | 116.945 | 128.5024 | 100.6412 | 115.6077 | 82.1603 | 286.3038 | 281.9833 | 317.3756 | 345.6412 | 184.0335 | 279.4809 | 287.0167 | 349.2871 | 399.9617 | 261.0718 | 289.177 | 411.3923
1404 | 17.4384 | 164.0635 | 240.6243 | 268.7966 | 199.1539 | 166.0678 | 188.9548 | 207.3983 | 177.4887 | 186.7557 | 136.2217 | 121.5748 | 174.8192 | 251.6243 | 158.1991 | 191.8687 | 225.9322 | 189.4661 | 178.0226 | 167.2896 | 207.1539 | 194.8913 | 396.5565 | 396.7513 | 189.4435 | 279.0904 | 273.0904 | 260.7287 | 331.0452 | 148.8235 | 128.6878 | 163.4478 | 156.3757 | 159.4661 | 142.1087 | 150.8009 | 94.8461 | 309.1991 | 319.3122 | 349.8009 | 290.3574 | 150.3983 | 151.733 | 145.0226 | 128.8192 | 76.8235 | 69.1539 | 85.3574 | 116.4887 | 306.2443 | 276.4661 | 327.113 | 203.1539 | 164.9591 | 168.8461 | 170.2896 | 125.3531 | 207.9365 | 208.9774 | 300.7557 | 184.1765 | 146.8644 | 160.887 | 192.3757 | 170.2443 | 185.6652 | 239.9322 | 283.9139 | 288.6469 | 200.4252 | 187.6652 | 324.0183 | 247.5113 | 247.7557 | 230.4844 | 239.267 | 325.9774 | 208.7374 | 190.0409 | 218.3122 | 248.6426 | 183.5339 | 171.9774 | 173.5339 | 186.4887 | 235.4435 | 246.9322 | 201.9096 | 204.38 | 140.3122 | 129.4661 | 123.5748 | 150.1539 | 325.6652 | 343.1991 | 380.0904 | 439.5113 | 354.8461 | 331.3348 | 322.1539 | 368.1539 | 138.7557 | 79.3122 | 72.9322 | 138.7557 | 110.3983 | 125.1991 | 122.8009 | 132.0409 | 126.4252 | 107.4435 | 108.0452 | 131.1313 | 70.5339 | 86.733 | 77.4435 | 127.3122 | 108.2217 | 111.6017 | 108.5565 | 85.7557 | 282.0452 | 272.733 | 316.3983 | 355.6243 | 201.0678 | 285.38 | 297.733 | 332.8461 | 414.9365 | 266.7966 | 307.6426 | 414.0452
1405 | 17.4494 | 189.2566 | 224.2566 | 288.8141 | 201.4778 | 158.7965 | 194.2301 | 219.9734 | 182.8496 | 182.1504 | 139.2301 | 143.4602 | 196.177 | 228.8406 | 168.0797 | 174.9734 | 245.5222 | 198.6195 | 173.9734 | 165.8496 | 214.3539 | 172.6637 | 391.0531 | 431.7965 | 190.584 | 266.7434 | 263.0531 | 290.8406 | 326.5398 | 142.593 | 143.9293 | 137.7699 | 167.6106 | 163.7434 | 163.3098 | 151.4426 | 86.1062 | 305.4778 | 308.6902 | 335.3008 | 282.416 | 159.3805 | 152.9734 | 144.5664 | 153.5133 | 78.2921 | 80.7168 | 69.2035 | 122.6195 | 298.2742 | 280.7434 | 315.0797 | 208.0442 | 148.2921 | 155.4867 | 166.5398 | 139.8672 | 190.9558 | 203.5309 | 303.3363 | 197.0797 | 168.8496 | 168.3008 | 212.0797 | 162.7875 | 200.0797 | 245.1504 | 269.4426 | 263.3539 | 182.6461 | 198.4867 | 343.4512 | 235.9911 | 237.2477 | 265.7434 | 236.4867 | 326.4336 | 181.8762 | 216.2035 | 209.4867 | 264.2832 | 176.6902 | 172.6902 | 169.7699 | 185.9469 | 244.5398 | 248.8141 | 212.9734 | 189.0531 | 134.5664 | 136.0531 | 142.1238 | 155.8141 | 345.2125 | 357.1859 | 368.5133 | 435.9469 | 343.5398 | 325.6461 | 333.4602 | 394.3453 | 137.7434 | 71 | 77.3805 | 137.7434 | 124 | 124.0442 | 119.0797 | 156.1504 | 108.6461 | 114.4867 | 105.593 | 142.3805 | 70.1062 | 85.407 | 83.9734 | 111.8141 | 116.3629 | 103.6461 | 101.7699 | 87.3098 | 283.1859 | 276.5398 | 327.4336 | 347.9824 | 196.6195 | 280.5753 | 301.7965 | 328.9824 | 401.8054 | 294 | 318.1504 | 408.2566
1406 | 17.4604 | 187.0418 | 225.5379 | 262.6214 | 189.0184 | 166.1514 | 187.4282 | 210.0078 | 181.8981 | 173.4204 | 133.7728 | 134.1592 | 187.7311 | 215.5718 | 170.8146 | 172 | 265.7806 | 209.5796 | 167.7728 | 174.3107 | 207.1436 | 170.5301 | 397.8068 | 430.5457 | 175.765 | 275.6815 | 265.5039 | 266.2767 | 325.8903 | 148.1592 | 153.2428 | 138.6553 | 151.7311 | 172.1436 | 169.4621 | 171.1175 | 97.4621 | 281.7233 | 313.0262 | 324.833 | 303.9582 | 153.799 | 155.9165 | 158.7233 | 159.5718 | 105.8903 | 77 | 67.4961 | 129.2767 | 272.1932 | 276.7728 | 330.1854 | 184.1932 | 181.5379 | 162.0679 | 171 | 130.6475 | 222.6058 | 191.1019 | 300.3603 | 199.2767 | 184.034 | 155.1436 | 222.8825 | 141.6893 | 201.4699 | 241.5301 | 296.6475 | 245.9243 | 191.201 | 200.2272 | 323.4465 | 214.799 | 205.235 | 260.1592 | 229.6214 | 325.4543 | 195.1932 | 209.3864 | 208.5379 | 273.6475 | 169.7311 | 169.4961 | 176.8408 | 178.3864 | 247.9243 | 230.9582 | 215.9165 | 196.8825 | 138.235 | 141.7311 | 127.7311 | 136.8825 | 364.6971 | 387.1436 | 370.2689 | 439.1436 | 345.3107 | 350.3369 | 332.4961 | 426.0078 | 132.966 | 74.765 | 67.4961 | 132.966 | 125.8825 | 108.5301 | 126.9243 | 137.7389 | 129.5718 | 117.5718 | 116.8068 | 134.1097 | 80.6553 | 74.1932 | 80.4621 | 97.1854 | 128.1932 | 118.6553 | 110.9922 | 88.8903 | 295.1253 | 295.7912 | 315.4282 | 366.4699 | 204.8903 | 305.167 | 303.5039 | 363.6058 | 431.966 | 294.8068 | 310.4961 | 415.4543
1407 | 17.4714 | 175.0629 | 230.6559 | 257.2803 | 230.407 | 173.6873 | 175.2812 | 197.1869 | 158.8428 | 183.0934 | 132.0305 | 120.7188 | 190.0934 | 222.938 | 160.4681 | 171.7188 | 238.8437 | 213.6864 | 158.7188 | 175.4367 | 228.0629 | 189.8751 | 399.4376 | 416.9371 | 183.469 | 287.6577 | 259.6559 | 260.9362 | 316.2183 | 129.938 | 135.5301 | 139.7808 | 150.4376 | 182.3765 | 159.8446 | 168.8122 | 99.6559 | 301.9686 | 343 | 355.5939 | 309.7502 | 166.5633 | 172 | 174.7502 | 166.938 | 92 | 82.3432 | 75.5939 | 120.5615 | 276.1869 | 272.4995 | 348.9686 | 187.062 | 182.1563 | 188.2183 | 165.6568 | 140.938 | 233.0638 | 218.6568 | 274.9991 | 187.7493 | 188.2507 | 173.8131 | 219.2821 | 146.469 | 189.7179 | 261.7188 | 314.2498 | 254.8113 | 178.4367 | 200.5633 | 354.0943 | 235.7188 | 199.2183 | 242.7817 | 216.4061 | 340.469 | 191.3127 | 210.062 | 209.7188 | 291.2498 | 169.8437 | 182.9371 | 186.0629 | 175.4061 | 254.5615 | 246.3118 | 242.124 | 195.5319 | 129.4061 | 139.8751 | 134.0305 | 142.2812 | 345.1258 | 405.2507 | 370.7188 | 456.1258 | 334.9066 | 364.0009 | 342 | 408.9686 | 122.8437 | 80.5005 | 74.469 | 122.8437 | 133.2498 | 129 | 124.8437 | 118.2812 | 147.062 | 128.5939 | 117.8751 | 140.9695 | 82.6244 | 75.9371 | 79 | 113.4376 | 121.7817 | 113.5939 | 121.2821 | 80.9057 | 275.4995 | 323.345 | 293.7197 | 356.9677 | 190.4376 | 284.8113 | 288.6568 | 377.7197 | 420.7188 | 288.0009 | 313.8131 | 436.6559
1408 | 17.4824 | 167.6416 | 242.3481 | 282.7611 | 209.5324 | 174.4778 | 176.2935 | 214.1741 | 184.8805 | 188.6519 | 145.3038 | 121.1741 | 200.0546 | 217.9248 | 172.0103 | 168.9454 | 240.413 | 232.1946 | 165.9248 | 186.14 | 222.4027 | 183.1843 | 400.9351 | 426.7065 | 177.2935 | 248.3584 | 282.2081 | 292.5324 | 333.4573 | 120.8157 | 160.5973 | 147.3038 | 156.0443 | 147.6416 | 149.86 | 176.8805 | 108.413 | 302.3038 | 337.7168 | 348.1843 | 311.3378 | 153.5119 | 168.7714 | 165.1843 | 151.3584 | 82.3141 | 88.9557 | 70.2389 | 140.8908 | 294.6416 | 295.1535 | 350.4778 | 200.8805 | 177.3584 | 187.553 | 161.9794 | 127.7065 | 209.6859 | 203.8259 | 291.5427 | 202.0649 | 182.1535 | 164.4573 | 211.6756 | 147.3378 | 206.4233 | 255.7168 | 313.5427 | 286.4676 | 195.3038 | 189.273 | 337.5324 | 244.3924 | 215.5767 | 230.8259 | 234.3821 | 341.3378 | 191.6962 | 222.1195 | 204.3038 | 302.5767 | 174.3481 | 192.7065 | 170.4233 | 182.8157 | 276.0649 | 264.1195 | 262.7168 | 181.9351 | 126.8362 | 141.6962 | 146.7168 | 144.1741 | 334.3924 | 386.5324 | 381.447 | 446.5665 | 331.9897 | 338.9454 | 351.3924 | 414.8805 | 135.5665 | 71.0546 | 73.5767 | 135.5665 | 133.7168 | 133.6962 | 130.2286 | 124.8703 | 153.8362 | 124.1195 | 117.3481 | 115.6622 | 88.4881 | 83.0649 | 81.6416 | 110.5324 | 116.5767 | 111.4676 | 104.2832 | 96.2935 | 290.2286 | 298.0443 | 288.7848 | 371.4984 | 194.5767 | 309.7168 | 282.6313 | 357.2286 | 429.3924 | 271.7508 | 302.1092 | 441.3038
1409 | 17.4934 | 167.2723 | 247.0826 | 290.058 | 206.3058 | 178.2322 | 185.2567 | 214.2478 | 184.5536 | 188.5045 | 138.2813 | 127.058 | 200.7187 | 237.6116 | 164.5871 | 169.1987 | 242.6696 | 206.1071 | 182.5536 | 165.5045 | 232.6942 | 179.8103 | 419.3147 | 425.3884 | 186.8683 | 250.8683 | 322.2567 | 294.8103 | 343.4955 | 137.6451 | 146.6362 | 139.9755 | 160.1562 | 158.5871 | 170.8013 | 175.942 | 100.5781 | 300.1741 | 327.4464 | 334.1071 | 319.8259 | 166.0491 | 163.7522 | 160.5871 | 138.8839 | 67.2567 | 67.7187 | 65.3884 | 131.6116 | 301.6116 | 308.2388 | 345.6696 | 204.2232 | 168.2478 | 160.6116 | 191.8103 | 124.2478 | 231.7768 | 197.942 | 279.9755 | 201.1161 | 200.7433 | 184.8929 | 236.2969 | 168.058 | 199.1161 | 242.6942 | 293.1071 | 296.1161 | 185.529 | 207.2723 | 333.3884 | 267.9174 | 237.2567 | 230.4464 | 270.8348 | 359 | 198.1071 | 211.7433 | 201.8683 | 302.8504 | 187.3393 | 188.0245 | 159.7522 | 189.8594 | 271.1406 | 262.6116 | 253.3638 | 196.9509 | 120.5625 | 146.8839 | 135.2232 | 149.1406 | 353.6362 | 385.1406 | 400.7433 | 462.5201 | 345.2232 | 345.9264 | 362.3795 | 416.3884 | 156.7187 | 85.4799 | 82.4129 | 156.7187 | 121.9174 | 132.462 | 141.6696 | 133.4955 | 135.942 | 121.3884 | 119.6362 | 105.0335 | 71.6696 | 81.8103 | 88 | 107.6116 | 121.4375 | 116.2232 | 119.1406 | 96.6942 | 295.5536 | 305.2143 | 323.6027 | 338.0826 | 211.9755 | 297.6116 | 299.471 | 361.0245 | 449.2478 | 295.3393 | 319.0335 | 441.9264
1410 | 17.5044 | 155.1809 | 249.4533 | 299.5448 | 210.499 | 202.1829 | 193.5029 | 217.5428 | 181.5448 | 195.2743 | 145.7733 | 128.2743 | 198.6343 | 227.5486 | 174.0457 | 174.6381 | 252.9086 | 205.8152 | 170.0019 | 171.6381 | 228.8191 | 187.5924 | 441.8191 | 417.6381 | 202.5467 | 271.9543 | 331.1391 | 304.1829 | 333.5448 | 148.6381 | 137.0914 | 151.5448 | 134.499 | 167.4095 | 159.9543 | 171.2267 | 94.2705 | 322.5905 | 328.8648 | 327.1352 | 302.1809 | 148.6362 | 167.4552 | 177.9981 | 131.0438 | 87.2724 | 62.4533 | 66.2267 | 128.8648 | 297.5924 | 285.2705 | 351.7733 | 211.7714 | 167.4076 | 166.1352 | 191.64 | 123.4076 | 220.4114 | 196.4076 | 289.9543 | 217.5448 | 179.7733 | 196 | 227.2686 | 177.5448 | 205.3657 | 238.501 | 288.9981 | 303.32 | 187.0914 | 194.8629 | 348.2228 | 271.5905 | 245.1848 | 243.3162 | 269.2743 | 362.1809 | 201.3143 | 207.3143 | 219.7733 | 286.0419 | 199.8648 | 187.2705 | 170.1352 | 177.6838 | 270.5924 | 268.4533 | 261.9543 | 215.3181 | 134.2743 | 139.0438 | 142.1352 | 154 | 340.9105 | 398.5886 | 379.4552 | 446.4076 | 349.5905 | 370.7733 | 344.2705 | 419.4533 | 151.7714 | 100.32 | 73.9086 | 151.7714 | 128.1352 | 110.3619 | 143.9562 | 131.4971 | 125.8191 | 117.4552 | 115.499 | 124.1809 | 80 | 88.6381 | 87.6819 | 110.2724 | 105.1809 | 124.0895 | 116.6838 | 96.6362 | 306.5409 | 292.2248 | 304.4533 | 342.3619 | 215.9105 | 310.1333 | 304.5886 | 353.2724 | 437.9105 | 308.8191 | 326.0933 | 471.5448
1411 | 17.5154 | 163.3215 | 258.6696 | 301.6961 | 220.9735 | 187.7227 | 179.59 | 224.4542 | 191.295 | 194.2331 | 140.0088 | 119.9646 | 211.3658 | 208.9469 | 168.3658 | 159.0177 | 260.6342 | 210.8288 | 158.3039 | 168.5724 | 233.2331 | 181.2862 | 410.8818 | 414.242 | 199.9381 | 272.0177 | 312.3127 | 302.9381 | 341.9735 | 128.3923 | 137.6254 | 159.3127 | 155.5458 | 167.6342 | 169.9293 | 168.4011 | 104.7315 | 320.7138 | 332.8851 | 340.9558 | 312.6342 | 151.3215 | 171.2508 | 184.0796 | 142.3481 | 90.3304 | 69.3569 | 67.3658 | 123.3039 | 287.3215 | 306.9646 | 363.5193 | 214.0707 | 177.6785 | 172.6873 | 159.3658 | 140.9469 | 213.5635 | 211.6342 | 302.5724 | 215.7315 | 178.3039 | 190.7138 | 256.0354 | 181.6785 | 198.9116 | 228.6873 | 298.3834 | 296.5989 | 188.2862 | 194.0619 | 363.5072 | 277.6431 | 233.5635 | 251.7315 | 257.6608 | 366.6873 | 220.4365 | 230.0707 | 215 | 297.1503 | 185.0531 | 205.9912 | 179.9912 | 173.8939 | 265.9381 | 281.3039 | 253.0973 | 210.3834 | 129.5989 | 157.6166 | 147.0354 | 157.6342 | 331.3215 | 417 | 376.6431 | 457.6696 | 346.0619 | 367.9823 | 352.7492 | 426.6873 | 153.0796 | 87.3215 | 78.6608 | 153.0796 | 127.7492 | 117.3923 | 139.2597 | 151.0265 | 124.6166 | 123.8939 | 125.6431 | 126.3746 | 86.2773 | 79.295 | 88.6519 | 110.0265 | 115.9646 | 133.0177 | 108.2685 | 95.6873 | 332.7492 | 311.0088 | 308.0531 | 360.6254 | 200.0442 | 327.0619 | 326.9646 | 361.2862 | 442.1978 | 305.9646 | 309.3127 | 472.3746
1412 | 17.5264 | 163.2586 | 269.3083 | 292.9439 | 218.4331 | 181.5109 | 202.1402 | 205.2243 | 184.9222 | 213.9159 | 140.0561 | 128 | 214.2803 | 224.0841 | 160.3146 | 161.5109 | 276.5669 | 188.7351 | 165.6854 | 175.9502 | 249.4891 | 188.5732 | 383.9313 | 433.6293 | 202.349 | 270.0561 | 312.8879 | 322.8255 | 356.9096 | 129.7351 | 146.3146 | 161.6013 | 170.6636 | 167.4611 | 180.1745 | 159.1682 | 91.7134 | 309.6573 | 347.349 | 350.6573 | 315.8879 | 158.7975 | 172.5235 | 165.3146 | 141.0561 | 91.3427 | 68.4548 | 65.1402 | 129.3146 | 298.9096 | 325.623 | 393.5732 | 212.7351 | 175.6854 | 176.2243 | 159.5389 | 146.1465 | 230.4891 | 216.2586 | 321.2586 | 199.2866 | 197.3364 | 172.8038 | 248.3427 | 180.3707 | 215.2306 | 239.7631 | 287.6854 | 313.3707 | 205.8535 | 191.3083 | 339.4205 | 269.9783 | 243.9783 | 252.0778 | 266.5389 | 374.3364 | 197 | 210.2306 | 214.6573 | 273.623 | 187.3924 | 206.6293 | 174.4611 | 190.4611 | 282.3987 | 285.9439 | 247.7912 | 210.3083 | 141.5732 | 163.8598 | 138.3146 | 152.3209 | 345.9937 | 408.0904 | 381.6573 | 465.5669 | 340.5669 | 364.8038 | 342.5669 | 423.028 | 148.3644 | 92.0561 | 80.6854 | 148.3644 | 114.1402 | 109.1402 | 142.6356 | 136.7757 | 131.5452 | 143.8879 | 132.028 | 121.7975 | 85.6356 | 84.2306 | 88.9159 | 105.7134 | 129.1402 | 134.4828 | 121.6293 | 88.2586 | 317.4268 | 302.4891 | 304.5669 | 373.4268 | 198.2243 | 318.8255 | 332.2586 | 371.6573 | 463.7477 | 308.1745 | 311.2586 | 466.4268
1413 | 17.5374 | 164.0345 | 269.3469 | 295.7444 | 200.5801 | 187.9655 | 206.6755 | 221 | 161.6795 | 201.9655 | 146.4848 | 124.8053 | 225.1258 | 230.7099 | 161.1298 | 165.4808 | 280.9311 | 224.0994 | 171.9696 | 156.8093 | 231.3205 | 181.2901 | 415.351 | 428.1603 | 179.645 | 270.4503 | 296.4808 | 332.9351 | 365.4808 | 156.2252 | 146.0649 | 154.5152 | 158.7139 | 157.6146 | 173.2596 | 174.5497 | 91.8053 | 310.7748 | 327.4848 | 343.6106 | 308.355 | 177.2292 | 149.2941 | 171.0994 | 143.2252 | 89.5152 | 72.7404 | 71.0954 | 122.3205 | 322.3894 | 336.7708 | 378.1258 | 225.3814 | 173.0954 | 188.4503 | 180.7444 | 133 | 225.4543 | 208.1603 | 327.7139 | 193.5152 | 212.4808 | 170.004 | 246.1603 | 185.4848 | 219.359 | 266.4503 | 301.7789 | 310.3205 | 212.9655 | 214.0649 | 359.7708 | 255.2941 | 227.5192 | 267.1562 | 273.5457 | 380.6066 | 209.7789 | 201.5801 | 214.7099 | 287.9655 | 208.2556 | 196.1907 | 174.1987 | 188.7789 | 285.5801 | 286.2596 | 281.4199 | 233.0649 | 139.2596 | 149.7404 | 144.4543 | 141.4888 | 371.1298 | 385.6755 | 394.8438 | 480.5801 | 353.0954 | 354.9046 | 352.9655 | 421.8053 | 169.7404 | 94.5801 | 80.9351 | 169.7404 | 124.355 | 119.7099 | 125.5192 | 128.0994 | 131.1643 | 144.8742 | 139.3245 | 140.5841 | 70.2941 | 84.0994 | 78.7404 | 107.2252 | 141.1298 | 142.5152 | 117.9351 | 80.8702 | 322.9351 | 299.2292 | 321.355 | 376.4503 | 209.0304 | 337.8093 | 320.2556 | 371.355 | 455.5841 | 308.004 | 317.359 | 472.2901
1414 | 17.5484 | 170.5929 | 239.5487 | 323.0576 | 204.2434 | 177.7345 | 205.8142 | 227.9779 | 172.3673 | 189.1637 | 152.1018 | 120.8363 | 218.5487 | 244.1196 | 164.6327 | 159.1815 | 265.6107 | 231.1239 | 191.2833 | 173.2036 | 223.5708 | 185.5089 | 405.916 | 429.6107 | 192.2213 | 265.1018 | 292.0178 | 329.531 | 352.5708 | 150.0619 | 154.6107 | 150 | 177.7964 | 168.2036 | 182.469 | 170.3495 | 89.6726 | 325.385 | 344.4868 | 341.916 | 311.9381 | 179.6769 | 159.9203 | 185.8363 | 138.8982 | 83.1637 | 65.7345 | 68.4071 | 114.5708 | 329.9601 | 319.6327 | 374.1196 | 203.6107 | 167.1018 | 181.6327 | 181.2477 | 134.1018 | 229.084 | 200.4292 | 331.3716 | 199.6505 | 212.0576 | 185.2256 | 255.3231 | 195.5089 | 227.8804 | 262.2036 | 332.7124 | 311.0178 | 199.7964 | 222.6107 | 340.0619 | 269.5929 | 244.1018 | 263.301 | 263.3053 | 347.531 | 231.1637 | 207.0797 | 217.1018 | 293.8939 | 194.531 | 196.6904 | 200 | 216.0399 | 295.1196 | 297.3053 | 288.0399 | 230.9601 | 145.1637 | 150.447 | 158 | 159.0221 | 373.531 | 380.4071 | 397.5972 | 478.7345 | 360.3231 | 350.2477 | 342.3673 | 415.6327 | 162 | 90.8982 | 84.5089 | 162 | 125.3673 | 126.1416 | 131.8185 | 146.5089 | 146 | 151.084 | 139.8185 | 156.8982 | 90.8363 | 93.6947 | 79.8142 | 106.5708 | 133.9822 | 135.7964 | 117.1018 | 82.5089 | 329.0797 | 311.9601 | 311.7167 | 366.6947 | 207.7124 | 364.1196 | 310.5708 | 359.5132 | 456.4733 | 323.5929 | 338 | 474.3053
1415 | 17.5594 | 157.3067 | 248.7885 | 330.7816 | 222.9706 | 175.584 | 220.2409 | 236.9636 | 176.7955 | 195.1092 | 148.3067 | 126.6569 | 227.7885 | 270.3137 | 163.6204 | 179.9342 | 262.5771 | 238.4593 | 193.4384 | 180.0364 | 226.1022 | 193.1022 | 433.2479 | 441.7591 | 194.5043 | 285.2184 | 303.409 | 319.409 | 366.1092 | 140.0658 | 154.2339 | 157.9706 | 163.7521 | 170.4818 | 178.5476 | 163.7661 | 102.8319 | 330.5771 | 371.6933 | 359 | 328.7661 | 163.3501 | 163.4593 | 176.8544 | 137 | 87.591 | 71.1751 | 76.7591 | 127.7297 | 328.6933 | 326.591 | 384.3726 | 221.0728 | 178.4888 | 169.2339 | 171.6274 | 145.4888 | 221.8683 | 200.9342 | 315.4229 | 217.2409 | 209.9133 | 185.7297 | 268.6135 | 197.409 | 205.8613 | 275.8683 | 348.6569 | 325.0658 | 191.0658 | 225.2703 | 341.0728 | 266.9342 | 251.3137 | 293.5476 | 257.2339 | 352.591 | 222.3067 | 216.8249 | 216.7227 | 313.5112 | 199.591 | 222 | 213.6638 | 221.1022 | 319.0364 | 313.2479 | 291.9636 | 221.7227 | 140.4818 | 162.3431 | 154.9636 | 154.9706 | 383.5252 | 408.4957 | 386.7072 | 483.416 | 373.6135 | 330 | 355.1456 | 406.6499 | 163.1387 | 86.3431 | 90.584 | 163.1387 | 124.8613 | 126.6499 | 122.0728 | 154.1022 | 138.409 | 145.3865 | 116.7885 | 142.4748 | 92.1022 | 87.6204 | 87.7885 | 118.591 | 118.416 | 134.6569 | 120.5182 | 88.584 | 339.9636 | 303.4818 | 303.8683 | 355.3067 | 214.1681 | 352.7382 | 322.2115 | 342.5546 | 449.6344 | 313.7227 | 339.8978 | 478.4818
1416 | 17.5704 | 160.1471 | 247.5797 | 311.1879 | 236 | 173.5266 | 217.0408 | 241.796 | 186.1347 | 213.5266 | 148.0123 | 123.555 | 219.5266 | 277.4326 | 173.5797 | 180.9469 | 253.1063 | 247.0161 | 170.7429 | 175.204 | 231.2287 | 198.6205 | 445.2979 | 433.5958 | 176.2818 | 303.6366 | 291.7837 | 315.2287 | 376.2979 | 142.1879 | 138.9184 | 162.3795 | 163.0655 | 169.5674 | 176.2287 | 184.9592 | 109.6897 | 321.89 | 360.2324 | 347.6366 | 344.8653 | 184.445 | 183.7713 | 165.2287 | 148.7552 | 99.6082 | 76.4326 | 79.9592 | 152.9716 | 334.4734 | 325.2856 | 379.5389 | 219.2448 | 181.4611 | 151.9592 | 185.4734 | 153.555 | 231.3387 | 211.7429 | 336.649 | 203.0692 | 182.8084 | 200.3918 | 265.2818 | 188.1347 | 198.5142 | 287.6897 | 355.7429 | 314.649 | 186.1347 | 220.796 | 345.1224 | 264.8121 | 258.4326 | 276.3387 | 239.9592 | 361.4734 | 214.9592 | 219.0123 | 218.0939 | 298.3918 | 214.7429 | 215.7305 | 223.0692 | 209.3795 | 324 | 320.5958 | 279.9469 | 241.5105 | 143.4858 | 166.6205 | 147.2571 | 153.8937 | 394.9877 | 441.6082 | 411.2324 | 491.7429 | 361.2695 | 348.0247 | 372.649 | 411.0247 | 168.1347 | 82.3918 | 87.351 | 168.1347 | 126.5266 | 116.5266 | 135.9184 | 153.351 | 134.6205 | 148.7182 | 127.335 | 128.6613 | 85.4734 | 90.9184 | 78.7429 | 127.4734 | 119.2979 | 143.7021 | 128.8776 | 85.7429 | 325.5958 | 304.9184 | 312.1632 | 346 | 208.0939 | 314.89 | 342.6205 | 350.3918 | 485 | 315.8776 | 336.7305 | 458.759
1417 | 17.5814 | 181.0207 | 260.7669 | 316.8967 | 226.705 | 176.1711 | 211.1711 | 242.4512 | 191 | 212.1297 | 159.5752 | 123.3157 | 217.3215 | 264.8967 | 180.705 | 180.9116 | 275.4041 | 209.938 | 178.2331 | 175.7612 | 253.295 | 205.1297 | 428.705 | 419.8083 | 194.3421 | 281.5545 | 299.062 | 323.1504 | 369.4041 | 150.7256 | 148.2331 | 144.5545 | 188.8083 | 179.6785 | 184.9587 | 183.5545 | 103.4041 | 334.8967 | 330.4248 | 340.1033 | 327.876 | 179.8347 | 169.7876 | 175.5752 | 166.5959 | 96.1711 | 74.4041 | 85.4248 | 155.8554 | 317.6843 | 315.7198 | 396.9793 | 213.5281 | 163.6165 | 170.357 | 180 | 143.6165 | 236.7612 | 210.7463 | 329.3628 | 199.5694 | 207.9587 | 198.9793 | 276.4719 | 196.6372 | 200.9116 | 281.4041 | 354.7463 | 326.357 | 201.1033 | 225.4926 | 347.3157 | 270.0149 | 240.2388 | 268.8289 | 242.5959 | 362.062 | 207.0884 | 227.3421 | 230.8496 | 306.2744 | 219 | 212.4661 | 200.9793 | 204.4868 | 325.6165 | 327.0149 | 287.186 | 273.9587 | 149.1711 | 164.643 | 155.5281 | 138.4661 | 377.7669 | 422.814 | 393.5488 | 491.9587 | 381.5074 | 379.2331 | 378.6992 | 443.8496 | 169.3628 | 82.5959 | 97.4661 | 169.3628 | 138.8703 | 123.2124 | 154.5281 | 158.2124 | 145.1711 | 140.7612 | 154.3215 | 144.6579 | 84.8496 | 97.8083 | 84.2124 | 119.1711 | 124.9322 | 148.1711 | 137.5959 | 90.4041 | 321.9116 | 298.4719 | 315.357 | 339.938 | 225.6992 | 363.0562 | 345.4926 | 351.4041 | 476.9174 | 331.062 | 333.062 | 461.5017
1418 | 17.5924 | 188.9971 | 247.6715 | 307.4135 | 219.2464 | 179.0792 | 202.4193 | 235.9912 | 199.7449 | 204.0821 | 157.2493 | 133.7536 | 206 | 255.4135 | 171.9971 | 194.5015 | 301.8183 | 211.827 | 189.6628 | 189.8387 | 253.6745 | 197.0821 | 421.2464 | 426.8299 | 189.7478 | 289.5748 | 306.7507 | 324.3285 | 383.3255 | 138.3343 | 141.7566 | 142.9971 | 182.9208 | 184.3372 | 180.6657 | 169.5044 | 102.3343 | 315.4193 | 328.1701 | 348.7536 | 314.1642 | 155.5836 | 163.0029 | 179.4956 | 167.2493 | 90.7507 | 77.0821 | 88.1672 | 139.5807 | 315.5748 | 360.9004 | 391.085 | 239.4956 | 166.4164 | 173.8445 | 190.4106 | 146.8328 | 240.4281 | 226.3197 | 319.4193 | 217.3372 | 207.4135 | 196 | 281.3197 | 205.3314 | 216.5836 | 286.5807 | 349.0821 | 342.7536 | 214.7507 | 213.3314 | 364.8328 | 286.8416 | 224.4956 | 276.3314 | 255.742 | 368.9179 | 201.4106 | 218.1672 | 234.2522 | 297.4281 | 211.088 | 213.2551 | 201.7478 | 214.5015 | 315.5073 | 342.1759 | 311.085 | 269.6657 | 137.088 | 149.0792 | 166.5044 | 142.1701 | 387.1584 | 423.4839 | 397.9266 | 503.0734 | 374.5131 | 390.6628 | 392.1672 | 443.9208 | 162.3343 | 89.912 | 96.5894 | 162.3343 | 151.915 | 127.0821 | 157.5923 | 157.9179 | 145.9971 | 155.2551 | 140.5015 | 150.9179 | 92.8328 | 91.9208 | 94.3285 | 119.1642 | 147.3314 | 131.5073 | 141.5807 | 85.5865 | 335.9179 | 292.9091 | 328.0058 | 346.4077 | 224.1759 | 370.8504 | 339.1613 | 344.9208 | 457.9208 | 336.2522 | 354.4927 | 505.8328
1419 | 17.6034 | 200.2871 | 255.7209 | 308.998 | 232.2871 | 198.8604 | 199.4297 | 238.994 | 202.997 | 224.5773 | 160.7149 | 129.2871 | 217.1466 | 261.2851 | 173.4267 | 188.8564 | 302.8454 | 231.7159 | 200.7149 | 180.8584 | 244.003 | 207.7169 | 417.1386 | 427.2831 | 191.1416 | 297.1386 | 311.4307 | 320.5663 | 396.8554 | 148.4327 | 128.1436 | 147.8564 | 179.4307 | 182.2881 | 179.999 | 161.5723 | 93.998 | 321.004 | 341.007 | 355.005 | 323.001 | 164.8604 | 162.002 | 187.4277 | 176.7169 | 90.2861 | 80 | 80.998 | 142.999 | 329.998 | 381.5653 | 389.5723 | 242.7119 | 175.5743 | 161.5763 | 189.5663 | 150.5723 | 219.002 | 236.7068 | 329.7199 | 208.4287 | 211.999 | 198.1436 | 305.8544 | 216.8594 | 221.1446 | 286.998 | 357.5733 | 320.2811 | 220.7169 | 224.4307 | 357.4257 | 279.4327 | 239.2871 | 271.5683 | 258.2801 | 367.2861 | 206.997 | 213.5703 | 227.2851 | 296.1516 | 206.8594 | 205.2871 | 205.7139 | 227.2911 | 304.8594 | 320.8574 | 290.7089 | 253.5653 | 130.2871 | 150.8544 | 160.2871 | 140.8594 | 388.1366 | 442.2791 | 387.2901 | 511.1376 | 354.1466 | 379.8504 | 381.1396 | 439.1446 | 155.2841 | 93.2831 | 88.8594 | 155.2841 | 152.5703 | 139.003 | 135.5713 | 149.4267 | 154.2861 | 145.1436 | 133.5703 | 140.2831 | 87.5693 | 88.8594 | 97.8544 | 120.7129 | 137.4237 | 124.2891 | 138.5683 | 75.8564 | 316.7089 | 312.7149 | 332.8634 | 368.8574 | 215.2901 | 339.8604 | 361.003 | 339.2871 | 466.4347 | 327.1416 | 362.997 | 485.5643
1420 | 17.6144 | 197.6209 | 277.1489 | 308.4971 | 225.6518 | 207.1489 | 216.7892 | 228.2302 | 189.677 | 241.8259 | 157.2669 | 142.0561 | 224.9439 | 268.2611 | 157.1799 | 177.1799 | 270.9942 | 235.795 | 199.031 | 180.2669 | 263.0561 | 203.4468 | 399.559 | 428.9071 | 193.9381 | 278.236 | 314.2669 | 314.5532 | 382.5029 | 158.3849 | 133.646 | 151.6151 | 176.9749 | 196.4101 | 177.764 | 171.6151 | 93.9381 | 336.559 | 366.354 | 370.1489 | 324.795 | 181.087 | 174.8511 | 191.9381 | 175.5339 | 88.913 | 77.354 | 80.9381 | 142.087 | 321.559 | 355.677 | 407.5532 | 227.2669 | 179.5029 | 184.087 | 185.7582 | 148.7079 | 241.1122 | 222.4352 | 332.5958 | 203.7079 | 202.7079 | 204.087 | 290.1489 | 219.8259 | 228.441 | 286.9381 | 356.1799 | 294.5029 | 230.559 | 231.677 | 355.7021 | 282.7698 | 241.472 | 257.677 | 235.118 | 357.9749 | 203.3791 | 215.1741 | 223.236 | 314.2418 | 210.7079 | 220.2611 | 208.8511 | 237.0929 | 304.2979 | 322.882 | 286.6402 | 228.118 | 132.913 | 149.2611 | 166.882 | 161.9071 | 369.6151 | 415.118 | 406.764 | 493.087 | 366.354 | 375.4913 | 375.4971 | 449.087 | 163.1122 | 85.646 | 93.5899 | 163.1122 | 149.323 | 147.031 | 133.677 | 147.2921 | 142.7698 | 140.5029 | 130.323 | 133.087 | 85.6151 | 98.441 | 82.1489 | 126.0251 | 124.1741 | 141.087 | 129.087 | 77.8511 | 303.8201 | 314.118 | 342.5648 | 368.236 | 243.1431 | 367.1122 | 381.3791 | 344.118 | 470.2727 | 332.1431 | 360.7021 | 446.8569
1421 | 17.6254 | 192.4396 | 260.5742 | 322.533 | 209.0934 | 188.761 | 231.1868 | 237.7473 | 197.9725 | 226.7473 | 154.9863 | 147.8269 | 218.2665 | 275.9203 | 167.7582 | 176.8791 | 294.0797 | 229.8269 | 198.9863 | 184.7857 | 260.6676 | 189.0797 | 408.5192 | 438.1071 | 204.9066 | 291.8654 | 298.8407 | 330.2802 | 379.7994 | 152.7198 | 139.7198 | 148.6676 | 169.7994 | 194.7473 | 175.467 | 175.0137 | 107.1731 | 339.6264 | 366.6264 | 376.6923 | 321.0934 | 175.0275 | 180.0934 | 189.761 | 161.9066 | 89.533 | 70.8269 | 92.8132 | 143.7335 | 320.0934 | 354.4533 | 417.3874 | 225.4396 | 175.4396 | 189.3599 | 205.1868 | 140.1868 | 249.9341 | 235.6538 | 302.5604 | 188.3874 | 206.4258 | 205.2802 | 285.8132 | 207.467 | 221.7473 | 290.6538 | 339.1071 | 313.5577 | 222.7473 | 223.2006 | 361.6538 | 270.533 | 253.3187 | 251.467 | 255.3049 | 362.5852 | 216.2665 | 221.1868 | 215.1071 | 300.2527 | 205.3599 | 231.0934 | 212.2802 | 219.7335 | 304.6923 | 307.6813 | 319.4258 | 232.8929 | 129 | 160.0934 | 159.8407 | 162.9478 | 370.7473 | 413.5467 | 418.0659 | 506.9725 | 380.2253 | 390.8544 | 400.4121 | 449.8269 | 180.5467 | 87.1868 | 85.2802 | 180.5467 | 143.7473 | 138.3736 | 129.7335 | 145.8407 | 136.4258 | 140.0659 | 140.1593 | 134.7335 | 103.5192 | 94.467 | 85.0659 | 118.6951 | 137.4396 | 138.2006 | 139.3462 | 83.0934 | 310.467 | 298.4945 | 322.0934 | 372.7994 | 242.4011 | 377.7473 | 375.6813 | 354.3324 | 452.1593 | 342.2802 | 382.9725 | 457.1181
1422 | 17.6364 | 183.17 | 268.9514 | 318.2227 | 215.915 | 174 | 228.0688 | 243.4332 | 218.3806 | 206.3603 | 157.7409 | 142.1376 | 209.8259 | 268.6721 | 187.9474 | 195.5871 | 291.5506 | 231.1215 | 211.0526 | 205.0526 | 222.7894 | 194 | 417.2065 | 423.085 | 202.7409 | 305.4129 | 303.8988 | 306.5141 | 378.6882 | 147.6194 | 146.2591 | 141.9838 | 180.3279 | 179.4818 | 187.2267 | 177.8462 | 106.3441 | 348.6559 | 371 | 388.7409 | 326.0526 | 166.7247 | 194.3644 | 186.1741 | 159.7409 | 94.0688 | 70.7247 | 88.0162 | 138.2065 | 310.1538 | 356.8624 | 407 | 234.7935 | 181.0688 | 190.0688 | 207.1903 | 136.6032 | 234.3441 | 228.4656 | 305.3118 | 202.6761 | 197.1174 | 217.3644 | 284.7409 | 203.8624 | 221.3806 | 283.9312 | 350.1579 | 348.9312 | 213.5344 | 220.5871 | 344.6882 | 279.7247 | 279.0526 | 252.9838 | 274.0162 | 388.7247 | 209.2227 | 227.8462 | 210.7935 | 291.7733 | 209.7935 | 216.0324 | 215.9838 | 215.1376 | 299.5141 | 304.2956 | 314.3077 | 234.4129 | 145.2956 | 157.6032 | 163.0365 | 151.1215 | 358.2753 | 413 | 433.1903 | 499.4453 | 399.1376 | 373.6032 | 408.7571 | 445.0688 | 174.8785 | 93.3806 | 74.085 | 174.8785 | 132.6721 | 134.9312 | 127.4656 | 137.3968 | 142.4818 | 151 | 152.7935 | 128.7409 | 110.3441 | 82.9474 | 92.2753 | 107.4332 | 142.6032 | 131.8624 | 144.2065 | 88.9838 | 311.9838 | 276.8097 | 307.9635 | 368.4291 | 226.9312 | 358.2915 | 340.6356 | 376.3806 | 446.17 | 352.9676 | 370.7894 | 469.6356
1423 | 17.6474 | 165.9558 | 271.0176 | 304.6106 | 225.9558 | 176.8673 | 230.823 | 245.6636 | 206.9733 | 192.5221 | 155.4336 | 128.5309 | 208.947 | 267.9115 | 187.6903 | 191.4867 | 278.6991 | 238.9115 | 199.3627 | 204.8318 | 209.3364 | 183.0088 | 408.2655 | 414.9558 | 202.3452 | 298.0442 | 299.9733 | 308.3276 | 355.5752 | 139.9115 | 155.2567 | 145.6991 | 170.5752 | 171.0442 | 193.1415 | 186.0885 | 100.5221 | 343.0088 | 351.407 | 385 | 343.9912 | 174.823 | 186.6724 | 179.407 | 159.823 | 95.3894 | 75 | 80.9558 | 130.6991 | 316.2479 | 344.6636 | 410.823 | 238.4779 | 169.9645 | 181.8318 | 197.1858 | 138.823 | 218.9645 | 230.4336 | 331.77 | 231 | 195.4161 | 215.8849 | 283.3894 | 200.7433 | 247.2479 | 294.5576 | 346.1503 | 338.5309 | 222.0797 | 233.2124 | 345.947 | 272.5309 | 254.46 | 281.0709 | 279.8585 | 378.1858 | 193.2212 | 224.6194 | 222.1239 | 277.2567 | 210.1327 | 220.0709 | 225.9115 | 204.8761 | 281.2567 | 312.9645 | 294.9558 | 223.6991 | 156.3539 | 151.2212 | 166.053 | 154.1327 | 352.5664 | 418.7345 | 438.4779 | 463.6194 | 381.7079 | 380.1239 | 406.5576 | 444 | 156.0973 | 104.4336 | 73.1239 | 156.0973 | 130.9558 | 136 | 128 | 146.6461 | 136.9115 | 151 | 148.3539 | 127.3894 | 98.7876 | 83.6461 | 95.1682 | 114.9203 | 132.3982 | 135.4336 | 138.1327 | 96.5221 | 307.0973 | 292.0709 | 290.1327 | 368.3364 | 218.9203 | 335.9645 | 336.8585 | 375.0088 | 445.2036 | 371.3894 | 357.3364 | 441.4867
1424 | 17.6584 | 172.882 | 251.9017 | 309.3525 | 217.6868 | 178.0393 | 231.0786 | 237.8427 | 199.1953 | 191.5098 | 147.1966 | 118.941 | 209.7458 | 255.2949 | 196.4705 | 169.7064 | 285.2739 | 227.2752 | 198.6068 | 185.7261 | 217.6475 | 183.2542 | 409.8427 | 420.4115 | 196.1966 | 307.2936 | 282.3919 | 329.1376 | 370.4298 | 142 | 161 | 147.3722 | 158.4705 | 176.3722 | 178.9803 | 184.4902 | 110.2936 | 337.3722 | 338.3329 | 375.6868 | 340.3146 | 174.5885 | 181.5478 | 172.7051 | 160.5688 | 97.0197 | 82.3525 | 76.6081 | 150.0183 | 338.4115 | 357.92 | 403.7261 | 233.6081 | 159.4902 | 176.4508 | 192.7641 | 143.4902 | 228.5871 | 243.7641 | 323.5112 | 217.2752 | 202.3736 | 195.4902 | 265.4129 | 199.4115 | 238.0407 | 300.0983 | 314.5885 | 327.9607 | 208.962 | 230.6868 | 348.7064 | 262.4508 | 233.2346 | 279.962 | 292.059 | 350.7261 | 207.6068 | 218.8231 | 235.9017 | 279.5688 | 205.0393 | 224.3539 | 210.8441 | 208.2542 | 286.0197 | 301.0197 | 311.1953 | 210.6671 | 148.4902 | 146 | 162.3132 | 171.5871 | 348.5492 | 411.7654 | 446.3525 | 470.0773 | 376.7641 | 395.3722 | 401.3146 | 467.5281 | 140.5295 | 96.6868 | 78.0786 | 140.5295 | 142.7837 | 149.2346 | 138.7837 | 161.3722 | 134.0983 | 148.5492 | 131.6671 | 125.5885 | 89.059 | 96.4115 | 104.4705 | 110.3919 | 119.5688 | 132.5885 | 129.6081 | 104.3329 | 306.2346 | 308.118 | 289.4508 | 367.3342 | 218.8034 | 326.4705 | 345.6278 | 359.5688 | 452.7261 | 350.4719 | 372.5098 | 437.8427
1425 | 17.6694 | 185.532 | 272.5763 | 318.5074 | 214.532 | 185.5468 | 218.4582 | 225.3991 | 197.9163 | 189.9951 | 156.0935 | 135.5664 | 198.5074 | 237.9902 | 201.0148 | 187.6403 | 282.4286 | 231.0886 | 196.8917 | 176.0098 | 226.5812 | 185.4483 | 410.9655 | 442.5861 | 179.9705 | 314.9852 | 301.064 | 323.5074 | 388.4434 | 153.0541 | 139.8966 | 150.9852 | 171.0541 | 170.4385 | 187.5369 | 177.9655 | 108.4385 | 342.9951 | 361.5714 | 368.0098 | 344.6009 | 161.4582 | 200.4779 | 176.9459 | 152.4779 | 82.936 | 87.9902 | 67.9852 | 162.4385 | 320.8917 | 368.404 | 379.4385 | 229.5074 | 155.4779 | 191.5615 | 198.468 | 140.4828 | 238.4434 | 247.9606 | 290.468 | 200.4877 | 188.5271 | 186.4532 | 253.5468 | 198.4729 | 215.0246 | 292.4877 | 321.5566 | 310.404 | 194.5468 | 225.0197 | 339.0049 | 253.9459 | 240.9705 | 246.9557 | 277.4434 | 365.6304 | 222.4729 | 220.4631 | 230.4483 | 284.5418 | 186.4188 | 197.9606 | 204.5566 | 213.4631 | 273.4434 | 294.4729 | 305.8966 | 199.4877 | 156.5369 | 150.5221 | 168.9852 | 160.8425 | 346.5025 | 391.468 | 444.9557 | 474.9163 | 376.4385 | 403.0049 | 396.0541 | 484.4631 | 143.5221 | 96.0443 | 83.5468 | 143.5221 | 138.4237 | 145.4139 | 143.468 | 163.9803 | 136.0345 | 155.5468 | 130.0345 | 125.0197 | 92.532 | 91.4532 | 96.9557 | 113.9902 | 124.0394 | 125.9902 | 129.0246 | 95.4139 | 322.5123 | 312.0492 | 304.5615 | 374.6206 | 220.9606 | 318.9557 | 333.9754 | 364.5418 | 452.0541 | 351.123 | 350.8966 | 463.0837
1426 | 17.6804 | 174.4984 | 291.6033 | 306.1017 | 221.5148 | 185.7345 | 222.354 | 218.8983 | 193.8688 | 188.4852 | 159.5574 | 137.6755 | 207.2065 | 247.3245 | 200.9115 | 196.9542 | 277.2655 | 246.941 | 180.6623 | 190.354 | 213.1444 | 181.6918 | 429.2228 | 454.8525 | 176.5738 | 293.9837 | 315.5443 | 313.6755 | 372.3672 | 154.7345 | 128.7803 | 154.177 | 174.2787 | 158 | 188.3082 | 189.5311 | 93.4262 | 325.5279 | 373.4262 | 372.5738 | 342.8657 | 148.882 | 197.5443 | 165.4852 | 150.5738 | 83.3835 | 77.2492 | 81.4721 | 146.3967 | 315.9868 | 357.236 | 361.8525 | 226.882 | 156.6623 | 201.941 | 182.2197 | 143.6918 | 234.7213 | 234.8525 | 293.7803 | 187.1902 | 191.4262 | 200.1638 | 253.2197 | 190.4262 | 240.0753 | 282.7935 | 334.5443 | 304.8983 | 195.2492 | 216.1312 | 340.5148 | 247.118 | 242.2065 | 238.5148 | 266 | 381.705 | 229.354 | 201.1607 | 207.646 | 302.2655 | 178.236 | 181.2492 | 213.941 | 195.1902 | 272.8098 | 268.9247 | 287.059 | 205.7508 | 154.7345 | 148.823 | 168.5738 | 131.059 | 334.1312 | 372.646 | 410.7772 | 469.3245 | 368.6328 | 417.3835 | 409.5738 | 461.0427 | 146.9705 | 97.2787 | 90.941 | 146.9705 | 127.118 | 130.059 | 138.0295 | 157.941 | 138.882 | 159.8525 | 137.5148 | 137.236 | 90.2492 | 84.5738 | 86.4557 | 115.6033 | 132.5148 | 132.236 | 132.9705 | 84.177 | 318.3082 | 298.8362 | 302.0722 | 388.705 | 202.705 | 297.1312 | 319.7345 | 360.1312 | 448.0722 | 355.9247 | 346.9868 | 452.2034
1427 | 17.6914 | 169.5949 | 292.3648 | 296.6893 | 218.8377 | 199.1358 | 225.0403 | 226.7296 | 201.2566 | 190.1082 | 148.6893 | 152.8251 | 214 | 251.1484 | 193.7837 | 181.6352 | 280.2025 | 242.8918 | 183.3648 | 193.0403 | 198.7031 | 206.9735 | 428.756 | 416.2693 | 179.527 | 294.9195 | 309.6214 | 316.7031 | 385.9195 | 157.0679 | 135.8918 | 135.756 | 161.2025 | 167.4868 | 178.3107 | 190.6617 | 86.7837 | 323.1761 | 376.7975 | 356.5535 | 343.0817 | 146.0541 | 185.8239 | 167.6352 | 149.8377 | 91.2566 | 82.1761 | 90.6755 | 150.3786 | 313.0265 | 351.8239 | 364.9056 | 219.3107 | 155.6755 | 189.4591 | 187.7572 | 146.8377 | 227.7698 | 225.2566 | 296.1484 | 191.7572 | 184.7837 | 206.7158 | 226.6076 | 184.8377 | 243.7158 | 286.0138 | 315.9724 | 310.0944 | 185.4189 | 206.1082 | 343.6352 | 245.7296 | 252.6893 | 253.7572 | 264.9459 | 365.6755 | 225.1887 | 208.9735 | 207.5949 | 299.4051 | 179.6755 | 181.9597 | 207.7837 | 183.9459 | 243.9988 | 262.122 | 280.5673 | 206.6214 | 130.1887 | 145.6352 | 161.5132 | 131.9459 | 327.2704 | 398.9471 | 397.0138 | 474.2025 | 370.8918 | 412.6076 | 424.122 | 453.3786 | 146.527 | 90 | 102.7031 | 146.527 | 126.7837 | 137.7975 | 146.3786 | 137.0265 | 130.2566 | 145.5132 | 148.5409 | 137.0944 | 76.7296 | 86.473 | 73.9321 | 123.2163 | 124.0403 | 134.7296 | 119.878 | 84.2025 | 317.7975 | 302.2981 | 292.7434 | 360.0265 | 195.1082 | 311.3521 | 312.0541 | 348 | 444.5409 | 350.7031 | 366.1623 | 431.7975
1428 | 17.7024 | 187.5507 | 265.1906 | 311.8653 | 222.472 | 215.2247 | 222.3933 | 220.9213 | 194.8427 | 181.2133 | 160.09 | 164.1347 | 192.966 | 247.1573 | 194.854 | 184 | 276.0787 | 242.0787 | 169.674 | 189.3147 | 221.7867 | 191.1006 | 425.3374 | 410.27 | 182.1573 | 302.91 | 298.146 | 327.382 | 382.0446 | 144.348 | 120.5167 | 125.8653 | 163.5507 | 158.7413 | 183.09 | 175.4607 | 90.5507 | 349.0227 | 350.1793 | 344.854 | 341.5053 | 152.3933 | 169.146 | 179.708 | 142.146 | 83.764 | 81.5953 | 77.9887 | 141.3593 | 297.0787 | 341.618 | 362.292 | 219.236 | 150 | 178.382 | 184.82 | 127.82 | 210.6853 | 236.1014 | 272.7413 | 197.4493 | 197.7194 | 191.9213 | 213.1573 | 175.528 | 255.3487 | 285.292 | 302.854 | 308.3933 | 196.4047 | 188.584 | 346 | 238.8427 | 256 | 262.146 | 252.674 | 363.236 | 212.0787 | 217.3707 | 212.6067 | 258.7973 | 173.4607 | 192.1573 | 199.146 | 187.3147 | 229.4947 | 258.438 | 256.2806 | 197.8427 | 132.18 | 134.5167 | 157.854 | 143.944 | 339.5507 | 390.932 | 400.0673 | 463.0673 | 374.3933 | 418.5734 | 422.5953 | 460 | 146.4607 | 80.8313 | 109.0673 | 146.4607 | 117.6067 | 145.6967 | 138.9773 | 134.562 | 126.5393 | 126.2133 | 134.2693 | 138.09 | 76.3147 | 86 | 78.0227 | 122.146 | 119.7753 | 126.2247 | 114.9327 | 78.4607 | 314.91 | 312.5953 | 305.09 | 343.5393 | 189.4493 | 310.73 | 300.5953 | 345.3033 | 434.584 | 359.764 | 343.112 | 439.1573
1429 | 17.7134 | 195.1033 | 246.2065 | 317.587 | 217.5191 | 214.7581 | 212.1061 | 231.0325 | 193 | 178.6195 | 164.7935 | 158.413 | 190.4455 | 265.5488 | 179.6931 | 197.7906 | 254.935 | 222.0382 | 175.4455 | 193.5516 | 228.2419 | 172.9293 | 440.4484 | 421.6224 | 179.587 | 284.9675 | 301.1711 | 316.0707 | 360.0354 | 130.5163 | 115.6195 | 135.4484 | 166.6902 | 162.2036 | 171.2447 | 179.413 | 104.7228 | 340.5899 | 330.7228 | 338.5191 | 331.7228 | 160.8614 | 169.9646 | 189.6549 | 136.3451 | 80.4484 | 83.1358 | 77.5163 | 130.7581 | 284.2094 | 346.8614 | 355.1033 | 209.8642 | 162.1358 | 165.4159 | 177.0679 | 122.826 | 199.8289 | 257.826 | 263.5163 | 197.0679 | 184.9704 | 206.4455 | 200.1061 | 156.2094 | 256.5899 | 272.0354 | 299.2772 | 317.413 | 197.6224 | 190.2036 | 312.3509 | 235.3451 | 245.5191 | 269.5841 | 254.0325 | 367.6549 | 196.4512 | 207.7935 | 196.9675 | 256.7199 | 182.9293 | 173.0382 | 194.4484 | 194.3098 | 245.3451 | 254.2744 | 242.5516 | 197.1033 | 128.348 | 139.5488 | 150.4159 | 162.7228 | 359.7906 | 368.239 | 394.4484 | 462.413 | 374.0354 | 424.8642 | 413.6549 | 493.0975 | 143.7935 | 85.1358 | 89.1061 | 143.7935 | 115.7581 | 144.1386 | 133.6195 | 149.2065 | 127.5516 | 131.8939 | 136.6166 | 140.0354 | 75.587 | 93.7228 | 80.6224 | 121.8614 | 115.8289 | 118.0354 | 115.587 | 83.5163 | 312.9646 | 270.0057 | 326.8939 | 368.8231 | 194.5841 | 292.2065 | 310.9618 | 335.8289 | 410.8289 | 339.348 | 341.9618 | 444.3098
1430 | 17.7244 | 205.5866 | 275.068 | 298.2104 | 205.0526 | 190.4959 | 225.3762 | 249.0227 | 188.4887 | 180.6165 | 161.8722 | 162 | 208.0752 | 271.6691 | 183.7371 | 216.8948 | 258.9927 | 204.4361 | 185.1804 | 201.3309 | 215.2856 | 169.7217 | 411.2403 | 412.8722 | 182.767 | 282.2031 | 297.4134 | 309.1278 | 346.9773 | 131.6165 | 129.4587 | 143.4587 | 160.6165 | 151.6238 | 169.534 | 181.3083 | 113.8196 | 329.6392 | 347.1505 | 342.9701 | 336.3083 | 156.1052 | 164.9773 | 179.7217 | 147.9701 | 77.1804 | 92.4361 | 79.7443 | 124.5413 | 296.6845 | 324.6238 | 354.8722 | 215.2257 | 166.3608 | 172.3536 | 171.8495 | 126.0526 | 197.9474 | 244.1351 | 287.1731 | 199.7443 | 151.1578 | 215.6165 | 200.9701 | 159.6618 | 242.8196 | 276.4587 | 295.2557 | 302.9547 | 187.7443 | 208.9474 | 294.5866 | 239.0752 | 248.8422 | 260.3908 | 255.6691 | 348.699 | 190.3309 | 202.6165 | 206.0453 | 286.1505 | 178.5938 | 165.0227 | 196.2557 | 194.1804 | 228.2104 | 242.3908 | 245.2557 | 183.9021 | 124.1505 | 138.9021 | 136.4887 | 159.4134 | 363.1052 | 367.1578 | 399.0752 | 432.165 | 350.8269 | 431.3536 | 418.3835 | 486.729 | 135.233 | 91.6165 | 74.1804 | 135.233 | 123.0752 | 149.4587 | 129.9773 | 138.5938 | 124.0526 | 142.3083 | 143.1578 | 129.7969 | 79.3309 | 85.9021 | 79.767 | 118.7969 | 110 | 121.8948 | 113.1278 | 85.7443 | 317 | 265.4288 | 318.699 | 364.7516 | 202.8722 | 284.9773 | 303.7516 | 325.4887 | 419.0979 | 344.1731 | 384.9401 | 418.2403
1431 | 17.7354 | 195.1376 | 292.5428 | 283.119 | 212.3717 | 176.3383 | 245 | 229.4945 | 186.1524 | 180.881 | 150.0521 | 168.3383 | 196.4424 | 249.8662 | 190.3569 | 220.6952 | 276.5762 | 218.4052 | 174.3569 | 186.2565 | 200.0855 | 168.4907 | 376.3234 | 397.5948 | 195.4907 | 274.3234 | 296.9145 | 301.3569 | 378.6059 | 131.881 | 130.8141 | 137.8997 | 160.3048 | 151.0148 | 182.881 | 176.5428 | 115.4238 | 328.2379 | 346.3569 | 340.8997 | 340.1859 | 152.881 | 175.2862 | 170.4238 | 149.9331 | 85.3717 | 82.2045 | 80.881 | 121.4572 | 259.1079 | 306.7286 | 337.29 | 202.2231 | 176.4052 | 182.119 | 160.5428 | 117.8141 | 218.2862 | 226.8476 | 254.751 | 184.171 | 159.2862 | 206.6617 | 204.6617 | 166.9666 | 255.9479 | 285.3048 | 298.7286 | 288.4238 | 191.7621 | 192.9852 | 318.1338 | 246.4572 | 239.5614 | 236.7807 | 253.4572 | 343.3717 | 206.9479 | 201.7286 | 220 | 281.8997 | 163.8141 | 170.2714 | 179.5614 | 192.5762 | 230.9814 | 225.6952 | 238.3569 | 178.1859 | 123.9331 | 149.8959 | 133.5762 | 146.2379 | 333.3755 | 377.0148 | 385.7138 | 427.8959 | 340.0334 | 409.4276 | 419.2714 | 467.3383 | 134.0334 | 87.2714 | 81.7955 | 134.0334 | 114.8997 | 157.1524 | 116.0855 | 140.5241 | 124.4572 | 132.3569 | 128.8141 | 121.5428 | 81.5428 | 79.6097 | 83.2714 | 108.2379 | 106.5428 | 125.6952 | 113.4238 | 82.8476 | 303.7472 | 262.409 | 303 | 356.9479 | 192.2045 | 277.4238 | 280.9666 | 312.7807 | 462.8438 | 346.8997 | 362.3607 | 402.9145
1432 | 17.7464 | 194.947 | 304.124 | 287.4735 | 217.646 | 170.9955 | 217.3406 | 208.823 | 182.8805 | 178.8805 | 160.2434 | 162.9955 | 177.3495 | 237.4115 | 199.301 | 210.1725 | 268.761 | 211.9336 | 165.646 | 172.177 | 216.1859 | 182.8275 | 387.0089 | 404.0664 | 193.938 | 306.7834 | 297.292 | 295 | 388.1681 | 128.115 | 125.2345 | 142.5355 | 160.823 | 157.938 | 167.9336 | 178.708 | 109.7035 | 319.8805 | 342.9425 | 329.646 | 344 | 152.0575 | 192.1195 | 170 | 148.1195 | 83.584 | 78.531 | 88.2965 | 113.407 | 234.4779 | 314.4735 | 329.1195 | 176.3495 | 171.699 | 177.646 | 174.4779 | 132.8319 | 222.1725 | 226.5885 | 248.8364 | 183.416 | 156.1105 | 201.4115 | 194.115 | 169.2965 | 252.8185 | 279.3495 | 290.584 | 288.5885 | 189.5265 | 168.407 | 305.8716 | 256.062 | 227.7655 | 227.646 | 258.354 | 346.292 | 203.23 | 212.416 | 200.5796 | 280.6505 | 168.8275 | 188.4204 | 180.7125 | 175.3451 | 216.6946 | 225.7655 | 253.7744 | 182.5885 | 118 | 140.1061 | 133.4115 | 131.9955 | 326.0045 | 370.4026 | 383.593 | 422.2256 | 341.823 | 375.1725 | 406.8185 | 466.7035 | 141.708 | 91.2965 | 94.885 | 141.708 | 113.6505 | 143.876 | 113.354 | 136.053 | 121.115 | 121.292 | 130.2965 | 140.1859 | 85.4735 | 81.823 | 89.6505 | 105.7655 | 93.407 | 103.9911 | 117.1195 | 82.5885 | 282.8185 | 245.646 | 294.761 | 349.1105 | 198.5355 | 290.5355 | 273.292 | 326.5974 | 429.3272 | 338.5885 | 339.8275 | 405.646
1433 | 17.7574 | 208.8063 | 300.7826 | 305.8182 | 217.2016 | 185.0276 | 216.6245 | 214.0079 | 184.8063 | 195.2253 | 155.5771 | 156 | 178.0079 | 228.589 | 192.5771 | 202.7984 | 246.1779 | 198.7984 | 164 | 180.8103 | 208.7708 | 185.3953 | 392.589 | 422.6087 | 181.7905 | 295.1463 | 307.8182 | 310.6205 | 367.7905 | 139.6205 | 137.2174 | 133.9763 | 149.1858 | 154.8024 | 158.4032 | 178.9961 | 102.3953 | 305.585 | 359.4189 | 328 | 353.6126 | 154.8063 | 190.1937 | 170.6008 | 158.8103 | 79.4032 | 87.6087 | 92.6008 | 118.0158 | 254.411 | 314.7945 | 336.2055 | 189.0237 | 151.3874 | 173.5968 | 186 | 135.1858 | 205.1858 | 218.589 | 251.5771 | 185.1937 | 157.6166 | 181.174 | 175.5811 | 168.1937 | 248.0039 | 278.2055 | 289.4071 | 299.2134 | 185 | 182.6284 | 290.8024 | 235.1661 | 215.7826 | 230.8063 | 247.9842 | 353.2134 | 167.9645 | 214.1937 | 197.8142 | 277.589 | 158.7787 | 192.9882 | 181.3874 | 172.0118 | 199.6008 | 213.1818 | 243.7669 | 185.4032 | 123.4071 | 140.6245 | 148.0197 | 137.6166 | 316.7787 | 354.7984 | 403.0158 | 437.6403 | 355.4189 | 381.6166 | 400.8024 | 463 | 128.1779 | 83.585 | 78.5732 | 128.1779 | 115.9961 | 122.585 | 125.2134 | 127 | 124.2095 | 121.0039 | 127.3913 | 135.174 | 79.7866 | 84.6047 | 85.3874 | 112.4071 | 108.2292 | 101.415 | 118.7984 | 83 | 277.4032 | 274.0395 | 290.2016 | 342.2055 | 184.5692 | 288.585 | 278.411 | 324.7787 | 401.6166 | 322.1779 | 362.2213 | 410.6087
1434 | 17.7684 | 193.6076 | 305.4878 | 300.2861 | 211.8692 | 184.2861 | 208.7684 | 216.1608 | 177.5776 | 180.4114 | 142.7738 | 154.7738 | 179.5477 | 256.1063 | 180.3869 | 217.327 | 258.9101 | 198 | 162.1608 | 178.03 | 209.327 | 179.3215 | 412.1363 | 395.7329 | 173.9346 | 288.2316 | 305.9646 | 293.4114 | 363.6131 | 148.1608 | 137.4169 | 136.1008 | 153.0354 | 151.0954 | 180.8447 | 183.1308 | 103.0654 | 310.2616 | 343.2506 | 344.5532 | 343.4468 | 182.5232 | 172.8992 | 174.0654 | 158.4823 | 76.0954 | 82.1907 | 81.9646 | 132.1308 | 253.2561 | 320.5831 | 330.4169 | 197.3215 | 164.0055 | 152.3814 | 177.4169 | 119.4169 | 204.7439 | 224.6485 | 245.5177 | 176.4823 | 151.2861 | 181.4878 | 177.6485 | 155.8038 | 254.9046 | 260.1553 | 282.5776 | 292.5122 | 192.357 | 200.6785 | 306.7139 | 211.2561 | 202.7084 | 228.4823 | 236.3215 | 349.5776 | 190.4633 | 204.8692 | 200.7084 | 267.7084 | 149.8392 | 163.0899 | 172.5477 | 185.97 | 220.8447 | 210.7439 | 229.4523 | 171.673 | 123.9346 | 140.1253 | 158.6131 | 142.3215 | 318.2616 | 354 | 388.9291 | 440.2207 | 318.4059 | 394.9046 | 365.2152 | 455.03 | 134.7793 | 82.1308 | 73.5831 | 134.7793 | 113.3869 | 114.3869 | 121.5776 | 111.673 | 116.7384 | 130.97 | 120.5477 | 111.5776 | 79.7439 | 72.8992 | 86.357 | 115.3869 | 115.643 | 111.4523 | 106.3515 | 89.1308 | 274.7084 | 282.3515 | 267.7029 | 356.0354 | 178.8093 | 296.9401 | 293.8093 | 323.8093 | 403.2561 | 318.97 | 363.1907 | 410.0954
1435 | 17.7794 | 182 | 293.9881 | 279.7433 | 204.2478 | 175 | 192.2478 | 231.885 | 187.2597 | 168.003 | 163.2626 | 160.2537 | 170.4955 | 251.9851 | 200.0119 | 207.6135 | 242.7314 | 191.1209 | 166.003 | 166.7463 | 210.8702 | 166.9941 | 392.9792 | 412.8969 | 180.1298 | 267.8553 | 313.3835 | 275.3746 | 352.1179 | 127.6135 | 127.6224 | 145 | 153.7463 | 173.0149 | 167.7344 | 175.1179 | 105 | 299.2389 | 339.5074 | 337.4896 | 333.6254 | 166.1061 | 163.3746 | 171.6224 | 158.7522 | 82.3806 | 87.8821 | 67.4955 | 122.8672 | 247.7493 | 310.3657 | 326.2507 | 198.7522 | 168.4955 | 156.2597 | 172 | 127.1328 | 207.7493 | 232.6254 | 263.3835 | 175.5015 | 168.2656 | 178.1179 | 165.6135 | 160.6313 | 241.115 | 220.109 | 273.4985 | 282.1239 | 176.3627 | 196.7463 | 298.4896 | 224.5104 | 214.3835 | 209.3657 | 225.8702 | 331.1179 | 187.8583 | 199.7493 | 212.3835 | 295.6432 | 146.6224 | 156.1298 | 172.2507 | 183.4955 | 208.9851 | 203.1179 | 216.6165 | 155.7463 | 128.2537 | 133.2507 | 144.6165 | 133.1209 | 309.115 | 367.7582 | 404.3925 | 389.6016 | 312.1387 | 388.6194 | 394.5312 | 454.3776 | 138.4955 | 82.2478 | 73.3717 | 138.4955 | 107.997 | 111.4985 | 121.8791 | 118.885 | 114.6283 | 127.2448 | 129.6313 | 127.5134 | 89.6283 | 70.2537 | 72.2389 | 104.9941 | 97.8672 | 107.3717 | 93.3717 | 84.2448 | 271.3746 | 253.7374 | 287.3954 | 337.3598 | 178.1209 | 246.9644 | 286.8672 | 330.6254 | 392.7463 | 321.4985 | 333.8613 | 408.2507
1436 | 17.7904 | 203.0429 | 287.1013 | 253.8701 | 215.3909 | 162.8844 | 200.2026 | 222.8701 | 193.8117 | 162.7104 | 142.8416 | 153.1597 | 160.261 | 252.9416 | 193.5078 | 190.261 | 233.5649 | 190.826 | 165.174 | 163 | 198.3481 | 159.087 | 384.1156 | 406.9429 | 181.174 | 265.2169 | 292.6675 | 296.0429 | 330.9714 | 138.9558 | 120.5364 | 146.2753 | 155.1013 | 173.3338 | 178.1442 | 173.739 | 102.4494 | 313.5065 | 343.174 | 314.2468 | 366.5208 | 170.5935 | 167.4636 | 155.6091 | 153.3481 | 64.4065 | 87.3481 | 71.9273 | 141.7818 | 241.261 | 277.4065 | 337.8403 | 190.1597 | 151.8844 | 160.8987 | 158.6091 | 129.261 | 200.6234 | 222.7974 | 254.1455 | 184.0143 | 175.7104 | 181.8403 | 189.7091 | 158.0727 | 243.1156 | 229.5065 | 293.6805 | 280.3623 | 168.4636 | 194.2753 | 284.174 | 222.2468 | 226.826 | 205.1013 | 219.7247 | 321.4494 | 198.4195 | 190.7104 | 212.7104 | 263.6247 | 155.4779 | 169.2896 | 165.9857 | 178.3623 | 204.8403 | 229.1584 | 215.6519 | 147.5364 | 112.8701 | 125.7104 | 137.2753 | 133.4636 | 321.3182 | 360.6961 | 371.6247 | 417.6507 | 320.8987 | 415.5208 | 401.1312 | 448.7104 | 144.8403 | 83.1883 | 76.3766 | 144.8403 | 122.8545 | 125.3039 | 105.5948 | 119.4351 | 116.7247 | 115.6234 | 118.1455 | 139.087 | 80.2468 | 80.3766 | 71.2026 | 102.1883 | 92.5506 | 112.9273 | 90 | 77.087 | 296.5065 | 257.5792 | 269.7403 | 336.6662 | 176.5506 | 240.3325 | 270.7104 | 317.6091 | 389 | 335.3039 | 339.1299 | 398.7974
1437 | 17.8014 | 215 | 279.6007 | 243.65 | 208.0511 | 172.2488 | 211.1996 | 191.8514 | 182.2507 | 158.65 | 134.3993 | 153.4995 | 156.35 | 267.1996 | 194.0492 | 194.7994 | 253.299 | 189.1003 | 157.1504 | 160.4002 | 204.3993 | 166.4494 | 385.8004 | 390.35 | 176.4002 | 258.7512 | 282.4995 | 314.4995 | 335.3491 | 142.2507 | 143.9981 | 136.6007 | 156.7001 | 156.5506 | 180.6509 | 160.1012 | 102.9499 | 310.4513 | 339.0501 | 341.4475 | 340.8032 | 176.8505 | 169.35 | 143.4503 | 147.0501 | 66.5989 | 90.7994 | 79.5998 | 149.2006 | 262.0483 | 267.8997 | 335.5506 | 192.4494 | 151.4995 | 168.3993 | 158.1495 | 125.35 | 197 | 213.7502 | 268.0483 | 189.9499 | 159.9508 | 184.7502 | 210 | 151.0501 | 243.5005 | 227.1012 | 280.6518 | 275.4503 | 166.4503 | 204.7493 | 292.3993 | 231.8988 | 219.2507 | 233.3482 | 228.7493 | 315.4503 | 191.9016 | 184.0501 | 201.5005 | 241.4995 | 172.3993 | 170.1003 | 165.8997 | 171.5005 | 207.1003 | 224.6017 | 198.5515 | 156.0492 | 107.1996 | 128.1495 | 130.8505 | 128.8505 | 315.1514 | 344.8505 | 371.598 | 430.7512 | 339.4485 | 422.9508 | 399.9991 | 453.0993 | 145.8004 | 72.0009 | 77.4002 | 145.8004 | 121.9508 | 111.9016 | 110.8988 | 125.0492 | 123.7994 | 116.5497 | 109.9499 | 125.0009 | 64.5506 | 73.6007 | 75.7001 | 91.6509 | 109.5989 | 111.5005 | 98.4494 | 76 | 284.3519 | 252.0511 | 279.1978 | 346.2999 | 176.0501 | 242.0511 | 265.35 | 301.5506 | 374.701 | 321.9016 | 310.3528 | 407.299
1438 | 17.8124 | 221.6224 | 258.7684 | 247.9735 | 203.9735 | 189.6092 | 207.3776 | 189.6092 | 163.7551 | 183.503 | 139.3378 | 166.9337 | 169.2449 | 233.5766 | 183.4439 | 193.702 | 250.4041 | 204.2184 | 156.649 | 160.3245 | 211.3245 | 176.298 | 369.7551 | 393.3112 | 182.9469 | 260.5959 | 252.8878 | 295.4837 | 304.9143 | 146.9337 | 146.0796 | 133.649 | 160.6357 | 150.6755 | 165.3908 | 167.5561 | 102.0133 | 333.45 | 336.6755 | 334.8347 | 318.9867 | 165.0531 | 170.9867 | 156.2316 | 154.6092 | 76.3245 | 92.351 | 74.3776 | 118.5102 | 270.3643 | 287.2184 | 293.9143 | 189.0531 | 160.9602 | 164.7286 | 161.3378 | 126.9867 | 189.0531 | 208.0265 | 270.4041 | 185.702 | 161.2847 | 171.7418 | 195.4306 | 152.649 | 239.3378 | 217.3378 | 251.7684 | 269.6888 | 171.2847 | 192.7816 | 309.9204 | 235.702 | 215.3245 | 241.702 | 219.4306 | 332.8673 | 185.298 | 181.6755 | 207.9337 | 258.2449 | 158.7949 | 162.702 | 171.9735 | 163.3643 | 203.6755 | 208.6888 | 191.6357 | 162 | 118.6092 | 140.6092 | 137.5959 | 130.9602 | 301.0265 | 345.6357 | 381.702 | 407.4306 | 354.3112 | 384.55 | 406.3378 | 439.4572 | 158.8939 | 71.6224 | 72.0265 | 158.8939 | 117.9867 | 117.8806 | 121.9867 | 113.7816 | 111.4439 | 110.3908 | 118.2847 | 115.351 | 71.9204 | 73.298 | 84.9337 | 104.2051 | 105.4174 | 99.3908 | 101.0133 | 72.6888 | 302.45 | 272.4765 | 290.0398 | 347.6622 | 190.8939 | 236.649 | 266.9867 | 315.5428 | 356.4041 | 304.702 | 306.2184 | 404.4041
1439 | 17.8234 | 206.113 | 263.1416 | 284.4013 | 214.0944 | 180.5093 | 201.9764 | 228.4013 | 163.0708 | 175.7639 | 131.5801 | 163.9056 | 169.2546 | 242.656 | 176.3491 | 187.6273 | 254.4435 | 186.6931 | 162.7217 | 170.4435 | 195.1366 | 182.3727 | 363.6745 | 399.7217 | 176.2074 | 244.4149 | 269.7268 | 315.0286 | 319.7268 | 138.5093 | 143.3727 | 138.3727 | 150.1838 | 162.1416 | 180.5615 | 158.4621 | 105.7217 | 317.6223 | 319.1366 | 336.8398 | 331.4671 | 146.8348 | 170.6509 | 154.5565 | 153.6037 | 77.6745 | 82.231 | 80.4435 | 109.0472 | 236.9714 | 275.0894 | 297.9342 | 170.1602 | 169.3963 | 161.3491 | 160.3255 | 125.9764 | 178.2546 | 219.4907 | 252.1838 | 177.6037 | 159.6037 | 168.6981 | 187.3255 | 147.9292 | 244.3963 | 220.3727 | 277.0522 | 274.7454 | 175 | 186.6981 | 328.1416 | 241.769 | 219.3727 | 232.2546 | 221.4435 | 336.2546 | 190.6981 | 193.8162 | 221.0944 | 274.4435 | 175.9814 | 157.3019 | 163.882 | 166.3963 | 197.6037 | 219.1416 | 204.118 | 152.5565 | 134.4671 | 138.9292 | 132.2074 | 128.6037 | 325.3069 | 356.0944 | 392.4907 | 427.656 | 326.3204 | 341.0186 | 382.3913 | 446.8634 | 146.7639 | 82.4199 | 78.769 | 146.7639 | 110.9056 | 124.9764 | 124.3491 | 123.2125 | 103.6745 | 105.3255 | 126.7217 | 129.5143 | 67.2074 | 67.231 | 81.231 | 109.2783 | 102.3727 | 107.1416 | 106.7454 | 80.4435 | 277.8533 | 259.344 | 280.2546 | 347.3255 | 186.8584 | 239.3491 | 271.3727 | 312.8584 | 382.7032 | 300.6509 | 341.3069 | 387.5329
1440 | 17.8344 | 190.8186 | 244.9611 | 288.6371 | 202.203 | 184.3024 | 197.5659 | 220.9611 | 192.0994 | 166 | 149.9785 | 151.0713 | 166.6868 | 223.7191 | 189.3629 | 187.3737 | 249.3844 | 187.7365 | 165.6868 | 164.6976 | 204.8575 | 188.8078 | 370.1814 | 389.6371 | 179.2419 | 261.4731 | 274.3239 | 314.8898 | 321.5766 | 131.3132 | 145 | 140 | 159.797 | 170.0605 | 170.082 | 167.1707 | 110.7473 | 316.4234 | 306.1922 | 343.3132 | 321.203 | 147.5551 | 166.5659 | 156.1814 | 136.5766 | 71.1317 | 75.2527 | 75.3844 | 116.4946 | 265.2701 | 250.5766 | 307.6263 | 184.9785 | 163.0713 | 163.3737 | 174.4234 | 137.3629 | 169.5054 | 203.3347 | 258.3629 | 192.8575 | 158.3737 | 172.0605 | 187 | 138.1317 | 235.3239 | 211.6976 | 261.406 | 282.121 | 170.879 | 181.1317 | 313.3952 | 241.1922 | 229.9287 | 229 | 215.6976 | 331.6263 | 175.5162 | 188.3239 | 192.0323 | 279.6868 | 191.7473 | 156 | 167.9287 | 165.5659 | 199.121 | 196.1533 | 216.5551 | 161.0497 | 127.6371 | 131.1922 | 139.3629 | 114.3239 | 300.2245 | 331.8401 | 369.4665 | 433.4449 | 292.7689 | 356.8467 | 390.918 | 417.9114 | 140.4341 | 88.0605 | 73.3844 | 140.4341 | 111.8078 | 117.1317 | 118.1317 | 129.9395 | 109.4946 | 129.7258 | 110.4557 | 116.3952 | 64.7473 | 82.918 | 76.3132 | 99.4449 | 99.1922 | 115.0605 | 114.8078 | 87.0605 | 277.2311 | 247.3737 | 274.2527 | 353.1814 | 182.3737 | 226.2635 | 251.7084 | 313.8683 | 382.8898 | 297.9395 | 324.4665 | 364.1425
1441 | 17.8454 | 192.1947 | 276.6462 | 275.3097 | 206.1859 | 185.5044 | 180.6194 | 205.8053 | 181.6282 | 165.3009 | 139.0265 | 140.7079 | 190.77 | 195.0176 | 181.7168 | 190.0974 | 231.7168 | 196.7965 | 157.6106 | 158.6018 | 211.6018 | 186.1062 | 378.5929 | 395.885 | 174.6106 | 299.1682 | 265.5044 | 295.4159 | 324.0885 | 133.0974 | 152.6903 | 144.8938 | 150.9203 | 158.4159 | 164.4956 | 186.1859 | 99.4159 | 307.6194 | 315.1859 | 310.8406 | 316.0974 | 144.7079 | 127.2476 | 159 | 131.3982 | 65.2035 | 69.1062 | 66.1062 | 118.3009 | 252.8406 | 268.4691 | 303.5044 | 189.4071 | 154.1062 | 166.0974 | 183.0974 | 138.1062 | 191.4691 | 194.3982 | 274.4868 | 181.4247 | 168.7876 | 196.77 | 212.1682 | 151.7788 | 231.3982 | 229.3718 | 266.5753 | 264.4247 | 172.4956 | 175.2035 | 304 | 236.2035 | 215.8229 | 227.6018 | 227.7788 | 347.7788 | 174.9912 | 171.8141 | 189.5841 | 268.8141 | 148.2565 | 155.3009 | 167.1062 | 173.7876 | 194.7079 | 213.7612 | 223.4956 | 139.0353 | 118.5044 | 130.3982 | 136.6106 | 120.885 | 280.9026 | 326.6903 | 364.3894 | 409.7256 | 310.5665 | 381.885 | 401.3982 | 417.7788 | 145.4956 | 78.5132 | 71.0974 | 145.4956 | 109.8053 | 114 | 131.7788 | 113.6194 | 109.2035 | 122.1238 | 108.9912 | 111.1947 | 70.1947 | 85.0088 | 82.9912 | 103.6903 | 90.0088 | 111.1062 | 117 | 68.4247 | 277.6106 | 270.3718 | 273.6991 | 344.8141 | 177.4071 | 252.1594 | 254.5841 | 320.4956 | 387.1859 | 288.6106 | 329.8762 | 386.0621
1442 | 17.8564 | 197.557 | 252.3155 | 230.7383 | 217.4027 | 178.3087 | 216.6845 | 206.1342 | 156.349 | 149.349 | 131.4228 | 142.9799 | 189.6175 | 227.4161 | 174.5772 | 181.7517 | 216.7517 | 185.9061 | 156.1342 | 148.0403 | 201.7517 | 190.4027 | 377.443 | 393.8859 | 188.0738 | 284.3893 | 259.7315 | 285.7315 | 321.5973 | 134.7114 | 151.7315 | 129.9262 | 157.5168 | 167.2282 | 162.443 | 164.6779 | 98.2685 | 320.9194 | 310.0403 | 336.1275 | 338.3423 | 129.9061 | 154.3959 | 141.9262 | 129.1544 | 72.5369 | 74.8255 | 66.1342 | 110.1745 | 224.0604 | 266.9061 | 301.2886 | 199.8054 | 140.6175 | 157.0403 | 184 | 124.6175 | 190.6175 | 210.651 | 270.4631 | 159.4832 | 169.443 | 174.9866 | 184.5839 | 154.0201 | 224.8859 | 241.8456 | 271.443 | 263.1141 | 178.2685 | 158.349 | 282.6577 | 240.6913 | 225.7852 | 229.1342 | 222.9061 | 351.443 | 192.9396 | 176.9597 | 192.1544 | 243.3691 | 163.1476 | 149.3087 | 162.8658 | 180.8456 | 199.1141 | 195.698 | 214.3289 | 146.9194 | 134.0738 | 131 | 147.2282 | 121.7315 | 303.4765 | 329.2886 | 355.9061 | 414.5168 | 319.1544 | 372.7718 | 382.0806 | 410.7718 | 160.5168 | 81.8255 | 65.5973 | 160.5168 | 108 | 119.6913 | 122.6376 | 114.8255 | 114.4027 | 113.2886 | 110.5772 | 117.9799 | 83.3825 | 87.6913 | 78.8859 | 109.8456 | 90.557 | 93.349 | 99.2148 | 87.7449 | 278.9799 | 265.0604 | 292.4966 | 332.1745 | 184.2483 | 239.6779 | 260 | 346.1879 | 372.0806 | 293.5369 | 358.6309 | 385.4832
1443 | 17.8674 | 214.9214 | 255.9873 | 245.8428 | 199.0127 | 163.6971 | 187.6832 | 205.5526 | 168.0925 | 146.6185 | 144.3029 | 147.8948 | 175.5919 | 223.4601 | 181.237 | 177.2763 | 224.5792 | 197.6451 | 176.5399 | 156.3029 | 204.5133 | 178.526 | 379.6185 | 395.3422 | 187.763 | 263.1445 | 278.9873 | 276.0393 | 335.6451 | 137.8948 | 136.9734 | 130.237 | 139.1179 | 162.8682 | 168.9607 | 161.237 | 94.2104 | 295.7098 | 314.6844 | 317.9861 | 307.9202 | 139.474 | 157.526 | 129.2104 | 138.8555 | 79.6185 | 74.3815 | 75.6844 | 130.1584 | 235.3688 | 272.1318 | 280.0127 | 178.2231 | 138.1711 | 165.3029 | 157.2231 | 140.9873 | 191.7896 | 213.3815 | 275.6844 | 169.1977 | 172.3422 | 159.8289 | 176.237 | 152.7237 | 235.0266 | 227.8023 | 262.0393 | 261.6578 | 172.763 | 167.9214 | 282.6844 | 271.948 | 233.7237 | 224.2104 | 223.0659 | 350 | 171.4994 | 199.8162 | 198.237 | 245.1318 | 153.8416 | 165.0925 | 167.0659 | 145.815 | 183.1838 | 208.3295 | 208.5526 | 161.5133 | 129.4208 | 134.6185 | 150.8289 | 150.3954 | 299.2497 | 329 | 345.9341 | 412.0393 | 335.3688 | 349.6312 | 383.4081 | 417.3029 | 131.9861 | 71.2497 | 66.6185 | 131.9861 | 127.5399 | 122 | 111.6578 | 128.8555 | 117 | 110.1052 | 119.4081 | 112.0393 | 72.0786 | 85.6578 | 71.6578 | 118.237 | 105.7503 | 101.474 | 97.7896 | 77.289 | 264.3549 | 264.7896 | 295.6578 | 326.8289 | 190.1711 | 210.9075 | 280.2636 | 350.2104 | 364 | 279.6312 | 308.3803 | 398.8162
1444 | 17.8784 | 203.3362 | 266.208 | 267.5679 | 190.264 | 164.1519 | 168.528 | 204.264 | 163.9601 | 157.5679 | 135.0162 | 146.056 | 195.5517 | 233.6638 | 186.944 | 172.5841 | 232.68 | 183.3923 | 166.3362 | 163.208 | 208.472 | 170.056 | 371.4321 | 395.528 | 181.32 | 245.7522 | 283.32 | 310.5355 | 340.528 | 128.6961 | 141.5679 | 129.32 | 155.7597 | 147.9601 | 148.4483 | 164 | 84.6401 | 270.4882 | 326.8319 | 327.8157 | 288.5841 | 133.2241 | 152 | 124.056 | 137.1121 | 82.472 | 68.5841 | 66.4882 | 113.2403 | 219.9203 | 289.9838 | 290.3998 | 189.3437 | 144.8879 | 177.3599 | 176.4396 | 134.2802 | 199.8879 | 206.1121 | 276.792 | 163.2241 | 154.1282 | 150.9041 | 181.208 | 150.792 | 200.9925 | 226.4159 | 278.8718 | 276.9278 | 168.528 | 176.944 | 275.6961 | 243.9925 | 251.6638 | 222.736 | 228.68 | 351.472 | 177.9278 | 201.1121 | 188.4882 | 229.1282 | 151.6239 | 158.7522 | 160.1681 | 171.0075 | 213.5355 | 197.3923 | 220.5118 | 177.9838 | 136.7759 | 119.8082 | 135.2802 | 141.3923 | 295.472 | 334.8879 | 344 | 416.3599 | 318.4483 | 351.8319 | 362.7123 | 412.4321 | 132.9838 | 67.472 | 70.944 | 132.9838 | 115.8642 | 124.208 | 128.3998 | 124.9041 | 111.8481 | 109 | 119.32 | 117.0959 | 60.1121 | 83.264 | 73.68 | 111.4321 | 100.6961 | 103.32 | 91.9041 | 79.3039 | 284.4957 | 262.5841 | 294 | 346.6077 | 182.9041 | 217.9838 | 248.9925 | 316.3524 | 377.9838 | 299.4957 | 311.7597 | 370.6724
1445 | 17.8894 | 220.1967 | 258.7689 | 248.5515 | 192.2448 | 163.7552 | 166.5034 | 192.0275 | 160.7483 | 166.238 | 136.7345 | 134.5241 | 186.0412 | 210.0687 | 171.5378 | 172.4966 | 250.4622 | 178.9931 | 158.5034 | 157.2655 | 203.762 | 172.7414 | 333.579 | 418.1967 | 184.4897 | 232.0206 | 251.3204 | 287.3342 | 325.7826 | 122.7552 | 149.4897 | 141.4691 | 156.7689 | 150.7345 | 159.4553 | 168.4897 | 86.4897 | 265.2517 | 301.5653 | 350.9656 | 307.2036 | 152.945 | 154.2448 | 158.1692 | 154.4553 | 88.9863 | 72.9863 | 79.9588 | 112.2311 | 257.6452 | 265.817 | 287.2723 | 173.8033 | 155.9794 | 150.817 | 162.3067 | 137.2311 | 198.2586 | 193.5241 | 251.3067 | 184.4416 | 148.4966 | 142.0137 | 176.762 | 161.2242 | 214.6864 | 225.7552 | 300.9656 | 263.5447 | 183.7139 | 180.2448 | 289.2105 | 244.9656 | 243.0344 | 229.7345 | 224.0137 | 310.8445 | 176.5172 | 181.7895 | 198.9656 | 213.7689 | 145.0206 | 161.4828 | 168.2242 | 190.9863 | 215.7758 | 195.238 | 207.0412 | 171.7758 | 135.0137 | 126.7208 | 141.9725 | 141.4828 | 291.5103 | 339.9931 | 356.7208 | 427.9794 | 282.3136 | 375.9519 | 367.4691 | 404.5103 | 152.2174 | 75.4828 | 69.7552 | 152.2174 | 128.4553 | 123.5034 | 123.0275 | 129.4828 | 112.2448 | 117.9794 | 127.7277 | 110.2723 | 69.2242 | 80.0069 | 66.0206 | 98.2723 | 105.2311 | 102 | 103.2174 | 82.2517 | 283.5241 | 275.9656 | 243.8651 | 362.2311 | 177.0069 | 235.7208 | 240.238 | 328.945 | 395.7208 | 298.5241 | 302.293 | 381.1967
1446 | 17.9004 | 214.3097 | 263.6057 | 271.4229 | 213.5355 | 181.2538 | 174.3663 | 194.0846 | 163.2817 | 152.7885 | 143.5634 | 128.7183 | 183.0423 | 225.8595 | 156.8731 | 190.4932 | 256.7606 | 186.8452 | 180.8172 | 161.0846 | 211.1269 | 156.5068 | 334.9297 | 421.4366 | 167.7462 | 235.8452 | 273.7046 | 281.8452 | 315.676 | 119.7183 | 146.4366 | 149.0423 | 157.8029 | 154.5211 | 154.5914 | 172.2817 | 82.676 | 284.0143 | 304.9297 | 301.2387 | 326.9297 | 145.7885 | 156.5211 | 150.2251 | 133.6194 | 75.7885 | 68.1548 | 85.2394 | 118.8029 | 248.6617 | 267.4086 | 289.324 | 181.972 | 143.7885 | 143.2817 | 156.2817 | 132.3943 | 209.9297 | 206.7326 | 253.648 | 169.4222 | 168.7749 | 149.1269 | 171.1971 | 152.8308 | 222.4789 | 232.6057 | 256.5628 | 266.8874 | 182.1548 | 188.6057 | 290.4366 | 246.1971 | 206.8165 | 229.7183 | 235.6903 | 322.8595 | 172.4789 | 176 | 210.8452 | 232.296 | 141.2394 | 178.4509 | 155.2674 | 185.3943 | 192.9857 | 217.5355 | 224.5778 | 145.9434 | 132.2394 | 131 | 156.648 | 146.2817 | 280.8731 | 299.9291 | 380.0143 | 418.8308 | 311.7892 | 333.5628 | 358.3097 | 380.1828 | 153.9577 | 77.2394 | 81.9297 | 153.9577 | 123.5914 | 123 | 105.3097 | 119.8308 | 117.5634 | 117.9577 | 108.9434 | 110.8029 | 79.0846 | 100.296 | 75.9297 | 104.8874 | 92.7885 | 98.9577 | 97.352 | 75.1548 | 299.0143 | 283.2817 | 258.9441 | 330.7742 | 168.3943 | 231.6337 | 264.0566 | 318.7462 | 379.4645 | 286.6337 | 322.6201 | 394.324
1447 | 17.9114 | 169.0382 | 248.2241 | 261.6784 | 217.6814 | 170.7699 | 163.0884 | 174.3598 | 154.7256 | 171.9588 | 117.1769 | 138.0472 | 169.3156 | 225.4985 | 160.9558 | 187.4985 | 246.1799 | 207.5487 | 160.1769 | 173.0472 | 233.3216 | 171.8673 | 368.3688 | 409.1799 | 184.4131 | 245.7286 | 274.7256 | 275.4985 | 330.2301 | 117.4543 | 146.5457 | 146.1357 | 159.7729 | 154.2271 | 153.3186 | 158.3156 | 88.7286 | 276.0884 | 321.3658 | 310.2774 | 300.0854 | 154.9116 | 172.4573 | 157.9116 | 126.5457 | 77.9558 | 74.5015 | 71.0884 | 119.2271 | 213.9498 | 263.2714 | 277.0884 | 204.003 | 140.5457 | 140.9085 | 150.0442 | 116.0884 | 203.1799 | 202.7256 | 263.9558 | 183.6402 | 147.9498 | 163.593 | 203.2332 | 152.8643 | 201.9055 | 231.1357 | 255.6844 | 232.9025 | 187.7286 | 180.9528 | 302.1387 | 231.8613 | 214.7759 | 231.3186 | 229.9528 | 350.3216 | 181.2744 | 178.3186 | 216.8643 | 239 | 154.1387 | 179.9085 | 161.593 | 184.5457 | 214.0502 | 206.997 | 224.2714 | 159.0945 | 119.6342 | 127.9085 | 125.994 | 124.5869 | 256.3598 | 317.9146 | 355.0854 | 416.5457 | 316.2714 | 320.3186 | 354.7729 | 392.3216 | 128.2684 | 87.0472 | 86 | 128.2684 | 113.0442 | 107.5427 | 111.0472 | 121.41 | 114.3628 | 115.4543 | 98.9085 | 118.9558 | 76.3628 | 84.5869 | 80 | 104.1357 | 99.593 | 107.2744 | 87.8171 | 89.2301 | 288.7699 | 247.6754 | 249.6784 | 367.9176 | 159.0442 | 229.7729 | 252.4513 | 343.1417 | 377.6372 | 290.1829 | 313.997 | 387.4985
1448 | 17.9224 | 216.6718 | 237.5039 | 230.0899 | 214.6445 | 151.0821 | 189.621 | 213.539 | 159.8516 | 175.0742 | 108.2148 | 138.6445 | 168.1406 | 221.4297 | 164.5703 | 185 | 243 | 208.2891 | 146.5039 | 172.0742 | 219.3711 | 168.5781 | 377.7852 | 362.8165 | 194.1406 | 231.5117 | 243.7344 | 280.8516 | 313.8008 | 112.2891 | 114.0235 | 132.4375 | 160 | 133.586 | 154 | 169.7031 | 89.4297 | 297.9101 | 310.8672 | 328.207 | 302.7773 | 151.1484 | 163.6524 | 154.9336 | 154.4804 | 90.9922 | 84.0664 | 76.4219 | 122.1406 | 218.5625 | 251.5781 | 270.6445 | 166.6016 | 152.7773 | 150.9922 | 158.207 | 131.6289 | 210.9922 | 186.6524 | 267.5703 | 190 | 158.0586 | 160.7188 | 180.0235 | 161.0664 | 214.8437 | 229.2148 | 274.1328 | 240.0586 | 186.8594 | 169.3633 | 283.2305 | 240.5625 | 205.0821 | 223.3633 | 237.9922 | 352.8594 | 189.4961 | 190.7773 | 184.2383 | 229.5781 | 150.1484 | 178.2148 | 164.2148 | 185 | 207.8672 | 212.207 | 207.8672 | 157.9336 | 114.4297 | 133.2812 | 134.0586 | 135.2734 | 312.0272 | 319.9336 | 359.3476 | 398.9414 | 290.4453 | 329.6367 | 350.2891 | 410.5625 | 132.7773 | 79.793 | 74.2227 | 132.7773 | 119.6367 | 112.4219 | 106.1484 | 138.7031 | 120.8516 | 115 | 111.3476 | 121.7852 | 64.0078 | 75.6445 | 77.6445 | 97.5039 | 102.2148 | 108.4297 | 91.4961 | 91.6445 | 274.5781 | 240.1406 | 276.1914 | 369.4375 | 177.414 | 219.0078 | 242.2891 | 324.5196 | 364.0821 | 296.7109 | 305.0742 | 389.7109
1449 | 17.9334 | 199.5073 | 220.0511 | 259.6825 | 195.6587 | 168.531 | 182.0511 | 186.9251 | 174.7591 | 170.0128 | 134.3157 | 136.4051 | 192.1259 | 224.1898 | 148.2536 | 185.7974 | 216.6843 | 176.697 | 151.3796 | 160.6332 | 219.5821 | 162.8102 | 375.6077 | 339.8358 | 183.0383 | 240.5566 | 255.9362 | 282.2026 | 326.3413 | 126.9489 | 117.7591 | 135.3796 | 149.6332 | 145.5438 | 150.0128 | 162.8358 | 92.9872 | 308.1898 | 307.7974 | 348.5438 | 314.7719 | 163.354 | 160 | 163.3668 | 150.8358 | 82.8358 | 76.4306 | 80.5949 | 111.0383 | 240.3413 | 261.7591 | 281.9617 | 172.5438 | 164.7719 | 138.8485 | 171.3668 | 122.646 | 204.4306 | 188.5821 | 252.0511 | 181.2281 | 159.8102 | 147.0383 | 194.1259 | 163.7974 | 211.2281 | 221.823 | 253.2792 | 261.7464 | 164.469 | 157.4306 | 306.5055 | 239.2153 | 209.7719 | 186.7098 | 241 | 340.0383 | 187.0128 | 183.6332 | 200.5183 | 237.3668 | 127.2664 | 162.0511 | 175.1642 | 181.0128 | 222.3413 | 229.354 | 187.2536 | 162.3796 | 115.5949 | 119.8485 | 144.5821 | 127.2409 | 303.4817 | 298.0638 | 347.8485 | 406.7591 | 308.7208 | 348.7464 | 354.5821 | 376.52 | 156.7336 | 78.5949 | 71.7974 | 156.7336 | 134.7591 | 111.8102 | 118.354 | 142.2026 | 125.3923 | 121.3796 | 126.9617 | 123.5949 | 78.5438 | 73.4051 | 84.177 | 105.5694 | 94.823 | 111.9872 | 81.0383 | 82.2281 | 281.5694 | 272.8978 | 307.3285 | 354.0383 | 183 | 243.9106 | 238.6077 | 328.9617 | 357.6077 | 287.6332 | 292.0383 | 403.7591
1450 | 17.9444 | 177.2345 | 243.5311 | 262.5221 | 184.5221 | 153.7566 | 181.2389 | 196.4337 | 150.4689 | 203.009 | 139.3805 | 130.3318 | 155.0839 | 235.5266 | 156.9558 | 183.5708 | 248.8674 | 186.0045 | 166.7655 | 166.0974 | 227.4779 | 184.6726 | 383.0974 | 349.1461 | 186.4779 | 215.6592 | 266.6682 | 269.8539 | 339.9071 | 128.5708 | 120.1903 | 138.6195 | 131.615 | 142.7124 | 135.2345 | 151.9026 | 94 | 285.5176 | 307.1903 | 323.8495 | 318.6195 | 160.5221 | 137.3274 | 153.8539 | 152.0487 | 75.9513 | 82.0974 | 77.7611 | 125.8142 | 214.23 | 266.6195 | 304.4337 | 181.8584 | 160.5221 | 170.6284 | 165.9026 | 136.0045 | 184.9955 | 187.5708 | 242.3318 | 174.9513 | 156.5708 | 165.0532 | 203.2389 | 168.8584 | 222.7655 | 218.3805 | 242.9513 | 287.0532 | 167.9071 | 182.5311 | 258.9379 | 227.4734 | 216.0487 | 217.6771 | 233.7124 | 346.7168 | 183.5708 | 185.0487 | 206.1903 | 254.5753 | 156.8187 | 171.7655 | 139.1326 | 171.9026 | 195.4203 | 197.3716 | 206.4824 | 164 | 126.5266 | 115.1903 | 141.1416 | 131.2876 | 302.6682 | 310.0045 | 352.0974 | 403.5221 | 306.0929 | 335.1858 | 354.3805 | 384.0045 | 141.7566 | 74.9513 | 65.5221 | 141.7566 | 131.5221 | 115.8584 | 119.5708 | 118.5176 | 118.7124 | 130.2876 | 127.5708 | 128.8584 | 74.9026 | 90.0045 | 78.7124 | 121.7655 | 105.9558 | 94.376 | 84.4779 | 71.9026 | 296.1461 | 239.7034 | 295.1858 | 401.2037 | 165.9955 | 246.9513 | 241.2389 | 331.1903 | 371.5753 | 270.4247 | 307.624 | 369.7521
1451 | 17.9554 | 183.8643 | 227.8052 | 279.0846 | 206.0168 | 152.2881 | 168.0255 | 187.0255 | 161.2626 | 175.6528 | 122.5594 | 135.5762 | 170.4829 | 214.9832 | 143.5594 | 180.5339 | 294.1692 | 205.6185 | 149.4493 | 168.822 | 235.5762 | 173.5594 | 361.9832 | 375.8388 | 189.6441 | 223.7965 | 247.4493 | 262.0678 | 309.1189 | 123.8899 | 133.1525 | 116.8052 | 128.822 | 141 | 152.5507 | 159.8643 | 92.3559 | 291.5084 | 283.1612 | 301.3815 | 313.2458 | 153.2458 | 150.9066 | 139.4916 | 131.6273 | 69.2458 | 79.0678 | 72.8899 | 137.3983 | 222.6185 | 262.8899 | 289.2713 | 167.3815 | 167.2203 | 145.2968 | 154.1357 | 141.6441 | 178.5339 | 214.1269 | 241 | 172.3559 | 179.8388 | 168.3559 | 213.8643 | 161.7797 | 207.9154 | 208.9577 | 246.9322 | 276.3815 | 188.9066 | 180.7797 | 280.5252 | 225 | 240.8388 | 227 | 230.3559 | 350.6441 | 173.1357 | 176.1357 | 206.822 | 268.6864 | 172.3983 | 167.6017 | 167.8133 | 165.8899 | 216.771 | 189 | 202.9577 | 164.822 | 133.9322 | 104.3136 | 120.2713 | 136.2881 | 292.4916 | 311.5339 | 363.8643 | 410.2203 | 310.5762 | 360.593 | 405.7878 | 396.2203 | 140.2881 | 73.178 | 70.5762 | 140.2881 | 132.4661 | 107.1357 | 123.1101 | 118.7542 | 125.2203 | 131.178 | 115.4916 | 119.3136 | 90.2626 | 76.7374 | 74.5339 | 106.0934 | 107.3559 | 99.0423 | 91.7542 | 83.1525 | 292.4238 | 255.4829 | 293.4661 | 362.8563 | 175.1525 | 222.9832 | 260.0846 | 292.3647 | 362.6696 | 308.1014 | 308.7119 | 397.1692
1452 | 17.9664 | 186 | 245.5266 | 270.4852 | 201.8225 | 153.8343 | 173.3432 | 181.497 | 179.1834 | 182.1834 | 135.6864 | 141.1716 | 186.8462 | 205.8284 | 165.0296 | 190.8462 | 284.4793 | 196.4852 | 150.0059 | 172.3373 | 201.9586 | 158.3195 | 336.142 | 367.6509 | 176.6509 | 268.716 | 275.5385 | 294.3728 | 296.9941 | 162.213 | 140.1716 | 132.858 | 145.6864 | 147.6746 | 149.4911 | 161.1657 | 88.6627 | 278.9823 | 298.0237 | 328.0355 | 322.8462 | 143.6568 | 146.6568 | 144.5089 | 130.3373 | 86.355 | 76.3314 | 76.1716 | 126.4852 | 246.0237 | 241.9763 | 284.1657 | 155.6568 | 158.1538 | 168.0355 | 144.4911 | 122.8107 | 198.858 | 204.1479 | 256.8521 | 192.0237 | 176.6568 | 157.9882 | 177.6213 | 189.2012 | 219.8521 | 259.5621 | 258.8462 | 280.5089 | 198.8402 | 165.6509 | 305.5207 | 237.5148 | 237.6568 | 222.8284 | 216.6509 | 324.3018 | 184.3491 | 179.0059 | 198.6568 | 236.7929 | 148.1361 | 165.1657 | 161.8166 | 156.6568 | 221.3314 | 194.0059 | 206.8402 | 154.1538 | 136.6686 | 110.3432 | 113.497 | 131.1598 | 300.0118 | 320.1775 | 369.3373 | 399.4852 | 306.9941 | 317.7752 | 369.4438 | 362.9586 | 133.4911 | 88.0177 | 69.497 | 133.4911 | 112.9763 | 113.3432 | 119.8284 | 120 | 117.8225 | 112.645 | 123.0118 | 111.9941 | 94.8343 | 75.503 | 79.8402 | 89.4852 | 122.8521 | 102.6686 | 99.6746 | 84.3314 | 273.4793 | 266.8402 | 272.3077 | 356.1716 | 180.503 | 260.5503 | 239.8047 | 333.2248 | 361.6686 | 306.9882 | 291.3136 | 383.3077
1453 | 17.9774 | 186.8466 | 208.5161 | 240.0619 | 200 | 158.2331 | 185.1593 | 186.9263 | 178.6136 | 175.6873 | 136.4602 | 122.528 | 181.3805 | 253.2568 | 167.4602 | 180.3008 | 262.3746 | 201.6195 | 163.6992 | 167.9203 | 226.3246 | 177.1653 | 339.6195 | 371.7729 | 178.2331 | 249.0619 | 278.6136 | 305.2331 | 328.1712 | 127.6695 | 137.6136 | 137.8466 | 148.1534 | 179.478 | 158.1593 | 156.7669 | 90.5398 | 270.0737 | 297.7669 | 323.8407 | 302.1415 | 130.1475 | 150.9263 | 147.6932 | 137.7729 | 84.9203 | 74.3068 | 74.4602 | 124 | 230.528 | 234.6136 | 269.6076 | 161.6195 | 150.0737 | 155.3746 | 139.6136 | 124.0797 | 216.5458 | 195.9203 | 243.9144 | 190.9203 | 169.9203 | 161.9263 | 197.9381 | 138.2771 | 197.6017 | 240.3686 | 272.8525 | 290.4661 | 165.289 | 174.8525 | 303.0737 | 225.6076 | 224.1475 | 239.7788 | 212.3068 | 355.4042 | 195.4661 | 163.9144 | 197 | 246.0856 | 147.2331 | 160.7669 | 164.9263 | 180.3983 | 226.0797 | 269.5017 | 220.6992 | 175.7051 | 121.761 | 121.3127 | 124.8525 | 126.6136 | 327.3983 | 315.2271 | 353.9144 | 380.0678 | 296.6873 | 329.1653 | 364.2331 | 383.0915 | 142.1593 | 73.2212 | 76.6195 | 142.1593 | 131.0119 | 124.3127 | 105.4542 | 127.6195 | 104.9941 | 122.5458 | 125.8466 | 111 | 94.1534 | 70.9203 | 73.3805 | 92.9263 | 123.4602 | 90.3008 | 103.5398 | 78.9203 | 291.1653 | 264.6136 | 268 | 356.1534 | 163.2212 | 227.5161 | 245.1593 | 334.5339 | 356.9203 | 303.3068 | 288 | 390.8525
1454 | 17.9884 | 201.6013 | 220.7547 | 220.3987 | 173.3741 | 173.6013 | 164.6686 | 183.7055 | 176.2822 | 202.3437 | 131.7055 | 128.4479 | 188.589 | 241.3864 | 167.8589 | 205.4848 | 249.5521 | 196.1288 | 171.1534 | 163.5644 | 228.5644 | 171.5521 | 349.589 | 380.7301 | 196.178 | 231.1165 | 266.8343 | 293.1165 | 317.6809 | 123.4356 | 124.9754 | 147.4479 | 137.6932 | 163.5275 | 157.4233 | 153.4233 | 84.1288 | 299.9204 | 309.0246 | 346.0492 | 299.7178 | 147.7547 | 156.2945 | 156.589 | 134.7055 | 75.411 | 82.589 | 58.5398 | 128.2945 | 224.4233 | 259.767 | 271.2945 | 169.8712 | 154.1534 | 146.8466 | 131.2699 | 119.8466 | 194.0919 | 190.7055 | 273.6383 | 204.6013 | 162.9877 | 171.589 | 196.1288 | 156.3437 | 202.4479 | 205.7974 | 259.5398 | 288.5644 | 194.215 | 181.2945 | 301.1411 | 218.7055 | 223.7178 | 204.3494 | 224.0246 | 346.5398 | 173.8097 | 191.0615 | 213.3191 | 261.0246 | 163.4602 | 170.3068 | 171.1534 | 180.7055 | 215.8343 | 210.8523 | 217.8466 | 172.2699 | 132.7424 | 117.8466 | 132.1534 | 119.9877 | 310.5275 | 320.0123 | 353.5767 | 413.9328 | 293.2822 | 339.0123 | 349.5398 | 359.6563 | 147.4356 | 83.7424 | 71.1288 | 147.4356 | 135 | 110.5398 | 116.7424 | 149.6136 | 124.4725 | 142.178 | 131.1534 | 116.1534 | 89.7055 | 76.8712 | 77.1534 | 114.6136 | 105.822 | 112.9081 | 98.8466 | 84.0123 | 291.5644 | 267.4356 | 281.7424 | 351.7055 | 173.7424 | 266.3807 | 241.8466 | 345.0246 | 361.1534 | 266.9261 | 288.8589 | 364.6563
1455 | 17.9994 | 186.5762 | 231.8407 | 244.1358 | 195.1358 | 169.9017 | 178.4238 | 184.7424 | 192.5527 | 173.8947 | 124.9017 | 125.644 | 190 | 274.5901 | 154.0609 | 189.9626 | 254.9695 | 179.3185 | 173.7424 | 148.1897 | 231.4848 | 175.2272 | 346.644 | 355.8642 | 190.2881 | 245.5527 | 268.4848 | 304.0679 | 321.9695 | 137.0679 | 143.9086 | 146.3864 | 156.0374 | 164.356 | 140.4473 | 146.0305 | 99.5527 | 284.0914 | 289.2202 | 322.993 | 298.2576 | 148.3864 | 159.6136 | 142.3185 | 146.1967 | 76.6136 | 63.0914 | 71.6815 | 133.356 | 264.946 | 246.5762 | 283.3255 | 183.1967 | 140.1897 | 147.7424 | 150.9086 | 122.4848 | 215.2646 | 186.5152 | 245.8947 | 189.5762 | 202.0749 | 169.5152 | 204.5831 | 175.8103 | 195.2881 | 222.7798 | 260.4848 | 296.7119 | 184.3185 | 179.3864 | 314.9391 | 235.4238 | 218.7728 | 210.1967 | 213.8033 | 315.2507 | 195.2646 | 181.1897 | 201.1897 | 266.4848 | 190.3934 | 158.9321 | 166.7728 | 185.2272 | 229.6815 | 219.9086 | 203.9321 | 176.2272 | 140.2272 | 123.9695 | 126.9017 | 120.7424 | 341.8477 | 311.4169 | 349.644 | 393.8642 | 317.3934 | 348.7119 | 375.7493 | 358.4848 | 141.9017 | 66.8338 | 82.1967 | 141.9017 | 151.5527 | 119.3255 | 112.9017 | 133.8338 | 133.2272 | 109.2811 | 108.4778 | 111.7728 | 106.4238 | 74.5152 | 88.4543 | 114.5152 | 107.356 | 110.0305 | 122.3934 | 75.4169 | 292.7424 | 287.1662 | 267.4473 | 368.4238 | 171.644 | 229.5692 | 254.0679 | 332.1897 | 360.2576 | 283.651 | 305.5527 | 383.5222
1456 | 18.0104 | 163.6799 | 212.6799 | 260.3688 | 190.1652 | 146.9129 | 182.767 | 197.3688 | 197.6505 | 173.4174 | 118.6991 | 121.4661 | 185.5826 | 248.1946 | 186.4559 | 176.3982 | 240.9808 | 188.4853 | 168.6991 | 162.7862 | 216.0973 | 150.379 | 368.0871 | 355.5339 | 177.5147 | 248.8835 | 291.0871 | 280.379 | 328.3009 | 150.4853 | 134.6312 | 158.3688 | 158.1165 | 156.1652 | 147.7183 | 182.9898 | 90.5147 | 302.2036 | 335.4751 | 304.8643 | 312.1357 | 159.4853 | 130.845 | 155.0192 | 125.9129 | 74.3495 | 62.6505 | 78.4174 | 126.0486 | 248.9129 | 248.4174 | 272.6312 | 160.2625 | 155.6697 | 153.3009 | 150.4661 | 139.7862 | 203.0973 | 193.9514 | 261.3201 | 202.0192 | 182.379 | 145.1459 | 190.9808 | 159.4468 | 184.2817 | 225.1165 | 283.0871 | 304.1844 | 205.8541 | 174.5826 | 279.8937 | 283.0577 | 231.2523 | 209.3495 | 224.3688 | 336.621 | 183.9808 | 172.8156 | 220.2036 | 224.5928 | 168.379 | 163.1844 | 177.4853 | 184.233 | 231.1165 | 216.8156 | 196.6991 | 175.233 | 149.8348 | 130.3009 | 129.5339 | 145.7375 | 313.4276 | 335.621 | 352.5339 | 403.2523 | 293.612 | 307.5928 | 372.0486 | 401.4072 | 155.1357 | 66.6505 | 69.8643 | 155.1357 | 150.4661 | 128.0679 | 134.9706 | 145.1357 | 134.8835 | 101.3495 | 113.8348 | 131.3201 | 89.5634 | 73.1165 | 79.3982 | 97.2138 | 113.3009 | 110.767 | 103.0294 | 80.1844 | 290.3495 | 254.6607 | 280.9027 | 352.4468 | 189.5532 | 253.9219 | 267.4853 | 360.9219 | 391.8054 | 282.5826 | 277.0781 | 405.5532
1457 | 18.0214 | 153.8338 | 205.5211 | 271.8535 | 201.5408 | 179.831 | 201.8113 | 183.7718 | 166.6479 | 208.0394 | 144.8732 | 141.6028 | 205.6028 | 273.5606 | 166.8141 | 174.1042 | 260.4986 | 205.2282 | 191.2901 | 187.3944 | 231.9155 | 181.9352 | 340.5437 | 361.3746 | 172.4169 | 268.707 | 287.7296 | 259.9803 | 345.1239 | 128.7099 | 119.5634 | 151.0423 | 155.3127 | 153.2085 | 151.6873 | 175.4592 | 90.7915 | 272.7521 | 306.169 | 327.1859 | 303.2507 | 144.8761 | 148.4986 | 161.4789 | 136.4366 | 79.3746 | 60.3127 | 74.5211 | 134.8535 | 242.4169 | 239.1465 | 301.4563 | 176.707 | 160.6873 | 136.0845 | 141.938 | 140.2085 | 201 | 207.5408 | 247.8761 | 191.4592 | 201.3944 | 165.2901 | 186.3127 | 167.7493 | 190.1662 | 259.0394 | 285.1042 | 292.4592 | 182.1268 | 197.2901 | 286.6451 | 213.7549 | 245.5408 | 234.9775 | 223.3127 | 383.893 | 188.2704 | 183.6451 | 170.1493 | 259.3099 | 190.0817 | 171.1662 | 185.2704 | 218.0394 | 254.2901 | 222.2704 | 232.7268 | 171.4169 | 129.5014 | 141.7493 | 125.5211 | 125.7099 | 317.9577 | 313.0225 | 351.2085 | 401.4169 | 273.8761 | 343.2056 | 353.9803 | 351.462 | 144.4592 | 66.1042 | 92.1859 | 144.4592 | 144.6254 | 122.7296 | 132.6254 | 129.0845 | 118.8761 | 129.6648 | 125.7493 | 136.6873 | 93.2704 | 85.5408 | 65.4592 | 118.2901 | 109.5211 | 107.4169 | 103.5831 | 75.6254 | 332.9972 | 248.2085 | 269.5634 | 374.1859 | 206.3324 | 206.9409 | 246.6056 | 320.2113 | 397.7915 | 271.2507 | 247.0225 | 391.8761
1458 | 18.0324 | 178.4258 | 208.6323 | 235.7693 | 187.5629 | 151.3096 | 195.8274 | 201.0694 | 180.2532 | 211.0919 | 145.2758 | 131.2871 | 238.8742 | 217.0677 | 169.4484 | 191.2532 | 258.4597 | 214.2645 | 167.6661 | 148.229 | 221.2871 | 201.4371 | 353.3452 | 380.1613 | 192.8855 | 269.1839 | 264.2984 | 301.5871 | 317.9419 | 129.6323 | 124.3565 | 136.379 | 135.9306 | 166.621 | 174.7016 | 174.9081 | 99.1726 | 274.4484 | 307.4484 | 321.8274 | 283.8387 | 135.7355 | 164.621 | 140.2064 | 147.0807 | 84.5403 | 89.9662 | 88.529 | 127.8274 | 226.5629 | 275.2307 | 255.0564 | 205.3339 | 151.9193 | 163.0581 | 158.2532 | 140.9081 | 205.5403 | 179.0338 | 263.2532 | 169.1145 | 184.9306 | 178.8968 | 186 | 176.2645 | 196.4484 | 221.229 | 278.6435 | 293.7242 | 186.2645 | 190.0113 | 300.7129 | 246.771 | 237.9193 | 231.6435 | 220.2758 | 352.6774 | 176.2871 | 199.529 | 214.8517 | 262.1839 | 193 | 166.5516 | 166.9306 | 216.5516 | 242.471 | 290.1969 | 234.2758 | 179.1726 | 135.1726 | 141.1839 | 120.4597 | 160.2307 | 319.9081 | 333.6097 | 367.3452 | 421.8855 | 298.3339 | 314.4016 | 359.2645 | 369.5178 | 144.8161 | 83.2532 | 87.7355 | 144.8161 | 147.6323 | 122.9081 | 149.2532 | 155.15 | 125.1726 | 142.9887 | 106.1145 | 140.6323 | 73.1145 | 67.0226 | 74.8968 | 92.85 | 102.6435 | 115.1726 | 100.3677 | 84.0807 | 265.3548 | 260.7129 | 271.6323 | 385.1726 | 180.758 | 252.7597 | 274.8742 | 351.3226 | 384.379 | 259.1032 | 283.9549 | 422.6904
1459 | 18.0434 | 193.885 | 229.2478 | 240.2832 | 206.2478 | 168.2478 | 181.1947 | 188.2743 | 179.2389 | 213.7611 | 145 | 113.4336 | 237.3982 | 234.0089 | 192.0886 | 206.8053 | 261.6814 | 214.0796 | 183.4071 | 160.5664 | 237.4868 | 172.6282 | 364.2036 | 374.6371 | 189.4779 | 260.7168 | 264.7611 | 300.4779 | 362.8585 | 118.0354 | 123.1593 | 119.3539 | 161.6107 | 165.2389 | 152.1504 | 154.7522 | 89.8761 | 284.2036 | 327.3275 | 355.9736 | 305.0089 | 170.8939 | 168.7611 | 164.6903 | 164.5664 | 78.5575 | 92.0796 | 66.9911 | 119.6371 | 234.2036 | 292.8053 | 275.77 | 184.9911 | 151 | 132.8672 | 144.3539 | 124.4336 | 240.9736 | 198.0886 | 270.5221 | 204.7346 | 193.1239 | 162.5132 | 194.2832 | 184.3629 | 181.3539 | 235.4071 | 277.0796 | 298.6018 | 192.5221 | 183.4779 | 298.3186 | 273.0886 | 228.7168 | 227.3186 | 259.5753 | 386.7346 | 204.4514 | 191.7964 | 198.8318 | 236.23 | 156.1857 | 194.5311 | 165 | 213.2389 | 263.0886 | 266.6282 | 243.2036 | 192.885 | 133.2389 | 152.0443 | 129.2036 | 147.4336 | 309.8761 | 315.7522 | 355.1947 | 435.0443 | 309.2832 | 305.4779 | 336.9911 | 366.4779 | 168.0089 | 63.8318 | 69.5132 | 168.0089 | 127.7522 | 143.2478 | 132.5929 | 164.4425 | 104.8318 | 120.0707 | 138.0532 | 130.8761 | 82.9646 | 79.7257 | 80.6018 | 129.5753 | 108.4425 | 111.3982 | 101.8407 | 83.1593 | 302.1771 | 296.0532 | 285.8053 | 380.4779 | 163.2743 | 266.2832 | 252.23 | 355.9204 | 385.7611 | 264.4425 | 278.7964 | 433.3629
1460 | 18.0544 | 187.5388 | 199.29 | 240.0674 | 191.2123 | 170 | 170.6735 | 170.2123 | 176.2021 | 198.145 | 150.5959 | 121.3265 | 259.3835 | 228.5388 | 205.1918 | 166.0309 | 258.2694 | 209.3368 | 185.9326 | 148.9429 | 230.6062 | 197.9794 | 370.5959 | 346.9532 | 183.4041 | 278.653 | 292.9794 | 325.1815 | 386.5856 | 121.6632 | 145.3835 | 119.8653 | 168.6632 | 159.4041 | 146.2694 | 158.5959 | 98.3265 | 299.9224 | 281.435 | 333.8185 | 295.8082 | 149.7512 | 148.4818 | 167 | 138.0206 | 78.9326 | 60.29 | 66.8653 | 112.4715 | 244.3265 | 258.5594 | 292.9224 | 188.5959 | 156.5959 | 153.3162 | 154.1918 | 131.3938 | 215.0879 | 196.2694 | 250.4818 | 210.7979 | 232.2386 | 155.4041 | 213.653 | 203.653 | 185.5959 | 222.0776 | 274.2021 | 309.2591 | 205.1244 | 177.4041 | 288.6735 | 234.8961 | 228.9326 | 245.653 | 233.1553 | 361.0879 | 201.4041 | 184.4715 | 203.5285 | 271.3059 | 185.6427 | 180.2123 | 163.1347 | 211.1347 | 264.0674 | 249.0776 | 208.5594 | 190.2694 | 138.5959 | 150.2021 | 143.9897 | 137.6062 | 296.8756 | 296.2797 | 327.8859 | 426.6735 | 304.4041 | 298.4715 | 331.2694 | 352.9429 | 171.8653 | 74.1244 | 75.4612 | 171.8653 | 130.6632 | 130.0776 | 140.3265 | 150.0776 | 142.1712 | 138.5182 | 120.4818 | 124.4041 | 96.1244 | 90.3265 | 91.2591 | 117.145 | 112.7306 | 111.9326 | 110.3938 | 71.8082 | 309.7306 | 289.6735 | 284.2021 | 355.7512 | 194.6427 | 242.7512 | 222.0206 | 360.6632 | 392.5285 | 272.4612 | 278.9326 | 410.6838
1461 | 18.0654 | 180.3854 | 242.3571 | 239.0551 | 191.8899 | 186.064 | 187.0089 | 207.7425 | 190.1741 | 232.9077 | 152.8899 | 133.3393 | 271.3943 | 247.8437 | 168.2025 | 174.3393 | 258.9449 | 236.4032 | 180.3304 | 196.1919 | 281.9716 | 193.3854 | 353.0462 | 361.064 | 196.2291 | 262.0462 | 242.0834 | 275.0284 | 361.5417 | 165.4672 | 139.4405 | 137.0089 | 139.7069 | 170.3393 | 136.5506 | 138.2113 | 89.5506 | 266.9822 | 279.8899 | 322.5506 | 335.6324 | 171.6235 | 147.9449 | 168.8899 | 161.5133 | 78.0551 | 75.0089 | 54.7158 | 129.9538 | 264.8437 | 244.6607 | 309.119 | 206.9538 | 172.119 | 165.3943 | 155 | 126.3304 | 212.0551 | 172.3765 | 259.3943 | 181.7069 | 205.7069 | 148.3854 | 234.8437 | 185.1563 | 188.8348 | 239.8988 | 280.6146 | 290.1563 | 206.9449 | 183.6146 | 271.936 | 270.7425 | 234.6696 | 218.6518 | 246.119 | 368.4494 | 209.5044 | 180.2202 | 170.9271 | 290.064 | 198.3943 | 162.936 | 154.4956 | 220.4494 | 268.7247 | 259.3393 | 240.9627 | 202.2842 | 150.3393 | 116.9271 | 128.936 | 128.4956 | 314.8988 | 326.183 | 327.8899 | 416.5506 | 268.0923 | 311.2291 | 325.3304 | 333.1012 | 147.4316 | 89.1741 | 71.2753 | 147.4316 | 158.4032 | 133.7247 | 135.3304 | 131.0462 | 117.5968 | 139.0551 | 125.6146 | 138.1741 | 94.1652 | 70.2113 | 99.5595 | 121.6696 | 105.4405 | 97.8259 | 101.5506 | 75.7247 | 305.2753 | 284.2753 | 257.5417 | 405.9716 | 191.3304 | 248.5595 | 259.6874 | 354.3854 | 377.881 | 286.2291 | 296.9538 | 414.6696
1462 | 18.0764 | 184.7861 | 220.112 | 241.8717 | 204.444 | 163.0693 | 200.444 | 168.8391 | 177.5988 | 238.8289 | 154.9145 | 123.4705 | 261.4705 | 261.444 | 151.6416 | 165.4277 | 268.5723 | 242.7861 | 195.3157 | 184.6416 | 235.2241 | 175.7699 | 371.1446 | 357.2139 | 192.2139 | 298.332 | 312.7067 | 326.5621 | 354.2566 | 135.4542 | 161.0163 | 137.0428 | 164.8024 | 187.2729 | 165.6741 | 153.2729 | 109.1018 | 299.4603 | 312.5458 | 336.3584 | 273.8656 | 159.5988 | 160.444 | 147.9409 | 161.0855 | 74.1711 | 64.5132 | 68.3584 | 135.7861 | 249.7271 | 235.3849 | 271.7108 | 194.5988 | 146.1976 | 148.7699 | 174.1446 | 141.3157 | 202.4277 | 181.5295 | 258.0855 | 198.1874 | 216.444 | 161.4012 | 211.112 | 218.4603 | 198.5723 | 256.3157 | 288.6578 | 314.8452 | 209.8717 | 183.0428 | 261.4277 | 278.7434 | 244.5723 | 220.8289 | 224.0265 | 373.7861 | 192.7699 | 202.0163 | 193.888 | 250.7963 | 197.0855 | 188.8024 | 176.9735 | 213.3422 | 296.7597 | 267.6578 | 199.9247 | 198.2139 | 128.0265 | 130.3157 | 135.6578 | 132.7861 | 338.0163 | 308.8554 | 325.1283 | 428.444 | 278.444 | 304.3422 | 353.7169 | 370.2892 | 157.4868 | 81.3849 | 77.7006 | 157.4868 | 148.5132 | 141.6578 | 131.1711 | 151.0591 | 139.9307 | 139.9572 | 113.556 | 140.9145 | 89.2139 | 93.888 | 91.3849 | 126.7861 | 120.3157 | 111.3584 | 96.2139 | 97.0591 | 320.3157 | 269.6416 | 265.5723 | 403.2566 | 178.556 | 259.5295 | 250.5132 | 396.1181 | 399.0163 | 262.112 | 269.2831 | 431.2729
1463 | 18.0874 | 179.1829 | 214.1524 | 270.1161 | 218.5733 | 177.5123 | 190.3353 | 185.4209 | 186.6952 | 276.8113 | 148.2134 | 114.2743 | 244.5181 | 263.939 | 164.5733 | 178.5733 | 275.7866 | 239.1219 | 186.3048 | 182.061 | 282.4456 | 196.3295 | 389.4514 | 334.701 | 192.9695 | 249.5849 | 263.6459 | 311.5486 | 353.0305 | 137.8781 | 158.1219 | 145.7257 | 170.8476 | 165.701 | 170.8781 | 160.7866 | 122.6038 | 304.8781 | 313.0305 | 321.4877 | 269.061 | 166.7562 | 140.6401 | 151.8476 | 150.3353 | 89.5123 | 76.6038 | 64.1524 | 112.7315 | 302.3236 | 294.1408 | 268.061 | 178.4877 | 145 | 150.9086 | 144.9449 | 166.238 | 239.8113 | 212.0551 | 267.6952 | 204.8171 | 248.0247 | 178.4819 | 260.4151 | 234.5428 | 171.8534 | 261.8476 | 267.6705 | 309.2134 | 213.8781 | 184.939 | 289.1161 | 262.5181 | 245.9695 | 221.9695 | 212.3353 | 397.2685 | 217.2075 | 198.1524 | 199.8476 | 284.8723 | 198.939 | 215.2075 | 193.5123 | 237.238 | 283.4572 | 271.8781 | 266.836 | 179.5791 | 141.5428 | 136.8171 | 121.4572 | 142.6952 | 326.3962 | 346.7808 | 326.939 | 376.6459 | 323.598 | 315.6343 | 294.8897 | 338.0668 | 158.9695 | 79.061 | 88.6647 | 158.9695 | 139.2743 | 163.3295 | 149.4209 | 191.7504 | 126.4572 | 124.4877 | 125.6038 | 161.3599 | 96.7562 | 92.0914 | 85.1829 | 103.7315 | 128.7562 | 112 | 110.5428 | 87.3353 | 325.8476 | 305.8418 | 288.299 | 347.7373 | 211.9332 | 241.5791 | 267.4514 | 369.8839 | 384.4877 | 271.6647 | 287.3904 | 429.0914
1464 | 18.0984 | 158.382 | 241.4907 | 262.1637 | 202.3092 | 172.1091 | 199.8181 | 186 | 209.5817 | 257.382 | 153.8909 | 113.0182 | 246.9454 | 275.7817 | 178.7454 | 147.582 | 274.0364 | 236.0546 | 211.5271 | 184.9454 | 315.418 | 197 | 368.4001 | 372.2907 | 188.0909 | 289.2361 | 290.4725 | 284.4911 | 370.6726 | 131.1273 | 138.3638 | 138.1455 | 171 | 204.2725 | 151.3638 | 152.1637 | 90.6002 | 296.1637 | 301.2183 | 319.0364 | 295.5089 | 172.8909 | 160.618 | 177.5271 | 151.9636 | 84.1091 | 85.8363 | 68.9091 | 108.0728 | 265.7093 | 228.255 | 242.4729 | 216.2907 | 165.618 | 146.0909 | 125.3456 | 144.4183 | 223.3274 | 187.4729 | 246.4001 | 245.2543 | 231.3274 | 175.0728 | 269.8545 | 227.1455 | 206.3452 | 262.9818 | 274.8545 | 337.4725 | 224.7999 | 193.8363 | 294.9091 | 282.618 | 275.4544 | 247.5271 | 235.5635 | 360.6911 | 197.382 | 195.0546 | 196.0728 | 291.8909 | 204.8909 | 205.2001 | 210.6908 | 255.6726 | 290.8545 | 297.5271 | 266.0546 | 227.1088 | 127.2728 | 140.9272 | 109.2183 | 137.1091 | 335.8181 | 315.6002 | 360.3816 | 458.454 | 296.5275 | 312.0728 | 341.1088 | 380.1997 | 159.9818 | 71.1455 | 79.1819 | 159.9818 | 176.3089 | 132.582 | 121.5275 | 172.382 | 136.7999 | 140.6908 | 119.1273 | 133.5275 | 87.1819 | 100.8363 | 89.9091 | 119.6908 | 125.0728 | 138.5089 | 109.0364 | 91.9091 | 331.8909 | 313.8727 | 295.8727 | 367.5999 | 202.2001 | 295.9815 | 240.5093 | 364.0909 | 380.0728 | 284.7636 | 288.9818 | 454.5271
1465 | 18.1094 | 180.8644 | 221.1238 | 241.1238 | 232.8172 | 183.9292 | 196.0236 | 187.9882 | 210.9941 | 278.8703 | 147.0413 | 172.6462 | 341.4398 | 294.888 | 175.0236 | 181.7936 | 309.7877 | 255.8821 | 209.0177 | 155.1769 | 247.4069 | 215.888 | 394.8408 | 342.1828 | 185.0177 | 293.9764 | 290.0059 | 282.0118 | 361.059 | 158.8349 | 156.888 | 139.9882 | 155.0944 | 183.1297 | 150.0059 | 158.9587 | 98.9469 | 316.8762 | 310.941 | 308.0649 | 291.0295 | 196.8585 | 163.9823 | 153.1474 | 163.9292 | 65.112 | 68.1061 | 71.9823 | 144.7818 | 309.7346 | 278.6934 | 267.8467 | 197.1179 | 174.9469 | 165.8821 | 140.9056 | 149.9646 | 230.9528 | 215.829 | 262.8998 | 235.0649 | 230.0059 | 178.9764 | 245.1474 | 233.9587 | 198.0531 | 270.9528 | 267.0472 | 372.7936 | 208.1002 | 191.0177 | 259.2123 | 291.9469 | 268.0472 | 245.0177 | 239.9764 | 404.7346 | 200.9764 | 188.0413 | 191.0295 | 297.9646 | 258.6816 | 220.9056 | 178.1946 | 236.1179 | 346.6698 | 317.8821 | 245.1238 | 225.0177 | 140.9174 | 128.0767 | 122.9174 | 142.9646 | 327.0531 | 319.9705 | 339.1297 | 432.1651 | 344.711 | 339.8349 | 341.0059 | 384.9764 | 171.9292 | 75.9705 | 88.941 | 171.9292 | 190.9174 | 159.8349 | 161.7582 | 159.0767 | 150.9174 | 129.0708 | 140.8703 | 169.7818 | 102.9056 | 84.1002 | 95.9646 | 141.8703 | 143.888 | 125.0826 | 113.9705 | 94.9823 | 340.9469 | 285.171 | 258.2241 | 405.7759 | 204.9823 | 292.0295 | 260.8762 | 395.8113 | 400.8762 | 224.3597 | 266.1356 | 430.1474
1466 | 18.1204 | 153.0384 | 224.0895 | 224.1598 | 222.0448 | 157.2046 | 208.9488 | 204.8721 | 195.1598 | 336.7954 | 161.9425 | 139.0639 | 360.9105 | 277.0895 | 141.2174 | 164.0767 | 286.0959 | 229.0192 | 219.8593 | 173.0064 | 307.0384 | 192.1215 | 409.8785 | 336.9041 | 198.211 | 325.0639 | 319.8977 | 292.9936 | 390.7442 | 146.9425 | 156.0064 | 152.8465 | 162.2622 | 168.2174 | 148.1982 | 180.8785 | 100.0895 | 288.1598 | 301.9936 | 276.0703 | 289.1726 | 180.1279 | 155.0575 | 159.0831 | 162.1343 | 70.1535 | 83.9169 | 78.9488 | 108.3453 | 314.1918 | 280.7442 | 276.0831 | 219.0959 | 189.0512 | 162.8785 | 162.9169 | 148.0128 | 251.2877 | 176.2238 | 244.1279 | 245.1918 | 249.2238 | 168.1854 | 249.3069 | 246.8338 | 220.9425 | 232.179 | 249.3261 | 372.1343 | 225.2813 | 181.1279 | 269.1598 | 302.9488 | 306.821 | 263.0895 | 228.0512 | 429.8018 | 223.9361 | 226.9169 | 191.9361 | 307 | 273.9105 | 216.1535 | 191.0831 | 276.9552 | 338.3709 | 296.1471 | 242.9361 | 247.968 | 153.8274 | 130.1023 | 141.968 | 178.9616 | 338.1087 | 370.8018 | 333.0064 | 453.0959 | 282.2174 | 283.2622 | 343.9616 | 405.8593 | 158.0639 | 95.8785 | 77.0064 | 158.0639 | 164.0703 | 136.9233 | 138.1471 | 170.0512 | 135.9936 | 129.0448 | 121.0448 | 159.1343 | 97.8913 | 96.9233 | 88.9552 | 115.0192 | 150.0384 | 135.9872 | 113.9169 | 97.9744 | 335.2685 | 326.9552 | 279 | 418.1982 | 177.2366 | 305.9105 | 237.2622 | 350.4092 | 415.8721 | 268.9744 | 257.1407 | 443.8913
1467 | 18.1314 | 159.6353 | 238.0561 | 248.6637 | 228.4581 | 188.9253 | 200.7758 | 184.9626 | 220.4298 | 304.7384 | 153.3924 | 149.0374 | 347.2055 | 291.5232 | 174.7571 | 175.5702 | 301.7287 | 232.5419 | 197.8131 | 174.355 | 313.4484 | 211.6166 | 391.7661 | 322.8782 | 230.6824 | 334.4021 | 303.3273 | 291.5702 | 351.2429 | 138.4484 | 156.5702 | 129.299 | 202.5142 | 201.701 | 179.0374 | 162.1121 | 113.6637 | 312.7571 | 301.1495 | 287.7848 | 315.4768 | 199.3087 | 163.9066 | 171.5142 | 182.645 | 94.0187 | 71.2242 | 71.1308 | 161.1779 | 344.2242 | 240.7758 | 289.0747 | 233.7197 | 196.9439 | 144.3363 | 149.9626 | 149.9253 | 295.9253 | 211.2616 | 264.299 | 274.9439 | 283.6263 | 197.1682 | 296.3834 | 221.9716 | 211.9253 | 260.1869 | 299.8318 | 393.1121 | 268.7758 | 200.8692 | 294.1682 | 295.9529 | 278.2152 | 276.7197 | 236.2616 | 400.009 | 214.4484 | 214.1495 | 182.355 | 308.3266 | 259.645 | 239.8505 | 203.9066 | 269.8692 | 395.5889 | 319.0561 | 233 | 243.5045 | 127.0934 | 145.645 | 136.8879 | 172.7758 | 355.9343 | 339.8131 | 333.9253 | 468 | 316.2616 | 323.9066 | 337.9813 | 383.2713 | 168.4858 | 77.2055 | 78.0747 | 168.4858 | 175.0187 | 125.5792 | 160.4394 | 178.355 | 135.71 | 136.3176 | 128.0374 | 179.9066 | 81.3363 | 85.0187 | 81.9253 | 118.3924 | 156.0561 | 134.1869 | 101.0561 | 94.355 | 377.1121 | 320.3176 | 279.7474 | 448.7197 | 214.1495 | 292.355 | 277.6263 | 414.1308 | 397.0277 | 265.1869 | 278.7945 | 427.3924
1468 | 18.1424 | 192.6283 | 241.6195 | 230.938 | 201.8586 | 185.9603 | 190.0532 | 183.5576 | 242.6283 | 291.2788 | 174.8364 | 150.7212 | 357.7522 | 319.6815 | 162.4956 | 153.5576 | 340.2168 | 261.2478 | 189.4869 | 192.3495 | 337 | 243.1946 | 430.4202 | 367.761 | 214.4027 | 303.062 | 269.239 | 270.1151 | 365.332 | 161.5663 | 134.6815 | 145.1549 | 177.6505 | 186.2168 | 180.0088 | 167.5354 | 95.8451 | 300.7434 | 309.062 | 328.3185 | 288.7744 | 164.0532 | 159.5885 | 147.1461 | 164.2788 | 94.8141 | 83.1549 | 77.7832 | 118.6505 | 356.4027 | 230.27 | 294.332 | 219.7744 | 193.5354 | 161.5663 | 148.4337 | 146.3717 | 291.1017 | 224.1636 | 279.969 | 270.9158 | 265.5798 | 207.0222 | 264.2478 | 273.5885 | 207.4115 | 270.4956 | 291.3407 | 399.4646 | 257.1859 | 194.6195 | 301.699 | 345.4734 | 239.571 | 262.5885 | 250.8054 | 453.9603 | 238.5885 | 222.0929 | 200.8761 | 378.1859 | 242.6108 | 232.1239 | 198.9071 | 264.7656 | 374.8983 | 322.6815 | 233.0929 | 270.031 | 132.3407 | 127.7744 | 131.1239 | 160.9071 | 404.4115 | 330.2478 | 330.6815 | 468.3717 | 330.7434 | 319.3407 | 337.2168 | 346.5178 | 194.4337 | 87.9071 | 81.8761 | 194.4337 | 176.8054 | 156.062 | 131.3407 | 197.7744 | 173.3717 | 152.4424 | 129.5044 | 174.6593 | 99.4337 | 85.938 | 78.7125 | 140.27 | 158.938 | 143.7212 | 104.2168 | 112.8451 | 383.7434 | 337.3717 | 318.0397 | 433.8451 | 222.6505 | 311.062 | 257.5044 | 420.9071 | 450.5973 | 274.5044 | 269.1461 | 449.5488
1469 | 18.1534 | 180.5241 | 261.4759 | 230.5576 | 260.1731 | 215.1779 | 222.2212 | 201.1298 | 231.6057 | 299.7404 | 200.8269 | 142 | 348.0962 | 340.0914 | 178 | 171.3894 | 348.2547 | 268.7837 | 234.7885 | 170.9616 | 336.9567 | 218.649 | 382.5144 | 330.3845 | 227.3894 | 305.3894 | 310.8124 | 304.8702 | 407.9086 | 149.2115 | 155.4375 | 148.3126 | 198.4327 | 193.2596 | 149.0433 | 152.5241 | 91.3894 | 325.1682 | 310.2645 | 308.2547 | 312.2212 | 197.8221 | 178 | 182.4375 | 172.7404 | 88.8269 | 87.2645 | 71.2163 | 138.2645 | 368.0481 | 270.8221 | 333.3607 | 243.6539 | 178.2212 | 148.8653 | 162.649 | 157.3077 | 264.6442 | 198.5625 | 279.9519 | 237.3029 | 314.8269 | 237.9616 | 272.2163 | 292.6923 | 190.1249 | 285.6971 | 299.1444 | 415.947 | 263.3894 | 213.5673 | 262.6874 | 329.6057 | 319.3943 | 279.3126 | 275.4808 | 483.4375 | 256.1779 | 225.649 | 197.5625 | 383.3943 | 292.6539 | 236.0433 | 195.9567 | 318.9616 | 402.1347 | 345.8172 | 236.8653 | 270.9567 | 143.4759 | 152.0433 | 134.8702 | 158.8221 | 385.4808 | 339.8172 | 351.6971 | 480.7788 | 352.4424 | 330.0433 | 342.8318 | 392.8751 | 208.9086 | 85.0433 | 78.2596 | 208.9086 | 202 | 159.1682 | 142.3461 | 220.7453 | 184.6539 | 135.8653 | 114.5192 | 164.7788 | 111.8751 | 83.2212 | 100.351 | 178.7885 | 157.9951 | 136.2547 | 111.3461 | 107.5673 | 405.8318 | 349.649 | 288.5192 | 429.2596 | 242.9135 | 313 | 242.7355 | 416.3558 | 437.6971 | 259.2163 | 304.6106 | 496.8318
1470 | 18.1644 | 169.8889 | 272.6667 | 263.4999 | 262.4999 | 196.9444 | 205.7223 | 204.3889 | 245.7223 | 294.6111 | 197.4445 | 143 | 305.6111 | 320.1111 | 178.1667 | 179.7222 | 376.9445 | 265.1667 | 210.9446 | 150.0001 | 333.6666 | 233.5556 | 417.2778 | 361.3889 | 235.0555 | 315.1112 | 371.4999 | 301.5 | 427.0555 | 175.7222 | 142.5555 | 112.1667 | 206.3888 | 199.4445 | 152.1112 | 142.9445 | 101.2778 | 350.4444 | 293.2222 | 333.0554 | 295.7778 | 215.2777 | 177.7778 | 169 | 167.9445 | 85.5 | 71.7778 | 75.7778 | 122.5 | 346.1111 | 289.5 | 276.5556 | 237.889 | 161.6111 | 169 | 176.2777 | 143.5556 | 302 | 210.6667 | 301.6667 | 244.8889 | 311.0556 | 215 | 277.4445 | 305.4999 | 213.611 | 278.5 | 238.389 | 458.2222 | 270.3333 | 202.4999 | 299.6111 | 341.0555 | 304.4444 | 246.0002 | 261.611 | 469.611 | 237.8333 | 238.7222 | 208.6666 | 366.6109 | 283.0555 | 237.6667 | 195.7223 | 294.1666 | 384.5556 | 385.5 | 256.3889 | 267.8888 | 154.2778 | 152.7778 | 132 | 176.5555 | 372.7777 | 378.3888 | 345.1667 | 498.2222 | 319.2223 | 328.3332 | 317.8889 | 367.2777 | 229.3333 | 86.3889 | 83.8889 | 229.3333 | 201.1111 | 185 | 149.6667 | 194.389 | 177 | 155 | 127.6667 | 183.2223 | 87.6111 | 67.2778 | 86.0556 | 151.9444 | 178.6666 | 162.8888 | 118.8333 | 99.889 | 382.1668 | 365.5001 | 299.4444 | 434 | 239.6666 | 313.1667 | 257.8333 | 379.3889 | 431.5 | 264.7223 | 295.1111 | 471.6667
1471 | 18.1754 | 169.2213 | 281.1503 | 239.5605 | 238.5605 | 197.7641 | 217.6608 | 209.0324 | 258.5428 | 305.0678 | 204.1858 | 159.0501 | 301.9854 | 326.8173 | 182.8998 | 173.9823 | 395.714 | 288.596 | 258.4217 | 183.3215 | 296.714 | 239.8789 | 441.0355 | 365.8288 | 218.3894 | 336.5783 | 345.3215 | 291.2359 | 414.8674 | 170.7965 | 135.2035 | 131.7286 | 178.6608 | 208.0856 | 188.8142 | 160.6284 | 122.7965 | 341.9499 | 283.2568 | 284.0856 | 308.7286 | 202.5251 | 174.5428 | 154.4426 | 184.0678 | 93.4572 | 85 | 72.4071 | 131.2035 | 331.882 | 296.1681 | 302.5929 | 268.1002 | 172.3392 | 185.8643 | 162.643 | 169.7286 | 320.4248 | 205.2714 | 310.761 | 261.8998 | 312.8142 | 217.8497 | 285.6785 | 264.5605 | 191.9676 | 271.1534 | 278.8643 | 442.4217 | 245.4603 | 180.0532 | 291.9144 | 311.1889 | 277.7965 | 311.3539 | 223.8497 | 447.357 | 232.8288 | 217.1357 | 185.1858 | 310.0177 | 250.882 | 249.4749 | 209.9676 | 262.7965 | 408.0825 | 377.882 | 263.6106 | 234.5104 | 157.7787 | 149.4071 | 132.1357 | 167.9144 | 352.9499 | 350.7965 | 348.4749 | 503.4248 | 337.0825 | 284.0177 | 333.5428 | 357.239 | 232.8288 | 93.7463 | 81.5929 | 232.8288 | 186.9499 | 184.1858 | 146.1033 | 215.8967 | 177.6106 | 155.0678 | 154.7109 | 203.4572 | 99.1534 | 88.8643 | 87.0678 | 152.6962 | 159.7996 | 144.4071 | 115.8643 | 130.3716 | 418.1858 | 389.643 | 290.882 | 417.1357 | 217.9499 | 315.0501 | 239.6284 | 368.9322 | 439.3215 | 275.9823 | 278.9823 | 482.3215
1472 | 18.1864 | 186.0801 | 241.3589 | 258.7561 | 259.9199 | 223.8432 | 212.7596 | 182 | 265.1986 | 304.6376 | 194.6829 | 145.439 | 339.2334 | 391.1568 | 209.7213 | 159.5192 | 429.1533 | 337.5192 | 225.8049 | 175.2822 | 334.9617 | 197.3206 | 462.9093 | 336.6411 | 212.8851 | 371.8397 | 336.6411 | 267.4007 | 466.7178 | 168.6411 | 137.3589 | 128.0801 | 172.2369 | 220.6759 | 198.2369 | 180.9164 | 118.7178 | 355.3206 | 328.2404 | 300.5227 | 303.6376 | 197.0418 | 180.9582 | 185.6759 | 184.1185 | 85.439 | 85.4007 | 78.8014 | 136.0836 | 344.561 | 273.1638 | 296.3589 | 241.8397 | 178.6028 | 182.0766 | 145.6028 | 161.5923 | 337.7561 | 209.7213 | 281.5227 | 286.878 | 319.5923 | 257.6411 | 295.6411 | 283.6759 | 220.122 | 288.9234 | 277 | 407.6411 | 289.554 | 224.1638 | 276.8397 | 353.6376 | 276.4425 | 273.4773 | 260.5157 | 466.7213 | 207.4879 | 218.9199 | 175.2021 | 322.4773 | 263.4007 | 257.7631 | 233.9547 | 262.5645 | 366.6376 | 390.4808 | 273.0418 | 265.3135 | 141.1603 | 154.1986 | 133.8397 | 152.1185 | 363.2752 | 351.9269 | 351.7142 | 526.6063 | 304.2928 | 298.3206 | 340.1986 | 388.0801 | 202.7596 | 102.4773 | 76.2404 | 202.7596 | 199.2787 | 173.4007 | 174.439 | 188.7631 | 187.122 | 155.1986 | 136.5192 | 197.2822 | 113.3972 | 87.8014 | 87.1986 | 174.7178 | 209.1951 | 151.4425 | 115.5227 | 108.561 | 404.8362 | 368.9965 | 300.7561 | 416.1149 | 234.3659 | 303.5227 | 259.1951 | 372.0872 | 430.4007 | 267.2091 | 264.9199 | 473.4007
1473 | 18.1974 | 188.294 | 236.1259 | 234.3865 | 260.0167 | 243.1514 | 210.9243 | 182.647 | 256.5546 | 288.2605 | 213.6303 | 140.0167 | 291.4709 | 372.5881 | 217.0757 | 156.6805 | 390.7307 | 330.8908 | 258.1849 | 189.7227 | 342.6725 | 201.4622 | 395.6805 | 344.3697 | 241.6557 | 372.3108 | 343.353 | 272.647 | 457.7307 | 175.8151 | 130.647 | 129.5546 | 155.8821 | 194.4622 | 176.7984 | 158.5713 | 107.6973 | 361.6805 | 329.1514 | 317.4454 | 288.647 | 207.4287 | 170.6638 | 160.2016 | 173.1681 | 79.8319 | 89.4454 | 86.8908 | 157.5968 | 349.6135 | 296.706 | 288.4454 | 239.5378 | 194.8741 | 162.3108 | 163.0757 | 113.4032 | 310.6135 | 220.8654 | 299.2773 | 270.8573 | 271.1259 | 268.42 | 302.353 | 257.9243 | 231.4287 | 313.7729 | 278.5713 | 414.1681 | 246.8573 | 246.5968 | 275.647 | 332.7314 | 291.2438 | 257.7562 | 234.4956 | 468.3449 | 254.286 | 219.294 | 191.0335 | 303.5211 | 264.4876 | 280.2184 | 199.9243 | 289.1346 | 350.5378 | 393.4119 | 285.0924 | 223.1849 | 142.7227 | 146.3865 | 132 | 142.5546 | 328.3027 | 393.0255 | 315.294 | 558.572 | 382.7314 | 303.3865 | 326.6557 | 387.9833 | 198.8908 | 86.0167 | 82.1427 | 198.8908 | 191.0924 | 178.5546 | 166.6135 | 207.5211 | 195.5633 | 146.0924 | 131.2773 | 208.58 | 94.0757 | 97.2773 | 79.3865 | 158.7984 | 174.7816 | 166.9833 | 133.0924 | 114.6303 | 380.9243 | 346.1849 | 275.3697 | 386.7897 | 269.8573 | 320.5378 | 227.4622 | 411.8828 | 434.6303 | 320.992 | 262.4287 | 475.7816
1474 | 18.2084 | 198.3819 | 256.2669 | 245.6961 | 270.6283 | 224.8953 | 218.0103 | 189 | 258.5442 | 277.6489 | 211.0472 | 149.4764 | 244.8008 | 394.963 | 207.1622 | 183 | 441.4497 | 319.5811 | 257.9056 | 187.4189 | 307.2772 | 205.7906 | 425.6653 | 345.2772 | 200.5708 | 392.9056 | 341.2936 | 284.1314 | 513.0575 | 172.3244 | 140.4558 | 135.9425 | 192.6386 | 196.1725 | 166.0472 | 173.3717 | 142.115 | 385.0678 | 312.3614 | 312.9425 | 301.9117 | 190.8481 | 183.963 | 172 | 168.037 | 90.4086 | 84.6283 | 75.5811 | 134 | 334.0103 | 287.2464 | 283.9425 | 237.4086 | 171.4867 | 183.7433 | 155.5708 | 138.4661 | 292.9158 | 247.0103 | 303.4661 | 242.8275 | 290.4292 | 300.5339 | 296.5236 | 264.8008 | 218.7228 | 338.8008 | 295.5708 | 401.2875 | 215.2669 | 227.2936 | 284.4086 | 281.2094 | 269.9158 | 276.2567 | 256.9056 | 409.7022 | 207.5811 | 231.4764 | 209.5339 | 293.1889 | 229.1622 | 295.2978 | 211.5133 | 256.078 | 339.9261 | 369.6756 | 289.4558 | 223.9528 | 137.2772 | 157.5914 | 133.3614 | 148 | 292.7331 | 354.2094 | 329.1519 | 470.2936 | 332.9897 | 315.0103 | 283.7906 | 380.8275 | 189.2567 | 95.8953 | 109.963 | 189.2567 | 188.4394 | 185.0472 | 152.6858 | 193.4189 | 151.6858 | 150.9795 | 132.9528 | 175.2094 | 86.3614 | 98.4292 | 91.4292 | 145.6386 | 153.5236 | 157.3142 | 135.6756 | 110.0575 | 387.9056 | 345.4867 | 282.1417 | 419.8111 | 237.0103 | 312.8583 | 234.9322 | 346.1889 | 426.078 | 266.922 | 245.4292 | 455.0472
1475 | 18.2194 | 175.5319 | 250.8191 | 228.3403 | 280.3295 | 229.9677 | 206.2656 | 189.234 | 230.4042 | 250.0531 | 222.1062 | 148.6274 | 224.5958 | 371.8083 | 200.8191 | 182.766 | 391.0747 | 314.7129 | 240.9361 | 192.4042 | 279.6705 | 202.8299 | 462.3826 | 323.7552 | 213.649 | 376.7552 | 376.1278 | 327.1809 | 503.7129 | 155.0747 | 170.234 | 140.8406 | 179.5958 | 173.8722 | 173.0108 | 165.3087 | 124.2448 | 360.351 | 327.2764 | 319.0108 | 357.1386 | 180.1809 | 164.5528 | 170.9469 | 191.8938 | 107.4896 | 88.3618 | 72.234 | 134.234 | 313.0215 | 318.3934 | 294.2233 | 253.7875 | 150.7021 | 174.0531 | 167.3618 | 149.7129 | 255.8938 | 231.2871 | 315.0639 | 260.6813 | 275.0108 | 284.6066 | 303.2233 | 244.2448 | 241.883 | 326.6698 | 307.3618 | 359.5958 | 207.4788 | 263.2979 | 304.0639 | 282.649 | 238.3934 | 284.4257 | 239.7021 | 458.8091 | 203.883 | 224.6597 | 198.6382 | 325.6066 | 218.8406 | 240.8083 | 231.7129 | 214.234 | 282.1493 | 376.6282 | 315.7236 | 212.6597 | 115.5212 | 139.1062 | 141.9576 | 147.2979 | 296.6597 | 363.6059 | 335.3726 | 500.2132 | 350 | 299.9892 | 288.2017 | 405.5851 | 198.4788 | 96.2871 | 82.9469 | 198.4788 | 160.9253 | 197.0424 | 150 | 194.8938 | 148.0639 | 183.8299 | 125.7552 | 174.5427 | 96.9469 | 86.4042 | 78.7021 | 133.883 | 157.0639 | 159.766 | 149.1809 | 101.649 | 367.8938 | 320.8406 | 308.883 | 383.649 | 216.6066 | 284.9469 | 259.5319 | 379.4257 | 387.5104 | 306.7767 | 233.9892 | 463.2979
1476 | 18.2304 | 173.8103 | 258.1637 | 243.5432 | 312.7414 | 272.5432 | 243.5432 | 193.1982 | 238.026 | 251.9227 | 236.707 | 169.1469 | 215.293 | 390.8018 | 206.4828 | 183.7154 | 369.5172 | 309.5257 | 247.3535 | 202.0949 | 241.5692 | 193.4828 | 490.5868 | 341.7843 | 209.4483 | 384.052 | 361.7414 | 318.6725 | 486.1124 | 131.5777 | 170.5777 | 122.0429 | 167.8363 | 188.3879 | 164.8531 | 146.293 | 109.9656 | 369.5945 | 349.9311 | 304.9051 | 330.5861 | 173.2242 | 211.4139 | 164.293 | 175.4828 | 80.2242 | 76 | 75.0344 | 134.707 | 276.1469 | 337.5348 | 304.0176 | 219.9311 | 156.5172 | 185.6205 | 154.3535 | 138.4483 | 241.1637 | 243.1982 | 306.1897 | 211.9656 | 253.1469 | 259.3879 | 316.7674 | 229.4483 | 246.1722 | 367.8194 | 299.9136 | 345.7674 | 193.1553 | 254.7674 | 296.2242 | 278.5777 | 258.1813 | 271.8531 | 243.5777 | 393.3795 | 204.5517 | 211.0084 | 207.8967 | 296.7674 | 196.681 | 257.345 | 221.4828 | 213.4139 | 241.8791 | 318.8018 | 290.8707 | 195.6465 | 125.0604 | 155.7758 | 122.3619 | 139.4139 | 278.2242 | 414.9227 | 314.5945 | 443.9571 | 349.7414 | 312.6725 | 333.319 | 411.681 | 187 | 106.6465 | 75.2586 | 187 | 181.1897 | 218.319 | 155.8187 | 175.8967 | 141.6465 | 171.5088 | 138.3535 | 154.0604 | 91.8447 | 95.1897 | 83.8707 | 131.8363 | 149.7414 | 154.3795 | 140.6725 | 100.5517 | 348.1209 | 298.8103 | 309.293 | 378.026 | 188.9311 | 278.0344 | 256 | 355.6036 | 409.2157 | 288.319 | 248.4828 | 450.371
1477 | 18.2414 | 185.1504 | 265.292 | 225.3009 | 309.1593 | 252.1769 | 227.5664 | 207.7168 | 214.6017 | 210.292 | 233.239 | 127.8143 | 217.2034 | 373.0265 | 205.5487 | 198.6017 | 368.1857 | 337.0265 | 243.8496 | 197.6991 | 208.5399 | 187.3097 | 417.0533 | 371.5929 | 197.1593 | 348.6727 | 360.1416 | 302.8584 | 444.1861 | 126.6727 | 160.0088 | 153.4513 | 157.8761 | 187.7433 | 209.0177 | 153.1593 | 103.8496 | 405.9291 | 342.5134 | 310.8584 | 353.8054 | 165.1681 | 193.292 | 174.416 | 171.8584 | 78.3983 | 75.1504 | 79.4513 | 126.708 | 236.2302 | 296.8319 | 262.947 | 206.2832 | 158.3009 | 215.239 | 147.7345 | 129.1328 | 249.2832 | 262.6727 | 295.6991 | 201.7433 | 207.1416 | 267.6638 | 298.6815 | 213.3362 | 276.0442 | 311.8054 | 330.3097 | 324.9912 | 176.1239 | 238.3806 | 294.2567 | 265.8849 | 219.2832 | 315.5929 | 231.0265 | 368.5752 | 212.4513 | 236.4424 | 188.1328 | 276.9823 | 207.1504 | 226.2832 | 219.9823 | 196.9912 | 262.053 | 304.4248 | 294.3895 | 195.8937 | 115.2567 | 156.4689 | 143.4424 | 122.8496 | 274.9823 | 370.3185 | 356.5929 | 410.4248 | 345.7345 | 303.2302 | 323.8407 | 423.708 | 183.6017 | 107.8849 | 75.7256 | 183.6017 | 164.7521 | 204.3097 | 192.7345 | 154.1504 | 144.0177 | 150.8319 | 137.3983 | 139.0177 | 108.1681 | 85.1239 | 83.4248 | 122.7256 | 143.3273 | 132.6903 | 123.8672 | 111.4248 | 314.7786 | 309.584 | 319.1328 | 357.1504 | 171.1769 | 285.2832 | 258.4071 | 321.0442 | 374.4513 | 279.4071 | 244.4336 | 401.8319
1478 | 18.2524 | 181.5389 | 259.1534 | 219.7705 | 295.6917 | 229.1539 | 233.4622 | 206.9233 | 194.3078 | 205.3845 | 262.0762 | 160.6933 | 178.6161 | 356.2311 | 217.9223 | 181.2316 | 336.5378 | 309.459 | 247.1534 | 208 | 226.9223 | 171.1539 | 471.9984 | 359.4622 | 186.7694 | 373.7684 | 366.6938 | 299.9995 | 499.843 | 151.3067 | 155.2311 | 138.3078 | 145.3078 | 164.4606 | 196.8461 | 147.386 | 109 | 355.9228 | 379.9995 | 304.614 | 384.3834 | 151.0777 | 188.2306 | 180.3839 | 165.1523 | 95.9223 | 72.4622 | 65.0772 | 132.0772 | 268.7679 | 316.3083 | 296.0762 | 201.8445 | 143.5373 | 237.459 | 140.1555 | 135.5383 | 246.5373 | 287.1528 | 302.4606 | 178.4606 | 208.3078 | 298.3067 | 318.0762 | 178.0767 | 250.7705 | 343.7684 | 311.23 | 329.1534 | 188.6922 | 264.0751 | 314.6155 | 245.7689 | 220.0767 | 297.1528 | 215.0005 | 366.1539 | 192.8451 | 224.0762 | 198.2316 | 286.0762 | 195.8445 | 223.6912 | 230.615 | 196.0751 | 226.1539 | 300.0751 | 316.8451 | 171 | 131.6145 | 128.3845 | 133.5383 | 130.0005 | 289.1544 | 350.4622 | 339.0762 | 416.3855 | 333.8466 | 338.0762 | 335.4617 | 429.6928 | 162.5383 | 92.3855 | 69.2311 | 162.5383 | 140.8466 | 189.8466 | 179.7694 | 150.0772 | 130.1534 | 165.8456 | 159.5389 | 126.0767 | 89.9228 | 96.3072 | 86.3078 | 118.2316 | 116.8466 | 145.7684 | 119.3083 | 113.5383 | 272.5389 | 296.5373 | 326.3078 | 353.385 | 150.77 | 282.3834 | 272.0767 | 292.0772 | 357.9228 | 296.6161 | 242.2311 | 412.3829
1479 | 18.2634 | 188.0059 | 255.3148 | 249.8221 | 280.175 | 226.1779 | 251.8192 | 210.6706 | 199.3236 | 196.0176 | 245.6676 | 184.6559 | 190.8397 | 355.3558 | 202.9941 | 198.6676 | 321.3382 | 259.4927 | 235.0059 | 207.3353 | 209.6676 | 165.5191 | 440.828 | 374.1632 | 190.0029 | 353.175 | 393.8456 | 291.3265 | 433.8309 | 130.0029 | 158.0117 | 137.8397 | 145.3382 | 143.6706 | 194.0059 | 172.8456 | 106.175 | 346.8426 | 370.3294 | 281.1486 | 361.1603 | 165.3412 | 186.825 | 166.3324 | 136.9883 | 81.6588 | 86.8309 | 69.8368 | 137.0029 | 237.0088 | 328.3353 | 281.828 | 168 | 122.1574 | 185.8309 | 169.6794 | 132.5015 | 220.172 | 262.3441 | 283.8309 | 166.478 | 206.6764 | 280.6588 | 305.1574 | 172.0029 | 275.5044 | 327.6618 | 295.1691 | 306.5367 | 187.3324 | 232.1486 | 311.1867 | 237.5044 | 214.6676 | 274.5044 | 231.9824 | 364.0088 | 171.828 | 201.8514 | 210.5161 | 263.3529 | 161.8338 | 201.1632 | 221.0059 | 161.1574 | 223.5103 | 263.3294 | 292.8368 | 164.5191 | 113.6647 | 123.6706 | 132.828 | 137.8426 | 291.694 | 361.3412 | 320.0088 | 427.1896 | 339.0147 | 331.1398 | 339.9941 | 436.3412 | 155.6794 | 106.347 | 79.8221 | 155.6794 | 151.3324 | 198.3382 | 182.3382 | 155.3353 | 116.6764 | 153.6676 | 164.8426 | 120.8338 | 82.6706 | 87.1662 | 88.1662 | 129.1867 | 126.3353 | 127.5015 | 131.1691 | 106.0147 | 275.6823 | 271.3353 | 321.1867 | 361.3324 | 163.8397 | 250.8514 | 268.3294 | 298.9971 | 349.3412 | 302.1926 | 250.4956 | 373.0088
1480 | 18.2744 | 179.0708 | 285.0368 | 266.6799 | 264.3938 | 198.9689 | 276.5354 | 202.0368 | 211.1813 | 166.7139 | 242.0368 | 199.2522 | 177.8584 | 312.3598 | 217.8187 | 195.7507 | 311.3938 | 273.0708 | 221.0737 | 206.8555 | 206.9292 | 138.711 | 449.6431 | 375.1813 | 183.7507 | 336.6799 | 370.6799 | 309.7818 | 439.9632 | 128.3909 | 141.7478 | 126.4646 | 132.5383 | 138.4278 | 180.6091 | 153.4278 | 96.8187 | 324.2182 | 374.1445 | 311.1077 | 384.8498 | 151.2861 | 204.7478 | 168.8923 | 152.1813 | 92.323 | 89.2153 | 63.6431 | 131.1077 | 220.2153 | 326.0708 | 294.3909 | 168.8923 | 133.7167 | 188.0368 | 149.7847 | 130.7139 | 210.8215 | 257.5636 | 286.3938 | 199.5383 | 184.1105 | 290.6091 | 315.4675 | 169.677 | 276.388 | 331.3598 | 291.5354 | 247.034 | 190.0708 | 255.1474 | 281.6402 | 229.4646 | 215.3201 | 265.2153 | 254.8612 | 340.4335 | 180.8215 | 179.5665 | 186.034 | 231.0708 | 159.5722 | 201.6459 | 213.677 | 177.8923 | 203.8584 | 268.5722 | 288.2493 | 131.4646 | 120.2124 | 119.1416 | 142 | 119.6799 | 243.3938 | 339.612 | 305.1785 | 384.0737 | 312.2153 | 366.2919 | 338.2211 | 431.745 | 139.5325 | 84.6062 | 98.2861 | 139.5325 | 147.8244 | 189.1077 | 171.8584 | 148.4306 | 101.7847 | 148.2522 | 147.9292 | 121.9632 | 75.4646 | 87.6431 | 92.5694 | 95 | 121.2153 | 121.966 | 129.4986 | 80.6431 | 247.9292 | 258.5412 | 287.3569 | 359.0737 | 154.0708 | 221.2493 | 267.8612 | 299.3598 | 329.0397 | 274.2407 | 261.7478 | 361.7478
1481 | 18.2854 | 184.763 | 279.4337 | 251.994 | 251.2772 | 171.538 | 266.6006 | 192.0462 | 189 | 167.1386 | 229.9478 | 187.1505 | 169.3353 | 288.3294 | 237.3294 | 193.2429 | 290.8376 | 269.9881 | 202.0462 | 219.8092 | 201.237 | 159.468 | 437.5082 | 353.3815 | 179.9076 | 321.994 | 358.0924 | 336.8033 | 438.1267 | 138.8495 | 157.6647 | 125.9076 | 122.6825 | 145.3815 | 162.7109 | 159.237 | 115.9478 | 300.6304 | 365.9598 | 307 | 441.9016 | 149.7168 | 215.5142 | 172.8092 | 132.4799 | 79.0462 | 80.8092 | 65.2429 | 125.4739 | 213.1446 | 329.0924 | 309.0462 | 172.237 | 116.4799 | 179.763 | 156.8554 | 133.8092 | 205.231 | 316.5663 | 281.0984 | 177.6647 | 158.3353 | 273.2832 | 281.225 | 179.3294 | 312.5201 | 305.4218 | 293.8092 | 265.769 | 184.5082 | 220.7168 | 304.7168 | 222.231 | 224.0924 | 251.0864 | 217.231 | 309.9657 | 175.8033 | 213.6587 | 201.1446 | 235.237 | 157.9598 | 177.5723 | 225.4277 | 175.7049 | 191.1386 | 250.9359 | 291.3294 | 131.2891 | 133.6647 | 127.2832 | 137.0402 | 112.0522 | 238.0984 | 299.0924 | 309.6244 | 368.4799 | 299.6125 | 325.0119 | 290.4858 | 469.3294 | 157.5604 | 91.0462 | 91.5663 | 157.5604 | 118.6587 | 188.4337 | 166.006 | 130.6647 | 110.3815 | 129.0924 | 142.237 | 129.8554 | 75.2891 | 83.5201 | 107.6647 | 92.5201 | 111.0924 | 110.0984 | 140.237 | 77.2832 | 242.0462 | 218.6244 | 291.2891 | 345.3875 | 156.5201 | 221.6587 | 246.4456 | 282.1967 | 307.4918 | 338.8897 | 276.1386 | 375.5663
1482 | 18.2964 | 193.0763 | 293.7969 | 240.0763 | 222.939 | 207.3298 | 214.8122 | 190.4366 | 188.3908 | 157.6397 | 217.4061 | 205.142 | 172.7664 | 273.142 | 224.1725 | 207.4061 | 242.3098 | 235.3908 | 197.7969 | 211.081 | 198.2031 | 138.2336 | 400.3098 | 350.7359 | 178.6597 | 299.1115 | 340.8627 | 315.5481 | 389.3603 | 117.4061 | 147.939 | 130.3451 | 162.5176 | 147.8122 | 158.0763 | 153.3603 | 101.3756 | 329.3451 | 382.5329 | 306.1878 | 424.878 | 158.4214 | 190.2336 | 172.6092 | 134.4719 | 76.2183 | 80.8122 | 77.9847 | 114.7359 | 216.5786 | 319.1725 | 300.7359 | 166.5634 | 129.4366 | 187.2641 | 151.7969 | 130.5634 | 186.4214 | 296.9237 | 292.2336 | 172.8122 | 161.1573 | 265.1878 | 244 | 166.7817 | 306.061 | 284.5786 | 293 | 280.5176 | 148.122 | 222.3146 | 306.9237 | 202 | 215.1878 | 235.5434 | 205.5281 | 343.939 | 155.3603 | 190.7969 | 205.7969 | 233.4214 | 171.2841 | 178.9847 | 217.9237 | 157.3251 | 180.6244 | 207.0458 | 278.7817 | 137.9542 | 129.8275 | 116.9542 | 117.6244 | 126.0153 | 244.7664 | 291.4061 | 319.1115 | 373.1115 | 279.1725 | 356.1573 | 314.9542 | 465.919 | 129.4519 | 86.3908 | 78.2183 | 129.4519 | 96.6092 | 193.4566 | 180.5634 | 128.858 | 114.2336 | 124.858 | 137.5786 | 124.1878 | 82.7664 | 73.2031 | 96.9237 | 85.4519 | 102.1878 | 119.4061 | 133.142 | 73.858 | 243.2793 | 230.3451 | 303.6397 | 358.5176 | 142.7512 | 200.4214 | 289.1115 | 302.2336 | 341.2383 | 289.5024 | 260.5481 | 361.2031
1483 | 18.3074 | 218.1682 | 294.9381 | 253.7551 | 211.3068 | 185.9852 | 221.6608 | 201.5074 | 187.0767 | 168 | 214.6932 | 193.0148 | 167.6608 | 246.3715 | 222.1682 | 211.3688 | 272.0767 | 226.5397 | 187.9556 | 186.2449 | 200.2921 | 146.7847 | 414.3567 | 337.0148 | 146.3068 | 276.2153 | 304.354 | 299.6312 | 376.4159 | 117.7079 | 131.1386 | 114.4926 | 141 | 152.2921 | 176.2773 | 149.2449 | 95 | 311.1238 | 359.1858 | 303.646 | 390.2006 | 158.4011 | 197.0619 | 169.6165 | 107.4603 | 84.0148 | 84 | 74.2153 | 102.3068 | 209.0619 | 305.7551 | 288.0915 | 157.4307 | 138.5693 | 198.1858 | 151.2153 | 120.354 | 190.9233 | 280.6608 | 299.7079 | 169.7551 | 149.5693 | 255.9704 | 234.7403 | 162.8614 | 330.829 | 277.4926 | 291.7079 | 259.2153 | 172.6784 | 202.3688 | 291.5222 | 206.7375 | 210.9233 | 262.4482 | 242.6608 | 332.5222 | 145.8614 | 184.4011 | 200.6932 | 238.7847 | 145.7375 | 170.9085 | 199.2921 | 185.9085 | 186.1386 | 214.0471 | 277.0148 | 131.3835 | 128.3863 | 105.646 | 122.7079 | 125.4778 | 241.8142 | 289.9852 | 302.4603 | 351.9381 | 272.646 | 349.9529 | 297.6165 | 489.8937 | 145.3688 | 83.5693 | 83.8614 | 145.3688 | 97.4926 | 157.9381 | 167.1238 | 134.525 | 115.2477 | 137.2006 | 133.9381 | 120.1386 | 73.354 | 77.6608 | 76.7847 | 99.8614 | 97.0619 | 119.0619 | 117.354 | 87.4926 | 268.9529 | 216.8614 | 306.8937 | 338.5074 | 132.3688 | 209.2301 | 259.9704 | 303.2477 | 302.0148 | 267.6608 | 242.4778 | 358.1238
1484 | 18.3184 | 230.3986 | 303.1382 | 232.3658 | 221.9921 | 173.8618 | 234.0895 | 204.7237 | 185.3092 | 168 | 195.5855 | 200.8131 | 183.504 | 210.6014 | 245.325 | 220.2276 | 262.5697 | 204.1382 | 157.0487 | 210.1382 | 198.6263 | 151.9185 | 369.4224 | 344.5855 | 159.496 | 273.8131 | 302.4553 | 295.0974 | 359.2605 | 116.8698 | 121.6263 | 117.4224 | 144.1869 | 154.9513 | 167.496 | 173.5934 | 98.6421 | 289.9434 | 341.4145 | 308.2763 | 382.0487 | 135.9513 | 179.3013 | 150.4553 | 129.6342 | 92.7237 | 82.6342 | 74.3171 | 118.2276 | 199.0408 | 288.2356 | 302.0895 | 160.1382 | 137.6829 | 171.5368 | 151.0895 | 121.6421 | 187.6829 | 286.4881 | 290.2197 | 149.7316 | 143.2197 | 227.8539 | 209.8777 | 164.1789 | 351.846 | 274.7316 | 289.7316 | 255.9026 | 149.6421 | 208.496 | 310.7316 | 221.0408 | 201.7644 | 221.0408 | 247.5776 | 344.675 | 155.1461 | 167.4145 | 186.3658 | 234.1302 | 156.3986 | 156.4553 | 201.496 | 169.4066 | 179.1302 | 199.6421 | 283.4473 | 150.0895 | 99.7316 | 111.1869 | 121.1869 | 112.8698 | 255.4881 | 275.5855 | 329.6421 | 361.9592 | 268.8539 | 361.3986 | 281.6421 | 454.8947 | 153.7724 | 79.9513 | 85.8618 | 153.7724 | 99.7395 | 158.6263 | 158.9185 | 106.6014 | 91.0974 | 125.8618 | 138.2684 | 116.7724 | 70.3171 | 86.9026 | 76 | 101.1789 | 92.2763 | 112.9105 | 112.496 | 84.0487 | 274.0249 | 222.9592 | 282.3658 | 347.4145 | 139.4066 | 215.0815 | 232.5368 | 285.9264 | 307.0815 | 280.5447 | 236.6987 | 340.5855
1485 | 18.3294 | 207.0403 | 306.7601 | 238.6795 | 200.2015 | 175.7582 | 236.9981 | 196.5202 | 169.8388 | 168 | 181.6008 | 190.4798 | 186.2418 | 230.8407 | 262.8791 | 223.6392 | 222.2783 | 206.0806 | 169.2783 | 212.0806 | 183.7179 | 165.762 | 390.8023 | 328.6814 | 155.5586 | 259.1612 | 297.0422 | 308.5202 | 328.9213 | 122.5624 | 104.7985 | 143.1209 | 151.1612 | 147.5202 | 150.8426 | 173.0019 | 111 | 271.7582 | 348.4414 | 316 | 380.3627 | 130.1996 | 155.476 | 152 | 127.1593 | 85.9597 | 80.6392 | 70.8004 | 121.6392 | 189.2399 | 292.0844 | 290.6027 | 160.881 | 140.2399 | 151.5605 | 152.7985 | 130.881 | 186.6411 | 271.0768 | 276.1593 | 158.2802 | 127.7198 | 207.4798 | 235.4414 | 155.8407 | 313.0403 | 273.6833 | 300.6795 | 239.3608 | 165.3589 | 196.881 | 314.2418 | 211 | 189.9981 | 203.5624 | 222.3589 | 323.2821 | 171.2015 | 173.2418 | 191.7198 | 214.2821 | 139.9981 | 153.4414 | 189.4011 | 157.5221 | 167.6795 | 203.1228 | 273.2783 | 145.0806 | 108.5202 | 119.8407 | 135.5989 | 123.1209 | 230.2399 | 268.7985 | 337.6814 | 373.4395 | 255.9175 | 341.6392 | 291.1209 | 425.6795 | 153 | 73.9597 | 81.2802 | 153 | 124.5605 | 138.1996 | 171.8023 | 123.4817 | 104.0403 | 123.4395 | 127.5605 | 121.7582 | 70.6392 | 71.3205 | 76 | 99.3186 | 102.6392 | 107.1228 | 109.7582 | 87.6411 | 236.9578 | 225.5624 | 297.5567 | 355.4011 | 132.0806 | 200.0384 | 221.1977 | 325.1593 | 287 | 274.2015 | 263.403 | 335.7179
1486 | 18.3404 | 209.0443 | 301.7123 | 243.7478 | 185.7566 | 186.1813 | 246.4158 | 204.8364 | 177.4601 | 160.4335 | 172.7389 | 192.5044 | 164.9734 | 229.8009 | 269.7301 | 228.7212 | 234.4601 | 193.1902 | 175.9114 | 196.4158 | 194.1548 | 143.9823 | 357.1548 | 324.5842 | 165.7123 | 241.7035 | 270.7123 | 311.7921 | 317.2611 | 104.8009 | 124.0354 | 134.7566 | 139.7566 | 141.2079 | 135.3408 | 160.8098 | 104.6946 | 281.4247 | 336.2699 | 312.469 | 360.3142 | 134.2522 | 187.6858 | 156.0354 | 137.0266 | 89.2522 | 87.9911 | 69.2699 | 125.7123 | 194.7921 | 247.7478 | 266.0177 | 155.7921 | 132.9291 | 148.2611 | 162.4513 | 124.531 | 172.2168 | 289.3804 | 276.6946 | 155.7478 | 135.0443 | 207.2345 | 219.9911 | 146.7301 | 309.7478 | 234.1902 | 295.407 | 242.3497 | 172.9734 | 173.8848 | 302.0532 | 210.7478 | 204.4601 | 184.2877 | 237.5399 | 306.7921 | 157.2611 | 148.9468 | 189.7123 | 202.0797 | 151.4335 | 140.5133 | 174.469 | 145.0709 | 175.0177 | 178.7832 | 288.2345 | 134.9646 | 105.7389 | 111.9911 | 140.4424 | 114.2522 | 229.9911 | 280.7212 | 326.2699 | 372.7035 | 266.1193 | 350.2522 | 290.8275 | 422.4247 | 155.5222 | 75.4867 | 78.2434 | 155.5222 | 117.9823 | 139.4778 | 152.2788 | 111.0443 | 105.7921 | 123.2079 | 127.2877 | 131.9291 | 78.4956 | 66.7566 | 75.4956 | 110.7301 | 114.0266 | 87.3231 | 123.9646 | 75.7389 | 256.2256 | 210.8275 | 322.6769 | 353.0797 | 133.8187 | 212.469 | 241.2256 | 324.1813 | 288.7655 | 256.7301 | 246.6196 | 358.0089
1487 | 18.3514 | 224.529 | 296.6707 | 251.9932 | 191.4386 | 161.1485 | 219.3225 | 227.3873 | 159.2389 | 140.116 | 170.3225 | 194.2645 | 157.058 | 235.8907 | 271.9932 | 214.7679 | 228.6707 | 176.7031 | 153.6707 | 175.4062 | 168.6127 | 141.4386 | 329.2321 | 350.0904 | 155.3806 | 226.2645 | 260.6451 | 319.1228 | 322.8515 | 116.4454 | 131.5034 | 136.7355 | 138.8259 | 128.058 | 160.1809 | 178.884 | 83.6194 | 245.9488 | 334.2713 | 306.7611 | 379.3549 | 128.1228 | 164.1485 | 170.6451 | 143.884 | 87.884 | 88.4386 | 79.1741 | 114.8515 | 207.1485 | 241.4454 | 270.942 | 166.5614 | 114.8191 | 148.8259 | 153.4062 | 123.6262 | 161.2645 | 245.2969 | 264.3481 | 151.0324 | 147.0904 | 197.7679 | 206.1552 | 142.6451 | 313.2321 | 212.942 | 255.2713 | 267.5358 | 158.4454 | 141.3481 | 313.9164 | 206.5614 | 191.529 | 204.9352 | 249.1485 | 318.6194 | 161.529 | 132.8515 | 182.5546 | 224.2389 | 132.1741 | 148.7031 | 164.7935 | 165.5614 | 186.2901 | 185.7679 | 266.6006 | 122.7355 | 106.2321 | 108.2065 | 122.058 | 117.9096 | 234.6707 | 273.6451 | 326.9164 | 360.1741 | 228.8772 | 342.8003 | 308.8003 | 399.9096 | 160.0904 | 71.2645 | 75.471 | 160.0904 | 112.7935 | 133.587 | 162.2645 | 122.5614 | 116.0324 | 114.0256 | 133.9164 | 114.6127 | 79.116 | 67.413 | 71.6194 | 105.058 | 118.2389 | 110.6775 | 117.2901 | 69.884 | 242.1809 | 230.1228 | 297.7355 | 382.116 | 153.5614 | 206.7611 | 226.6519 | 305.7611 | 296.9096 | 241.5358 | 274.8327 | 358.3549
1488 | 18.3624 | 232.3663 | 312.7409 | 274.232 | 201.5536 | 151.2944 | 214.6961 | 200.9729 | 147.9376 | 144.616 | 163.4817 | 200.8127 | 154.7409 | 209.7056 | 297.8304 | 217.6703 | 242.5265 | 178.9729 | 168.6337 | 197.6608 | 183.1248 | 146.0177 | 337.4016 | 339.7856 | 162 | 224.7856 | 266.3392 | 292.8033 | 328 | 104.5984 | 120.9376 | 138.768 | 128.8928 | 133.768 | 145.384 | 171.0624 | 77.2768 | 234.0542 | 314.4288 | 321.384 | 358.1601 | 113.9823 | 163.152 | 168.5889 | 141.5984 | 80.8928 | 89.4193 | 80.2497 | 117.232 | 195.6337 | 221.5712 | 280.7327 | 156.7232 | 129.8928 | 148.5807 | 170.1248 | 95.3216 | 166.152 | 232.7856 | 262.6242 | 138.0624 | 141.4912 | 179.6337 | 181.8127 | 143.9105 | 319.4641 | 213.5984 | 241.5183 | 245.384 | 144.9376 | 164.8127 | 291.1873 | 197.2768 | 199.3663 | 227.8304 | 240.1248 | 310.0624 | 158.848 | 151.5631 | 193.848 | 215.152 | 139.0624 | 154.2944 | 167.2768 | 149.6337 | 194.6337 | 169.848 | 227.9199 | 123.1072 | 112.4641 | 106 | 127.4912 | 117.4193 | 259.5984 | 267.1601 | 305.0177 | 363.4641 | 255.4735 | 321.9376 | 287.1072 | 392.4994 | 147.2944 | 73.384 | 72.616 | 147.2944 | 111.6784 | 130.2497 | 152.2049 | 115.2144 | 106.384 | 122.4641 | 115.8928 | 124.6961 | 80.2944 | 64.384 | 72.7503 | 104.4016 | 96.6961 | 106.1696 | 128.1248 | 70.6431 | 223.5088 | 216.2591 | 286.2049 | 378.8657 | 144.5536 | 220.8304 | 218.6879 | 314.5712 | 300.5712 | 225.1967 | 237.6784 | 349.616
1489 | 18.3734 | 241.1946 | 303.9145 | 277.1181 | 198.3746 | 144.7818 | 189.82 | 188.3836 | 155.3127 | 146.2036 | 140.0472 | 216.8673 | 147.0709 | 210.7256 | 289.5928 | 260.3363 | 223.5782 | 170.4308 | 155.7581 | 199.6873 | 173.2891 | 139.0709 | 322.2182 | 336.3127 | 159.1091 | 217.8436 | 254.1946 | 276.8909 | 323.3746 | 114.8673 | 122.5309 | 138.1946 | 127.4454 | 140.1327 | 142.351 | 168.8909 | 78.8673 | 265.3037 | 307.7874 | 317.1946 | 313.0383 | 125.2655 | 161.5692 | 149.2036 | 142.0383 | 75.3982 | 65.2655 | 67.5163 | 110.8673 | 185.0709 | 215.7818 | 304.4601 | 155.4218 | 129.8909 | 162.6164 | 160.8673 | 106.6018 | 179.6018 | 226.7109 | 213.7345 | 141.6726 | 144.82 | 164.4454 | 196.7109 | 128.8673 | 298.6401 | 217.5073 | 257.4365 | 248.1327 | 153.4691 | 178.5546 | 276.2891 | 197.4218 | 215.4218 | 222.7728 | 239.2508 | 310.2036 | 150.3127 | 171.6783 | 187.0472 | 223.1091 | 132.5546 | 137.9528 | 171.7581 | 134.7345 | 171.6401 | 165.9381 | 237.0383 | 134.6726 | 100.0236 | 111.2036 | 130.5309 | 101.0709 | 256.5692 | 244.009 | 296.3363 | 359.4072 | 267.1946 | 332.2036 | 285.9528 | 352.8673 | 136.7345 | 74.1091 | 71.8909 | 136.7345 | 107.6254 | 118.3836 | 124.5782 | 116.0855 | 110.8673 | 107.7109 | 118.7818 | 103.5782 | 72.9145 | 68.2891 | 94.7345 | 95 | 81.0709 | 98.5073 | 109.9055 | 85.1091 | 228.5635 | 227.3982 | 261.18 | 347.1946 | 143.9764 | 220.1091 | 247.4691 | 298.6637 | 281.1946 | 245.7345 | 231.8909 | 344.5546
1490 | 18.3844 | 223.2055 | 314.1917 | 289.3421 | 192.1234 | 155.0821 | 182.3014 | 203.8766 | 159.1917 | 157.1917 | 149.0959 | 212.0683 | 165.3152 | 190.8359 | 265.2055 | 254.7255 | 217.1641 | 194.9724 | 161.6848 | 194.9041 | 169.1779 | 152.7945 | 300.4655 | 357.6717 | 157.1234 | 227.9593 | 241.9317 | 293.9455 | 305.9724 | 112.7807 | 118.5207 | 122.6166 | 132.8083 | 132.2738 | 127.8083 | 167.26 | 87.6303 | 245.6303 | 265.1097 | 297.3972 | 289.1917 | 121.3014 | 133.411 | 168.329 | 121.2055 | 78.6441 | 65.2193 | 78.7669 | 109.9862 | 196.6848 | 226.6848 | 270.4517 | 150.9862 | 130.3697 | 149.2331 | 155.4655 | 116.0138 | 183.8903 | 216.3559 | 229.7538 | 159.0821 | 140.6166 | 168.3014 | 189.9724 | 131.9041 | 273.0545 | 201.589 | 238.9455 | 241.7807 | 156.8766 | 168.9724 | 280.0138 | 196 | 213.0959 | 203.1917 | 255.3697 | 322.3972 | 161.4248 | 145.4524 | 197 | 208.7669 | 129.6028 | 123.7807 | 177.9862 | 141.7124 | 158.1097 | 177.2738 | 210.1779 | 144.5476 | 100.4793 | 114.9586 | 123.5069 | 116 | 221.1779 | 260.6297 | 308.5069 | 366.3145 | 242.5752 | 333.849 | 278.7124 | 363.4386 | 140.0959 | 68.8083 | 74.1779 | 140.0959 | 115.6848 | 133.8766 | 129.6166 | 103.0959 | 112.6986 | 105.4931 | 129.3834 | 102.8903 | 76.26 | 72.0138 | 91.7669 | 95 | 92.3834 | 91.329 | 93.0138 | 79.5069 | 241.6297 | 216.4793 | 268.0959 | 331.9179 | 136.8903 | 204.26 | 248.4655 | 293.7262 | 271.9455 | 247.2876 | 248.3428 | 331.9586
1491 | 18.3954 | 229.764 | 304.9813 | 254.3539 | 202.7453 | 145.3727 | 181.4316 | 194.823 | 161.5871 | 150.4906 | 144.8043 | 201.0187 | 172.059 | 212.1367 | 254.8258 | 211.0403 | 203.0777 | 181 | 154 | 181.9438 | 165.2144 | 156.2734 | 283 | 388.9813 | 166.177 | 240.4345 | 246.0805 | 310.0993 | 287.2949 | 104.882 | 117.3942 | 134.1367 | 134.4906 | 127.1554 | 130.7453 | 159.8605 | 97.0403 | 252.8446 | 271.5684 | 311.3698 | 287.1957 | 121.059 | 138.059 | 179.8633 | 119.2949 | 92.8446 | 77.1957 | 66.7051 | 104.5684 | 193.3914 | 208.9625 | 260.1152 | 147.764 | 118.764 | 163.177 | 144.2734 | 122.6863 | 182.4504 | 201.5281 | 251.9813 | 149.3727 | 140.8446 | 159.2762 | 173.8043 | 136.1957 | 239.9625 | 198.1957 | 262.941 | 234.5094 | 153.1554 | 152.177 | 277.5899 | 196.3137 | 202.5309 | 188.3352 | 245.9597 | 321.6273 | 164.2762 | 168.5496 | 189.1582 | 202.6648 | 135.0777 | 118.7051 | 165.0403 | 146.7453 | 161.7453 | 166.1957 | 209.3511 | 126.1554 | 100.9785 | 99.9597 | 130.4504 | 112.236 | 217.8418 | 235.9785 | 306.6676 | 338.9972 | 231.882 | 300.1582 | 284.059 | 377.0403 | 143.3324 | 73.6273 | 66.764 | 143.3324 | 107.6863 | 122.3137 | 133.2949 | 110.0964 | 111.3727 | 102.3137 | 122.2547 | 103.3324 | 60.3914 | 69.9036 | 74.6863 | 94.6863 | 94.3511 | 99.7266 | 106.2734 | 81.118 | 213.2144 | 200.6676 | 266.3137 | 357.3109 | 137.3324 | 184.941 | 229.118 | 299.7856 | 288.4129 | 235.3539 | 270.3137 | 318.2144
1492 | 18.4064 | 235.3923 | 294.9779 | 250.149 | 196.7404 | 149.2154 | 170.5029 | 210.1269 | 168.5029 | 146.9558 | 139.674 | 209.7183 | 157.9394 | 210.199 | 218.8288 | 206.149 | 210.6962 | 175.4587 | 154 | 156.8231 | 179.6962 | 159.8952 | 296.0385 | 365.6135 | 163.2596 | 203.9558 | 224.5635 | 297.2817 | 274.0663 | 106.3923 | 101.1269 | 139.3702 | 136.1712 | 141.0885 | 130.6077 | 178.8731 | 84.6519 | 235.9779 | 264.2433 | 340.4365 | 284.8288 | 122.2596 | 132.4144 | 167.7846 | 106.3923 | 68.4808 | 78.4144 | 77.6519 | 107.0221 | 187.3539 | 192.8673 | 282.199 | 151.7625 | 128.6298 | 149.8288 | 156.3702 | 120.3702 | 178.2654 | 206.5914 | 223.0721 | 148.326 | 129.8452 | 133.4365 | 176.8231 | 135.1327 | 209.851 | 195.5029 | 258.8067 | 241.6298 | 161.8731 | 151.2154 | 265.801 | 198.3038 | 179.7625 | 179.6683 | 239.1164 | 322.674 | 145.9337 | 162.3038 | 177.2154 | 216.2212 | 140.7404 | 129.6519 | 155.5856 | 147.5856 | 170.4086 | 169.6962 | 225.1548 | 137.1548 | 112.9394 | 112.0221 | 133.1106 | 109.2154 | 239.2375 | 255.7625 | 299.2375 | 354.1164 | 238.6077 | 287.8894 | 295.0385 | 349.0058 | 142.9394 | 73.6962 | 76.6298 | 142.9394 | 105.3702 | 122.3481 | 125.9337 | 131.6298 | 104.4587 | 93.2212 | 120.4365 | 113.3702 | 75.2154 | 57.8231 | 76.6077 | 98.2375 | 108.525 | 91.8673 | 112.5029 | 84.4971 | 232.2596 | 196.1712 | 260.8067 | 371.8125 | 149.326 | 179.5029 | 221.4144 | 283.3481 | 287.6904 | 226.5856 | 257.3096 | 353.8836
1493 | 18.4174 | 231.353 | 288.5437 | 256.1759 | 186.2792 | 152.5584 | 163.1181 | 219.8672 | 149.6175 | 153.0295 | 133.588 | 201.2792 | 136.0443 | 198.765 | 204 | 224.0148 | 218.4858 | 155.8819 | 156.3678 | 166.5732 | 171.2497 | 137.3087 | 320.6322 | 349.4711 | 170 | 215.4416 | 230.9557 | 293.2054 | 279.8388 | 106.412 | 118.647 | 132.9557 | 139.4994 | 128.6027 | 136 | 162.0295 | 87.353 | 239.6913 | 242.3383 | 327.9557 | 270.3383 | 122.5732 | 123.9852 | 159.706 | 107.0885 | 60.0738 | 75.059 | 80.0148 | 101.9557 | 160.7503 | 208.7208 | 259.2645 | 142.6617 | 129.294 | 135.3383 | 154.0148 | 116.6617 | 159.412 | 184.353 | 202.5744 | 146.9705 | 150.4268 | 128.059 | 184.2054 | 122.6617 | 200.7798 | 178.647 | 242.9557 | 241.6175 | 150.1033 | 166.7355 | 276.5584 | 200.3235 | 180.8093 | 206.2792 | 203.0738 | 320.3087 | 149.6322 | 161.9557 | 185.6322 | 197.6913 | 137.3825 | 132.0148 | 165.0295 | 154.3235 | 178.4404 | 164.9705 | 209.9852 | 121.647 | 98.4858 | 112.0295 | 127.7355 | 113.9115 | 252.059 | 246.3235 | 306.3087 | 322.8093 | 243.6617 | 292.588 | 313.2054 | 315.294 | 126.1181 | 69.9852 | 77.9705 | 126.1181 | 105.0443 | 113.5584 | 129.294 | 133.647 | 93.3383 | 80.7798 | 116.412 | 112.0295 | 83.294 | 73.3235 | 83.6913 | 104.6322 | 85.9262 | 102.3087 | 92.941 | 92.2202 | 237.647 | 219.4563 | 248.3383 | 326.9852 | 147.9705 | 168.3973 | 220.765 | 296.5449 | 315.3235 | 226.5584 | 235.0885 | 371.6606
1494 | 18.4284 | 227.3397 | 271.2989 | 242.3614 | 179.7011 | 137.6495 | 176.1957 | 196.0408 | 144.7527 | 161.5571 | 125.8043 | 196.1033 | 151.4647 | 216.5571 | 204.3505 | 216.5353 | 234.4429 | 151.5679 | 161.3505 | 151.8967 | 164.5571 | 148.3723 | 313.8967 | 362.0408 | 164.7418 | 225.4429 | 229.9076 | 274.2473 | 302.6495 | 114.8967 | 114.5978 | 129.1033 | 123.6603 | 127.3614 | 139.5054 | 167.0516 | 89.6495 | 241.2473 | 250.0109 | 309.731 | 281.8668 | 113.856 | 130.7636 | 157.288 | 123.2065 | 74.9076 | 84.0516 | 80.9484 | 98.8043 | 184.269 | 212.144 | 250 | 150.7636 | 120.144 | 133.8967 | 149.6902 | 115.2989 | 166.8451 | 190.5054 | 222.0299 | 147.2065 | 151.0815 | 135.2989 | 178.2174 | 124.4538 | 205.1332 | 178.4538 | 233.4946 | 234.2989 | 156.6386 | 162.6793 | 263.4022 | 194.7935 | 200.3505 | 197.5978 | 215.1033 | 308.9375 | 148.1549 | 160.2065 | 183.8043 | 213.6196 | 137.3397 | 133.6495 | 161.288 | 150.5462 | 159.2174 | 168.712 | 208 | 124.2582 | 110.587 | 115.6495 | 129.9891 | 114.269 | 262.4538 | 241.8451 | 313.1658 | 335.6902 | 241.2473 | 284.4538 | 301.2582 | 323.3207 | 153.5679 | 70.8043 | 77.5054 | 153.5679 | 114.5054 | 101.8043 | 123.6495 | 127.144 | 102.413 | 101.9592 | 125.2473 | 112.144 | 75.5462 | 70.2473 | 86.2989 | 99.6495 | 80.2065 | 103.5571 | 98.3207 | 74.1957 | 237.4538 | 233.1957 | 237.7826 | 340.0734 | 148.9076 | 170.3397 | 234 | 325 | 304.1848 | 216.2065 | 245.9484 | 327.0516
1495 | 18.4394 | 217.9027 | 250.7698 | 257.0885 | 190.0709 | 137.3628 | 168.0973 | 186.8142 | 151.6283 | 164.5575 | 126.2832 | 188.7521 | 162.1061 | 212.3008 | 196.292 | 202.2655 | 232.8939 | 178.0797 | 152.9291 | 157.0709 | 174.8142 | 171.7345 | 303.1061 | 356.8054 | 148.469 | 220.9912 | 239 | 274.2655 | 303.8142 | 116.8054 | 117.4425 | 137.7168 | 126.5663 | 143.5398 | 145.2743 | 160.292 | 93.7168 | 233.2832 | 257.9203 | 290.2655 | 287.8495 | 123.5398 | 144.0973 | 139.7345 | 131.3628 | 78.1946 | 76.9291 | 77.5487 | 107.9912 | 192.4866 | 200.2832 | 252.177 | 151.3981 | 123.5222 | 132.5398 | 136.5487 | 120.1682 | 167.8939 | 187.5663 | 207.4337 | 155.7257 | 135.8939 | 132.1858 | 180.1326 | 130.0885 | 188.9912 | 184.8142 | 219.0176 | 230.0973 | 153.885 | 150.6992 | 267.0885 | 191.7168 | 196.646 | 199.7168 | 222.9912 | 295.0885 | 152.1858 | 168 | 189 | 228.6548 | 122.823 | 130.0973 | 147.3628 | 148.177 | 161.1326 | 181.5487 | 213.8054 | 130.3717 | 101.8939 | 110.646 | 118.0885 | 133.0088 | 256.8406 | 249.062 | 332.3628 | 317.1061 | 246.708 | 281.0176 | 298.3008 | 336.7521 | 160.1238 | 71.2832 | 77.1061 | 160.1238 | 116.2832 | 110.9912 | 121.5487 | 113.469 | 114.7345 | 114.8142 | 124.9027 | 110.0797 | 71 | 67 | 82.4602 | 98.6372 | 82.9203 | 102.9291 | 115.7433 | 71.177 | 232.9291 | 226.1858 | 226.0709 | 367.2743 | 150.0176 | 157.6372 | 222.3893 | 317.0176 | 283.0973 | 227.9912 | 249.0797 | 327.9115
1496 | 18.4504 | 209.4955 | 232.005 | 266.5025 | 201.4985 | 135.3741 | 164.8756 | 196.7523 | 160.6259 | 161.002 | 109.3721 | 168.6249 | 150.3751 | 196.002 | 188.1274 | 209.1254 | 231.1204 | 175.7462 | 141.1264 | 179.002 | 172.7482 | 158.1224 | 300.7533 | 371.5015 | 150.8796 | 227.2497 | 231.4975 | 286.7523 | 310.001 | 121.3731 | 118.7472 | 142.2487 | 114.5015 | 141.2477 | 144 | 154.3781 | 96.7482 | 229.1264 | 262.1294 | 315.1314 | 263.3751 | 129.8756 | 129.2467 | 143.003 | 124.8726 | 64.6239 | 66.6269 | 80.2518 | 114.6249 | 177.8786 | 190.4995 | 244.3711 | 134.8786 | 134.8726 | 140.7513 | 133.2497 | 120.1214 | 175.4985 | 162.7472 | 222.8746 | 158.5005 | 146.4995 | 129 | 149.2508 | 133.1254 | 197.1254 | 190.6259 | 213.2528 | 228.3761 | 160.6209 | 161.7462 | 266.7492 | 196.6239 | 193.8766 | 209.8756 | 228.8746 | 296.6249 | 152.3761 | 160.8726 | 186.3741 | 210.6239 | 126.5025 | 129.5015 | 143.1234 | 150.1244 | 145.2558 | 175.999 | 209.7452 | 129.2518 | 109.4985 | 109.002 | 121.1254 | 122.2487 | 233.3731 | 261.1214 | 315.3691 | 317.7543 | 261.6249 | 263.999 | 275.6249 | 314.3741 | 141.5025 | 67.1264 | 68.7513 | 141.5025 | 105.7492 | 113.8736 | 121.2508 | 108.7523 | 102.2477 | 114.999 | 123.998 | 112.2477 | 72.1254 | 66.6249 | 79.5005 | 87.1214 | 77.001 | 101.2548 | 102.6249 | 78.001 | 219.2508 | 216.2477 | 245.3761 | 363.3741 | 134.1244 | 165.6279 | 211.7533 | 319.8807 | 293.3802 | 233.1244 | 242.2447 | 332.002
1497 | 18.4614 | 185.8377 | 241.5033 | 279 | 188.5762 | 141.0728 | 157.1523 | 189.7914 | 156.4768 | 161.3146 | 103.1358 | 157.9272 | 154.4868 | 206.7748 | 185.2782 | 212.5497 | 199.8642 | 168.2351 | 142.5762 | 172.3411 | 175.6225 | 147.3245 | 316.2252 | 382.0993 | 170.5132 | 222.9007 | 210.0894 | 291.8013 | 296.7914 | 112.7748 | 118.947 | 127.0894 | 125.8742 | 124.5762 | 135.0894 | 166.1258 | 80.2517 | 233.2881 | 266.5662 | 335.3775 | 270.1987 | 129.5132 | 118.0364 | 159.1623 | 125.7848 | 64.4238 | 67.4768 | 83.9636 | 107.8013 | 201.8113 | 195.7483 | 232.7483 | 161.9106 | 129.5861 | 136.5397 | 132 | 104.7119 | 179.2351 | 156.7483 | 227.1987 | 154.4139 | 140.5132 | 141.0099 | 163.4603 | 142.3609 | 189.3146 | 176.7914 | 226.2252 | 223.9272 | 141.5497 | 132.5397 | 263 | 189.4503 | 208.5861 | 206.4139 | 221.5762 | 305.6854 | 164.1987 | 146.2881 | 176.1887 | 209.2616 | 131.6391 | 139.3245 | 141.5861 | 137.702 | 168.5762 | 170.6126 | 188.3245 | 131.8013 | 102 | 112.0265 | 122.2252 | 115.6126 | 225.5497 | 237.5762 | 297.6225 | 339.3874 | 254.8013 | 246.9901 | 277.3245 | 317.7483 | 150.9007 | 77.4868 | 79.649 | 150.9007 | 106.649 | 104.2881 | 119.5762 | 120.3874 | 104.1722 | 98.7649 | 117.0993 | 107.5861 | 78.2616 | 72.5861 | 80.0629 | 82.5596 | 93.6225 | 119.1887 | 101.6126 | 71.7649 | 229.1987 | 211.1987 | 250.2252 | 360.1623 | 133.3245 | 193.947 | 214.4404 | 323.2053 | 304.2782 | 223.0265 | 217.5497 | 338.1258
1498 | 18.4724 | 183.6003 | 214.202 | 273.0044 | 182.798 | 148.6061 | 143.1977 | 175.3895 | 144.1991 | 147.9985 | 116.3997 | 151.9927 | 158.4012 | 202.4041 | 162.7994 | 201.8096 | 196.9927 | 168.4129 | 138.7965 | 152.795 | 196.798 | 147.4026 | 312.202 | 364.2167 | 153.0088 | 219.5988 | 205.593 | 274.4056 | 266.4012 | 116.3982 | 127.8096 | 112.6003 | 133.9985 | 124.3939 | 126.1962 | 156.8023 | 78.7921 | 219.8067 | 240.9985 | 303.8096 | 279.2006 | 125.9985 | 126.3997 | 151.0073 | 136.8067 | 79.3953 | 55.9985 | 71.6032 | 100.3982 | 207.8023 | 214.3953 | 252.1947 | 158.0132 | 134.0044 | 123.9971 | 139.1947 | 116.1947 | 175.0161 | 172.9971 | 229.8053 | 144.7994 | 135.7994 | 145.6105 | 172.4056 | 153.2006 | 168.8038 | 162.7892 | 235.3924 | 207.2006 | 144.7994 | 134.3866 | 264.9985 | 189.798 | 203.4114 | 194.0015 | 212.6003 | 315.0044 | 154.814 | 144.798 | 164.202 | 210.0044 | 127.9942 | 136.6047 | 152.3997 | 131.3939 | 158.8009 | 165.2035 | 188.0029 | 124.798 | 107.1962 | 106.1962 | 114.205 | 116.1991 | 225.202 | 233.3968 | 318.798 | 330.8067 | 253.7936 | 226.4012 | 281.7994 | 336.795 | 144.8009 | 78.2035 | 85.4012 | 144.8009 | 117.9971 | 96.0029 | 107.8023 | 124.1977 | 114.2079 | 99.3866 | 122 | 114.4026 | 83.8009 | 81.0015 | 78.9985 | 102.8009 | 108.0029 | 104.4041 | 104.1977 | 62.7936 | 235.4026 | 228.1947 | 255.3953 | 349.6091 | 142.1962 | 207.6061 | 200.9942 | 304.7848 | 290.593 | 208.0029 | 243.183 | 340.7936
1499 | 18.4834 | 176.82 | 205.056 | 254.524 | 186.176 | 140.532 | 152.532 | 185.052 | 143.116 | 144.408 | 128.948 | 164.94 | 153.528 | 191.528 | 158.644 | 190.24 | 208.292 | 163.4239 | 149.296 | 157.644 | 196.056 | 145.708 | 322.832 | 330.412 | 133.352 | 230.24 | 220 | 263 | 277.596 | 115.88 | 102.232 | 115.708 | 136.176 | 134.94 | 131.528 | 156.532 | 93.764 | 217.536 | 239.056 | 293.476 | 271.408 | 125.704 | 127 | 145.888 | 125.06 | 80.76 | 70.124 | 62.116 | 99.056 | 193.524 | 214.936 | 263 | 138.416 | 120.056 | 134.944 | 149.176 | 117.524 | 144.06 | 164.992 | 217.352 | 148.06 | 155.1279 | 133.064 | 156.056 | 142.112 | 165.944 | 178.82 | 232.0481 | 207.648 | 148.06 | 157.468 | 258.524 | 186.584 | 183.708 | 182.348 | 214.472 | 298.996 | 143.3039 | 142.408 | 154.644 | 201 | 131.76 | 124.116 | 151.352 | 152.24 | 162.356 | 157.176 | 189.416 | 131.06 | 113.764 | 114 | 106.472 | 110.584 | 225.12 | 249.888 | 319.292 | 319.06 | 258.76 | 238.42 | 283.824 | 353.18 | 140.528 | 72.236 | 76.408 | 140.528 | 115.76 | 101.948 | 104.236 | 133.532 | 102.944 | 115.052 | 118.292 | 111.472 | 74.996 | 78.412 | 80.352 | 98.94 | 100.352 | 97.236 | 99.936 | 77.236 | 228.764 | 230.348 | 262.94 | 343.772 | 136.404 | 188.82 | 209.292 | 332.704 | 284.8121 | 207.768 | 265.228 | 342.052
1500 | 18.4944 | 160.3628 | 208.6072 | 240.5757 | 183.3028 | 146.5757 | 150.5142 | 173.5158 | 143.2429 | 140.9401 | 137.0899 | 151.3928 | 151.6972 | 188 | 152.6972 | 199.03 | 209.3643 | 184.6656 | 157.3943 | 145.7571 | 180.9385 | 148.03 | 319.8155 | 351.3659 | 142.0315 | 229.6956 | 225.9401 | 268.0915 | 293.6057 | 115.0915 | 97.8186 | 128.2129 | 135.8486 | 134.5457 | 122.0599 | 157.4842 | 99.2129 | 237.3943 | 236.6672 | 304.8786 | 265.8186 | 125.6672 | 125.7271 | 163.8186 | 126.3028 | 66.03 | 74.5442 | 68.183 | 97.5158 | 174.4842 | 213.8801 | 256.2114 | 146.03 | 116.8186 | 139.4543 | 135.2713 | 111.6372 | 148.6972 | 143.7271 | 206.0899 | 152.6972 | 159.3612 | 136.6656 | 154.9401 | 117.817 | 168.7571 | 172.97 | 215.6072 | 210 | 146.7571 | 146.7571 | 250.0915 | 182.3643 | 182.2114 | 181.8801 | 215.03 | 286.8786 | 157.3928 | 138.9401 | 151.6672 | 203.5457 | 132.7287 | 127.2129 | 149.4243 | 144.907 | 169.5757 | 166.183 | 193.6357 | 125.9385 | 119.6372 | 112.7271 | 114.2429 | 109.7587 | 226.7571 | 255.9385 | 312.3028 | 313.0899 | 240.6357 | 258.1514 | 277.3628 | 349.2713 | 140.3943 | 68.4842 | 73.3643 | 140.3943 | 100.1814 | 116.03 | 112.3643 | 130.2413 | 107.8786 | 95.8786 | 118.9401 | 110.7571 | 64.1514 | 81.5457 | 75.8786 | 96.4243 | 98 | 103.2429 | 90.3943 | 79.8486 | 227.4243 | 214.6057 | 251.0899 | 350.5442 | 120.8186 | 171.5142 | 203.1514 | 335.2129 | 271.2744 | 206.2413 | 218.1483 | 327.5457
1501 | 18.5054 | 161.3511 | 211.2123 | 239.1268 | 183.6195 | 142.9439 | 145.1121 | 177.0708 | 142.8879 | 147.6903 | 125 | 140.9292 | 150.9439 | 194.1121 | 168.0974 | 199.3658 | 209.705 | 165.0147 | 170.295 | 151.3511 | 172.6195 | 153.1682 | 306.6461 | 354.9881 | 166.1682 | 215.1829 | 222.6489 | 261.0294 | 276.3392 | 110.6489 | 109.5487 | 136.2537 | 126.9586 | 132.761 | 121.8585 | 149.5634 | 95.0294 | 227.5928 | 234.705 | 293.2684 | 277.9853 | 120.6489 | 126.1829 | 157.6489 | 127.4926 | 64.6195 | 62.9292 | 81.1268 | 108.0561 | 187.9559 | 216.9026 | 260.9706 | 147.239 | 120.6903 | 143.4219 | 132.4072 | 123.7463 | 157.5487 | 153.3511 | 206.6608 | 150.6342 | 136.1121 | 124 | 166.9292 | 118.4072 | 167.3658 | 172.6342 | 230 | 206.9439 | 140.1268 | 138.8171 | 238.227 | 186.1976 | 182.1682 | 196.6903 | 215.3658 | 296.2243 | 143.4366 | 154.4219 | 167.1682 | 210.9292 | 148.2537 | 140.9292 | 151.3097 | 128.1682 | 165.0708 | 166.9026 | 184.1268 | 124.6048 | 124.761 | 112.7463 | 123.0561 | 117.761 | 231.0414 | 241.5074 | 307.3805 | 314.9706 | 225.8879 | 258.3097 | 269.1829 | 320.6489 | 139.3244 | 66.239 | 85.0561 | 139.3244 | 100.239 | 106.3244 | 115.761 | 109.5781 | 105.8732 | 93 | 109.1002 | 104.5634 | 66.4926 | 80.1976 | 71.6903 | 97.8732 | 101.4926 | 97.6489 | 95.4366 | 80.1268 | 226.2537 | 206.5074 | 235.0708 | 347.6608 | 127.3097 | 173.0974 | 214.2243 | 315.3125 | 292 | 199.5487 | 205.6314 | 346.7169
1502 | 18.5164 | 160.2895 | 195.1023 | 246.9774 | 190.1421 | 137.6534 | 154.7955 | 184.1195 | 148.4662 | 141.9602 | 129.9376 | 138.3692 | 150.591 | 197.9602 | 170.1872 | 211.2842 | 195.8579 | 151.0624 | 165.2895 | 157.4714 | 176.4489 | 161.8579 | 318.9827 | 349.1368 | 162.7384 | 224.2842 | 214.2842 | 260.506 | 262.2444 | 103.182 | 116.6534 | 133.1023 | 117.1421 | 122.409 | 136.4887 | 149.8977 | 78.3466 | 214.8353 | 232.5286 | 294.5684 | 288.7955 | 110.9376 | 121.3692 | 145.2444 | 128.8579 | 71.591 | 67.5511 | 80.8977 | 98.085 | 204.2271 | 199.1023 | 264.6361 | 151.7556 | 131.5684 | 149.8579 | 151 | 119.2669 | 171.3865 | 168 | 213.1248 | 154.8752 | 131.8805 | 121.3068 | 171.1023 | 133.8579 | 169.8579 | 173.2842 | 226.8579 | 204.7955 | 133.1647 | 145.8752 | 221.6308 | 184.0398 | 181.8805 | 188.7158 | 227.7331 | 320.0797 | 143.5511 | 161.7556 | 180.7955 | 208.818 | 134.3293 | 149.591 | 149.8579 | 143.1421 | 164.4887 | 157.6308 | 190.182 | 123.085 | 117.1023 | 111.8579 | 127.8977 | 115.9376 | 230.5338 | 224.8805 | 291.4316 | 321.3293 | 225.182 | 246.9827 | 255.3917 | 310.9376 | 131.1023 | 71.2045 | 83.1647 | 131.1023 | 115.0797 | 103.4887 | 112.1421 | 109.8752 | 104.3068 | 96.591 | 90.4887 | 103.5511 | 67.409 | 76.2444 | 72.2444 | 96.7556 | 97.4714 | 83.4489 | 98.2444 | 79 | 230.7331 | 206.4887 | 231.7955 | 367.591 | 133.4887 | 180.5737 | 221.0226 | 297.6481 | 295.591 | 206.2045 | 230.5737 | 354.8805
1503 | 18.5274 | 148.8398 | 198.6116 | 259.3835 | 182.4709 | 146.6067 | 158.3835 | 190.6991 | 152.5486 | 142.0728 | 127.2379 | 145.6019 | 142.0874 | 202.6844 | 152.4612 | 219.4612 | 184.6214 | 157.4514 | 148.3058 | 151.1505 | 171.0049 | 152.466 | 297.9272 | 367.5486 | 142.6165 | 225.5437 | 214.6214 | 270.0098 | 284.83 | 123.5291 | 121.9174 | 138.4563 | 122.3835 | 121.6893 | 136.9272 | 143.6214 | 79.0777 | 224.767 | 234.2379 | 300.3156 | 270.2477 | 119.3058 | 116.1505 | 149.3835 | 114.8544 | 78.767 | 67 | 79.6942 | 88.3786 | 185.9272 | 197.5388 | 244.233 | 157.7621 | 144.6942 | 140.466 | 148.233 | 111.9223 | 171.5437 | 168.9223 | 202.3835 | 159.1602 | 112.8495 | 132.7572 | 168.6165 | 135.0728 | 172.9174 | 169.9321 | 220.6942 | 211.6116 | 127.8447 | 152.9272 | 243.9126 | 194.0728 | 172.0728 | 181 | 227.6991 | 314.3205 | 153.6067 | 164.534 | 183.4612 | 199.0777 | 123.6893 | 142.9321 | 149.6893 | 147.4612 | 160.3156 | 176.2233 | 187.4709 | 123.063 | 112.3107 | 101.0826 | 127.6165 | 117.3884 | 226.4514 | 221.5291 | 289.3737 | 308.8447 | 239.0728 | 229.6165 | 253.4465 | 319.767 | 132.767 | 67.1553 | 77.3835 | 132.767 | 116.6991 | 98.3933 | 114.1553 | 115.0826 | 103.3058 | 106.0728 | 99.6893 | 108.0728 | 70.8398 | 67.0098 | 77.3058 | 97.6893 | 92.0728 | 86.767 | 100.5388 | 81.3058 | 242.6893 | 203.2379 | 235.8447 | 363.6991 | 141.3058 | 173.1505 | 213.6893 | 317.466 | 289.8544 | 223.3688 | 223.6116 | 339.5388
1504 | 18.5384 | 147.584 | 199.2655 | 243.0087 | 174.208 | 155.6858 | 161.8938 | 185.2611 | 149.8364 | 139.9513 | 117.9469 | 156.0044 | 144.9913 | 192.6946 | 142.1106 | 196.8008 | 185.4691 | 174.8407 | 133.0087 | 152.6858 | 168.7345 | 154.0487 | 296.7345 | 342.1106 | 142.4204 | 200.9557 | 199.8451 | 258.8407 | 292.8495 | 129.4247 | 126.2124 | 145.5265 | 132.5222 | 130.7345 | 128.6327 | 140.6814 | 78.4735 | 221.8451 | 232.0487 | 304.624 | 255.4691 | 124.8407 | 114.3716 | 151.4735 | 109.1549 | 79.6327 | 59.8982 | 74.1593 | 98.4204 | 178.5796 | 205.9956 | 237.6858 | 158.0575 | 128.2698 | 133.0531 | 141.6858 | 115.3673 | 158.3186 | 168.1062 | 204.9469 | 154.7345 | 119.8364 | 147.6327 | 160.3716 | 140.0531 | 169.6371 | 174.5178 | 206.6371 | 219.3673 | 142.7833 | 141.792 | 242.9026 | 201.8938 | 168.5309 | 188.1018 | 226.5222 | 289.8451 | 157.4778 | 170.0531 | 178.3186 | 207.9425 | 122.3186 | 142.3098 | 149.2655 | 141.8451 | 164.1505 | 174.6902 | 176.3673 | 133.0575 | 108.4735 | 100.1018 | 122.2124 | 119.1018 | 240.0531 | 236.4204 | 284.168 | 315.7345 | 245.9469 | 232.2611 | 287.3098 | 313.5309 | 134.5796 | 69.7345 | 79.9469 | 134.5796 | 107.4735 | 94.5222 | 118.1549 | 119.3098 | 112.1549 | 110.1062 | 101.1593 | 117.3142 | 75.7389 | 64.8364 | 74.792 | 93.0044 | 105.5753 | 94.7345 | 100 | 84.4735 | 246.5265 | 215.2523 | 248.416 | 345.4778 | 139.2655 | 175.6327 | 206.6371 | 311.4735 | 276.1062 | 225.168 | 236.5753 | 337.1062
1505 | 18.5494 | 146.6864 | 193.5143 | 235.1436 | 189.7149 | 148.114 | 167.886 | 191.943 | 153.7708 | 144.1151 | 122.4002 | 134.8279 | 161.114 | 178.2292 | 146.0296 | 171 | 194.057 | 168.2851 | 121.7434 | 141.2281 | 173.0285 | 152.3136 | 307.3432 | 315.4276 | 146.4287 | 208.2302 | 194.4572 | 252.2851 | 296.887 | 113.2281 | 107.9134 | 136.7423 | 142.9715 | 136.4857 | 130.8575 | 150.886 | 83.8575 | 201.3991 | 229.8279 | 301.113 | 251.9134 | 121.6853 | 113.8004 | 147.1425 | 116.4857 | 76.5143 | 63.1721 | 72.943 | 104.8575 | 173.5998 | 209.1996 | 223.3136 | 156.7149 | 114.943 | 133.4572 | 147.7149 | 118.4857 | 153.4572 | 156.2851 | 198.2281 | 154.1711 | 141.7434 | 144.5143 | 150.0855 | 134.6283 | 163.3147 | 183.3136 | 204.2006 | 218.114 | 146.7994 | 136.9715 | 247.3443 | 196.7138 | 177.9155 | 178.9989 | 218.9989 | 290.2862 | 143.1425 | 165.114 | 172.9715 | 220.4572 | 118.9145 | 139.4847 | 136.7138 | 138.886 | 170.7994 | 164.886 | 171.2281 | 122.9715 | 112.4002 | 116.7434 | 127.2292 | 119.2281 | 228.3136 | 236.5428 | 271.2006 | 323.9145 | 257.2006 | 239.9145 | 304.4002 | 318.0581 | 122.7994 | 68.6853 | 76.6283 | 122.7994 | 106.057 | 100.5998 | 117.7138 | 128.1425 | 113.6568 | 104.5998 | 100.4287 | 115.6568 | 75.8575 | 77.0285 | 68.0285 | 90.2862 | 109.6283 | 87.3706 | 94.6568 | 83.057 | 239.6853 | 221.5417 | 248.3421 | 345.7149 | 136.4287 | 181.7434 | 189.6283 | 321.2292 | 273.5143 | 221.4298 | 232.3706 | 321.3991
1506 | 18.5604 | 159.4941 | 190.5098 | 240.0314 | 198.0079 | 136.5295 | 168.0157 | 182.0471 | 150.4745 | 150.0314 | 125.0039 | 137.9411 | 154.5098 | 187.0039 | 151.5452 | 176.4784 | 203.9529 | 167.4627 | 135.4823 | 139.9725 | 182.4588 | 143.9882 | 317.9804 | 325.3998 | 155.9725 | 219.0668 | 205.4627 | 253.9529 | 308.5412 | 106.4941 | 97.9804 | 131.4863 | 136.5295 | 128.5334 | 137.9921 | 163.4666 | 89.9961 | 231.8233 | 215.0118 | 268.0432 | 250.447 | 117.9882 | 116.5177 | 133.5334 | 114.5098 | 76.498 | 73.5059 | 75.498 | 112.4902 | 178.4666 | 192.5373 | 218.9725 | 151.0628 | 119.9882 | 128.5255 | 152.0236 | 127.9647 | 152.0118 | 158.4509 | 196.9725 | 143.5177 | 126.0982 | 137.5255 | 147.498 | 121.5334 | 166.0157 | 176.4823 | 205.5373 | 217.9843 | 146.4588 | 142.4823 | 232.1532 | 189.9961 | 185.055 | 173.4509 | 218.9293 | 298.9961 | 151.447 | 160.0039 | 157.5648 | 220.5059 | 128.4745 | 138.9332 | 138.4627 | 143.498 | 149.5412 | 165.5137 | 171.9647 | 118.5216 | 114.0079 | 118.5295 | 124.0511 | 123.9489 | 225.9647 | 231.0157 | 276.5216 | 330.4863 | 247.5805 | 221.0864 | 291.5648 | 317.0668 | 126.9411 | 63.9921 | 76.9804 | 126.9411 | 106.4902 | 100.9843 | 112.4902 | 123.498 | 107.0039 | 106.9764 | 103.9961 | 121.4548 | 67.055 | 75.502 | 75.9647 | 95.0118 | 102.5098 | 89.4391 | 89.9961 | 71.5373 | 233 | 207.9764 | 241.9764 | 361.4784 | 143.4823 | 195.4823 | 207.9097 | 333.4902 | 268.5177 | 226.0628 | 224.9764 | 336.3841
1507 | 18.5714 | 159.9793 | 193.1033 | 230.469 | 183.7521 | 144.3098 | 149.7108 | 172.5516 | 146.2831 | 132.8141 | 121.4484 | 137.6902 | 149.4484 | 183.9587 | 146.1239 | 170.7728 | 212.938 | 172.9174 | 135.4071 | 153.1239 | 192.4897 | 163.3305 | 316.8761 | 355.5929 | 166.5723 | 218.8408 | 213.9793 | 256.3038 | 287.2831 | 103.9174 | 109.1239 | 129.8967 | 144.3098 | 129.6962 | 135.9174 | 153.1179 | 96.1033 | 246.3804 | 224.2479 | 270.7788 | 276.2479 | 132.7375 | 117.6136 | 133.1652 | 119.6549 | 73.4277 | 69.4484 | 70.8967 | 109.8967 | 202.8201 | 180.9587 | 206.6076 | 162.0473 | 124.531 | 121.4897 | 140.3864 | 121.6902 | 146.4484 | 173.0413 | 195.8348 | 154.8201 | 118.3511 | 134.5723 | 154.1239 | 117.5929 | 155.3657 | 176.4071 | 210.7995 | 217.9174 | 166.6962 | 143.938 | 212.9028 | 193.5516 | 173.5516 | 187.531 | 219.6489 | 312.7582 | 162.4484 | 147.7728 | 161.9234 | 211.8554 | 135 | 140.6902 | 136.7728 | 139.9174 | 162.9854 | 153.3244 | 193.2479 | 118.6136 | 107.4071 | 106.9174 | 116.1033 | 124.2418 | 229.8967 | 235.6756 | 272.0207 | 330.938 | 217.3038 | 217.3718 | 267.3451 | 297.9587 | 132.8141 | 60.8967 | 75.8761 | 132.8141 | 110.531 | 96.8348 | 118.5723 | 115.8348 | 100.8967 | 104.8348 | 106.531 | 117.1799 | 51.469 | 69.3864 | 70.7108 | 89.4484 | 93.3657 | 98.3657 | 92.531 | 73.7375 | 228.4071 | 215.0207 | 237.2831 | 364.9587 | 140.8554 | 187.2418 | 204.463 | 321.2005 | 267.5723 | 214.0826 | 234.5723 | 338.4423
1508 | 18.5824 | 165.2714 | 183.8893 | 228.4774 | 165.206 | 148.5476 | 150.6333 | 175 | 135.4774 | 127.1357 | 117.4321 | 124.5679 | 149.0905 | 194.5428 | 148.8643 | 158.9548 | 195.8893 | 162.206 | 141.975 | 154.8191 | 198.794 | 160.7084 | 317.794 | 347.4572 | 163.206 | 224.5476 | 230.769 | 267.9047 | 274.3869 | 118.8143 | 112.9095 | 125 | 148.025 | 141.6131 | 133.5679 | 147.5428 | 86.8893 | 241.5631 | 230.2512 | 291.8393 | 271.2762 | 133.025 | 136.0655 | 130.025 | 117.5929 | 65.2964 | 69.6131 | 75.9297 | 114.4071 | 200.7536 | 206.1761 | 207.3369 | 154.0298 | 128.6131 | 123.0905 | 131.3417 | 112.2262 | 148.7036 | 167.1607 | 180.6833 | 155.3667 | 136.5679 | 138.5226 | 150.5929 | 120.9548 | 152.6583 | 177.7488 | 220.8191 | 218.7036 | 165.025 | 155.1107 | 225.7286 | 194.9548 | 189.0655 | 175.9345 | 201.4321 | 308.7988 | 164.1809 | 146.9297 | 176.7738 | 194.5476 | 131.3417 | 128.0905 | 132.2714 | 146.975 | 173.9798 | 148.6583 | 187.231 | 127.1357 | 97.7738 | 94.1155 | 110.5476 | 114.0905 | 219.2512 | 241.3869 | 311.1511 | 323.8191 | 222.1107 | 214.6179 | 272.5428 | 305.4071 | 120.8643 | 65.9297 | 75.2262 | 120.8643 | 114.6131 | 102.5881 | 119.3869 | 111.1357 | 108.0202 | 95.9548 | 112.7036 | 111.4071 | 60.4524 | 65.0452 | 70.0655 | 83.3417 | 92.7488 | 91.4774 | 99.2262 | 76.6381 | 225.0452 | 206.5929 | 213.9345 | 352.025 | 134.5226 | 171.8643 | 203.0404 | 307 | 275.7036 | 212.2512 | 238.5226 | 334.4726
1509 | 18.5934 | 167.7905 | 165.6509 | 237.6284 | 178.2567 | 140.6047 | 161.8604 | 181.9538 | 145.1633 | 134.2793 | 114.3953 | 128.6745 | 168.652 | 204.3953 | 143.3255 | 166.5585 | 187.2793 | 156.5349 | 146.6509 | 154.7444 | 188.9527 | 146.1396 | 315.9764 | 336.5349 | 160.7444 | 211.7905 | 227.2084 | 244.7422 | 275.7444 | 118.8829 | 115.2793 | 126.0698 | 137.4651 | 138.6509 | 126.9764 | 156.3255 | 78.2793 | 246.8829 | 212.1622 | 284.5574 | 262.4189 | 123 | 138.3716 | 122.6745 | 112.2095 | 68.4887 | 70.3953 | 75.9065 | 108.5574 | 198.9076 | 221.3716 | 224.4651 | 136.3727 | 126.1858 | 114.3018 | 132.2793 | 110.5811 | 155.1396 | 154.116 | 193.2567 | 146.8142 | 129.4415 | 135.7905 | 156.4426 | 129.6284 | 161.884 | 173.3716 | 198.2782 | 222.4651 | 148.0462 | 172.2793 | 234.4426 | 170.9989 | 201 | 177.9076 | 204.2793 | 300.9538 | 169.0698 | 157.0698 | 164.5113 | 210.1407 | 128 | 130 | 134.2556 | 141.4876 | 154.4415 | 165.9076 | 173.1396 | 136.4189 | 98.3491 | 90.8142 | 97.7905 | 109.5811 | 223.0935 | 239.5349 | 313.9054 | 311.4415 | 223.232 | 216.8615 | 275.4415 | 301.1622 | 128.6982 | 73.9302 | 63.4178 | 128.6982 | 115.9302 | 106.9764 | 119.6745 | 109.1858 | 108.8367 | 99.8142 | 103.0924 | 113.0462 | 70 | 68.6745 | 87.884 | 90.6982 | 87.8367 | 104.3727 | 95.4415 | 77.0236 | 223.8604 | 199.6047 | 216.9775 | 349.4887 | 132.8604 | 174.884 | 206.8131 | 296.3018 | 276.7905 | 207.5349 | 251.8378 | 345.3716
1510 | 18.6044 | 169.3776 | 171.944 | 237.792 | 173.112 | 154.5856 | 153.9808 | 172.6784 | 156.736 | 135.264 | 113 | 124.9808 | 170.8672 | 190.9616 | 146.472 | 171.8112 | 198.6608 | 156.4528 | 157.8688 | 144.8672 | 173.9056 | 154.0192 | 298.056 | 326.6032 | 156.8864 | 211.1888 | 223.1328 | 236.4544 | 266.9616 | 112.7552 | 104.7728 | 131.3776 | 140.2832 | 135.6416 | 116.1696 | 158.9248 | 79.8112 | 216.6784 | 205.9248 | 271.4528 | 264.7168 | 127.9248 | 133.8304 | 137.0384 | 115 | 79.9248 | 68.4528 | 70.736 | 91.6224 | 199.5088 | 192.2064 | 245.3408 | 148.5472 | 128.0192 | 111.0192 | 131.6224 | 114.8496 | 160.8304 | 155.5664 | 208.3584 | 152.6416 | 121.4528 | 140.6608 | 162.4336 | 133.0752 | 175.208 | 187.9824 | 195.1328 | 230.7552 | 140.9056 | 166.6976 | 231.1312 | 163.0384 | 195.528 | 185.0752 | 201.4336 | 304.8112 | 157.4144 | 153.0752 | 159.6416 | 219.9616 | 136.208 | 125.0752 | 135.9248 | 142.7744 | 171.624 | 174.1696 | 166.2448 | 126.1312 | 101.9056 | 90.6224 | 114.6992 | 113.8496 | 220.6032 | 234.528 | 277.792 | 319.3216 | 225.0384 | 230.6224 | 270.1888 | 290 | 121.584 | 72.4528 | 64.8688 | 121.584 | 107.8864 | 113.68 | 113.792 | 111.0192 | 104.3776 | 100.7168 | 103.0384 | 101.528 | 64.528 | 78.1136 | 77.1312 | 102.736 | 89.3968 | 99.584 | 99.4912 | 71.4144 | 225.2832 | 209.208 | 237.6608 | 369.1328 | 126.0752 | 180.5472 | 194.2832 | 302.8688 | 276.1888 | 211.8304 | 261.3584 | 335.9056
1511 | 18.6154 | 157.8937 | 186.5949 | 224.8456 | 162.8329 | 142.6203 | 151.7975 | 168.3924 | 163.4759 | 133 | 111.3215 | 122.238 | 156.6431 | 179.3215 | 152.119 | 170.5595 | 206.6785 | 163.8329 | 153.6558 | 142.3924 | 176.3569 | 156.2025 | 304.1063 | 339.8202 | 157.1544 | 221.3924 | 211.0609 | 233.7039 | 268.7493 | 118.881 | 107.5468 | 126.0481 | 135.7266 | 140.9164 | 126.9873 | 158.5241 | 77.4405 | 218.5468 | 216.1544 | 268.762 | 265.9164 | 128.0836 | 127.4886 | 146.4405 | 115.5595 | 76.7266 | 66.3215 | 68.5241 | 108.7012 | 179.3697 | 181.7139 | 236.7039 | 150.119 | 146.4278 | 119.9164 | 141.4278 | 117.9646 | 166.238 | 157.0836 | 203.0836 | 136.0963 | 129.9519 | 151.4759 | 156.4405 | 127.881 | 184.7975 | 181.8229 | 194.2507 | 234.6431 | 142.238 | 148.9292 | 224.1544 | 193.7012 | 194.9164 | 172.6076 | 205.5114 | 317.5468 | 159.8683 | 153.4759 | 173.3088 | 212.7975 | 132.3697 | 124.9164 | 136.0836 | 149.9646 | 172.9773 | 162.0481 | 160.6785 | 110.762 | 106.0354 | 97.0708 | 110.8583 | 116.4051 | 213.6785 | 237.2734 | 277.7139 | 309.0609 | 216.5368 | 243.7847 | 272 | 305.6658 | 128.1417 | 69.762 | 77.4405 | 128.1417 | 96.4051 | 121.6431 | 116.5114 | 105.9292 | 99.0481 | 110.8683 | 102.9292 | 99.7975 | 74.5468 | 73.3697 | 69.5949 | 95.4886 | 98.5595 | 101.6658 | 90.5368 | 74.9873 | 222.4051 | 206.4886 | 225.5368 | 374.9646 | 123.119 | 189.9519 | 202.0354 | 336.1417 | 295.9037 | 217.238 | 241.5368 | 331.6431
1512 | 18.6264 | 161.4381 | 208.1535 | 232.5792 | 173.5743 | 132.4332 | 142.5644 | 161.2772 | 156.7079 | 145.0074 | 119.7203 | 118.854 | 155.7153 | 179.1436 | 151.8564 | 164.7104 | 209.7153 | 164.2822 | 133.1386 | 144.9975 | 174.4257 | 159.7178 | 316.5718 | 348.995 | 164.8589 | 206.2723 | 206.151 | 223.8639 | 271.7104 | 113.4257 | 119.5718 | 119 | 140.2921 | 139.4257 | 127.7079 | 144.1361 | 81.5743 | 263.1634 | 219.8564 | 279.0074 | 254.1337 | 122.1411 | 120 | 129.9901 | 117.1436 | 78.4332 | 69.5743 | 76.4356 | 113.5644 | 175.5743 | 188.7153 | 223.4332 | 147.5693 | 129.5544 | 136.151 | 130.844 | 105.995 | 157.7079 | 158.0025 | 208.5767 | 143.7277 | 125.5644 | 147.5668 | 154.2847 | 129.2871 | 166.9876 | 169.4332 | 209.5866 | 214.2747 | 134.8515 | 144.4307 | 231.2871 | 203.7104 | 176.844 | 179.151 | 201.5644 | 304.9851 | 156.849 | 152.4257 | 185.8589 | 190.4134 | 129.7178 | 130.2871 | 147.8663 | 143.1411 | 162.5743 | 152.1411 | 172.8639 | 115.2896 | 114.5743 | 107.8589 | 103.005 | 116.0025 | 211.5693 | 243.5718 | 283.5718 | 304.151 | 225.4431 | 248.4233 | 277.7178 | 309.995 | 148.0025 | 72.0025 | 86.7203 | 148.0025 | 113.156 | 111.5668 | 111.9926 | 116.297 | 108.0099 | 112.995 | 104.7203 | 103.7153 | 73.9926 | 62.1411 | 67.1386 | 89.7153 | 98.4282 | 100.849 | 88.5792 | 86 | 215.7129 | 192.1386 | 240.7327 | 375.5743 | 144.0124 | 188.4233 | 195.1361 | 321.6955 | 281.9827 | 205.2772 | 228.1436 | 321.5668
1513 | 18.6374 | 150.2211 | 200.5574 | 229.7345 | 179.3363 | 143.2567 | 164.9559 | 167.9381 | 146.6637 | 151.6637 | 119.407 | 105.6548 | 151.7433 | 179.4159 | 145.7433 | 160.5841 | 211.5841 | 172.2655 | 136.177 | 144.9204 | 184.4337 | 158.1593 | 303.5663 | 359.3541 | 165.2478 | 201.0974 | 220.6726 | 233.7522 | 284.5222 | 119.3452 | 106.5663 | 122.5044 | 143.9115 | 136 | 120.5841 | 130.7433 | 63.9734 | 262.8848 | 219.5841 | 283.9115 | 245.9204 | 122.3363 | 124.6726 | 124.4248 | 117.4159 | 85.1681 | 72.4159 | 90.8407 | 116.0974 | 174.9115 | 190 | 220.2389 | 147.3363 | 119.5133 | 133.7345 | 126.6815 | 111.6815 | 153.5044 | 158.6637 | 218.5044 | 167.593 | 128.6815 | 139.6637 | 145.9911 | 125.7433 | 153.7522 | 175 | 215.7345 | 222.6107 | 122.7433 | 149.9204 | 240.593 | 195.4956 | 177.3541 | 194.2567 | 188.9115 | 303.3541 | 149.9204 | 160.0974 | 171.6459 | 198.2833 | 131.0796 | 135.5044 | 139.1415 | 142.7522 | 167.1681 | 162.8496 | 180.4159 | 124.6726 | 107.4867 | 114.6726 | 114.2567 | 128.3452 | 214.8407 | 256.2655 | 271.7345 | 318.0885 | 224.8937 | 262.4426 | 277.3274 | 307.5044 | 134.0619 | 74.4159 | 79.3982 | 134.0619 | 128.4159 | 107.7522 | 118.1859 | 120.6548 | 112.407 | 89.4778 | 112.5841 | 93.3185 | 57.407 | 68.7611 | 68.4248 | 89.8319 | 93.9115 | 99.177 | 98.5841 | 92.4248 | 211.6637 | 190.5044 | 238.4689 | 362.0619 | 147.9026 | 189.593 | 182.9115 | 304.8407 | 279.1063 | 183.3185 | 236.593 | 321.2567
1514 | 18.6484 | 152.4873 | 194.9455 | 230.5418 | 167.88 | 154.1564 | 158.9709 | 166.88 | 135.4582 | 161.3309 | 114 | 124.6437 | 143.2291 | 192.7164 | 137.8254 | 144.3018 | 207.2291 | 166.6872 | 146.1746 | 140.44 | 190.4218 | 134.9163 | 297.5782 | 346.5491 | 160.4218 | 215.56 | 216.2473 | 254.0837 | 288.6509 | 118.2473 | 105.9455 | 123.8073 | 130.2654 | 141.3673 | 125.7709 | 148.4691 | 64.5055 | 231.88 | 233.12 | 266.6872 | 248 | 124.5964 | 140.5237 | 132.5782 | 124.7527 | 74.0727 | 76.7709 | 85.4582 | 122.8073 | 186.9091 | 200.7346 | 208.6327 | 138.2654 | 132.3673 | 130.9636 | 132.6145 | 118.2109 | 156 | 154.6145 | 206.6872 | 164.0545 | 135.2109 | 140.3855 | 146.9636 | 131.5418 | 162.7527 | 189.3128 | 212.9636 | 233.8254 | 150.011 | 153.1927 | 255.5418 | 203.1382 | 179.4582 | 190.2473 | 200.3128 | 310.8254 | 147.2291 | 148.9163 | 164.7709 | 219.3855 | 122.44 | 151.12 | 133.3491 | 136.8436 | 169.1927 | 180.9455 | 172.8436 | 129.1927 | 104.7709 | 114.4218 | 119.1927 | 122.4763 | 216.0182 | 250.0909 | 261.2109 | 319.2109 | 231.68 | 264.4763 | 248.9527 | 310 | 129.7527 | 67.44 | 63.0364 | 129.7527 | 115.4763 | 101.8436 | 123.0364 | 118.7709 | 102.2291 | 97.8728 | 118.3673 | 92.7527 | 62.1382 | 69.0364 | 63.4582 | 83.6327 | 89.2109 | 93.0727 | 95.4218 | 88.0545 | 197.4763 | 199.56 | 227.56 | 351.7891 | 146.56 | 180.0909 | 178.8073 | 297.0727 | 277.6872 | 182.1564 | 230.0727 | 342.2946
1515 | 18.6594 | 158.3048 | 206.4779 | 224.2182 | 185.7808 | 149.8702 | 145.4779 | 171.3904 | 146.651 | 155.6096 | 112.7827 | 140.7827 | 155.2164 | 197.7394 | 133.174 | 157.3462 | 194.2615 | 161.8692 | 140.4788 | 139.6519 | 180.0875 | 134.1731 | 321.9116 | 328.7827 | 160.4346 | 210.8702 | 217.0866 | 251.175 | 271.4355 | 119.0866 | 115.6519 | 123 | 119.3481 | 143.7827 | 127.7827 | 156.9134 | 85.3471 | 230.3039 | 255.9991 | 246.0442 | 231.5663 | 112.2182 | 137.4355 | 144.7385 | 123.3048 | 72.0866 | 85.4779 | 77.174 | 108.0009 | 184.8269 | 183.0451 | 203.174 | 163.2587 | 135.3913 | 124.0442 | 110.9144 | 109.0875 | 153.5654 | 144.4788 | 198.826 | 156.174 | 130.9567 | 138.3481 | 138.8269 | 128.8702 | 158.2615 | 173.4365 | 209.6961 | 214.5663 | 152.7403 | 149.7394 | 254.6961 | 183.8278 | 169.9567 | 178.3048 | 208.7827 | 306.7827 | 134.8702 | 134.7827 | 168.6087 | 218.5654 | 118 | 130.1759 | 137.7827 | 138.0866 | 153.5663 | 183.9567 | 164.9567 | 122.6961 | 100.6961 | 109.5654 | 107.2182 | 109.7394 | 210.9567 | 237.5654 | 270.3471 | 322.2606 | 245.6087 | 231.0451 | 270.3453 | 305.1308 | 132.5654 | 75.7818 | 66.3039 | 132.5654 | 116.1298 | 101.8692 | 120.826 | 108.6096 | 94.1308 | 104.6961 | 112.2615 | 108.9558 | 69 | 67.4346 | 59.4346 | 83.0433 | 90.4346 | 71.6096 | 89.9567 | 91.7385 | 202.3904 | 195.4788 | 219.2182 | 336.5663 | 136.3923 | 163.9134 | 204.7808 | 301.1731 | 275.9125 | 188.2173 | 208.6096 | 357.6519
1516 | 18.6704 | 147.7905 | 209.3791 | 194.8878 | 186.8554 | 146.4838 | 160.798 | 183.7257 | 144.3392 | 152.5885 | 120.6933 | 119.5088 | 156.9277 | 186.3067 | 132.6209 | 161.202 | 197.3142 | 162.5162 | 138.1047 | 143.2419 | 179.3466 | 142.6209 | 315.5088 | 324.8953 | 160.7581 | 208.7257 | 215.4115 | 237.2743 | 277.0399 | 108.7182 | 111.1696 | 112.4439 | 143.0798 | 149.2095 | 126.3791 | 154.8628 | 95.7257 | 221.3392 | 250.0973 | 233.4115 | 204.8554 | 117.6608 | 129.3791 | 146.0324 | 119.6209 | 77.8628 | 70.3716 | 67.3067 | 94.0324 | 180.1047 | 176.9351 | 206.9676 | 149.8479 | 123.823 | 120.1047 | 109.177 | 114.5561 | 150.7581 | 153.9027 | 220.4912 | 148.7905 | 112.2344 | 134.7581 | 139.6933 | 118.0324 | 143.3067 | 176.8703 | 205 | 214.5561 | 147.8304 | 143.2743 | 222.6783 | 180.0399 | 182.9027 | 174 | 201.7905 | 302.2743 | 140.798 | 140.2095 | 167.7581 | 198.3716 | 123.5885 | 113.4838 | 136.3791 | 146.3466 | 147.9676 | 182 | 163.6209 | 129.798 | 99.7257 | 96.202 | 105.2095 | 119.4189 | 226.3865 | 220.4763 | 254.6459 | 331.2095 | 237.9277 | 217.4838 | 264.5736 | 302.6209 | 140.3142 | 74.6858 | 80.3142 | 140.3142 | 110.823 | 117.4189 | 112.6858 | 107.4514 | 104.0399 | 92.5486 | 104.7581 | 112.2743 | 69.6209 | 64.0324 | 64.1047 | 86.2419 | 98.8304 | 84.7331 | 89.8628 | 85.5811 | 233.975 | 199.9351 | 223.4189 | 336.5561 | 146.2493 | 166.8304 | 210.202 | 331.975 | 275.4115 | 198.9351 | 219.8703 | 372.4189
1517 | 18.6814 | 133.8682 | 200.1347 | 196.5645 | 171.9341 | 133.4355 | 173.6991 | 181.5673 | 144.8653 | 137.6361 | 129.7994 | 116.4986 | 146.9341 | 175.9341 | 153.2636 | 159.6991 | 212.4986 | 174.298 | 138.7335 | 140.8338 | 176.9341 | 134.7679 | 309.3324 | 329.9656 | 148.6017 | 204.0344 | 220.2321 | 229.9341 | 290.0659 | 125.9627 | 103.2006 | 118.6648 | 155.2006 | 155.533 | 124.1003 | 154.7335 | 96.1003 | 216.7994 | 243.533 | 254.6963 | 234.894 | 124.1003 | 120.7679 | 146.7994 | 106.0688 | 86.5989 | 61.533 | 64.533 | 96.0659 | 170.6017 | 188.6991 | 215.0659 | 144.8309 | 124.5989 | 122.6332 | 122.3324 | 107.0688 | 162.0315 | 151.6017 | 194.3725 | 138.6676 | 126.6963 | 137.1662 | 148.1662 | 133.9971 | 150.6648 | 172.5358 | 210.6991 | 221 | 140.6017 | 131.5014 | 228.1633 | 190.533 | 198.9656 | 186.6648 | 204.9656 | 305.0659 | 138.5014 | 156.6648 | 172.0659 | 193.9656 | 147.8968 | 121.9656 | 142.9656 | 150.8997 | 153.533 | 176.9341 | 181.0974 | 141.4327 | 89.9685 | 99.1318 | 110.2665 | 114.3352 | 210.404 | 232.5301 | 259.9971 | 336.8997 | 233 | 220.8997 | 277.0286 | 301.7335 | 125.7364 | 75.3324 | 74.6017 | 125.7364 | 105.8997 | 124.3668 | 107 | 113.2665 | 100.6017 | 89.8997 | 101.467 | 116.3324 | 77.5989 | 62.2665 | 67.8997 | 99.6648 | 97.3009 | 94.8338 | 94.1662 | 95.0974 | 220.7708 | 190.8023 | 231.6332 | 357.5645 | 151.0344 | 179.2321 | 197.9341 | 320.0373 | 288.4642 | 205 | 238.3324 | 373.6676
1518 | 18.6924 | 136.3919 | 204.0374 | 212.1008 | 172.2276 | 152.1122 | 159.5707 | 169.3171 | 148.0634 | 145.0748 | 126.8358 | 141.3659 | 159.4927 | 177.5187 | 140.4699 | 153.3171 | 227.0374 | 174.8992 | 144.4553 | 133.1902 | 193.3659 | 137.7463 | 312.3545 | 326.2537 | 143.9366 | 199.3545 | 207.5707 | 230.8732 | 285.8992 | 127.4439 | 96.4813 | 119.5447 | 142.026 | 155.7089 | 126.8732 | 146.8992 | 90.4813 | 218.3545 | 223.6975 | 248.9886 | 262.5187 | 124.9366 | 122.4553 | 138.6715 | 112.2016 | 83.2537 | 63.6455 | 65.3545 | 100.9366 | 162.0634 | 196.5187 | 220.5821 | 141.735 | 125.1268 | 128.8098 | 133.7463 | 111.265 | 163.1902 | 139.1902 | 196.5301 | 139.5187 | 142.9366 | 120.2797 | 155.1642 | 141.7724 | 152.1902 | 178.4927 | 211.4179 | 222.2911 | 133.3545 | 137.6195 | 228.735 | 184.8992 | 189.4439 | 197.8732 | 188.9886 | 300.8992 | 142.6829 | 153.6715 | 161.4439 | 198 | 142.5707 | 124.0634 | 130.8618 | 147.4813 | 158.2276 | 170.7724 | 178.0894 | 127.8618 | 82.3545 | 107.5821 | 107.1268 | 118.6195 | 213.0748 | 252.7463 | 254.8618 | 316.6975 | 229.1268 | 212.9626 | 263.3691 | 302.9366 | 148.8586 | 79 | 81.5561 | 148.8586 | 106.3545 | 118.8358 | 118.6195 | 115.9366 | 95.9366 | 92.9366 | 116.7837 | 108.3805 | 72.3171 | 69.4553 | 65.1268 | 92.1528 | 102.3919 | 89.7724 | 97.2911 | 103.7089 | 216.6195 | 193.6195 | 234.5821 | 355.026 | 141.8358 | 193.6829 | 192.4179 | 303.7089 | 288.9626 | 200.4813 | 228.4439 | 328.0406
1519 | 18.7034 | 134.4218 | 205.7109 | 229.8142 | 162.1592 | 149.2389 | 163.1564 | 167.9469 | 149.6313 | 145.1327 | 125.9735 | 129.6341 | 159.4483 | 187.8938 | 135.5517 | 133.528 | 214.8967 | 172.9735 | 132.8701 | 129.3422 | 200.0531 | 138.7109 | 310.6844 | 318.0531 | 148.9469 | 196.3687 | 194.0531 | 234.9735 | 282.6578 | 98.9498 | 109.7877 | 118.6313 | 141.9204 | 140.528 | 129 | 143.6578 | 88 | 232.472 | 227.7877 | 240.6313 | 245.2654 | 124.6844 | 119.4218 | 129.0531 | 120.6578 | 69.1327 | 71.8938 | 61.0531 | 103.9735 | 180.7346 | 180.581 | 223.9735 | 155.3924 | 123.6578 | 140.5517 | 139.9735 | 110.1062 | 165.2626 | 139.9469 | 194.8701 | 149.8938 | 129.528 | 134.3393 | 159.3156 | 121.581 | 150.9735 | 194.236 | 195.528 | 209.1858 | 132.3422 | 157.1564 | 208.1858 | 175.7374 | 188.5782 | 181.581 | 197.7346 | 322.655 | 131.5545 | 144.7109 | 161.8938 | 182.8701 | 146.1564 | 130.8938 | 120.6844 | 148.2891 | 152.1062 | 175.5782 | 178.8938 | 115.7109 | 103.7081 | 101.1062 | 111.5782 | 135.5251 | 216.4218 | 224.7668 | 274.9441 | 326.0502 | 218.4483 | 219.8408 | 258.4986 | 288.8701 | 141.0294 | 73.0796 | 65.9763 | 141.0294 | 97.4483 | 114.0265 | 111.8436 | 106.4749 | 96.3422 | 86.764 | 110.2123 | 109.236 | 80.1564 | 69.7109 | 62.3422 | 84 | 96.4749 | 86.6844 | 96.0265 | 83.9232 | 220.3687 | 194.7374 | 222.8436 | 366.7611 | 132.4218 | 181.2389 | 194.9469 | 293.7905 | 281.3687 | 200.6313 | 223.6313 | 329.9972
1520 | 18.7144 | 145.0723 | 193.2782 | 219.5771 | 172.7528 | 133.2989 | 155.2576 | 170 | 136.2576 | 148.7114 | 116.9483 | 122.6908 | 161.031 | 200.7114 | 130.6184 | 146.1033 | 213.3712 | 158.5874 | 141.0827 | 147.093 | 187.9483 | 142.3609 | 301.2886 | 305.2782 | 140.2782 | 213.7631 | 200.0413 | 235.3299 | 281.6598 | 115.4745 | 105.938 | 126.031 | 130.2576 | 142.3816 | 132.3506 | 154.0517 | 84.6494 | 221.907 | 251.4126 | 241.3299 | 251.7528 | 124 | 122.7011 | 152.4642 | 116.3092 | 71.3712 | 63.938 | 60.3402 | 99.6391 | 200.3816 | 170.3299 | 220.3092 | 147.907 | 126.0103 | 134.2782 | 145.6908 | 115.3816 | 157.2782 | 170.1447 | 189.0103 | 164.7218 | 132.0517 | 116.1749 | 157.3196 | 138.1343 | 150.6598 | 185.2679 | 179.9587 | 210.0413 | 135.3506 | 145.2369 | 216.4126 | 185.062 | 170.5565 | 184.062 | 197.9483 | 314.5565 | 133.7218 | 149.031 | 163.9897 | 195.7734 | 126.1956 | 138.3506 | 131.3919 | 136.5978 | 150.0103 | 166.938 | 176.2989 | 120.7011 | 102.2679 | 108.3919 | 110.9793 | 123.5771 | 206.2989 | 219.3919 | 243.0922 | 334.9897 | 227.4022 | 238.062 | 252.938 | 283.0103 | 139.062 | 72.6804 | 71.4229 | 139.062 | 97.6908 | 105.6288 | 100.9793 | 103.0103 | 95.3299 | 109.1343 | 101.3299 | 121.7114 | 83.6494 | 72.6908 | 57.3092 | 89.3609 | 89.6598 | 99.4022 | 92.3196 | 78.0207 | 207.6081 | 216.7941 | 228.062 | 378.0103 | 115.5978 | 172.6701 | 188.2886 | 323.8458 | 275.9793 | 197.3092 | 230.3609 | 351.0413
1521 | 18.7254 | 134.94 | 208.4721 | 219.1888 | 165.3048 | 138.8712 | 161.648 | 167.9528 | 152.8841 | 146.8584 | 125.648 | 117.4936 | 151.0344 | 205.6824 | 133.5064 | 149.8584 | 208.1288 | 157.824 | 156.5064 | 157.3648 | 172.0816 | 145.6824 | 309.9656 | 300.6824 | 141.824 | 225.0472 | 203.3176 | 222.0344 | 284.412 | 124.9056 | 104.7768 | 143.3304 | 129.1416 | 135.3992 | 132.6352 | 154.2232 | 85.7296 | 220.5064 | 234.4335 | 254.648 | 247.7168 | 134.9184 | 127.3648 | 146.5751 | 122.1888 | 77.7296 | 83.2489 | 65.0944 | 94.9528 | 196.1288 | 166.588 | 217.3176 | 150.6008 | 128.3648 | 119.4464 | 139.1288 | 125.4592 | 156.7768 | 156.7039 | 185.2232 | 159.764 | 141.7768 | 131.7081 | 160.7768 | 142.1288 | 150.6824 | 195.3776 | 194.4249 | 215.3648 | 153.3776 | 146.9184 | 212.3992 | 204.648 | 161.3648 | 182.4936 | 209.3776 | 298.5408 | 154.6952 | 157.4592 | 171.1888 | 230.5665 | 120.5064 | 120.2103 | 131.5408 | 121.1288 | 148.2704 | 154.176 | 152.5279 | 110.352 | 89.176 | 107.176 | 115.824 | 117.7296 | 204.3648 | 223.6352 | 245.9313 | 321.7168 | 223.764 | 237.8584 | 278.3905 | 305.8369 | 152.5064 | 72.6352 | 69.764 | 152.5064 | 100 | 106.7768 | 109.9184 | 122.4249 | 97.7296 | 109.0344 | 92.8112 | 108.2575 | 83.3648 | 59.3048 | 54.3176 | 93.3648 | 98.5536 | 86.2103 | 92.3648 | 78.6352 | 204.0472 | 213.3048 | 218.9872 | 372.176 | 111.7296 | 172.3176 | 188.7768 | 345.7768 | 245.3391 | 197.0472 | 217.9872 | 337.94
1522 | 18.7364 | 153.3983 | 198.5486 | 226.4735 | 165.6416 | 137.4425 | 154.1061 | 165.6106 | 150.1061 | 140.5265 | 118.1061 | 112.6106 | 145.6947 | 195.5796 | 139.7788 | 158.1151 | 203.3053 | 167.2522 | 146.8008 | 152.4425 | 169.7788 | 152.9469 | 311.8318 | 303.7788 | 154.7257 | 223.2212 | 214.1151 | 236.1461 | 278.3584 | 113.969 | 115.3363 | 111.7919 | 134.0841 | 143.8939 | 116.0221 | 142.2743 | 92.5575 | 233.7257 | 216.7478 | 253.3584 | 231.5796 | 108.7388 | 134.2522 | 147.7257 | 126 | 80.3894 | 65.1282 | 69.7788 | 98.1682 | 182.969 | 176.8098 | 201.7168 | 176.8408 | 135.9469 | 121.2522 | 146.8098 | 121.7478 | 157.6106 | 166.9249 | 195.5045 | 168.1992 | 128.0221 | 146.6947 | 157.4425 | 128.969 | 164.1992 | 171.7388 | 196.0531 | 203.4955 | 150.5796 | 153.3894 | 230.6196 | 190.1592 | 184.9249 | 163.7168 | 196.1592 | 309.1992 | 168.2522 | 144.7168 | 163.1902 | 219.0751 | 135.8098 | 119.3363 | 142.1992 | 143.3983 | 160.8939 | 141.2743 | 165.2302 | 106.7788 | 93.6416 | 97.7478 | 105.8008 | 119 | 205 | 232.031 | 258.6947 | 332.6726 | 230.8098 | 230.8318 | 249.2343 | 293.0751 | 150.4425 | 68.5265 | 59.4425 | 150.4425 | 101.3894 | 100.6637 | 121.9469 | 120.5796 | 92.0531 | 92.5796 | 95.2522 | 104.1682 | 74.2743 | 63.1151 | 57.4735 | 87.7478 | 110.6416 | 95.7567 | 96.4735 | 74.5265 | 207.7788 | 228.2302 | 231.4514 | 391.2302 | 130.3673 | 183.8098 | 168.7698 | 336.1902 | 267.4293 | 188.969 | 218.6416 | 330
1523 | 18.7474 | 146.6184 | 196.3629 | 215.2743 | 143.5486 | 154.0886 | 158.6048 | 157.9302 | 149.6559 | 133.3441 | 112.707 | 107.0511 | 147.414 | 205.1397 | 155.1397 | 151.6883 | 198.0511 | 171.414 | 142.414 | 141.5162 | 172.414 | 158.8279 | 311.414 | 288.0324 | 156.1721 | 235.4327 | 216.1721 | 219.5486 | 292.6746 | 110 | 119 | 99.9489 | 130.0511 | 140.1022 | 123.8467 | 138 | 93.586 | 202.6508 | 220.3629 | 250 | 262.3492 | 116.9165 | 123.5673 | 140.6883 | 100.5486 | 73.9302 | 59.7768 | 61.8092 | 83.0324 | 175.4651 | 201.7956 | 199.9489 | 158.4276 | 124.1533 | 126.121 | 130.0835 | 129.6048 | 149.9302 | 165.6883 | 222.2095 | 159.1533 | 142.2095 | 134.9813 | 161.3629 | 129.2419 | 165.0511 | 176.3816 | 184.5162 | 223.4514 | 141.0511 | 158.9489 | 212.0137 | 183.8279 | 201.3629 | 178.2095 | 203.9676 | 319.2419 | 183.0187 | 150.7257 | 147.3952 | 200.5162 | 128.9813 | 119.4651 | 143.0511 | 140.1533 | 174.7768 | 152.5536 | 167.9302 | 105.879 | 101.2419 | 107.7257 | 97.1721 | 128.8978 | 214.1908 | 238.8279 | 244.1533 | 334.3441 | 231.0511 | 227.586 | 262.5213 | 327.54 | 133.1533 | 69.8279 | 68.3117 | 133.1533 | 96.3441 | 105.4838 | 122.879 | 113.1721 | 101.7257 | 100.7257 | 98 | 110.2419 | 61.5162 | 74.3629 | 72.4327 | 93.4838 | 109.7581 | 98.3441 | 91.6371 | 69.4651 | 212.5349 | 253.5536 | 223.7394 | 379.0835 | 123.1533 | 180.5162 | 167.4838 | 327.4651 | 270.2743 | 194.1908 | 209.9813 | 353.3305
1524 | 18.7584 | 153.3855 | 201.1578 | 232.2975 | 155.2975 | 136.8254 | 157.9651 | 157.1578 | 159.912 | 143.2277 | 114.4386 | 115.0699 | 163.1048 | 189.4217 | 145.8952 | 168.5783 | 180.8952 | 161.2109 | 120.7025 | 130.8072 | 191.7012 | 140.5796 | 311.2807 | 291.7891 | 149.9651 | 221.5796 | 186.9483 | 197.4916 | 272.6675 | 108.5614 | 108.9301 | 106.3157 | 113.6145 | 142.4735 | 109.141 | 146.6313 | 81.4916 | 203.8241 | 198.5446 | 272.2975 | 256.141 | 146.859 | 135.9819 | 133.8422 | 116.6132 | 67.4036 | 61.5614 | 66.6313 | 89.6662 | 182.6313 | 209.2807 | 216.3855 | 146.2807 | 124.4735 | 132.7542 | 138.2626 | 121.7723 | 157.7891 | 160.2807 | 213.7374 | 147.2458 | 135.1759 | 135.0349 | 165.4386 | 146.8241 | 152.9301 | 187.5964 | 204.0168 | 209.5446 | 133.9651 | 137.9832 | 210.7891 | 201.5433 | 191.053 | 180.5265 | 201.6494 | 297.2639 | 158.5097 | 131.5446 | 159.5433 | 211.3855 | 119.6843 | 128.0699 | 143.1578 | 136.8771 | 162.8952 | 155.4036 | 167.1578 | 124.4204 | 102.2807 | 107.9651 | 103.912 | 112.8603 | 226.6313 | 240 | 229.3687 | 325.5265 | 236.912 | 215.4916 | 266.6494 | 327.2988 | 116.2109 | 72.4386 | 73 | 116.2109 | 99.0349 | 94.6145 | 122.7193 | 113.4386 | 90.4567 | 108.8771 | 97.2807 | 104.8072 | 58 | 84.912 | 66.4916 | 84.053 | 103.6843 | 89.5265 | 107.7012 | 81.6662 | 260.7529 | 216.8435 | 214.8422 | 366.4036 | 119.8771 | 181.3157 | 186.1048 | 338.947 | 270.7542 | 195.8422 | 193.4916 | 336.3868
1525 | 18.7694 | 158.2684 | 172.7374 | 218.3214 | 166.9263 | 129.9263 | 153.0737 | 152.879 | 152.0265 | 150.1947 | 104.7581 | 108.7581 | 160.2212 | 192.7051 | 147.3157 | 154.3214 | 201.3364 | 160.6578 | 117.121 | 148.4839 | 163.8848 | 141.7788 | 313.1947 | 313.5576 | 151.6578 | 201.5634 | 191.3629 | 208.3629 | 271.5104 | 123.3629 | 114.5104 | 116.7788 | 127.7523 | 125.9793 | 116.3629 | 139.0265 | 77 | 207.8053 | 216.0679 | 241.4954 | 255.3157 | 128.6636 | 140.0737 | 139.5841 | 116.7581 | 78.4366 | 71.9735 | 64.1475 | 105.2419 | 186 | 196.5634 | 214.6843 | 127.7108 | 129.1947 | 117.4424 | 129.6371 | 122.121 | 148.8318 | 158.5369 | 209.1947 | 147.9263 | 139.2419 | 128.9528 | 169.6578 | 144.2212 | 146.8053 | 166.3214 | 184.4689 | 214.6313 | 140.7788 | 156.0679 | 228.1682 | 196.2949 | 174.2949 | 178.7316 | 208.2419 | 308.4839 | 157.2419 | 130.7788 | 172.5841 | 219.9263 | 128.9735 | 132.7316 | 156.4366 | 149.7051 | 161.3894 | 175.947 | 156.2949 | 132.7316 | 112.2419 | 104.5369 | 113.5841 | 122.5576 | 226.3422 | 207.8111 | 242.826 | 323 | 240 | 213.1947 | 263.7316 | 299.5161 | 120.0472 | 78.8525 | 74.4631 | 120.0472 | 101.7316 | 106.5576 | 117.1475 | 107.4159 | 100.826 | 101.9528 | 101.3894 | 115.1682 | 58.7316 | 75.5634 | 73.7051 | 81.1947 | 108.5841 | 73.8318 | 94.5161 | 81.879 | 253.3952 | 234.3836 | 236.6786 | 359.1475 | 151.7258 | 164.7108 | 182.4896 | 327.9055 | 255.4424 | 211.826 | 218.9942 | 328.9263
1526 | 18.7804 | 152.793 | 181.3403 | 211.4877 | 151.6351 | 140.6701 | 139.3544 | 163.6456 | 146.5123 | 141.3299 | 111.414 | 110.207 | 142.1229 | 199.9754 | 144.793 | 143.0246 | 196.1229 | 157.5369 | 138.3403 | 141.8667 | 177.7789 | 130.8667 | 311.7684 | 299.9159 | 161.1824 | 205.1824 | 197 | 189.4527 | 311.449 | 111.8913 | 120.2316 | 108.8421 | 135.7439 | 129.414 | 128.6947 | 144.6701 | 83.6947 | 219.6456 | 233.4386 | 239.6456 | 268.414 | 134.621 | 121.1474 | 147.207 | 116.7193 | 74.8176 | 74.5123 | 62.7439 | 114.207 | 175.586 | 168.1966 | 213.4877 | 135.1333 | 116.6106 | 117.6947 | 123.2561 | 115.0737 | 149.207 | 159.4877 | 191.4036 | 146.7684 | 137.0491 | 122.2807 | 161.3299 | 123.8913 | 152.6947 | 179.5719 | 198.5473 | 236.3649 | 140.2807 | 168.2316 | 195.8071 | 171.172 | 190.0841 | 183.4631 | 185.965 | 297.4036 | 149.0983 | 134 | 145.2457 | 218.0246 | 136.7193 | 134.4877 | 146.8667 | 146.5614 | 164.4877 | 184.7439 | 158.6947 | 123.3299 | 98.8913 | 99.5369 | 106.3299 | 112.6351 | 222.0246 | 206.414 | 261.6456 | 299.1966 | 228.8421 | 245.242 | 248.379 | 276.379 | 144.5719 | 67.6106 | 72.0246 | 144.5719 | 103.4877 | 95.1474 | 113.5123 | 96.0737 | 105.5123 | 92.3053 | 113.414 | 102.8913 | 55.2807 | 68.0246 | 72.793 | 90.9263 | 96.1229 | 80.1579 | 104.8526 | 76.2807 | 235.8176 | 227.6597 | 245.7439 | 368.1579 | 147.379 | 179.5719 | 183.4631 | 315.3053 | 273.5473 | 201.6351 | 215.1229 | 335.9509
1527 | 18.7914 | 160.0738 | 173.6332 | 199.9016 | 180.0267 | 149.293 | 159.1967 | 175.5615 | 152.8053 | 142.5369 | 112.7316 | 107.4631 | 150.6107 | 194.1947 | 135.4385 | 142 | 188.7316 | 166.5861 | 146.5123 | 152.8791 | 178.6824 | 126.7561 | 303.4385 | 282.4139 | 169.293 | 208 | 203.8053 | 219.5636 | 268.8012 | 102.9754 | 126.293 | 110.293 | 155.6599 | 133.7561 | 122.6824 | 132.1209 | 95.8299 | 209.6332 | 216.3401 | 263.6599 | 244.7787 | 132.4385 | 120.293 | 134.6332 | 133.8791 | 77.293 | 80.0492 | 79.6353 | 104.6578 | 175.0246 | 183.4406 | 192.0717 | 137.7316 | 121.8299 | 132.0984 | 112.4139 | 120.3176 | 158.5615 | 172.8545 | 197.8545 | 139.1947 | 135.7561 | 129.3176 | 151.1947 | 129.3422 | 171.6353 | 171.1209 | 219.8545 | 223.0963 | 140.5123 | 156.9016 | 202.6599 | 198.7828 | 181.877 | 176.6824 | 195.9037 | 323.5144 | 148.7807 | 169.539 | 143.3176 | 220.7807 | 126.6578 | 142.5615 | 138.9754 | 138.707 | 161.9754 | 175.9262 | 177.6353 | 120 | 99.8053 | 106.3176 | 105.2684 | 117.5861 | 250.4898 | 234.9529 | 247.8524 | 340.1496 | 223.4877 | 215.168 | 258.123 | 288.3914 | 148.9754 | 78.123 | 78.5615 | 148.9754 | 110.8053 | 107.9037 | 116.0246 | 99.8053 | 111.8053 | 94.5369 | 106.4139 | 108.3422 | 56.2684 | 73.8053 | 90.6599 | 93.2439 | 92.5123 | 85.5123 | 86.8033 | 71.9754 | 205.7787 | 236.8791 | 237.6824 | 358.3893 | 151.8299 | 177.9262 | 185 | 324.3422 | 270.6578 | 205.0738 | 203.1947 | 362.1967
1528 | 18.8024 | 159.1578 | 183.6003 | 228.2745 | 167.9468 | 139.4734 | 142.4099 | 160.3259 | 148.8525 | 140.1578 | 112 | 101.3894 | 138.8628 | 183.5472 | 150.6741 | 158.1372 | 191.8422 | 161.5472 | 133.1681 | 141.0943 | 204.4323 | 144.6741 | 291.0103 | 305.1269 | 161.0103 | 199.5472 | 214.4528 | 231.2316 | 284.0429 | 115.0635 | 125.6947 | 101.2418 | 130.494 | 160.1269 | 129.9897 | 139.295 | 92.084 | 244.9588 | 222.295 | 253.8628 | 243.6844 | 131.5369 | 121.2316 | 136.916 | 118.2521 | 84.3791 | 75.084 | 62.7153 | 105.6106 | 192.1372 | 190.2316 | 219.5798 | 130.084 | 120.3894 | 132.1578 | 106.6947 | 112.2418 | 156.3894 | 160.8628 | 205.8422 | 150.0635 | 131.3894 | 122.7787 | 137.4734 | 146.8319 | 160.8628 | 176.7582 | 194.0309 | 207.0103 | 141.7684 | 139.1681 | 199.0103 | 200.7787 | 163.1681 | 177.0738 | 213.5266 | 296.3465 | 152.3053 | 149.494 | 151.3791 | 210.4734 | 126.0738 | 109.6518 | 122.6312 | 136.2316 | 174.8319 | 155.3259 | 186.8422 | 117.6947 | 86.6312 | 99.7787 | 99.084 | 124.0738 | 236.9468 | 233.0103 | 232.7787 | 339.8628 | 223.7684 | 224.2847 | 233.7993 | 303.9897 | 150.3053 | 70.705 | 66.3997 | 150.3053 | 95.3259 | 94.789 | 122.3791 | 98.1578 | 107.084 | 99.8422 | 99.1578 | 93.5575 | 60.0738 | 69.8525 | 76.2521 | 93.7684 | 92.2316 | 83.6947 | 88.2213 | 72.5369 | 213.1372 | 236.6209 | 241.1475 | 342.4734 | 154.2316 | 194.9794 | 198.0635 | 303.4099 | 280.8319 | 219.5266 | 214.0635 | 380.7582
1529 | 18.8134 | 139.2625 | 189.5615 | 215.3588 | 172.711 | 134.4385 | 152.1761 | 149.5349 | 149.3422 | 139 | 98.7276 | 109.3688 | 141.0266 | 211.4485 | 132.5781 | 154.412 | 196.9037 | 169.1495 | 131.3422 | 139.1229 | 198.1661 | 145.3156 | 330.9402 | 305.9734 | 165.8073 | 204.0266 | 217 | 213.0432 | 300.2458 | 116.6578 | 111.7276 | 110.4917 | 122.5615 | 139.8937 | 137.6844 | 151.588 | 88.4385 | 229.6744 | 233.0266 | 260.711 | 238.9734 | 134.3422 | 145.9834 | 124.9469 | 129.9568 | 78.1927 | 74.5615 | 63.8073 | 111.6844 | 176.701 | 176.7276 | 205.0166 | 168.598 | 127.588 | 135.6844 | 109.1229 | 117.588 | 154.2193 | 171.6146 | 203.8771 | 150.0963 | 120.6312 | 144.2026 | 111.3588 | 148.6578 | 145.8505 | 173.7542 | 208.4219 | 224.299 | 127.9469 | 119.3854 | 213.9568 | 191.7542 | 188.6678 | 179.5615 | 213.0963 | 302.9568 | 160.0266 | 160.299 | 167.8339 | 205.4385 | 123.0963 | 125.5448 | 119.5615 | 136.7807 | 171.9734 | 155.4651 | 184.0963 | 131.8339 | 103.8605 | 96.2193 | 107.1495 | 107.0432 | 254.2026 | 237.0266 | 233.1229 | 300.6479 | 222.4385 | 233.3422 | 240.6146 | 305.4385 | 143.9734 | 49.0432 | 80.7375 | 143.9734 | 103.2724 | 103.0531 | 120.0963 | 106.3688 | 96.412 | 100.2193 | 95.6578 | 92.6844 | 73.4917 | 66.4385 | 60.6312 | 96.3422 | 91.2193 | 89.2458 | 91 | 81.588 | 235.1761 | 212.3588 | 242.2193 | 360.8605 | 132.1395 | 171.3322 | 201.2193 | 315.5183 | 266.2625 | 185.5249 | 203.1661 | 363.701
1530 | 18.8244 | 144.3093 | 213.7906 | 223.2744 | 175.207 | 152.2395 | 170.4814 | 155.1372 | 166.6535 | 142.1721 | 105.3093 | 114.3791 | 153.3093 | 196.2094 | 146.6186 | 174.2046 | 206.7232 | 159.3117 | 131.207 | 152.6883 | 211.4465 | 147.1721 | 314.4512 | 333.3418 | 166.414 | 217.1023 | 205.8977 | 220.6883 | 269.4861 | 107.2768 | 103.2419 | 119.5512 | 126.1721 | 132.793 | 121.5535 | 128.6233 | 98.3093 | 223.586 | 207.2442 | 222.7629 | 229.8628 | 137.3791 | 121.2791 | 122.586 | 130.2419 | 74.414 | 63.1047 | 72.3442 | 124.8953 | 183.6883 | 197.5837 | 201.1721 | 158.5884 | 133.9651 | 133.0349 | 121.8953 | 127.9302 | 168.2744 | 159.3465 | 193.4837 | 154.5512 | 123.5512 | 138.3117 | 144.6511 | 129.7605 | 147.7581 | 159.3117 | 194.3814 | 228.414 | 143.8255 | 164.1673 | 235.6535 | 178.8977 | 162.9001 | 172.8628 | 200.8977 | 324.6535 | 136.6233 | 139.8303 | 155.3465 | 216.1023 | 141.8255 | 133.793 | 127.9302 | 142.5512 | 174.7581 | 164.1372 | 180.6209 | 119.3465 | 99.6907 | 96.793 | 118.7232 | 109.9302 | 235.6233 | 228.6907 | 230.8279 | 325.8929 | 230.7232 | 214.9675 | 247.3791 | 289.1396 | 125.3465 | 74.1348 | 71.7256 | 125.3465 | 108.586 | 91.1396 | 108.6907 | 108.207 | 110.6535 | 84.1396 | 104.5163 | 101.9302 | 73.0349 | 68.3791 | 67.5163 | 86.6907 | 94.1721 | 82.2768 | 91 | 72.1047 | 220.1745 | 216.3093 | 228.5186 | 362.2419 | 128.3791 | 183.6186 | 203.3791 | 326.5512 | 290.3418 | 197.9976 | 213.2744 | 344.5186
1531 | 18.8354 | 142.9734 | 212.7522 | 207.6725 | 166.9469 | 136.8672 | 152.2566 | 173.1063 | 177.4425 | 138.1681 | 118.4691 | 110.1681 | 142.3097 | 190 | 128.6459 | 150.2034 | 204.9734 | 170.4956 | 140.6638 | 147.1416 | 175.7344 | 160.885 | 302.9734 | 318.4512 | 145.0619 | 203.8937 | 207.0266 | 222.3894 | 294.8231 | 100.9734 | 104.4159 | 118.5841 | 133.4425 | 123.3362 | 137.9381 | 123.6106 | 88.115 | 220.7787 | 200 | 224.885 | 233.6372 | 125.115 | 117.8319 | 124.6106 | 127.3894 | 73.1947 | 64.8319 | 76.4159 | 118.3362 | 193.4425 | 192.7256 | 212.4691 | 186.8231 | 136.6106 | 127.9734 | 124.1947 | 115.5044 | 155.8937 | 168.6903 | 217.5753 | 151.9734 | 128.2213 | 145.4691 | 167.0797 | 118.5575 | 161.885 | 186.6019 | 183.3628 | 206.2566 | 152.2213 | 157.6725 | 221.4778 | 202.5753 | 166.0797 | 197.5753 | 174.6459 | 325.7787 | 140.4691 | 111.2566 | 151 | 217.3894 | 127.6725 | 138.0266 | 130.8053 | 138.3628 | 180.8319 | 139.4247 | 169.5309 | 121.4425 | 106.6638 | 97.8053 | 95.23 | 111.1947 | 237.8584 | 228.4159 | 240.4691 | 346.2744 | 242.6638 | 234.9647 | 269.7434 | 293.0531 | 155.6284 | 73.1416 | 76.0531 | 155.6284 | 107.3894 | 94.2478 | 117.2744 | 105.5841 | 98.0884 | 86.4425 | 94.115 | 102.3894 | 74.4159 | 64.1681 | 78.8584 | 87.2213 | 89.3628 | 82.4159 | 100.6638 | 79.4691 | 222.2478 | 209.3362 | 231.4425 | 398.8497 | 142.6903 | 184.1681 | 206.4159 | 312.6991 | 265.7875 | 198.3628 | 184.7875 | 340.1947
1532 | 18.8464 | 137.912 | 182.3839 | 214.4464 | 181.3521 | 141.8112 | 158.4464 | 181.088 | 185.5408 | 160.7105 | 125.9056 | 121.2641 | 152.0817 | 186.7296 | 136.8993 | 147.9056 | 204.8176 | 178.9056 | 120.1071 | 147.4528 | 166 | 137.8367 | 300.3648 | 299.1007 | 148.9936 | 204.088 | 203.0944 | 198.2895 | 272.7487 | 101.6352 | 106.6352 | 121.2704 | 124.3712 | 129.176 | 138.0944 | 128.9056 | 102.1697 | 238.8049 | 223.7105 | 214.9183 | 257.8929 | 150.6161 | 136.9873 | 137.2641 | 116.3712 | 73.8176 | 66.8176 | 60.6479 | 111.912 | 181.9183 | 207.9873 | 215 | 171.2895 | 142.7232 | 133.5408 | 148.5281 | 116.088 | 156.088 | 185.8993 | 192.1135 | 151.8176 | 134.7232 | 140.6416 | 154.4655 | 143.1633 | 129.843 | 182.5536 | 195.0817 | 203.4528 | 135.0127 | 148.0944 | 224.3584 | 203.2768 | 192.7105 | 186.8303 | 188.6225 | 324.1824 | 141.3648 | 138.7041 | 155.088 | 224.3584 | 144.2577 | 152.8993 | 130.1824 | 128.824 | 150.9311 | 159.1633 | 165.3648 | 131.176 | 114.7232 | 111.0817 | 99.6288 | 116.7232 | 218.7423 | 222.4592 | 242.1824 | 316.2959 | 234.3712 | 224.6479 | 250.4719 | 283.5536 | 131.2959 | 76.7232 | 70.6416 | 131.2959 | 107.8176 | 83.7359 | 115.0944 | 86.1951 | 113.6225 | 92.9056 | 94.2704 | 99.5472 | 66.0064 | 66.2704 | 63.0127 | 102.7169 | 110.8929 | 92.8112 | 91.5536 | 90.176 | 211.7359 | 211.088 | 238.7232 | 379.3839 | 145.1824 | 183 | 190.6479 | 327.8049 | 266.9936 | 174.9247 | 188.4464 | 363.7105
1533 | 18.8574 | 144.4691 | 185.1288 | 194.5928 | 166.7423 | 118.2733 | 151.0413 | 168.7217 | 157.1238 | 146.0825 | 125.3402 | 126.4897 | 155 | 204.2577 | 130.0413 | 153.9794 | 206.6598 | 170.0413 | 140.7267 | 141.3608 | 182.5979 | 143.6185 | 304.1495 | 310.9381 | 154.3196 | 207.4897 | 215.2783 | 226.1958 | 255.2114 | 110.2989 | 125.2577 | 117.0206 | 125.3196 | 119.3815 | 131.1907 | 139.1288 | 78.6135 | 221.4227 | 224.0206 | 261.7937 | 237.2733 | 139.5722 | 114.4434 | 131.7011 | 108.1907 | 74.8299 | 65.3402 | 72.768 | 120.9587 | 192.2783 | 216.1495 | 225.7886 | 142.763 | 133.2114 | 141.6392 | 125.7836 | 113.6804 | 143.7217 | 142.5259 | 198.2783 | 166.9381 | 131.8505 | 147.2989 | 147.6804 | 134.0619 | 149.3865 | 194.9381 | 210.4484 | 218.9381 | 158.3865 | 149.4897 | 222.6804 | 194.5309 | 184.7217 | 178.8505 | 193 | 315.7011 | 150.9587 | 136.0413 | 162.6392 | 220.1907 | 153.9794 | 141.8918 | 140.7886 | 112.0619 | 155.6185 | 155.0413 | 166.6598 | 116.4021 | 94.4227 | 94.9124 | 108.6392 | 104.7217 | 204.0413 | 226.8093 | 268.5566 | 338.8762 | 245.2783 | 230.1288 | 246.6598 | 320.8349 | 131.4691 | 77.1701 | 71.4897 | 131.4691 | 120.4484 | 100.0876 | 115.6598 | 108.5566 | 107.2114 | 95.6598 | 90.8505 | 115.5979 | 69.8093 | 60.3608 | 73.1082 | 96.0413 | 93.5928 | 86.7011 | 99.7886 | 74.5722 | 214.8093 | 203.7011 | 230.8712 | 389.5979 | 130.8918 | 189.6392 | 187 | 307.1032 | 278.1288 | 183.2783 | 210.0876 | 350.7423
1534 | 18.8684 | 139.2625 | 202.1593 | 211.0546 | 171.4219 | 147.5296 | 149.8422 | 155.8938 | 137.525 | 142.8422 | 100.5766 | 124.4734 | 139.8407 | 172.6282 | 141.475 | 146.5781 | 211.2109 | 144.4188 | 158.6328 | 146.7375 | 192.7375 | 146 | 307.5266 | 317.3687 | 162.5797 | 205.4734 | 212.1047 | 228.7891 | 288.3718 | 96.8407 | 100.3657 | 106.736 | 112.525 | 117.8422 | 112.3141 | 141 | 76.3687 | 195.1032 | 241.5281 | 220.6267 | 259.7921 | 135.1578 | 111.5266 | 130 | 97.736 | 84.264 | 76.7906 | 68.4203 | 113.736 | 207.6328 | 200.9985 | 214.525 | 136.3156 | 125.9469 | 137.9469 | 121.6844 | 109.6313 | 153.6328 | 171.7405 | 178.261 | 159.0516 | 136.0531 | 137.2094 | 168.8968 | 121.736 | 150.7891 | 168.5235 | 191.9454 | 205.9985 | 136.2079 | 148.3156 | 204.3141 | 198.0531 | 187.8953 | 188.9484 | 203.1062 | 328.3172 | 155.5266 | 147.475 | 153.0516 | 226.5797 | 128.0501 | 151.6328 | 132.8938 | 127.5281 | 177.3703 | 162.264 | 172.0531 | 133.2125 | 103.475 | 94.3687 | 99.0516 | 126.4234 | 222.2125 | 226.3156 | 241.1548 | 320.5766 | 226.9454 | 217.6828 | 240.2625 | 286.8907 | 138.0531 | 72.7891 | 77.0531 | 138.0531 | 101.1032 | 103.1578 | 107.5781 | 124.9484 | 107.5266 | 103.5797 | 106.8437 | 102.1563 | 68.4734 | 73.3172 | 79.3687 | 94.8422 | 89.8422 | 97.6328 | 97.7891 | 85.3172 | 240.4234 | 227.2656 | 215.525 | 350.8907 | 130.5266 | 175.8407 | 197.1062 | 326.4234 | 253.8922 | 211.2656 | 226.6328 | 357.1062
1535 | 18.8794 | 145.6903 | 201.5821 | 215 | 183.2537 | 146.3097 | 155.9813 | 166.8172 | 148.6716 | 153.2537 | 125.0522 | 131.6903 | 167.7612 | 184.7985 | 148.2724 | 168.0709 | 211.1455 | 145.1269 | 145.6194 | 148 | 202.5448 | 126.347 | 320.8172 | 283.8209 | 159.7276 | 198.1642 | 232.3619 | 215.1828 | 283.0373 | 104.2537 | 119.7798 | 126.3619 | 130.5074 | 135.944 | 137.1977 | 129.0373 | 90.6716 | 213.2164 | 254.3992 | 234.0709 | 244.4926 | 146.9627 | 114.5634 | 112.9105 | 132.7425 | 73.1828 | 73.0187 | 69.5634 | 102.6008 | 194.6194 | 191.1642 | 199.1828 | 141.1269 | 136.1082 | 139.5634 | 137.3806 | 113.2724 | 165.3992 | 164.4739 | 182.5448 | 161.2724 | 139.5634 | 131.5821 | 144.8023 | 95.2202 | 164.5261 | 168.9813 | 191.4179 | 225.2164 | 136.1269 | 142.0187 | 193.3097 | 177.6381 | 181.3097 | 200.3992 | 183.6381 | 304.5112 | 151.7276 | 156.8358 | 155.2724 | 221.1642 | 129.8358 | 150.5821 | 138.6903 | 134.4179 | 169.8918 | 168.2724 | 184.1082 | 134.4366 | 94.0373 | 112.944 | 111.5261 | 130.1455 | 232.8358 | 206.347 | 241.8358 | 318.5634 | 255.4701 | 207.3097 | 217.6381 | 282.4179 | 127.0373 | 59.1828 | 70.3097 | 127.0373 | 120.0709 | 88.4739 | 125.653 | 121.8731 | 108.8545 | 100.7276 | 115.9813 | 104.9813 | 71.4179 | 76.8545 | 68.8918 | 84.7463 | 104.5261 | 77.7836 | 92.7276 | 84.5821 | 222.7836 | 226.0187 | 210.4366 | 359.2351 | 118.1828 | 208.0335 | 193.0187 | 319.8918 | 274.6343 | 183.5299 | 212.7649 | 345.3284
1536 | 18.8904 | 143.5329 | 190.5988 | 220.2006 | 194.5344 | 146.7335 | 144.8653 | 171.6003 | 157.9341 | 151.5329 | 111.7979 | 120.8653 | 156.5314 | 204.4686 | 137.732 | 152.9311 | 220.5344 | 134.732 | 137.7994 | 158.4012 | 171.9296 | 142.0689 | 299.5973 | 306.6033 | 145.9985 | 188.3323 | 213.464 | 240.7366 | 248.0628 | 98.1991 | 112.732 | 116.9985 | 126.1991 | 129.4656 | 133.3323 | 141.7351 | 100.8009 | 243.8698 | 246.4656 | 247.5344 | 258.3354 | 109.9955 | 121.0674 | 131.6692 | 133.7994 | 71.8668 | 82.4012 | 92.536 | 105.3338 | 198.0674 | 169.1976 | 222.1363 | 136.7994 | 139.7335 | 141.7335 | 125.2649 | 122.6677 | 173.9341 | 186.2695 | 197.8683 | 162 | 155.6018 | 136.2006 | 118.3308 | 126.5375 | 145.3308 | 167.3997 | 210.2021 | 237.6677 | 145.6677 | 140.1332 | 205.8683 | 178.3338 | 217.271 | 187.2649 | 193.0015 | 303.4671 | 151 | 145.8653 | 161.2006 | 236.4686 | 127.5329 | 127.464 | 134.7994 | 145.4012 | 177.5344 | 157.732 | 181.6662 | 116.6646 | 99.8009 | 115.1332 | 106.1991 | 130 | 238.3338 | 251.539 | 218.7305 | 326.8009 | 221.9955 | 219.0015 | 252.1378 | 285.6003 | 121.5329 | 70.8683 | 74.2006 | 121.5329 | 114.4656 | 105.9357 | 109.0643 | 120.1332 | 95.9985 | 105.2006 | 104.8653 | 97.3323 | 65.9326 | 55.3308 | 70.4671 | 101.2021 | 120.0015 | 92.2021 | 88.5329 | 77.9326 | 261.4716 | 205.9311 | 211.7335 | 373.268 | 141.1363 | 205.3323 | 174.6646 | 310.1991 | 283.3338 | 210.9372 | 180.5299 | 334.3323
1537 | 18.9014 | 153.5487 | 192.5162 | 213.0885 | 179.2979 | 158.4278 | 156.1859 | 142.9911 | 153.7256 | 154.5162 | 123.944 | 125.1534 | 155.7581 | 204.3628 | 143.9115 | 156.1534 | 207.056 | 176.0739 | 154.5812 | 158.2419 | 184.5812 | 165.2184 | 288.0885 | 329.4602 | 161.5812 | 215.1299 | 224.944 | 248.5162 | 279.0414 | 104.0325 | 122.4278 | 114.1209 | 146.9765 | 128.8791 | 140.7906 | 133.4513 | 75.6282 | 233.056 | 266.0974 | 244.6047 | 239.9026 | 131.2508 | 107.056 | 135.8791 | 133 | 66.7256 | 76.0885 | 89.8466 | 95.4513 | 190.2094 | 189.7346 | 195.2329 | 149.1859 | 137.3628 | 126.1769 | 130.0325 | 108.1769 | 182.9115 | 170.6607 | 189.4513 | 147.9351 | 143.056 | 144.0325 | 142.2508 | 129.3628 | 155.1859 | 198.6461 | 191.9026 | 218.7816 | 142.6047 | 144.3953 | 211.5162 | 171.9675 | 181.6842 | 206.0974 | 194.1209 | 340.9205 | 172.9765 | 156.3068 | 164.6372 | 226.6932 | 134.0325 | 136.3068 | 125.2094 | 177.7671 | 161.4188 | 176.2184 | 173.9675 | 129.8231 | 105.3953 | 98.2979 | 114.6696 | 122.9675 | 234.6047 | 242.2979 | 243.1299 | 328.8791 | 217.7581 | 231.5487 | 261.5162 | 275.4513 | 129.7906 | 73 | 73.2419 | 129.7906 | 107.7256 | 101.0885 | 120.0649 | 111.2094 | 110.7021 | 99.8466 | 107.3953 | 104.7906 | 73.7906 | 52.8791 | 78.0325 | 89.9351 | 88.5958 | 89.7256 | 109.9765 | 84.0325 | 209.9821 | 221.4602 | 236.6136 | 390.8231 | 141.4838 | 211.9115 | 196.6136 | 305.4838 | 218.0706 | 212.4838 | 200.6136 | 327.7256
1538 | 18.9124 | 143.4125 | 187.6519 | 208.4346 | 170.7606 | 156.4346 | 160.674 | 159.501 | 153.8913 | 141.6298 | 121.5433 | 121.5433 | 153.326 | 193.3038 | 152.1308 | 169.4789 | 221.0443 | 160.6076 | 157 | 176.7183 | 190.5654 | 157.3038 | 319.9799 | 329.326 | 152.4125 | 187.8028 | 195.8028 | 234.7384 | 287.5654 | 105 | 121.326 | 103.3038 | 119.6941 | 139.6962 | 133.0865 | 132 | 87.1529 | 227.4346 | 238.6941 | 264.501 | 234.326 | 129.6519 | 133.5232 | 127.0865 | 141.0221 | 75.8048 | 76.7827 | 84.5433 | 98.4567 | 210.3924 | 206.3702 | 184.7606 | 151.8913 | 128.9779 | 168.5675 | 140.8048 | 117.5875 | 166.173 | 163.5433 | 195.1308 | 154.9135 | 157.0443 | 132.5211 | 150.4567 | 158.4145 | 154.326 | 195.8692 | 209.501 | 232.0443 | 142.8913 | 138.7606 | 233.3924 | 169.2173 | 174.2173 | 214.3481 | 208.2616 | 319.2595 | 159.0644 | 158 | 152.5211 | 208.0644 | 113.6076 | 146.0221 | 129.3481 | 140.9979 | 161.674 | 154.0422 | 154.2817 | 123.0865 | 106 | 128.9799 | 113.326 | 110.4125 | 239.3481 | 221.2817 | 227.3903 | 323.6519 | 226.0221 | 222.3038 | 231.6941 | 303.4145 | 140.8048 | 59.6298 | 75.674 | 140.8048 | 100.7606 | 102.674 | 116.6519 | 115.3481 | 97.8471 | 108.8048 | 102.6519 | 113.1308 | 70.5433 | 61.0221 | 71.8692 | 77.3038 | 96.4789 | 87.2173 | 112.1087 | 76.0865 | 239.4367 | 224 | 221.2817 | 376.0644 | 141.8913 | 197.8471 | 194.6519 | 325.501 | 235.7405 | 176.346 | 202.2173 | 327.8913
1539 | 18.9234 | 134.7709 | 222.242 | 209.8073 | 205.242 | 141.5417 | 152.8672 | 133.0835 | 144.0599 | 165.3019 | 111.9636 | 111.0599 | 152.0964 | 162.1798 | 149.3854 | 161.9636 | 222.0964 | 166.1328 | 149.7709 | 165.4454 | 181.0599 | 134.3126 | 338.4583 | 345.2655 | 146.4818 | 216.5311 | 215.4947 | 230.2891 | 277.1563 | 116.7473 | 102.9272 | 116.4583 | 134.9765 | 124.7345 | 114.8308 | 118.4454 | 92.6146 | 219.7709 | 243.1328 | 234.4689 | 256.591 | 128.0964 | 108.0835 | 119.6745 | 126.6381 | 82.4218 | 63.4454 | 93.0364 | 103.5182 | 184.9871 | 177.2762 | 217.4347 | 170.0728 | 144.2655 | 140.5653 | 114.8908 | 116.2891 | 161.2891 | 151.2527 | 185.1563 | 149.6745 | 130.9871 | 141.8437 | 162.7473 | 141.2163 | 155.8073 | 164.2762 | 207.4818 | 236.7109 | 145.7109 | 151.5546 | 237.8073 | 175.3255 | 176.7109 | 182.4689 | 208.1927 | 325.0364 | 156.0964 | 163.4218 | 153.7109 | 220.4583 | 111.9036 | 136.1563 | 157.1092 | 127.8672 | 174.651 | 186.242 | 180.0129 | 113.8672 | 121.3619 | 95.0471 | 122.0364 | 116.2291 | 256.2655 | 244.3019 | 215.0599 | 318.4818 | 227.9036 | 208.349 | 246.0728 | 308.8073 | 147.4218 | 57.0964 | 57.9272 | 147.4218 | 101.8073 | 112.0364 | 112.3854 | 119.6146 | 96 | 84.6981 | 120.9765 | 97.7345 | 71.8073 | 63.8073 | 71.9036 | 94.9765 | 94.3854 | 88.8073 | 110.1927 | 76.8073 | 222.3126 | 215.8672 | 239.7837 | 383.9401 | 145.6146 | 189.6745 | 220.2056 | 333.4218 | 286.893 | 182.8437 | 209.2291 | 332.5182
1540 | 18.9344 | 153.2346 | 200.3539 | 234.7302 | 170.5309 | 174.8054 | 152 | 156.562 | 132.0088 | 163.4203 | 117.4115 | 137.478 | 145.5885 | 202.9647 | 136.177 | 128.0264 | 207.3451 | 177.9912 | 144.4203 | 128.2787 | 190.9912 | 155.8142 | 297.8671 | 323.1858 | 161.5708 | 232.823 | 186.8583 | 238.2434 | 275.0841 | 118 | 118.4027 | 126.2434 | 138.8319 | 155.0576 | 110.2522 | 131.6549 | 85.6725 | 217.1681 | 239.4203 | 244.739 | 266.3275 | 120.6725 | 129.7302 | 118.0841 | 121.3363 | 77.5044 | 72.0753 | 72.0176 | 106.7478 | 193.9071 | 179.4956 | 211.8407 | 172 | 137.7566 | 110.438 | 126.6549 | 100.4292 | 148.177 | 154.5797 | 216.9736 | 157.2434 | 151.8142 | 114.6061 | 146.5973 | 143.5797 | 151.4203 | 168.3275 | 204.2522 | 243.4115 | 155.1593 | 161.2434 | 227.0088 | 219.9647 | 190.739 | 186.3275 | 180.522 | 317.7566 | 160.5797 | 144.7654 | 143.0088 | 234.823 | 131.2346 | 122.177 | 137.1017 | 135.2434 | 152.1858 | 178.0929 | 181.1681 | 134.9824 | 89.1105 | 119.3939 | 104.6814 | 135.3186 | 263.4956 | 222.2698 | 245.1417 | 325.3275 | 206.0176 | 186.8495 | 224.1858 | 316.3275 | 143.4203 | 67.9912 | 56 | 143.4203 | 108.4115 | 109.3363 | 112 | 105.3451 | 107.9071 | 86.5797 | 114.7566 | 92.3363 | 55.5132 | 65.8319 | 71.0841 | 91.5044 | 90.3363 | 99.0753 | 98.0929 | 73.3363 | 226.4115 | 236.9824 | 234.6725 | 368.5132 | 125.8495 | 191.7478 | 187.2787 | 342.2434 | 263.6061 | 186.7478 | 184.3539 | 337.5797
1541 | 18.9454 | 162.4258 | 201.7129 | 210.0815 | 178.1387 | 141.7993 | 171.4927 | 157.1436 | 149.5645 | 171.354 | 131.9233 | 116.7944 | 148.7129 | 188.4355 | 142.4258 | 168.6265 | 209.7129 | 160.4355 | 134.7178 | 179.7652 | 204.9951 | 177.4927 | 300.4976 | 330.2822 | 164.8564 | 225.646 | 209.9902 | 230.646 | 285.2105 | 110.5742 | 136.708 | 109.3638 | 132.5024 | 143.1484 | 106.2871 | 126.5024 | 84.0718 | 208.646 | 237.1436 | 241.3589 | 267.9282 | 114.4307 | 132.9282 | 113.3589 | 120.0718 | 64.9331 | 65.5742 | 77.4258 | 102.3589 | 181.0767 | 198.5645 | 229.5645 | 140.4404 | 144.4258 | 125.6362 | 115.0049 | 112.9233 | 161.8516 | 139.2202 | 231.1387 | 167.2822 | 143.7895 | 127.7798 | 139.4307 | 131.9331 | 154.7129 | 200.5596 | 210.4976 | 231.9331 | 162.4976 | 149.9331 | 232.4976 | 223.0718 | 200.354 | 213.9185 | 209.5596 | 319.7847 | 191.6314 | 147.6411 | 157.7798 | 233.2153 | 135.7847 | 135.8516 | 133.1436 | 158.2774 | 172.2774 | 182.5693 | 179.1436 | 159.2774 | 109.2056 | 97.8662 | 115.9951 | 118.4355 | 228.7275 | 241.3491 | 256.354 | 358.4878 | 202.1436 | 214.7031 | 211.7895 | 316.0718 | 136.5024 | 68.0718 | 78.2774 | 136.5024 | 110.8564 | 99.7178 | 117.5693 | 115.1387 | 96.0049 | 110.2056 | 99.1484 | 108.708 | 89.2725 | 79.9233 | 78.4258 | 104.9233 | 103.9233 | 83.292 | 91.4307 | 84.1387 | 228.8564 | 216.7226 | 246.0669 | 379.9951 | 124 | 182.7178 | 196.0669 | 332.7895 | 261 | 191.6411 | 198.708 | 344.4976
1542 | 18.9564 | 140.4277 | 230.2154 | 225.8698 | 179 | 139 | 151.3682 | 158.881 | 155.7026 | 159.7733 | 116.0707 | 146.9775 | 144.2974 | 197.3457 | 149.5836 | 147.5466 | 232.5723 | 190.9775 | 122.7138 | 160.4871 | 180.6061 | 177.119 | 279.3682 | 326.2974 | 167.8215 | 209.9518 | 219.5241 | 213.0707 | 300.1077 | 123.1672 | 104.1415 | 130.5723 | 132.9405 | 150.4646 | 121.0482 | 159.8585 | 83.0595 | 239.9775 | 243.5836 | 233.4759 | 261.4164 | 140.3344 | 117.9518 | 138.3939 | 114.3569 | 74.3457 | 83.8103 | 73.2974 | 105.7621 | 207.2749 | 191.5354 | 214.0707 | 129.5354 | 140.2974 | 132.6431 | 143.1559 | 107.4164 | 147.9518 | 162.4534 | 218.8328 | 176.4646 | 136.4164 | 125.2379 | 154.0482 | 136.6431 | 155 | 174.7846 | 203.4759 | 223.4759 | 167.7026 | 155.5836 | 262.1559 | 181.6174 | 160.5579 | 189.6656 | 194.1302 | 317.1785 | 160.1415 | 139.5354 | 154.2974 | 208.5466 | 119.0707 | 137 | 124.5354 | 144.9518 | 179.6431 | 200.8698 | 180.881 | 125.2605 | 108.1785 | 100.7026 | 103.8328 | 148.9775 | 230.7026 | 246.7621 | 235.3682 | 320.5579 | 217.0482 | 218.881 | 235.4534 | 321.6431 | 132.2379 | 62.3569 | 68.7138 | 132.2379 | 119.4646 | 98.0595 | 102.9518 | 111.2974 | 109.1077 | 113.881 | 114.9293 | 111.881 | 69.4277 | 75.3569 | 62.0707 | 81.5466 | 88.0707 | 85.7621 | 112.6318 | 69.9518 | 228.0595 | 201.8328 | 228.1897 | 364.0707 | 114.5949 | 157.5466 | 198.881 | 361.1559 | 295.799 | 197.6431 | 214.1077 | 345.9405
1543 | 18.9674 | 144.7168 | 216.7551 | 211.7551 | 192.3393 | 141.8584 | 146.1888 | 168.5281 | 176.9617 | 151.3776 | 145.4898 | 128.0383 | 147.8112 | 178.9439 | 118.5574 | 142.1888 | 237.8112 | 181.5663 | 147.7257 | 126.6046 | 175.1888 | 154.1327 | 312.301 | 324.0944 | 168 | 194.7079 | 260.9706 | 228.1977 | 291.4719 | 128.764 | 108.6696 | 127.236 | 151.1033 | 125.2743 | 115.3304 | 126.7462 | 90.6224 | 244.8584 | 255.4337 | 259.6785 | 264.8112 | 117.2271 | 121.764 | 131.4248 | 124.4809 | 67.3776 | 74.5191 | 86.3393 | 95.5191 | 185.1799 | 212.9145 | 207.2832 | 149.9617 | 116.1799 | 140.6224 | 127.8495 | 128.9145 | 160.3393 | 155.4248 | 196.0855 | 159.8495 | 130.2832 | 124.0472 | 175.9617 | 130.3304 | 171.1977 | 180.6224 | 223.0089 | 239.1977 | 152.7551 | 146.4719 | 224.935 | 183.764 | 204.6875 | 177.5191 | 193.9528 | 319.8584 | 162.764 | 157.1033 | 172.1033 | 218.4337 | 143.7257 | 142.7168 | 158.301 | 162.1033 | 180.9528 | 177.2271 | 183.8584 | 145.8673 | 108 | 105.764 | 115.3865 | 129.0855 | 254.8201 | 269.8673 | 209.2271 | 318.9528 | 236.1033 | 224.7168 | 245.5752 | 326.764 | 133.9056 | 70.5752 | 67.0472 | 133.9056 | 111.4248 | 101.8112 | 118.1977 | 130.0561 | 108.0944 | 100.6607 | 105.5191 | 112 | 69.9056 | 68.3304 | 82.9145 | 89.5281 | 113.6785 | 98.3865 | 99.7079 | 76.6224 | 219.4248 | 248.6403 | 261.301 | 348.7079 | 145.4426 | 174.1033 | 165.6518 | 351.5663 | 275.1327 | 202.764 | 216.9056 | 368.8673
1544 | 18.9784 | 151.7557 | 198.6283 | 222.5812 | 182.3839 | 140.0698 | 159.5113 | 171.8953 | 150.0122 | 139.4188 | 127.6981 | 127 | 125.8028 | 227.2199 | 154.6388 | 169.0227 | 217.733 | 164.5934 | 170.2321 | 138.5113 | 175 | 157.8255 | 318.8255 | 354.8831 | 171.8604 | 207.5113 | 215.7103 | 230.9302 | 296.7906 | 112.5934 | 117.6859 | 137.6161 | 155.8604 | 130.7557 | 104.3839 | 123.0698 | 92.9302 | 237.2792 | 242.4887 | 261 | 230.2565 | 129.5113 | 135.5113 | 131.9651 | 117.2792 | 57.349 | 73.0349 | 58.0471 | 113.3368 | 188.8255 | 198.5585 | 225.3368 | 162.5812 | 137.1972 | 124.5934 | 131.8255 | 110.6981 | 159.0698 | 176.2321 | 193.0698 | 161.8953 | 121.3141 | 151.9878 | 153.8377 | 149.3019 | 147.8726 | 185.8255 | 215.3141 | 233.2443 | 147.1745 | 160.4764 | 229.7557 | 195.5812 | 189.6283 | 177 | 200.7557 | 313.2443 | 186.1623 | 120.3612 | 156.5934 | 217.0698 | 134.3839 | 136.2443 | 134.9075 | 136.9424 | 190.651 | 176 | 178.2094 | 146.0349 | 101.2443 | 104.0698 | 114.0698 | 108.6981 | 224.1518 | 229.5009 | 272.6614 | 312.2443 | 256.3019 | 211.4887 | 244.0698 | 302.8726 | 134.9651 | 58.4537 | 89.1972 | 134.9651 | 114.8604 | 109.7208 | 109.349 | 127.1396 | 107.0349 | 88.4188 | 109.8255 | 112.9651 | 77.7208 | 53.5236 | 70.4887 | 76.4887 | 89.9075 | 90.3141 | 105.7557 | 74.1047 | 227.6859 | 225.9075 | 203.164 | 375.0227 | 141.2094 | 165.349 | 191.9878 | 336.5236 | 288.4764 | 180.8028 | 183.2216 | 340.082
1545 | 18.9894 | 127.5653 | 167.701 | 224.9548 | 153.6558 | 135.1131 | 155.1131 | 176.8869 | 148.0226 | 158.5477 | 119.1809 | 117.2261 | 168.005 | 206.5201 | 123.7462 | 156.3166 | 205.2714 | 144.4523 | 134.8367 | 144.8643 | 176.9548 | 150.1809 | 328.7739 | 320.8141 | 173.9548 | 217.7739 | 238.4347 | 199.7236 | 289.1809 | 125.6834 | 102.3618 | 127.2487 | 142.3166 | 145.6608 | 147.005 | 148.4121 | 73.4523 | 214.5201 | 260.5704 | 262.9548 | 234.8643 | 121.2035 | 103.7462 | 134.9322 | 125.7965 | 84.3668 | 78.8643 | 90.2312 | 120.8417 | 196.8191 | 202.8869 | 224.0452 | 137.5879 | 145.8191 | 147.4573 | 133.9548 | 132.4799 | 164.8643 | 173.0905 | 197.8869 | 152.2261 | 126.8643 | 132.4749 | 154.9548 | 124.5879 | 168.5025 | 182.0905 | 220.8643 | 218.3392 | 147.9774 | 155.1357 | 229.0226 | 201.8643 | 205.6156 | 221.9598 | 205.8869 | 332.5477 | 170.3844 | 175.6885 | 150.1357 | 209.1809 | 156.4799 | 142.8417 | 137.9095 | 151.6382 | 160.701 | 169.1583 | 169.2035 | 119.6105 | 93.1809 | 119.6382 | 93.4749 | 110.9322 | 212.2487 | 231.9095 | 223.1985 | 345.2312 | 241.3618 | 243.2538 | 261.593 | 306.8869 | 142.8191 | 66.7965 | 76.3166 | 142.8191 | 107.1809 | 90.4523 | 140.2764 | 103.5427 | 94.294 | 116.3442 | 93.3844 | 102.2487 | 78 | 74.5025 | 82.706 | 79.9095 | 98.7739 | 101.7286 | 106.9774 | 76.9322 | 226.0452 | 260.1859 | 241.0729 | 400.4347 | 122.4296 | 185.5251 | 210.593 | 347.7286 | 275.3166 | 207.3668 | 198.6156 | 319.4523
1546 | 19.0004 | 139.8658 | 194.711 | 204.2168 | 168.8349 | 128.0723 | 152.031 | 192.8349 | 171.7523 | 141.1858 | 121.969 | 125.9071 | 174.9381 | 188.1858 | 164.5665 | 152.0413 | 207.969 | 197.4426 | 125.0929 | 154.8968 | 174.031 | 158.9071 | 320.0929 | 332.8658 | 146.289 | 210.0826 | 205.3509 | 218.7936 | 294.9381 | 112.1445 | 108.9277 | 115.1239 | 140.0206 | 142.0413 | 134.1445 | 125.2477 | 89.8245 | 237.7523 | 240.2168 | 263 | 246.8761 | 129.9071 | 128.7316 | 115.2064 | 115.1135 | 82.031 | 58.2168 | 54.3819 | 95.2684 | 204.9174 | 180.2374 | 224 | 140.9587 | 140.0619 | 147.0103 | 152.8039 | 119.1445 | 189.742 | 151.2271 | 209.8761 | 163.8761 | 133.9277 | 137.9381 | 169.8452 | 122.0206 | 147.2271 | 183.9794 | 228.9174 | 221.9587 | 184.6181 | 141.1445 | 214.1548 | 204.969 | 204.0206 | 202.2168 | 174.3303 | 379.5149 | 150.2064 | 141.3716 | 151.9794 | 220.8761 | 127.3097 | 132.1135 | 152.8452 | 148.0413 | 150.1032 | 169 | 172.9587 | 113.0619 | 112.7936 | 96.2477 | 101.9071 | 131.7832 | 207.0516 | 226.0619 | 214.0826 | 321.258 | 200.4232 | 217.2787 | 258.0413 | 290.1755 | 149.9277 | 64.031 | 74.0206 | 149.9277 | 117.8865 | 103.8555 | 122.1961 | 126.7523 | 96.969 | 99.1858 | 97.9484 | 109.9174 | 74.0413 | 69.0619 | 82.0103 | 86.9277 | 90.0929 | 94.0826 | 111.9484 | 87.8865 | 229.9587 | 236.258 | 228.1445 | 378.2374 | 129.9174 | 188.969 | 203.0826 | 362.8452 | 239.3716 | 188.2064 | 210.8761 | 341.7626
1547 | 19.0114 | 150.9862 | 183.002 | 245.8877 | 189.0177 | 155.002 | 145.0354 | 151.0414 | 153.9921 | 141.0571 | 132.9567 | 113.0158 | 154.9862 | 222.9823 | 142.9882 | 155.9586 | 198.0453 | 188.998 | 146.0098 | 158.0039 | 183.9882 | 154.0118 | 295.0729 | 341.002 | 169.0217 | 214.002 | 246.002 | 250.0138 | 299.0394 | 128.0276 | 128.002 | 108.0394 | 135.0079 | 120.0473 | 143.0098 | 126.0591 | 56.0748 | 230.0197 | 248.9567 | 246.0354 | 220.0867 | 128.0354 | 119.0236 | 135.9902 | 116 | 85.9764 | 68.9764 | 66.0039 | 128.935 | 193.9862 | 186.9862 | 211.0039 | 159.9961 | 158.9626 | 146.9527 | 154.9705 | 127.9842 | 144.0256 | 151.0138 | 195.0886 | 174.0039 | 126.998 | 153.9449 | 173.9685 | 116.0847 | 132.0118 | 186.0118 | 192.0492 | 232.0335 | 147.0492 | 147.9842 | 216.0512 | 211.9783 | 176.0374 | 197.9724 | 196.0197 | 320.0177 | 165.9508 | 158.9882 | 166.9902 | 224 | 162.9291 | 130.9783 | 133.9665 | 171.9705 | 184.9646 | 171.0059 | 210.9606 | 106.0354 | 110.0039 | 87.0256 | 100.0177 | 112.0197 | 214.0098 | 255.9626 | 238.002 | 354.9882 | 233.9724 | 222.0374 | 208.063 | 303.0532 | 130.0571 | 76.0433 | 72.0098 | 130.0571 | 105.0138 | 104.0079 | 124.0217 | 102.0236 | 89.0315 | 94.0039 | 109.9724 | 112.0118 | 58.0354 | 66.0256 | 84.9449 | 87.0335 | 111.9389 | 103.0335 | 110.9527 | 77.9842 | 231.9212 | 258.9567 | 260.8956 | 405.8877 | 122.0453 | 191.9882 | 184.0177 | 327.065 | 278.9685 | 185.998 | 198.0276 | 371.9567
1548 | 19.0224 | 143.9572 | 183.9715 | 189.6275 | 197.8859 | 155.7148 | 163.0285 | 171.7433 | 150.0285 | 169.8146 | 111.0856 | 120.9857 | 148.1711 | 213.5151 | 137.5847 | 135.0285 | 220.8859 | 187.6007 | 150.8574 | 160 | 178.2282 | 159.8004 | 331.5151 | 341.7576 | 179.9715 | 214.8289 | 246.9144 | 256.6292 | 318.3868 | 141.6435 | 128.9572 | 127.9144 | 139 | 144.1711 | 147.6007 | 155.8146 | 93.8004 | 239.8004 | 227.4278 | 264 | 263.7005 | 145.4724 | 130.6577 | 131.2709 | 116.0428 | 73.8574 | 57.3708 | 68.0713 | 96.1996 | 186.9144 | 180.328 | 213.2852 | 158.3565 | 139.8431 | 122.9715 | 139.7291 | 119.8859 | 157.1141 | 158.4849 | 239.7148 | 175.7718 | 126.1426 | 126.1569 | 157.8431 | 158.9002 | 138.1996 | 191.8717 | 216.6863 | 248.672 | 171.3725 | 140.2852 | 241.615 | 201.2424 | 194.943 | 184.2424 | 206.0856 | 329.3137 | 141.2709 | 153.0285 | 162.0428 | 224.6702 | 127.3708 | 120.385 | 117.385 | 156.7291 | 166.8574 | 174 | 190.8431 | 124.0856 | 112.0713 | 100.0285 | 108.8289 | 122.3565 | 219.3565 | 237.0143 | 239.0998 | 348.5864 | 220.3708 | 240.5009 | 240.1569 | 329.9287 | 158.9002 | 97.5437 | 76.8717 | 158.9002 | 112.1426 | 108.0713 | 134.5437 | 114 | 105.1996 | 95.8574 | 95.9857 | 117.7861 | 75.9287 | 78.9002 | 57.3137 | 103.9572 | 81.2995 | 119.3868 | 87.1711 | 70.0143 | 192.4991 | 237.5134 | 208.6845 | 349.6417 | 144.9857 | 186.4278 | 193.1141 | 359.3583 | 263.2567 | 185.4136 | 212.057 | 350.1996
1549 | 19.0334 | 141.0797 | 181.8672 | 232.4159 | 189.7876 | 136.1328 | 164.8407 | 154.239 | 151.6283 | 156.8672 | 117.531 | 119.7079 | 160.0531 | 180.5576 | 177.4159 | 136.5486 | 213.1593 | 160.2921 | 141.7169 | 160.1328 | 193.5221 | 146.6372 | 298.6372 | 325.5576 | 177.761 | 203.3717 | 239.8052 | 231.1062 | 276.9824 | 117.1859 | 126.2124 | 122.0266 | 138.7876 | 156.239 | 120.6638 | 142.761 | 80.2921 | 226.2124 | 255.9114 | 263.8407 | 243.6638 | 109.7169 | 107.5045 | 149.469 | 119.4248 | 64.3983 | 82.4159 | 73.1593 | 109.9734 | 181.1328 | 202.3893 | 232.2566 | 183.0531 | 128.8407 | 121.7169 | 121.2655 | 111.6814 | 165.2655 | 191.0972 | 219.7876 | 159.9734 | 136.1593 | 137 | 147.6903 | 152.1859 | 151.6283 | 182.9734 | 195.6107 | 225.761 | 128.8231 | 159.761 | 214.9734 | 216.9379 | 190.8672 | 200.8938 | 213.3276 | 350.4424 | 159.9734 | 155.5576 | 164.2831 | 269.1148 | 152.4424 | 146.5221 | 144.1062 | 138.7434 | 157.1328 | 173.9734 | 179.6814 | 129.761 | 117.1328 | 102.0266 | 97.6107 | 145.8583 | 243.6814 | 237.5221 | 245.9203 | 320.3983 | 245.4424 | 206.4248 | 249.8848 | 325.3452 | 152.1328 | 66.6107 | 67.7876 | 152.1328 | 121.7079 | 112.761 | 103.1062 | 114.1328 | 118.3097 | 87.0886 | 94.761 | 103.6638 | 71.3717 | 71.6548 | 78.469 | 100.7876 | 101.8672 | 77.3452 | 98.3362 | 71.5045 | 226.7079 | 272.0707 | 255.2831 | 394.5841 | 143.8938 | 215.23 | 201.1593 | 316.1683 | 280.4955 | 213.7345 | 215.6283 | 363.7345
1550 | 19.0444 | 143.262 | 177.8934 | 211.7769 | 181.8835 | 141.2331 | 158.7669 | 163.3107 | 138.6603 | 151.5339 | 136.4562 | 110.3983 | 161.8058 | 200.9612 | 155.6504 | 121.3206 | 219 | 170.8835 | 167.0289 | 164.6504 | 177.2818 | 169.6504 | 322.4273 | 344.9901 | 168.8835 | 217.7769 | 195.9223 | 235.6215 | 312.7281 | 124 | 133.3008 | 122.9223 | 131.3496 | 164.3397 | 144.8835 | 134.1165 | 90.6893 | 234.8934 | 216.5438 | 257.6116 | 268.8545 | 135.0678 | 125.5727 | 129.7281 | 135 | 79.5438 | 61.738 | 78.7669 | 109.6603 | 185.7281 | 180.5438 | 204.2231 | 184.1066 | 123.1554 | 147.3008 | 130.9223 | 100.3107 | 174.5339 | 159.243 | 212.4273 | 159.0777 | 142.2719 | 136.6504 | 171.5628 | 158.3397 | 140.0587 | 182 | 218.5438 | 217.1942 | 157.9513 | 150.8835 | 214.3107 | 178.6603 | 186.8157 | 198.0876 | 260.3298 | 330.9322 | 159.4273 | 175.1455 | 139.2041 | 202.2529 | 131.9223 | 129.2331 | 147.6504 | 165.0289 | 162.738 | 173.0388 | 168.9322 | 121.2719 | 121.1066 | 103.4273 | 119.6504 | 104.6603 | 231.5339 | 221.6702 | 242.1066 | 335.8545 | 224.495 | 222.4273 | 208.8446 | 337.0289 | 155.7959 | 88.2231 | 59.7669 | 155.7959 | 110.4562 | 104.5438 | 107.505 | 118.8835 | 93.6992 | 125.8736 | 86.6603 | 127.2231 | 84.5339 | 59.2719 | 59.3884 | 93.1554 | 97.6215 | 90.7769 | 74.8934 | 89.6116 | 217.8256 | 237.1843 | 230.6702 | 415.3008 | 140.6992 | 186.9612 | 207.505 | 358.8835 | 262.1165 | 203.2231 | 203.0487 | 354.3496
1551 | 19.0554 | 125.6136 | 200.3579 | 229.8239 | 178.4375 | 147.1534 | 154.7897 | 170.7955 | 155.2045 | 140.8693 | 121.9262 | 143.0569 | 156.4887 | 200.4602 | 147.5625 | 154.0511 | 221.1987 | 168.5625 | 143.767 | 155.8466 | 207.3126 | 160.8977 | 331.4149 | 321.4317 | 166.4602 | 236.3353 | 194.8693 | 249.5683 | 306 | 124.3068 | 115.2841 | 121.0511 | 139.233 | 147.7955 | 141.7955 | 136.6932 | 83.9715 | 256.0285 | 229.0796 | 247.0285 | 289.3864 | 113.0227 | 116.4829 | 123.0511 | 134.7955 | 92.233 | 79.6421 | 73.8181 | 125.4887 | 179.5625 | 194.8693 | 188.2725 | 162.1534 | 127.3579 | 130.9204 | 128.3353 | 108.3068 | 163.1023 | 188.3637 | 223.4602 | 160.6421 | 149.7159 | 129.6363 | 136.8693 | 141.8977 | 189.6705 | 183.0227 | 231.0285 | 222.5113 | 135.1703 | 149.3806 | 222 | 197.1987 | 205.466 | 224.0796 | 217.5683 | 354.3579 | 169.8977 | 153.3864 | 169.1534 | 257.233 | 128.5171 | 135.1023 | 139.3579 | 141.2557 | 180.8977 | 173.7443 | 190.9773 | 128.8693 | 100.2783 | 113.1819 | 110.5398 | 121.2045 | 221.8919 | 261.7728 | 222.0453 | 356.233 | 211.9489 | 231.6705 | 206.4317 | 312.7955 | 127.1761 | 68.7955 | 55.3806 | 127.1761 | 98.5851 | 117.6932 | 119.3353 | 115.8977 | 111.767 | 99.5851 | 103.5113 | 108.8693 | 73.4602 | 66.8181 | 69.1023 | 97.767 | 112.5909 | 110.4602 | 96.7955 | 79.9489 | 261.8751 | 217.7159 | 269.9035 | 398.9204 | 157.0796 | 184.8751 | 217.6479 | 355.5398 | 264.5909 | 184.0511 | 229.2557 | 361.5683
1552 | 19.0664 | 138.0148 | 205.4778 | 209.7759 | 168.2537 | 150.3171 | 186.5412 | 167.8245 | 159.1268 | 157.3805 | 103.0148 | 108.7125 | 149.3467 | 209.5708 | 157.7463 | 156.2051 | 260.2875 | 177.8584 | 156.4144 | 152.6195 | 177.0148 | 160.0148 | 306.0592 | 346.351 | 174.7463 | 223.4926 | 212.2051 | 224.2051 | 306.5074 | 130.0634 | 102.5074 | 122.3805 | 125.8879 | 144.6342 | 137.4292 | 131.6977 | 100.5412 | 239.5222 | 213.8393 | 229.1903 | 277.4926 | 131.6195 | 141.907 | 124.6342 | 131.1268 | 78.0634 | 72.9366 | 88.6829 | 117.0782 | 190.2537 | 209.9218 | 247.0338 | 165.444 | 132.4778 | 147.4292 | 117.9662 | 114.8879 | 164.8732 | 158.9514 | 230.7315 | 154.5708 | 161.5412 | 158.0338 | 154.2051 | 138.9218 | 165.3805 | 201.2389 | 214.649 | 231.556 | 192.6047 | 172.463 | 220.5412 | 236.7315 | 178.7125 | 207.3171 | 192.7125 | 359.9852 | 166.8584 | 142.5708 | 171.6829 | 241.4778 | 102.5222 | 137.6342 | 146.5708 | 146.5074 | 179.444 | 168.3023 | 171.9366 | 143.0973 | 123.9366 | 98.8245 | 102.444 | 124.9366 | 255.9852 | 238.8732 | 257.463 | 339.9704 | 210.7463 | 205.4778 | 232.6195 | 308.6195 | 148.1755 | 65.6977 | 80.4926 | 148.1755 | 127.4292 | 112.0634 | 108.4588 | 115.0782 | 124.6047 | 127.0973 | 113.3805 | 124.556 | 81.556 | 80.4778 | 71.1903 | 111.3023 | 105.0634 | 117.4778 | 94.0148 | 79.5074 | 240.7463 | 231.8245 | 232.7759 | 413.3362 | 139.8732 | 165.5856 | 177.0444 | 347.3171 | 260.4884 | 185.3171 | 231.1459 | 334.3023
1553 | 19.0774 | 151.94 | 186.1043 | 234.6371 | 173.06 | 153.8643 | 165.4543 | 178.5614 | 162.8929 | 165.0443 | 117.3186 | 133.6971 | 183.0914 | 217.3943 | 153.5457 | 173.3186 | 234.47 | 163.4986 | 136.0286 | 148.1357 | 191.47 | 172.4257 | 363.5771 | 324.12 | 171.8329 | 219.1043 | 229.47 | 241.6214 | 313.6214 | 129.8643 | 109.6214 | 128.1514 | 137.7128 | 153.1671 | 127.5614 | 143.5143 | 79.6814 | 261.0914 | 239.3343 | 234.4229 | 269.2428 | 128.2715 | 113.3472 | 133.6214 | 132.0157 | 78.5457 | 70.6371 | 82.7885 | 132.2428 | 195.1357 | 193.5457 | 218.9843 | 171.1671 | 110.1514 | 139.3786 | 147 | 127.9243 | 163.3786 | 173.53 | 213.06 | 163.9843 | 139.2428 | 129.2271 | 170.94 | 123.9086 | 168.2742 | 190.3786 | 238.3943 | 226.59 | 172.1514 | 135.2271 | 203.3158 | 218.4543 | 206.1958 | 211.47 | 219.59 | 346.3629 | 150.8329 | 152.59 | 168.7415 | 220.59 | 125.9086 | 145.2585 | 156.1357 | 154.6814 | 184.5614 | 158.9086 | 172.8929 | 118.0286 | 121.94 | 110.6214 | 110.06 | 125.5143 | 242.8929 | 238.7415 | 219.7729 | 313.7101 | 207.6814 | 185.12 | 225.6371 | 306.4072 | 137.3629 | 76.1514 | 71.4857 | 137.3629 | 119.3786 | 111.41 | 126.8199 | 128.5771 | 104.7572 | 99.53 | 117.94 | 119.1357 | 74.9243 | 59.06 | 74.6057 | 100.5457 | 103.9557 | 95.8329 | 107.7885 | 85.7128 | 236.47 | 242.8643 | 258.0914 | 377.8615 | 138.53 | 190.1357 | 218.2899 | 351.9243 | 307.7599 | 190.9086 | 189.9086 | 325.2872
1554 | 19.0884 | 140.4081 | 222.1519 | 218.9681 | 187.6721 | 142.0242 | 171.176 | 160.1199 | 183.1838 | 187.9758 | 110.5841 | 128.0639 | 172.352 | 211.2321 | 147.472 | 165.616 | 227.736 | 189.8637 | 168.5997 | 161.56 | 185.352 | 144.1682 | 334.528 | 347.7118 | 182.528 | 253.5677 | 223 | 236.2959 | 310.2321 | 115.1199 | 105.9681 | 129.1199 | 126.0483 | 142.2959 | 112.5522 | 159.8878 | 88.8801 | 247.0078 | 219.616 | 265.7601 | 261.3201 | 152.4797 | 141.736 | 129.352 | 120.528 | 72.648 | 54.7921 | 67.472 | 120.5358 | 206.6238 | 189.5841 | 231 | 161.616 | 113.4081 | 143.384 | 146.912 | 126.0319 | 166.5039 | 180.8801 | 225.736 | 174.3279 | 131.4961 | 132.264 | 158.44 | 137.0242 | 137.9042 | 195.264 | 228.7118 | 246.176 | 173.0319 | 140.1122 | 256.44 | 222.7679 | 231.1519 | 205.44 | 237.1519 | 362.472 | 159.5919 | 171.384 | 189.648 | 239.1199 | 137.44 | 123.912 | 168.5919 | 161.4159 | 169.0242 | 171.2321 | 194.6799 | 153.3279 | 107.2399 | 106 | 119.7438 | 142.5039 | 263.8878 | 257.6238 | 219.5522 | 355.9198 | 216.088 | 211.528 | 209.44 | 345.1838 | 153.912 | 77.0319 | 54.4961 | 153.912 | 124.9681 | 94.3762 | 90.0242 | 101.7282 | 112.5919 | 105.1199 | 103.6799 | 132.0319 | 73.2959 | 71.912 | 81.824 | 96.3201 | 80.5841 | 106.528 | 94.4081 | 71.2321 | 228.5919 | 228.1199 | 248.1441 | 424.472 | 145.264 | 203.472 | 175.3762 | 350.472 | 257.5841 | 202.352 | 198.8878 | 343.4642
1555 | 19.0994 | 152.8937 | 202.5988 | 232.4101 | 203.1946 | 162.6991 | 173 | 154.6107 | 154 | 169.5074 | 127.5015 | 109.3039 | 178.006 | 221.8937 | 143.5045 | 171.6991 | 225.1003 | 148.8054 | 125.2066 | 154.4926 | 189 | 176.2979 | 340.5015 | 325.4042 | 185.0914 | 216.8967 | 222.2979 | 232.1092 | 318.4866 | 107.003 | 114.8967 | 120.2006 | 166.1946 | 136.9057 | 137.5958 | 136.6961 | 95.994 | 217.6048 | 228.2066 | 280.289 | 273.5958 | 117.2036 | 136.8937 | 131.997 | 126.4012 | 69.5015 | 63.6018 | 62.2006 | 97.6077 | 185.5104 | 204.7964 | 228.8937 | 164.6902 | 128.7021 | 137.6018 | 143.0914 | 116.1003 | 152.1033 | 187.0914 | 222.2979 | 160.2095 | 146.1976 | 136.003 | 164.003 | 154.4896 | 168.2006 | 194.3893 | 205.4012 | 247.7994 | 157.9851 | 163.003 | 258.6932 | 209.7994 | 212.8024 | 206.289 | 213.4866 | 358.5045 | 146.3039 | 163.3952 | 185.6991 | 229.2979 | 141.4985 | 124.9057 | 153.7994 | 142.7934 | 187.994 | 184.1003 | 181 | 134.7964 | 90.3039 | 106.8024 | 87.9057 | 124.994 | 242.4012 | 232.4985 | 247.7051 | 327.0089 | 216.3982 | 217.8024 | 214.6018 | 319.1092 | 152.6991 | 68.6048 | 69.1976 | 152.6991 | 133.8967 | 117.997 | 110.7994 | 132.1063 | 99.4042 | 96.6018 | 92.2066 | 121.5988 | 66.6018 | 69.4955 | 80.2006 | 109.0973 | 101.0119 | 112.3009 | 107.8967 | 82.2949 | 219.5164 | 221.708 | 258.6961 | 416.994 | 147.7994 | 198.9027 | 199.7994 | 349.714 | 273.1976 | 205.1976 | 178.003 | 367.3952
1556 | 19.1104 | 132.086 | 200.2385 | 260.1853 | 188.1259 | 160.9007 | 169.9601 | 188.4637 | 160.7553 | 185.4698 | 131.8874 | 121.4504 | 193.523 | 203.7881 | 155.7615 | 169.9007 | 226.9007 | 167.8147 | 144.0993 | 133.3511 | 191.2518 | 166.9601 | 349.2784 | 340.8147 | 161.0266 | 205.7615 | 221.6294 | 256.3972 | 279.3511 | 116.3378 | 108.1525 | 119.7482 | 151.1259 | 152.7615 | 126.0133 | 126.914 | 78.5629 | 233.914 | 243.0461 | 246.2119 | 268.0931 | 127.7748 | 116.8741 | 124.3511 | 131.3511 | 73.4371 | 69.4504 | 65.8014 | 119.1986 | 215.6489 | 192.7615 | 210.7881 | 139.5036 | 134.8874 | 140.7482 | 118.3511 | 116.6622 | 160.4238 | 162.0133 | 217.914 | 185.0594 | 140.914 | 146.6888 | 170.9734 | 124.4637 | 170.7881 | 164.7021 | 207.4238 | 245.4371 | 120.5435 | 170.3112 | 239.2385 | 206.8741 | 219.5496 | 175.9273 | 177.6161 | 373.0133 | 159.6888 | 151.1392 | 182.8874 | 225.8147 | 139.8147 | 140.3112 | 150.6489 | 123.1126 | 170.6755 | 181.9601 | 182.0133 | 126.477 | 102.6755 | 112.6489 | 105.3378 | 109.2518 | 245.7748 | 228.9007 | 261.4238 | 355.8014 | 209.9867 | 225 | 218.9867 | 345.8741 | 147.523 | 79.9601 | 63.0133 | 147.523 | 121.6356 | 109.9007 | 110.5762 | 148.4105 | 109.6356 | 101.3245 | 111.7748 | 117.7748 | 73.0133 | 57.8014 | 81.2119 | 100.7748 | 133.172 | 114.0993 | 98.4504 | 67.9007 | 266.8608 | 244.8741 | 248.4637 | 401.1392 | 144.7615 | 204.8608 | 197.3245 | 385.1321 | 266.3378 | 197.7748 | 188.1259 | 357.6161
1557 | 19.1214 | 122.8694 | 213.6215 | 237.8751 | 197.2498 | 167.2507 | 147.2488 | 198.1268 | 207.8809 | 136.8722 | 131.1249 | 125.2498 | 171.7522 | 209.2507 | 146.999 | 174.877 | 234.7493 | 183.756 | 137.1249 | 147.4966 | 209.4995 | 143.2498 | 379.2536 | 357.1306 | 175.8761 | 198.6225 | 261.2546 | 227.7502 | 289.7502 | 120.9981 | 130.3775 | 105.8713 | 158.8761 | 142.3756 | 135.1239 | 142.999 | 84.9981 | 243.3804 | 208.6206 | 243.7464 | 316.8818 | 125.3756 | 111.4966 | 136.3737 | 141.2507 | 67.2517 | 73.1249 | 79.7502 | 105.1249 | 203.0019 | 187.121 | 217.999 | 173.1268 | 134.2498 | 122.6254 | 129.999 | 115.3737 | 151.3718 | 167.3775 | 230.2526 | 157.9962 | 154.5014 | 159.1258 | 159.4957 | 136.3746 | 175.3766 | 187.7483 | 200.245 | 243.9971 | 175.3794 | 156.8751 | 253.3708 | 202.2459 | 214.2517 | 201.6225 | 211.5024 | 377.0038 | 171.8761 | 164.2536 | 179.0029 | 243.2546 | 158.5034 | 129.8722 | 142.4976 | 129.3698 | 176 | 162.4957 | 190.3746 | 146.7493 | 106.8761 | 104.123 | 107.2507 | 125.5014 | 241.3775 | 236.8742 | 250.1239 | 370.6244 | 205.8713 | 224.1258 | 217.6186 | 333.1287 | 129.4986 | 56.4995 | 70.7502 | 129.4986 | 105.6225 | 115.6263 | 120.2526 | 128.8742 | 95.4957 | 100.6206 | 108.5014 | 115.2507 | 79.7512 | 72.2498 | 74.6254 | 99.7493 | 103.2498 | 107 | 100.0019 | 76.2488 | 251.7483 | 238.4976 | 257.3775 | 414.5034 | 138.2469 | 190.9971 | 193.7483 | 331.7493 | 271.4976 | 195.2507 | 199.8722 | 389.6292
1558 | 19.1324 | 159.3362 | 233.5132 | 238.2345 | 199.823 | 158.022 | 154.4911 | 182.9424 | 165.6859 | 154.5309 | 132.4115 | 127.9602 | 155.177 | 207.4293 | 149.199 | 161.0974 | 241.0974 | 142.7434 | 141.4293 | 166.7875 | 213 | 145.5487 | 353.5487 | 318.4691 | 165.2566 | 217.6859 | 228.823 | 230.1151 | 291.2921 | 133.7655 | 112.9602 | 131.4911 | 149.7655 | 140.3319 | 144.7832 | 148.2168 | 96.6681 | 208.2523 | 237.9822 | 266.022 | 266.8628 | 121.9602 | 134.4911 | 144.9026 | 135.7257 | 60.3497 | 73.3141 | 78.5487 | 109.2921 | 185.5707 | 215.5885 | 224.5885 | 158.6283 | 135.3141 | 122.6062 | 137 | 126.0974 | 174.4513 | 150.0974 | 215.5266 | 185.0398 | 145.9204 | 152.7257 | 187.5707 | 139 | 164.9602 | 198.6283 | 231.5132 | 265.2743 | 143.6859 | 156.2743 | 278.9247 | 229.3938 | 205.4293 | 220.1372 | 198.801 | 352.5664 | 162.3539 | 143.0753 | 162.2523 | 213.1549 | 134.2743 | 147.5309 | 158.2168 | 164.3938 | 177.6461 | 191.2566 | 194.3717 | 152.823 | 99.6859 | 116.6681 | 105.2921 | 114.0398 | 226.2921 | 245.0576 | 252.376 | 375.5487 | 232.8628 | 218 | 258.7875 | 310.3895 | 137.9424 | 60.9602 | 70.6461 | 137.9424 | 122.9026 | 105.5885 | 101.5487 | 133.0796 | 124.3938 | 131.4911 | 98.823 | 109.7257 | 70.8628 | 73.7257 | 71.7257 | 102.3938 | 104.3141 | 104.6681 | 87.2345 | 83.3539 | 267.2921 | 253.9424 | 243.6638 | 387.6681 | 159.0796 | 207.199 | 206 | 341.3895 | 287.4911 | 193.5664 | 217.2566 | 359.823
1559 | 19.1434 | 133.4376 | 196.9075 | 247.943 | 201.2634 | 136.886 | 145.7473 | 171.7935 | 172.0925 | 144.2312 | 132.9075 | 131.9075 | 154.3344 | 226.0108 | 122.1387 | 169.3452 | 245.4591 | 163.7366 | 164.943 | 136.3914 | 212.7011 | 154.6796 | 355.057 | 340.8753 | 154.5763 | 223.7935 | 230.2634 | 251.5161 | 311.2527 | 125.7011 | 117.057 | 120.6548 | 141.7011 | 155.2989 | 156.4484 | 139.6903 | 84.0925 | 235.8968 | 225.2774 | 243.9892 | 266 | 130.5409 | 126.0462 | 137.8505 | 134.1495 | 90.4129 | 71.0925 | 79.6086 | 131.0462 | 166.8398 | 213.2989 | 223.943 | 152.6903 | 131.5978 | 140.943 | 136.4022 | 130.7581 | 172.7935 | 158.3452 | 238.7688 | 180.943 | 159.7935 | 151 | 168.8398 | 140.943 | 169.2065 | 194.6333 | 201.1849 | 268.0462 | 150.8398 | 157.8505 | 240.1247 | 214.7935 | 222.6656 | 218.3097 | 227.057 | 371.8613 | 155.886 | 171.114 | 184.3667 | 234.6441 | 137.1957 | 134.3914 | 145.9538 | 149.4946 | 190.2419 | 174.2989 | 198.2172 | 156.9538 | 103.2527 | 102 | 122.114 | 109.3452 | 246.1032 | 251.7226 | 220.1387 | 373.9183 | 230.9538 | 214.2634 | 228.2419 | 331.1247 | 124.7011 | 66.2527 | 79.5054 | 124.7011 | 115.5516 | 102.4022 | 108.2882 | 121.5978 | 110.3914 | 120.2065 | 104.8968 | 109.3452 | 69.5516 | 73.3452 | 68.057 | 88.2419 | 98.8043 | 94.6333 | 95 | 77.0355 | 284.114 | 243.3914 | 269.2634 | 381.6688 | 147.1495 | 181.9322 | 205.1032 | 370.943 | 277.2527 | 213.4591 | 201.943 | 360.9645
1560 | 19.1544 | 150.7355 | 192.2789 | 256.5737 | 181.4558 | 155.7649 | 155.8231 | 183.6175 | 181.5737 | 174 | 121.3235 | 122.4263 | 178.0295 | 205.9267 | 138.8677 | 171.8239 | 230.1912 | 189.0438 | 175.0295 | 155.4996 | 210.353 | 187.4853 | 344.4853 | 375.9562 | 191.1912 | 230.6032 | 214.6614 | 226 | 309.1028 | 124.9705 | 107.6175 | 115.588 | 146.6319 | 157.4853 | 158.0295 | 151.7649 | 95.5147 | 240.1912 | 258.0884 | 264.3968 | 267.7793 | 143.8677 | 134.4263 | 136.353 | 138.2351 | 69.353 | 81.059 | 68.2645 | 131.9857 | 190.6024 | 215.3235 | 236.7793 | 167.353 | 139.3673 | 148.6032 | 137.6908 | 118.1618 | 183 | 164.706 | 208.1912 | 190.706 | 165.3092 | 148.2502 | 184.1912 | 150.5442 | 160.4558 | 216.6327 | 219.0151 | 270.1179 | 165.2207 | 153.1179 | 205.8526 | 227.2645 | 196.8677 | 207.0438 | 215.5147 | 349.7944 | 171.5299 | 151.4263 | 160.2645 | 240.7649 | 143.1912 | 147.6765 | 143.3968 | 156.5442 | 198.3092 | 177.4558 | 173.2645 | 149.3825 | 100.941 | 103.1323 | 103.8821 | 116.5147 | 242.294 | 217.941 | 237.8382 | 348.0733 | 230.1761 | 198.9849 | 238.0884 | 298.7936 | 124.294 | 61.353 | 73.4263 | 124.294 | 113.4853 | 101.4263 | 126.0295 | 125 | 119.1474 | 107.412 | 107.7355 | 117.9705 | 68.4558 | 81.1618 | 62.6614 | 101.3235 | 93.7793 | 117.6032 | 95.1618 | 99.0295 | 264.1028 | 260.5586 | 249.4558 | 423.7211 | 146.8677 | 212.7498 | 200.8088 | 384.588 | 271.5442 | 202.0733 | 214.4558 | 347.7052
1561 | 19.1654 | 143.8732 | 227.7374 | 244.3481 | 189 | 142.1328 | 185.3038 | 187.8229 | 169.6962 | 175.2183 | 126.3068 | 136.5664 | 175.7847 | 192.6107 | 138.3954 | 146.2183 | 231.9175 | 204.2565 | 165.8229 | 178.2565 | 208.7404 | 187.2153 | 347.174 | 350.8229 | 185.0855 | 216.6549 | 241.9115 | 231.9175 | 323.1328 | 128.4336 | 118.7847 | 128.6519 | 177.6932 | 156.6932 | 151.7817 | 135.6962 | 91.6076 | 240.3511 | 245.4366 | 274.8642 | 282.5694 | 138.3481 | 143.5191 | 133.5221 | 149.6046 | 68.0885 | 73.2626 | 80.6962 | 104.2596 | 223.5604 | 213.5191 | 237.4718 | 161.9115 | 158.5191 | 130.4779 | 155.9085 | 120.9145 | 177.9527 | 160.6107 | 208.8732 | 182.9557 | 145.003 | 138.6992 | 176.3451 | 142.3038 | 170.7847 | 195.3924 | 193.7435 | 252.9588 | 155.3038 | 142.9205 | 231 | 237.6076 | 196.5694 | 226.4336 | 213.6962 | 333.7435 | 142.003 | 161.3893 | 171.1298 | 231.832 | 138.6519 | 147.7404 | 156.3008 | 150.3924 | 176.4366 | 189.8732 | 184.4779 | 139.0855 | 109.4336 | 105.6932 | 123.1298 | 116.0885 | 252.3068 | 228.5221 | 242.5694 | 369.5694 | 249.3421 | 228.4336 | 222.6519 | 337.829 | 133.6107 | 58.8702 | 84.6076 | 133.6107 | 115.826 | 112.0855 | 120.4779 | 124.4779 | 95.7404 | 93.6519 | 97.7404 | 118.4749 | 75.826 | 80.0855 | 87.4749 | 104.2183 | 105.0885 | 101.0443 | 97.3924 | 91.2153 | 270.3008 | 266.6429 | 260.1328 | 388.2626 | 140.6519 | 218.4718 | 205.174 | 391.5604 | 266.0885 | 211.7344 | 224.6107 | 392.0825
1562 | 19.1764 | 162.9006 | 214.9503 | 242.2733 | 191.9814 | 160.354 | 175.6646 | 171.7267 | 177.472 | 177.4596 | 141.8137 | 140.646 | 190.3043 | 205.9317 | 158.5963 | 149.5776 | 256.646 | 179.236 | 146.3727 | 148.0186 | 214.9503 | 170.8323 | 351.1677 | 339.1987 | 176.559 | 235.4907 | 229.5776 | 255.7143 | 336.5093 | 126.4037 | 132 | 126.2547 | 159.9503 | 144.354 | 146.3727 | 140.8634 | 85.1863 | 256.3416 | 254.205 | 241 | 304.528 | 138.6957 | 129.4224 | 130.4099 | 122.6957 | 75.5776 | 83.2547 | 83.2547 | 102.7081 | 198.2174 | 201.0994 | 209.2981 | 153.8634 | 139.5776 | 130.4224 | 130.913 | 127.764 | 158.6584 | 176.1677 | 230.323 | 179.4907 | 158.969 | 161 | 157.6957 | 143.472 | 176.5466 | 196.7826 | 224.913 | 268.3043 | 153.118 | 180.4969 | 237.1491 | 229.5093 | 222.8137 | 224.7764 | 219.9814 | 357.4596 | 160.8137 | 154.9627 | 167.559 | 269.9317 | 137.1863 | 153.205 | 136.795 | 154.0186 | 188.3727 | 204.4286 | 184.795 | 130.559 | 106.6584 | 97.0807 | 116.205 | 126 | 257.5652 | 228.205 | 267.882 | 395.0683 | 218.031 | 220.441 | 219.323 | 345.3789 | 148.9814 | 67.6584 | 75.3913 | 148.9814 | 113.5093 | 108.5901 | 122.6584 | 126.6584 | 103.6273 | 96.6584 | 106.9317 | 100.5404 | 74.0683 | 72.3043 | 69.5404 | 110.9317 | 112.9503 | 105.6273 | 101.205 | 80.9814 | 253.9627 | 217.6894 | 272.0186 | 400.4907 | 138.441 | 184.8944 | 203.5776 | 364.913 | 282.1491 | 186.6584 | 228.4286 | 370.5217
1563 | 19.1874 | 144.4041 | 207.9863 | 225.0069 | 205.5959 | 169.2398 | 152.9657 | 188.5959 | 197.1986 | 168.1574 | 141.3973 | 129.178 | 191.233 | 214.5684 | 140.3629 | 143.767 | 244.7808 | 189.3973 | 148.5959 | 139.767 | 207.3904 | 156.4041 | 364.8014 | 372.4179 | 182.5753 | 238.6234 | 220.1918 | 247.5272 | 330.1986 | 142.8495 | 129.4179 | 124.3904 | 154.178 | 158.2055 | 142.2398 | 145.0275 | 91.1643 | 240.767 | 238.2261 | 245.5684 | 279.2398 | 135.3835 | 136.822 | 113.5478 | 118.589 | 65.9931 | 75.8288 | 76.2261 | 126.8357 | 218.8151 | 213.0412 | 251.4797 | 165.3766 | 131.1849 | 138.0206 | 149.1918 | 123.1643 | 179.1986 | 187.4179 | 217.2398 | 185.6027 | 135.8082 | 159.0137 | 156.767 | 159.4247 | 148.767 | 176.7808 | 241.6027 | 274 | 157.2055 | 139 | 257.644 | 220.2192 | 225.5753 | 238.8769 | 232.8014 | 345.5753 | 158.411 | 175.4385 | 169.4041 | 267.8426 | 141.5753 | 143.3835 | 149.7945 | 147.1574 | 191.9863 | 171.1712 | 200.178 | 136.5753 | 122.8288 | 126.6096 | 104.3973 | 125.8014 | 216.3698 | 218.7808 | 257.8357 | 389.411 | 240 | 228.3286 | 211.8014 | 302.5753 | 158.4247 | 85.4179 | 63.8014 | 158.4247 | 109.5821 | 136.178 | 142.2055 | 143.4247 | 104.9794 | 117.9931 | 108.8151 | 113.6165 | 75.3629 | 74.4247 | 83.411 | 109.6371 | 105.3904 | 103.4041 | 90.3904 | 91.4179 | 277.0206 | 257.2192 | 257.2124 | 403.2261 | 140.9657 | 212.0206 | 194.3904 | 367.6989 | 300.6577 | 204.6165 | 192.3904 | 395.6234
1564 | 19.1984 | 141.1563 | 211.7817 | 231.124 | 206.3127 | 149.5856 | 170.6799 | 189.3127 | 191.6725 | 189.2035 | 140.0472 | 137.3598 | 173.2655 | 231.3127 | 162.2109 | 161.9454 | 259.1638 | 194.3747 | 153.531 | 161.531 | 210.6799 | 154.4218 | 367.3747 | 366.2183 | 199.9528 | 226.4764 | 220.9926 | 285.1489 | 331.2109 | 120.9603 | 115.6253 | 127.0472 | 164.6799 | 157.531 | 124.2729 | 136.3821 | 104.407 | 263.5856 | 225.7419 | 256.8288 | 269.0769 | 148.5856 | 125.4764 | 141.0472 | 125.8437 | 70.2109 | 61.9528 | 62.2655 | 111.9057 | 207.6799 | 191.7419 | 213.4367 | 173.5162 | 142.9528 | 134.2258 | 150.8362 | 139.1489 | 180.6328 | 173.6253 | 198.6402 | 186.9528 | 134.5856 | 156.9057 | 179.1638 | 147.7419 | 165.8908 | 189.4764 | 234.5162 | 271.6799 | 151.8908 | 141.7419 | 229.2581 | 212.3747 | 233.4616 | 199.6948 | 227.9926 | 360.2109 | 154.531 | 153.2109 | 161.938 | 243 | 154.5236 | 153.2109 | 153.6328 | 171.3672 | 194.7271 | 183.3598 | 202.876 | 146.3598 | 103.8908 | 115.5162 | 108.531 | 123.7345 | 236.9528 | 240.9677 | 245.2258 | 371.8214 | 241.6873 | 260.2978 | 215.8511 | 326.0694 | 148.2183 | 74.5782 | 64.2655 | 148.2183 | 121.0546 | 139.5087 | 132.4616 | 128.5782 | 111.7271 | 116.938 | 105.4838 | 106.7419 | 92.3598 | 64.4293 | 75.9454 | 92.4293 | 111.8437 | 101 | 98.7419 | 80.5782 | 269.0074 | 248.531 | 246.8362 | 388.2109 | 164.3747 | 206.5384 | 202.7419 | 316.9603 | 274.3821 | 195.6328 | 198.2109 | 375.6725
1565 | 19.2094 | 141.5713 | 197.7857 | 258.286 | 204.241 | 160.0983 | 168.2498 | 185.9017 | 172.4642 | 180.357 | 131.2321 | 118.6607 | 178.2232 | 224.3304 | 159.8751 | 158.6696 | 263.8573 | 205.884 | 160.5447 | 164.9734 | 201.5536 | 157.116 | 377.3215 | 375.527 | 204.3037 | 226.1966 | 213.1427 | 270.7945 | 332.6696 | 140.7502 | 105.6785 | 122.473 | 152.8751 | 164.9911 | 141.5447 | 158.8662 | 82.473 | 262.7147 | 234.2143 | 233.3481 | 302.5986 | 158.4287 | 129.2143 | 127.5447 | 128.3304 | 71 | 71.2143 | 65.884 | 129.3126 | 204.5802 | 200.6607 | 251.473 | 155.2321 | 150.4287 | 164.3304 | 139.6696 | 118.7679 | 180.5447 | 163.9017 | 214.884 | 195.5447 | 147.7768 | 172.9734 | 188.3215 | 158.6696 | 157.5536 | 197.9017 | 226.7229 | 266.3481 | 152.4819 | 157.1338 | 229.4908 | 221.2055 | 210.5625 | 216.5358 | 213.4464 | 364.5713 | 161.7679 | 155.5625 | 155.7229 | 241.4375 | 150.116 | 159.5802 | 150.8662 | 166.5447 | 182.1427 | 173.8123 | 174.1515 | 128.3304 | 97.116 | 98.5713 | 115.7679 | 123.9106 | 248 | 277.759 | 272.4287 | 338.4819 | 243.3126 | 224 | 232.884 | 346.3843 | 167.3481 | 74.116 | 68.7768 | 167.3481 | 129.9106 | 104.4464 | 114.2498 | 123.2055 | 93.5625 | 106.9283 | 122.8751 | 110.3037 | 76.7857 | 76.759 | 74.0089 | 107.4375 | 119.0178 | 103.0089 | 108.7768 | 78.3304 | 275.7413 | 251.7502 | 235.2232 | 380.7413 | 170.3037 | 226.5358 | 213.2232 | 346.3481 | 305.5713 | 200.0089 | 200.116 | 357.5802
1566 | 19.2204 | 153.529 | 210.123 | 225.239 | 214.29 | 148.587 | 190.7065 | 191.29 | 187.058 | 196.174 | 138.058 | 116.5905 | 181.355 | 223.1775 | 151.5905 | 166.8875 | 236.645 | 195.6415 | 150.587 | 146.826 | 207.3005 | 164.768 | 359.7575 | 352.7645 | 180.2935 | 219.4235 | 238.942 | 275.7575 | 335 | 123.239 | 117.5905 | 140.5255 | 139.4095 | 158.2355 | 134.884 | 143.5905 | 101.232 | 224.0685 | 240.246 | 253.007 | 259.1265 | 140.1125 | 128.4165 | 123.0035 | 130.4745 | 72.1775 | 62.4095 | 62.471 | 115.1195 | 215.7575 | 201.181 | 267.1705 | 161.3515 | 144.123 | 154.2285 | 139.174 | 111 | 175.0615 | 176.355 | 211 | 186.2935 | 147.942 | 146.1125 | 183.0035 | 160.058 | 161.652 | 204.232 | 255.928 | 272.819 | 179.4095 | 176.6485 | 257.7575 | 208 | 220.9455 | 200.529 | 215.2355 | 375.587 | 152.587 | 165.71 | 193.877 | 234.116 | 158.71 | 174.0545 | 137.9455 | 158.471 | 200.877 | 201 | 197.7575 | 131.652 | 103.355 | 109.116 | 106.587 | 134.6415 | 247.058 | 261.8225 | 259.058 | 375.065 | 231.0035 | 227.0615 | 231.5905 | 310.304 | 171.464 | 79.6485 | 68.7065 | 171.464 | 138.051 | 108.826 | 133.4095 | 118.2425 | 101.5905 | 129.993 | 111.7645 | 95.949 | 81.3515 | 64.355 | 76.0545 | 97.5255 | 128.5255 | 106.703 | 101.877 | 80.9455 | 252.8875 | 236.594 | 236 | 367.072 | 151.0035 | 206.0545 | 214.2355 | 353.761 | 308.3445 | 207.471 | 207.5325 | 382.652
1567 | 19.2314 | 152.4956 | 225.0442 | 234.0442 | 207.9205 | 142.5133 | 185.5663 | 176.9912 | 184.9912 | 184.0265 | 136.7345 | 129.2124 | 190.4867 | 218.3273 | 158.0177 | 182.5309 | 243.3718 | 187.3895 | 146.4956 | 159.7169 | 222.0795 | 176.2566 | 352.3277 | 352.2478 | 176.0177 | 254.0795 | 242.7434 | 253.2124 | 329.3008 | 136.7522 | 124.0177 | 130.4602 | 139.1771 | 163.4602 | 140.7522 | 152.4956 | 92.4691 | 257.5221 | 274.3008 | 278.0088 | 294.7257 | 125.9558 | 143.5221 | 137.4867 | 141.5309 | 76.4956 | 54.7434 | 67.2213 | 120.5133 | 190.9823 | 220.2301 | 236.8229 | 154.761 | 164 | 135.1948 | 139.1682 | 117.1948 | 189.2124 | 184.2478 | 213.4779 | 175.5751 | 152.7345 | 134.6815 | 185.8406 | 152.292 | 180.7434 | 189.2743 | 210.4691 | 253.0265 | 174.9646 | 172.584 | 231 | 217.6638 | 232.5398 | 208.416 | 219.9646 | 371.7434 | 156.6727 | 176.7876 | 175.4867 | 229.9823 | 166.3185 | 157.5044 | 150.2831 | 154.5486 | 179.761 | 201.2478 | 176.4514 | 146.5309 | 110.0088 | 105.4779 | 106.2124 | 119.2036 | 249.947 | 253.0442 | 250.7964 | 393.0442 | 243.7522 | 230.3097 | 239.5044 | 345.761 | 150.9646 | 85.2478 | 73.9735 | 150.9646 | 109 | 120.9735 | 125.9912 | 135.8318 | 107.5221 | 106.2301 | 113.9735 | 120.7876 | 78.7257 | 69.5221 | 62.9735 | 83.9912 | 111.2743 | 100.4602 | 85.2213 | 96.0088 | 277.947 | 258.9912 | 245.416 | 411.5398 | 162.0177 | 195.947 | 212.5221 | 348.9205 | 279.2301 | 202.5574 | 217.761 | 394.5574
1568 | 19.2424 | 151.3992 | 214.1613 | 220.8206 | 219.9153 | 141.6411 | 162.7399 | 181.3004 | 190.3407 | 179.1613 | 144.3407 | 142 | 194.2198 | 203.7037 | 153.8206 | 174.3407 | 279.4576 | 217.9778 | 146.6996 | 166.3185 | 200.9819 | 181.8024 | 394.6149 | 365.744 | 168.4395 | 229.8609 | 241.6189 | 265.7399 | 310.4395 | 132.6189 | 122.4214 | 142.4395 | 163.4798 | 176.4798 | 144.6815 | 152.1794 | 102.2198 | 247.1391 | 259.0807 | 277.8609 | 294.9375 | 138.1391 | 140.4214 | 137.3185 | 129.4395 | 76.6996 | 58.8206 | 77.5202 | 116 | 204.2823 | 221.0181 | 215.9012 | 151.2601 | 162.1794 | 151.3992 | 165.6996 | 134.6996 | 191.0766 | 183.6996 | 219.4395 | 160.4033 | 158 | 157.4798 | 172.4255 | 139.5605 | 183.5202 | 181.5202 | 226.2016 | 241.9193 | 179.9778 | 147.3811 | 245.5646 | 245.1794 | 221.3407 | 234.1391 | 235.3811 | 382.5827 | 178.3185 | 160.0988 | 175.0585 | 237.8206 | 144.3004 | 153.3992 | 139.5202 | 142.1613 | 181.7218 | 205.6411 | 193 | 138.3407 | 110.121 | 107.0181 | 117.4395 | 131.6189 | 264.879 | 246.0625 | 241.242 | 375.1391 | 248.7218 | 215.4617 | 241.3811 | 341.2198 | 159.6189 | 83.6593 | 81.9597 | 159.6189 | 115.242 | 126.6189 | 125.3589 | 112.6815 | 103.3811 | 114.0403 | 121.1794 | 111.3811 | 87.2601 | 66.6815 | 70.1794 | 87.7802 | 102.7399 | 115.0403 | 98.9012 | 88.8387 | 294.4395 | 262.2601 | 270.0988 | 413.6049 | 155.2198 | 209.0585 | 209.1613 | 365.8568 | 287.8206 | 187.9415 | 215.0403 | 379.6815
1569 | 19.2534 | 141.821 | 225.1829 | 220.8249 | 179.4514 | 147.9144 | 162.5447 | 190.9922 | 194.821 | 183.3736 | 147.7315 | 133.8288 | 192 | 232.3697 | 152.4631 | 175.0078 | 263.1751 | 202.0856 | 150.8988 | 155.7237 | 219.0895 | 192.5604 | 350.179 | 384.2958 | 171.0817 | 236.0934 | 230.0934 | 264.7276 | 323.1595 | 127.6303 | 133.3658 | 133.0973 | 165.2685 | 164.9222 | 150.1012 | 148.0895 | 100 | 246.3541 | 264.4553 | 278.1012 | 286.1634 | 130 | 147.2802 | 124.8171 | 127.7237 | 78.1751 | 66.4514 | 79 | 120.0856 | 218.7354 | 204 | 228.6342 | 160.7159 | 149.9183 | 143.7276 | 157.0973 | 135.6303 | 164.6303 | 174.5525 | 210.0973 | 189.1829 | 158.5447 | 155.7276 | 204.2958 | 145.3619 | 182.0039 | 177.5525 | 233.6459 | 246.0778 | 164.358 | 155.6381 | 283.4592 | 233.1907 | 229.0895 | 218.1012 | 243.0934 | 398.0117 | 164.7276 | 147.3658 | 165.6303 | 250.8988 | 146.6381 | 145.7276 | 149.4435 | 160.2646 | 196.9105 | 206.1946 | 195.179 | 145.5447 | 118.1829 | 96.5369 | 119.8093 | 115.7354 | 252.1868 | 269.2724 | 248.3775 | 380.6186 | 256.0117 | 235.358 | 246.9144 | 339.8171 | 150.8171 | 75.0934 | 77.9105 | 150.8171 | 126.1907 | 114.821 | 121.8093 | 126 | 118.4475 | 116.1829 | 119.2685 | 120.4553 | 84.4592 | 82.4514 | 66.0895 | 97.0817 | 106.0856 | 114.1868 | 108.9105 | 80.8132 | 295.1751 | 264.0895 | 264.4475 | 451.1907 | 157.0856 | 196.3619 | 218.821 | 343.2646 | 302.8054 | 199.5486 | 217.1829 | 386.7354
1570 | 19.2644 | 138.8466 | 221.5767 | 219.2405 | 191.6933 | 143.5473 | 167.7006 | 205.8613 | 183.5914 | 177.8025 | 139.2847 | 120.5399 | 185.4527 | 225.9779 | 142.1166 | 153.8761 | 273.8687 | 212.4307 | 169.438 | 162.146 | 224.2773 | 168.2626 | 356.854 | 346.9632 | 186.8687 | 233.2773 | 234.9632 | 268.8393 | 356.7595 | 139.146 | 130.4233 | 118.8613 | 170.0147 | 147.1092 | 137.2552 | 147.8687 | 96.8687 | 267.1387 | 261.0074 | 254.7227 | 317.4527 | 137.4013 | 139.4013 | 127 | 141.8319 | 87.73 | 77.8393 | 77.5767 | 131 | 206.5546 | 204.854 | 232.7153 | 178.8761 | 136.6933 | 139.8687 | 142.5767 | 143.73 | 173.584 | 158.0074 | 203.5473 | 186.146 | 159.4307 | 155.854 | 171.2332 | 147.292 | 173.7153 | 170.1166 | 215.1387 | 277.27 | 166.8908 | 153.4233 | 269.1607 | 217.2773 | 239.1166 | 203.5473 | 240.562 | 380.5473 | 168.27 | 150.1166 | 174.0147 | 271.146 | 141.5767 | 147.2773 | 175.416 | 177.2847 | 192.5767 | 176.0147 | 197.2994 | 145.8613 | 114.5767 | 115.9926 | 135.1534 | 111.8099 | 243.708 | 271.708 | 235.0872 | 407.8908 | 232 | 245.8613 | 228.8834 | 338.8687 | 153.5693 | 70 | 73.292 | 153.5693 | 109.708 | 113.5546 | 134.3067 | 129.7006 | 124.1754 | 117.7006 | 130.2773 | 119.854 | 73.5767 | 84.1607 | 69.2847 | 110.0221 | 111.3067 | 104.854 | 106 | 88.7227 | 296.7595 | 273.5473 | 282.416 | 434.1387 | 161.4527 | 198.8613 | 211.8613 | 360.2847 | 323.5914 | 198.1239 | 213.8613 | 384.5178
1571 | 19.2754 | 145.703 | 220.6726 | 249.4061 | 202.7335 | 158.2183 | 178.7817 | 198.8426 | 173.2665 | 204.6066 | 139.1091 | 136.9518 | 176.4239 | 249.9086 | 160.5939 | 147.236 | 260.9518 | 211.8909 | 162.703 | 161.4848 | 221.9822 | 189.3452 | 356.0305 | 383.2843 | 182.5635 | 244.9391 | 255.2487 | 276.5457 | 333.1269 | 138.1878 | 133.703 | 124.3147 | 157.6244 | 170 | 156.5152 | 144.7513 | 86.0305 | 268.8122 | 261.6421 | 256.1269 | 298.3274 | 147.3883 | 154.7335 | 133.533 | 153.0609 | 77.8909 | 85.2487 | 68.0609 | 132.1878 | 215.6244 | 199.8731 | 229.3274 | 161.8426 | 152.1878 | 133.7817 | 141.9695 | 135.9695 | 164.1091 | 158.9391 | 222.0787 | 188.4543 | 161.2665 | 149.3883 | 203.4365 | 139.7335 | 174.4848 | 193.9391 | 219.4848 | 283.2487 | 153.7995 | 154.3274 | 259.0787 | 220.0305 | 252.5457 | 215.8426 | 237.2005 | 391.9518 | 184.9391 | 169.4848 | 169.6421 | 267.8122 | 140.0787 | 154.1878 | 174.4975 | 173.8426 | 193.1574 | 171.9391 | 186.8122 | 141.2183 | 114.8604 | 117.1396 | 120.2792 | 144.5939 | 244.1396 | 273.0305 | 268.2487 | 385.8909 | 237.0482 | 240.9213 | 210.703 | 327.4365 | 156.1878 | 69.4061 | 71.0787 | 156.1878 | 112.8122 | 125.8909 | 120.8909 | 139.8909 | 106.7513 | 120.764 | 132.1396 | 125.2665 | 73.2665 | 77.3452 | 69.703 | 95.297 | 99.6726 | 104.3274 | 107.4848 | 81.8122 | 268.9695 | 281.9822 | 287.4365 | 433.7335 | 143.5152 | 196.5939 | 202.17 | 363.9695 | 304.3579 | 219.0178 | 213.9695 | 419.8122
1572 | 19.2864 | 151.1849 | 222.3613 | 253.2572 | 203.0404 | 154.3092 | 177.1243 | 189.6185 | 189.6589 | 180.3092 | 136.6908 | 124.0723 | 191.8353 | 262.9596 | 164.1647 | 162.3092 | 257.3496 | 215.237 | 158.9075 | 166.5462 | 214.4421 | 202 | 352.0925 | 396.3932 | 191.3092 | 249.4134 | 244.9277 | 270.6387 | 353.0925 | 140.3815 | 130.526 | 135.0319 | 155.0925 | 161.3411 | 143.1041 | 160.9479 | 82.4017 | 267.8555 | 264.7949 | 275.7832 | 287.6706 | 132.5664 | 148.237 | 149.9277 | 142.7109 | 80.6185 | 71.8353 | 61.7311 | 133.1445 | 208.763 | 186.4017 | 229.4942 | 147.0521 | 144.1764 | 135.2168 | 148.0723 | 140.526 | 167.8757 | 175.1647 | 223.2891 | 199 | 160.9596 | 135.8034 | 193.1445 | 131.6908 | 175.8353 | 207.0723 | 215.8873 | 270.4538 | 178.8555 | 152.3294 | 252.2487 | 217.3294 | 243.8555 | 204.1445 | 224.5144 | 375.0521 | 193.1243 | 170.8353 | 175.2689 | 268.7109 | 146.237 | 155.1445 | 148.9075 | 163.3815 | 198.6706 | 185.3815 | 190.1849 | 138.237 | 124.9277 | 110.1647 | 111.2774 | 146.3092 | 234.3815 | 263.8353 | 258.5462 | 394.1849 | 245.2891 | 236 | 220.2051 | 323.0202 | 154.9798 | 72.3294 | 76.9277 | 154.9798 | 110.3092 | 124.2891 | 124.237 | 137.3613 | 122.0202 | 108.474 | 117.4336 | 130.2168 | 83.1647 | 85.0521 | 69.3092 | 95.6908 | 108.474 | 96.1445 | 111 | 78.0723 | 292.0808 | 262.6908 | 272.1966 | 435.8959 | 134.4336 | 206.3496 | 188.1647 | 366.0521 | 301.8236 | 231.8151 | 212.6504 | 410.5058
1573 | 19.2974 | 160.1769 | 212.3215 | 262.4277 | 228.0738 | 155.3215 | 174.9323 | 190.3569 | 197.7462 | 186.1446 | 136.9646 | 125.5369 | 189.2508 | 244.3569 | 161.9262 | 164.6077 | 273.4277 | 218.6431 | 156.18 | 169.3569 | 237.7492 | 199.1062 | 364.4661 | 369.2154 | 191.6785 | 229.2861 | 244.4277 | 279.3923 | 361.9292 | 135.1416 | 119.5339 | 121.7523 | 159.4277 | 147.4661 | 134.5045 | 169.7492 | 92.2508 | 264.9262 | 244.2508 | 284.2508 | 292.7907 | 141.4631 | 142.3185 | 148.7846 | 153.5723 | 80.3923 | 73.4307 | 78.0354 | 130.2861 | 202.4631 | 199.7877 | 244 | 147.2538 | 131.8969 | 127.7139 | 140.2123 | 131.1416 | 186.4661 | 184.5015 | 221.1092 | 198.6785 | 142.3569 | 152.6077 | 188.3569 | 135.8231 | 175.5015 | 205.3215 | 210.9323 | 267.3215 | 181.3923 | 168.7523 | 233.7169 | 228.9292 | 245.4277 | 215.7553 | 247.6785 | 368.8231 | 177.1062 | 176.2892 | 163.1446 | 283.1092 | 161.2184 | 148.1062 | 146.8231 | 161.0354 | 188.6785 | 188.5015 | 202.7139 | 141.9646 | 126.0354 | 116.6077 | 125.2477 | 140.2477 | 252.9353 | 251.2831 | 267.1446 | 404.7846 | 244.7169 | 247.8969 | 242.0354 | 339.0738 | 154.0384 | 83.9292 | 78.3569 | 154.0384 | 115.8231 | 115.6785 | 124.7492 | 141.7907 | 121.0324 | 109.82 | 112.0738 | 125.9292 | 79.9616 | 74.3215 | 72.2508 | 99.5015 | 110.4631 | 94.8938 | 112.6077 | 77.2861 | 318.3539 | 257.1769 | 257.3215 | 455.7492 | 130.36 | 225 | 202.3246 | 359.5015 | 318.2831 | 217.0354 | 193.6785 | 412.7199
1574 | 19.3084 | 150.6677 | 213.3338 | 251.6588 | 223.638 | 154.6647 | 196.3309 | 184.9941 | 188.3546 | 190.9911 | 136.997 | 133.6677 | 201.3442 | 243 | 153.6766 | 170.6706 | 266.3309 | 217.3294 | 170.6751 | 165.6632 | 222.3175 | 188.9941 | 370.9926 | 371.3264 | 193.3368 | 241.681 | 240.0015 | 275.6662 | 360.3249 | 131.0059 | 106.9985 | 130.3249 | 156.3368 | 158.6662 | 157.9926 | 154.9852 | 92.9941 | 260.0148 | 253.6736 | 288.6662 | 339.0223 | 136.003 | 125.3353 | 144.3368 | 151.322 | 79.003 | 85.3309 | 74.3309 | 128.3264 | 205.0148 | 209.6632 | 250.0089 | 165.6588 | 152.6602 | 135.0148 | 136.0119 | 120.6632 | 202.0059 | 187.9911 | 230.9955 | 189.9881 | 155.0208 | 155.6662 | 186.3323 | 155.3472 | 192.0104 | 198.6558 | 228.9926 | 269.3353 | 168.319 | 174.6543 | 248.6647 | 236.0045 | 239.6662 | 241.322 | 245.6647 | 388.3472 | 181.0148 | 184.9807 | 175.0015 | 286.6528 | 172.3116 | 149.0104 | 150.3234 | 163.3398 | 197.681 | 200.6706 | 206.0089 | 148.0059 | 125.0015 | 117.3294 | 115.0059 | 132.6766 | 279.9777 | 236.3353 | 263.3116 | 380.9748 | 261 | 260.3145 | 235.3264 | 348.6588 | 167.9955 | 84.3279 | 78.3353 | 167.9955 | 125.6662 | 115 | 117.997 | 165.9896 | 101.3353 | 108.678 | 129.003 | 127.6632 | 73.6825 | 72.3294 | 72.3264 | 113.0059 | 101.3309 | 93.9896 | 109.9911 | 84.0059 | 298 | 247.3338 | 259.0015 | 441.319 | 152.3398 | 219.9926 | 210.6499 | 373.3398 | 306.0059 | 222.3442 | 204.3501 | 431.9763
1575 | 19.3194 | 159.9607 | 211.9233 | 235.4621 | 187.0767 | 161.6913 | 198.5379 | 179.769 | 204.5397 | 190.4219 | 141.3069 | 130.5005 | 205.2703 | 245.0767 | 173.2301 | 175.3078 | 265.769 | 215.1151 | 179.231 | 165.8457 | 213.4603 | 177.5388 | 373.4603 | 372.0374 | 200.0767 | 261 | 253.4219 | 269.4621 | 345.1927 | 139.6922 | 117.4603 | 120.7306 | 156.5005 | 162.8457 | 147.3078 | 139.1534 | 89.8457 | 272.3854 | 246.3863 | 293.5379 | 358.6164 | 138.9616 | 142.1909 | 144.5005 | 141.5379 | 84.7306 | 77.1543 | 68.5772 | 126.6146 | 231.2301 | 202.5772 | 250.9242 | 167.1142 | 147.8073 | 149.4621 | 147.8466 | 118.0767 | 212.0767 | 183.2685 | 235.7297 | 177.8073 | 178.5005 | 163.9991 | 182.231 | 167.4237 | 194.5781 | 184.3461 | 225.2301 | 274.4228 | 156.2685 | 156.6155 | 257.0758 | 243.0384 | 240.3845 | 224.6155 | 242.3078 | 398.6931 | 199.2694 | 162.1151 | 180.8073 | 267.3078 | 154.4219 | 167.8457 | 138.7306 | 181.6913 | 202.463 | 201.8466 | 202.0785 | 151.8466 | 118.6931 | 111.3078 | 121.9616 | 138.3854 | 248.9616 | 247.3069 | 235.0384 | 367.0749 | 266.884 | 235.3461 | 242.6137 | 339.7306 | 160.2694 | 74.9233 | 76.5005 | 160.2694 | 128.8073 | 115.3461 | 121.2685 | 144.7315 | 113.3452 | 117.0776 | 122.9626 | 127.1534 | 88.7699 | 73.2301 | 60.231 | 121.6922 | 96.6155 | 87.6146 | 99.3845 | 83.347 | 301.4612 | 258.0374 | 260.6539 | 419.5772 | 163.4228 | 212.769 | 198.0374 | 384.4228 | 307.4237 | 216.9251 | 219.3854 | 399.3078
1576 | 19.3304 | 175.7168 | 223.4116 | 232.1682 | 193.5089 | 181.4336 | 196.4557 | 184.2832 | 184.5841 | 191.5706 | 155.5664 | 127.3761 | 193.9602 | 247.5664 | 181.7743 | 170.7743 | 276.3761 | 224.2257 | 176.8673 | 172.4911 | 238.792 | 170.6416 | 388.3982 | 398.885 | 194.3229 | 265.3009 | 270.3407 | 264.0177 | 349.8275 | 136.6991 | 134.1902 | 123.6416 | 153.0177 | 171.2832 | 156.3939 | 150.2257 | 95.0575 | 253.6991 | 227.9028 | 295.3982 | 325.022 | 139.8673 | 156.0972 | 143.885 | 147.3407 | 82.6239 | 66.208 | 72.6018 | 130.2477 | 233.6813 | 193.3407 | 239.6771 | 188.9248 | 153.9248 | 136.1327 | 146.8098 | 121.2832 | 215.2832 | 193.7743 | 244.1725 | 181.0575 | 178.6018 | 177.7743 | 178.4336 | 153.9248 | 171.5664 | 188.9425 | 233.0575 | 277.208 | 165.6991 | 155.792 | 273.3407 | 243.9248 | 237.2654 | 219.1327 | 240.2832 | 393.3939 | 192.4159 | 171.2257 | 183.6991 | 270.3009 | 168.115 | 168.3982 | 153.4691 | 190.3229 | 179.3009 | 195.0752 | 186.3364 | 152.9602 | 104.4336 | 114.3009 | 119.2832 | 133.6771 | 256.3187 | 254.7566 | 236.6416 | 405.3584 | 262.9468 | 224.5309 | 262.8893 | 336.1902 | 153.7743 | 72.4336 | 72.3009 | 153.7743 | 133.4911 | 122.093 | 130.3407 | 130.6416 | 135.3009 | 111.5266 | 106.8673 | 134.6416 | 79.8673 | 85.3584 | 63.9602 | 120.4911 | 105.4691 | 102 | 101.6416 | 69.5089 | 309.0752 | 279.5089 | 254.2654 | 418.9425 | 161.907 | 207.6061 | 210.1902 | 383.9823 | 294.2832 | 190.4691 | 202.1327 | 390.8318
1577 | 19.3414 | 174.0344 | 235.8013 | 253.5261 | 200.9656 | 181.0344 | 172.8879 | 182.6293 | 186.5516 | 176.8714 | 150.0344 | 137.3707 | 198.5083 | 253.1554 | 177.4828 | 168.3363 | 279.6981 | 220.7325 | 182 | 163.9223 | 229.0255 | 178.1554 | 383.0089 | 406.3274 | 186.2675 | 260.7669 | 263.1121 | 263.3618 | 360.6204 | 137.5261 | 120.8879 | 134.8624 | 162 | 169.6293 | 168.3274 | 155.6293 | 92.4484 | 261.1987 | 262.4917 | 275.5516 | 306.6726 | 146.4828 | 130.0344 | 151.3274 | 141.9656 | 72.6293 | 67.0777 | 85.4051 | 115.8191 | 225.8968 | 202.0522 | 259.5949 | 183.6637 | 152 | 130.2586 | 161.5949 | 123.3363 | 208.4395 | 178.3618 | 215.2153 | 170.6637 | 189.9223 | 167.5516 | 187.3019 | 162.3796 | 173.8191 | 200.4484 | 225.6293 | 269.1809 | 170.2331 | 157.8879 | 257.9567 | 242 | 229.5949 | 206.9567 | 245.6726 | 397.5427 | 186 | 178.8535 | 175.2586 | 270.2153 | 158.0777 | 157.4484 | 171.7758 | 175.2242 | 192.5605 | 199.2242 | 206.8102 | 160.4739 | 112.5605 | 112.7325 | 115.7758 | 148.4051 | 273 | 249.2675 | 237.8535 | 391.172 | 241.1898 | 209.5605 | 254.4917 | 338.4573 | 160.9745 | 73.1465 | 77.0344 | 160.9745 | 129.3707 | 135.9656 | 117.9223 | 128.5172 | 126.3185 | 125 | 111.6293 | 126.586 | 80.9223 | 87.1121 | 74.0688 | 119.707 | 121.1809 | 102 | 104.3363 | 83.6382 | 300.9911 | 278.8102 | 260.3108 | 428.2242 | 155.8191 | 196.7847 | 204.3019 | 385.3796 | 299.6726 | 214.3363 | 201.8191 | 384.3019
1578 | 19.3524 | 170.532 | 221.4898 | 266.085 | 208.298 | 171.7871 | 171 | 179.3191 | 193.5537 | 202.1272 | 141.5531 | 141.0639 | 186.2768 | 250.9143 | 167.7443 | 183.9578 | 269.383 | 203.851 | 174.3402 | 160.0639 | 219.2557 | 192 | 393.4891 | 402.2768 | 202.766 | 253.0217 | 263.085 | 256.5748 | 352.1918 | 148.9361 | 123.5959 | 148.8721 | 162.766 | 172.0639 | 152.0211 | 157.298 | 99.617 | 282.4041 | 268.5953 | 260.7232 | 313.4041 | 149.766 | 128.8299 | 146.1279 | 142.0211 | 79.2768 | 71.3191 | 82.149 | 133.5748 | 221.4674 | 227.9361 | 250.5953 | 175.3191 | 154.298 | 140.1068 | 159.4891 | 135.5109 | 205.8939 | 171.1918 | 208.5109 | 177.6387 | 176.4891 | 164.5959 | 201.0639 | 175.0211 | 179.702 | 204.5531 | 238.0217 | 271.7238 | 181.4252 | 147.1912 | 264.3836 | 253.8728 | 237.4469 | 206.8728 | 259.298 | 382.1918 | 186.766 | 172.4252 | 172.851 | 266.1918 | 160.7871 | 168.0639 | 165.3191 | 175.5531 | 192.4252 | 199.9361 | 203.7449 | 152.1279 | 126.5959 | 106.9578 | 109.702 | 137.8721 | 278.7449 | 258.4891 | 243.298 | 379.4047 | 249.234 | 235.0857 | 259.0633 | 363.766 | 165.3824 | 74.9789 | 71.234 | 165.3824 | 138.8088 | 143.298 | 114.4469 | 131.3619 | 119.4476 | 124.234 | 112.915 | 123.9578 | 81.2768 | 84.0211 | 69.7449 | 117.1912 | 125.2557 | 98.9361 | 110.383 | 96.5531 | 292.4258 | 273.0639 | 278.0422 | 435.8299 | 164.383 | 212.298 | 221.1279 | 390.7443 | 310.234 | 220.766 | 204.2551 | 396.915
1579 | 19.3634 | 167.8613 | 241.5811 | 254.3038 | 213.1858 | 173.3009 | 172.9764 | 177.767 | 206.7226 | 183.3953 | 139.9528 | 151.534 | 194.4425 | 228.5575 | 157.0915 | 190.9085 | 282.6491 | 197.6519 | 163.1858 | 166.1858 | 246.9057 | 182.1179 | 385.3717 | 400.9557 | 192.5368 | 272.8142 | 263.1623 | 260.5368 | 365.8142 | 146.767 | 119.1415 | 130.6991 | 165.1858 | 177.7906 | 147.9528 | 151.1179 | 94.6519 | 267.4896 | 258.0679 | 266.7877 | 305.6047 | 144.6755 | 130.2566 | 143.3509 | 135.9764 | 86.2566 | 70.5575 | 77.2802 | 142.6755 | 210.1594 | 225.767 | 227.8142 | 166.6519 | 166.3009 | 147.3274 | 148.6047 | 146 | 206.7226 | 181.6519 | 239.95 | 203.534 | 164.0236 | 170.0236 | 191.3745 | 174.1151 | 185.0915 | 203.3481 | 247.9321 | 290.1858 | 174.3009 | 145.2538 | 279.3745 | 257.584 | 250.1151 | 239.4396 | 253.9085 | 382.3745 | 179.6991 | 160.1623 | 180.4868 | 286.1387 | 157.5575 | 174.1858 | 161.3953 | 174.7434 | 196.3689 | 202.9057 | 218.1387 | 159.233 | 124.1179 | 122.2094 | 112.7198 | 131.9528 | 282.8613 | 254.7198 | 245.4189 | 401.7906 | 258.6726 | 242.817 | 235.7906 | 342.8641 | 144.9057 | 79.4425 | 69.2094 | 144.9057 | 146.2802 | 142.2566 | 128.3009 | 138.0236 | 144.5132 | 123.3953 | 119.9528 | 126.9557 | 87.466 | 75.2094 | 74.6519 | 102.6047 | 137.4896 | 99.1387 | 107.8377 | 90.6047 | 299.2802 | 264.1651 | 263.1858 | 426.584 | 173.6962 | 216.7906 | 230.6047 | 381.2774 | 301.8849 | 214.0943 | 209.6255 | 394.8613
1580 | 19.3744 | 165.2985 | 243.5924 | 243.4454 | 219.0757 | 175.4031 | 181.2985 | 185.2606 | 188.5546 | 186.853 | 154.9666 | 155.5167 | 193.4833 | 248.2651 | 171 | 188.0045 | 295.4788 | 193.853 | 167.0379 | 161.8864 | 243.5122 | 191.0469 | 392.2606 | 385.8909 | 184.7818 | 265.7015 | 272.0757 | 262.5636 | 364 | 150.5924 | 117.303 | 122.8909 | 174.7439 | 183.0757 | 151.147 | 144.9666 | 92.0757 | 268.6348 | 270.5167 | 293.1849 | 315.1894 | 143.1515 | 129.1136 | 132.5591 | 134.5167 | 82.6682 | 78.4454 | 72.706 | 130.147 | 238.147 | 227.9621 | 224.7773 | 168.9666 | 175.3318 | 137.1136 | 149.2227 | 143.5546 | 201.1894 | 177.0379 | 263.0334 | 201.4031 | 167.1136 | 155.9955 | 173.4833 | 178.4788 | 192.4788 | 205.1091 | 237.9288 | 296.4833 | 183.3318 | 159.9243 | 262.706 | 238.0757 | 261 | 260.4076 | 236.3318 | 369.7818 | 173.5212 | 171.9288 | 188.4788 | 274.0289 | 166.2606 | 169.0712 | 160.3697 | 187.706 | 218.6637 | 210.5167 | 214.1803 | 161.9288 | 122.4499 | 119.7773 | 115.2561 | 139.6303 | 263.1803 | 273.1515 | 250.3363 | 401.3697 | 268.0712 | 225.5969 | 247.1894 | 323.2651 | 161.8909 | 71.5546 | 72.8909 | 161.8909 | 140.8909 | 136.6303 | 148.3363 | 135.8151 | 126.3318 | 122.7773 | 115.4031 | 120.8575 | 79.8151 | 72.3697 | 72.8909 | 102.8151 | 121.5167 | 115.5591 | 105.0379 | 87.5546 | 306.5257 | 263.3787 | 272.7439 | 412.3742 | 175.5879 | 223.2985 | 216.9576 | 377.0289 | 306.0802 | 206.8909 | 232.8151 | 393.5212
1581 | 19.3854 | 163.3431 | 236.702 | 243.2212 | 209.8849 | 163.7788 | 189 | 184.1219 | 191.7178 | 185.9616 | 154.9842 | 149 | 214.2754 | 249.6862 | 171.4199 | 185.1061 | 283.4808 | 206.3973 | 162.4424 | 159.298 | 227.7178 | 209.6253 | 400.3589 | 402.237 | 208.237 | 253.8013 | 282.6185 | 280.5034 | 364.4199 | 155.0384 | 142.6795 | 131.2596 | 173.8239 | 181.8623 | 146.5801 | 149.1828 | 98.8397 | 287.4199 | 274.9165 | 289.0609 | 322.4424 | 155.4199 | 144.7178 | 144.1603 | 123.3815 | 89.2596 | 81.5801 | 76.3815 | 126.4199 | 237.3589 | 215.3431 | 227.1987 | 181.5801 | 164.5418 | 140.5418 | 149.7404 | 147.9774 | 211.8013 | 190.9165 | 240.763 | 186 | 180.1987 | 162.2529 | 189.6569 | 164.3815 | 182.1603 | 202.6185 | 244.6411 | 299.6411 | 183.0384 | 162.8172 | 281.4966 | 242.3205 | 254.702 | 260.1603 | 228.9007 | 400.3747 | 179.2212 | 178.6411 | 176.4808 | 258.8556 | 161.763 | 158.5801 | 161.7788 | 188.0226 | 212.1987 | 197.2822 | 194.4808 | 162.763 | 130.2438 | 118 | 109.0384 | 139.4808 | 246.8397 | 267.3657 | 250.0835 | 399.8397 | 260.0993 | 250.237 | 261.5801 | 338.5418 | 171.9391 | 73.8781 | 77.9007 | 171.9391 | 146.7404 | 136.0609 | 157.7404 | 143.298 | 121.4199 | 123.0993 | 108.3973 | 134.702 | 75.1219 | 80.4966 | 82.9391 | 108.6185 | 112.9007 | 125.0609 | 106.7404 | 82.3205 | 323.1828 | 304.7562 | 282.3205 | 430.1987 | 173.6953 | 225.5418 | 205.1377 | 366.4741 | 321.2822 | 220.298 | 223.0835 | 403.8397
1582 | 19.3964 | 152.5929 | 232.7537 | 247.9396 | 208.0149 | 172.8894 | 185.5427 | 180.3216 | 201.8643 | 184.618 | 146.0753 | 149 | 222.3318 | 235.1004 | 173.2965 | 162.8894 | 280.4322 | 212.382 | 162.3718 | 171.0251 | 249.5325 | 187 | 394.6282 | 417.1608 | 205.4424 | 251.8894 | 264.7992 | 273.3467 | 358.0855 | 158.5427 | 139.382 | 125.6533 | 163.0502 | 179.3467 | 155.5074 | 134.8392 | 103.0251 | 284.5427 | 271.8494 | 291.0502 | 318.4823 | 158.1608 | 154.8643 | 144.2965 | 125.211 | 86.6784 | 80.5678 | 73.4573 | 128.7286 | 242 | 208.0502 | 238.618 | 168.0353 | 161.3216 | 134.2965 | 148 | 144.7388 | 198.6533 | 202.1106 | 223.2463 | 188.5929 | 172.171 | 183.0604 | 201.2714 | 155.8392 | 173.6533 | 210.2965 | 240.4322 | 311.4219 | 182.6533 | 174.5678 | 282.171 | 236.9749 | 258.5325 | 252.9498 | 223.8392 | 419.1106 | 189.5576 | 177.4573 | 185.9647 | 284.0502 | 155.9145 | 159.7286 | 159.2212 | 178.7286 | 223.618 | 199.2361 | 201.3718 | 162.1004 | 123.0753 | 126.211 | 118.5929 | 132.1106 | 256.6431 | 245.5929 | 252.2863 | 388.0353 | 262.1357 | 250.4675 | 261.8643 | 336.1859 | 173.4071 | 77.2463 | 73.7035 | 173.4071 | 133.7639 | 135.4573 | 140.8745 | 142.9247 | 128.0502 | 127.2965 | 113.0855 | 132.9145 | 71.7537 | 93.7035 | 85.7035 | 113.2714 | 109.5678 | 116.6784 | 109.3216 | 86.4823 | 323.9647 | 298.0957 | 283.8894 | 432.5427 | 190.1106 | 220.5929 | 218.8894 | 387.2463 | 328.8541 | 229.4322 | 220.1004 | 414.0753
1583 | 19.4074 | 165.3334 | 231.8888 | 272.5556 | 246.3334 | 175.7778 | 183.6667 | 188.6667 | 199.4444 | 198.6667 | 157.1111 | 145 | 205.2222 | 257.2223 | 172.3333 | 170.6667 | 284.1111 | 207.6667 | 184.4445 | 176.3333 | 246.8888 | 186.1111 | 378.7778 | 395.5554 | 193.8889 | 268.1112 | 245.0001 | 280.7778 | 341.8888 | 145.9999 | 136.8889 | 118.6667 | 165.2222 | 185.4444 | 177.7778 | 141.3334 | 104.3333 | 276.8889 | 269.7779 | 292.7777 | 316.7777 | 162.3333 | 152 | 142.8889 | 137.7778 | 81.8889 | 84.4445 | 74 | 127 | 234.4444 | 221.7778 | 255.3334 | 159.4445 | 158.1111 | 124.8888 | 149.3333 | 138.4445 | 194.3334 | 189.4444 | 221.4445 | 188.8889 | 157.5556 | 170 | 202.1111 | 157.8889 | 171.1112 | 211.5556 | 233.4444 | 324.9999 | 179.6667 | 175.7778 | 268.4445 | 239.6667 | 261.6666 | 249.8889 | 242.3335 | 410.8889 | 203.3333 | 181.1111 | 201.6667 | 283.5555 | 159.2222 | 168.2223 | 153.4445 | 167.2221 | 222.5555 | 215.3333 | 215.4444 | 170.8888 | 134.1111 | 128.1111 | 109.111 | 145.6668 | 272 | 260.5556 | 262.9999 | 391.4446 | 255.2222 | 232.6667 | 256.3333 | 331.4444 | 172.2222 | 72.3333 | 74.6667 | 172.2222 | 133.2223 | 142.6667 | 129.8889 | 130.5556 | 124.4444 | 131.6667 | 115.1112 | 130.8888 | 83.7778 | 80.4444 | 79.1111 | 121.6667 | 109.8889 | 107 | 111 | 87.8888 | 338.3333 | 278.2223 | 301.4445 | 436.0001 | 178.3333 | 226.2222 | 228.4445 | 379.6667 | 352.7778 | 224.2222 | 239.5556 | 423.3333
1584 | 19.4184 | 174.2596 | 215.2354 | 283.8028 | 235.5957 | 178.0241 | 189.7404 | 192.9135 | 198.6539 | 224.8128 | 161.9376 | 135.4326 | 206.1972 | 270.4809 | 163.9759 | 168.5191 | 282.5191 | 216.4809 | 186.7787 | 174.8028 | 244.9234 | 197.3319 | 397.9234 | 372.3078 | 205 | 283.827 | 265.9376 | 283.4809 | 343.505 | 146.0482 | 141.6298 | 128.8511 | 158.4567 | 174.1489 | 174.4709 | 148.4326 | 107.8511 | 287.1589 | 286.495 | 288.6539 | 322.2695 | 158.0624 | 148.3702 | 131.6922 | 150.505 | 74.322 | 84.5191 | 66.7163 | 131.1348 | 232.3078 | 221.2354 | 267 | 179.5915 | 155.678 | 116.4809 | 144.6057 | 159.0241 | 200.0865 | 180.4567 | 242.9858 | 198.7022 | 173.8752 | 172.7404 | 210.5674 | 162.173 | 194.6156 | 211.9135 | 245.01 | 301.2978 | 202.1589 | 183.4809 | 289.5291 | 231.1007 | 254.5915 | 239.1007 | 275.8511 | 414.3078 | 179.0142 | 187.3078 | 203.1972 | 261.1248 | 166.6156 | 165.0383 | 172.7022 | 158.2213 | 225.7022 | 212.8894 | 203.6681 | 166.5674 | 131.173 | 122.4809 | 109.8993 | 153.0142 | 279.7404 | 268.6057 | 238.4085 | 432.0724 | 262.6156 | 238.7646 | 253.0241 | 327.0482 | 175.9135 | 70.8028 | 79.3702 | 175.9135 | 145.5191 | 147.3702 | 143.7404 | 130 | 116.8511 | 135.4567 | 121.6922 | 126.1348 | 91.4567 | 71.9376 | 75.7404 | 122.6681 | 115.5674 | 104.2837 | 110.5674 | 75.8028 | 328.6057 | 299.2213 | 298.2737 | 446 | 173.6539 | 217.5815 | 233.5433 | 372.5191 | 349.0624 | 221.5674 | 231.2113 | 417.2596
1585 | 19.4294 | 179.4425 | 224.2921 | 287.0354 | 216.1328 | 178.8407 | 194.4071 | 197.2832 | 201.5929 | 228.6459 | 167.1239 | 136.5664 | 223.1328 | 284.5045 | 174.416 | 171.3806 | 286.7877 | 230.9735 | 180.2832 | 164.9026 | 240.5131 | 207.7787 | 413.2123 | 382.8761 | 208.7522 | 288.3451 | 273.9381 | 291.4071 | 359.7522 | 143.9291 | 137.1858 | 133.2478 | 157.5929 | 182.8851 | 163.5045 | 154.2567 | 115.531 | 304.9381 | 285.2567 | 293 | 344.9381 | 151.469 | 154.6903 | 133.3186 | 154.5575 | 77.1328 | 76.1239 | 68.6903 | 132.6194 | 248.5045 | 220.9116 | 254.8052 | 182.0884 | 162.2478 | 144.5754 | 146.3806 | 147.6459 | 204.6284 | 180.531 | 251.2123 | 206.5575 | 181.9026 | 174.5929 | 208.9646 | 168.4425 | 195.584 | 203.1504 | 266.7168 | 286.8761 | 201.1769 | 189.0619 | 284.0795 | 233.5134 | 251.9291 | 251.8321 | 267.584 | 424.8761 | 178.2921 | 199.7522 | 188.2388 | 264.3541 | 176.0265 | 168.416 | 173.522 | 171.2832 | 239.1858 | 206.5929 | 202.1328 | 160.7433 | 128.531 | 122.9026 | 132.3451 | 138.2213 | 283.9381 | 249.7433 | 242.947 | 404.9732 | 275.7787 | 239.0884 | 279.1683 | 342.7522 | 176.531 | 82.0089 | 78.1858 | 176.531 | 132.9026 | 149.469 | 141.8407 | 136.5664 | 118.4336 | 127.5575 | 124.2567 | 137.9381 | 82.1504 | 75.7168 | 78.0619 | 112.2213 | 111.6194 | 114.0354 | 118.3451 | 87.0089 | 326.1593 | 293.9911 | 282.0974 | 454.4425 | 181.7522 | 213.8496 | 228.7787 | 361.7787 | 335.4336 | 226.0619 | 215.531 | 402.2743
1586 | 19.4404 | 186.1121 | 237.2989 | 283.4483 | 215.5977 | 180.2198 | 195.0374 | 204.8506 | 198.556 | 216.2198 | 164.5187 | 138.2242 | 239.9253 | 282.4109 | 173.7472 | 179.1121 | 287.4483 | 228.4483 | 185.9253 | 158 | 226.2155 | 211.1824 | 404.3692 | 385.9626 | 210.5934 | 292.9253 | 288.9583 | 278.0791 | 377.477 | 134.2572 | 130.1494 | 133.8132 | 153.5934 | 189.4857 | 161.8923 | 168.9253 | 108.2615 | 307.444 | 290.2989 | 292.5187 | 335.8923 | 164.5143 | 150.8549 | 153.1824 | 141.1494 | 82.8549 | 75.9253 | 69.6681 | 130.1824 | 262.7758 | 220.4109 | 238.1121 | 184.5143 | 170.033 | 156.967 | 147.8549 | 140.5143 | 204.2615 | 185.2945 | 233.7055 | 199.8879 | 179.8132 | 182.6264 | 190.4109 | 171.2615 | 181.8132 | 205.477 | 276.9583 | 299.5891 | 184.3692 | 172.5977 | 263.7385 | 250.3362 | 252.8462 | 276.0043 | 241.2989 | 416.4109 | 196.1121 | 191.0043 | 175.1824 | 277.2615 | 166.5187 | 181.7055 | 163.7011 | 185.1077 | 222.0417 | 206.444 | 214.1121 | 144.1494 | 117.4109 | 121.2945 | 127.7802 | 140.5934 | 285.4813 | 253.8462 | 265.8132 | 394.4309 | 280.1451 | 232.8506 | 271.6768 | 341.2242 | 173.1121 | 85.523 | 78.8506 | 173.1121 | 136.1077 | 143.2615 | 141.2945 | 159.4023 | 112.9253 | 130.5143 | 137 | 134.1868 | 77.2572 | 70.0747 | 78.9253 | 113.1494 | 105.3319 | 110.4483 | 115.7055 | 96.2989 | 327.1868 | 288.5517 | 289.9626 | 476.033 | 182.1494 | 223.0747 | 218.7055 | 381.5473 | 323.1868 | 209.5977 | 212.1121 | 403.4396
1587 | 19.4514 | 178.5575 | 245.2892 | 269.0256 | 202.9489 | 189.4808 | 189.5575 | 199.1278 | 199.4681 | 222.5192 | 158.0767 | 131.5064 | 228.1789 | 273.9617 | 163.3658 | 168.1022 | 282.8978 | 230.3147 | 178.6214 | 160.4681 | 261.3275 | 214.5703 | 411.4553 | 385.0256 | 205.5319 | 294.5064 | 296.1022 | 272.3658 | 375.7108 | 138.0383 | 126.9872 | 135.0511 | 166.7956 | 176 | 152.4808 | 164.5831 | 99.5192 | 301.1022 | 283.4808 | 284.1022 | 321.0511 | 186.885 | 143.4681 | 155.5831 | 141.9361 | 69.5447 | 79.9744 | 69.9361 | 133.0767 | 255.1789 | 218.8722 | 229.5703 | 186.1533 | 161.2811 | 157.8211 | 147.8722 | 151.0383 | 202.4297 | 176.1917 | 236.885 | 203.4936 | 171.1789 | 185.5958 | 197.2636 | 166.4681 | 197.8595 | 205.1917 | 283.115 | 312.9744 | 187.0128 | 162.9106 | 269.0256 | 221.7492 | 281.3658 | 254.1022 | 240.4042 | 430.1742 | 192.5064 | 171.0767 | 190.9106 | 276.9106 | 181.302 | 178.9617 | 171.0128 | 181.6853 | 218.8467 | 224.2892 | 225.3147 | 145.9233 | 124.2636 | 128.9744 | 126.4169 | 148.3658 | 279.5831 | 274.4681 | 268.5319 | 423.6853 | 285.0639 | 236.5064 | 258.7445 | 349.3147 | 159.6342 | 71.4936 | 76.5831 | 159.6342 | 147.4936 | 134.5192 | 148.9744 | 167.6086 | 122.8978 | 148.4425 | 135.0256 | 140.3531 | 84.9872 | 76.3914 | 81 | 121.8339 | 118.885 | 118.7317 | 101.6086 | 87.5064 | 334.8339 | 288.6597 | 287.0511 | 480.115 | 185.4042 | 214.5831 | 219.9106 | 397.6342 | 317.5064 | 199.4169 | 219.8595 | 402.7108
1588 | 19.4624 | 179.5649 | 242.2698 | 281.1653 | 214.0826 | 184.4233 | 189.5531 | 201.6476 | 209.0826 | 232.1299 | 159.5885 | 130.4941 | 212.4823 | 293.6948 | 172.9882 | 162.0236 | 285.4351 | 237.4115 | 187.2125 | 163.5059 | 266.3997 | 206.9764 | 396.7875 | 396.6594 | 206.0354 | 291.9764 | 281.4233 | 279.4587 | 373.1653 | 143.0944 | 136.6003 | 127.9646 | 185.0236 | 184.6003 | 157.0354 | 151.9292 | 101.0354 | 305.6476 | 288.5413 | 284.0944 | 311.4351 | 172.7285 | 154.6003 | 154.059 | 149.5295 | 73.0826 | 86.0472 | 79.5531 | 124.4705 | 242.0118 | 228.4941 | 234.624 | 183.1062 | 159.2361 | 157.3288 | 159.0118 | 148.5059 | 211.0354 | 173.6476 | 242.9646 | 186.7993 | 171.6712 | 172.941 | 210.941 | 164.4469 | 209.5059 | 189.4941 | 271.9764 | 326.1299 | 192.0708 | 157.3524 | 272.059 | 225.7774 | 279.8583 | 233.8583 | 236.8701 | 445.376 | 185.4233 | 179.1653 | 194.4587 | 290.0708 | 174.2344 | 179.4705 | 172.5295 | 177.624 | 237.5767 | 231.3879 | 255.683 | 165.6594 | 136.9292 | 120.8819 | 133.5059 | 160.0118 | 278.5649 | 273.9646 | 255.376 | 410.5177 | 286.5767 | 240.0472 | 260.7875 | 364.5059 | 169.2361 | 77.059 | 76.5767 | 169.2361 | 153.5649 | 136.5413 | 150.4941 | 148.8819 | 138.5885 | 142.8819 | 125.9174 | 146.941 | 85.4941 | 79.941 | 81.5059 | 125.3879 | 118.8938 | 129.376 | 104.6358 | 85.9882 | 340.9174 | 289.1535 | 288.5649 | 453.2934 | 177.8229 | 221.1535 | 247.2361 | 379.9174 | 314.9764 | 211.5649 | 238.5885 | 390.0472
1589 | 19.4734 | 168.936 | 234.4003 | 278.936 | 233.4366 | 185.7729 | 194.5182 | 204.6725 | 214.4454 | 220.1995 | 157.1543 | 131.5546 | 205.2999 | 294.9724 | 175.1092 | 162.4454 | 280.5182 | 235.7001 | 192.5634 | 161.9272 | 261.1092 | 200.3363 | 396.1004 | 415.1819 | 212.1092 | 295.1819 | 284.3275 | 268.7453 | 388.0364 | 148.9272 | 136.1907 | 133.2911 | 174.5634 | 175.3814 | 170.3639 | 146 | 106.0728 | 301.4178 | 298.2183 | 289.9272 | 333.4643 | 172.2823 | 161.4454 | 152.2635 | 150.4454 | 79.4818 | 85.8545 | 79.8545 | 133.9185 | 241.4454 | 225.409 | 253.2911 | 202.3639 | 170.1907 | 143.5182 | 163.6274 | 146.409 | 216.0728 | 190.6637 | 232.2271 | 187.1004 | 181.3363 | 161.2635 | 211.1819 | 164.1455 | 200.6725 | 205.064 | 280.8821 | 312.1267 | 195.9272 | 167.2823 | 279.0728 | 255.2183 | 287.6914 | 216.8181 | 250.3551 | 436.5546 | 186.7729 | 194.5546 | 185.2999 | 277.345 | 167.0276 | 173.3726 | 181.7365 | 169.8632 | 224.8268 | 225.1092 | 248.7177 | 170.7089 | 115.9724 | 117.2183 | 128.8181 | 150.6361 | 281.409 | 278.2547 | 255.8821 | 410.9636 | 291.9636 | 255.4003 | 254.9185 | 354.1179 | 182.2635 | 80.4454 | 79.3726 | 182.2635 | 163.6637 | 142.591 | 140.6725 | 140.5546 | 139.2635 | 128.3363 | 121.591 | 136.8181 | 87.0728 | 76.0364 | 79.409 | 108.7453 | 126.064 | 127.2911 | 111.8545 | 95.3639 | 332.9636 | 297.8545 | 292.9636 | 439.1092 | 176.9912 | 237.6274 | 242.6449 | 398.3915 | 313 | 228.9185 | 232.527 | 416.2823
1590 | 19.4844 | 158.7744 | 233.3293 | 269.8353 | 227.494 | 191.4085 | 207.2011 | 189.1037 | 208.2256 | 232.951 | 159.1402 | 134.061 | 195.2256 | 274.1037 | 176.939 | 172.6707 | 285.7744 | 252.1402 | 186.8353 | 163.1829 | 253.9208 | 186.4513 | 410.939 | 416.2866 | 204.9208 | 301.061 | 276.555 | 255.1037 | 375.2074 | 145.4085 | 128 | 155.8536 | 163 | 175.9755 | 169.3903 | 156.6097 | 103.2256 | 279.1037 | 294.9818 | 293.3048 | 348.9208 | 177.0855 | 161.061 | 148.1219 | 148.4695 | 83.2439 | 82.5305 | 71.2256 | 131.2074 | 238.4085 | 232.0182 | 249.8598 | 207.7561 | 170.4878 | 141.8781 | 161.1647 | 157.7926 | 212.6952 | 198.1829 | 237.7316 | 200.8781 | 175.939 | 156.5915 | 201.6769 | 161.1037 | 189.8781 | 224.2439 | 261.8233 | 302.7926 | 193.4695 | 183.2256 | 284.7134 | 258.8781 | 299.1037 | 240.6462 | 264.5122 | 459.2194 | 205.1402 | 199.1829 | 192.7316 | 276.9755 | 191.0792 | 170 | 167.3111 | 169.9755 | 212.3048 | 238.6097 | 243.3841 | 155.5732 | 103.061 | 118.2256 | 118.6952 | 143.1219 | 274.7561 | 276.5122 | 254.3293 | 422.7316 | 269.2501 | 252.7379 | 270.7744 | 342.939 | 165.9208 | 74.7561 | 68.5732 | 165.9208 | 169.5915 | 161.445 | 136.7744 | 132.9818 | 144.1402 | 130.8963 | 124.5305 | 142.0792 | 84.7561 | 77.5305 | 84.9573 | 108.8963 | 132.5122 | 131.2866 | 113.3048 | 98.6342 | 351.6279 | 305.1402 | 290.4085 | 437.2256 | 176.7379 | 254.2621 | 224.2439 | 386.988 | 317.2439 | 243.7134 | 224.7744 | 440.1829
1591 | 19.4954 | 165.1711 | 240.9116 | 289.1122 | 232.0827 | 187.2861 | 211.4867 | 202.5399 | 205.6283 | 228.8051 | 183.6549 | 142.5989 | 212.1683 | 273.4278 | 176.5428 | 184.0856 | 286.7433 | 244.6312 | 193.6283 | 183.3688 | 236.8584 | 195.3688 | 420.8555 | 399.9734 | 209.0266 | 302.5428 | 260.3717 | 285.366 | 363.5428 | 135.8584 | 124.7433 | 141.3184 | 173.8555 | 189.9144 | 160.5428 | 164.9144 | 102.7994 | 285.4838 | 283.7433 | 287.1445 | 353.5694 | 163 | 158.7433 | 147.8289 | 156.6844 | 87 | 73.7728 | 69.7139 | 126.0561 | 249.4838 | 218.8317 | 261.711 | 186.0884 | 170.9439 | 151.3983 | 141.3451 | 152.0884 | 209.0856 | 189.0589 | 244.6578 | 200.0856 | 185.8555 | 175.3688 | 208.5399 | 166.3983 | 196.1416 | 229.6283 | 243.5989 | 297.6312 | 191.3717 | 185.5133 | 280.9439 | 270.0266 | 273.4601 | 262.2006 | 260.2567 | 467.6873 | 206.8584 | 196.5722 | 191.5162 | 281.1445 | 173.404 | 176.5133 | 163.6549 | 190.4278 | 224.5989 | 231.7167 | 240.174 | 158.2272 | 111.5989 | 121.0561 | 104.23 | 141.7433 | 274.7994 | 283.1122 | 234.7728 | 419.8878 | 274.9677 | 241.5428 | 281.5133 | 338.2006 | 152.6578 | 74.7994 | 78.2833 | 152.6578 | 171.5428 | 146.6901 | 139.9144 | 131.5133 | 141.5162 | 146.2272 | 133.1416 | 143.9439 | 78.2861 | 79.0856 | 81.6873 | 104.6873 | 121.7433 | 126.3717 | 126.1416 | 97.3422 | 358.6873 | 307.9439 | 283.5722 | 460.6816 | 161.7433 | 269.7994 | 239.9411 | 393.4515 | 321 | 233.4306 | 218.1445 | 423.4601
1592 | 19.5064 | 178.6564 | 236.5727 | 264.9207 | 236.2335 | 195.5463 | 206 | 190.6916 | 201.4493 | 209.9912 | 181.5683 | 139.5639 | 228.3348 | 291.3216 | 171.4493 | 180.0044 | 289.5507 | 234.9956 | 208.8767 | 193.5595 | 249.4273 | 200.0088 | 434.9956 | 403.9868 | 222.2203 | 315.7665 | 269.5463 | 292.6872 | 368.4405 | 144.5419 | 132.5463 | 131.3216 | 172.4537 | 177.3436 | 167.1057 | 157.3392 | 94.3436 | 327.0837 | 307.6432 | 287.9912 | 350.5683 | 163 | 152.1145 | 141.0044 | 152.8987 | 93.6608 | 80.4317 | 72 | 128.6696 | 282.7577 | 235.9692 | 276.6696 | 186.5419 | 172.2159 | 158.2247 | 135.6608 | 154.2071 | 224.4317 | 196.7621 | 277.0793 | 212.6564 | 191.1145 | 179.4537 | 212.7885 | 168.7841 | 210.2159 | 223.7841 | 278.3084 | 304.6476 | 196.1057 | 182.1189 | 290.5419 | 256.022 | 264.2159 | 259.5551 | 243.5727 | 455.6696 | 212.2115 | 187.5595 | 186.5507 | 277.5507 | 153.2203 | 172.0088 | 168.7841 | 191.1145 | 235.4405 | 242.978 | 238.7665 | 162.1145 | 129.6564 | 122.5595 | 104.326 | 136.7797 | 271.3392 | 285.5639 | 263.0793 | 411.8855 | 270.4669 | 255.3216 | 269.7929 | 340.5507 | 169.5375 | 73.5595 | 98.6608 | 169.5375 | 159.7885 | 146.978 | 144.5507 | 137.5551 | 136.5507 | 138.4581 | 134.4493 | 141.8855 | 84.8811 | 85.5507 | 79.1057 | 110.9868 | 120.6652 | 125 | 128.0044 | 93.2291 | 354.9956 | 303.1101 | 289.5463 | 469.5683 | 171.2115 | 265.7841 | 238.8987 | 401.6872 | 328.2159 | 218.1145 | 227.3172 | 419.7665
1593 | 19.5174 | 189.7021 | 228.943 | 264.3679 | 232.8083 | 192.0855 | 205.4326 | 188.4249 | 213.4534 | 209.4896 | 167.1632 | 156.3964 | 231.5389 | 303.1347 | 172.1062 | 181.1062 | 285.4896 | 231.057 | 211.4896 | 178.0855 | 261.5959 | 197.6736 | 420.8445 | 393.3057 | 214.9223 | 318.6243 | 276.8653 | 266.6243 | 361.2202 | 139.4119 | 142.1347 | 143.1347 | 183.8575 | 192.3964 | 178.8083 | 153.7021 | 93.5104 | 342.3549 | 298.93 | 302.8083 | 353.1839 | 162.4326 | 151.8368 | 144.943 | 159.3187 | 85.9508 | 81.7876 | 73.1347 | 129.9715 | 262.987 | 236.1995 | 268.8938 | 206.943 | 184.8083 | 156 | 144.9715 | 164.0285 | 232.0285 | 196.2487 | 288.9793 | 214.057 | 198.7798 | 158.5181 | 210.943 | 173.2124 | 200.114 | 229.9145 | 265.2565 | 325.4041 | 198.1632 | 184.6451 | 292.2202 | 268.9068 | 267.7306 | 249.2202 | 251.2617 | 461.5104 | 214.057 | 199.886 | 192.7021 | 274.057 | 162.943 | 156.057 | 164.1347 | 193.1062 | 240.7021 | 247.114 | 236.5181 | 159 | 136.7306 | 119 | 124.6166 | 137.8368 | 274.5104 | 265.5181 | 276.114 | 418.9715 | 272.127 | 264.2979 | 240.114 | 375.6373 | 181.4611 | 70 | 86.4119 | 181.4611 | 147.1632 | 158.4896 | 167.1555 | 151.0492 | 157.4534 | 120.3264 | 137.943 | 148.9715 | 84.057 | 92.8368 | 82.8653 | 129.8083 | 127.6736 | 140.886 | 129.1062 | 92.1062 | 365.2409 | 286.9793 | 293.4611 | 427.3627 | 183.2694 | 260 | 228.2979 | 388.1062 | 335.1347 | 234.2902 | 230.6528 | 441.9145
1594 | 19.5284 | 198.5354 | 232.1018 | 272.6681 | 245.5354 | 189.9558 | 189.9291 | 210.9558 | 213.8274 | 198.3628 | 166.7389 | 165.7256 | 239.0575 | 304.5796 | 180.6372 | 183.2611 | 293.4912 | 235.4337 | 197.4646 | 152.7698 | 263.0575 | 215.9116 | 407 | 405.5621 | 227.7081 | 306.0442 | 280.6372 | 283.0267 | 365.1726 | 151.5046 | 144.1593 | 133.5663 | 175.814 | 205.2035 | 158.4954 | 156.7389 | 96.5221 | 351.8098 | 277.7389 | 305.6814 | 361.6814 | 163.1593 | 138.9291 | 141.146 | 164.6239 | 81.2168 | 87.9116 | 72.8407 | 141.1151 | 240.3761 | 206.8849 | 266.7389 | 203.7256 | 178.9867 | 164.1151 | 139.3053 | 140.1326 | 218.5663 | 190.7965 | 264.9867 | 211.4779 | 202.3628 | 160.593 | 215.8407 | 182.7389 | 184.5221 | 212.3362 | 257.1284 | 315.8274 | 190.7832 | 180.9867 | 268.9291 | 274.8716 | 275.2744 | 258.3895 | 273.6372 | 476.1151 | 214.3761 | 217.2168 | 200.3761 | 285.9691 | 175.3761 | 170.8674 | 175.4337 | 199.8982 | 252.4337 | 233.2611 | 237.4337 | 172.3319 | 129.2035 | 131.1726 | 128.6239 | 141.1593 | 294.3319 | 262.9558 | 240.2344 | 422.5796 | 287.2611 | 262.4204 | 230.3186 | 378.1326 | 173.9425 | 72.8982 | 77.0575 | 173.9425 | 152.5354 | 150.2611 | 158.3937 | 151.7256 | 158.407 | 116 | 139.3628 | 152.5796 | 78 | 88.6239 | 69.2477 | 128.0442 | 130.8407 | 130.3937 | 121.9867 | 89.6239 | 390.2877 | 303.5621 | 314.9249 | 432.9824 | 185.5796 | 244.3495 | 218.3053 | 414.6063 | 339.4779 | 253.0575 | 221.4204 | 431.2919
1595 | 19.5394 | 202.2242 | 230.5919 | 254.1209 | 240.937 | 200.9194 | 166.5693 | 195.2821 | 209.7355 | 212.1663 | 168 | 167.2872 | 242 | 296.1209 | 178.0806 | 182 | 315.8388 | 242.4081 | 197.3275 | 165.8615 | 267.7758 | 223.6322 | 407.5919 | 416.937 | 211.2191 | 302.1839 | 313.597 | 289.529 | 358.0176 | 163.4887 | 145 | 144.9421 | 178.534 | 198.6322 | 158.9421 | 168.6549 | 101.1033 | 346.2418 | 293.2066 | 289.7934 | 342.8337 | 166.9597 | 139.2469 | 148.2066 | 158.8161 | 83.8161 | 95.0403 | 81.471 | 125.0982 | 243.8161 | 230.0051 | 278.063 | 200.5516 | 179.2872 | 156.9773 | 137.7355 | 143.9421 | 221.6549 | 187.3048 | 267.6549 | 200.3854 | 197.2242 | 166.0403 | 212.0403 | 194.063 | 201.534 | 215.4937 | 257.9773 | 307 | 185.2242 | 187.2066 | 273.9824 | 274.6146 | 287.3275 | 268.6322 | 251.5466 | 444.7078 | 214.8564 | 211.529 | 191.3854 | 299.5919 | 173.4887 | 184.2242 | 174.7128 | 196.0806 | 260 | 240.8791 | 253.2872 | 165.4257 | 129.7355 | 139.4081 | 114.529 | 144.9597 | 307.5516 | 278.6549 | 256.7406 | 420.0403 | 279.4887 | 264.9597 | 258.6373 | 363.5919 | 161.529 | 77.9597 | 70.529 | 161.529 | 149.1209 | 148.4081 | 149.6952 | 137.3048 | 133.4257 | 117.1839 | 130.6725 | 147.0806 | 79.1839 | 91.6952 | 77.1663 | 114.1209 | 136.5113 | 124.6549 | 114 | 89.7355 | 357.6448 | 306.6499 | 327.6725 | 474.126 | 188.3678 | 255.4937 | 204.8967 | 421.529 | 343.1839 | 237.6499 | 221 | 421.7355
1596 | 19.5504 | 205.8338 | 232.2085 | 281.2327 | 235.8127 | 201.9788 | 188.8158 | 191.8762 | 218.4381 | 222.7915 | 160.145 | 150.6435 | 251.0635 | 304.5015 | 180.0423 | 185.6254 | 302.8519 | 244.4169 | 201 | 183 | 279.8762 | 207.4985 | 408.6042 | 425.71 | 208.1269 | 278.2265 | 303.7884 | 279.3746 | 363.9185 | 162.0212 | 146.8127 | 137.4561 | 181.1238 | 193.3746 | 185.8973 | 153.0392 | 98.7492 | 338.2296 | 285.4773 | 292.0846 | 322.5408 | 162.9577 | 147.6042 | 153.1662 | 163.4381 | 80.5831 | 76.0815 | 84.3746 | 122.0846 | 251.4592 | 259.6677 | 278.3535 | 199.9788 | 186.8127 | 135.3111 | 149.4592 | 151.5619 | 238.0635 | 182.6042 | 269.5619 | 218.7946 | 205.6677 | 186.9608 | 220.8762 | 194.9577 | 219.8338 | 211.6646 | 239.9365 | 316.0635 | 190.0423 | 197 | 287.4169 | 274.3323 | 278.9154 | 275.4592 | 267.8369 | 446.7946 | 206.5619 | 189.29 | 202.5439 | 308.4592 | 175.6465 | 185.4169 | 153.8942 | 181.1238 | 270.8762 | 232.4985 | 248.1238 | 161.2508 | 113.6646 | 135.9788 | 123.1058 | 134.9154 | 279.1842 | 276.3323 | 267.3111 | 413.7704 | 282.855 | 282.71 | 272.1662 | 359.7704 | 170.1058 | 77.5831 | 64 | 170.1058 | 144.2085 | 147.3958 | 158.6254 | 148.9185 | 124.4169 | 128.8762 | 124.5831 | 155.6889 | 79.3958 | 102.4381 | 92.0212 | 117.6677 | 134.9577 | 130.7915 | 118.8338 | 91.7915 | 346.7523 | 305.4804 | 316.145 | 468.6646 | 199.6677 | 248.0392 | 221.7523 | 408.3535 | 330.1027 | 221.9788 | 230.0635 | 444.3354
1597 | 19.5614 | 203.4513 | 236.6991 | 277.7226 | 239.4661 | 209.2478 | 203.4513 | 204.5487 | 202.2713 | 224.4661 | 161.7817 | 155.7287 | 242.82 | 319.5487 | 183.3835 | 194.7817 | 317.36 | 245.3835 | 199.767 | 189.1652 | 280.8348 | 212.18 | 456.4722 | 400.2417 | 214.4513 | 289.127 | 281.5339 | 310.9087 | 392.2626 | 162.1504 | 143.6843 | 148.6609 | 177.0826 | 189.767 | 189.4513 | 152.7965 | 98.3157 | 317.5722 | 286.3304 | 298.767 | 340.4278 | 165.1652 | 157.8644 | 148.767 | 174.3983 | 85.7817 | 73.0148 | 71.5191 | 138.0148 | 247.7522 | 248.1209 | 288.18 | 188.1356 | 171.9704 | 149.8113 | 166.0974 | 152.3157 | 231.6696 | 184.8496 | 273.3983 | 221.5869 | 200.9026 | 186.5869 | 220.6017 | 184.8348 | 234.7139 | 208.2478 | 261.127 | 305.9704 | 211.2626 | 182.2035 | 287.1504 | 254.4365 | 280.2478 | 270.5191 | 286.3009 | 458.8348 | 205.4661 | 187.0148 | 199.9704 | 298.5869 | 175.6843 | 180.8348 | 169.8939 | 203.5931 | 260.7374 | 232.8644 | 258.8791 | 156.7522 | 118.2626 | 128.4513 | 133.6165 | 139.3304 | 283.6609 | 266.3009 | 265.7817 | 427.6461 | 304.6461 | 257.8582 | 261.6017 | 348.3687 | 170.7522 | 84.0148 | 72.0148 | 170.7522 | 155.4809 | 149.4661 | 146.82 | 156.9174 | 129.0826 | 124.9026 | 132.2478 | 157.2183 | 84.5487 | 90.5869 | 95.8496 | 112.9026 | 131 | 121.9852 | 123.233 | 88.5339 | 356.0678 | 313 | 320.2478 | 441.2035 | 198.6017 | 237.3157 | 254.4957 | 421.8791 | 330.8644 | 248.9765 | 229.8348 | 411.3774
1598 | 19.5724 | 187.4237 | 245.2881 | 273.3186 | 252.3186 | 203.0526 | 202.5153 | 190.3932 | 204.4627 | 226 | 157.8254 | 164.2271 | 225.1966 | 309.7949 | 183.6288 | 188.939 | 333.1136 | 265.122 | 199 | 193.6288 | 254.3627 | 205.2797 | 423.7475 | 415.1 | 216.0305 | 304.7424 | 277.4847 | 310.6204 | 384.7644 | 169.1746 | 144.1441 | 158.9695 | 170.1966 | 185.2271 | 188.5153 | 168.2881 | 101.6288 | 355.8203 | 312.2356 | 312.4627 | 332.9915 | 170.8864 | 160.8559 | 161.8339 | 165.7949 | 93.1441 | 80.5153 | 70.0305 | 127.9085 | 253.3186 | 237 | 298.8864 | 205.2661 | 181.4932 | 148.2797 | 153.5068 | 151.8559 | 222.1136 | 200.4627 | 259.1356 | 223.9474 | 199.0305 | 188.9474 | 220.4017 | 185.4017 | 233.8254 | 219.0305 | 274.2271 | 307.9474 | 192.4458 | 176.7729 | 276.5678 | 260.4627 | 293.5458 | 265.2576 | 287.7729 | 434.878 | 200.7119 | 173.1356 | 195.6593 | 295.2881 | 179.9169 | 177.6288 | 170.2797 | 197.4763 | 261.9474 | 239.6288 | 252.3932 | 158.5458 | 117.0526 | 130.0305 | 122.0526 | 150.7729 | 278.878 | 236.3322 | 264.9695 | 444.2881 | 296.1356 | 239.1441 | 279.6373 | 349.9169 | 181.3491 | 75.1661 | 75.1136 | 181.3491 | 163.2576 | 144.0831 | 151.8339 | 161.9169 | 136.6593 | 130.5763 | 132.3407 | 149.2271 | 82.9695 | 86.0305 | 91.9695 | 115.4322 | 131.6288 | 122.6593 | 123.3712 | 95.1746 | 347.9695 | 302.3102 | 314.0526 | 472.8729 | 190.8559 | 243.7729 | 259.7119 | 443.061 | 326.3102 | 245.6204 | 233.5458 | 394.9474
1599 | 19.5834 | 165.8957 | 264.3865 | 281.9233 | 236.5614 | 197.9233 | 190.5368 | 196.0276 | 214.9233 | 224.0767 | 170.3098 | 153.6656 | 227.0521 | 334.6963 | 183.3589 | 186.8466 | 321.7423 | 248.7914 | 207.9755 | 181.1779 | 268.5675 | 212.0276 | 407.5153 | 442.7699 | 209.3834 | 296.3067 | 297.7975 | 290.9479 | 370.1534 | 153.4847 | 146.6411 | 143.819 | 169.4877 | 197.1043 | 181.0245 | 157.8957 | 102.6411 | 360.7147 | 287.0982 | 310.1012 | 338.592 | 177.1288 | 156.4356 | 162.9479 | 156.7178 | 82.819 | 83.2822 | 67.8712 | 140.7975 | 269.6166 | 238.2822 | 282.6902 | 204.8957 | 171.2025 | 147.3344 | 163.1564 | 153.2055 | 217.1534 | 212.8466 | 258.2577 | 226.5123 | 185.9724 | 188.3067 | 223.6411 | 190.5644 | 213.6135 | 228.411 | 273.2822 | 307.3067 | 173.0767 | 198.8742 | 281.8988 | 265.7945 | 296.0767 | 289.0798 | 277.1779 | 440.9509 | 205.3344 | 189.5675 | 191.9479 | 290.0245 | 182.0767 | 165.1779 | 177.0276 | 183 | 286.9509 | 243.2055 | 236.2301 | 154.0245 | 129.8742 | 114.408 | 122.6933 | 138.2546 | 272.1288 | 255.2607 | 274.8221 | 433.2546 | 284.3589 | 255.1043 | 276.3313 | 352.7178 | 182.0245 | 78.5399 | 100.2853 | 182.0245 | 160.1534 | 141.9233 | 140.7669 | 167.2822 | 126.5368 | 146.3344 | 130.9233 | 154.6933 | 84.4877 | 74.2546 | 78.7423 | 113.9479 | 126.8712 | 129.2055 | 124.9233 | 91.0245 | 353.9755 | 306.8988 | 298.6656 | 468.7914 | 198.6166 | 231.2546 | 263.6933 | 431.6902 | 319.3589 | 230.4356 | 231.589 | 394.3067
1600 | 19.5944 | 176.9485 | 250.7846 | 281.6932 | 245.562 | 212.7213 | 193.4543 | 214.1475 | 197.0515 | 218.4262 | 165.6253 | 140.5059 | 231 | 328.8244 | 177.1194 | 192.9204 | 333.6417 | 249.3349 | 220.8407 | 199.4824 | 285.9602 | 204.6651 | 418.3466 | 440.1194 | 209.8806 | 318.7892 | 329.6019 | 285.6932 | 390.2154 | 142.6534 | 143.733 | 146.8009 | 175.267 | 199.1194 | 162.9321 | 166.9883 | 105.6136 | 319.8642 | 286.6019 | 303.3466 | 337.8923 | 173.466 | 159.5738 | 137.438 | 153.3864 | 85.1475 | 83.3466 | 70.2272 | 137.9719 | 256.0515 | 250.1077 | 285.4145 | 199.6136 | 177.9485 | 169.6417 | 163.2389 | 146.5059 | 218.267 | 209.7728 | 271.1873 | 221.3864 | 182.8806 | 163.0914 | 229.2272 | 210.9485 | 192.5855 | 217.6253 | 293.6019 | 315.4145 | 192.9086 | 187.7846 | 280.8127 | 274.4543 | 295 | 273.2506 | 273.267 | 454.267 | 190.3981 | 193.8923 | 178.1991 | 280.4262 | 183.6136 | 163.8806 | 163.7846 | 185.6136 | 281.9719 | 258.7213 | 238.2272 | 160.1077 | 142.9602 | 108.5738 | 117.1991 | 133.9204 | 268.466 | 288.7213 | 268.2787 | 432.8407 | 272.2389 | 270.8407 | 274.8407 | 373.562 | 179.6136 | 86.9602 | 101.9321 | 179.6136 | 169.1077 | 151.4941 | 130 | 137.9438 | 126.1873 | 145.1194 | 133.9602 | 143.9719 | 90.9204 | 78.4145 | 78.2272 | 104.7728 | 120.0796 | 122.5059 | 129.9204 | 82.0796 | 374.0281 | 305.1593 | 329.2833 | 475.8688 | 190.9321 | 230.1873 | 277.8009 | 408.9321 | 320.9602 | 237.4941 | 220.8127 | 391.9602
1601 | 19.6054 | 181.6745 | 255.6422 | 298.3078 | 242.3549 | 216.0059 | 193.0088 | 213.6745 | 210.9647 | 228.648 | 156.6686 | 155.3049 | 210.3637 | 294.698 | 183.3196 | 200.3255 | 327.6892 | 262.6598 | 211.6863 | 197.6892 | 290.3284 | 206.651 | 463.2667 | 409.0412 | 206.3431 | 303.0441 | 318.698 | 288.9941 | 402.6657 | 155.648 | 143.3314 | 134.6922 | 174.3373 | 188.0118 | 171.9765 | 176 | 104.3373 | 306.6657 | 294.3373 | 290.352 | 339.9882 | 168.0029 | 179.3078 | 140.6451 | 150.0029 | 85.3402 | 83.6657 | 72.3343 | 140.651 | 273.9588 | 250.0088 | 283.3461 | 193.6775 | 170.6922 | 167.0176 | 153.0059 | 136.6745 | 208.0176 | 222.3108 | 269.0088 | 196.701 | 191.9912 | 177.2961 | 210.698 | 211.6804 | 208.6333 | 217.9882 | 283.3637 | 313.3461 | 201.3373 | 185.9853 | 287.651 | 265.3549 | 293.6686 | 273.3108 | 278.9941 | 459.3284 | 185.9912 | 186.6686 | 185.648 | 287.3196 | 194.3196 | 173.6569 | 191.9412 | 202.3108 | 277.9941 | 262.6716 | 268.9588 | 165.3343 | 124.6951 | 125.6451 | 119.3225 | 142.6569 | 280.9765 | 280.6892 | 246.3549 | 445.6539 | 278.648 | 264.349 | 268.349 | 380.3402 | 186.3255 | 86.0029 | 93.0029 | 186.3255 | 167.6775 | 138.6922 | 137.3225 | 148.6275 | 135.3255 | 142.6657 | 131.0059 | 149.3137 | 103.651 | 78.3431 | 83.6627 | 109.3225 | 118 | 114.6716 | 128.0059 | 95.9765 | 366.0235 | 329.6245 | 322.7039 | 464.701 | 172.0176 | 240.6569 | 255.0412 | 399.3373 | 333.9824 | 265.9647 | 230.3137 | 406.3137
1602 | 19.6164 | 212.2209 | 254.5083 | 292.0845 | 240.4238 | 205.8643 | 213.0512 | 200.1524 | 206.0506 | 244.4917 | 166.1697 | 160.9321 | 225.7631 | 313.831 | 170.3726 | 201.644 | 333.5595 | 271.7798 | 209.0679 | 169.6607 | 263.525 | 202.5083 | 465.6607 | 402.4578 | 221.9834 | 290.0339 | 303.2542 | 286.9321 | 396.8982 | 158.6101 | 154.1697 | 147.6953 | 175.7119 | 182.644 | 180 | 172.6101 | 109.1018 | 336.831 | 304.5256 | 297.5595 | 353.4917 | 194.1191 | 164.9488 | 140.5422 | 163.2375 | 86.3899 | 81.9661 | 75.3899 | 131.7625 | 261.5589 | 244.2881 | 265.4405 | 202.2036 | 180.3054 | 167.1018 | 144.2202 | 144.1697 | 227.0852 | 227.9661 | 270.7458 | 176.1863 | 178.7285 | 179.1524 | 217.6274 | 196.1524 | 235.5935 | 222.678 | 283.1697 | 340.187 | 214.9155 | 173.3726 | 310.6274 | 260.0339 | 295.7119 | 274.2202 | 266.0845 | 452.8643 | 203.2375 | 183.2881 | 184.5422 | 319.1191 | 179.3387 | 181.7458 | 174.7112 | 191.6946 | 256.9488 | 257.6101 | 274.8643 | 183.3054 | 129.9155 | 122.8303 | 110.7964 | 147.356 | 303.9155 | 256.7285 | 257.3054 | 422.8809 | 282.9661 | 285.4411 | 287.4072 | 367.8303 | 189.678 | 86.356 | 94.0339 | 189.678 | 166.0339 | 145.5935 | 155.9155 | 149.7964 | 135.2881 | 126.7285 | 135.1018 | 148.5422 | 95.4405 | 76.356 | 81.6101 | 106.8982 | 121.3899 | 124.5256 | 132.7798 | 102.644 | 350.5422 | 308.7452 | 307.9661 | 458.4238 | 178.8816 | 233.1524 | 252.5256 | 382.4065 | 351.5256 | 274.6101 | 241.0679 | 434.0173
1603 | 19.6274 | 217.646 | 276.2302 | 304.3275 | 236.7876 | 219.2567 | 219.8584 | 239.8674 | 200.0708 | 235.1946 | 185.4956 | 176.947 | 236.6195 | 290.7255 | 188.2302 | 187.8849 | 342.7611 | 264.646 | 195.8141 | 166.2124 | 281.7523 | 217.3275 | 448.407 | 405.3097 | 231 | 311.7081 | 320.3275 | 301.5664 | 382.2654 | 170.8054 | 168.6638 | 147.646 | 191.4956 | 188.2124 | 172.407 | 146.8406 | 98.1946 | 320.6282 | 314.8319 | 295.0265 | 367.6638 | 198.7168 | 165.0443 | 170.8232 | 172.0708 | 99.0443 | 82.3805 | 60.4336 | 128.4513 | 260.0443 | 265.7789 | 286.6107 | 185.2211 | 173.8141 | 170.6903 | 143.7611 | 149 | 212.7698 | 218.0265 | 273.6903 | 211.3453 | 184.1151 | 189.8762 | 209.4336 | 211.0178 | 237.4779 | 227.8319 | 280.407 | 339.1238 | 211.646 | 191.9205 | 284.4866 | 264.4513 | 298.3805 | 295.8497 | 279.0178 | 447.6195 | 211.3805 | 195.8054 | 184.4513 | 294.7255 | 165.1681 | 161.9114 | 169.4248 | 193.354 | 239.0973 | 256.6903 | 275.1416 | 191.3097 | 121.1238 | 131.1151 | 122.2567 | 141.7876 | 311.6903 | 285.5842 | 252.8849 | 422.4248 | 264.7433 | 274.5309 | 270.699 | 367.8319 | 183.0973 | 82.1681 | 80.5044 | 183.0973 | 155.9557 | 147.4779 | 153.3362 | 167.4691 | 126.407 | 144.5399 | 140.7611 | 129.8141 | 105.5664 | 77.6903 | 74.4779 | 111.593 | 138.8762 | 122.407 | 120.1238 | 102.6903 | 343.5487 | 305.1151 | 323.5664 | 462.3805 | 201.5664 | 243.8762 | 214.5131 | 386.7346 | 329.3893 | 232.9644 | 231.9557 | 399.8228
1604 | 19.6384 | 200.3539 | 267.6258 | 276.4669 | 241.7282 | 214.3539 | 216.5949 | 230.4925 | 204.5128 | 208.6258 | 169.5181 | 154.7899 | 239.5128 | 313.3433 | 196.4872 | 198.8614 | 317.3027 | 276.8614 | 196.0256 | 172.5128 | 281.9488 | 191.5746 | 433.0565 | 408.5128 | 223.9744 | 323.8102 | 283.3336 | 297.759 | 384.0256 | 181.9179 | 144.8976 | 134.5693 | 198.7026 | 193.8102 | 167.5949 | 164.1024 | 115.1845 | 321.7537 | 315.5949 | 282.5693 | 357.2463 | 173.2207 | 174.9179 | 172.6514 | 166.677 | 93.4616 | 73.1642 | 69.1589 | 116.6514 | 246.0309 | 255.6258 | 283.5437 | 200.9947 | 159.2718 | 163.9744 | 148.5128 | 151.1077 | 217.8614 | 202.759 | 271.8923 | 216.3539 | 206.8614 | 187.1642 | 210.4307 | 202.436 | 236.4051 | 224.3795 | 266.4616 | 312.3283 | 205.5949 | 196.9744 | 294.9947 | 266.7026 | 250.5234 | 290.1386 | 261.3027 | 465.2665 | 216.2154 | 194.9744 | 177.5693 | 287.1333 | 173.5384 | 178.6973 | 164.4616 | 202.2154 | 237.4051 | 248.5693 | 278.4051 | 175.5437 | 118.9179 | 118.7335 | 112.436 | 142.5128 | 297.9488 | 273.1951 | 260.3486 | 392.872 | 264.0256 | 249.2463 | 279.0768 | 384.0512 | 189.8358 | 77.1898 | 78.9179 | 189.8358 | 141.8667 | 149.9179 | 149.7026 | 156.9232 | 137.7537 | 137.7335 | 142 | 134.241 | 87.7079 | 76.5949 | 83.241 | 104.4616 | 134.0565 | 121.8102 | 118.6205 | 98.0821 | 343.4051 | 327.1589 | 331 | 459.4872 | 190.0309 | 239.0565 | 220.9947 | 387.7846 | 347.2101 | 221.323 | 230.5128 | 429.1792
1605 | 19.6494 | 212.8713 | 278.301 | 264.5743 | 239.2812 | 193.9881 | 196.699 | 218.7149 | 208.1446 | 220.0158 | 159.2851 | 157.7267 | 228.1327 | 302.8357 | 186.4218 | 190.9881 | 317.4376 | 256.1208 | 190.4218 | 171.1406 | 265.2772 | 204.1643 | 420.1366 | 430.7307 | 214.5663 | 314.9921 | 286.1604 | 281.5624 | 367.699 | 156.1208 | 161.594 | 127.4257 | 189.7069 | 190.7109 | 166.2851 | 171.7109 | 118.5663 | 338.7228 | 317.8594 | 278.2851 | 342.4218 | 155.4218 | 168.4218 | 159.1366 | 162.5703 | 88.2851 | 77.5782 | 67.4218 | 131.7307 | 261.5901 | 244.1406 | 310.598 | 188.4099 | 166.7228 | 158.1406 | 162.8673 | 162.7228 | 200.6951 | 209.4376 | 265.996 | 213.1446 | 205.4218 | 176.5663 | 200.4178 | 187.1366 | 244.8634 | 217.7109 | 274.1525 | 317.5822 | 179.9802 | 172.5545 | 280.9802 | 235.5466 | 275.7504 | 280.1406 | 289.3168 | 491.5861 | 205.1327 | 172.699 | 168.996 | 280.9921 | 178 | 188.5703 | 176.4416 | 200.4257 | 267.3089 | 236.4218 | 291.8673 | 171.8594 | 136.7267 | 114.5743 | 120.0119 | 136.8515 | 286.2812 | 240.9842 | 271.7188 | 398.7307 | 269.1446 | 246.5743 | 304.4416 | 377.8475 | 186.8515 | 82.4337 | 78.8554 | 186.8515 | 140.1446 | 154.8594 | 142.8515 | 168.7307 | 138.996 | 157.1643 | 122.699 | 146.8634 | 87.7228 | 87.4376 | 90.8594 | 116.4416 | 126.1406 | 126.5743 | 137.4416 | 85.2772 | 338.996 | 349.0119 | 295.9723 | 451.5663 | 172.7069 | 226.8515 | 239.8634 | 389.5743 | 354.996 | 221.1406 | 242.7228 | 449.2851
1606 | 19.6604 | 227.2714 | 275.82 | 264.5457 | 233.3643 | 185.8186 | 219.5399 | 231.3614 | 216.2714 | 220.0014 | 179.36 | 166.9086 | 233.18 | 315.2685 | 201.9057 | 208.2685 | 343.0871 | 241.3643 | 197.9071 | 196.1771 | 260.2729 | 205.5472 | 444.6314 | 438.2729 | 214.1814 | 297.9115 | 307.3614 | 300.9957 | 378.9057 | 157.3614 | 158.4572 | 136.9071 | 169.2757 | 179.5472 | 174.7257 | 172.9086 | 104.3657 | 368.45 | 317.5457 | 308.5399 | 345.5443 | 170.9057 | 159.1829 | 164.7257 | 145.2757 | 85.0914 | 87.5443 | 82.1785 | 117.4586 | 252.0943 | 236.4557 | 297.8229 | 191.3614 | 162.2743 | 164.2714 | 155.6386 | 152.4572 | 217.1771 | 192.1858 | 269.09 | 192.9129 | 203.4543 | 187.0885 | 233.5385 | 175.2743 | 229.8215 | 213.8186 | 279 | 316.4557 | 188.1785 | 159.6372 | 284.0885 | 268.0828 | 275.0943 | 276.8186 | 255.5544 | 461.9158 | 203.6357 | 168.6357 | 193.9042 | 276.2729 | 172.1829 | 178.5472 | 193.18 | 198.2729 | 271.7286 | 223.5472 | 257.0072 | 178.09 | 132.82 | 105.82 | 108.5486 | 141.2714 | 288.3628 | 243.18 | 276.9086 | 412.09 | 264.91 | 265.4514 | 288.4586 | 373 | 182.5457 | 77.7286 | 75.0914 | 182.5457 | 162.0871 | 151.6372 | 163.2685 | 155.9129 | 126.0929 | 155.6386 | 128.0885 | 152.1814 | 89.8186 | 84.0014 | 91.2729 | 117.91 | 125 | 143.9971 | 122.9129 | 84.6357 | 339.1814 | 320.8243 | 295.8157 | 453.3628 | 163.1829 | 249.45 | 235.7286 | 379.3657 | 334.0943 | 247.6314 | 233.1843 | 435.1843
1607 | 19.6714 | 220.3874 | 246.1199 | 285.4435 | 242.3521 | 186.4789 | 200.6833 | 210.8593 | 204.9508 | 233.2675 | 187 | 161.3451 | 211.1199 | 328.4366 | 205.3028 | 204.4296 | 341.3874 | 258.4858 | 187.2113 | 188.2536 | 261.4789 | 226.4013 | 474.2252 | 392.1552 | 222.3943 | 307.0492 | 288.3382 | 304.824 | 394.8732 | 140.5565 | 144.8662 | 149.1338 | 205.8871 | 187.8309 | 180.9577 | 171.0423 | 95.5634 | 353.5987 | 320.6972 | 325.9155 | 360.5704 | 169.1268 | 157 | 165.7817 | 157.4858 | 87.6972 | 95.176 | 89.7394 | 135.6196 | 253.1338 | 253.9647 | 285.0423 | 207.8309 | 156.0423 | 169.2183 | 142.8662 | 155.8732 | 229.2183 | 195.0915 | 262.1268 | 218.2745 | 180.3382 | 202.3521 | 230.2536 | 195.6618 | 242.2252 | 218.176 | 271.6057 | 311.7817 | 188.3451 | 175.0069 | 277.9085 | 283.5211 | 270.2183 | 287.0915 | 261.4013 | 465.0069 | 187.9931 | 187.7464 | 196.6057 | 292.2675 | 181.8309 | 186.8309 | 188.8662 | 182.4719 | 251.2536 | 250.3167 | 238.3028 | 168.1691 | 123.0845 | 122.7041 | 113.8309 | 138.0845 | 298.8732 | 273.6196 | 252.8593 | 396.9931 | 300.7111 | 273.4789 | 269.6479 | 375.9577 | 181.2606 | 73.5211 | 68.0845 | 181.2606 | 158.9085 | 154.4366 | 159.4296 | 136.1691 | 135.3098 | 142.8662 | 148.5281 | 154.4789 | 91.9577 | 83.2183 | 72.5142 | 102.6902 | 138.3098 | 119.6833 | 115.7394 | 88.9577 | 351.0915 | 293.2113 | 298.0423 | 457.2183 | 183.183 | 260.4789 | 231.5211 | 369.824 | 315.9085 | 254.3028 | 239.8309 | 435.176
1608 | 19.6824 | 221.5103 | 279.0965 | 295.2552 | 246 | 203.5379 | 222.3241 | 205.7586 | 224.0552 | 229.2276 | 158.4345 | 158.2483 | 230.5655 | 308.9517 | 215.2759 | 189.4759 | 360.5517 | 250.7173 | 194.7793 | 197.0345 | 249.9724 | 202.6759 | 456.9448 | 409.8276 | 234.0207 | 304.4828 | 315.331 | 290.9724 | 390.4828 | 162.3172 | 144.2552 | 143.731 | 171.3862 | 189.7448 | 179.7448 | 148.2 | 111.2896 | 327.7104 | 300.2 | 303.9448 | 355.2276 | 173.5172 | 165.269 | 159.7379 | 159.4897 | 92.0069 | 86.4759 | 86.2414 | 129.2138 | 256 | 242.2069 | 287.7586 | 182.6828 | 171.5379 | 163.9862 | 146.0138 | 153.7379 | 226.2414 | 195.9931 | 243.2138 | 188.6552 | 200.5655 | 189.4621 | 233.0207 | 179.1931 | 231.7104 | 200.4552 | 246.4483 | 324.531 | 171.7173 | 176.4897 | 274.7517 | 277.7379 | 285.2827 | 259.4276 | 276.0138 | 460.4759 | 187.2621 | 186.4828 | 184.9793 | 284.469 | 180.7379 | 188.7448 | 171.7173 | 173.2414 | 214.1793 | 243.7104 | 256.5448 | 179.7862 | 111.2276 | 140.5241 | 136.0414 | 143.5172 | 274.9379 | 273.9793 | 269.5586 | 409.0414 | 292.2 | 274 | 274.269 | 395.7931 | 186.2621 | 70.7448 | 79.531 | 186.2621 | 147.4828 | 131.9448 | 151.9931 | 142.5241 | 144.5103 | 136.9931 | 139.7173 | 137.7104 | 105.0276 | 92.269 | 76.5241 | 116.0414 | 138.4897 | 117.269 | 125.0207 | 86.2414 | 361.7655 | 275.9724 | 286.4759 | 454.2414 | 184.9862 | 218.1518 | 230.2483 | 381.531 | 336.0552 | 243.9793 | 234.2276 | 400.1655
1609 | 19.6934 | 205.4277 | 289.9439 | 287.5958 | 236.8318 | 223.5163 | 207.7875 | 202.4159 | 230.472 | 208.4277 | 177.2685 | 151.8879 | 229.3038 | 302 | 212.8879 | 208.1564 | 322.1591 | 258.9883 | 211.2242 | 187.4837 | 263.5723 | 202.6962 | 448.236 | 448.5045 | 219.4277 | 317.2803 | 296.4394 | 292.3481 | 374.2477 | 155.9557 | 138.1239 | 146.3481 | 173.3364 | 195.1121 | 191.2242 | 152.4602 | 117 | 303.6636 | 299.8761 | 294.472 | 357.3481 | 164.5398 | 158.8318 | 139.6636 | 160.292 | 92.236 | 85.292 | 85 | 124 | 261.3481 | 219.1917 | 274.4837 | 188.2803 | 178.528 | 171.9322 | 158.6962 | 164.9883 | 237.2242 | 198.8201 | 256.3364 | 193.3364 | 212.292 | 185.528 | 221.4837 | 177.1121 | 243.5723 | 199.3481 | 242.8201 | 355.7405 | 196.0326 | 190.2803 | 246.7315 | 247.7315 | 274.7197 | 265.8083 | 274.9439 | 466.1682 | 199.6962 | 203.8644 | 185.8201 | 283.8201 | 178.236 | 161.2595 | 178.4602 | 187.2803 | 254.425 | 225.7758 | 235.4955 | 163.6076 | 117.1682 | 128.7197 | 122.8996 | 149.8201 | 266 | 252.6636 | 253.5515 | 419.5841 | 269.7197 | 263.3038 | 255.6076 | 380.6076 | 181.8879 | 79.1682 | 88.5841 | 181.8879 | 160.2803 | 152.2685 | 162.4602 | 145.236 | 133.7758 | 136.764 | 139.5841 | 131.236 | 78.4394 | 91.9439 | 77.708 | 111.3038 | 130.1239 | 118.472 | 114.2477 | 87.292 | 362.472 | 289.5723 | 285.292 | 469.0443 | 171.5398 | 231.5045 | 231.528 | 398.2242 | 325.6636 | 236.4159 | 260.7966 | 388
1610 | 19.7044 | 193.0132 | 245.5269 | 273.3554 | 233.2237 | 203.9345 | 193.0132 | 208.7631 | 220.6843 | 214.6446 | 178.2369 | 167.855 | 219.7896 | 318.3024 | 204.0132 | 215 | 320.4209 | 272.3157 | 203.3554 | 171.3554 | 274.4341 | 192.8028 | 399.869 | 426.7241 | 221.7631 | 305.6976 | 306.4076 | 301.7631 | 384.7498 | 151 | 153.855 | 144.8948 | 172.7896 | 200.1052 | 201.2104 | 142.8028 | 110.7896 | 299.5526 | 289.5791 | 287.0132 | 352.0132 | 180.4076 | 148.2369 | 153.4076 | 168.7631 | 85.0132 | 87.5526 | 74.908 | 126.3289 | 249.0265 | 237.2892 | 260.6843 | 196.8815 | 156.4871 | 154.0397 | 161.2237 | 165.1185 | 223.145 | 207.7631 | 260.4474 | 202.1052 | 189.7108 | 184.4474 | 213.8948 | 197.6313 | 225.7108 | 206.4341 | 263.4076 | 318.9743 | 191.8028 | 200.4341 | 283.0257 | 245.9868 | 276.9868 | 232.1847 | 260.0265 | 472.1052 | 192.1317 | 175.0662 | 182.3422 | 272.5791 | 178 | 170.0787 | 185.1052 | 192.7763 | 268.4474 | 253.0522 | 252.6181 | 168.6446 | 106.0265 | 124.7763 | 134.0787 | 151.7763 | 276.092 | 259.4209 | 257.6446 | 424.8815 | 276.6446 | 267.7631 | 280.8285 | 378.6578 | 178.4474 | 78.8948 | 82.2369 | 178.4474 | 159.5659 | 154.0132 | 172.2104 | 141.8948 | 133.8815 | 131.5659 | 137.1185 | 153.5129 | 81.4209 | 93.3289 | 70.0132 | 129.7366 | 124.1185 | 126.5394 | 112.3289 | 95.7631 | 362.7763 | 292.6711 | 305.4076 | 456.145 | 173.4341 | 215.9345 | 245.9735 | 399.6711 | 342.5129 | 255.9603 | 228.0795 | 395.7631
1611 | 19.7154 | 202.829 | 260.6009 | 270.7886 | 243.2518 | 209.6175 | 206.7719 | 233.8693 | 236.9263 | 201.4395 | 177.5772 | 176.1544 | 220.3658 | 289.8789 | 194.114 | 213.4228 | 345.2921 | 232.4158 | 188.9596 | 163.2684 | 285.4632 | 198.4632 | 449.8763 | 414.8456 | 229.5202 | 292.3254 | 298.5939 | 289.8053 | 400.0404 | 148.6342 | 145.5939 | 146.3658 | 172.5772 | 172.6105 | 177.7649 | 161.0807 | 86.1307 | 319.7149 | 312.0237 | 293.6746 | 367.3491 | 161.5535 | 144.4228 | 158.2114 | 181.2518 | 76.6912 | 80.9026 | 60.9596 | 126.2114 | 235.5368 | 232.171 | 284.0237 | 196.4228 | 137.3825 | 162.1947 | 164.1544 | 160.8456 | 226.6746 | 216.3088 | 243.4395 | 218.7719 | 182.2114 | 161.1307 | 208.2684 | 195.9026 | 223.7316 | 217.4632 | 265.8456 | 374.6079 | 198.2518 | 186.2281 | 278.6509 | 263.7719 | 266.3825 | 238.3491 | 241.8053 | 495.8693 | 178.7482 | 178.4061 | 201.5035 | 269.7886 | 156.7079 | 178.943 | 156.0333 | 185.114 | 262.4798 | 218.6272 | 226.8789 | 155.4395 | 139.8526 | 117.9026 | 121.6509 | 152.7886 | 278.2114 | 268.5202 | 272.0404 | 418.114 | 288.6746 | 265.2684 | 282.114 | 360.2851 | 163.0167 | 74.057 | 82.3658 | 163.0167 | 165.0974 | 160.6746 | 164.5368 | 147.3088 | 129.4798 | 119.7482 | 138.3658 | 141.0737 | 75.5368 | 94.7886 | 82.9833 | 121.0167 | 151.3895 | 127.4228 | 120.0974 | 95.6342 | 361.4228 | 291.2114 | 300.7482 | 465.9833 | 159.2281 | 233.4465 | 242.9026 | 397.4228 | 340.3254 | 251.7482 | 218.3658 | 424.0237
1612 | 19.7264 | 214.0089 | 273.6062 | 259.7876 | 249.2036 | 208.9956 | 215.8053 | 239.1991 | 218.7743 | 202.6062 | 180.4027 | 172.9956 | 220.1991 | 306.6284 | 203.2124 | 210.5973 | 318.1636 | 249.832 | 210.0266 | 186.8275 | 278.3893 | 194.5929 | 450.9822 | 403.5885 | 232.6018 | 306.8186 | 302.208 | 290.8053 | 390.9867 | 158.4115 | 138.7964 | 144.5973 | 183.4115 | 181.8186 | 170.1991 | 161.3938 | 101.624 | 328.2036 | 318.1991 | 302.4071 | 354.3805 | 159.8053 | 139.1947 | 149.1902 | 155.9689 | 85.4115 | 72.5929 | 65.208 | 134.8098 | 233.8009 | 228.1991 | 290.1991 | 196.8009 | 147.6151 | 176.4115 | 144.1769 | 146.3849 | 228.1991 | 201.9822 | 206.1636 | 209.3849 | 182 | 192.6417 | 202.9956 | 194.8009 | 245.0222 | 231.2124 | 244.1769 | 338.5396 | 217.0178 | 194.0133 | 262.7876 | 233.5618 | 262.1991 | 234.9911 | 260.4249 | 465.96 | 181.6062 | 190.8098 | 186.9778 | 258.7876 | 160.6107 | 155.1725 | 172.8275 | 190.208 | 232.9689 | 224.6195 | 234.0178 | 167.0178 | 121.168 | 136.0222 | 147.4337 | 130.5751 | 291.6151 | 241.168 | 227.7476 | 395.9778 | 285.3938 | 272.8098 | 255.1725 | 407.0577 | 182.2257 | 77.0044 | 90.208 | 182.2257 | 157.3893 | 160.5973 | 135.5707 | 157.0089 | 142.416 | 128.2124 | 125.3849 | 140.8053 | 89.0178 | 91.7964 | 69.3805 | 102.584 | 146.1858 | 114.9867 | 117.9956 | 94.1991 | 341.7787 | 279.7876 | 316.4204 | 495.6284 | 155.8009 | 246.4071 | 225.7831 | 384.1858 | 330.792 | 245.7964 | 210.1902 | 409.376
1613 | 19.7374 | 235.5163 | 284.7581 | 291.9667 | 276.8349 | 192.5496 | 225.945 | 202.4069 | 227.6372 | 196.6814 | 163.1101 | 185.0109 | 221.6264 | 313 | 193.8023 | 212.4395 | 315.6923 | 267.5713 | 199.7364 | 196.2527 | 272.7473 | 193 | 399.0224 | 412.3845 | 213.4837 | 305.3077 | 308.0659 | 299.3186 | 424.5931 | 145.5496 | 151.0109 | 122.8573 | 180.3077 | 165.6705 | 174.0659 | 144.5496 | 104.5605 | 320.8682 | 291.9783 | 290.9891 | 344.3077 | 179.7031 | 155.8899 | 165.7031 | 170.9558 | 94.5054 | 71.8132 | 71.8791 | 139.4395 | 235.6264 | 235.3186 | 274.5496 | 168.5387 | 152.6264 | 152.9783 | 155.2636 | 151.945 | 223.9341 | 219.9558 | 228.0876 | 211.6923 | 181.1868 | 185.7364 | 200.3736 | 189.3077 | 240.2419 | 236.4395 | 274.7799 | 363.2194 | 194.1651 | 192.9341 | 276.2636 | 236.3845 | 279.0768 | 272.0326 | 259.4946 | 460.2527 | 202.5163 | 190.5605 | 200.7031 | 291.7799 | 178.4504 | 152.2527 | 172.4946 | 189.5605 | 226 | 251.769 | 235.5605 | 169.3736 | 118.8791 | 132.8682 | 128.9783 | 134.7581 | 279.5496 | 236.4395 | 243.6481 | 396.6923 | 275.055 | 288.0109 | 261.1977 | 372.8356 | 170.9232 | 88.5713 | 84.6814 | 170.9232 | 141.176 | 173.0109 | 150.1427 | 163.0659 | 126.4837 | 120.4287 | 128.5054 | 139.5605 | 79.176 | 89.3736 | 76.3845 | 113.6372 | 154.3845 | 129.8899 | 110.4946 | 95.6264 | 345.945 | 281.0659 | 281.9674 | 483.2969 | 184.4613 | 231.7364 | 219.5605 | 376.1209 | 335.5054 | 231.9891 | 238.9008 | 433.2744
1614 | 19.7484 | 218.5378 | 268.8397 | 279.3633 | 249.1557 | 227.7971 | 235.4292 | 195.6509 | 214.4906 | 177.6651 | 163.9528 | 164.0614 | 239.3349 | 297.3161 | 204.2075 | 192.3633 | 321.1273 | 243.585 | 193.5236 | 197 | 262.0944 | 185.5708 | 424.3207 | 418.3019 | 236.2405 | 304.8255 | 280.1086 | 289.4434 | 421.4434 | 160.9858 | 144.0944 | 146.066 | 178.1745 | 179.1603 | 169.2217 | 173.1933 | 114.7311 | 330.5566 | 281.8727 | 335.0518 | 329.7925 | 156.7595 | 160 | 135.3303 | 142.1557 | 101.7783 | 83.5566 | 81.2547 | 134.2217 | 241.7783 | 237.8255 | 283.382 | 170.2547 | 145.5708 | 165.1603 | 147.4434 | 138.3161 | 225.4764 | 209.3161 | 235.8255 | 185.7595 | 169.4434 | 200.1603 | 218.1603 | 173.1415 | 224.7925 | 223.7925 | 254.1086 | 304.3115 | 192.9528 | 194.4764 | 265.1415 | 245.6038 | 252.4576 | 249.6321 | 277.8113 | 443.6651 | 184.7123 | 198.2547 | 157.1227 | 291.7453 | 172.0944 | 148.8727 | 175.9528 | 174.967 | 214.4434 | 217.3774 | 210.2359 | 185.5094 | 113.3962 | 132.6509 | 99.0614 | 142.7783 | 286.7311 | 264.2405 | 242.5708 | 412.8585 | 277.1273 | 257.1557 | 263.1745 | 376.033 | 171.1273 | 86.0472 | 83 | 171.1273 | 142.9528 | 162.7925 | 164.9056 | 149.967 | 134.382 | 137.8113 | 125.8727 | 129.0944 | 81.7783 | 87.3491 | 88.9056 | 121.1273 | 148.7453 | 118.3161 | 106.5236 | 91.8727 | 345.5236 | 300.1603 | 291.1603 | 454.2359 | 200.0802 | 221.3962 | 226.4292 | 387.382 | 335.3491 | 221.5708 | 237.7453 | 421.8397
1615 | 19.7594 | 214 | 290.9706 | 285.8909 | 236.9734 | 185.7343 | 199.3008 | 196.8378 | 227.7552 | 191.593 | 145.7315 | 155.649 | 236.2979 | 323.3216 | 214.5398 | 193.0266 | 332.0531 | 239.6755 | 193 | 157.6252 | 246.5958 | 178.1357 | 455.4573 | 439.944 | 217.7049 | 265.6252 | 277.351 | 305.4307 | 414.8112 | 147.407 | 159.593 | 147.8112 | 172.1357 | 179.649 | 141.1916 | 173.2979 | 110.2979 | 327.9734 | 308.6461 | 330.7581 | 316.1091 | 151 | 155.8112 | 123.8112 | 149.2419 | 82.056 | 86 | 76.2979 | 122.1091 | 239.649 | 244.7021 | 254.1651 | 194.6195 | 159.0797 | 162.2979 | 124.056 | 149.2419 | 213.4336 | 206 | 230.1357 | 215.186 | 177.0531 | 189.7581 | 191.8406 | 175.0266 | 257.186 | 204.2448 | 265.593 | 322.6727 | 198.1888 | 180.7581 | 280.4307 | 236.9469 | 232.5958 | 263.1063 | 240.1119 | 452.5664 | 208.4839 | 187.4336 | 173.8084 | 305.9175 | 174.1888 | 171.4573 | 182.8643 | 192.1063 | 211.1622 | 230.7818 | 247.7258 | 183.1357 | 126.2419 | 117.0825 | 121.6461 | 152.3776 | 293.1888 | 227.2742 | 274.5104 | 409.2979 | 288.0531 | 247.4867 | 262.1622 | 385.7021 | 156.0825 | 75.7846 | 95.5664 | 156.0825 | 134.7846 | 137.3805 | 149.407 | 141.9734 | 120.2448 | 133.6224 | 145.944 | 127.8378 | 82.1622 | 88.6755 | 82.6224 | 101.8937 | 121.0294 | 127.5664 | 125.2685 | 102.7287 | 329.9203 | 281.3805 | 301.7021 | 460.7287 | 176.8672 | 221.6755 | 253.1328 | 413.4573 | 303.1651 | 209.9469 | 232.649 | 417.1622
1616 | 19.7704 | 252.2523 | 285.7994 | 281.1996 | 243.6505 | 195.5512 | 223.4519 | 201.2503 | 255.6515 | 188.1996 | 164.1014 | 181.3516 | 218.849 | 325.5998 | 198.999 | 195.7001 | 344.2006 | 239.1499 | 158.1479 | 169.5512 | 275.4519 | 196.5512 | 409.8469 | 438.0496 | 206.1996 | 269.9007 | 279.7001 | 298.7994 | 394.4488 | 144.8501 | 132.3981 | 150.4002 | 182.0507 | 167.0993 | 163.2017 | 157.5491 | 120.9007 | 332.9504 | 294.4488 | 301.6484 | 345.4519 | 175.6515 | 161.8004 | 139.151 | 163.9007 | 82.2503 | 77.4995 | 65.6494 | 138.7011 | 240.7001 | 228.1489 | 259.9007 | 194.7497 | 158.5998 | 145.6991 | 149.7518 | 143.4995 | 205.8997 | 198.3495 | 259.6019 | 199.8986 | 179.8501 | 193.8004 | 217.4519 | 171.7497 | 224.0476 | 218.001 | 299.6019 | 312.849 | 211.7508 | 171.1996 | 289.9504 | 272.4023 | 235.9504 | 249.1489 | 236.2503 | 457.5502 | 190.1985 | 162.0486 | 178.1499 | 254.5967 | 177.5502 | 181.1003 | 151.698 | 177.2989 | 222.9007 | 263.0517 | 244.0993 | 171.7994 | 123.0496 | 105.4995 | 121.8997 | 142.0993 | 280.3992 | 252.152 | 273.3495 | 426.7011 | 275.5491 | 259.7508 | 299.4023 | 345.3474 | 185.302 | 76.5502 | 77.5987 | 185.302 | 155.1014 | 145.7508 | 142.5998 | 163.1014 | 143.3516 | 126.0496 | 136.3992 | 155.2017 | 79.4498 | 88.1499 | 76.7497 | 117.5512 | 121.1003 | 126.5998 | 121.3495 | 99.8997 | 321.0496 | 295.7011 | 273.2482 | 455.3495 | 199.2017 | 204.1489 | 261.4002 | 406.0993 | 310.6008 | 235.2017 | 229.4498 | 436.5512
1617 | 19.7814 | 215.8828 | 291.7611 | 265.3402 | 255.3481 | 210.2105 | 221.2389 | 203.7247 | 235.8544 | 197.3481 | 178.3481 | 148.0569 | 209.9636 | 310.3402 | 183.0649 | 213.2469 | 324.4414 | 215.7167 | 157.1741 | 173.8623 | 237.8828 | 222.4209 | 398.413 | 394.7452 | 207.587 | 278.0364 | 289.4858 | 296.1377 | 426.3561 | 150.1741 | 127 | 130.3038 | 182.2753 | 190.008 | 168.8623 | 154.1377 | 115.2389 | 328.8259 | 297.8987 | 288.3766 | 363.0728 | 170.5142 | 145.7531 | 149.7611 | 132.3686 | 76.1013 | 72.5506 | 73.4858 | 130.7895 | 216.8544 | 231.0364 | 243.0285 | 208.6598 | 167.4858 | 153.3481 | 140.3402 | 145.4494 | 199.8259 | 210.7975 | 289.1456 | 198.587 | 180.8623 | 172.5791 | 221.2753 | 172.7247 | 215.2753 | 203.7531 | 267.0569 | 318.6234 | 177.7816 | 170.8623 | 303.9351 | 245.3686 | 268.9067 | 239.1013 | 235.2753 | 444.2025 | 170.4778 | 175.2469 | 173.6883 | 247.587 | 166.7895 | 179.413 | 171.008 | 199.8703 | 231.0364 | 236.0933 | 248.8987 | 165.6883 | 134.0728 | 121.2469 | 117.5506 | 139.1377 | 261.6155 | 245.0649 | 266.8259 | 394.6439 | 262.6519 | 266.3117 | 270.6439 | 375.0808 | 168.5791 | 77.8623 | 67.1013 | 168.5791 | 159 | 146.2753 | 144.587 | 151.4778 | 142.8259 | 131.8987 | 108.992 | 144.4778 | 75.5506 | 81.9636 | 75.1377 | 113.2389 | 136.6598 | 135.4858 | 123.4494 | 89.5142 | 321.7247 | 294.6883 | 294.7326 | 460.8987 | 184.1661 | 219.9715 | 259.413 | 394.5142 | 297.4778 | 230.5142 | 254.008 | 428.7895
1618 | 19.7924 | 243.1107 | 312.242 | 266.4985 | 232.5102 | 210.2507 | 220.8746 | 197.0029 | 218.0059 | 202.4985 | 170.379 | 140.2507 | 215.1224 | 287.8834 | 193.2449 | 195.8834 | 319.2507 | 229.4927 | 165.8717 | 206.3615 | 256.3644 | 221.6268 | 427.7376 | 378.379 | 229.8659 | 286.8717 | 276.1312 | 316.9912 | 403.1371 | 153.6239 | 159.3615 | 123.5015 | 155.761 | 189.6268 | 162.8776 | 164.4956 | 106.1283 | 307.8834 | 291.1283 | 314.1137 | 332.6385 | 154.1312 | 169.239 | 140.5044 | 147.9912 | 102.1137 | 75.4985 | 71.5015 | 122.8776 | 226.9941 | 236.3732 | 253.9941 | 190.0088 | 177.7463 | 160.2478 | 122.2566 | 157.3702 | 218.242 | 196.382 | 260.6385 | 202.4985 | 164.382 | 166.3761 | 180.7668 | 178.2478 | 234.242 | 213.2449 | 271.4956 | 306.0059 | 186.8688 | 173.6239 | 242.1517 | 259.242 | 232.8922 | 213.5102 | 237.6239 | 482.2332 | 178.4956 | 176.2507 | 171.2507 | 265.4927 | 171.1224 | 192.1195 | 168.8776 | 193.5044 | 238.9971 | 235.3732 | 212.3907 | 174.621 | 105.3878 | 114.379 | 140.6151 | 153.8688 | 272.9941 | 240.3761 | 258.1283 | 390.7493 | 276.7434 | 248.6327 | 259.7522 | 356.5102 | 184.242 | 78.8746 | 75.621 | 184.242 | 150.2537 | 161.7434 | 146.7493 | 148.1254 | 146.3732 | 141.7463 | 113.7463 | 128.8805 | 90.7434 | 79.2507 | 75 | 121.621 | 128.5044 | 116.0088 | 114.379 | 101.1195 | 317.6268 | 311.4927 | 304.2478 | 485.6151 | 182.7493 | 201.1341 | 187.28 | 426.2361 | 293.2507 | 224.6268 | 256.2507 | 440.1195
1619 | 19.8034 | 213.4099 | 252.0283 | 276.7562 | 276.007 | 188.7138 | 205.9223 | 200.4346 | 215.1131 | 199.4523 | 171.6608 | 176.3639 | 253.2509 | 281.4523 | 184.3569 | 202.7562 | 314.5654 | 227.5654 | 178.53 | 188.8269 | 232.5053 | 204.1484 | 416.9223 | 403.6078 | 213.4876 | 276.47 | 269.5654 | 280.0884 | 393.6784 | 139.8092 | 157.7915 | 132.7562 | 159.9823 | 198.7562 | 168.2085 | 168.6608 | 105.8869 | 299.6784 | 285.5654 | 300.2615 | 339.53 | 155.5477 | 168.5654 | 149.6431 | 131.4876 | 98.9046 | 75.1131 | 87.8516 | 148.6078 | 229.8869 | 243.2085 | 280.8339 | 172.8092 | 152.3922 | 138.8269 | 127.0954 | 144.8092 | 224.5477 | 209.0777 | 265.7562 | 210.0954 | 167.3216 | 182.8516 | 188.3039 | 168.3569 | 215.7138 | 212.3392 | 250.8269 | 294.2438 | 200.53 | 194.3993 | 256.947 | 265.5477 | 275.7809 | 243.7032 | 238.8869 | 458.7315 | 170.2438 | 195.5124 | 184.3039 | 296.3816 | 157.8092 | 171.8269 | 170.6608 | 184.0177 | 226.6961 | 215.6007 | 238.9293 | 152.9399 | 110.7562 | 104.1307 | 120.053 | 139.1484 | 269.6784 | 243.5477 | 264.9823 | 397.2085 | 279 | 285.9116 | 250.1307 | 389.3639 | 193.2085 | 77.2261 | 70.7915 | 193.2085 | 150.7739 | 155.1307 | 187.7986 | 151.5477 | 122.1661 | 133.2438 | 131.8516 | 145.6254 | 82.3569 | 103.8339 | 91.8516 | 129.2085 | 150.0601 | 97.9223 | 97.9223 | 99.4523 | 345.3816 | 279.4099 | 293.47 | 490.7739 | 164.3746 | 225.4947 | 224.007 | 395.523 | 337.3463 | 232.8693 | 230.2792 | 434.0177
1620 | 19.8144 | 219.7906 | 282.6663 | 272.6047 | 265.8141 | 185.1008 | 200.4031 | 210.8914 | 206.9071 | 185.5118 | 180.0929 | 151.3259 | 226.5274 | 293.589 | 221.6663 | 178.8219 | 330.1859 | 235.9922 | 173.7055 | 201.2867 | 255.0773 | 190.3102 | 390.7212 | 423.1859 | 216.3953 | 273.2016 | 278.8914 | 285.7906 | 353.4345 | 138.8992 | 141.7133 | 171.7671 | 178.9843 | 200 | 180.6898 | 160.0078 | 117.6898 | 324.1781 | 325.4647 | 273.7212 | 353.589 | 149.7055 | 169.7984 | 142.0078 | 158.6741 | 90.8063 | 75.8992 | 80.1086 | 123.2251 | 235.3953 | 219.7212 | 268.7133 | 172.7984 | 167.8835 | 164.7749 | 125.3024 | 159.1859 | 210.6125 | 180.4266 | 247.2172 | 192.1165 | 176.9922 | 176.0078 | 207.9843 | 153.5118 | 239.0773 | 236.2788 | 260.589 | 307.3875 | 167.8298 | 175.4188 | 263.5969 | 238.1243 | 278.4031 | 248.8992 | 271.3718 | 477.4804 | 178.8914 | 177.318 | 186.8992 | 287.411 | 156.8992 | 152.8141 | 171 | 164.1165 | 212.411 | 208.5039 | 208.8298 | 159.8914 | 108.4031 | 101.2016 | 120.5969 | 140.5969 | 277.0929 | 246.6976 | 243.5196 | 395.3024 | 280.7984 | 255.0313 | 249 | 381.411 | 151.7368 | 80.5969 | 70.8992 | 151.7368 | 135.7133 | 153.1008 | 138.1478 | 151.1008 | 126.1937 | 110.4188 | 132.2016 | 160.589 | 85.4961 | 95.3102 | 82.3102 | 111.1165 | 104.4423 | 92.4031 | 96.8992 | 97.2016 | 340.9071 | 291.1859 | 289.3024 | 433.4502 | 165.5969 | 204.7212 | 213.8141 | 381.1086 | 295.3415 | 206.1243 | 211.7133 | 455.4804
1621 | 19.8254 | 227.3806 | 273.3274 | 242.8318 | 248.5044 | 209.6107 | 180.8583 | 210.177 | 231.5222 | 190.3806 | 170.9734 | 160.7611 | 246.6992 | 299.5575 | 211.4159 | 203.3452 | 312.8583 | 240.646 | 177.5575 | 172.0087 | 249.7964 | 195.3806 | 409.8762 | 441.4071 | 195.1238 | 292.1417 | 249.0087 | 287.9115 | 387.2833 | 140.823 | 137.2655 | 132.2476 | 183.7345 | 186.3274 | 182.9115 | 169.0266 | 109.8849 | 314.2389 | 296.2742 | 299.2567 | 368.6726 | 147.177 | 160.8849 | 168.3452 | 133.7433 | 68.1238 | 87.8496 | 80.823 | 135.4956 | 211.3893 | 246.1682 | 278.8496 | 196.6992 | 157.2389 | 145.2123 | 149.6107 | 150.0619 | 198.9734 | 205.2567 | 231.3274 | 192.7345 | 183.469 | 166.7964 | 195.4159 | 187.5488 | 225.5929 | 223.5044 | 266.5575 | 295.3274 | 194.9913 | 188.4956 | 275.8496 | 277.8408 | 269.7964 | 245.354 | 295.9647 | 468.1504 | 160.8583 | 175 | 175.1504 | 263.2123 | 176.1417 | 180.1682 | 147.3008 | 182.0532 | 225.5841 | 229.8762 | 216.8496 | 163.7345 | 106.177 | 121.9647 | 106.4159 | 152.8496 | 266.1504 | 229.6814 | 257.4071 | 361.2742 | 265.5044 | 278.3452 | 277.2567 | 347.1857 | 168.8762 | 85.5575 | 72.823 | 168.8762 | 160.4337 | 135.6814 | 170.2833 | 151.9115 | 151.6107 | 139.9028 | 108.3008 | 171.1151 | 67.7699 | 85.7964 | 101.9647 | 115.3806 | 137.2833 | 98.3806 | 113.4071 | 81.5044 | 328.1504 | 288.4425 | 297.2036 | 496.2746 | 176.0266 | 212.9381 | 221.1151 | 349.0087 | 353.8055 | 207.5575 | 227.3186 | 404.2211
1622 | 19.8364 | 235.3904 | 272 | 260.3236 | 210.0481 | 188.905 | 198.3998 | 217.3904 | 210.905 | 201.1618 | 194.9426 | 149.9906 | 207.4291 | 309.238 | 230.3236 | 207.8476 | 324.857 | 241.9238 | 165.9906 | 183.7808 | 245.3048 | 204.3142 | 447.1043 | 431.9144 | 217.0188 | 278.2954 | 261.7046 | 286.1524 | 372.524 | 160.3998 | 146.238 | 170.4947 | 178.4572 | 181.3048 | 162.6764 | 154.2954 | 121.9332 | 330.5522 | 261.5909 | 295.5334 | 364.4572 | 158.0856 | 158.1524 | 149.7526 | 161.4853 | 90.0188 | 85.3048 | 82.8476 | 119.4478 | 227.476 | 243.4572 | 249.5147 | 184.2192 | 156 | 137.4572 | 164.9332 | 149 | 181.3716 | 189.524 | 228.1524 | 204.0856 | 172.9144 | 167.8476 | 181.9906 | 152.2005 | 236.0094 | 202.6002 | 265.1524 | 314.3236 | 184.143 | 172.4478 | 264.9906 | 251.5147 | 258.8382 | 257.0094 | 240.7245 | 469.7714 | 174.7046 | 175 | 164.762 | 270.238 | 192.7808 | 169.143 | 178.2567 | 209.8664 | 217.762 | 214.4478 | 238.3236 | 151.9906 | 114.3142 | 107.3716 | 106.8476 | 138.2954 | 260.381 | 239.0856 | 238.6764 | 410.6565 | 268.619 | 272.6858 | 246.7433 | 380.0281 | 174.6952 | 81.381 | 81.3142 | 174.6952 | 147.2954 | 131.2286 | 146.286 | 134.4478 | 126.286 | 125.4478 | 147.5709 | 142.4385 | 84.476 | 74.8382 | 79.9812 | 106.762 | 133.6096 | 103.619 | 100.2192 | 105.8664 | 320.5334 | 303.7046 | 291.5334 | 492.8382 | 172.381 | 210.3048 | 201.6764 | 377.4091 | 299.9531 | 213.5428 | 228.0762 | 419.3236
1623 | 19.8474 | 230.3835 | 258.8947 | 256.3835 | 255.6767 | 189.8083 | 211.2331 | 184.3008 | 207.1278 | 207.6165 | 153.0038 | 164.9135 | 216.1692 | 311.8722 | 221.703 | 192.0865 | 285.7482 | 235.4474 | 169.6804 | 196.2331 | 229.0865 | 205 | 405.0677 | 392.6203 | 193.7256 | 278.8722 | 289.2105 | 270.0865 | 359.7669 | 147.0226 | 141.3835 | 152.4699 | 182.6804 | 169.7669 | 151.6391 | 148.3196 | 108.0226 | 311.406 | 291.7631 | 305.297 | 321.876 | 158.0639 | 164.5526 | 143.3196 | 144.3421 | 75.1504 | 85 | 65.2143 | 125.4887 | 233.6804 | 199.9399 | 232.9587 | 178.3196 | 147.5752 | 166.0188 | 156.6391 | 149.9361 | 211.827 | 186.1278 | 248.594 | 211.5526 | 172 | 172.6804 | 215.6353 | 172.4022 | 220.1504 | 204.7444 | 218.1955 | 310.3835 | 188.6165 | 179.4248 | 265.8722 | 261.1692 | 260.8083 | 237.406 | 242.5526 | 455.0226 | 163.8308 | 168.4474 | 177.1053 | 267.2556 | 171.5339 | 169.8722 | 158.5339 | 173.6203 | 223.5526 | 214.8722 | 218.4699 | 171.594 | 113.1278 | 134.0827 | 110.7444 | 119.2143 | 250.6391 | 257.7857 | 237.9361 | 437.4661 | 259.6391 | 230.8121 | 239.3196 | 356.7895 | 180.6165 | 81 | 78.2556 | 180.6165 | 141.3196 | 140.3609 | 146.8083 | 172.3158 | 109.9587 | 120.2556 | 120.109 | 140.9361 | 63.5339 | 88.9774 | 67.703 | 112.5526 | 120.8308 | 127.4022 | 107.4248 | 93.9587 | 320 | 290.9587 | 292.8722 | 462.0451 | 177.6165 | 203.4474 | 241.1879 | 412.7631 | 319.3383 | 241.1466 | 222.3835 | 447.2105
1624 | 19.8584 | 217.6711 | 265.587 | 289.1933 | 229.6002 | 181.4646 | 219.587 | 189.587 | 184.2389 | 215.587 | 170.8644 | 137.5486 | 202.7743 | 307.2581 | 214.3613 | 217.5547 | 286.7935 | 237.8451 | 153.8775 | 202.6903 | 264.0385 | 218.2773 | 438.0385 | 373.8775 | 205.2773 | 312.1933 | 277.7227 | 302.1933 | 398.832 | 171.6063 | 145.7419 | 134.8775 | 148.8583 | 163.3097 | 175.6579 | 170.7611 | 102.2581 | 317.587 | 303.4838 | 291.7743 | 346.503 | 149.4646 | 159.3097 | 163.8644 | 140.1549 | 96.7611 | 93.5354 | 95.2966 | 122.2065 | 227.3613 | 215.0192 | 238.6387 | 164.7227 | 166.916 | 166.1032 | 151.2581 | 134.8259 | 191.2389 | 198.3289 | 240.5162 | 214.8451 | 161.5678 | 170.1549 | 197.1356 | 163.5678 | 217.1032 | 206.8968 | 255.7804 | 273.9615 | 154.8583 | 176.2065 | 269.7935 | 276.2257 | 256.2581 | 247.3806 | 249.6387 | 454 | 196.1933 | 162.3097 | 187.4838 | 263.2065 | 159.5678 | 165.2581 | 160.7935 | 201.3482 | 200.2905 | 217.8451 | 217 | 160.6711 | 115.8451 | 118.9292 | 126.1741 | 110.413 | 262.3289 | 258.0516 | 245.587 | 420.0324 | 259.9484 | 229.8968 | 271.2449 | 328.4453 | 168.6711 | 84.7935 | 79.8968 | 168.6711 | 138.1549 | 142.8968 | 143.2065 | 133.2713 | 113.7419 | 120.9484 | 136.0192 | 123.9292 | 87.6063 | 83.3613 | 83.1225 | 93.084 | 131.3806 | 127.1032 | 108 | 88.2581 | 326.6387 | 302.3289 | 279.7227 | 480.8644 | 151.4453 | 232.3998 | 213.6518 | 378.0132 | 305.8259 | 221.1872 | 212.5162 | 414.8583
1625 | 19.8694 | 217.9607 | 283.2921 | 246.8091 | 263.5448 | 162.7472 | 218.0787 | 217.8595 | 217.5842 | 202.5506 | 153.7472 | 142.7247 | 203.9213 | 277.2192 | 210.1573 | 209.3933 | 311.0168 | 215.9046 | 193.3482 | 179.9439 | 250.6293 | 191.1405 | 405.4158 | 375.882 | 209.8427 | 278.4552 | 272.1966 | 279.0225 | 390.4326 | 144.1799 | 141.1966 | 131.118 | 175.8201 | 152.4326 | 161.6293 | 170.0787 | 110.646 | 322.8034 | 287.6686 | 291 | 336.4719 | 165.3314 | 164.764 | 159.236 | 135.1966 | 88.3933 | 86.3146 | 82.5899 | 121.0393 | 236.6067 | 221.764 | 229.3933 | 161.118 | 152.6293 | 140.0619 | 147.1573 | 147.4494 | 190 | 207.646 | 255.3707 | 210.1966 | 179.2528 | 167.118 | 210.4101 | 157.236 | 225.646 | 207.9607 | 244.5506 | 282.5674 | 173.1741 | 188.4887 | 270.9607 | 242.4158 | 263.6854 | 270.0954 | 262.4887 | 453.0393 | 177.8259 | 158.1573 | 184.1573 | 248.5899 | 150.354 | 148.6686 | 173.4887 | 172.2585 | 224.9381 | 171.8878 | 222.764 | 155.1966 | 125.6067 | 127.6067 | 125.0787 | 131.1348 | 246.6686 | 257.0393 | 239.2753 | 385.3765 | 279.2134 | 248.2528 | 262.4326 | 395.2077 | 169.9213 | 74.4326 | 68.4719 | 169.9213 | 123.5899 | 138.1966 | 131.4719 | 136.764 | 114 | 116.1966 | 117.7866 | 137.4101 | 86.118 | 81.0787 | 84 | 106.4101 | 132.9607 | 138.5281 | 109.9213 | 113.9381 | 320.2753 | 284.7472 | 288.6067 | 454.1405 | 170.1741 | 209.0225 | 243.7022 | 398.0954 | 318.4494 | 223.8427 | 208.1573 | 383.2192
1626 | 19.8804 | 214.1081 | 278.1622 | 264.4593 | 249.4326 | 197.0267 | 191.73 | 208.2974 | 220.9459 | 197.1352 | 192.8915 | 143 | 219.5674 | 303.243 | 189.5678 | 195.3785 | 290.5948 | 241.27 | 173.5948 | 185.8107 | 281.1348 | 172.4867 | 371.8922 | 383.7837 | 202.2163 | 279.9189 | 301.1889 | 292.5944 | 381.2433 | 144.9459 | 136.1352 | 132.9459 | 158.5137 | 184.1078 | 167.8107 | 188.4863 | 111.973 | 300.6219 | 274.3515 | 271.5407 | 364.2159 | 156.2704 | 148.4596 | 162.8919 | 149.5944 | 95.7837 | 84.0541 | 67.4056 | 132.6756 | 225.3244 | 255.0807 | 266.9455 | 177.5404 | 150.0541 | 136.0811 | 169.3781 | 151.8919 | 225.0267 | 185.6219 | 248.2163 | 187.6219 | 199.4593 | 190.3511 | 179.8652 | 174.5133 | 217.2433 | 189.5137 | 266.3781 | 307.3241 | 182.7567 | 195.8107 | 247.6489 | 241 | 222.1626 | 249.5948 | 263 | 479.27 | 189.6485 | 171.6215 | 190.8107 | 272.3241 | 185.0267 | 166.4863 | 162.3244 | 171.973 | 204.5948 | 193.3511 | 183.1085 | 165.7026 | 112.3785 | 91.0274 | 105.5407 | 129.0811 | 266.4322 | 230.73 | 279.8644 | 401.5133 | 268.3244 | 291.8104 | 290.2159 | 384.3785 | 174.8648 | 78.8648 | 83.5674 | 174.8648 | 148.297 | 136.0541 | 142.6756 | 142.8378 | 127.6215 | 124.7567 | 137.4322 | 139.9459 | 89.8919 | 66.4056 | 103.4593 | 124.5133 | 121.3244 | 106.8922 | 109.027 | 94.5678 | 333.6215 | 275.2433 | 269.5407 | 472.4593 | 145.703 | 234.27 | 206.0815 | 373.703 | 290.7841 | 232.7567 | 208.973 | 398.5404
1627 | 19.8914 | 218.9263 | 266.177 | 240.3687 | 230.2802 | 198.9853 | 191 | 194.2064 | 215.0885 | 190.1032 | 160.5014 | 176.4986 | 226.8968 | 272.4719 | 201.8083 | 209.7788 | 276.2064 | 244.9558 | 182.8525 | 184.0295 | 307.6166 | 180.8673 | 359.177 | 386.9558 | 229.5871 | 297.7346 | 259.6341 | 283.1475 | 362.2802 | 146.9705 | 126.1475 | 155.6608 | 166.8673 | 176.1327 | 168 | 175.2064 | 100.177 | 303.941 | 293.7051 | 292.6756 | 343.3244 | 142.2064 | 143.0737 | 142.3097 | 128.3244 | 102.8968 | 85.9705 | 75.8673 | 142.8525 | 243.7198 | 239.2507 | 263.0737 | 158.2949 | 159.8525 | 155.7051 | 152.2654 | 162.8378 | 200.3834 | 200.7641 | 228.2949 | 204.7346 | 173.3982 | 165.3834 | 217.4249 | 189.7788 | 224.882 | 202.7936 | 236.4571 | 284.3539 | 172.1622 | 178.2654 | 258.823 | 271.5429 | 237.7493 | 262.7936 | 248.2212 | 455.3687 | 160.4424 | 171.0147 | 191.9853 | 266.1032 | 149.5456 | 146.3097 | 156.0885 | 175.941 | 201.0442 | 203.8525 | 222.3954 | 154.177 | 115.941 | 112.6608 | 105 | 123.0885 | 230.5456 | 226.059 | 263.2654 | 375.3982 | 248.2949 | 294.9705 | 270.3097 | 354.4424 | 181.8968 | 69.1475 | 79.0737 | 181.8968 | 148.0147 | 141.9115 | 138.0737 | 154.823 | 120.118 | 129.9263 | 130.118 | 129.1622 | 87.0442 | 88.6608 | 90.2064 | 99.3834 | 117.059 | 132.6018 | 127.7198 | 94 | 329.0737 | 288.7936 | 291.6608 | 487.7788 | 167.6608 | 224.1622 | 208.941 | 400.5871 | 337.2922 | 221.177 | 210.9705 | 389.1475
1628 | 19.9024 | 214.0123 | 239.0663 | 245.9853 | 268.9042 | 193.0147 | 195.9877 | 196.9926 | 188.0663 | 204.9632 | 160.9975 | 154.0565 | 180.1154 | 256.0393 | 209.9804 | 194.0393 | 319.892 | 195.1228 | 192.9754 | 189.9853 | 230.1915 | 189.9779 | 399.8993 | 415.9288 | 198.0786 | 278.0491 | 279.9484 | 305.9435 | 406.8895 | 142.0123 | 158.919 | 145.027 | 150.0417 | 159.0417 | 152.0393 | 163.0295 | 107.9804 | 320.9583 | 293.0025 | 302.9754 | 387.8895 | 152.973 | 158.9607 | 178.9092 | 162.9141 | 106.9902 | 80.0147 | 68.0196 | 105.0933 | 194.1228 | 209.0737 | 278.9607 | 204.8846 | 163.9902 | 137.0467 | 147.0123 | 143.0491 | 184.0393 | 171.0737 | 265.9067 | 207.9926 | 181.9779 | 170.9853 | 194.0589 | 177.0319 | 210.0368 | 194.0221 | 246.973 | 301.9558 | 199.9313 | 135.1056 | 270.9705 | 242.0737 | 238 | 236.0663 | 257.9754 | 432.0565 | 183.9411 | 180.9754 | 159.081 | 260.0147 | 164.9607 | 152.9828 | 165.9754 | 195.9509 | 195.0147 | 216.9681 | 232.9754 | 157.9902 | 123.9804 | 114.9951 | 113.9779 | 119.0098 | 260.9239 | 215.027 | 239.0589 | 404.9263 | 240.0196 | 281.0344 | 256.0344 | 376.9435 | 152.0737 | 85.9583 | 79 | 152.0737 | 148.9975 | 136.0147 | 132.0147 | 154.0025 | 143.9411 | 133.9902 | 150.9484 | 118.027 | 75.0295 | 89 | 73.0417 | 94.0123 | 107.0246 | 127.0147 | 115.0319 | 100.9828 | 325.0098 | 308.9509 | 289.0074 | 466.054 | 156.0295 | 234.973 | 202.0172 | 383.0442 | 287.1252 | 196.0614 | 220.9754 | 364.0614
1629 | 19.9134 | 227.941 | 259.7639 | 227.9311 | 247.1279 | 195.8918 | 206.2066 | 192.941 | 205.0492 | 195.6754 | 165.0295 | 143.2656 | 188.9803 | 274.4721 | 214.823 | 189.0984 | 314.7148 | 234.3902 | 172.1082 | 169.1377 | 231.9311 | 173.9607 | 410.6656 | 391.2852 | 195.9115 | 257.3049 | 306.7049 | 281.8033 | 403.1377 | 141.1082 | 133.0393 | 134 | 176.8033 | 174.9902 | 182.6853 | 169.882 | 115.8721 | 333.6459 | 268.0393 | 285.0492 | 336.2951 | 151.9902 | 174.7344 | 164.8918 | 147.9607 | 92.0295 | 74.1574 | 83.0197 | 137.8131 | 227.9311 | 221.9115 | 248.1574 | 180.8918 | 174.8525 | 140.8918 | 144 | 142.9016 | 184.2459 | 187.2459 | 248.0787 | 196.8721 | 166.8426 | 176.7836 | 166.3049 | 177.9902 | 240.7738 | 209.0689 | 206.3639 | 302.1574 | 160.0984 | 160.9902 | 218.8459 | 232.2557 | 233.541 | 236.9508 | 240.7934 | 433.3639 | 165.2164 | 193.823 | 162.2262 | 262.6262 | 155.2164 | 153.1574 | 150.0984 | 175.2164 | 198.8426 | 195.9115 | 228.823 | 186.7541 | 115.0295 | 126.7934 | 116.941 | 120.8918 | 256.0098 | 259.8525 | 233.1672 | 373.118 | 246.2262 | 262.8721 | 239.5115 | 376.7049 | 180.6951 | 71.9902 | 82.8131 | 180.6951 | 122.1672 | 140.8918 | 149.1279 | 129.1672 | 136.8525 | 125.9705 | 133.9902 | 154.6262 | 75.9213 | 71.0787 | 80.059 | 106.9803 | 119.0098 | 97.1574 | 127.9115 | 95.9115 | 291.1574 | 281.459 | 276.2361 | 495.118 | 183.8426 | 206.3246 | 241.7738 | 437.0853 | 305.8328 | 198.9607 | 227.9508 | 418.882
1630 | 19.9244 | 221.0264 | 237.3055 | 221.531 | 259.3583 | 184.9336 | 226.4026 | 186.9336 | 209.5575 | 163.8851 | 167.8672 | 170.3098 | 187.0221 | 321.0707 | 196.646 | 199 | 285.6017 | 174.1506 | 183.177 | 182.8451 | 225.7081 | 170.0664 | 376.8894 | 420.0664 | 187.531 | 286.6281 | 276.3583 | 262.4425 | 416.2919 | 151.4911 | 136.9557 | 133.9779 | 157.4204 | 174.2434 | 151.5089 | 158.1328 | 103.1328 | 298.0885 | 272.0443 | 290.4868 | 365.7124 | 150.823 | 148.5532 | 153.2698 | 143.7345 | 94.4468 | 89.2919 | 84.7345 | 119.2213 | 221 | 213.2213 | 263.1592 | 169.6238 | 159.3583 | 130.5089 | 143.6902 | 133.1328 | 208.3583 | 211.6681 | 255.8009 | 183.4911 | 151.3983 | 154.6902 | 196.823 | 176.8451 | 217.9115 | 215.4468 | 243.3319 | 317.0043 | 169.5353 | 160.6417 | 302.6281 | 257.7124 | 287.0707 | 232.1328 | 220.3983 | 469.2256 | 186.7124 | 175.7787 | 184.7787 | 226.4161 | 176.0264 | 168.9115 | 159.7787 | 196.4247 | 183.7744 | 186.3804 | 211.4204 | 161.646 | 117.8451 | 106.1328 | 110.9336 | 110.8851 | 256.8672 | 245.3983 | 249.5796 | 385 | 269.2876 | 249.8451 | 289.8494 | 347.7302 | 150.5089 | 71.1106 | 64.0885 | 150.5089 | 139.2655 | 130.5089 | 161.7345 | 146.0664 | 122.0885 | 123.177 | 133.0885 | 117.0443 | 68.0885 | 79.1991 | 85.646 | 105.0664 | 120.1991 | 112.6902 | 118.7345 | 87.177 | 306.9779 | 228.0178 | 300 | 506.0928 | 167.4911 | 238.823 | 218.6238 | 346.5046 | 289.5532 | 195.6417 | 222.6238 | 407.4204
1631 | 19.9354 | 180.7878 | 293.3135 | 245.5507 | 231.2753 | 181.4493 | 200.0688 | 184.5507 | 189.9312 | 201.6233 | 161.8623 | 183.1051 | 188.0344 | 279.9656 | 181.8604 | 198.8279 | 269.067 | 224.9312 | 190.4493 | 175.7247 | 255.6921 | 172.8279 | 372.826 | 423.2409 | 210.2772 | 227.9962 | 248.5163 | 281.8623 | 385.5851 | 129 | 134.7591 | 133.2409 | 175.4837 | 183.6921 | 174.3098 | 163.5526 | 108.6558 | 301.6902 | 273.1396 | 311.2428 | 353.0688 | 143.3442 | 172.3461 | 122.1358 | 133.1014 | 70.1721 | 58.7228 | 73 | 129.4819 | 220.2772 | 222.1051 | 228.0651 | 154.6176 | 130.7247 | 152.3461 | 130.7572 | 138.9312 | 180.1033 | 196.4493 | 247.9981 | 162.239 | 168.8967 | 142.1014 | 189.1033 | 169.5181 | 214.4474 | 191.3786 | 257.9312 | 273 | 150.2735 | 187.6233 | 242.9637 | 244.6214 | 246.0344 | 237.3116 | 236.7954 | 435.6884 | 173.7935 | 166.1377 | 174.8623 | 287.9675 | 134.067 | 164.0019 | 150.2753 | 170.0019 | 217.587 | 159.6195 | 229.587 | 146.5163 | 111.6539 | 112.7228 | 108.2065 | 149.4493 | 251.5507 | 262.1396 | 232.1014 | 386.5144 | 281.174 | 242.2428 | 238.4493 | 379.4493 | 172.4149 | 76.5163 | 68.7228 | 172.4149 | 149.8298 | 152.5181 | 149.5181 | 148.9312 | 125.6558 | 131.2065 | 136.6902 | 119 | 72 | 87.6214 | 70.6539 | 107.8279 | 128.1396 | 99.1033 | 107.5163 | 94.6558 | 305.6214 | 271.8642 | 298.5544 | 465.9312 | 145.2065 | 230.2772 | 201.3805 | 412.5181 | 314 | 223.7935 | 205.7935 | 425.3461
1632 | 19.9464 | 256.5244 | 246.8408 | 262.7749 | 239.4204 | 165.5329 | 201.5329 | 199.673 | 188.2803 | 164.7749 | 158.2803 | 157.8599 | 188.1125 | 279.8875 | 205.4862 | 194.3737 | 298.4862 | 223.6072 | 175.6539 | 168.1868 | 219.327 | 168.0467 | 396.5138 | 428.1784 | 190.2335 | 283.6263 | 264.0276 | 278.6539 | 400.3652 | 128.7197 | 128.327 | 140.0934 | 160.4395 | 147.3737 | 182.2059 | 151.1868 | 98.7665 | 294.6348 | 249.327 | 290.2335 | 353.879 | 153.6539 | 153.7197 | 153.5329 | 162.552 | 74.4862 | 79.4671 | 73 | 142.5329 | 200.0934 | 198.1677 | 284.0382 | 199.2994 | 125.0552 | 135.9618 | 150.5987 | 136.7197 | 182.673 | 180.5796 | 274.7856 | 195.1784 | 165.9533 | 173 | 192.6072 | 155.8599 | 226.1592 | 202.1401 | 255.2059 | 272.8132 | 185.0191 | 150.2611 | 270.0467 | 233.6263 | 245.8323 | 218.8875 | 204.9618 | 454.0191 | 168.2335 | 169.0658 | 171.0934 | 259.8408 | 164.0934 | 137.0743 | 158.1868 | 142.9342 | 206.2335 | 177.8875 | 217.6263 | 159.8323 | 129.5329 | 131.3185 | 114.6072 | 134.3737 | 265.8323 | 238.0934 | 263.4204 | 426.0106 | 257.5329 | 222.121 | 226.0828 | 364.9342 | 156.7006 | 89.7856 | 87.879 | 156.7006 | 117.2803 | 138.7665 | 136.0467 | 146.9066 | 115.9066 | 135.879 | 128.9809 | 119.4204 | 72.0934 | 77.2335 | 88.2994 | 103.327 | 103.7665 | 102.0934 | 120.879 | 85.1868 | 295 | 240.9342 | 259.448 | 462.7856 | 151.6072 | 209.4862 | 183.8132 | 398.9066 | 312.4119 | 217.8132 | 199.8599 | 405.9257
1633 | 19.9574 | 207.18 | 264.413 | 296.817 | 247.174 | 158.065 | 191.351 | 193.059 | 192.82 | 198.935 | 163.233 | 153.702 | 171.475 | 296.171 | 196.888 | 201.528 | 286.466 | 235.233 | 188.292 | 172.354 | 225.941 | 168.41 | 406.41 | 392.062 | 195.177 | 275.882 | 283.466 | 290.879 | 369.537 | 122.528 | 135.413 | 141.705 | 149.413 | 156.593 | 165.705 | 154.941 | 93.528 | 324.755 | 256.885 | 293.997 | 330.056 | 166.882 | 149.003 | 145.18 | 134.062 | 65.416 | 89.295 | 73.354 | 132.764 | 204.006 | 221.351 | 244.826 | 184.882 | 166.174 | 174.112 | 122.944 | 131.826 | 176.177 | 173.829 | 249.292 | 159.829 | 164.41 | 173.885 | 204.528 | 153.354 | 209 | 205.236 | 238.764 | 269 | 165.59 | 174.879 | 270.41 | 225.056 | 222.944 | 235.938 | 243.171 | 433.469 | 172.764 | 151.121 | 173.472 | 275.879 | 164.525 | 158.233 | 161.941 | 163.239 | 207.873 | 195.115 | 210.472 | 135.823 | 118.997 | 97.354 | 126.056 | 141.174 | 241.882 | 239.882 | 269.581 | 369.425 | 249.18 | 244.469 | 286.522 | 384.531 | 171.295 | 65.118 | 66.239 | 171.295 | 124.062 | 134.059 | 135.702 | 144.469 | 115.18 | 113.826 | 147.171 | 128.059 | 73.764 | 81.528 | 74.118 | 109.351 | 99.767 | 104.59 | 98.295 | 88.233 | 296.298 | 260.177 | 289.472 | 438.767 | 163.764 | 200.003 | 181.829 | 397.826 | 282.124 | 213.587 | 196.941 | 384.531
1634 | 19.9684 | 226 | 269.1464 | 262 | 234.6416 | 189.289 | 180.2158 | 195.7823 | 173.7861 | 167.4258 | 150.0019 | 135.9229 | 194.8574 | 265.7168 | 223.5067 | 194.0713 | 262.8555 | 221.5742 | 177.2139 | 176.5742 | 227.7804 | 159.9981 | 395.1464 | 406.1484 | 194.6493 | 272.289 | 258.3584 | 273.9981 | 407.2909 | 115.2852 | 140.5029 | 137.0713 | 156.7842 | 182.7842 | 160.2158 | 153.7861 | 87.3545 | 274.711 | 269.289 | 277.7148 | 315.9981 | 164.2158 | 162.9326 | 163.0019 | 150.7148 | 87.3584 | 93.9287 | 78.8574 | 129.7842 | 232.7206 | 211.2852 | 256.7881 | 182.7148 | 152.9287 | 144 | 136.4316 | 144.0732 | 179.9981 | 201.3603 | 236.8593 | 188.2871 | 156.8536 | 185.6474 | 195.5029 | 158.501 | 209.2139 | 209 | 237.21 | 270.1407 | 175.3565 | 158.5684 | 262.4258 | 212.8516 | 237.2158 | 220.5684 | 216.21 | 425.9981 | 168.1445 | 169.5703 | 179.5742 | 256.5742 | 142.1407 | 146.9981 | 160.8574 | 182.4297 | 160.9942 | 182.5684 | 218.2852 | 133.7842 | 104.0694 | 103.499 | 112.2832 | 127.0019 | 239.2871 | 237.7148 | 234.1368 | 435.4374 | 266.0751 | 236 | 245.289 | 391.8593 | 173.8613 | 67.5703 | 84.289 | 173.8613 | 142.711 | 134.5723 | 114.501 | 136.3565 | 133.5723 | 127.0713 | 117.3584 | 129.2139 | 71.0694 | 73.7148 | 75.7861 | 98.9287 | 112.0713 | 113.8574 | 104.9961 | 76.1426 | 316.2158 | 260.9326 | 297.9268 | 452.8536 | 159.1445 | 215.0732 | 211.9268 | 410.7148 | 312.578 | 205.7168 | 196.9268 | 393.2139
1635 | 19.9794 | 224.6627 | 244.4149 | 263.7552 | 243.0836 | 169.0896 | 172.2567 | 218.3313 | 171.8358 | 186.6687 | 138.5045 | 171.409 | 190.9104 | 251.9254 | 179.003 | 192.994 | 275.0866 | 201.997 | 177.7433 | 157.7493 | 256.9791 | 172.1642 | 371.8358 | 371.3403 | 153.006 | 250.5015 | 251.003 | 282.3194 | 374.3403 | 119.9194 | 124.009 | 137.0806 | 167.5851 | 191.6627 | 149.5821 | 153.591 | 105 | 296.7493 | 250.9284 | 274.4179 | 330.2537 | 153.9164 | 138.2567 | 150.1672 | 144.994 | 78.2478 | 92.2478 | 78.1701 | 138.3284 | 193.1761 | 213.997 | 242.2537 | 177.7463 | 149.8269 | 144.5851 | 118.0896 | 130.8299 | 192.8328 | 179.9164 | 225.0925 | 181.7582 | 178.0746 | 154.9164 | 176.6687 | 153.6716 | 211.1642 | 208.4149 | 264.6627 | 285.3343 | 177.6597 | 177.1612 | 280.5821 | 247.4925 | 226.4149 | 238.994 | 243.7463 | 435.0716 | 158.7552 | 176.8328 | 162.7552 | 237.5821 | 156.0806 | 161.1701 | 159.5851 | 175.4179 | 200.1672 | 202.8328 | 220.3284 | 145.0866 | 117.7493 | 111.003 | 127.1612 | 114.2507 | 231.5045 | 235.0866 | 273.6627 | 380.5134 | 242.9224 | 236.2507 | 224.0866 | 375.997 | 147.7552 | 73.997 | 62.3343 | 147.7552 | 161.3224 | 126.5761 | 113.4328 | 140.4149 | 125.5761 | 126.8299 | 109.2507 | 132.5851 | 85 | 71.0866 | 73.1672 | 98.4179 | 112.6657 | 110.3284 | 128.9104 | 79.0866 | 303.9104 | 237.0925 | 306.7403 | 475.9134 | 148.5851 | 203.2507 | 223.4985 | 405.5791 | 271.5104 | 190.3373 | 208.8328 | 393.8269
1636 | 19.9904 | 208.3702 | 238.5752 | 281.9454 | 245.3422 | 187.7816 | 197.767 | 210.6165 | 183.1092 | 191.8908 | 153.3702 | 142.8083 | 172.0267 | 283.233 | 189.4248 | 171.8629 | 285.0825 | 192.6298 | 153.8629 | 155.5752 | 185.1772 | 163.2876 | 380.0413 | 401.506 | 176.7257 | 256.3835 | 263.0546 | 236.6565 | 397.2197 | 129.0413 | 154.1238 | 128.0546 | 171.2197 | 176.233 | 145.1917 | 179.7949 | 104.9041 | 292.5619 | 288.9187 | 277.466 | 343.6165 | 153.4794 | 160.5073 | 152.9587 | 124.3422 | 72.2051 | 85.2464 | 87.8362 | 119.6165 | 225.2197 | 203.0959 | 254.8495 | 166.0133 | 125.6165 | 151.3835 | 143.7803 | 118.4794 | 191.5752 | 180.3422 | 257.3702 | 212.767 | 146.8629 | 154.2876 | 185.8216 | 170.5619 | 202.2876 | 201.4248 | 248.2743 | 283.8228 | 153.534 | 159.4927 | 276.6711 | 220.4794 | 221.1505 | 218.2464 | 232.6844 | 395.3556 | 174.4527 | 175.3835 | 182.0959 | 231.8495 | 146.2876 | 170.2609 | 166 | 178.8495 | 200.0825 | 199.8495 | 199.6032 | 156.8083 | 115.1917 | 122.4794 | 105.5619 | 118.7257 | 250.2051 | 244.699 | 260.1505 | 425.9733 | 261.5073 | 240.6298 | 237.0546 | 366.0546 | 164.699 | 64.4381 | 67.3422 | 164.699 | 124.1638 | 105.1772 | 171.466 | 134.1917 | 101.0133 | 113.3289 | 112.6711 | 137.1784 | 83.466 | 82.7124 | 73.9454 | 103.4794 | 111.0133 | 92.3835 | 106.4794 | 88.1238 | 281.4794 | 268.3156 | 269.7536 | 464.1917 | 156.6298 | 205.6165 | 217.8083 | 393.4514 | 309.6165 | 205.2876 | 206.9041 | 376.2318
1637 | 20.0014 | 196.5694 | 244.1082 | 274.0551 | 253.3489 | 152.0021 | 204.8918 | 207.9735 | 201.053 | 172.3245 | 140.7041 | 142.298 | 209.5673 | 271.4571 | 183.8918 | 181.6225 | 264.8367 | 208.3245 | 162 | 162.7306 | 232.3775 | 167.4061 | 371.9735 | 439.9979 | 188.8898 | 256.7551 | 275.1633 | 289.0265 | 369.4857 | 119.4592 | 127.5408 | 135.7286 | 148.4613 | 168.0265 | 150.4613 | 161.1633 | 104.1082 | 278.8918 | 243.3796 | 266.5694 | 339.4592 | 156.9184 | 136.0265 | 161.5673 | 135.5939 | 91.3775 | 93.8632 | 60.8123 | 117.8388 | 200.7306 | 203.6755 | 240.8632 | 182.6204 | 124.2714 | 154.6755 | 166.5387 | 124.6225 | 194.5122 | 191.4857 | 241.351 | 218.5939 | 153.1612 | 156.2429 | 201.5939 | 158.6225 | 203.5939 | 195.351 | 234.0551 | 238.7857 | 168.3245 | 180.6204 | 281.3775 | 226.0816 | 231.1347 | 227.8367 | 255.2959 | 424.1877 | 149.3286 | 179.7571 | 183.649 | 225.2184 | 148.8918 | 135.1082 | 166.4327 | 167.351 | 182 | 187.4857 | 180.8939 | 155.4327 | 117.649 | 126.8918 | 95.0286 | 133.1612 | 265.4836 | 225.2714 | 266.8898 | 390.1633 | 238.7571 | 252.9714 | 249.4878 | 375.4592 | 146.2449 | 77.5673 | 80.7571 | 146.2449 | 149.6734 | 150.4306 | 154.9449 | 138.3796 | 120.5408 | 108.6225 | 119.649 | 120.649 | 68.7837 | 78.702 | 64.8653 | 108.5408 | 129.351 | 99.1368 | 111.649 | 62.4061 | 288.4878 | 245.4878 | 261.7857 | 462.7551 | 173.298 | 204.1633 | 217.4061 | 425.1347 | 301.7816 | 208 | 206.4327 | 431.9714
1638 | 20.0124 | 221.7841 | 242.2295 | 289.5545 | 216.2046 | 183.9159 | 206.0477 | 215.8795 | 185.9636 | 171.3863 | 161.1932 | 154.325 | 201.5432 | 237.4932 | 184.0841 | 195.4818 | 248.6864 | 212.6864 | 159.2295 | 177 | 217.4341 | 177.0727 | 378.0727 | 406.7591 | 154.6386 | 234.8182 | 289.5068 | 277.1454 | 355.9159 | 115.7227 | 133.6864 | 118.6864 | 175.7954 | 162.0477 | 177.675 | 176.1091 | 105.8432 | 277.3977 | 261.1932 | 291.1818 | 336.4455 | 147.1568 | 130.8909 | 157.0364 | 123.7591 | 76.1932 | 68.6023 | 82.3136 | 134.2046 | 213.675 | 202.9273 | 221.7455 | 165.1682 | 143.6023 | 150.675 | 138.1205 | 136.3136 | 174.1205 | 181.0477 | 233.7113 | 204.7113 | 139.5659 | 152.65 | 191.0841 | 168.8682 | 193.5659 | 191.6864 | 249.3136 | 271.0841 | 171.1205 | 160.5182 | 268.1205 | 233.3136 | 225.8068 | 212.7705 | 237.6978 | 399.0727 | 212.3386 | 186.9636 | 189.3614 | 260.2409 | 149.9273 | 137.2046 | 164.5795 | 164.1682 | 183.325 | 177.0477 | 214.4455 | 156.4704 | 120.1091 | 124.5545 | 118.0727 | 119.8068 | 219.6386 | 244.1205 | 231.0727 | 406.7954 | 245.1205 | 231.7341 | 265.7841 | 366.6636 | 167.2181 | 74.7227 | 85.7954 | 167.2181 | 115.0841 | 120.1568 | 138.9636 | 156.9159 | 123.9159 | 122 | 126.325 | 125.7591 | 69.0477 | 69.4455 | 72 | 111.5545 | 122.1932 | 120.9045 | 115.9159 | 74.2409 | 301.5318 | 262.6273 | 294.9273 | 444.4319 | 178.3386 | 221.1568 | 224.4227 | 412.0977 | 270.2523 | 205.8318 | 207.7113 | 403.025
1639 | 20.0234 | 193.1239 | 225.7169 | 274.7521 | 225.9292 | 178.1239 | 216.4866 | 216.4602 | 191.8053 | 145.7965 | 149.1947 | 162.0088 | 170.593 | 271.0975 | 190.6725 | 197.5398 | 260.6018 | 224.8673 | 140.4602 | 173.9468 | 208.3894 | 162.6018 | 362.938 | 405 | 158.6373 | 266.4158 | 263.469 | 256.301 | 349.8584 | 123.7877 | 142.8141 | 131.6637 | 165.5398 | 174.2123 | 166.1416 | 154.0708 | 108.9468 | 278.7081 | 250.9204 | 261.1771 | 347.531 | 139.2743 | 150.5398 | 152.6549 | 123.8584 | 65.7877 | 74.9912 | 72.2567 | 141.6725 | 201.0796 | 217.531 | 272.593 | 181.9292 | 144.2831 | 140.3363 | 138.8673 | 126.6549 | 175.1327 | 196.531 | 219.5222 | 187.0708 | 150.0708 | 171.2035 | 196.4778 | 145.6018 | 207.6549 | 202.938 | 239.5222 | 282.4514 | 170.8053 | 155.938 | 265.2831 | 221.6637 | 235.8141 | 232.2035 | 265.8937 | 388.4514 | 167.8141 | 187.8937 | 191.6018 | 259.6106 | 164.2655 | 147.062 | 129.1151 | 178.8053 | 191.0088 | 190.8053 | 222.3451 | 139.4602 | 100.8584 | 111.8761 | 105.8584 | 134.4602 | 219.9204 | 242.3451 | 223.5046 | 396.1416 | 244.1416 | 255.5486 | 240.8408 | 338.3187 | 128.9824 | 78.6725 | 77.9292 | 128.9824 | 123.7965 | 114.3982 | 147.3275 | 148.8673 | 114.6725 | 121.7345 | 135.8673 | 124.3982 | 83.6018 | 79.8673 | 75.1859 | 101.1327 | 109.1327 | 94.0532 | 105.0796 | 76.7965 | 250.3187 | 241.7169 | 306.2123 | 501.6194 | 138.9912 | 213.5398 | 193.1239 | 372.1859 | 292.0088 | 190 | 195.6461 | 376.7877
1640 | 20.0344 | 208.3053 | 252.1756 | 254.2519 | 234.4656 | 188.6641 | 187.2901 | 222.3741 | 183.7099 | 151.7252 | 154.6794 | 152.3359 | 180.2748 | 312.6871 | 183.3053 | 192.2061 | 260.1756 | 221.0993 | 153.4809 | 159.3664 | 220.3741 | 159.1298 | 356.5801 | 401.084 | 203.4504 | 250.3969 | 258.5496 | 295.8779 | 367.5114 | 138.8092 | 125.6259 | 136 | 154.5496 | 152.9542 | 152.8397 | 149.0153 | 89.7252 | 311.3588 | 266.3741 | 297.0305 | 347.3741 | 129.7404 | 140.2748 | 164.1985 | 133.9695 | 82.9847 | 86.8397 | 87.4351 | 129.9542 | 180.9847 | 220.1298 | 280.5344 | 188 | 121.0153 | 140.6412 | 140.1756 | 144.8702 | 173.8244 | 197.8244 | 235.855 | 176.6489 | 146.4656 | 162.9542 | 183.0458 | 144.5954 | 220.0687 | 196.145 | 256.4351 | 300.5038 | 164.7404 | 152.0458 | 241.8702 | 224.2596 | 218.1908 | 223.9542 | 226.145 | 413.1756 | 150.3359 | 153.3664 | 186.5344 | 244.5801 | 161.939 | 152.1145 | 153.0993 | 168.8244 | 180.7557 | 181.6947 | 207.1756 | 145.5191 | 110.2443 | 99.5954 | 118.7252 | 142.1145 | 239.2901 | 226.4504 | 260.0535 | 385.8855 | 233.8855 | 232.855 | 273.0535 | 356.229 | 153.6947 | 71.4504 | 82.5496 | 153.6947 | 127.1145 | 117.2901 | 152.084 | 146.1145 | 104.5496 | 118.4046 | 134.7099 | 126.4199 | 79.8397 | 79.145 | 95.855 | 102.4351 | 110.2901 | 111.9695 | 90.2061 | 80.5496 | 267.7939 | 263.2443 | 287.2748 | 470.5344 | 152 | 203.5649 | 208.1603 | 396.0458 | 280.8855 | 193.4809 | 223.0993 | 392.9695
1641 | 20.0454 | 217.416 | 263.8987 | 290.7069 | 248.056 | 169.6293 | 191.3599 | 204.1466 | 181.8427 | 158.3599 | 143.5173 | 167.5948 | 178.0453 | 243.6401 | 186.2801 | 209.9202 | 269.2241 | 203.2134 | 172.4267 | 182.9762 | 201.8319 | 156.5173 | 357.4827 | 385.2371 | 212.3147 | 286.4482 | 247.9547 | 248.3147 | 397.7974 | 121.7306 | 145.4267 | 133.7974 | 146.472 | 141.7651 | 146.1573 | 154.3707 | 94.6293 | 281.6401 | 245.6293 | 306.9547 | 329.3039 | 137.9547 | 140.4052 | 132.8319 | 120.7414 | 78.1013 | 82.2026 | 88.4267 | 117.1918 | 179.7199 | 216.7306 | 262.2241 | 185.1681 | 134.0798 | 163.9095 | 149.5388 | 147.1681 | 165.2479 | 188.9332 | 232.0108 | 154.7759 | 162.5733 | 146.416 | 202.1013 | 155.2586 | 195.0905 | 201.0905 | 259.944 | 275.416 | 171.6961 | 169.2134 | 247.472 | 214.472 | 237.7866 | 209.6186 | 223.6961 | 428.0453 | 170.944 | 178.8642 | 171.528 | 247.2134 | 140.0453 | 145.0905 | 136.7866 | 156.7866 | 192.5948 | 175.2586 | 217.6401 | 127.5173 | 97.7759 | 106.4827 | 124.5733 | 130.056 | 238.1681 | 236.5733 | 228.6186 | 394.0108 | 247.3599 | 230.7414 | 238.7866 | 338.5625 | 151.8212 | 75.9095 | 74 | 151.8212 | 120.8772 | 116.1681 | 130.1013 | 135.944 | 98.5173 | 112.6186 | 129.6961 | 125.8319 | 74.1013 | 80.6293 | 92.9547 | 104.2134 | 113.2586 | 101.5733 | 113.7414 | 76.7199 | 248.6638 | 247 | 279.5388 | 455.528 | 153.2586 | 201.944 | 215 | 410.8534 | 289.1681 | 212.2694 | 207.2586 | 380.8427
1642 | 20.0564 | 220.4203 | 248.8405 | 261.4101 | 240.4557 | 173 | 203.4911 | 217.2684 | 182.357 | 170.4911 | 155.4734 | 146.8835 | 185.0988 | 235.5797 | 167.2759 | 173.357 | 247.643 | 197.643 | 166.2051 | 143.562 | 219.7139 | 170.5089 | 348.919 | 419.081 | 207.3848 | 253.6329 | 242.0886 | 251.8658 | 387.8658 | 130.1519 | 136.8304 | 127.5974 | 146.9646 | 175.2861 | 145.1342 | 153.8835 | 94.6076 | 268.4911 | 254.0886 | 295.4911 | 338.6886 | 131.919 | 157.5545 | 146.9823 | 118.7493 | 81.1165 | 91.1165 | 78.9823 | 139.5722 | 203.643 | 225.6608 | 240.1342 | 170.8481 | 163.519 | 143.357 | 135.919 | 125.7241 | 189.1873 | 203.8937 | 219.562 | 176.6962 | 170.1519 | 157.7316 | 204.6076 | 166.2405 | 214.9646 | 216.2152 | 262.4557 | 282.6608 | 153.1519 | 165.6785 | 249.3215 | 214.7949 | 237.7595 | 229.3392 | 211.0886 | 435.2684 | 174.9823 | 150.2405 | 173.2405 | 239.438 | 148.7949 | 159.0279 | 142.0177 | 159.9823 | 166.7949 | 185.9012 | 212.124 | 139.8127 | 115.7949 | 97.2405 | 129.2684 | 125.5089 | 224.0177 | 245.6785 | 245.7949 | 376.643 | 259.3215 | 228.4101 | 249.6152 | 365.2051 | 179.6076 | 57.0532 | 75.1873 | 179.6076 | 149.0354 | 105.0709 | 137.8658 | 139.9823 | 108.7772 | 125.7241 | 112.8481 | 141.8481 | 77.2861 | 83.6608 | 81.6608 | 98.8127 | 120.8481 | 111.6962 | 106.1519 | 99.2861 | 275.2684 | 246.4911 | 263.0354 | 460.6329 | 157.1165 | 205.1342 | 213.8127 | 391.7949 | 280.1063 | 203.1696 | 212.4734 | 381.357
1643 | 20.0674 | 198.7276 | 206.9106 | 282.4543 | 226.1809 | 174.2734 | 196.0884 | 196.0894 | 188.6362 | 166.7266 | 149.6372 | 156.0884 | 164.7266 | 254.7256 | 194.1799 | 179.2724 | 236.0884 | 187.7256 | 175.9085 | 165.7297 | 229.7256 | 168.4522 | 371.5447 | 397.7287 | 178.2744 | 286.6341 | 262.8161 | 263.1829 | 388.0864 | 125.8181 | 134.1809 | 150.998 | 137.5457 | 164.9106 | 135.6362 | 165.2713 | 80.9106 | 262.1799 | 278.4543 | 292.6362 | 309.8242 | 150.5427 | 130.9116 | 138.5437 | 139.2713 | 67.3638 | 74.2713 | 72.7266 | 116.2744 | 198.2734 | 224.5457 | 233.9106 | 172.999 | 126.9116 | 144.9065 | 160.7276 | 97.7287 | 190.6341 | 176.9106 | 236.2724 | 180.998 | 159.997 | 171.4522 | 189.2734 | 180.2693 | 205.5457 | 190.0945 | 249.2724 | 276.998 | 149.5457 | 166.0894 | 246.9096 | 204.7276 | 215.7266 | 225.1789 | 230.3608 | 413.3618 | 167.6352 | 166.998 | 196.5467 | 224.5467 | 137.6362 | 124.1839 | 152.4573 | 156.0915 | 159.6382 | 203.1809 | 235.9035 | 136.001 | 106.8191 | 116.9085 | 111.9096 | 128.1819 | 228.6362 | 242.6331 | 235 | 377.0945 | 248.7236 | 242.3598 | 267.9929 | 366.9045 | 164.4553 | 72.9096 | 80.0904 | 164.4553 | 154.998 | 121.1799 | 146.4543 | 134.0904 | 94.0915 | 98.0925 | 122.0935 | 141.4522 | 65.6372 | 80.7266 | 80.3638 | 96.0925 | 118.6331 | 115.998 | 104.9106 | 86 | 257.5457 | 245.0915 | 272.2724 | 495.2713 | 144.2734 | 199.2744 | 207.2724 | 383.1829 | 296.1748 | 199.2703 | 207.1829 | 389.8191
1644 | 20.0784 | 201.8058 | 224.0782 | 279.029 | 220.1362 | 178.6406 | 177.6898 | 184.1652 | 187.1942 | 164.3014 | 152.5334 | 139.2434 | 162.3014 | 237.087 | 178.5536 | 178.1362 | 218.6608 | 179.4087 | 166.3392 | 181.3682 | 212.8638 | 147.6116 | 375.2927 | 410.5826 | 190.087 | 262.145 | 250.9913 | 273.7768 | 359.2637 | 121.3102 | 124.8348 | 133.7768 | 139.029 | 169.0667 | 137.4956 | 150.2812 | 92.8348 | 252.9623 | 274.058 | 284.5913 | 349.2117 | 126.971 | 144.8435 | 125.913 | 129.7188 | 69.5826 | 69.9623 | 69.913 | 127.5044 | 200.3102 | 225.058 | 249.3304 | 169.4376 | 146.2812 | 128.6029 | 156.8145 | 113.3014 | 170.8058 | 180.6783 | 232.4174 | 169.0202 | 133.5826 | 151.3884 | 189.3682 | 148.2522 | 202.7566 | 221.8058 | 243.4754 | 263.2724 | 150.8348 | 162.5159 | 247.5044 | 213.4376 | 210 | 200.942 | 209.4666 | 395.2522 | 158.3594 | 150.1652 | 205.4754 | 236.1942 | 133.6696 | 143.971 | 169.6203 | 156.7276 | 176.3884 | 193.4464 | 195.1275 | 145.3884 | 111.8928 | 109.8638 | 117.5536 | 133.6608 | 228.7478 | 220.6029 | 233.8348 | 399.4841 | 216.5826 | 210.7188 | 213.3507 | 337.8463 | 164.0087 | 77.3884 | 77.5536 | 164.0087 | 132.145 | 108.0782 | 144.971 | 128.6406 | 101.5246 | 109.2812 | 140.3392 | 120.4174 | 69.8928 | 73.6406 | 78.5044 | 108.8348 | 94.2724 | 99.942 | 116.6406 | 83.8638 | 268.3507 | 251.942 | 264.7276 | 483.3884 | 148.4464 | 210.116 | 208.2724 | 392.2232 | 244.2927 | 181.3014 | 218.3594 | 394.8928
1645 | 20.0894 | 201 | 236.2507 | 277.4779 | 228.7935 | 176.7169 | 191.3481 | 192.5104 | 183.4572 | 170.5663 | 138.5428 | 159.5546 | 177.0325 | 225.413 | 185.6195 | 189.0649 | 235.5221 | 202.6961 | 151.5428 | 160.2714 | 206.0649 | 147.0325 | 405.4039 | 418.3688 | 178.413 | 239.3156 | 272.7728 | 276.7935 | 394.0234 | 107.8584 | 119.587 | 138.6519 | 139.5428 | 133.239 | 149.1091 | 141.2714 | 88.6195 | 284.3481 | 267.239 | 268.7376 | 294.6844 | 134.9234 | 116.5753 | 140.2714 | 135.0117 | 75.5104 | 98.4572 | 79.7286 | 114.0325 | 185.826 | 216.7935 | 258.3805 | 184.3598 | 133.3481 | 151.2507 | 139.3896 | 131.5428 | 169.7935 | 152.3246 | 229.174 | 189.0442 | 137.4454 | 151.5546 | 169.9234 | 141.413 | 188.4337 | 221.6195 | 237.7493 | 271.7078 | 146.2065 | 195.3922 | 231.761 | 237.6519 | 210.6195 | 205.2831 | 225.3156 | 390.6844 | 159.0442 | 152.1091 | 195.413 | 235.5546 | 126.4779 | 144.4896 | 137.5546 | 139.413 | 179.6519 | 186.174 | 222.6754 | 147.826 | 100.239 | 106.1623 | 118.2182 | 142.0559 | 237.8584 | 244.4896 | 231.8909 | 364.1299 | 222.5104 | 222.826 | 247.174 | 380.8052 | 137.6844 | 77.1416 | 82.9675 | 137.6844 | 114.8909 | 121.4896 | 146.3156 | 126.5104 | 106.0117 | 102.5428 | 124.3039 | 115.9351 | 58.6519 | 66.3481 | 69.5753 | 100.6961 | 111.587 | 110.4779 | 106.2507 | 78.0974 | 298.6637 | 259.7935 | 254.0208 | 480.8701 | 141.7935 | 197.9234 | 213.1975 | 386.1091 | 286.7935 | 192.3156 | 222.3481 | 381.587
1646 | 20.1004 | 202.5315 | 221.9381 | 289.1878 | 223.4055 | 183.7801 | 185.2188 | 200.3116 | 164.2497 | 140.0309 | 152.0619 | 153.1249 | 175.3116 | 234.6574 | 187.1249 | 197 | 229.9381 | 193.5005 | 167.0309 | 170.9691 | 212.0309 | 152.3127 | 355.4685 | 419.5283 | 182.6254 | 227.2497 | 245.6564 | 284.5326 | 355.0011 | 118.2818 | 121.2818 | 142.8122 | 149.7801 | 146.9701 | 138 | 150.4376 | 91.6564 | 276.4685 | 268.7182 | 304.7492 | 311.1259 | 148.2497 | 138.2188 | 148.7812 | 112.6564 | 87.0309 | 82.0939 | 74.2818 | 117.3436 | 190.9691 | 213.5934 | 254.2497 | 155.6564 | 127.6564 | 133.2188 | 184.1558 | 143.0928 | 163.968 | 193.7503 | 222.4995 | 168.4685 | 142.1249 | 151.032 | 189.5945 | 141.4685 | 214.8122 | 221.1558 | 253.3746 | 243.2199 | 154.8762 | 163.8741 | 234.6574 | 241.8122 | 210.5934 | 200.6265 | 225.2828 | 405.5945 | 142.1878 | 140.3436 | 193.7182 | 235.6884 | 136.2188 | 133.8442 | 135.9381 | 140.7812 | 187.3127 | 186.7193 | 186.5005 | 151.8751 | 103.9061 | 120.968 | 98.1878 | 110.0641 | 241.7182 | 225.5304 | 238.4055 | 375.1868 | 237.5315 | 229.2818 | 244.2188 | 353.937 | 150.6254 | 70.2188 | 77.9061 | 150.6254 | 125.7812 | 111.063 | 140.3757 | 137.1558 | 91.9701 | 116.2807 | 134.8122 | 111.7193 | 66.9691 | 67.2199 | 84.4365 | 88.6564 | 108.0309 | 114.3116 | 95.8762 | 84.3116 | 269.7481 | 261.6254 | 281.8751 | 462.3746 | 139.6873 | 216.063 | 171.0011 | 377.4066 | 280.7492 | 184.8442 | 214.2497 | 377.5934
1647 | 20.1114 | 204.765 | 235.5379 | 297.2311 | 211.6215 | 159.6175 | 188.3107 | 184.6215 | 166.0876 | 138.3147 | 144.239 | 143.761 | 160.3147 | 249.2909 | 178.4542 | 195.6136 | 236.6056 | 201.6096 | 163.235 | 173.3785 | 197.6056 | 159.7729 | 353.235 | 372.8764 | 188.765 | 229.7808 | 243.1474 | 303.2152 | 374.6853 | 131.3864 | 130.2271 | 133.6136 | 129.7769 | 175.0717 | 140.5418 | 147.8407 | 96.0797 | 267.996 | 263.239 | 276.757 | 336.9283 | 136.9164 | 131.6056 | 147.3068 | 118.004 | 79.5379 | 84.1514 | 82.765 | 117.5418 | 200.3107 | 203.3825 | 246.3825 | 163.7769 | 134.3904 | 130.765 | 164.5259 | 119.2271 | 158.7091 | 200 | 214.6215 | 161.3825 | 141.3107 | 165.5339 | 198.3745 | 136.2311 | 199.8367 | 208.6893 | 236.8367 | 270.8486 | 173.9881 | 149.3944 | 252.2949 | 231.6893 | 207.7769 | 232.6893 | 258.3028 | 409.0598 | 154.1593 | 144.0079 | 180.6136 | 252.765 | 140.235 | 140.765 | 147.4582 | 146.9322 | 189.2271 | 196.5339 | 194.3785 | 146.6813 | 99.5379 | 103.9243 | 111.5458 | 140.2231 | 224.2231 | 207.3864 | 230.5498 | 359.1554 | 242.0757 | 233.3745 | 237.6056 | 346.6972 | 159.0757 | 72.3864 | 73.3068 | 159.0757 | 122.6893 | 115.9921 | 150.6853 | 121.2231 | 116.8367 | 105.5418 | 124.2271 | 120.8407 | 74.2311 | 88.1474 | 64.4661 | 95.1593 | 105.8526 | 97.004 | 118.2231 | 67.9283 | 237.0916 | 264.761 | 281.765 | 456.004 | 132.6893 | 220.9921 | 197.3864 | 385.5379 | 261.7689 | 198.235 | 211.4661 | 381.7091
1648 | 20.1224 | 198.1372 | 228.1592 | 294.5929 | 227.4867 | 160.6194 | 191.1327 | 198.2964 | 185.2123 | 161.354 | 162.0796 | 129.3275 | 178.9735 | 213.8142 | 164.3806 | 190.5133 | 213.2168 | 188.9469 | 175.7036 | 160.677 | 180.5443 | 174.1107 | 362.7832 | 435.6149 | 177.7566 | 254.1327 | 231.8673 | 267.4071 | 355.2964 | 132.1061 | 124.1638 | 125.1061 | 147.2964 | 162.4867 | 148.2699 | 145.9249 | 102.7566 | 257.5133 | 282.5398 | 255.2479 | 340.5087 | 127.7877 | 114.5443 | 143.2964 | 127.5133 | 80.9204 | 77.5133 | 72.2434 | 124.0531 | 203.376 | 202.1372 | 248.0576 | 180.323 | 143.646 | 123.6505 | 130.8097 | 114.3806 | 189.8052 | 191.2388 | 237.301 | 157.2168 | 149.7301 | 151.323 | 177.3806 | 140.6505 | 181.354 | 200.0265 | 219.0841 | 273.8362 | 144.7301 | 169.1327 | 229.4913 | 219.6194 | 221.646 | 233.0311 | 241.1327 | 369.4071 | 166.5398 | 163.2699 | 175.0265 | 239.323 | 153.4337 | 135.354 | 138.7566 | 167.0796 | 181.2168 | 185.7301 | 182.4071 | 125.8142 | 101.6505 | 108.7036 | 126.3495 | 115.4026 | 213.5178 | 213.9469 | 256.0796 | 361.0531 | 239.4867 | 208.2434 | 221.5178 | 359.7301 | 157.2168 | 75.2964 | 72.4602 | 157.2168 | 116.2168 | 101.1107 | 136.6505 | 109.7877 | 98.354 | 114.9735 | 110.1327 | 113.8142 | 72.323 | 74.677 | 77.4602 | 111.677 | 119.8097 | 103.3495 | 103.624 | 81.2434 | 270.3495 | 246.9204 | 283.1683 | 455.5353 | 131.0841 | 201.9735 | 206.1372 | 390.8142 | 264.8939 | 203.8893 | 220.323 | 421.323
1649 | 20.1334 | 207.1651 | 214.8575 | 280.9321 | 225.1651 | 150.5792 | 172.6765 | 194.2783 | 170.4593 | 172.1199 | 150.9547 | 146.6991 | 177.3982 | 234.7443 | 176.2104 | 191.5566 | 232.0838 | 196.0226 | 175.8869 | 174.6245 | 202.9774 | 184.3755 | 357.7217 | 393.233 | 183.3914 | 237.4661 | 240.0973 | 284.1358 | 344.6312 | 124.0905 | 131.9547 | 121.6923 | 144.5566 | 159.6538 | 143.1878 | 170.4728 | 99.9547 | 252.1651 | 268.0294 | 287.3235 | 310.4887 | 145.6086 | 140.0453 | 140.5566 | 126.767 | 86.6765 | 72.4208 | 71.2104 | 119.2783 | 190.0226 | 215.5113 | 259.5181 | 170.7217 | 128.2783 | 136.2783 | 121.9547 | 129.2783 | 161.6471 | 167.8349 | 253.8642 | 170.715 | 152 | 149.3914 | 191.767 | 147.3982 | 196.2104 | 195.2104 | 231.8642 | 269.1131 | 142.9095 | 158.0905 | 249.5407 | 212.3914 | 201.6765 | 251.8642 | 225.233 | 391.2489 | 161.233 | 155.3755 | 170.2104 | 224.8642 | 152.052 | 148.1651 | 140.8122 | 159.0226 | 179.1199 | 188 | 200.9254 | 142.9095 | 108.9095 | 115.7896 | 117.5566 | 102.3462 | 241.7896 | 224.3462 | 250.3235 | 361.9027 | 235.1199 | 211.8122 | 242.3755 | 353.5633 | 154.0973 | 68.4661 | 75.9774 | 154.0973 | 117.9547 | 118.5339 | 147.4887 | 132.9774 | 111.4208 | 114.1651 | 99.3462 | 132.6991 | 62.2104 | 77.3755 | 82.767 | 116.9321 | 114.2783 | 97.6245 | 114.1651 | 81.7443 | 264.3688 | 257.7896 | 315.1199 | 433.7376 | 146.1651 | 208.0679 | 216.6991 | 403.5633 | 269.3303 | 191.8349 | 209.1878 | 410.6991
1650 | 20.1444 | 192.7846 | 224.7128 | 273.8756 | 206.7654 | 159.3923 | 162.4641 | 194.7846 | 194.2487 | 157.1436 | 146.6077 | 138.3923 | 160.7846 | 228.641 | 175.8231 | 189.6218 | 272.2487 | 196.5885 | 160.4641 | 183.4975 | 201.8756 | 167.7321 | 362.0384 | 393.4115 | 193.1577 | 241.3064 | 225.359 | 294.4975 | 334.359 | 137.8756 | 121.7128 | 145.9282 | 148.2487 | 148.8756 | 134.7321 | 139.091 | 92.3923 | 240.7321 | 241.4115 | 297.8038 | 315.6987 | 129.1436 | 142.5167 | 154.6987 | 129.5359 | 84.2346 | 65.7513 | 73.2346 | 125.6795 | 194.0718 | 217.8038 | 241.0577 | 164.5885 | 134.6795 | 145.6269 | 123.7705 | 132.1962 | 179.3923 | 175.2487 | 234.3923 | 188.55 | 150.3923 | 164.7846 | 194 | 145.7897 | 191.2679 | 195.6269 | 227.3974 | 286.6795 | 135.8231 | 165.9808 | 231.1102 | 226.9808 | 198.6987 | 232.6603 | 235.3256 | 404.7077 | 158.1962 | 144.6269 | 165.5359 | 215.3064 | 132.1628 | 142.359 | 150.0718 | 166.8231 | 163.8756 | 190.4115 | 215.9616 | 128.3205 | 100.4833 | 117.7846 | 120.7128 | 114.7321 | 236.2821 | 235.1244 | 264.0192 | 369.1385 | 215.5885 | 222.9474 | 228.4115 | 335.4308 | 136.6795 | 65.0718 | 75.1436 | 136.6795 | 114.1436 | 124.6077 | 143.5885 | 129.7321 | 99.1244 | 111.0384 | 114.1436 | 124.3923 | 64.5025 | 65.0192 | 75.3539 | 104.7846 | 110.7654 | 110.7846 | 107.2872 | 81.8038 | 271.4257 | 261.3923 | 299.8756 | 466.0718 | 137.6795 | 215.3923 | 210.5359 | 386.7705 | 242.7128 | 202.1962 | 208.2346 | 412.0384
1651 | 20.1554 | 189.6048 | 222.8084 | 268.9114 | 188.121 | 162.2862 | 164.6431 | 190.7641 | 181.7581 | 157.9557 | 153.8024 | 139.3245 | 155.9233 | 214.5605 | 185.9174 | 166.2419 | 261.7198 | 184.9557 | 156.7581 | 155.4335 | 212.6048 | 162.2359 | 385.121 | 409.8084 | 164.2359 | 248.5869 | 244.7641 | 264.7521 | 340.5928 | 139.6371 | 113.6431 | 140.7581 | 133.5162 | 149.7964 | 136.2419 | 153 | 93.3245 | 236.6371 | 244.0767 | 297.7581 | 326.1533 | 135 | 134.121 | 166.8347 | 132.4012 | 96.7138 | 74.879 | 90.1976 | 129.9174 | 197.5605 | 221.9617 | 261.6695 | 163.0443 | 144.8024 | 160.4395 | 148.3629 | 131.121 | 171.6371 | 166.9617 | 235.6048 | 164.4838 | 147.121 | 159.6431 | 190.0767 | 171.2743 | 191.1593 | 208.7581 | 263.1593 | 285.0324 | 153.7641 | 154.7198 | 249.5545 | 217.6814 | 214.4778 | 224.7964 | 266.7581 | 361.8084 | 157.4012 | 155.5162 | 173.4455 | 241.3629 | 148.4395 | 150.5545 | 152.7198 | 176.9174 | 174.0443 | 194.1976 | 192.5988 | 123.2419 | 109.4395 | 109.2802 | 108.1593 | 114 | 212.6048 | 228.5988 | 255.6635 | 348.8527 | 210.9617 | 227.0767 | 233.879 | 352.1593 | 144.8407 | 72.4838 | 77.6371 | 144.8407 | 120.5605 | 127.3186 | 142.6048 | 129 | 103.5281 | 124.0324 | 121.9617 | 123.6431 | 76.0324 | 72.9174 | 61.3305 | 90.6755 | 95.7641 | 101.4395 | 115.0767 | 82.3186 | 250.8024 | 260.9233 | 312.8467 | 456.9497 | 140.5162 | 201.1916 | 212.8407 | 400.7138 | 241.3688 | 198.879 | 218.1916 | 423.9114
1652 | 20.1664 | 199.612 | 239.2241 | 264.0622 | 195.0954 | 174.5643 | 173.7075 | 191.2988 | 175.7075 | 156.4087 | 152.8092 | 158.8506 | 168.9253 | 212.4896 | 175.3672 | 175.5104 | 252.8092 | 167.9046 | 158.0207 | 148.8713 | 219.9793 | 153.8029 | 378.639 | 420.081 | 156.973 | 219.2386 | 251.7345 | 241.5851 | 321.0477 | 130.7075 | 124.7552 | 133.5373 | 123.7552 | 139.2925 | 148.4357 | 153 | 103.7822 | 233.2655 | 236.9253 | 285.2718 | 305.7075 | 149.0414 | 135.5373 | 144.5643 | 136.2925 | 80.3402 | 72.5851 | 81.2448 | 122.8776 | 193.1494 | 216.4689 | 234.388 | 171.612 | 146.1494 | 157.8299 | 151.4419 | 134.5851 | 160.0747 | 173.1701 | 245.9046 | 180.6805 | 147.3672 | 162.5643 | 183.8029 | 157.2178 | 189 | 201.2448 | 269.7759 | 262.5851 | 163.3672 | 155.1701 | 231.5373 | 215.2718 | 205.3672 | 215.4627 | 255.7345 | 389.3402 | 160.4149 | 148.6805 | 195.2655 | 248.2448 | 142.612 | 131.9523 | 149.3672 | 170.1701 | 186.4149 | 183.4896 | 189.8776 | 123.7345 | 106.5373 | 109.4149 | 110.388 | 118.6805 | 224.3672 | 218.5643 | 222.9667 | 382.612 | 220.0954 | 212.3195 | 229.8299 | 348.2448 | 147.5851 | 76.1017 | 74.2655 | 147.5851 | 119.0747 | 125.0477 | 154.6598 | 126.9523 | 118.9316 | 102.4627 | 122.3195 | 126.8569 | 57.973 | 62.9523 | 86.8569 | 81.973 | 106.5373 | 100.8776 | 105.5851 | 75.9523 | 260.9253 | 263.3942 | 328.027 | 435.946 | 125.7822 | 185.0684 | 211.7822 | 386.0954 | 264.7822 | 199.8029 | 195.0477 | 388.054
1653 | 20.1774 | 191.1337 | 230.4977 | 297.3204 | 199.2084 | 157.7807 | 180.3157 | 179.4603 | 173.4759 | 171.5023 | 138.4868 | 159.5879 | 173.8663 | 210.0964 | 170.2193 | 182.1711 | 237.8771 | 168.0591 | 164.3313 | 172.9891 | 203 | 164.2193 | 358.1337 | 390.0964 | 176.0109 | 242.7807 | 238.6952 | 258.5459 | 340.9362 | 125.7325 | 116.5023 | 118.112 | 124.9518 | 136.3421 | 156.3795 | 153.6096 | 93.5614 | 246.3048 | 243.6952 | 277.535 | 301.0373 | 168.6734 | 131.3904 | 132.9627 | 131.818 | 83.2566 | 72.1711 | 77.9145 | 124.3904 | 186.0109 | 195.8771 | 233.7216 | 162.8289 | 134.4386 | 159.2675 | 138.6096 | 137.2193 | 156.0482 | 179.0482 | 234.4759 | 185.2939 | 147.0964 | 154.3157 | 191.7807 | 162.5614 | 186.5614 | 205.8398 | 272.1011 | 258.5132 | 157.6096 | 161.0482 | 238.6687 | 206.6205 | 206.0109 | 220.8289 | 240.5614 | 383.4168 | 158.6952 | 160.9145 | 189.1011 | 237.5988 | 135.353 | 136.4495 | 147.5723 | 169.9518 | 182.8663 | 174.3904 | 192.9145 | 112.2193 | 101.4759 | 104.0373 | 120.6952 | 123.9036 | 217.3904 | 205.7434 | 261.8663 | 370.1711 | 227.2566 | 202.2193 | 228.2193 | 349.7916 | 149.6096 | 59.4386 | 86.3904 | 149.6096 | 115.6579 | 122.0747 | 141.0747 | 130.535 | 98.8289 | 109.6579 | 113.7434 | 112.3048 | 72.7434 | 58.6096 | 73.2193 | 97.353 | 106.6579 | 102.3904 | 100.5988 | 74.9627 | 254.2831 | 244.3904 | 308.6843 | 449.0264 | 125.6205 | 206.6952 | 207.353 | 403.3157 | 258.8289 | 205.9518 | 203.6579 | 365.2193
1654 | 20.1884 | 191.5604 | 244.0957 | 275.0754 | 190.1225 | 161.0251 | 188.438 | 210.0235 | 168.4144 | 144.9027 | 147.5118 | 135.3422 | 164.8775 | 221.146 | 177.1225 | 179.9027 | 251.2198 | 194.8053 | 149.5134 | 159.854 | 209.9765 | 161.292 | 364.2684 | 406.5369 | 186.292 | 231.438 | 238.6342 | 293.3658 | 371.1947 | 119.1711 | 100.3674 | 101.7818 | 135.1241 | 127.3658 | 140.9027 | 160.3909 | 93.3909 | 245.0973 | 244.5856 | 293.1947 | 294.2198 | 136.6342 | 132.5369 | 142.6342 | 120.9513 | 69.7064 | 66.0973 | 74.292 | 119.5118 | 194.0722 | 208.5856 | 216.8791 | 164.146 | 136.5118 | 165.5118 | 142.2198 | 139.0487 | 165.854 | 181.5604 | 234.8053 | 168.4144 | 150.8524 | 161.8037 | 194.0738 | 151.6091 | 179.0973 | 222.5118 | 235.3422 | 254.8791 | 152.0235 | 166.0973 | 264.9027 | 217.3406 | 208.9984 | 224.0487 | 234.3658 | 384.9059 | 162.4396 | 159.5118 | 162.4898 | 219.1947 | 144.5856 | 151.9749 | 158.9513 | 162.0487 | 180.5369 | 178.0738 | 195.3171 | 117.854 | 93.5604 | 96.2684 | 121.5856 | 117.6107 | 225.5134 | 218.9765 | 248.438 | 375.8305 | 228.9278 | 211.0251 | 236.0738 | 355.3893 | 150.0487 | 71.3422 | 87.2198 | 150.0487 | 119.8775 | 105.2684 | 127.4396 | 136.3642 | 109.0251 | 115.146 | 116.8289 | 115.854 | 72.3406 | 64.7567 | 77.2684 | 98.0235 | 110.5604 | 101.3171 | 89.4882 | 83.0487 | 234.0989 | 247.7567 | 299.2936 | 436.3674 | 141.0973 | 202.1947 | 214.3658 | 406.0471 | 262.0487 | 196.7802 | 216.7567 | 369.5856
1655 | 20.1994 | 181.0118 | 215.6234 | 290.283 | 202.294 | 171.6942 | 176.5881 | 201.2587 | 159.0707 | 156.9057 | 141.9764 | 153.6116 | 165.5527 | 216.3176 | 182.3766 | 194.5409 | 249.0825 | 195.7531 | 174.941 | 167.647 | 217.4237 | 149.6588 | 361.7295 | 413.647 | 174.9882 | 217.6116 | 233.4119 | 291.0118 | 359.0472 | 126.9882 | 124.0354 | 131.2587 | 162.7767 | 133.9057 | 139.7295 | 167.3884 | 100.7178 | 247.2587 | 254.2587 | 288.9528 | 300.9764 | 138.6588 | 135.0354 | 149.2705 | 126.294 | 64.2115 | 60.6824 | 65.6234 | 115.294 | 186.8821 | 215.294 | 229.4119 | 154.706 | 130.9764 | 151.0825 | 152.5999 | 131.7295 | 168.706 | 177.2705 | 247.6116 | 165 | 132.9528 | 144.3058 | 209.941 | 140 | 181.2587 | 213.0236 | 235.4945 | 270.0472 | 136.0118 | 170.2351 | 256.1533 | 207.2822 | 189.8939 | 219.6942 | 221.8003 | 435.1179 | 167.3884 | 157.5999 | 196.353 | 230.7642 | 148 | 138.6706 | 162.3176 | 156.3766 | 179.4119 | 181.0943 | 201.9293 | 127.6234 | 88.941 | 102.294 | 117.4237 | 132.353 | 230.8475 | 231.0354 | 233.294 | 373.283 | 244.353 | 221.6942 | 240.7413 | 347.2351 | 146.0236 | 77.0943 | 83.4355 | 146.0236 | 121.8703 | 111.9528 | 131.0707 | 123.4702 | 122.3294 | 112.9528 | 113.294 | 118.3766 | 68.8703 | 69.4001 | 69.7885 | 86.294 | 105.941 | 106.6116 | 107.5409 | 81.9882 | 255.3648 | 249.1061 | 314.0118 | 473.2115 | 136.6706 | 194.9882 | 214.647 | 376.3294 | 261.647 | 191.6706 | 221.7295 | 376.9528
1656 | 20.2104 | 179 | 221.2915 | 247.8088 | 209 | 167.6834 | 186.2414 | 209.8997 | 156.5674 | 165.4827 | 139.8495 | 155.8244 | 187.7336 | 220.3668 | 166.1159 | 207.4576 | 235.3417 | 177.5831 | 179.8746 | 167.5831 | 198.9248 | 154.0752 | 345.5329 | 419.3919 | 174.6081 | 219.8244 | 218.6332 | 286.2664 | 351 | 126.9498 | 123.8088 | 139.558 | 149.8088 | 149.3166 | 138.2257 | 153.558 | 90.9749 | 261.7085 | 254.0156 | 278.5485 | 305 | 138.2915 | 127.2915 | 151.4576 | 134.0251 | 81.5831 | 63.8088 | 71.9749 | 114.4827 | 204.2664 | 220.9749 | 214.2915 | 162.0095 | 132.9498 | 141.7837 | 160.9248 | 121.3417 | 167.1254 | 179.1159 | 245.7241 | 165.6834 | 138.2664 | 157.2007 | 216.2414 | 139.6583 | 193.3166 | 210.3668 | 221.6678 | 256.8746 | 136.7336 | 176.0658 | 246.4013 | 221.1254 | 204.5831 | 212.6081 | 206.859 | 410.558 | 165.1756 | 169.3417 | 181.0407 | 269.9749 | 144.2414 | 139.0251 | 160.8997 | 144.8997 | 175.9091 | 176.9593 | 214.6834 | 130.8997 | 96.6081 | 104.558 | 108.4922 | 122.5078 | 207.1756 | 215.4326 | 238.6332 | 327.7336 | 231.0909 | 215.9749 | 221.1661 | 350.6739 | 148.8339 | 64.7586 | 68.7336 | 148.8339 | 132.7241 | 119.3166 | 124.8088 | 137.5234 | 118.558 | 112.4576 | 121.3668 | 111.6834 | 81.0909 | 61.7586 | 63.6929 | 90.6081 | 111.2163 | 104.0407 | 122.8495 | 82.2915 | 252.3919 | 234.7586 | 297.3071 | 475.8902 | 145.5674 | 197.6834 | 220.0502 | 381.442 | 259.5329 | 203.6427 | 216.1254 | 388.7586
1657 | 20.2214 | 186.7877 | 222.2478 | 258.646 | 219.2656 | 165.4601 | 172.9822 | 201.646 | 177.1239 | 154.1859 | 135.0797 | 142 | 184.8583 | 219.4601 | 152.8938 | 193.8319 | 233.8761 | 176.3098 | 156.5486 | 161 | 205.7435 | 156.1062 | 347.7435 | 423.292 | 171.062 | 210.6018 | 211.7522 | 290.9116 | 352.7699 | 120.8761 | 126.8583 | 134.8938 | 158.1681 | 151.5399 | 145.3804 | 143.5841 | 92.8938 | 257.5663 | 240.6815 | 251.1417 | 307.8319 | 131.8141 | 125.7699 | 138.8938 | 132.4601 | 76.4159 | 76.062 | 72.1239 | 106.7257 | 199.354 | 216.8761 | 217.7257 | 179.7964 | 132.1859 | 143.9203 | 146.8583 | 121.646 | 173.1062 | 164.1239 | 219.7522 | 170.8938 | 139.3717 | 168.0442 | 196.6105 | 139.354 | 187.3982 | 215.1239 | 239.1504 | 248.062 | 146.9558 | 155 | 265.2565 | 221.4424 | 195.8761 | 217.9116 | 214.5663 | 398.8141 | 167.3185 | 165.0442 | 162.9558 | 258.4424 | 142.6638 | 145.2478 | 156.5399 | 141.6018 | 181.9025 | 196.4601 | 214.5841 | 119.8141 | 102.6196 | 98.4779 | 116.7522 | 114.2478 | 220.646 | 207.2656 | 243.0797 | 342.5576 | 213.062 | 217.1859 | 222.8674 | 326.4159 | 165.5399 | 64.5663 | 76.4779 | 165.5399 | 114.5399 | 114.4601 | 135.646 | 115.2214 | 115.6638 | 98.8319 | 125.4159 | 111.2301 | 66.9558 | 73.9558 | 83.8583 | 89.8938 | 106.2478 | 88.4337 | 109.5841 | 75.8141 | 257.354 | 235.9822 | 267.938 | 455.1594 | 155.1504 | 192.2743 | 209.4866 | 388.2301 | 257.4956 | 212.0707 | 221.7522 | 396.708
1658 | 20.2324 | 198.8023 | 221.2094 | 254.3372 | 232.1396 | 155.7035 | 171.8895 | 198.4361 | 174.0407 | 166.593 | 146.1686 | 148.9593 | 173.2965 | 215.9302 | 166.2267 | 191.3082 | 227.8023 | 187.8314 | 139.6918 | 160.6337 | 212.8838 | 149.7326 | 363.6745 | 429.3255 | 175.9302 | 213.5059 | 202.1686 | 306.8023 | 352.3372 | 121.029 | 113.4651 | 137.6047 | 148.2384 | 155.8023 | 127.8314 | 142.4651 | 100.7326 | 254.622 | 253.5756 | 271.1279 | 302.7443 | 128.9302 | 130.8314 | 141.2384 | 116.8431 | 79.3663 | 75.0698 | 73.4361 | 122.6628 | 201.8314 | 211.1686 | 239.8895 | 171.4244 | 138.7326 | 132.4651 | 139.6918 | 124.2965 | 157.9419 | 163.2384 | 222.2557 | 168.4768 | 148.0698 | 155.3372 | 177.2965 | 138.1686 | 177.1686 | 210.5756 | 246.0754 | 247.8023 | 144.2792 | 160.1279 | 234.3082 | 212.4244 | 190.5349 | 230.1396 | 226.6393 | 383.8431 | 137.2094 | 157.4651 | 176.3953 | 241 | 146.7733 | 151.1686 | 165.5639 | 146.3372 | 165.3255 | 190 | 203.5756 | 133.0464 | 115.407 | 103 | 104.6047 | 114.3082 | 224.029 | 224.1686 | 239.5176 | 351.5756 | 220.8604 | 218.2384 | 246.7035 | 330.8314 | 159.9129 | 56.4941 | 76.2384 | 159.9129 | 117.7035 | 112.3953 | 145.2557 | 133.4768 | 116.1105 | 103.6337 | 121.0407 | 119.7208 | 71.2384 | 77.1396 | 72.6628 | 92.2384 | 112.9012 | 96.1396 | 99.3082 | 73.6628 | 251.7733 | 238.5522 | 274.7906 | 482.4651 | 141.5639 | 177.4361 | 193.2557 | 387.5639 | 260.1803 | 194.7208 | 216.9302 | 388.1105
1659 | 20.2434 | 192.7286 | 215.4642 | 248.7571 | 214.4287 | 140.5358 | 181.0501 | 197.7857 | 160.4857 | 160.9645 | 145.2714 | 151.1572 | 173.7001 | 217.9715 | 181.9214 | 174.2143 | 224.3786 | 186.8358 | 148.8358 | 169.4642 | 187.2714 | 152.8928 | 367.2784 | 431.7787 | 176.8358 | 208.1642 | 212.2499 | 277.2574 | 356.2213 | 120.4287 | 114.4857 | 135.6785 | 144.1642 | 146.3215 | 129.1072 | 144.6214 | 102 | 268.4572 | 254.1427 | 280 | 284.2429 | 137.7857 | 124.5358 | 131.1072 | 112.357 | 79.2429 | 73.0285 | 66.7286 | 124.8358 | 209.9214 | 212.1642 | 261.5428 | 176.5358 | 132.4287 | 141.8142 | 155.2714 | 126.5929 | 145.5143 | 167.4927 | 229.0215 | 151.6214 | 144.1357 | 154.6785 | 190.5713 | 127.0501 | 181.5713 | 196 | 260.636 | 259.5213 | 140.2784 | 162.1858 | 236.1427 | 219.0501 | 195.9499 | 217.7286 | 266.0786 | 369.1357 | 131.8428 | 163.7857 | 178.3215 | 228.8858 | 131.4572 | 145.7286 | 162.0501 | 137.35 | 175.35 | 186.5929 | 204.5213 | 156.1572 | 106.65 | 101.1072 | 103.4356 | 104.4071 | 217.3716 | 223.2714 | 219.7071 | 355.1712 | 226.8358 | 217.1928 | 241 | 337.0285 | 144.6785 | 61.8358 | 66.1072 | 144.6785 | 111.2429 | 115.4572 | 154.6715 | 120.7857 | 102.4071 | 101.4857 | 107.8643 | 139.6214 | 69.0571 | 72.2999 | 80.8928 | 95.7356 | 111.3786 | 88.2714 | 90.9214 | 82.2714 | 236.0786 | 245.343 | 286.2143 | 476.6715 | 144.1072 | 177.9214 | 198.8858 | 396.9214 | 255.0355 | 200.993 | 218.5929 | 372.5143
1660 | 20.2544 | 197.208 | 230.2183 | 248.0457 | 203.5634 | 129.2994 | 172.6903 | 197.3554 | 162.6903 | 146.68 | 152.5177 | 141.2537 | 174.7715 | 210.264 | 180.2285 | 173.0812 | 225.7817 | 173.7463 | 151.7715 | 177.1828 | 204.2892 | 150.1371 | 390.8731 | 415.7817 | 175.8629 | 211.4823 | 232.8274 | 257.6698 | 355.8022 | 115.4263 | 113.1726 | 134.4468 | 147.0914 | 144.0355 | 131.0812 | 161.1978 | 100.0457 | 255.528 | 266.8834 | 281.9543 | 293.1623 | 133.4468 | 114.8629 | 131.5177 | 142.7258 | 83.472 | 72.9188 | 64.9543 | 127.3806 | 214.0914 | 218.6091 | 264.7006 | 161 | 120.7817 | 136.9748 | 150.0103 | 125.2994 | 149.5634 | 185.6548 | 216.0812 | 157.6446 | 139.7463 | 162.0457 | 192.4468 | 131.5886 | 177.5839 | 200.6903 | 219.7817 | 286.1726 | 148.6297 | 155.6903 | 241.8479 | 204.4366 | 207.4366 | 206.9645 | 247.4468 | 382.7258 | 160.5075 | 166.0914 | 176.8172 | 225.4161 | 125.5634 | 139.264 | 152.5177 | 143.1623 | 173.3452 | 187.2537 | 222.9645 | 143.9086 | 116.8629 | 101.1269 | 105.0103 | 119.7715 | 198.5177 | 223.4823 | 240.2537 | 366.6297 | 220.7817 | 220.0103 | 239.8274 | 348.6446 | 151.6548 | 61.6446 | 68.0812 | 151.6548 | 114.2994 | 103.3097 | 138.6548 | 123.0914 | 109.5634 | 97.8274 | 107.5634 | 121.0205 | 78.6091 | 73.5737 | 81.6548 | 108.3554 | 104.9645 | 91.2183 | 98.2183 | 83.264 | 228.3909 | 223.472 | 287.0355 | 464.9543 | 137.0914 | 184.4366 | 193.0709 | 407.736 | 267.7565 | 185.8172 | 204.792 | 379.6903
1661 | 20.2654 | 195.4866 | 228.1937 | 263.5449 | 199.5496 | 142.8535 | 175.5654 | 180.9528 | 163.5496 | 171.2252 | 154.7748 | 153.822 | 175.0835 | 206.8063 | 169.4346 | 183.0157 | 238.6913 | 176.9008 | 162.9685 | 158.9528 | 216.937 | 147.4504 | 385.5969 | 411.7591 | 185 | 218.4976 | 221.7276 | 268.811 | 337.8693 | 122.1622 | 113.3874 | 134.126 | 137.7748 | 141.6961 | 132.1465 | 180.3402 | 95.3874 | 237.7276 | 252.4346 | 293.8693 | 304.5654 | 119.822 | 125.6126 | 145.063 | 147.6961 | 85.9528 | 62.9843 | 66.7906 | 130.7118 | 203.1622 | 206.7118 | 251.1465 | 158.178 | 119.9843 | 123.6913 | 138.0157 | 120.3087 | 147.5654 | 177.1622 | 221.5811 | 175.2567 | 135.6441 | 164.2409 | 175.5969 | 160.6126 | 164.0945 | 205.9843 | 219.7906 | 277.115 | 130.4819 | 150.5024 | 243.4457 | 206.8378 | 198.9528 | 199.2252 | 237.0472 | 396.9685 | 166.6961 | 155.9685 | 182.9528 | 235.2772 | 121.9528 | 136.2094 | 150.3402 | 144.8898 | 177 | 185.7118 | 208.7748 | 145.1622 | 118.3402 | 103.5811 | 98.2567 | 128.1465 | 205.2094 | 222.4346 | 236.6961 | 356.5449 | 212.3244 | 211.2409 | 240.822 | 362.6283 | 146.7906 | 69.178 | 73.5811 | 146.7906 | 116.9685 | 98.822 | 135 | 118.6126 | 107.5654 | 92.7748 | 110.8063 | 102.2724 | 75.178 | 65.2252 | 69.5339 | 98.4031 | 105.2882 | 90.4031 | 99.822 | 71.7433 | 237.8693 | 229.9843 | 289.1307 | 478.885 | 131 | 182 | 208.9528 | 405.9528 | 248.8378 | 200.0157 | 203.2882 | 391.0315
1662 | 20.2764 | 184.386 | 210.7978 | 265.2427 | 198.7243 | 150.5074 | 173.9853 | 167.8456 | 150.6765 | 171.0147 | 147.0919 | 155.9228 | 161.3088 | 210.9081 | 177.2463 | 182.261 | 251.0147 | 203.8934 | 158.2463 | 156.647 | 206.2463 | 160.3235 | 367.9669 | 409.8162 | 182.9228 | 228.2757 | 212.5698 | 242.9706 | 323.9964 | 124.1397 | 115.5698 | 143.0441 | 126.7684 | 134.8787 | 120.3382 | 159.8603 | 91.3382 | 227.3235 | 258.1691 | 293.2905 | 303.4007 | 126.0772 | 130.9081 | 157 | 128.8309 | 78.6618 | 61.2463 | 66.2463 | 117.739 | 194.3382 | 201.2169 | 236.8456 | 166.8787 | 127.386 | 140.5846 | 148.4779 | 110.0625 | 155.125 | 166.261 | 224.2316 | 186 | 149.4007 | 152.8897 | 176.6618 | 146.3824 | 179.9228 | 193.0294 | 223.8162 | 251.8603 | 145.5846 | 153.1875 | 270.0589 | 218.9853 | 201.6323 | 203.5846 | 254.7243 | 377.7059 | 158.6323 | 147.0919 | 171.0919 | 217.8162 | 122.1397 | 135.0919 | 137.7537 | 138.9559 | 178.2463 | 169.8309 | 201.9228 | 126.3677 | 99.1066 | 100.4154 | 111.0772 | 119.6618 | 203.6765 | 225.6765 | 217.8309 | 384.4154 | 216.6948 | 216.0919 | 237.9375 | 359.2757 | 149.9853 | 68.3235 | 69.5846 | 149.9853 | 111.4154 | 100.9228 | 128.7684 | 125.9853 | 103.4926 | 97.1397 | 121.1397 | 118.4926 | 62.6912 | 68.7537 | 63.647 | 102.739 | 123.2463 | 91.8309 | 104.8309 | 67.2316 | 250.5846 | 223.261 | 272.6765 | 478.7978 | 135.9853 | 189.8934 | 208.7243 | 392.4302 | 265.1397 | 197.5993 | 215.0147 | 414.0331
1663 | 20.2874 | 196.4336 | 208.5311 | 234.5605 | 206.7227 | 144.0059 | 182.2832 | 176.1151 | 164.3925 | 159.7227 | 145.1387 | 148.7227 | 181.1269 | 215.8555 | 183.2773 | 179.1387 | 237.584 | 195.7521 | 151.4454 | 160.8672 | 198.5899 | 171.416 | 350.416 | 419.7109 | 182.1387 | 218.5664 | 219.8555 | 260.4219 | 297.121 | 124.1681 | 110.879 | 122.8555 | 128.2656 | 147.8672 | 129.121 | 149.9941 | 94.5605 | 237.9882 | 245.0235 | 272.2891 | 314.4219 | 129.8555 | 126.4454 | 148.0176 | 124.0118 | 83.9941 | 65.5664 | 68 | 120.8613 | 196.2773 | 207.5899 | 233.1387 | 179.4395 | 132.5899 | 142.5664 | 155.1504 | 118.2891 | 168.584 | 173.4042 | 249.2538 | 179.584 | 145.4512 | 146.1034 | 183.705 | 124.7109 | 179.9941 | 176.8613 | 234.9941 | 239.8555 | 147.9941 | 163.3243 | 230.0235 | 220.0118 | 207.7344 | 200.4336 | 257.1622 | 370.5488 | 161.3126 | 149.8437 | 164.4336 | 225.5723 | 140.1328 | 131 | 138.5664 | 162.5664 | 185.9882 | 176.9882 | 188.7344 | 114.705 | 99.121 | 103.9941 | 118.2773 | 124.5664 | 214.3983 | 215.8672 | 216.4336 | 366.6075 | 232.7344 | 227.8261 | 222.0118 | 339.7286 | 153.5781 | 74.7109 | 70.7109 | 153.5781 | 102.5899 | 95.4336 | 120.8555 | 133.8555 | 105.2891 | 96.7403 | 127.584 | 124.5664 | 62.9765 | 66.5723 | 73 | 113.9882 | 119.8672 | 102.8378 | 104.2891 | 81.1328 | 236.7403 | 226.1269 | 272.2773 | 455.1445 | 134.8731 | 190.3067 | 215.0118 | 384.2891 | 274.5781 | 188.2891 | 214.8437 | 413.6193
1664 | 20.2984 | 191.2399 | 224.1194 | 257.4008 | 199.6799 | 141.3201 | 183.12 | 190.5998 | 171.3991 | 154.88 | 141.3997 | 147.4002 | 184.0796 | 196.3189 | 180.1995 | 179.3599 | 230.7601 | 161.6396 | 156.7209 | 154.44 | 187.3201 | 160.199 | 348.4395 | 418.0398 | 178.3997 | 210.5594 | 215.2797 | 253.5594 | 318.1603 | 105.6799 | 110.0807 | 128.4002 | 139.7999 | 135.2797 | 143.56 | 155.3201 | 95.3997 | 259.6407 | 229.88 | 283.2007 | 288.6384 | 122.6396 | 127.7607 | 143.0404 | 120.8402 | 84.9198 | 67.2399 | 69.3201 | 110.9596 | 189.2393 | 198.0802 | 224.9994 | 172.44 | 127.4804 | 146.88 | 147.7999 | 129.6407 | 171.8806 | 182.5196 | 241.9585 | 163.9596 | 126.9198 | 156.2791 | 181.1995 | 137.1205 | 172.9994 | 183.6805 | 231.0796 | 227.7993 | 153.7601 | 141.8402 | 226.3206 | 213.3201 | 202.3604 | 195.6799 | 244 | 386.88 | 149.6003 | 151.0796 | 152.6396 | 226.3201 | 148.7601 | 138.0404 | 138.0398 | 160.2797 | 186.5196 | 188.9602 | 183.3604 | 128.4804 | 107.8397 | 104.4798 | 120.0398 | 119.6396 | 223.5594 | 219.5605 | 234.1211 | 347.7203 | 212.8397 | 238.4395 | 226.3206 | 333.9602 | 143.2797 | 72.1598 | 75.6401 | 143.2797 | 102.3206 | 95.5202 | 125.0802 | 131.0398 | 101.6799 | 96.1609 | 120.7601 | 121.8397 | 76.7999 | 68.6401 | 72.56 | 109.2393 | 111.2399 | 99.7192 | 103.7601 | 79.6396 | 232.2007 | 215.4389 | 270.9596 | 475.5611 | 124 | 192.2808 | 210.5202 | 396.9608 | 271.3201 | 199.2007 | 219.1598 | 373.0796
1665 | 20.3094 | 184.9292 | 219.4032 | 269.3107 | 197.0954 | 144.8093 | 176.5722 | 188.1662 | 155.3324 | 151.4769 | 127.1199 | 150.2861 | 173.5477 | 181.7602 | 169.4523 | 176 | 256.9726 | 163.2125 | 163.0736 | 164.9509 | 196.237 | 141.7847 | 329.1199 | 416.6185 | 172.7139 | 209.5694 | 214.7847 | 250.3078 | 341.3787 | 111.237 | 124.3815 | 126.3107 | 133.8338 | 131.6185 | 147.0708 | 152.9292 | 86.0954 | 253.932 | 232.8093 | 271.0519 | 277.1388 | 115.6185 | 150.3324 | 144.763 | 128.8093 | 69.2398 | 65.9046 | 75.0708 | 96.5722 | 183.4032 | 200.1907 | 216.4278 | 173.4523 | 136.5477 | 136.6922 | 131.8829 | 124.8366 | 178.8584 | 168.0245 | 207.7139 | 164.0463 | 124.3569 | 141.3324 | 174.9755 | 150.4523 | 173.474 | 191.9292 | 219.6431 | 220.9509 | 160.9755 | 142.5722 | 236.5722 | 205.9537 | 212.0954 | 196.2616 | 240.8338 | 381.6676 | 160.7139 | 146.4278 | 164.6157 | 232.5231 | 148.7384 | 153.3324 | 135.2616 | 147.5722 | 158.0028 | 184.0491 | 198.5231 | 143.8801 | 94.5722 | 102.7139 | 117.2616 | 122.5694 | 209.9046 | 221.6922 | 243.3598 | 355.9046 | 219.9264 | 219.5722 | 239.7384 | 329.0491 | 151.3787 | 71.8801 | 72.2153 | 151.3787 | 112.1199 | 111.7602 | 134.8801 | 124.1907 | 98.6431 | 119.8093 | 119.3815 | 109.0245 | 76.2616 | 68.8338 | 73.8093 | 94.8093 | 102.6676 | 88.0463 | 103.7384 | 69.9046 | 254.8801 | 204.2833 | 264.2616 | 487.169 | 122.1907 | 211.1907 | 200.526 | 410.1199 | 283.4032 | 207.406 | 209.3815 | 390.1388
1666 | 20.3204 | 180 | 238.9691 | 248.8495 | 196.4646 | 148.8584 | 162.1017 | 189.2124 | 160.677 | 153.8983 | 122.323 | 143.7478 | 156.2742 | 197.8584 | 171.3938 | 182.0399 | 253.0131 | 178.9292 | 136.6371 | 177 | 200.8186 | 147.677 | 340.5842 | 401.4866 | 179.7168 | 208.345 | 219.7478 | 268.646 | 353.9601 | 122.7876 | 122.7876 | 121.1814 | 139.2921 | 133.2124 | 145.9601 | 145.2124 | 85.4646 | 235.7567 | 229.8893 | 255.0488 | 297.4557 | 124.646 | 145.4557 | 146.6859 | 131.4646 | 70.7966 | 66.0708 | 80 | 103.9381 | 199.7168 | 196.1416 | 216.0309 | 167.0309 | 139.2522 | 126.7168 | 132.7258 | 106.3629 | 164.8186 | 167.1107 | 204.9601 | 165.7079 | 130.1814 | 149.9292 | 175.8893 | 147.2832 | 177.0619 | 190.2522 | 211.9601 | 246.4735 | 152.5973 | 144.8275 | 242.5443 | 192.677 | 212.3938 | 190.6371 | 230.0309 | 376.7876 | 154.2433 | 150.677 | 174.345 | 221.2742 | 149.2522 | 150.3141 | 146.8275 | 140.0708 | 153.5842 | 183.6151 | 203.5354 | 151 | 93.5753 | 102.2832 | 121.8584 | 129.2433 | 216.5753 | 220.0797 | 233.1505 | 363.9691 | 231.4955 | 212.0708 | 233.2832 | 336.978 | 163.4955 | 73.8893 | 65.3938 | 163.4955 | 113.3629 | 110.2034 | 132.2433 | 120.1416 | 102.5753 | 109.4557 | 114.0708 | 113.6859 | 74.354 | 67.323 | 76 | 98.3938 | 98.7168 | 96.677 | 107.0399 | 71.9292 | 250.3849 | 213.7787 | 271.1814 | 477.7567 | 121.3938 | 192.7388 | 182.5354 | 403 | 290.8893 | 201.1814 | 203.1416 | 427.1814
1667 | 20.3314 | 178.5693 | 231.7393 | 245.5379 | 181.2624 | 153.3845 | 162.061 | 191.2459 | 165.1534 | 152.9852 | 136.4455 | 139.5231 | 153.6914 | 199.5231 | 174.9076 | 191.3845 | 226.9224 | 182.8614 | 140.83 | 176.0462 | 199.8152 | 143.0924 | 346.2624 | 376.6155 | 166.3549 | 208.5065 | 211.231 | 253.4935 | 343.6617 | 117.8928 | 126 | 126.9538 | 135.6931 | 134.2921 | 130.4159 | 143.9076 | 89.8152 | 242.231 | 247.8448 | 271.2772 | 280.6155 | 131.4307 | 127.1848 | 137.5397 | 131.5231 | 82.8004 | 73.1072 | 69.5083 | 118.0924 | 198.2772 | 202.1072 | 217.5379 | 153.7541 | 127.2624 | 129.5693 | 148.1848 | 120.2921 | 163.3383 | 174.4307 | 201.8152 | 154.0462 | 128.8004 | 151.4769 | 163.3235 | 140.6155 | 157.1848 | 191.1386 | 208.3383 | 267.7228 | 144.1072 | 159.2921 | 236.9704 | 204.7838 | 203.0314 | 201.5065 | 217.7079 | 371.4159 | 138.7079 | 140.8466 | 166.8762 | 217.7376 | 140.6007 | 145.1534 | 156.0462 | 145.1996 | 166.8928 | 189.3697 | 202.5231 | 142.4159 | 106.6765 | 99.9076 | 116.3697 | 126.1072 | 225.3845 | 222.0776 | 234.5545 | 375.3383 | 230.1996 | 202.8928 | 225.6617 | 341.4011 | 161.2459 | 70.8614 | 61.7541 | 161.2459 | 120.6155 | 102.4917 | 125.2921 | 114.1848 | 97.5545 | 91.1848 | 107.7541 | 116.939 | 68.0462 | 68.0924 | 78.3845 | 101.4307 | 107.769 | 89.7079 | 111.6155 | 71.5693 | 234.1386 | 217.2607 | 251.2014 | 490.9076 | 127.769 | 172.5693 | 190.5841 | 392.0314 | 281.6617 | 195.5083 | 205.769 | 437.2459
1668 | 20.3424 | 197.0567 | 224.2946 | 248.6622 | 188.4593 | 154.5324 | 159.7703 | 195.0433 | 153.9217 | 143.4676 | 144.5973 | 139.4892 | 169.5541 | 196.5541 | 183.3594 | 189.5973 | 233.1514 | 188.9351 | 153.5541 | 177.9351 | 209.3811 | 144.9135 | 319.7054 | 380.3594 | 147 | 216.3244 | 218.7189 | 247.6324 | 326.792 | 106.5541 | 117.1946 | 130.4459 | 122.9351 | 144.3811 | 125.4027 | 145.0216 | 86.6622 | 224.7703 | 269.5973 | 272.827 | 287.2946 | 135.9351 | 122.5108 | 116.4459 | 125.1298 | 70.1298 | 76.6189 | 66.8054 | 109.6406 | 188.5757 | 192.9001 | 228 | 155.827 | 110.9784 | 135.3378 | 135.1946 | 126.9567 | 164.5541 | 174.5324 | 210.8919 | 170.6107 | 130.8054 | 147.1081 | 150.1298 | 135.5541 | 148.5973 | 196.8054 | 214.9351 | 252.4758 | 145.173 | 160.5757 | 237.9918 | 216.1081 | 188.5541 | 212.2595 | 224.7405 | 387.9269 | 138.8919 | 144.7405 | 174.6406 | 231.0865 | 140.208 | 141.2595 | 151.5757 | 166.6756 | 165.3378 | 182.9567 | 197.5973 | 140.3378 | 110.5757 | 102 | 108.4892 | 138.4243 | 227.0216 | 206.0649 | 211.2379 | 368.2379 | 236.0216 | 198.4027 | 228.8486 | 324.9135 | 167.9784 | 68.6189 | 64.8054 | 167.9784 | 113.5973 | 117.9135 | 137.827 | 119.7837 | 83.0433 | 95.3162 | 113.2513 | 107 | 70.4243 | 71.8702 | 76.5973 | 106.9135 | 99.792 | 84.0216 | 106.0649 | 71.9567 | 228.5541 | 229.1081 | 235.2513 | 485.173 | 129.0865 | 164.1514 | 196.0865 | 380.9784 | 271.1514 | 184.9784 | 212.4676 | 433.2163
1669 | 20.3534 | 189.9172 | 224.4719 | 227.4601 | 208.4837 | 164.0325 | 151.5133 | 197.5133 | 143.5281 | 166.0621 | 138.4956 | 135.4867 | 159.4778 | 196.5074 | 180.9734 | 180.9882 | 241.5458 | 178.4601 | 147.9911 | 179.4956 | 209.4837 | 143.4837 | 353.1271 | 374.9645 | 154.0207 | 202.5044 | 225.4808 | 266.5044 | 333.0591 | 108.5133 | 120.0355 | 123.4719 | 131.5163 | 144.4837 | 140.531 | 142.4956 | 85.0177 | 221.0266 | 248.4512 | 260.9409 | 292.4867 | 126.4631 | 123.003 | 128.0266 | 110.4749 | 67.0089 | 73.5074 | 74.4956 | 100.4926 | 182.4926 | 198.5724 | 221.4808 | 159.9882 | 117.0148 | 139.9911 | 136.531 | 118.469 | 160.997 | 178.0148 | 209.4808 | 169.9438 | 142.0059 | 136.9852 | 159.5458 | 135.5074 | 152.0236 | 192.9616 | 224.5192 | 245.5458 | 139.5074 | 147.4719 | 241.9409 | 221.0296 | 182.4897 | 201.5044 | 234.4926 | 402.4631 | 149.5163 | 149.9793 | 170.5074 | 249.5665 | 144.4719 | 134.5163 | 150.5074 | 157.9291 | 169.9911 | 186.5044 | 196.0089 | 142.4837 | 107.5015 | 95.4808 | 108.4985 | 134.9793 | 231.5163 | 206.0089 | 212.0355 | 356.4985 | 237.0059 | 193.9734 | 232.4897 | 352.068 | 158.469 | 61.4956 | 77.0089 | 158.469 | 112.0089 | 115.4808 | 127.4453 | 131.003 | 86.5163 | 97.4808 | 105.4424 | 121.5429 | 63.9704 | 74.4897 | 65.4808 | 108.9941 | 90.0118 | 89.0177 | 103 | 76.5074 | 229.0089 | 210.4601 | 252.5163 | 467.9734 | 133.5251 | 171.0414 | 197.5163 | 394.5369 | 254.9734 | 188.5074 | 216.5074 | 400.4335
1670 | 20.3644 | 153.7797 | 218.5964 | 245.8537 | 213.7891 | 153.9355 | 152.4036 | 191.7246 | 169.9544 | 156.1739 | 129.8072 | 131 | 161.2478 | 199.5138 | 170.9725 | 179.0551 | 249.2935 | 178.8718 | 141.9174 | 156.4217 | 193.7246 | 135.4312 | 383.1558 | 360.4312 | 165.1102 | 208.1102 | 216.9449 | 246.4217 | 352.4862 | 110.4312 | 122.7522 | 127.358 | 125.1833 | 144.6515 | 150.4862 | 137.4036 | 96.1377 | 218.6971 | 225.321 | 251.7891 | 295.7065 | 110.9174 | 124.5138 | 142.6515 | 107.1377 | 68.9725 | 76 | 80.1928 | 106.7341 | 184.1102 | 211.1833 | 215.5138 | 158.0551 | 123.0275 | 136.4862 | 145.4587 | 123.9268 | 164.1102 | 169.642 | 208.6515 | 165.3855 | 142.9725 | 143.8167 | 170.8898 | 145.7065 | 174.8993 | 182.0551 | 206.8529 | 263.5688 | 137.8898 | 133.8898 | 246.6609 | 228.4312 | 183.6239 | 200.9449 | 237.1377 | 383.321 | 149.3485 | 146.5964 | 188.9268 | 245.7978 | 136.0275 | 132.2935 | 145.8072 | 140.679 | 166.4862 | 189.0275 | 187.1833 | 135.4587 | 105.4312 | 93.1102 | 105.4312 | 122.8623 | 239.5688 | 215.679 | 242.4957 | 354.4587 | 221.0181 | 201.4406 | 227.9725 | 347.2565 | 151.5964 | 65.6515 | 68.1833 | 151.5964 | 110.3761 | 120.8167 | 113.1102 | 129.4312 | 93.5413 | 98.1928 | 92.1652 | 127.2659 | 58.1102 | 69.9725 | 66.7065 | 95.6971 | 98.1102 | 90.8898 | 104.0275 | 72.321 | 249.9819 | 205.2203 | 252.3485 | 447.6971 | 146.6239 | 181.9174 | 188.6145 | 404.9449 | 269.6334 | 193.0551 | 204.1007 | 388.7891
1671 | 20.3754 | 150.2085 | 218.8436 | 257.0618 | 212.9527 | 137.6824 | 161.6254 | 175.1612 | 173.9006 | 139.6254 | 123.5261 | 144.6776 | 164.7394 | 203.6824 | 166.3176 | 168.9006 | 244.1042 | 187.2655 | 140.5782 | 150.2036 | 188.7345 | 136.6824 | 349.4315 | 355.3697 | 173.7345 | 206.7394 | 223.417 | 224.9479 | 348.8436 | 102.7394 | 114.5261 | 134.2133 | 136.6206 | 141.583 | 138.9527 | 139.7867 | 97.8436 | 236.4072 | 220.0521 | 247.2703 | 298.7915 | 123.7818 | 117.1091 | 144.3176 | 116.2085 | 66.9479 | 79.6824 | 85.9479 | 119.7345 | 187.4739 | 199.4739 | 216 | 158.4218 | 124.5261 | 136 | 132.4267 | 129.5309 | 174.3127 | 154.8958 | 200.3224 | 171.6352 | 146.2085 | 146.583 | 170.6824 | 147.2133 | 173.2182 | 178.2133 | 212.9333 | 274.417 | 146.6254 | 148.9382 | 243.697 | 225.4739 | 198.5212 | 193.2133 | 242.5261 | 369.1091 | 151.8909 | 146.8436 | 172.9625 | 227.3127 | 128.583 | 135.5212 | 133.7394 | 147 | 163.8958 | 186.8436 | 174.4218 | 135.5782 | 106.6824 | 102.7867 | 101.4218 | 120.6303 | 226.2182 | 227.2606 | 229.4885 | 335.64 | 205.0521 | 205.9527 | 232.2606 | 312.057 | 147.1091 | 71.5261 | 68.5733 | 147.1091 | 100.7394 | 124.6352 | 121.2085 | 125.4218 | 86.583 | 99.7394 | 111.6776 | 112.6873 | 64.6303 | 81.6254 | 74 | 88.6824 | 105.6824 | 88.0521 | 101.8436 | 65.4739 | 239.1188 | 211.4218 | 241.2133 | 448.0473 | 141.5309 | 180.5782 | 183.943 | 403.5261 | 267.8012 | 199.2085 | 194.7345 | 399.5261
1672 | 20.3864 | 156.6918 | 209.0014 | 230.848 | 191.6946 | 140.4616 | 157.3878 | 172.7685 | 179.1505 | 156.9219 | 125.6151 | 148.3864 | 170.767 | 202.6932 | 157.6165 | 154.2315 | 227.6961 | 164.1577 | 143.0767 | 160.3082 | 182.7713 | 157.2272 | 336.4602 | 354.6151 | 159.6961 | 196.0781 | 224.0014 | 243.919 | 333.0796 | 104.9986 | 105.848 | 122.0014 | 155.3849 | 148.5355 | 126.3082 | 150.9219 | 106.3054 | 244.2344 | 243.6107 | 248.5355 | 298.2301 | 133.1548 | 115.9219 | 140.4616 | 118.9233 | 67.6151 | 79.2315 | 76.9247 | 117.5398 | 168.1577 | 209.2287 | 206.848 | 169.9204 | 133.6136 | 133.3082 | 134.9204 | 131.2287 | 178.3849 | 169.1505 | 201.9971 | 153.6946 | 154.3068 | 146.5369 | 175.0767 | 148.9986 | 162.7685 | 170.8466 | 221.2344 | 259.3893 | 145.6179 | 143.9276 | 220 | 211.0029 | 203.1548 | 182.6165 | 219.3126 | 375.4588 | 148.233 | 139.6932 | 158.4588 | 239.9986 | 125.3068 | 143.9233 | 138.6903 | 160.9971 | 159.8466 | 187.7685 | 175.1534 | 138.6151 | 111.6151 | 99.9247 | 95.6932 | 123 | 201.233 | 225.0014 | 219.2272 | 353.9161 | 221.6122 | 209.9971 | 243.9986 | 334.686 | 144.8452 | 69.3082 | 75.2315 | 144.8452 | 103.5369 | 127.6903 | 115.3097 | 128.3068 | 91.3821 | 117.0724 | 116.4631 | 111.8452 | 68.6151 | 75.2344 | 80.4602 | 93.6151 | 103.6165 | 91.6918 | 100.6151 | 76.8437 | 218.8452 | 207.3082 | 238.6918 | 464.9986 | 130.3082 | 182.5384 | 198.4602 | 384.6194 | 239.0014 | 215.9204 | 199 | 383.3111
1673 | 20.3974 | 166.709 | 228.8802 | 229.6077 | 207.8802 | 155.9686 | 155.5635 | 186.4622 | 176.4807 | 157.0442 | 130.3039 | 139.8987 | 167.3352 | 194.5949 | 156.3039 | 152.1013 | 215.8545 | 159.0128 | 149.5064 | 161.3039 | 181.709 | 153.2782 | 333.7404 | 362.0571 | 157.2154 | 207.5193 | 223.5064 | 245.0314 | 321.4494 | 108.7974 | 98.5506 | 116 | 144.0884 | 151.0884 | 135.5635 | 143.8859 | 102.2339 | 233.4051 | 233.2653 | 258.7974 | 308.1583 | 138.3609 | 134.7661 | 129.5378 | 122.4051 | 80.5635 | 81.5064 | 76.6077 | 108.6961 | 163.5635 | 195.4237 | 214.418 | 181 | 141.5506 | 132.6519 | 150.4051 | 128.9872 | 163.7846 | 181.3481 | 205.1897 | 140.6961 | 156.3481 | 148.5949 | 166.0884 | 162.1583 | 161.0442 | 177.8103 | 204.8987 | 253.7661 | 144.9116 | 143.7218 | 223.8545 | 193.4936 | 198.4494 | 192.3167 | 211.1141 | 391.9558 | 143.4051 | 154.1712 | 155.1327 | 228.9301 | 130.6519 | 133.6391 | 144.2468 | 147.1198 | 165.1583 | 186.0442 | 181.4051 | 135.0442 | 117.4051 | 94.6519 | 103.0128 | 120.7974 | 209.0699 | 210.1897 | 226.291 | 342.8031 | 224.5378 | 220.3481 | 241.1897 | 346.1327 | 151.7532 | 66.4494 | 77.5064 | 151.7532 | 116.0571 | 122.7846 | 119.6648 | 132 | 105.8545 | 121.1327 | 102.8417 | 110.4936 | 75.5064 | 60.5506 | 83.7974 | 97.2026 | 96.2468 | 93.4494 | 97.0442 | 74.8859 | 227.9558 | 203.3481 | 244.3039 | 468.2468 | 145.0699 | 179.6961 | 199.0442 | 387.8231 | 271.545 | 207.1769 | 199 | 382.7661
1674 | 20.4084 | 161.1784 | 242.6819 | 227.6819 | 205.3568 | 163.9336 | 167.2517 | 177.5491 | 160.1853 | 151.3112 | 131.3112 | 131.1189 | 172.7621 | 199.444 | 164.0664 | 167.6363 | 224.0664 | 170.8147 | 154 | 160.0595 | 193.0664 | 145.3181 | 332.0664 | 354.6155 | 165.7483 | 218.1853 | 211.6749 | 226.3707 | 341.2656 | 97.9931 | 104.6293 | 117.6888 | 137.1259 | 149.1923 | 134.8672 | 129.4965 | 98.3181 | 253.888 | 221.3845 | 262.6293 | 317.3776 | 140.3707 | 137.556 | 123.2517 | 120.3707 | 86.6224 | 82.0595 | 76.9336 | 104.3112 | 178.6293 | 184.8811 | 219.1189 | 174.2448 | 135.8077 | 152.5768 | 142.1784 | 124.5035 | 156.9405 | 189.5699 | 210.9475 | 147.0069 | 143.1784 | 146.1259 | 148.9931 | 149.9861 | 166.5699 | 182.1853 | 205.1259 | 250.9267 | 141.1784 | 164.0664 | 231.5035 | 198.5699 | 205.8811 | 207.9405 | 220.4371 | 376.8603 | 156.5699 | 147.0456 | 159.3251 | 223.444 | 139.3181 | 123.1853 | 150.444 | 136.1328 | 168.7483 | 178.0595 | 188.3776 | 136.6293 | 111.4301 | 103.8811 | 108.6224 | 111.1189 | 211.1784 | 213.6957 | 225.0664 | 328.0139 | 210.9336 | 203.8008 | 238.5035 | 340.1923 | 160.7552 | 73.8811 | 78.6224 | 160.7552 | 105.8008 | 112 | 115.3637 | 122.4301 | 99.9931 | 105.6224 | 106.5699 | 134.1467 | 69.8672 | 68.3181 | 67.9267 | 97.8741 | 98.5035 | 92.4371 | 90.7483 | 68.9405 | 230.3112 | 213.2587 | 244.7483 | 455.8672 | 142.6749 | 178.1259 | 184.8672 | 387.6749 | 278.7939 | 183.6819 | 197.3112 | 394.5629
1675 | 20.4194 | 153.7257 | 216.8672 | 223.1504 | 203.6815 | 149.646 | 167.8496 | 166.3009 | 155.1239 | 152.3009 | 135.7522 | 139.9558 | 178.6725 | 202.5487 | 179.5045 | 170.3716 | 233.177 | 166.6725 | 147.6725 | 157.5752 | 197.5752 | 146.9734 | 320.4689 | 365.1328 | 173.2036 | 209.0973 | 221.4337 | 242.7081 | 343.7698 | 88.6991 | 111.8761 | 114.9734 | 141.4513 | 148.823 | 121.2478 | 120.823 | 104 | 235.3627 | 229.2743 | 263.5487 | 296.991 | 137.1504 | 128.1239 | 128.4513 | 107.3716 | 75.3716 | 76.6991 | 63.2212 | 107.6018 | 187.0266 | 193.8761 | 208.3982 | 148.8672 | 134.4513 | 141.6901 | 137.0266 | 129.1504 | 157.6991 | 169.9645 | 212.9469 | 160.9027 | 139.7522 | 142.9734 | 165.009 | 149.2302 | 164.7964 | 190.354 | 207.7257 | 244.0266 | 138.3275 | 153.0442 | 229.823 | 216.9293 | 201.646 | 191.4424 | 220.5752 | 371.2036 | 153.646 | 143.8054 | 163.9469 | 227.1239 | 151.3275 | 125.0266 | 144.9203 | 138.823 | 177.9293 | 150.8406 | 185.2478 | 145.6018 | 101.6991 | 103.6725 | 122.1063 | 107.8761 | 193.3716 | 219.6991 | 224.3982 | 341.4248 | 220.8054 | 193.1504 | 239.1239 | 352.4779 | 163.1239 | 80.5752 | 65.646 | 163.1239 | 110.1063 | 98.7698 | 117.8054 | 113.2743 | 106.2302 | 114.5045 | 109.3982 | 130.1149 | 68.3275 | 78.0266 | 72.5311 | 105.0531 | 95.0973 | 87.3982 | 96.4779 | 81.2036 | 231.8761 | 213.9469 | 256.8054 | 461.8054 | 123.6725 | 183.0266 | 194.2567 | 410.0887 | 260 | 176.2743 | 193.6991 | 416.2833
1676 | 20.4304 | 172.0379 | 217.8629 | 222.2375 | 197.1995 | 142 | 164.0624 | 165.6499 | 152.4125 | 139.3121 | 139.4125 | 153.9376 | 159.3121 | 221.1506 | 176.2497 | 172.8127 | 232.2999 | 163.7625 | 148.2876 | 145.6622 | 205.6377 | 152.8127 | 326.0379 | 382.3501 | 161.787 | 193.4247 | 235.6499 | 245.0747 | 324.5998 | 103.4504 | 109.8874 | 115.5251 | 127.5496 | 141.4749 | 128.1628 | 121.1126 | 85.1995 | 224.3378 | 224.4749 | 258.6499 | 288.8127 | 139.7625 | 118.3623 | 128.6499 | 105.7001 | 80.7503 | 73.825 | 68.9253 | 108.0624 | 198.8127 | 200.7001 | 206.7625 | 150.4504 | 133.4749 | 123 | 142.3501 | 137.0502 | 170.6879 | 150 | 203.4749 | 161.2999 | 133.4247 | 136.4749 | 181.4125 | 151.3623 | 164.4627 | 195.0624 | 210.7625 | 241.7124 | 156.5129 | 151.4002 | 231.8752 | 209.7246 | 198.1126 | 188.1628 | 212.1873 | 379.7625 | 148.9376 | 146.9498 | 167.9878 | 222.6499 | 144.8372 | 132.7001 | 127.1873 | 148.5129 | 174.8372 | 161.7883 | 176.2999 | 142.5373 | 107.0502 | 99.5875 | 125.7748 | 121.1628 | 203.4504 | 209.7748 | 203.3745 | 352.1628 | 215.1371 | 206.9253 | 236.4125 | 348.5998 | 167.4627 | 72.1873 | 69.1628 | 167.4627 | 127.2876 | 89 | 128.5875 | 120.2252 | 113.6499 | 104.787 | 103.0624 | 112.7625 | 67.1248 | 75.7124 | 79.2999 | 93.9621 | 107.0379 | 99.2754 | 99.0624 | 86.2375 | 221.0747 | 213.2876 | 248.787 | 463.1873 | 121.9376 | 179.5373 | 200.5373 | 395.4614 | 254.1248 | 163.8372 | 189.0624 | 409.6744
1677 | 20.4414 | 172.6037 | 207.407 | 232.3963 | 194.1967 | 145.5988 | 154.8023 | 165.7994 | 157.9981 | 137.3973 | 127.0039 | 154.2006 | 161.5959 | 222.2045 | 162.6017 | 179.5998 | 224.2015 | 172.1977 | 147.8014 | 129.8023 | 193.0058 | 150.6027 | 351.7955 | 337.8149 | 169.3934 | 194.9971 | 217.2054 | 249.7955 | 322.1986 | 114.4002 | 111.1986 | 122.1986 | 122.5979 | 146.1977 | 127.0029 | 124.5998 | 79.7975 | 234.1996 | 244.7926 | 256.4002 | 291.4012 | 130.2035 | 115.3992 | 137.1967 | 119.7965 | 83.6017 | 70.001 | 75.6008 | 108.3992 | 199.6017 | 207.5988 | 210.999 | 152.4031 | 123.8023 | 135.5959 | 146.3992 | 126.405 | 179.8004 | 155.3983 | 193.8023 | 150.8023 | 136.1967 | 125.0029 | 183.3992 | 144.2006 | 166.6008 | 178.0048 | 213.7994 | 235.6008 | 136.6095 | 144.8043 | 204.8101 | 198.4002 | 200.4002 | 202.5979 | 200.0019 | 352.8091 | 149.2006 | 152.1967 | 180.3983 | 228.7975 | 125.0039 | 138.999 | 135.3954 | 141.8052 | 175.9971 | 169.4041 | 180.7975 | 121.2054 | 92.8062 | 100 | 120 | 120.0029 | 214.4002 | 198.6017 | 211.3934 | 379.7936 | 214.1967 | 208.8023 | 221.6046 | 354.5959 | 161.8033 | 63.001 | 66.8033 | 161.8033 | 123.2025 | 107.594 | 131.3992 | 134.3973 | 101.2035 | 93.1996 | 107.5979 | 118.7985 | 76.1957 | 73.7994 | 80.7985 | 96.3954 | 107.0039 | 108.2006 | 102.3983 | 89.7985 | 256.9855 | 200.2054 | 244.9971 | 461.1986 | 132.3973 | 197.7926 | 205.5969 | 382.595 | 257.1977 | 191.3886 | 184.6008 | 413.1938
1678 | 20.4524 | 153.3701 | 203.4057 | 221.6371 | 184.4734 | 149.8363 | 168.3629 | 176.7936 | 162.6121 | 137.4911 | 105.5339 | 143.8185 | 166.3274 | 210.5516 | 162.0605 | 181.8363 | 238.7508 | 181.8968 | 154.1815 | 140.9145 | 192.6299 | 138.879 | 331.6193 | 337.6049 | 184.8363 | 214.4661 | 215.1815 | 226.5588 | 326.8363 | 105.4306 | 126.854 | 131.7331 | 131.2847 | 136.3097 | 137.5266 | 130.5089 | 93.5694 | 233.7758 | 234.2919 | 248.0427 | 279.2064 | 117.4911 | 118.8363 | 134.2064 | 125.1637 | 72.0427 | 68.6121 | 77.6726 | 100.8185 | 176.4129 | 194.0855 | 224.0178 | 153.9573 | 134.3024 | 131.7581 | 135.7581 | 127.6299 | 168.5943 | 177.3629 | 180.4306 | 153.9573 | 141.1637 | 153.2775 | 170.3097 | 138.7153 | 169.2847 | 175.9573 | 207.6548 | 240.121 | 121.2847 | 153.7508 | 220.1137 | 183.9218 | 202.4484 | 213.7331 | 200.2847 | 384.8041 | 137.5943 | 143.0855 | 155.8435 | 233.3879 | 111.4911 | 141 | 148.0605 | 131.6121 | 182 | 161 | 179.879 | 124.6903 | 91.6299 | 100 | 121.2242 | 121.3452 | 213.3879 | 214.5871 | 229.8968 | 369.1282 | 222.8363 | 205.8363 | 219.9573 | 369.7331 | 152.5516 | 70.1815 | 67.9573 | 152.5516 | 113.7153 | 113.879 | 123.2064 | 127.7581 | 102.5694 | 91.5516 | 118.121 | 123.8363 | 73.3701 | 76.8363 | 79.7153 | 91.9218 | 112.4661 | 89.8613 | 101.7153 | 74.025 | 252.7225 | 211.0355 | 237.5339 | 479.9145 | 134.9395 | 188.5161 | 195.4734 | 396.6726 | 253.4306 | 189.904 | 194.6299 | 393.5588
1679 | 20.4634 | 164.1073 | 193.1415 | 192.5171 | 184.6146 | 146.6293 | 180.6244 | 176.7561 | 157.3805 | 155.2292 | 116.9804 | 132.3805 | 160.2537 | 210.8732 | 157.7561 | 173.0098 | 233.1415 | 169.3903 | 147.5122 | 152.2488 | 187.5122 | 143.1171 | 314.2488 | 354.5024 | 182.2537 | 206.7659 | 204.1415 | 212.8683 | 334.2439 | 93.1317 | 123.761 | 131.6293 | 131.5024 | 135.1171 | 151.7463 | 143.9902 | 87.1366 | 223.0098 | 209.2586 | 246.1219 | 262.3854 | 117.7463 | 114.3805 | 142.3609 | 133.3658 | 74.4927 | 63.3805 | 78.7512 | 97.4976 | 181.4829 | 198.3609 | 226.0049 | 159 | 140.2537 | 121.5024 | 136.1171 | 120.0147 | 170.7414 | 167.1464 | 195.6048 | 167.7414 | 136.8781 | 149.4001 | 158.5024 | 131.6293 | 170.7512 | 170.3854 | 202.3756 | 233.3854 | 135.239 | 136.2781 | 215.5269 | 189.3609 | 187.1415 | 208.0098 | 210.4927 | 403.8878 | 127.2537 | 146.1122 | 150.4878 | 218.6391 | 127.9804 | 136.0049 | 132.5171 | 138.8683 | 173.2586 | 162.2488 | 198.478 | 130.8781 | 95.8781 | 106.8683 | 113.2586 | 122.2537 | 217.9951 | 202.6488 | 242.9902 | 346.5073 | 224 | 215.7414 | 217.5073 | 354.0196 | 147.8781 | 62.2635 | 68.0049 | 147.8781 | 124.7365 | 115.6195 | 125.7414 | 133.1122 | 103.6293 | 103.7365 | 121.3756 | 126.2488 | 64.7512 | 60.5171 | 75.7512 | 89.2439 | 117.2537 | 87.7414 | 94.6293 | 68.8683 | 222.2586 | 234.3658 | 212.7659 | 462.5269 | 139.8683 | 164.8829 | 187.4976 | 409.6146 | 254.2439 | 174.6244 | 188.2635 | 385.4878
1680 | 20.4744 | 171.1799 | 186.8201 | 219.5641 | 201.8236 | 160.5537 | 166.9931 | 165.9965 | 173.1004 | 168.6367 | 130.91 | 138.5502 | 171.3703 | 212.6367 | 155.91 | 167 | 231.2768 | 156.8166 | 145.0934 | 149.8166 | 192.7336 | 150.5467 | 323.9135 | 331.9896 | 178.7266 | 209.7336 | 188.9066 | 207.4498 | 338 | 92.1834 | 126.5502 | 129 | 131.91 | 131.7232 | 138.7197 | 142.3598 | 87.6402 | 229.0969 | 227.5571 | 230.8097 | 267.4602 | 114.2699 | 119.2768 | 136.3564 | 133.2699 | 76.4533 | 70.8236 | 62.7197 | 100.91 | 182.4498 | 207.6367 | 221.09 | 167.9135 | 129.7232 | 120 | 128.9031 | 117.3668 | 156.2629 | 158.4568 | 213.0934 | 172.3633 | 140.7301 | 139.0934 | 165.9135 | 132.1834 | 170.6367 | 158.2699 | 194.9965 | 222.5432 | 138.1799 | 147.6506 | 202.1834 | 194.1799 | 201.8305 | 181.6263 | 200.3564 | 372.4429 | 159.3807 | 141.2664 | 150.3598 | 201.7232 | 126.6297 | 140.6402 | 140.4637 | 143 | 152.7197 | 155.3598 | 197.3564 | 112.4463 | 109.917 | 112.91 | 112.4568 | 106.6297 | 215.2699 | 216.6506 | 239.4498 | 347.7301 | 218.2699 | 218.4533 | 230.827 | 324.173 | 160.0069 | 64.8236 | 67.5467 | 160.0069 | 125.3598 | 113.2699 | 121.4498 | 119.9896 | 99.7266 | 118.3668 | 110.8131 | 114.9031 | 63.3633 | 65.917 | 66.7232 | 89.8166 | 94.6263 | 98.7301 | 93.91 | 80.0035 | 220.8201 | 234.9066 | 213.8236 | 453.6402 | 129.3564 | 165.0934 | 190.2734 | 406.4498 | 260.5467 | 187.0969 | 188.2768 | 368.1695
1681 | 20.4854 | 162.5103 | 189 | 241 | 194.3716 | 158.3185 | 167.4366 | 166.5428 | 161.286 | 151.4778 | 137.8407 | 138.8082 | 168.6165 | 200.0206 | 140.1268 | 161.8082 | 224.3185 | 177.714 | 147.351 | 133.7226 | 200.649 | 149.4041 | 340.0325 | 350.5016 | 176.7021 | 204.0206 | 211.3099 | 240.2893 | 320.4778 | 102.4366 | 109.286 | 121.8613 | 130.4041 | 127.649 | 140.3835 | 152.9263 | 90.7021 | 210.0412 | 242.2979 | 250.2037 | 275.9469 | 122.6815 | 119.4572 | 113.7226 | 125.4572 | 63.9675 | 79.5959 | 64.3835 | 110.4366 | 177 | 186.584 | 229.7346 | 169.1062 | 125 | 120 | 118.1062 | 117.7551 | 159.2774 | 162.9469 | 182.9025 | 168.1062 | 138.8082 | 151.0856 | 159.9675 | 141.1387 | 176.8407 | 173.8201 | 198.1387 | 245.3099 | 140.5428 | 155.5634 | 222.1712 | 187.4572 | 201.7226 | 168.0531 | 199.7876 | 363.5428 | 157.584 | 141.1387 | 147.2979 | 216.4691 | 133.2774 | 143.7021 | 135.4247 | 139.7551 | 151.7876 | 163.9794 | 188.351 | 117.9263 | 103.4247 | 100.3716 | 96.8288 | 109.3835 | 202.9144 | 223.9144 | 238.5428 | 356.8407 | 215 | 223.4897 | 228.6696 | 329.5753 | 155.0206 | 72.9469 | 66.4041 | 155.0206 | 109.3185 | 118.4366 | 128.3304 | 115.7876 | 99.649 | 110.3185 | 109.5428 | 125.5222 | 60.4041 | 63.9675 | 71.7346 | 99.6815 | 100.5222 | 90.3185 | 109.9263 | 84 | 218.4572 | 222.2654 | 205.7226 | 472.2774 | 143.065 | 178.3835 | 180.6165 | 395.1593 | 249.6696 | 194.649 | 202.0856 | 397.481
1682 | 20.4964 | 149.7423 | 194.9514 | 219.1784 | 188.3225 | 152 | 173.3225 | 154.4522 | 156.9351 | 131.4198 | 129.0973 | 140.6288 | 160.355 | 196.9676 | 145.5478 | 165.6126 | 243.1279 | 168.1784 | 143.6937 | 134.5964 | 190.4198 | 155.9351 | 332.1135 | 351.791 | 170.7099 | 216.8378 | 217.7261 | 235.8721 | 328.1928 | 96.4198 | 102.2901 | 123.9514 | 134.9514 | 127.3387 | 133.436 | 151.7423 | 83.3874 | 201.9514 | 233.0811 | 249.4685 | 285.5964 | 128.3387 | 128.9027 | 129.1442 | 138.209 | 67.9189 | 83.645 | 68.6775 | 110.3712 | 178.3225 | 173.6775 | 229.0486 | 138.5657 | 115.0811 | 112.7261 | 126.5802 | 126.5802 | 172.9514 | 167.9676 | 194.0955 | 167.9838 | 128.7261 | 138.8072 | 176.4829 | 158.2252 | 162.1459 | 170.7748 | 194.0649 | 261.645 | 135.7261 | 159.5964 | 230.6775 | 198.8865 | 190.6937 | 186.8378 | 206.645 | 395.7568 | 140.0486 | 139.7099 | 141.3874 | 237.5802 | 133.0649 | 152.2577 | 128.9838 | 142.6288 | 158.645 | 177.6126 | 191.3063 | 123.355 | 106.2414 | 105.564 | 106.8378 | 105.0811 | 201.3063 | 202.4685 | 235.0486 | 369.9189 | 195.8234 | 216.4198 | 232.5802 | 325.436 | 156.5964 | 68.0486 | 63.6775 | 156.5964 | 106.9676 | 115.0649 | 124.4198 | 127.9351 | 100.6613 | 94.7423 | 119.2739 | 123.0973 | 64.2901 | 72.5478 | 72.3712 | 98.7261 | 91.4847 | 94.5802 | 119.9838 | 81.355 | 230.5478 | 220.3063 | 213.8703 | 460.8234 | 129.2108 | 181.355 | 169.7099 | 417.1279 | 240.355 | 196.3225 | 193.7748 | 386.6144
1683 | 20.5074 | 152.4091 | 192.6116 | 218.7768 | 199.7768 | 156.7149 | 167.938 | 148.3471 | 155.6116 | 142.8388 | 137.8182 | 123.467 | 173.8182 | 205.7355 | 160.4091 | 148.1199 | 237.1405 | 163.062 | 144.0207 | 139 | 199.8182 | 155.9587 | 316.9174 | 336.5909 | 169.3471 | 215.5496 | 207.938 | 225.0207 | 348.4504 | 109.1859 | 107.0207 | 105.4463 | 152.8182 | 133.062 | 135.0826 | 133.5289 | 81.3471 | 205.6736 | 239.4504 | 246.3884 | 288.6529 | 112.5083 | 121.5289 | 128.2025 | 125.7934 | 72.3264 | 70.1818 | 57.8967 | 102.6116 | 191.7975 | 174.3471 | 219.938 | 151.6157 | 133.5743 | 107.6529 | 137.3884 | 129.3058 | 148.3843 | 168.6529 | 195.5289 | 168.3264 | 120.9587 | 134.7149 | 179.2438 | 152.8761 | 146.2645 | 173.4297 | 192.0207 | 239.4257 | 125.2645 | 162.6529 | 225.9587 | 199.2645 | 176.8761 | 182.8554 | 207.3264 | 422.4504 | 147.7768 | 140.3678 | 158.8801 | 238.2851 | 137.0826 | 139.4876 | 137.4091 | 164.533 | 171.4504 | 170.2232 | 191.6529 | 120.6529 | 102.5703 | 104.5909 | 105.5496 | 108.0826 | 207.7149 | 204.7768 | 215.8347 | 360.1818 | 198.1239 | 205.6116 | 214.4257 | 300.8141 | 151.5703 | 68.3678 | 68.3884 | 151.5703 | 115.7355 | 107.6322 | 108.2232 | 140.7562 | 94.938 | 105.4917 | 109.5289 | 114.9793 | 64.3058 | 63.1612 | 76.7355 | 92.9793 | 99.5124 | 96.6322 | 111.5703 | 71.2438 | 223.8554 | 212.5703 | 221.3264 | 483.3305 | 128.1239 | 169.2232 | 187.8801 | 376.1158 | 224.8554 | 197.6736 | 192.3884 | 388.533
1684 | 20.5184 | 145.0265 | 205.0886 | 225.3717 | 189.2256 | 141.8539 | 165.6858 | 166.1461 | 177.6904 | 142.7699 | 142.2566 | 138.0045 | 173.4557 | 209 | 157.1416 | 157.2036 | 217.0265 | 159.1416 | 133.3407 | 147.2301 | 181.6238 | 149.8849 | 328.0886 | 332.3142 | 163.8274 | 196.969 | 202.2566 | 221.1991 | 332.7388 | 115.2566 | 111.4292 | 110.7744 | 145.5973 | 130.5132 | 126.6548 | 122.1991 | 80.6283 | 200.5132 | 234.7123 | 235.969 | 259.8804 | 128.3187 | 117.0575 | 116.5132 | 111.1991 | 79.5443 | 68.4868 | 57.8009 | 113.031 | 191.8274 | 183.2301 | 227.2877 | 160.8849 | 132.6548 | 121.4027 | 131.084 | 118.3982 | 150.0886 | 149.4822 | 202.031 | 175.5443 | 132.031 | 126.0265 | 164.0265 | 140.1416 | 153.9735 | 183.4868 | 196.4292 | 242.4027 | 128.1726 | 146.2256 | 220.5708 | 209.7169 | 181.9735 | 177.3717 | 225.5178 | 412.9114 | 150.2566 | 151.6018 | 167.6283 | 214.053 | 149.916 | 135.8009 | 138.9425 | 174.6858 | 170.8274 | 176.6593 | 194.4292 | 120 | 103.4868 | 106.4868 | 93.1416 | 115.4292 | 202.4557 | 199.0265 | 219.6593 | 342.0265 | 200.5708 | 203 | 198.8849 | 326.978 | 162.7744 | 69.3142 | 66.8849 | 162.7744 | 105.969 | 106 | 108.4868 | 120.9955 | 102.9735 | 99.2831 | 109.1726 | 120.8584 | 62.3142 | 72.1461 | 72.4557 | 97.4868 | 93.5973 | 89.5132 | 105.6283 | 76.6018 | 220.4292 | 210.0575 | 224.4292 | 477.053 | 129.8849 | 179.0886 | 198.6858 | 395.9515 | 263.9515 | 192.5132 | 195 | 409.6593
1685 | 20.5294 | 151.1702 | 200.1317 | 229.4907 | 195.9627 | 135.3963 | 166.6981 | 163.528 | 175.7354 | 134.8112 | 133.3205 | 153.1888 | 176.2832 | 220.8683 | 157.4907 | 148.5466 | 214.7925 | 162.9813 | 149.6422 | 156.5851 | 183.9627 | 156.3776 | 315.4522 | 336.1888 | 152.6224 | 191 | 221.944 | 216.9056 | 334.1702 | 115.3963 | 131.1515 | 107.528 | 133.4149 | 136.3776 | 126.5851 | 130.472 | 79.3019 | 204.9813 | 225.8112 | 227.9056 | 256.0758 | 135.5093 | 112.4149 | 118.1888 | 122.2646 | 77.4149 | 73.0944 | 69.0758 | 114.1131 | 175.0373 | 181.4149 | 211.056 | 147.1317 | 126.3019 | 119.6224 | 121.4149 | 118.8869 | 159.0944 | 154.9627 | 198.9242 | 176.9056 | 124.7354 | 127.9813 | 167.3776 | 141.8869 | 147.1317 | 172.7354 | 185.4336 | 237.8298 | 129.6037 | 139 | 233.6608 | 202.0373 | 191.1888 | 194.7552 | 200.4895 | 380.169 | 153.8869 | 158.0944 | 156.528 | 229.1329 | 132.3578 | 132.4149 | 131.7168 | 149.169 | 159.6224 | 159.6597 | 179.2448 | 127.6795 | 108.7925 | 94.3392 | 103.2646 | 117 | 208.0758 | 207.2646 | 223.6037 | 342.5851 | 189.9242 | 194.6224 | 222.831 | 317.8671 | 164.4149 | 76.6795 | 63.6037 | 164.4149 | 116.7552 | 110.1888 | 119.3776 | 129.5478 | 105.9056 | 99.9813 | 105.7168 | 127.4907 | 70.3776 | 86.9813 | 68.3019 | 100.6981 | 93.9813 | 94.6795 | 98.0187 | 78.9056 | 224.7925 | 205.4149 | 227.3963 | 472.5851 | 135.6795 | 167.1504 | 187.8298 | 399.5466 | 243.1119 | 196.2832 | 211.7552 | 408.7168
1686 | 20.5404 | 146.054 | 194.2896 | 213.2394 | 176.4247 | 153.7606 | 154.9228 | 166.8147 | 160.054 | 159.2858 | 124.3166 | 149.3166 | 162.6602 | 215.3436 | 158.2896 | 149.5251 | 223.1043 | 175.946 | 150.4749 | 150.4749 | 181.4749 | 151.4749 | 331.1545 | 345.1043 | 158.2355 | 207.3398 | 216.7915 | 220.2626 | 335.4479 | 106.054 | 114.8455 | 117.9189 | 131 | 126.5019 | 126.1583 | 124.3436 | 71.1853 | 200.8957 | 221.8687 | 234.1043 | 269.2355 | 127.6062 | 107.1583 | 124.973 | 124.4479 | 72.8687 | 69.7374 | 73 | 111.2896 | 189.6023 | 182.5521 | 235.39 | 164.7336 | 113.9228 | 136.6023 | 120.4209 | 122.4209 | 144.3706 | 170.2355 | 206.3668 | 165.3436 | 145.2858 | 144.4981 | 146.1351 | 148.2626 | 164.0232 | 164.8687 | 188.5251 | 255.7336 | 124.7374 | 133.3166 | 230.054 | 208.7877 | 175.9498 | 196.3166 | 195.2355 | 377.5251 | 159.5521 | 145.5019 | 135.6602 | 223.6602 | 127.2626 | 127.1583 | 131.1313 | 132.3166 | 151.027 | 164.9189 | 162.054 | 130.2896 | 113.5521 | 97.2355 | 98.3436 | 121.973 | 231.1815 | 211.5791 | 216.6062 | 349.2626 | 203.7606 | 200.946 | 242.5251 | 322.6294 | 164.1313 | 68.6332 | 65.8417 | 164.1313 | 116.1853 | 99.9228 | 108.7915 | 123.7915 | 95.7645 | 96.6062 | 88.0811 | 129 | 69.027 | 75.0811 | 64.4479 | 87.5019 | 98.4209 | 81.6602 | 98.5521 | 82.973 | 241.6294 | 199.4479 | 229.4209 | 465.7645 | 134.027 | 170.3668 | 182.2896 | 379.2123 | 235.6564 | 198.2896 | 223.973 | 398.8957
1687 | 20.5514 | 165.1269 | 197.6136 | 206 | 186.9588 | 155.941 | 158.6726 | 166.3864 | 151.6637 | 147.5958 | 127.7817 | 141.941 | 166.118 | 208.8318 | 142.8229 | 167.4543 | 213.7138 | 166.9911 | 142.6637 | 155.6726 | 181.6136 | 132.8229 | 305.5869 | 332.1002 | 159.8318 | 218.3363 | 214.6136 | 227.059 | 316.6548 | 114.2862 | 108.6136 | 124 | 143.2862 | 139.7906 | 119.2183 | 117.1091 | 89.6815 | 207.3953 | 207.9911 | 237 | 296.1269 | 132.2272 | 120.4543 | 132.7817 | 117.941 | 75.6136 | 76.6726 | 67.2183 | 111.7227 | 196.5546 | 195.5635 | 217.2005 | 169.6637 | 122.7316 | 131.9911 | 124.6136 | 100.5958 | 169.7995 | 150.8731 | 200.882 | 168.9499 | 127.814 | 155.059 | 152.6225 | 142.7728 | 165.0501 | 182.0679 | 195.6136 | 253.4365 | 128.7817 | 144.7316 | 239.0089 | 227.0089 | 184.1771 | 212.7906 | 199 | 386.059 | 148.7138 | 150.118 | 134.7817 | 205.4365 | 125.3864 | 118.7728 | 134.1682 | 144.4543 | 174.2951 | 165.941 | 163.7817 | 132.1682 | 101.2684 | 111.8408 | 107.7316 | 105.9321 | 212.2595 | 226.8998 | 224.118 | 332.9321 | 219.6726 | 201.3864 | 229.3775 | 318.882 | 167.1682 | 69.7817 | 71.3363 | 167.1682 | 100.7138 | 105.118 | 123.2361 | 115.1091 | 84.7728 | 87.4955 | 100.0679 | 118.1592 | 74.2272 | 53.8229 | 73.8408 | 84.8909 | 101.1682 | 84.3953 | 99.2773 | 93.6726 | 233.5457 | 213.8998 | 209.0412 | 459.1091 | 130.5546 | 160.5457 | 170.4365 | 389.1771 | 245.059 | 193.6637 | 207.2094 | 388.0501
1688 | 20.5624 | 159.2998 | 196.795 | 219.2301 | 196.47 | 150.325 | 160.53 | 160.59 | 163.2301 | 155.9052 | 127.06 | 137.06 | 156.0348 | 213.1451 | 145.8201 | 166.3849 | 219.2901 | 159.06 | 143.205 | 158.265 | 196.9652 | 142.4352 | 319.1953 | 325.265 | 153.1199 | 214.1199 | 196.8898 | 226.795 | 305.59 | 122.675 | 116.6151 | 110.0348 | 145.795 | 127.8898 | 145.6653 | 113.06 | 104.6151 | 214.675 | 242.6904 | 229.6499 | 284.4198 | 118.0948 | 123.06 | 126.1799 | 116 | 88.7601 | 65.2998 | 63.53 | 109.06 | 175.4198 | 190.4449 | 198.3849 | 164.325 | 118.4449 | 134.3501 | 126.735 | 114.0502 | 159.2148 | 149.3501 | 189.6499 | 158.7699 | 138.3152 | 140.8298 | 156.06 | 138.53 | 148.7699 | 177.2399 | 207.2901 | 260.7601 | 132.47 | 136.7699 | 213.1297 | 200.3947 | 193.675 | 184.7196 | 205.6151 | 382.1199 | 141.06 | 149.59 | 153.1702 | 243.6305 | 125.47 | 125.5551 | 147.4951 | 161.0251 | 169.2998 | 150.7699 | 177.0251 | 121.9749 | 96 | 118.205 | 107.1199 | 113.7002 | 204.205 | 215.3598 | 197.1297 | 367.1604 | 220.06 | 216.1702 | 195.8047 | 318.675 | 159.9149 | 78.6151 | 82.5551 | 159.9149 | 122.4603 | 100.1799 | 128.06 | 114 | 87.8801 | 90.1451 | 101.1199 | 124.2901 | 77.735 | 52.41 | 80.205 | 104.3752 | 97.59 | 93.1451 | 99 | 72.7447 | 233.1451 | 211.9149 | 208.3501 | 456.53 | 132.94 | 155 | 182.9052 | 411.1702 | 261.7002 | 195.675 | 187.5049 | 399.7002
1689 | 20.5734 | 167.4514 | 205.7149 | 206.0648 | 184.2959 | 150.4946 | 170.4622 | 161.9892 | 151.5594 | 146.3067 | 123.7581 | 145.7149 | 157.7257 | 221.7257 | 162.4514 | 164.7473 | 223.7473 | 143.054 | 148.4838 | 146.0432 | 200.5054 | 142.0216 | 320.0324 | 335.4622 | 166.6933 | 188.8337 | 214.6609 | 240.1987 | 295.0324 | 128.4838 | 116.0108 | 102.7581 | 127.0648 | 132.2095 | 127.6026 | 154.596 | 98.7797 | 245.892 | 206.1837 | 235.9676 | 270.2743 | 126.1987 | 137.6933 | 118.5162 | 130.946 | 83.2851 | 62.9892 | 80.1879 | 115.473 | 177.7149 | 188.4946 | 221.4082 | 163 | 122.473 | 128.0324 | 128.4946 | 111.5378 | 163.7041 | 151.2527 | 194.473 | 161.473 | 128.0648 | 127.527 | 152.0108 | 145.473 | 144 | 181.9676 | 181.8553 | 233.6134 | 142.7149 | 137.2311 | 225.1663 | 202.4514 | 191.2635 | 173.4946 | 207.2527 | 384.4838 | 148.2203 | 134.5486 | 156.7581 | 234.3391 | 133.473 | 117.7905 | 153.4946 | 159.7689 | 161.7581 | 172.1555 | 163.8121 | 123.9784 | 101.2311 | 96.581 | 96.0324 | 107.0432 | 206.4946 | 214.2311 | 211.4082 | 382 | 210.7797 | 216.0216 | 206.9244 | 328.2203 | 159.2419 | 58.581 | 70.3067 | 159.2419 | 106.5918 | 97 | 118.0324 | 119.9784 | 84.0216 | 99.473 | 111.7041 | 116.0432 | 79.4946 | 81.6501 | 72.7797 | 104.2743 | 97.4946 | 89.7689 | 107.2203 | 70.7257 | 230.5162 | 183.5918 | 230.4298 | 474.6825 | 132.5054 | 185.6393 | 186.0108 | 415.5054 | 260.2743 | 190.2743 | 193.4622 | 381.0864
1690 | 20.5844 | 158.3274 | 180.1356 | 212.913 | 183.798 | 140.3657 | 147.4143 | 180.4706 | 155.8747 | 150.8747 | 126.798 | 129.2507 | 156.9616 | 195.1356 | 156.3657 | 158.9233 | 232.3555 | 169.9028 | 152.2788 | 149.5959 | 184.0486 | 142.2788 | 295.7315 | 307.8568 | 156.0486 | 212.9028 | 210.087 | 247.2788 | 294.2788 | 111.0102 | 118.798 | 117.1918 | 120.2404 | 131.4425 | 133.9514 | 140.1356 | 117.2685 | 218.0205 | 211.0281 | 236.7212 | 249.0102 | 119.6061 | 127.0486 | 136.7493 | 113.2123 | 79.2404 | 73.1151 | 77.6445 | 121.798 | 174.9233 | 163.1739 | 188.2225 | 145.5294 | 125 | 118.9233 | 125.202 | 111.0384 | 184.711 | 162.3939 | 178.7698 | 151.8466 | 125.798 | 142.4706 | 164.6726 | 138.1253 | 147.0384 | 159.9335 | 229.7288 | 241.9898 | 144.4808 | 151.913 | 225.4041 | 213.0767 | 161.1356 | 183.8747 | 198.6445 | 368.5294 | 135.8082 | 151.2685 | 176.5089 | 232.8363 | 140.5575 | 114.7596 | 139.5678 | 168.6343 | 165.5575 | 194.6726 | 181.5473 | 141.1918 | 90.087 | 103.4322 | 113.5089 | 104.5192 | 210.798 | 203.8466 | 248.1048 | 358.4527 | 223.9514 | 218.5575 | 206.4041 | 326.4425 | 168.3555 | 81.3836 | 71.0767 | 168.3555 | 98 | 100.0384 | 110.4425 | 112.1253 | 97.1918 | 101.2404 | 104.6061 | 130.2302 | 67.087 | 69.7315 | 73.798 | 90.6061 | 92.6829 | 106.2302 | 107.7212 | 81.3555 | 232.0384 | 177.2788 | 230.9233 | 468.8466 | 136.5575 | 181.5678 | 197.913 | 426.3939 | 258 | 190.2788 | 187.8849 | 374.5192
1691 | 20.5954 | 158.6313 | 193.3846 | 210.0531 | 198.8939 | 135.4562 | 162.9283 | 177.5093 | 146.6499 | 148.5968 | 117.1937 | 127.6313 | 170.6657 | 200.6657 | 142.9655 | 163.9469 | 217.2467 | 163.7905 | 154.5438 | 128.0717 | 185.9469 | 144.5438 | 286.6844 | 327.3315 | 149.4562 | 209.878 | 183.6154 | 225.6154 | 321.244 | 119.6657 | 111.5093 | 101.1592 | 116.9125 | 126.1406 | 140.5438 | 131 | 96.9842 | 228.3846 | 204.878 | 246.8063 | 271.5596 | 142.244 | 118.1406 | 146.0875 | 127.6127 | 75.1406 | 63.6499 | 83.4907 | 114.5093 | 177.6313 | 184.3315 | 206.6471 | 150.0345 | 117.2812 | 119.3156 | 124 | 140.5596 | 178.4218 | 159.0531 | 194.6127 | 151.0875 | 132.4032 | 134.1061 | 157.4218 | 160.4721 | 148 | 183.6471 | 204.7747 | 254.9469 | 150.175 | 146.7374 | 223 | 202.6499 | 174.3846 | 190.0875 | 185.1937 | 382.297 | 157.244 | 151.0531 | 176.0531 | 247.3501 | 113.4404 | 115.7719 | 129.5968 | 157.3343 | 169.3156 | 185.1061 | 180.5093 | 125.1592 | 87.5438 | 107.2281 | 108.4218 | 108.8594 | 216.6313 | 206.175 | 230.756 | 361.8063 | 220.5093 | 211.5093 | 217.122 | 343.5251 | 156.3343 | 68.6154 | 83.8063 | 156.3343 | 103.4032 | 106.4032 | 125.9814 | 122.8939 | 106.6313 | 99.4562 | 109.4907 | 117.4749 | 69.9469 | 71.4907 | 77.3156 | 91.6313 | 95.6313 | 97.3343 | 97.7374 | 77.0531 | 220.6499 | 218.1378 | 235.9469 | 462.6844 | 134.9125 | 195.5251 | 174.2123 | 406.0717 | 233.2998 | 193.3156 | 195.8063 | 354.9311
1692 | 20.6064 | 144.3166 | 198.4317 | 199.3741 | 164.5756 | 154.6043 | 152.7482 | 170.295 | 172.0144 | 136.0216 | 118.705 | 157.2302 | 167.9425 | 210.4892 | 169.0144 | 167.5683 | 201.8058 | 171.5468 | 145.5899 | 124.1367 | 206.036 | 148.9209 | 299.3309 | 312.4748 | 150.5683 | 199.7266 | 218.5611 | 219 | 307.8274 | 126.3525 | 95.6691 | 116.1726 | 146.5827 | 137.5468 | 126.8849 | 156.8777 | 81.1583 | 216.1799 | 228.446 | 244.5108 | 276.8633 | 124.9065 | 126.4101 | 123.4748 | 123.8058 | 86.5468 | 70.1942 | 67.1799 | 108.0791 | 193.1151 | 204.7626 | 190.9065 | 162.4101 | 115 | 120 | 156.9352 | 125.4748 | 146.7698 | 164.8417 | 195.5108 | 146.5108 | 137.1367 | 122.1583 | 157.1367 | 156.2374 | 151.1367 | 172.6115 | 207.6834 | 243.6691 | 139.4532 | 149.4892 | 210.4532 | 190.3741 | 187.2734 | 183.1583 | 205.5252 | 398.9784 | 157.9425 | 142.7266 | 132.4389 | 240.0216 | 132.446 | 143.446 | 138.1942 | 160.0575 | 177.8417 | 169.2374 | 190.5468 | 132.3309 | 102.1151 | 98.3741 | 116.7626 | 125.6834 | 210.9425 | 218.9784 | 200.259 | 354.8058 | 192.9065 | 196.4532 | 238.2518 | 349 | 177.0935 | 72.1942 | 80.7266 | 177.0935 | 100.295 | 104.8633 | 123.9425 | 121.5108 | 111.1367 | 115.4676 | 107.295 | 109.6475 | 72.7842 | 66.1583 | 73.295 | 98.4892 | 109.5468 | 95.3525 | 80.8849 | 84.4101 | 220.1367 | 193.9281 | 213.6907 | 458.0791 | 126.1583 | 200.2158 | 202.8561 | 391.9425 | 252.6619 | 181.4532 | 195.8633 | 399.9712
1693 | 20.6174 | 157.5222 | 180.4778 | 207.354 | 188.2479 | 152.0354 | 155.1682 | 158.646 | 162.4778 | 152.9116 | 116.0177 | 129.1592 | 185.1151 | 185.7167 | 154.699 | 162.4248 | 224.4868 | 150.3096 | 134.2389 | 137.7434 | 187.9026 | 132.4778 | 318.1328 | 338.655 | 153.3894 | 188.4424 | 199.7344 | 227.7611 | 314.7434 | 125.4071 | 107.9293 | 114.0354 | 140.6637 | 131.4424 | 120.6106 | 125.7698 | 86.9646 | 209.4071 | 216.8849 | 231.8495 | 268.8318 | 124.3717 | 117.0531 | 131.3363 | 137.7257 | 63.7167 | 79.3717 | 89.0798 | 101.4248 | 173.1061 | 203.2212 | 181.6106 | 160.2212 | 119.7788 | 135.1328 | 146.0884 | 111.8318 | 147.7611 | 143.1061 | 186.8318 | 165.708 | 138.7965 | 136.7257 | 166.7611 | 141.8495 | 149.6106 | 154.2566 | 190.4955 | 257.5222 | 169.4515 | 147.8141 | 210.9823 | 206.3187 | 186.6106 | 208.8762 | 219.1682 | 402.7965 | 154.4071 | 140.2035 | 141.708 | 230.6283 | 120.0884 | 126.3096 | 134.6283 | 140.4955 | 180.7965 | 171.5752 | 171.699 | 142.3717 | 100.4248 | 96.7965 | 108.8495 | 127.6106 | 205.8141 | 206.8672 | 226.655 | 370.3187 | 221.0444 | 206.5399 | 241.4071 | 360.9469 | 171.2566 | 65.4424 | 75.0177 | 171.2566 | 115.7257 | 97.6283 | 119.6106 | 122.3894 | 104.0354 | 115.2212 | 110.7788 | 112.1859 | 86.5399 | 65.5929 | 79.9646 | 96.8141 | 96.2743 | 103.9646 | 78.5929 | 79.0354 | 232.9469 | 203.9116 | 241.2479 | 460.1859 | 139.9293 | 184.8672 | 182.7344 | 412.301 | 266.3717 | 189.1505 | 190.2212 | 427.5399
1694 | 20.6284 | 140.1636 | 177.6175 | 217.2788 | 184.0599 | 149.1912 | 164.2788 | 156 | 160.4263 | 146.6774 | 119.8525 | 111.1037 | 200.5138 | 221.864 | 157.8963 | 177.1751 | 203.5024 | 170.6889 | 135.235 | 139.3825 | 177.9562 | 140.1313 | 301.7811 | 321.1199 | 151.5737 | 194.0875 | 196.8525 | 214.6337 | 349.5414 | 112.8687 | 123.3226 | 124.9401 | 125.6774 | 120.9125 | 132.9401 | 132.1751 | 94.6613 | 231.6451 | 216.8525 | 240.3226 | 271.8525 | 125.1912 | 122.8963 | 136.6175 | 133.1037 | 74.7926 | 64.0162 | 72.5461 | 93.53 | 193.6889 | 180.9724 | 205.2626 | 133.9286 | 118.5737 | 116.3549 | 126.4424 | 99.4862 | 171.0276 | 145.0875 | 185 | 164.53 | 143.8525 | 129.6774 | 152.0162 | 139 | 175.6889 | 167.9838 | 222.2027 | 257.9562 | 159.3987 | 154.2788 | 223.3226 | 208.5737 | 196.5138 | 189.3111 | 188.6498 | 379.5461 | 146.7212 | 150.5138 | 134.0599 | 224.1475 | 139.2626 | 118.3825 | 129.765 | 142.2788 | 199.6013 | 174.6175 | 189.4539 | 124.5899 | 102.235 | 102.6613 | 110.8525 | 118.9125 | 188.0162 | 197.3387 | 222.0599 | 358.8249 | 238.8963 | 204.3387 | 218.3549 | 334.0761 | 165.5737 | 74.3226 | 65.1037 | 165.5737 | 100.5899 | 109.7488 | 97.9724 | 122.1912 | 94.7212 | 92.1636 | 97.4424 | 117.8525 | 71.3987 | 60.3387 | 65.0162 | 87.1037 | 96.8525 | 89.8249 | 92.7488 | 80.235 | 220.6337 | 203.3387 | 245.9562 | 489.3064 | 157.7488 | 180.1912 | 204.1152 | 392.1199 | 303.5852 | 190.3825 | 190.6175 | 457.6889
1695 | 20.6394 | 146.4946 | 181.2842 | 224.7473 | 166.2212 | 144.8948 | 161.8948 | 148.6106 | 155.2527 | 152.2105 | 128.3894 | 125.4209 | 157.0215 | 202.4424 | 126.3371 | 150.6213 | 218.3472 | 153.1897 | 145.0315 | 163.6314 | 199.1682 | 143 | 315.884 | 303.5054 | 189.5891 | 187.7895 | 211.9578 | 218.3894 | 338.116 | 120.6736 | 115.3264 | 118.1475 | 138.5999 | 137.884 | 129.4316 | 126.9685 | 98.4631 | 212.3686 | 241.8103 | 238.8948 | 266.4316 | 121.7158 | 114.3264 | 130.4316 | 111.2949 | 76.5369 | 63.2842 | 62.0737 | 103.4946 | 200.821 | 170.2527 | 209.3579 | 150.9893 | 136.884 | 125.7788 | 117.2527 | 115.063 | 167.7895 | 139.6106 | 216.1997 | 175.3157 | 150.7473 | 145.063 | 138.1475 | 151.3157 | 157.3686 | 175.2842 | 212.937 | 244.6843 | 153.3579 | 157.6421 | 216.9685 | 191.5791 | 186.6843 | 193.6736 | 193.3157 | 374 | 138.4316 | 159.5684 | 153.9893 | 229.5684 | 144.179 | 136.884 | 146.242 | 141.5369 | 181.8318 | 170.0737 | 169.5476 | 102.758 | 94.7895 | 94.1475 | 98.0422 | 108.7895 | 216.0208 | 200.1052 | 191.0845 | 370.5999 | 212.2634 | 190.6843 | 225.3157 | 330.2842 | 178.1367 | 82.7473 | 69.5684 | 178.1367 | 117.3472 | 107.2527 | 109.4209 | 138.4209 | 94.6421 | 93.5684 | 110.4209 | 128.8525 | 69.4631 | 76.242 | 60.179 | 88.2842 | 87.3264 | 102.4209 | 101.7473 | 82.6421 | 220.2842 | 206.1052 | 237.6106 | 499.2842 | 141.2949 | 171.7895 | 184.7265 | 414.7366 | 266.0215 | 216.2735 | 206.5999 | 409.811
1696 | 20.6504 | 146.5 | 197.0001 | 200.1666 | 174.6667 | 130.6666 | 133.4999 | 159.5 | 172.3334 | 151.5 | 145.8334 | 114.8333 | 137 | 198.5 | 148.1668 | 162.3334 | 224.6667 | 165.5001 | 130.3333 | 152.3333 | 214 | 134.6666 | 317.5 | 300.1667 | 161.3332 | 212.6668 | 219.1667 | 204.1666 | 330.6667 | 94.6666 | 107.1666 | 111.8333 | 126.1666 | 133.6666 | 127.1667 | 127.5 | 87.3333 | 223.6667 | 215.3332 | 225.5 | 278.3334 | 115.1666 | 121.1667 | 139.8334 | 118.6667 | 81.8334 | 54.8333 | 56 | 105.1667 | 182.6666 | 189.8334 | 218.1667 | 169.3334 | 110.9999 | 108.9999 | 116 | 119.8333 | 159.3333 | 155.5001 | 192.1666 | 162.1666 | 151.1667 | 150.6667 | 158.5001 | 143.1666 | 159.5 | 192.6667 | 204 | 259.5001 | 143.8333 | 144.6666 | 198.3333 | 208.0001 | 205.6667 | 188.5 | 211.8334 | 374.8333 | 132 | 161.8333 | 167.3334 | 225.1666 | 130.6666 | 143.5 | 127.4999 | 147.6667 | 187 | 180.6667 | 157.3333 | 114.0001 | 98 | 104.5 | 105.8334 | 110.3333 | 191.3332 | 203.5 | 200.8334 | 361.5 | 217.6667 | 208.8334 | 205.4999 | 355.1668 | 165.9999 | 74.8333 | 61 | 165.9999 | 122.8333 | 107.6667 | 123 | 132 | 97.5 | 105.8334 | 122.3334 | 122.6666 | 69.1667 | 77.5 | 79.1667 | 97.3334 | 112.5001 | 101.8333 | 96.3333 | 73 | 220.1667 | 206.1667 | 238.5 | 489.1666 | 127 | 170.8333 | 194.8334 | 444.3334 | 259.3333 | 184.4999 | 175.8332 | 407.1667
1697 | 20.6614 | 149.3825 | 198.3087 | 232.2076 | 171.9262 | 154.2145 | 161.8251 | 167.9194 | 173.4631 | 172.9863 | 122.7855 | 128.9125 | 156.9863 | 197.3087 | 151.4631 | 166 | 235.1475 | 166.4631 | 157.4426 | 154.0738 | 219.3825 | 154.1407 | 343.2145 | 301.6913 | 170.0669 | 200.2418 | 204.7787 | 239.8989 | 319.0068 | 125.362 | 109.3825 | 105.9262 | 117.9262 | 128.6175 | 132.9194 | 132.2281 | 96.8388 | 200.7855 | 235.2145 | 231.4563 | 265.7787 | 133.4494 | 106.0875 | 120.0137 | 110.0068 | 64.3962 | 75.8319 | 73.6038 | 120.2213 | 176.4631 | 182.1612 | 206.47 | 155.9331 | 126.1407 | 112.6106 | 122.765 | 118.3087 | 152.9262 | 152.235 | 186.8456 | 155.6175 | 138.3156 | 135.7787 | 161.3087 | 156.2213 | 140.7049 | 167.2487 | 214.8388 | 275.6844 | 124.2418 | 152.1475 | 181.47 | 188.3224 | 210 | 192.0738 | 195.5506 | 374.1544 | 119.1612 | 140.0137 | 145.3224 | 232.4563 | 128.8456 | 150.765 | 114.5437 | 141.3894 | 183.9262 | 182.1544 | 179.6776 | 127.9932 | 108.3019 | 100.235 | 106.3087 | 127.0669 | 198.53 | 208.2281 | 227.6776 | 363.2281 | 222.5369 | 215.5369 | 211.9932 | 314.3361 | 180.7582 | 76.3825 | 79.2951 | 180.7582 | 113.6981 | 107.1544 | 103.8593 | 135.9194 | 107.3019 | 113.0738 | 115.5437 | 120.1544 | 58.0068 | 74.4631 | 88.9194 | 93.9262 | 119.6913 | 90.8525 | 89.0806 | 90.4494 | 221.6913 | 229.6776 | 231.3894 | 492.0738 | 138.53 | 190.4494 | 211.53 | 437.1612 | 259.1544 | 189.6844 | 201.9795 | 427.6038
1698 | 20.6724 | 131.1258 | 210.0109 | 222.6995 | 197.5955 | 147.8471 | 152.5575 | 160.4208 | 155.8416 | 164.9891 | 111.9946 | 143.1529 | 165.2896 | 174.6941 | 151 | 153.1312 | 215.552 | 151.4154 | 137.2624 | 138.6995 | 210.5629 | 158.8579 | 291.3774 | 332.0271 | 171.2842 | 202.1475 | 207.1475 | 225.552 | 300.6995 | 102.8308 | 132.3059 | 124.7321 | 126.4371 | 124.5683 | 132.2842 | 157.8796 | 101.5738 | 217.448 | 221.1258 | 239.0054 | 291.3113 | 136.1421 | 116.7267 | 122.8633 | 119.1529 | 77.3005 | 72.2787 | 64.9891 | 112.705 | 173.4262 | 204.0217 | 197.1367 | 171.8742 | 123.1367 | 126.8688 | 122.2842 | 128.295 | 170.8742 | 169.0163 | 187 | 173.0163 | 140.2896 | 121.8471 | 163.5738 | 177.8742 | 162.7376 | 174.0109 | 203.2733 | 241.9674 | 157.8905 | 172.8742 | 197.8742 | 209.7376 | 212.5738 | 181.8471 | 189.4262 | 374.8579 | 163.3276 | 146.295 | 157.3005 | 218.5575 | 152.1638 | 129.6941 | 108.7104 | 138.2842 | 201.0163 | 151.1149 | 193.4371 | 130 | 107.4262 | 129.0271 | 102.5683 | 105.1204 | 210.4371 | 196.1312 | 206.2624 | 361.4262 | 207.5575 | 222.8633 | 210.5683 | 313.7213 | 155.6887 | 67.5629 | 66.6995 | 155.6887 | 136.0217 | 114.7213 | 118.0163 | 134.4262 | 111.5738 | 99.4154 | 92.552 | 107.1312 | 68.8688 | 80.8633 | 79.705 | 81.8471 | 115.7104 | 87.2842 | 107.7321 | 69.9783 | 238.3005 | 198.8253 | 225.7104 | 515.3059 | 139.2842 | 179.4154 | 188.2624 | 385.2407 | 283.8796 | 198.8633 | 192.5575 | 370.0878
1699 | 20.6834 | 156.7169 | 206.7787 | 212.1681 | 195.0383 | 166.885 | 156.0915 | 164.2213 | 146.9085 | 139.5043 | 141.4573 | 127.5958 | 174.7021 | 197.1063 | 128.3746 | 152.7404 | 226.7936 | 174.2361 | 133 | 135.1298 | 189.8554 | 147.6873 | 310.7169 | 338.7404 | 146.6341 | 198.649 | 232.3659 | 255.068 | 324.1063 | 127.5871 | 118.5958 | 125.3894 | 146.2744 | 128.351 | 120.6873 | 139.3746 | 78.5043 | 234.0531 | 218.8702 | 223.466 | 244.6576 | 119.466 | 115.5192 | 129.2213 | 120.1298 | 75.649 | 76.2213 | 63 | 114.4808 | 182.5723 | 190.5958 | 228.1978 | 185.4425 | 114.1681 | 129 | 116.7787 | 108.2448 | 159.2064 | 153.7256 | 215.7169 | 186.4425 | 145.351 | 124.351 | 153.5575 | 173.1681 | 152.2064 | 170.7787 | 198.3894 | 235.1298 | 135.2831 | 148.1533 | 211.4425 | 200.9469 | 198.2064 | 192.1829 | 179.4277 | 400.2361 | 145.7639 | 157.5723 | 146.9469 | 233.4042 | 123.8022 | 137.3127 | 128.0148 | 144.9617 | 169.1916 | 155.5723 | 188.9085 | 131.7404 | 107.8702 | 85.2683 | 101.1298 | 108.8319 | 212.8702 | 200.0915 | 205.4808 | 339.2448 | 221.534 | 201.3746 | 226.534 | 323.7021 | 178.8467 | 86.885 | 78.7936 | 178.8467 | 113.8937 | 113.3894 | 134.0531 | 130.5192 | 103.2979 | 103.9617 | 98.5723 | 124.1446 | 88.4042 | 75.0383 | 63.2064 | 99.1446 | 100.2064 | 103.534 | 97.0766 | 75.5723 | 229.6873 | 233.0296 | 226.7404 | 486.8022 | 122.466 | 200.4957 | 169.2064 | 386.5723 | 266.2448 | 191.2979 | 193.4808 | 413.0828
1700 | 20.6944 | 161 | 190.9975 | 213.6475 | 181.645 | 164.705 | 150.8225 | 185.2975 | 139.8225 | 173.9475 | 129.2325 | 129.4125 | 164.5275 | 212.4725 | 155.005 | 160.9425 | 237.825 | 172.705 | 136.53 | 146.4725 | 194.06 | 153.06 | 314.1175 | 326.645 | 166.8275 | 220.0625 | 213.055 | 209.6975 | 315.645 | 115.2325 | 113.3525 | 125.8825 | 138.41 | 127.235 | 136.65 | 133.3525 | 89.12 | 181.285 | 196.9375 | 235.12 | 270.535 | 120.53 | 127.355 | 108.82 | 139.415 | 79.4125 | 81.4125 | 77.12 | 107.0575 | 187.53 | 188.8825 | 190.64 | 163.1725 | 146.535 | 137.825 | 123.9425 | 118.2375 | 158.765 | 180.1225 | 197.055 | 177.41 | 139.8225 | 147.945 | 151.1175 | 147.29 | 151.765 | 150.585 | 201.53 | 240.295 | 156.5925 | 150.1775 | 206.8225 | 217.5325 | 187.175 | 187.8225 | 172.705 | 363.405 | 172.8875 | 143.115 | 153.825 | 232.47 | 136.65 | 121.35 | 131 | 145.1175 | 185.18 | 183.475 | 157.1125 | 136.4125 | 92.9975 | 94.7675 | 103.6475 | 121.4725 | 205.94 | 190.41 | 206.8825 | 351.885 | 229.295 | 201.53 | 213.115 | 322.3525 | 162.7025 | 76.7625 | 71.2925 | 162.7025 | 109.1175 | 108.5875 | 125.41 | 117.645 | 96.705 | 106.765 | 103.53 | 123.47 | 66.29 | 74.8825 | 61.8825 | 94.0575 | 99.765 | 96.2925 | 102.06 | 79.6475 | 217.41 | 220.4675 | 221.705 | 493.4725 | 136.7675 | 196.0575 | 202.3 | 425.065 | 298.3 | 197.06 | 199.295 | 396.29
1701 | 20.7054 | 143.999 | 191.6844 | 218.474 | 183.5792 | 165.7896 | 151.7896 | 165.6302 | 155.1063 | 146.7875 | 136.8427 | 128.2104 | 179.1063 | 214.8948 | 151.8417 | 156.6312 | 254.2115 | 191.6854 | 139.6844 | 135.4729 | 209.3167 | 155.7896 | 300.5781 | 349.1594 | 150.3146 | 214.9469 | 210.8948 | 225.3698 | 325.6323 | 92.4198 | 92.4198 | 135.8427 | 123.5781 | 143.1063 | 141.6844 | 139.2635 | 96.3688 | 192.7906 | 250.3719 | 234.3156 | 278.5792 | 124.5792 | 126.3156 | 147.1604 | 128.5781 | 78.2104 | 68.5781 | 65.5781 | 115.8427 | 203.2115 | 188.1052 | 219.0021 | 143.8937 | 116.1031 | 139 | 119.6312 | 128.9479 | 162.5792 | 162.525 | 194 | 165.2625 | 140.7896 | 154.5792 | 146.526 | 160.1063 | 147.526 | 161.4219 | 190.3677 | 248.1583 | 143.8937 | 153.6844 | 216.7375 | 194.9458 | 195.8427 | 201.3167 | 201.5281 | 365.1583 | 157.3146 | 155.3167 | 137.1042 | 215.8937 | 123.7885 | 146.7385 | 135.474 | 147.6844 | 182.6312 | 167.3146 | 144.9469 | 147.7375 | 100.8427 | 86.2625 | 112.0531 | 129.2635 | 188.8937 | 218.5281 | 208.7896 | 370.1063 | 212.1042 | 208.2635 | 219.0531 | 321.1052 | 161.7896 | 84.8427 | 70 | 161.7896 | 106.3156 | 117.8427 | 134.7375 | 120.474 | 101.3688 | 98.9469 | 102.2104 | 122.1052 | 86.2646 | 67.8417 | 70.0531 | 98.3688 | 134.0021 | 97.6844 | 105.6844 | 84.474 | 226.7375 | 237.6854 | 227.2635 | 533.476 | 113.9458 | 183.3677 | 164.9448 | 415.6833 | 316.4219 | 183.6833 | 208.0531 | 428.7917
1702 | 20.7164 | 141.0929 | 201.9779 | 223.5354 | 172.2079 | 126.9954 | 152 | 168.4425 | 161.5354 | 153.885 | 133.4646 | 138.885 | 194.6062 | 196.8583 | 143.7433 | 156.9071 | 247.8363 | 169.5088 | 137.2788 | 134.9071 | 212.8142 | 139.6725 | 304.4425 | 324.7875 | 173.3983 | 185.8804 | 239.1196 | 208.0442 | 360.5621 | 112.6771 | 102.6992 | 124.3008 | 136.5133 | 148.6283 | 131.115 | 135.4646 | 86.115 | 204.9779 | 211.6459 | 234 | 274.4646 | 133.1637 | 114.2079 | 139.3008 | 139.6992 | 66.2079 | 62.4646 | 68.5354 | 89.7875 | 198.6504 | 183.4646 | 206.6725 | 170.1196 | 120.1637 | 120.8583 | 119 | 110.9513 | 182.0487 | 161.8142 | 214.8629 | 159.4646 | 115.6017 | 130.5088 | 142.3717 | 139.3229 | 116.1592 | 172.0708 | 194.4425 | 265.3275 | 158.3275 | 146.7433 | 214.3717 | 192 | 188.8363 | 187.5796 | 192.3008 | 369.6283 | 144.115 | 147.0221 | 151.3275 | 229.4204 | 134.6992 | 148.1858 | 150.5133 | 132.5796 | 165.6725 | 175.885 | 173.9337 | 120.8804 | 96.5575 | 92.2567 | 111.1858 | 134.5354 | 205.1417 | 198.4158 | 207.1858 | 325.7388 | 223.6062 | 231.6771 | 227.2567 | 327.3496 | 162.9071 | 76.0221 | 74.5354 | 162.9071 | 93.3008 | 109.0221 | 119.6725 | 108.3008 | 114.6992 | 107.0708 | 106.5354 | 124.7212 | 64.5088 | 75.1637 | 55.5796 | 93.5575 | 126.2079 | 105.2567 | 113.2567 | 83.1858 | 231.6283 | 217.3229 | 235.2567 | 503.5309 | 117.3496 | 157.5088 | 187.2125 | 415.8142 | 286.2521 | 180.1858 | 195.3938 | 424.8363
1703 | 20.7274 | 146.5162 | 214.9519 | 219.4031 | 201.3393 | 138.6294 | 140.9675 | 183.71 | 149.1288 | 147.645 | 133 | 137.2419 | 177.6125 | 172.935 | 144.8387 | 162.5162 | 261.71 | 178.0325 | 133.3225 | 151.5487 | 181.7413 | 171.0975 | 306.8387 | 326.5969 | 167.7256 | 205.9843 | 214.4188 | 268.5174 | 339.1769 | 130.6294 | 113.1937 | 112.8869 | 138 | 155.4356 | 144.71 | 146.0325 | 95.1131 | 235.42 | 206.0806 | 236.7581 | 294.2262 | 129.4031 | 108.4031 | 117.7738 | 120.7738 | 60.4031 | 77.6294 | 59.8063 | 93.4356 | 199.8387 | 188.5162 | 188.4513 | 156.4513 | 132.0325 | 109.8063 | 127.2744 | 115.4356 | 161.0157 | 148.2094 | 192.1769 | 189.3393 | 119.4356 | 166.6137 | 185.2106 | 175.6137 | 143.3393 | 166.5644 | 192.2419 | 240.3382 | 139.7738 | 147.8387 | 225.9519 | 214.065 | 196.2744 | 186 | 195.5969 | 371.8387 | 154.0325 | 139.5644 | 130.935 | 232.8387 | 140.5969 | 130.5319 | 144.645 | 143.8712 | 178.71 | 216.5331 | 182.5162 | 136.3875 | 108.8712 | 103.1131 | 94.4513 | 109.2575 | 196.8869 | 205.1937 | 215.2744 | 352.2587 | 207.5319 | 231.2419 | 217.8869 | 310.5319 | 162.0806 | 86.9519 | 71.3225 | 162.0806 | 109.4681 | 111.6775 | 125.355 | 111.5969 | 94.8544 | 115.355 | 124.4681 | 107.5319 | 60.1613 | 76.9194 | 69.6294 | 105.8712 | 105.6932 | 98.645 | 93.7738 | 91.2744 | 248.5487 | 233.3875 | 235.0806 | 485.29 | 147.42 | 205.5656 | 180.8063 | 445.42 | 302.3068 | 159.7738 | 190.3225 | 430.4356
1704 | 20.7384 | 140.4783 | 172.2118 | 211.5467 | 208.6583 | 134.41 | 128.82 | 186.8633 | 181.5399 | 173.0866 | 110.64 | 122.0934 | 198.36 | 191.4966 | 155.2483 | 159.2733 | 242.5034 | 184.59 | 147.9066 | 139.9567 | 206.9499 | 136.7334 | 336.8133 | 335.385 | 167.9317 | 207.0683 | 210.1367 | 214.3735 | 329.5467 | 135.7267 | 122.385 | 114.795 | 146.385 | 128.0501 | 157.18 | 124.64 | 82.9567 | 206.3234 | 215.3166 | 235.1367 | 308.1116 | 128.0683 | 117.3166 | 107.615 | 117.1367 | 80.4966 | 71.5467 | 76.7016 | 106.1116 | 186.9567 | 186.205 | 200.0433 | 147.5467 | 129.2733 | 139.7449 | 124.2733 | 110.41 | 159.9317 | 158.18 | 192.795 | 162.1867 | 138.6333 | 158.82 | 163.845 | 175.2733 | 171.155 | 169.7267 | 198.5217 | 263.155 | 151.0433 | 152.6583 | 206.5034 | 246.7449 | 208.18 | 208.36 | 178.2984 | 439.0799 | 176.4283 | 144.59 | 142.975 | 230.205 | 140.0683 | 129 | 131.8884 | 167.36 | 211.6766 | 220.9317 | 149.4601 | 149.18 | 123.0433 | 96.5467 | 85.5467 | 119.1116 | 229.5399 | 207.8633 | 190.845 | 357.795 | 219.0433 | 202.1184 | 220.7267 | 339.7449 | 170.385 | 74.9567 | 84.975 | 170.385 | 108.205 | 115.7267 | 128.795 | 122.2483 | 98.59 | 117.8633 | 131.59 | 133.9499 | 73.0433 | 66.7517 | 70.0683 | 105.1367 | 122.6333 | 105.4533 | 88.2733 | 66.845 | 230.435 | 232.205 | 213.5717 | 495.18 | 138.82 | 202.5467 | 202.36 | 411.6651 | 276.9817 | 196.1983 | 199.3166 | 431.9317
1705 | 20.7494 | 159.8231 | 189.767 | 211.944 | 188.233 | 156.6549 | 147.8231 | 155.8495 | 167.9527 | 179.7198 | 138.2626 | 116.2802 | 183.9527 | 188.2802 | 128.6253 | 168.4396 | 231.5604 | 189.7198 | 146.1681 | 143.7198 | 211.8319 | 161.3747 | 396.5813 | 336 | 162.3363 | 207.944 | 213.7758 | 203.3923 | 308.233 | 93.522 | 130.5516 | 123.4956 | 163.0473 | 147.711 | 119.2978 | 142.8231 | 74.4484 | 220.0473 | 224.4956 | 218.0088 | 258.9703 | 137.4396 | 131.2154 | 142.8703 | 121.7198 | 60.289 | 87.9912 | 69.5044 | 86.233 | 189.7758 | 186.944 | 201.944 | 172.4868 | 123.3363 | 115.5692 | 117.3923 | 118.4956 | 160.944 | 159 | 195.8319 | 172.2714 | 131.5044 | 124.0176 | 163.8879 | 168.3923 | 163.5604 | 175.6637 | 186.7286 | 261.2242 | 151.056 | 164.3275 | 192.7286 | 222.5692 | 216.5516 | 177.9055 | 202.4868 | 382.6429 | 160.0648 | 165.767 | 140.2242 | 265.8703 | 132.4484 | 136.5516 | 124.3923 | 156.7286 | 184.7374 | 169.0824 | 169.6549 | 139.6165 | 121.1681 | 96.944 | 100.1033 | 118.1121 | 203.6813 | 215.5516 | 200.3275 | 362.7198 | 245.4868 | 212.2714 | 240.8231 | 341.056 | 165.3363 | 72.1121 | 79.3923 | 165.3363 | 113.6637 | 118.8319 | 109.1769 | 145.6549 | 119.767 | 127.4396 | 105.5692 | 117.1209 | 68.3363 | 69.7758 | 82.2714 | 118.2154 | 111.7286 | 91.8407 | 106.8791 | 85.767 | 257.3187 | 215.9527 | 221.4396 | 475.233 | 136.1121 | 192.5604 | 169.0736 | 419.3835 | 306.1505 | 199.944 | 181.1209 | 449.9352
1706 | 20.7604 | 156.2188 | 188.1313 | 216.7812 | 211.8624 | 145.5688 | 149.9562 | 178.8624 | 175.6062 | 146.5314 | 127.5688 | 112.175 | 188.7375 | 212.8624 | 133.6937 | 164.2188 | 232.9125 | 188.0875 | 130.7 | 158.3437 | 207.2188 | 151.525 | 349.3189 | 315.9188 | 188.7749 | 181.2251 | 226.4312 | 213.5187 | 339.5124 | 124.4686 | 127.175 | 109.6563 | 150.6125 | 139.4375 | 137.0812 | 129.6563 | 93.1249 | 229.6062 | 219.2625 | 264.8123 | 284.6874 | 153.3 | 105.2251 | 119.1813 | 118.175 | 76.2124 | 62.2251 | 62.3063 | 115.5999 | 174.7 | 232.8999 | 209.65 | 170.175 | 144.0374 | 137.9062 | 129.4312 | 154.3811 | 187.7749 | 139.8751 | 185.4813 | 169.175 | 161.5999 | 140.1687 | 160.175 | 160.35 | 160.1313 | 183.65 | 207.0374 | 258.1313 | 158.65 | 139.1813 | 230.2499 | 218.1313 | 218.9125 | 173.1313 | 182.9626 | 395.2562 | 160.9125 | 144.0501 | 145.7375 | 240.2688 | 145.3875 | 137.9562 | 123.0438 | 146.4375 | 192.5625 | 203.2936 | 214.9874 | 121.7876 | 97.0938 | 103.6937 | 106.7375 | 128.5187 | 236.4249 | 233.2124 | 211.5187 | 347.7 | 217.3563 | 234.0374 | 221.9188 | 335.2625 | 156.3938 | 77.7375 | 58.9188 | 156.3938 | 109.2188 | 92.2251 | 133.8187 | 117.3563 | 100.9188 | 129.9125 | 143.2061 | 116.9562 | 77.5625 | 70 | 67.7 | 110.3938 | 112.9125 | 92.9125 | 107.0438 | 85.0875 | 244.6563 | 226.475 | 219.1313 | 451.0501 | 155.1249 | 188.175 | 183.2562 | 422.8687 | 327.1249 | 190.4375 | 204.8624 | 450.0438
1707 | 20.7714 | 142.4405 | 196.7168 | 230.5595 | 201.3775 | 159.5281 | 152.9056 | 154.818 | 186.6539 | 167.2764 | 119.2517 | 125.5595 | 193.8427 | 214.9685 | 146.591 | 146.5663 | 249.4651 | 188 | 152.2764 | 126.0697 | 227.3393 | 137.4405 | 358.6225 | 332.4337 | 156.1012 | 215.8359 | 197.9439 | 239.182 | 321.6292 | 116.3146 | 101.818 | 119.6539 | 136.4405 | 143.8427 | 121.5349 | 140.6225 | 92.0629 | 233.8742 | 224.8112 | 239.8809 | 288.9056 | 103.636 | 120.4651 | 120.9056 | 110.2517 | 68.2832 | 88.1191 | 48.4405 | 104.409 | 199.182 | 184.636 | 204.1888 | 169.0315 | 110.1326 | 125.4405 | 122.2517 | 125.9753 | 142.5102 | 144.8112 | 218.8988 | 162.2202 | 153.3146 | 143.9056 | 164.8427 | 168.7168 | 167.7483 | 157.8495 | 177.0068 | 267.6854 | 148.3461 | 151.5595 | 225.2202 | 228.6539 | 226.7483 | 203.0247 | 187.8112 | 421.182 | 146.4719 | 154.6225 | 163.4337 | 237.0629 | 134.3775 | 146.7168 | 156.8988 | 172.1505 | 208.4966 | 208.8742 | 175.3529 | 145.2134 | 104.7168 | 89.4719 | 99.2517 | 106.7236 | 211.8495 | 217.5349 | 237.182 | 360.5595 | 212.1258 | 213.6922 | 220.0315 | 355.3393 | 189.8988 | 73.1573 | 60.9056 | 189.8988 | 121.591 | 110.3708 | 112.7236 | 146.0247 | 115.4966 | 110.6292 | 102.3843 | 119.9056 | 77.0315 | 60.3146 | 67.9685 | 99.3461 | 106.2202 | 87.1888 | 91.5034 | 77.2517 | 244 | 246.3708 | 225.7798 | 502.3011 | 151.1573 | 193.8112 | 199.4966 | 443.3393 | 331.8742 | 213.2449 | 183.7236 | 435.4719
1708 | 20.7824 | 140.0383 | 187.1917 | 211.3834 | 197.0767 | 159.0192 | 142.2109 | 147.1342 | 147.7668 | 166.0383 | 144.5016 | 126 | 188.115 | 204.2109 | 160.7316 | 142.0767 | 256.8658 | 172.3067 | 146.1342 | 142.6549 | 195.6326 | 167.4057 | 272.687 | 321.2301 | 178.5399 | 207.1917 | 219.5591 | 202.7285 | 328.8466 | 127.7699 | 132.3865 | 127.8466 | 157.5782 | 157.7316 | 131.7891 | 126.2876 | 92 | 204.5751 | 229.9041 | 235.0767 | 291.9425 | 113.7699 | 119.0383 | 114.1342 | 144.329 | 76.8275 | 79.1917 | 85.2715 | 110.8658 | 193.1342 | 184.9617 | 210.8658 | 191.5591 | 120.7699 | 119.115 | 135.7316 | 114.2109 | 158.6549 | 163.6358 | 212.1534 | 191.4249 | 153 | 140.0767 | 162.0575 | 163.115 | 150.3451 | 178.5782 | 203.4632 | 281.7316 | 138.1917 | 156.9041 | 207.3451 | 223.115 | 207.3834 | 235.3865 | 220.3674 | 390.6135 | 122.4601 | 136.3642 | 169.885 | 215.4218 | 141.8466 | 147 | 140.3451 | 156.3259 | 205.0767 | 185.4601 | 163.2109 | 169.5399 | 120.6933 | 92.9233 | 103.9041 | 146.214 | 223.7508 | 221.9041 | 194.8435 | 370.8083 | 258.099 | 208.0959 | 241.5782 | 329.5176 | 186.0959 | 71.0383 | 89.444 | 186.0959 | 133.7699 | 113.9425 | 107.0959 | 120.5176 | 106.1917 | 114.9041 | 91.1917 | 107.2492 | 76.0192 | 65.885 | 69.9617 | 106.8466 | 111.885 | 117.4057 | 105.7124 | 78.9617 | 228.3067 | 219.5368 | 247.5782 | 493.2109 | 154.9233 | 225.3865 | 178.4218 | 449.885 | 325.1342 | 204.1917 | 213.4057 | 430.0959
1709 | 20.7934 | 152.9106 | 206.8624 | 175.2477 | 201.9656 | 160.9862 | 149.945 | 166.8624 | 167.8555 | 171.9587 | 149.9656 | 126.9931 | 182.0413 | 221.8762 | 140.1445 | 158.883 | 247.0688 | 158.0963 | 160.8968 | 141.0138 | 187.055 | 168.9931 | 349.4565 | 316.0344 | 172.0482 | 218.9174 | 230.9243 | 227.8211 | 335.9518 | 114.0963 | 91.289 | 107.1445 | 118.2752 | 147.0757 | 126.0413 | 118.055 | 87.0344 | 225.8486 | 214.1101 | 201.2339 | 296.9656 | 114.9931 | 114.0344 | 141.8074 | 110.2408 | 80.9725 | 74.0344 | 62.1651 | 142.7798 | 220.8074 | 205.8555 | 202.0619 | 171.1445 | 132.9174 | 135.883 | 118.1238 | 149.7523 | 178.8624 | 170.9518 | 174.2614 | 174.1238 | 137.1101 | 152.9106 | 166.9656 | 153.0688 | 141.0619 | 200.8486 | 191.0894 | 287.9587 | 172.7592 | 145.0826 | 194.0894 | 223 | 226.8624 | 197.2683 | 215.0413 | 398.9381 | 174.6354 | 136 | 149.1445 | 260.6835 | 138.0275 | 169.8418 | 144.9656 | 176.8555 | 206.9862 | 213.8005 | 163 | 117.3646 | 91.2064 | 105.9106 | 100.0275 | 108.2683 | 199.172 | 208.0963 | 203.9312 | 369.0138 | 257.0138 | 223.8899 | 241.0069 | 318.0757 | 179.0482 | 64.0482 | 81.0619 | 179.0482 | 145.9174 | 115.9862 | 128.8486 | 139.8624 | 106 | 105.0688 | 92.9862 | 121.8968 | 62.0963 | 67.9862 | 69.0069 | 96.0757 | 112 | 118.9931 | 105.0069 | 87.9381 | 267.7248 | 250.7798 | 210.2614 | 490.0206 | 147.055 | 220.0413 | 210.773 | 456.9518 | 315.0688 | 221.8762 | 198.1101 | 452.8418
1710 | 20.8044 | 148.0703 | 190.9784 | 187.1353 | 205.8485 | 147.0595 | 162.0433 | 157.9297 | 175.8918 | 184.9134 | 139.1082 | 112.9838 | 248.7727 | 201.1136 | 128.0649 | 174.8268 | 225.9892 | 175.0108 | 132.0487 | 139.9405 | 189.0974 | 174.9188 | 335.7836 | 316.1948 | 193.8972 | 194.0595 | 209.9026 | 238.816 | 356.7944 | 108.9838 | 107.0379 | 129.0216 | 146.8485 | 140.9892 | 127.0433 | 148.1299 | 86.0216 | 208.0433 | 194.2056 | 222.0433 | 287.9297 | 138.1028 | 134.0108 | 123.8701 | 136.9134 | 78.0054 | 70.9188 | 83.8485 | 117.0271 | 211.0812 | 203.8377 | 221.0271 | 181.0758 | 143.9675 | 125.1028 | 131.8755 | 118 | 179.9134 | 160.0974 | 222.9459 | 209.8864 | 180 | 160.881 | 176.0487 | 172.0541 | 177.8972 | 177.0649 | 219.9297 | 286.9134 | 152.0162 | 157.0108 | 200.0054 | 257.8268 | 226.8701 | 198.0974 | 196.9242 | 395.2922 | 138.1894 | 150.0325 | 157.9838 | 248 | 141.0379 | 152.9567 | 127.0271 | 151.0758 | 219.9242 | 219.8972 | 180.8593 | 177.881 | 95.1028 | 110.0487 | 99 | 139.8485 | 236.0216 | 224.0866 | 211.8864 | 373.9729 | 242.1028 | 222.1082 | 218.1515 | 343.8755 | 185.0054 | 92.9242 | 68.0433 | 185.0054 | 109.1461 | 114.9351 | 108.0649 | 148.9729 | 109.0325 | 109.0054 | 100.1028 | 132.9134 | 79.9459 | 56.0379 | 78.9729 | 110.0379 | 110.119 | 123.1299 | 124.9405 | 72.1245 | 263.0108 | 222.9946 | 209.1515 | 486.0054 | 138.0325 | 194 | 211.9892 | 408.1082 | 358.8539 | 209.1353 | 187.0487 | 438.0541
1711 | 20.8154 | 160.8761 | 187.177 | 212.3894 | 178.4957 | 157.9823 | 169.8584 | 145.2124 | 156.1947 | 168.7522 | 158.5397 | 110.0177 | 207.0708 | 221.9823 | 140.2301 | 142.9823 | 223.8938 | 177.2478 | 141.3717 | 129.1947 | 206.6637 | 159.8053 | 296.5134 | 351.6991 | 175.3186 | 205.4072 | 192.832 | 205.2655 | 319.5842 | 106.3894 | 113.5928 | 132.7699 | 119.8674 | 139.1416 | 134.9469 | 171.2742 | 89.7876 | 216.354 | 231.7168 | 229.8053 | 275.4603 | 156.4689 | 135.9115 | 100.1239 | 120.9823 | 78.823 | 56.4249 | 56.4072 | 121.646 | 226.2832 | 174.7258 | 225.5043 | 194.5397 | 137.8584 | 143.5751 | 108.7876 | 118.2832 | 164.478 | 177.6991 | 212.9469 | 189.2478 | 179.7522 | 139.3363 | 184.9823 | 182 | 158.9292 | 188.522 | 206.4866 | 271.5842 | 154.9823 | 158.9646 | 201.0708 | 226.0354 | 203.1416 | 215.3627 | 183.5134 | 448.4335 | 172.522 | 155.9292 | 155.1416 | 247.646 | 147.9469 | 144.9469 | 132.3363 | 164.8761 | 206.4072 | 201.8674 | 155.3717 | 155.646 | 113.823 | 118.6106 | 99.4426 | 111.8584 | 239.5928 | 239.5574 | 191.0531 | 369.0354 | 261.0708 | 241.7876 | 245.6991 | 320.8584 | 185.9115 | 79 | 75.8407 | 185.9115 | 135.6814 | 103.1062 | 119.6106 | 143.7345 | 114.5751 | 109.9646 | 119 | 116.9823 | 69.9292 | 63 | 73.9646 | 116.6637 | 131.6106 | 146.2919 | 113.5928 | 94.6991 | 265 | 221.9646 | 236.9469 | 487.177 | 144.0885 | 194.2301 | 210.2301 | 427.8407 | 331.4689 | 233.7522 | 196.0354 | 448.4072
1712 | 20.8264 | 154.06 | 197.2999 | 233.82 | 205.97 | 156.7301 | 162.2399 | 156.79 | 166.1302 | 154.88 | 132.3401 | 111.21 | 210.2502 | 221.4799 | 152.91 | 142.6599 | 219.7096 | 191.2399 | 160.9502 | 140.4499 | 188.09 | 149.18 | 325.12 | 334.94 | 192.6401 | 227.85 | 238.5501 | 220.5099 | 351.88 | 127.85 | 91.5099 | 119.88 | 167.3401 | 145.8003 | 131.97 | 130.6701 | 78.06 | 236.7498 | 216.6299 | 219.5699 | 301 | 126.91 | 131.3899 | 106.6701 | 120.4799 | 69.06 | 79.91 | 78.5801 | 102.2999 | 241.3102 | 214.2202 | 198.8998 | 169.5099 | 129.97 | 120.8398 | 97.3599 | 134.09 | 190.4001 | 160.88 | 209.4601 | 202.91 | 165.6701 | 157.5501 | 183.79 | 181.85 | 154.7601 | 162.5999 | 179.6796 | 303.2202 | 154.4799 | 156.7301 | 205.8098 | 227.79 | 210.6101 | 181.2597 | 212.2399 | 416.97 | 146.7798 | 151.97 | 162.5801 | 229.1697 | 146.1098 | 142.7498 | 150.4301 | 157.6401 | 228.1602 | 249.4001 | 176.8698 | 137.1397 | 103.82 | 97.3899 | 123.1602 | 104.9898 | 217.15 | 215.7498 | 194.7798 | 371.9898 | 264.5201 | 229.3401 | 228.97 | 314.1697 | 180.82 | 79.2699 | 66.4601 | 180.82 | 118.3599 | 109.2999 | 98.4799 | 129 | 91.6898 | 108.4799 | 118.3102 | 116.5999 | 66.4499 | 63.06 | 71.94 | 97.91 | 110.5399 | 106.82 | 90.7301 | 78.8398 | 264.2502 | 219.85 | 233.1602 | 496.94 | 149.3299 | 206.7301 | 221.6503 | 420.0198 | 301.91 | 219.6101 | 198.3599 | 470.7301
1713 | 20.8374 | 156.1691 | 206.8309 | 227.3657 | 205.2114 | 147.2812 | 169.7463 | 150.7188 | 139.3108 | 152.5222 | 112.1417 | 117.8309 | 186.592 | 237.1691 | 151.0148 | 163.9577 | 274.2389 | 198.9577 | 127.2537 | 153.816 | 192.48 | 154.112 | 328.4926 | 334.5645 | 180.1966 | 222.9577 | 224.7611 | 238.2685 | 349.0148 | 122.9577 | 107.5349 | 116.9725 | 146.3806 | 108.7337 | 131.9303 | 119.9577 | 81.0994 | 259.3086 | 236.8309 | 237.1543 | 298.8435 | 123.5772 | 143.8731 | 96.8457 | 135.7463 | 70.7886 | 76.408 | 65.8034 | 112.1269 | 219.7188 | 189.6765 | 227.3235 | 186.0423 | 129.2114 | 147.6617 | 109.7188 | 136.9577 | 171.2114 | 156.1543 | 192.8457 | 199.8731 | 154.7886 | 143.8457 | 177.6765 | 177.296 | 147.0846 | 181.8309 | 232.9852 | 279.8182 | 169.5349 | 148.2114 | 232.296 | 221.1269 | 199.184 | 220.9852 | 218.8583 | 416.2114 | 172 | 151.3383 | 148.9154 | 266.8309 | 180.4355 | 165.52 | 131.5772 | 146.6765 | 201.1691 | 229.704 | 204.2389 | 172.52 | 98.2537 | 109.408 | 96.7611 | 135.3509 | 220.816 | 239.6194 | 219.2389 | 404.3806 | 249.3806 | 208.4651 | 227.3657 | 352.296 | 174.6617 | 87.9154 | 50.3108 | 174.6617 | 129.8309 | 118.9577 | 113.9154 | 129.1269 | 114.2114 | 123.4926 | 96.0846 | 136 | 80.4503 | 65.888 | 70 | 96.1417 | 128.4228 | 102.0994 | 82.888 | 104.6892 | 241.1417 | 215.5497 | 206.3383 | 496.48 | 159.1966 | 197.8731 | 178.4228 | 451.6892 | 300.6491 | 207.8457 | 210.592 | 463.6491
1714 | 20.8484 | 159.5634 | 202.9454 | 211.8539 | 210.1637 | 131.4912 | 163.8363 | 166.1814 | 169.5458 | 187.1637 | 137.7817 | 116.2376 | 200.2183 | 238.2167 | 171.8716 | 162.3997 | 256.5088 | 195.8716 | 133.2729 | 127.7641 | 226.0546 | 136.4559 | 315.7993 | 365.9068 | 163.4735 | 223.1461 | 240.6356 | 265.9085 | 370.5265 | 121.9454 | 97.382 | 138.6725 | 153.3627 | 148.0546 | 152.618 | 119.3275 | 104.7993 | 221 | 234.2552 | 219.2007 | 251.9647 | 115.2552 | 140.7817 | 116.1637 | 129.2359 | 66.3275 | 63.7641 | 83.8363 | 115.0546 | 234.0353 | 204.6725 | 212.4366 | 185.6356 | 134.3275 | 139.0176 | 125.8363 | 135.0176 | 176.2729 | 138.2007 | 210.7993 | 197.5458 | 150.6003 | 162.3451 | 193.1092 | 183.0176 | 149.1092 | 179.9647 | 210.6003 | 319.6356 | 159.2183 | 152.8363 | 239.1637 | 224.9824 | 225.6725 | 199.5281 | 192.6725 | 420.5088 | 171.2905 | 158.8908 | 146.5088 | 263.4912 | 146.0369 | 134.0739 | 122.7095 | 161.6725 | 205 | 223.8186 | 186.7271 | 140.4735 | 104.5458 | 97.2552 | 112.6902 | 98.8732 | 194.4912 | 230.0176 | 202.9278 | 410.7078 | 259.1461 | 219.4912 | 212.618 | 357.581 | 168.2552 | 85.3451 | 79.8908 | 168.2552 | 127.5827 | 119.2007 | 111.6725 | 133.0369 | 118.7271 | 111.8908 | 98.6003 | 134.7993 | 68.1092 | 85.7817 | 70.1637 | 120.7448 | 137.1814 | 127.1092 | 103.1092 | 75.0546 | 265.6902 | 226.9631 | 214.3275 | 526.5072 | 141.0546 | 195.2729 | 187.5634 | 422.7095 | 336.3082 | 227.8186 | 223.3451 | 498.6533
1715 | 20.8594 | 150.3952 | 201.4651 | 193.8723 | 212.7325 | 141.4042 | 162.0699 | 153.7385 | 177.7964 | 188.9301 | 132.9301 | 152.7934 | 202.9301 | 192.4072 | 135.9361 | 153.8054 | 248.1337 | 160.4042 | 137.5319 | 142.2036 | 224.8603 | 178.0639 | 296.6048 | 297.0729 | 187.3283 | 243.4012 | 217.7355 | 244.4591 | 345.8693 | 120.4651 | 104.6687 | 132.3313 | 125.6018 | 148.5988 | 145.2645 | 126.3373 | 85.7385 | 221.6687 | 253.0579 | 239.5319 | 284.997 | 136.1976 | 137.2675 | 119.2006 | 117.0699 | 72.0669 | 75.8633 | 79.6627 | 118.3403 | 199.5319 | 199.4681 | 220.0669 | 165.0759 | 138.3952 | 123.1367 | 123.3343 | 117.7994 | 181.2006 | 159.4012 | 190.8024 | 206.5988 | 159.5958 | 151.5349 | 194.2645 | 167.3373 | 151.9361 | 211.9271 | 219.7295 | 296.2006 | 163.4012 | 151.7385 | 238.991 | 240.3952 | 218.3283 | 225.5319 | 186.6657 | 410.9301 | 158.6657 | 156.5319 | 139.8054 | 270.9301 | 162.3952 | 168.4621 | 135.8024 | 155.5988 | 205.0669 | 234.9241 | 184.6078 | 165.7325 | 113.3982 | 118.7325 | 90.6687 | 115.8723 | 204.3373 | 212.3313 | 219.2006 | 373.3403 | 278.3313 | 228.4012 | 205.9331 | 335.4741 | 189.6657 | 74.2006 | 76.9301 | 189.6657 | 152.6597 | 138.4621 | 107.1367 | 147.991 | 115.003 | 109.2645 | 108.5988 | 114.1337 | 70.5349 | 81.2645 | 72.5988 | 98.3982 | 122.9331 | 127.8633 | 105.9361 | 77.8054 | 242.7325 | 212.6777 | 221.3373 | 469.2096 | 142.2006 | 198.1277 | 180.5349 | 435.5349 | 308.5379 | 242.9361 | 213.4042 | 455.8603
1716 | 20.8704 | 129.5831 | 193.2876 | 220.0792 | 210.2665 | 160.6834 | 174.7836 | 178.1583 | 164.1663 | 173.5251 | 120.6914 | 121.7625 | 188.7124 | 224.496 | 149.6333 | 176.9419 | 252.9288 | 180.3166 | 131.0792 | 159.3166 | 197.0581 | 164.6042 | 320.2665 | 325.3667 | 164.2375 | 246.7044 | 225.7044 | 209.0792 | 358.3958 | 112.8417 | 113.1293 | 121.2585 | 134.6333 | 143.7124 | 135.8707 | 143.5752 | 108.1794 | 232.2665 | 213.1874 | 232.7625 | 272.504 | 124.6042 | 139.1002 | 120.8918 | 133.4248 | 72.6042 | 61.6623 | 61.2375 | 147.992 | 194.2665 | 207.2665 | 220.3667 | 205.0712 | 115.6042 | 138.1293 | 128.2375 | 116.4248 | 182.8918 | 164.2084 | 200.496 | 201 | 141.3457 | 158.1293 | 183.7625 | 184.3377 | 164.4459 | 184.1082 | 200.1794 | 299.0792 | 167.8918 | 174.5752 | 197.6333 | 218.1583 | 195.1583 | 218.6834 | 182.9499 | 398.4539 | 152.5752 | 150.8707 | 164.9208 | 255.1293 | 139.0501 | 146.8417 | 145.1794 | 149.6042 | 205.8417 | 193.1874 | 216.963 | 158.2796 | 104.5251 | 114.3667 | 99.1293 | 141.8417 | 223.1583 | 204.504 | 220.5752 | 405.6834 | 270.504 | 231.7044 | 207.6914 | 368.3377 | 184.9208 | 75.971 | 63.9789 | 184.9208 | 122.6122 | 118.504 | 123.4749 | 108.533 | 126.9419 | 100.3457 | 104.1874 | 118.3747 | 77.6042 | 73.4539 | 67.3958 | 91.2665 | 121.2084 | 111.7625 | 119.6333 | 100.8627 | 239.5541 | 260.2796 | 239.6834 | 517.4829 | 145.8707 | 171.1293 | 189.8206 | 443 | 327.7836 | 254.2585 | 233.4749 | 429.4829
1717 | 20.8814 | 147.0855 | 185.6401 | 220.5428 | 222.1651 | 157.5487 | 150.0059 | 178.6283 | 198.9881 | 167.6342 | 149.9881 | 118.7257 | 196.4513 | 209.3777 | 157.4572 | 152.2684 | 282.7934 | 183.5428 | 135.2007 | 161.7197 | 222.3717 | 158.9026 | 335 | 319.0974 | 166.7257 | 217.4394 | 201.9264 | 211.9204 | 363.823 | 111.9145 | 103.2743 | 101.823 | 141.5428 | 149.8052 | 146.2743 | 126.177 | 89.0119 | 246.3599 | 228.0059 | 229.0855 | 289.3599 | 121.0974 | 118.2803 | 108.9145 | 148.171 | 70.0119 | 81.2743 | 63.8171 | 113.7316 | 208.7257 | 222.0914 | 215.652 | 170.4513 | 112.6461 | 127.8171 | 132.0059 | 134.0974 | 169.3599 | 155.2743 | 182.4513 | 200.8171 | 158.1888 | 148.2743 | 178.5309 | 166.0059 | 159.3717 | 196.0855 | 182.3836 | 299.4513 | 152.8967 | 159.829 | 203.0796 | 218.9026 | 199.2862 | 215.3658 | 194.7316 | 424.8967 | 137.6461 | 160.0855 | 164.0914 | 242.5309 | 129.6461 | 143.7197 | 125.4632 | 145.9145 | 202.171 | 208.2803 | 175.6698 | 118.1093 | 98.6342 | 104.9881 | 89.0914 | 138.2625 | 222.8967 | 220.6283 | 206.3777 | 400.0796 | 287.2684 | 208.1093 | 237.171 | 348.5428 | 184.823 | 64.5487 | 85.6283 | 184.823 | 148.2447 | 135.0855 | 126.7138 | 139.6461 | 103.1829 | 118.1033 | 116.9026 | 143.1651 | 71.9941 | 99.7138 | 71.3599 | 105.5428 | 113.4632 | 108.5428 | 125.0796 | 75.5428 | 247.1888 | 216.6342 | 236.823 | 554.7079 | 153.0736 | 165.0237 | 209.6283 | 438.7019 | 300.9941 | 235.0059 | 237.0796 | 465.2447
1718 | 20.8924 | 139.1411 | 193.2449 | 215.1701 | 196.1785 | 161.0289 | 162.0748 | 163.2364 | 180.3486 | 167.0084 | 130.6224 | 116.5187 | 187.9916 | 231.7551 | 160.5476 | 146.2449 | 243.801 | 182.8384 | 163.7841 | 148.7925 | 230.813 | 147.585 | 334.3776 | 329.0663 | 164.1037 | 194.386 | 227.2654 | 230.4813 | 368.0579 | 121.9337 | 107.3486 | 110.2075 | 136.8963 | 129.7636 | 147.2364 | 120.5935 | 105.8878 | 238.9626 | 234.5766 | 220.415 | 283.8299 | 132.7262 | 132.7636 | 118.8299 | 131.1411 | 89.1701 | 85.1411 | 63.6598 | 118.8215 | 205.0663 | 223.7262 | 240.2364 | 164.585 | 128.1701 | 126.8299 | 138.7841 | 142.8215 | 164.9711 | 159.7636 | 179.2823 | 196.199 | 167.9252 | 151 | 156.386 | 174.7551 | 175.4898 | 187.6224 | 216.6224 | 295.8673 | 131.6888 | 174.1616 | 190.4319 | 208.6224 | 220.7551 | 217.7551 | 204.5935 | 402.1327 | 154.5187 | 150.585 | 164.4813 | 219.5561 | 146.5187 | 133.386 | 138.9626 | 154.0663 | 181.8215 | 220.3776 | 227.1242 | 149.1037 | 95.6224 | 85.6224 | 90.415 | 119.7551 | 202.1037 | 206.2738 | 228.2364 | 377.4729 | 278.4439 | 239.5187 | 219.1037 | 348.3571 | 191.7551 | 68.3402 | 72.9337 | 191.7551 | 104.7346 | 126.8299 | 105.6598 | 153.4065 | 105.8299 | 136.8215 | 107.1411 | 118.4234 | 63.1411 | 76.2738 | 65.9337 | 98.5102 | 125.4065 | 103.585 | 106.6224 | 72.3486 | 257.2364 | 214.4524 | 242.9252 | 521.8963 | 124.1037 | 204.7925 | 197.4524 | 397.0374 | 291.2075 | 242.9252 | 219.1411 | 416.0289
1719 | 20.9034 | 152.1328 | 205.9725 | 210.3206 | 215.3481 | 145.8565 | 175.8229 | 149.1603 | 192.3481 | 190.2153 | 138.6687 | 123.4366 | 163.0443 | 221.5802 | 148 | 156.884 | 263.9389 | 211.5908 | 138.2046 | 150.3649 | 186.4641 | 145.8565 | 324.9389 | 319.8397 | 166.3924 | 215.6962 | 196.3924 | 224.4916 | 338.5252 | 128.0275 | 118.1878 | 112 | 138.2046 | 144.0718 | 133.7405 | 142.884 | 82.1328 | 236.7298 | 207.6687 | 223.7679 | 291 | 136.3313 | 144.6351 | 127.7405 | 138.5634 | 81.5359 | 91.6351 | 77.0718 | 100.6962 | 198.6244 | 226.055 | 225.6962 | 161.2321 | 122.9725 | 132.4916 | 111.6962 | 126.7847 | 179.5634 | 175.4641 | 201.4366 | 171.884 | 151.1603 | 151 | 178.1603 | 161.9114 | 191.6626 | 192.5359 | 219.7954 | 308.171 | 132.0168 | 141.3481 | 239.8504 | 210.7511 | 206.9832 | 207.3481 | 224.6794 | 388.7847 | 157.3756 | 147.2321 | 160.116 | 236.5969 | 148.2153 | 152.9557 | 130.9282 | 148.4366 | 164.6244 | 208.0382 | 183.6076 | 149.3038 | 101.116 | 90.3038 | 98.0611 | 110.2763 | 201.4916 | 199.5359 | 214.3924 | 349.9008 | 265 | 246.0886 | 222.0886 | 387.8565 | 180.1878 | 57.3649 | 80.0718 | 180.1878 | 133.4031 | 132.0275 | 118.9557 | 130.884 | 113.3481 | 118.3481 | 115.7237 | 145.4473 | 72.884 | 68.8397 | 74 | 81.8733 | 103.8122 | 100.2321 | 111.4198 | 83.7679 | 238.287 | 227.3481 | 228.945 | 520.4198 | 126.6244 | 201.3756 | 208.7237 | 411.1053 | 291.9557 | 226.2763 | 230.8565 | 396.9114
1720 | 20.9144 | 169.3273 | 216.8937 | 224.2787 | 216.8451 | 158.5885 | 145.0266 | 157.4867 | 190.7654 | 167.4601 | 157.0752 | 140.9734 | 170.615 | 224.9734 | 149.5399 | 153.8185 | 232.1017 | 178.1283 | 155.6416 | 172.2787 | 194.6196 | 159.2301 | 294.8983 | 313.8231 | 176.8717 | 222.5399 | 209.5664 | 213.385 | 373.1283 | 111.5885 | 112.5133 | 109.5619 | 152.9469 | 136.2301 | 147 | 140.7168 | 105.1017 | 281.7122 | 226.3318 | 221.3584 | 287.5353 | 117.0266 | 119.6416 | 138.5619 | 123.0797 | 76.9734 | 65.3584 | 70 | 106.2566 | 209.0752 | 196.2566 | 228.6902 | 164.2832 | 138.8983 | 123.6948 | 118.1549 | 146.4601 | 162.5399 | 175.5353 | 214.9955 | 170.4867 | 156.5353 | 147.6636 | 187 | 148.3098 | 156.9028 | 188.1017 | 204.0266 | 274.5087 | 151.9203 | 145.5399 | 193.7699 | 187.2832 | 184.1283 | 206.407 | 202.5087 | 402.814 | 144.8717 | 148.8717 | 158.5619 | 260.792 | 131.2346 | 144.2301 | 139.4115 | 166.2301 | 177.7699 | 157.4381 | 173.2566 | 140.6636 | 100.8451 | 87.5664 | 127.0752 | 119.1017 | 195.5178 | 195.7434 | 226.5399 | 369.3052 | 261.2787 | 265.3364 | 235.9469 | 398.792 | 171.4336 | 81.8451 | 75.1815 | 171.4336 | 108.5664 | 119.0531 | 112.7965 | 129.2301 | 113.5619 | 121.8983 | 107.0266 | 124.3318 | 71.8717 | 62.3098 | 72.5885 | 117.0752 | 110.6416 | 102.1283 | 106.7434 | 82.8983 | 211.4115 | 229.2301 | 254.6371 | 513.4336 | 139.6416 | 186.3052 | 199.8983 | 441.4024 | 279.6371 | 240.3629 | 241.022 | 381.6416
1721 | 20.9254 | 127.9372 | 192.7965 | 201.9529 | 210.9686 | 149 | 153.2655 | 151.469 | 158.9065 | 155.8751 | 141.6096 | 133.7188 | 169.2655 | 218.8437 | 160.5625 | 142.2341 | 228.3904 | 179 | 160.2812 | 143.9065 | 221.9222 | 154.8437 | 301.9843 | 331.0314 | 172.4847 | 228.7816 | 226.0157 | 216.4218 | 370.7659 | 104.6716 | 115.7188 | 94.4061 | 138.7188 | 131.9843 | 145.5939 | 131.7188 | 99.2812 | 237.172 | 216.828 | 217.4218 | 264.5625 | 124.2812 | 124 | 123.1249 | 140.7659 | 71.828 | 64.5153 | 68.3126 | 109.5467 | 194.4533 | 192.3755 | 211.4533 | 169.9065 | 144.7188 | 139.3284 | 127.6874 | 132.3441 | 172.8594 | 155.3747 | 186.0621 | 169.3904 | 136.2341 | 129.6402 | 187.4218 | 163.4375 | 195.172 | 183.4061 | 212.6874 | 248.6402 | 131 | 156.4218 | 198.7188 | 196.4218 | 184.0157 | 186.6402 | 175.9372 | 362.1555 | 146.3904 | 148.4218 | 144.1092 | 238.0157 | 156.2035 | 139.8437 | 149.4218 | 159.4533 | 180.6096 | 174.7031 | 178.3747 | 121.6559 | 93.7031 | 105.8437 | 111.8908 | 113 | 230.3284 | 192.5939 | 236.5782 | 351.4218 | 240.359 | 283.641 | 223.5467 | 379.9529 | 155.4061 | 72.5939 | 88.3126 | 155.4061 | 126.2812 | 131.8751 | 125.9843 | 125.6874 | 96.2969 | 126.5939 | 112.7345 | 114.1249 | 70.5782 | 77.5782 | 62.2969 | 103.2969 | 113.4533 | 103.9843 | 105 | 89.5625 | 221.2812 | 225.4061 | 223.5939 | 496.4218 | 142.5939 | 173.4839 | 205.7188 | 384.4061 | 255.7651 | 265.7816 | 214.2812 | 390.4996
1722 | 20.9364 | 150.9297 | 214.1652 | 221.6819 | 222.2355 | 150.3761 | 159.471 | 142.8807 | 140.9174 | 149.9174 | 128.0703 | 127.4129 | 174.7064 | 224 | 166.4465 | 161.0826 | 243 | 178.6942 | 157.26 | 126.2232 | 196.0826 | 157.2477 | 313.1987 | 320.3761 | 153.2936 | 205.7523 | 219.2355 | 215.0245 | 349.6239 | 119.0123 | 115.529 | 104.529 | 137 | 139.9877 | 136.3884 | 132.2936 | 100.3884 | 226.1529 | 228.4587 | 222.9052 | 264.0245 | 127.2232 | 125.9877 | 130 | 118.0948 | 82.6942 | 86.1529 | 59.2232 | 115.4832 | 184.0826 | 162.8929 | 204.1407 | 152.9877 | 140.5535 | 125.4465 | 137.6942 | 114.2813 | 175.0581 | 173.2477 | 212.5658 | 182.6239 | 153.4006 | 155.7064 | 186.789 | 155.5658 | 184.4587 | 191.0826 | 221.7768 | 270.2722 | 137.422 | 158.0826 | 201.1284 | 197.471 | 178.1529 | 217.9052 | 197.4006 | 399.6148 | 155.8013 | 150.2355 | 154.5535 | 233.8348 | 133 | 142.7064 | 150.9297 | 151.3761 | 168.9052 | 177.471 | 195.0245 | 142.0123 | 99.3761 | 110.8471 | 99.1529 | 110.8593 | 211.8593 | 187.9755 | 229.5658 | 355.6819 | 261.4955 | 260.6697 | 232.3884 | 395.2477 | 162.9297 | 63.2355 | 77.5413 | 162.9297 | 127.3884 | 123.3181 | 132.1529 | 131.5658 | 94.9052 | 117.3884 | 103.0123 | 121.4587 | 69.0703 | 74.8471 | 60.4465 | 99.9174 | 102.9297 | 111.0703 | 105.7645 | 91.7768 | 223.9174 | 234.6116 | 212.4006 | 507.7155 | 132.6239 | 203.4832 | 204.3058 | 396.211 | 288.4006 | 237.7064 | 218.4465 | 410.8135
1723 | 20.9474 | 146.8171 | 204 | 227.2005 | 212.3834 | 159.8171 | 154.7995 | 178.6873 | 148.1475 | 155.9912 | 135.6519 | 105.9647 | 169.4338 | 219.3745 | 174.3834 | 151.2093 | 245.3127 | 178.1564 | 131.6608 | 135.4779 | 190.8348 | 140.3481 | 340.0177 | 325.0265 | 165.5044 | 216.0442 | 223.0945 | 201.5901 | 339.1917 | 106.3481 | 123.0088 | 114.9823 | 135.8436 | 153.3127 | 131.5132 | 146.8171 | 96.6696 | 227.9912 | 233.1475 | 236.6873 | 253.3984 | 129.5397 | 135.0177 | 129.8348 | 102.826 | 80.0088 | 83.8613 | 67.1564 | 99.6343 | 181.9735 | 187.8525 | 216 | 163.174 | 126.0088 | 137.5132 | 135.6696 | 135.5221 | 190.0177 | 160.8083 | 188.4956 | 177.3127 | 137.8436 | 145.9735 | 174.947 | 132.8171 | 184.1917 | 182.8613 | 210.2093 | 237.9558 | 170.8525 | 152.8348 | 224.826 | 195.2774 | 180.9823 | 228.2182 | 185.4779 | 342.0681 | 128.9823 | 147.1564 | 142.1564 | 242.6784 | 131.1829 | 127.1917 | 143.8436 | 158.8348 | 184.5044 | 169.3304 | 179.9382 | 131.6608 | 107 | 110.1652 | 100.6608 | 104.6166 | 197.8525 | 209.6696 | 204.174 | 356.2447 | 228.6873 | 275.7138 | 233.4603 | 387.599 | 161.6255 | 61.9735 | 74.3392 | 161.6255 | 122.1829 | 116.9735 | 133.3304 | 110.4691 | 110.0088 | 112.8436 | 97.6166 | 125.3216 | 73.3481 | 74.1652 | 72.3481 | 103.5132 | 96.0088 | 114.3569 | 110.1652 | 83.8436 | 236.2686 | 232.5486 | 206.2597 | 537.9938 | 126.6519 | 183.8348 | 202.3657 | 410.3657 | 271.5221 | 224.8348 | 230.0177 | 375.4603
1724 | 20.9584 | 159.9047 | 200.8053 | 207.8225 | 184 | 164.2079 | 172.5801 | 167 | 159.7099 | 159.0477 | 138.4979 | 122.4503 | 202.2079 | 198.1298 | 150.4371 | 130.9047 | 251.4979 | 182.2251 | 137.8398 | 143.0304 | 194.9696 | 138.3895 | 327.8702 | 316.567 | 164.4199 | 192.7749 | 257.9655 | 240.2079 | 324.645 | 99.2424 | 113.3854 | 118.788 | 128.9351 | 163.0477 | 122.9351 | 152.4503 | 90.5629 | 231.5801 | 238.3205 | 223.7576 | 307.0953 | 109.8398 | 122.6927 | 128.4675 | 111.7921 | 78.1947 | 69.2424 | 71.2251 | 118.8874 | 191.4979 | 175.4026 | 214.4026 | 158.1126 | 122.4199 | 126.0953 | 132.9351 | 126.0173 | 176.8053 | 173.5801 | 191.1775 | 182.078 | 132.7099 | 153.1906 | 199.143 | 143.2424 | 170.8874 | 167 | 190.7921 | 263 | 157.1602 | 152.355 | 224.9178 | 223.8053 | 188.8702 | 201.1298 | 196.4026 | 395.5152 | 139.7099 | 148.7403 | 149.5974 | 236.7099 | 124.1298 | 116.7272 | 140.6623 | 161.0173 | 179.6927 | 174.0173 | 208.433 | 134.1126 | 105.7576 | 107.9827 | 101.6927 | 129.8053 | 191.7921 | 210.1298 | 203.7272 | 322.2942 | 222.5021 | 250.4199 | 253.5152 | 416.5932 | 179.6754 | 76.645 | 69.5801 | 179.6754 | 113.355 | 128.8053 | 132.8702 | 128.5325 | 102.3377 | 108.0649 | 122.4503 | 127.2079 | 65.355 | 74.8225 | 66.4848 | 93.8702 | 92.0649 | 102.5974 | 110.1126 | 75.8702 | 268.5629 | 210.6795 | 237.1734 | 455.3418 | 137.1298 | 179.0953 | 179.5629 | 403.0041 | 261.1298 | 224.5325 | 216.9827 | 404.3895
1725 | 20.9694 | 173.2556 | 191.3138 | 210.416 | 185.8978 | 142 | 164.4818 | 161.3067 | 164.8978 | 152.2773 | 117.1751 | 111.124 | 180.3796 | 203.4454 | 170.6204 | 153.7444 | 228.6569 | 175.708 | 149.4818 | 125.3138 | 222.1751 | 164.7738 | 318.7591 | 343.7738 | 169.292 | 201.9489 | 224.3067 | 223.6933 | 322.6204 | 108.0364 | 95.0364 | 95.5035 | 123.0511 | 154.9489 | 119.3284 | 136.3796 | 77.4013 | 223.6716 | 214.6569 | 218.3796 | 291.4671 | 122.0511 | 118.2626 | 126.1898 | 131.1533 | 89.584 | 71.2045 | 64.3284 | 121.8613 | 166 | 168.1898 | 208.8978 | 156.8467 | 128.0511 | 110.0875 | 125.3431 | 138.292 | 166.5546 | 162.4454 | 190.6716 | 153.6058 | 138.0875 | 120.416 | 173.3284 | 147.1022 | 176.3284 | 165.1022 | 211.1022 | 269.6422 | 143.8102 | 153.2409 | 208.8467 | 205.3942 | 182.416 | 201.5112 | 191.4671 | 383.8102 | 146.0364 | 137.1604 | 158.8978 | 232.0294 | 130.2045 | 134.7591 | 150.3431 | 171.9636 | 170.3284 | 176.0441 | 180.8467 | 128.292 | 99.6204 | 98.1751 | 90.6204 | 113.8613 | 211.1533 | 219.0511 | 220.0511 | 368.4818 | 245.5329 | 253.9636 | 237.0658 | 377.1751 | 157.6569 | 73.8613 | 64.3284 | 157.6569 | 115.7591 | 116.2773 | 123.5693 | 131.1898 | 92.8467 | 113 | 104.8613 | 110.124 | 65.8613 | 72.6716 | 75.5329 | 88.3649 | 97.1898 | 107.5329 | 106.9489 | 66.3796 | 244.2045 | 233.7738 | 190.2262 | 487.3578 | 143.0147 | 162.1387 | 163.3649 | 448.3067 | 269.8613 | 221.3067 | 206.7955 | 427.168
1726 | 20.9804 | 154.2124 | 214.8083 | 222.9793 | 192.3835 | 136.7462 | 159.2124 | 140.8393 | 172.5959 | 160.3835 | 133.1814 | 131.7876 | 180.5855 | 189.2538 | 166.171 | 167.1297 | 237.9793 | 183.6062 | 142.1917 | 143.1504 | 235.1504 | 151.8083 | 321.5959 | 329.5959 | 166.0414 | 210.4455 | 205.2538 | 241.5338 | 322.6165 | 118.5752 | 109.2124 | 116.5752 | 121.4248 | 155.3628 | 126.2124 | 135.7773 | 106.3731 | 219.0103 | 226.6062 | 229.2951 | 293.9173 | 122.0414 | 139.9483 | 127.2021 | 118.1917 | 75.8083 | 58.031 | 66.5648 | 116.1917 | 173.0724 | 178.2951 | 215.3835 | 172.4352 | 122.7876 | 117.3835 | 117.2331 | 124.938 | 180.14 | 146.6372 | 188.8393 | 167.9897 | 145.5855 | 130.1504 | 176.171 | 141.6269 | 183.6165 | 157.8083 | 203.4041 | 287.2904 | 142.5959 | 144.7462 | 213.9276 | 192.2227 | 200.4352 | 165.4455 | 201.1917 | 384.0103 | 157.9897 | 166.7462 | 164.3731 | 201.8393 | 115.4145 | 130.1193 | 141.4248 | 159.8083 | 169.7359 | 135.1034 | 197.0414 | 121.6062 | 98.6062 | 116.8083 | 90.6165 | 111.829 | 202.2331 | 210.9586 | 217.4145 | 361.1917 | 228.938 | 241.4041 | 218.829 | 401.2641 | 165.5648 | 71.4248 | 69.1917 | 165.5648 | 121.2227 | 126 | 127.4145 | 129.9793 | 103.7876 | 108.7566 | 104.6476 | 131.1917 | 64.8393 | 69.4248 | 65.2021 | 103.5648 | 97.3938 | 98.2124 | 109.3835 | 69.6165 | 236.0827 | 217.171 | 210.4662 | 450.7152 | 129.4352 | 169.829 | 178.969 | 421.7773 | 265.4041 | 200.8393 | 220.171 | 394.1607
1727 | 20.9914 | 160.0718 | 214.3559 | 212.2133 | 187.2862 | 124.1456 | 162.0708 | 154.4995 | 167.3559 | 157.001 | 123.4277 | 131.0728 | 174.7856 | 210.6431 | 155.2144 | 146.8595 | 236.0021 | 183.4277 | 136.2133 | 131.2882 | 212.7846 | 156.7149 | 322.1436 | 327.1426 | 177.2841 | 227.3569 | 219.5692 | 213.6421 | 328.3569 | 107.0708 | 107.7836 | 105.9282 | 139.5744 | 141.6431 | 130.7856 | 127.6431 | 92.3559 | 226.6441 | 227.4995 | 262.1415 | 277.6482 | 133.0697 | 121.2862 | 127.3569 | 112.8564 | 77.7138 | 71.7149 | 58.1446 | 111.2851 | 197.3559 | 214.1426 | 213.7159 | 176.5682 | 123.5733 | 112.0718 | 124.4267 | 105.2872 | 156.1426 | 158.2133 | 197.9979 | 159.7128 | 143.8585 | 118.0738 | 168.8585 | 153.2862 | 195.359 | 163.3579 | 206.7149 | 250.7877 | 145.2872 | 126.7867 | 193.2882 | 201.4287 | 213.3569 | 185.5723 | 195.6421 | 382.641 | 152.4995 | 141.7128 | 152.9282 | 222.359 | 119.4995 | 103.7159 | 145.8554 | 165.3579 | 152.2882 | 169.8554 | 208.0697 | 128.0728 | 99.0708 | 118.7138 | 96.1426 | 125.3579 | 212.6421 | 195.6431 | 225.7867 | 356.9282 | 209.5015 | 247.9303 | 228.4995 | 408.4226 | 158.0021 | 80.1426 | 67.9292 | 158.0021 | 130.8574 | 124.4995 | 126.9979 | 124.3579 | 100.5005 | 98.6451 | 124.9292 | 125.2133 | 79.7856 | 76.6421 | 65.1426 | 94.5015 | 96.7149 | 101.4995 | 105.1436 | 75.3569 | 260.9262 | 202.3559 | 226.7826 | 439.0072 | 137.8554 | 174.3549 | 170.001 | 409.999 | 258.2113 | 210.2123 | 210.9292 | 388.2174
1728 | 21.0024 | 160.3736 | 202.5866 | 201.32 | 193.8132 | 153.32 | 157.0399 | 154.6663 | 157.6264 | 168.68 | 118.3064 | 143.5068 | 170.3736 | 212.32 | 159.6264 | 167.9203 | 249.3873 | 174.2267 | 129.1731 | 154.1469 | 199.6401 | 162.0934 | 315.9477 | 330.1196 | 160.68 | 226.1333 | 201.6663 | 210.4396 | 326 | 101.3599 | 94.2927 | 102 | 162.6139 | 141.5068 | 128.4134 | 124.1071 | 81.0399 | 237.7335 | 225.8531 | 250.7198 | 327.5205 | 120.8929 | 126.4533 | 119.3337 | 107.2665 | 75.0399 | 78 | 78.1333 | 104.32 | 188.533 | 205.3337 | 226.467 | 146.9863 | 139.467 | 117.1333 | 106.4932 | 123.7198 | 155.2665 | 148.2267 | 178.9465 | 145.3599 | 156.6401 | 144.4533 | 180.5068 | 156.4134 | 207.9738 | 173.3599 | 206.2005 | 269.2802 | 154.6538 | 136.0934 | 219.7733 | 205.9465 | 205.787 | 194.7733 | 185.8667 | 365.0399 | 142.0137 | 126.9066 | 146.5068 | 239.5068 | 123.2927 | 123.2665 | 131.3064 | 171.0535 | 170.8406 | 152.1333 | 192.7198 | 138.0137 | 94.7198 | 111.7335 | 99.1196 | 129.2403 | 202.8667 | 200.4932 | 231.467 | 351.4134 | 220.3337 | 254.6675 | 231.1594 | 352.3861 | 178.1868 | 73.8269 | 75.2267 | 178.1868 | 134.4533 | 113.5604 | 108.1731 | 133.2267 | 108.7198 | 118.9203 | 122.7073 | 112.0535 | 73.787 | 71.3998 | 64.7198 | 113.4932 | 103.9066 | 94.1868 | 112.32 | 71.6401 | 239.9863 | 197.8394 | 204.1059 | 504.6675 | 122.8531 | 157.0262 | 181.4533 | 397.8667 | 236.4259 | 187.6401 | 219.1333 | 416.1071
1729 | 21.0134 | 153.0178 | 207.7611 | 196.6106 | 200.9558 | 153.8673 | 162.3274 | 172.3274 | 171.4336 | 165.5044 | 127.2654 | 134.0442 | 158.7168 | 213.5841 | 168.8938 | 159.0178 | 224.3009 | 168.3096 | 141.1947 | 144.8231 | 200.0178 | 153.5044 | 269.4336 | 339.2919 | 165.1505 | 222.8318 | 217.4159 | 223.7433 | 327.6726 | 106.7168 | 120.5044 | 105.3452 | 132.0885 | 133.9558 | 122.0442 | 112.8408 | 91.3452 | 224.7433 | 233.2654 | 253.0707 | 299.8495 | 137.6548 | 130.8673 | 102.6284 | 117.1505 | 80.5664 | 75.8495 | 79.3716 | 107.9735 | 200.1769 | 184.0885 | 211.7168 | 163.2832 | 128.062 | 117.4159 | 115 | 128.6991 | 169.2125 | 151.8673 | 187.5664 | 156.9293 | 148.6548 | 146.2389 | 172.4779 | 154.8231 | 185.9115 | 177.0442 | 183 | 264.5399 | 138.7346 | 133.2389 | 213.0265 | 215.5221 | 192.7788 | 189.6991 | 187.4956 | 370.8053 | 132.4072 | 133.3452 | 141.1062 | 238.1682 | 149.2654 | 125.9822 | 147.1947 | 161.7168 | 147.062 | 152.8938 | 202 | 124.5841 | 105.9115 | 102.3274 | 111.1592 | 110.1327 | 196.8495 | 209.7168 | 211.4601 | 348.6284 | 204.1062 | 223.8851 | 253.2654 | 370.7255 | 166.7433 | 62.5221 | 77.6726 | 166.7433 | 134.5664 | 101.2125 | 117.8053 | 131.3716 | 117.7611 | 104.5221 | 99.3629 | 102.7168 | 56.7168 | 81.0885 | 72.5664 | 113.6371 | 106.5221 | 89.6726 | 109.7611 | 72.0178 | 260.1062 | 225.3716 | 240.5664 | 468.1505 | 133.6106 | 181.7168 | 188.9735 | 400.6904 | 265.7875 | 184.9115 | 215.8318 | 389.5486
1730 | 21.0244 | 165.4827 | 201.2217 | 189.2512 | 202.4926 | 159.734 | 154.2364 | 163.734 | 175.4975 | 145.9901 | 121.3104 | 130.4975 | 162.2562 | 222.7389 | 151.463 | 172.7389 | 237.2463 | 155.7931 | 142.4877 | 153.2315 | 209.468 | 141.0246 | 277.7685 | 312.803 | 179.0345 | 199.4975 | 206 | 217.5518 | 331.2414 | 112.266 | 101.7438 | 114.9951 | 129.266 | 136.2463 | 122.0148 | 140.2463 | 98.2315 | 214.2808 | 209.9754 | 232.7586 | 288.2266 | 128.7586 | 135.2266 | 117.0296 | 123.4975 | 74.2414 | 68.4975 | 71.2512 | 125.7438 | 185.9901 | 172.2217 | 218.2709 | 160.2463 | 139.4926 | 113.2857 | 112.4877 | 115.2512 | 182.9704 | 159.2414 | 182.7488 | 176.7389 | 137.4975 | 141.4729 | 175.7586 | 167.2512 | 192.5025 | 166.463 | 187.2709 | 269.4532 | 157.2512 | 137.266 | 202.2907 | 214.7537 | 205.9901 | 181.7783 | 206.7586 | 371.0197 | 153.2364 | 141.4877 | 147.9852 | 256.4778 | 129.4926 | 117.537 | 148.7389 | 162.4926 | 158.2414 | 149.7833 | 200.7438 | 136.7537 | 112.7537 | 97.5025 | 83.5025 | 101 | 193.7685 | 215.5173 | 195.7438 | 354.7389 | 209.9803 | 245.9852 | 241.7833 | 344.2808 | 150.5025 | 63.7636 | 80.4877 | 150.5025 | 124.4729 | 118.9901 | 117.7685 | 128.5271 | 114.2364 | 101.9951 | 119.7192 | 109.7734 | 59 | 76.2562 | 69.5074 | 92.5271 | 104.2463 | 93.4926 | 105.2315 | 88.2512 | 262.7143 | 215.7438 | 230.7241 | 473.4926 | 125.2463 | 180.734 | 198.2019 | 429.0345 | 243.4877 | 183.9655 | 198.2759 | 367.4926
1731 | 21.0354 | 159.7434 | 189.7975 | 189.2094 | 199.8447 | 149.2094 | 142.8377 | 158.2164 | 172.9459 | 147.3787 | 144.0333 | 126.8918 | 165.7365 | 215.473 | 140.3315 | 166.527 | 230.7836 | 179.3648 | 140.534 | 141.473 | 191.3176 | 147.8308 | 291.6352 | 340.2566 | 192.0943 | 200.6352 | 207.3176 | 239.2496 | 326 | 125.1623 | 108.8045 | 111.2094 | 140.8447 | 135.5811 | 128.8918 | 135.3648 | 89.1082 | 223.5132 | 194.7365 | 244.5881 | 272.1553 | 127.1484 | 123.3717 | 133.6824 | 117.5201 | 70.8447 | 67.2635 | 69.3648 | 117.5201 | 183.1623 | 160.7975 | 223.3578 | 156.6824 | 130.7836 | 126.8308 | 110.534 | 116.527 | 161.5742 | 149.2566 | 179.8918 | 168.1553 | 138.3717 | 123.4189 | 175.9931 | 164.0472 | 188.2025 | 149.7975 | 201.5811 | 244.9528 | 156.4189 | 144.1012 | 225.6824 | 213.3107 | 194.9931 | 195.5742 | 204.8849 | 386.6893 | 144.2094 | 134.7906 | 138.7365 | 239.0472 | 126.0541 | 139.7365 | 135.4119 | 156.9459 | 159.5881 | 168.4189 | 197.2635 | 135.3107 | 110.7836 | 100.0541 | 93.1692 | 102.3176 | 199.9931 | 223.8918 | 204.6491 | 341.4119 | 207.2233 | 232.2566 | 254.6213 | 362.7365 | 158.5881 | 69.8918 | 69.5742 | 158.5881 | 109.8447 | 116.6893 | 130.0541 | 139.7295 | 108.1082 | 98.2094 | 102.3176 | 121.3648 | 62.4258 | 76.0472 | 75.3176 | 107.6824 | 102 | 93.2164 | 95.8447 | 86.1012 | 247.061 | 215.9528 | 218.9528 | 474.534 | 124.5811 | 173.6352 | 177.1692 | 445.2566 | 239.6893 | 170.3787 | 210.0472 | 372.7504
1732 | 21.0464 | 166.3805 | 207.1033 | 192.5163 | 196.4498 | 151.6888 | 135.6549 | 173.3098 | 171.1033 | 167.4484 | 108.8614 | 122.6549 | 167.4823 | 212.6209 | 165.9321 | 163.8628 | 226.447 | 175.034 | 153.7935 | 138.0693 | 194.1726 | 132.9986 | 302.8614 | 314.8642 | 170.5516 | 205.2419 | 216.7921 | 210.3451 | 330.413 | 134.2758 | 129.3465 | 113.413 | 139.1046 | 140.5516 | 124.9307 | 136.2744 | 97.2065 | 198.1033 | 199.7921 | 253.6223 | 267.8274 | 115.447 | 124.2079 | 130.2758 | 110.2405 | 77.1033 | 69.1033 | 65.034 | 107.2065 | 190.3451 | 179.7581 | 203.9307 | 154.1033 | 118.7242 | 105.9307 | 120.7595 | 130.1358 | 154.2065 | 141.7921 | 180.0679 | 167.1372 | 144.1726 | 119.2758 | 163.1033 | 148.8628 | 177.7921 | 166 | 206.8274 | 259.8614 | 153.6549 | 133.5177 | 216.2079 | 206.3098 | 182.3791 | 186 | 186.9307 | 386.2432 | 146.413 | 135.8967 | 138.8274 | 241.2391 | 131.4823 | 138.4484 | 118.9307 | 151.5177 | 176.8967 | 168.1372 | 204.3437 | 128.034 | 98.7242 | 103.5516 | 109.9321 | 109.3098 | 192.3437 | 214.966 | 232.2772 | 318.8628 | 225.553 | 220.3791 | 241.9986 | 346.2786 | 168.4498 | 65.3791 | 63.5856 | 168.4498 | 118.034 | 118.1739 | 125.2772 | 132.1712 | 117.8614 | 96.8274 | 103.2419 | 118.413 | 73.1033 | 66.1033 | 77.3451 | 106.4823 | 104.4823 | 97.8288 | 107.3437 | 78.8274 | 260.1386 | 219.553 | 222.2772 | 494.1372 | 134.7921 | 179.0693 | 196.1386 | 433.1033 | 252.2065 | 182.1739 | 207.8261 | 392.8288
1733 | 21.0574 | 156.4907 | 203.0855 | 211.8978 | 183.7956 | 156.5093 | 148.3551 | 172.4925 | 168.2379 | 160.2379 | 121.0167 | 124.9833 | 182.3551 | 209 | 150.8643 | 155.3048 | 239.3736 | 177.8141 | 143.3903 | 137.0335 | 194.3048 | 153.7119 | 307.526 | 284.1524 | 171.7119 | 205.2026 | 227.3903 | 204.5595 | 339.4071 | 128.6617 | 109.7621 | 120.3903 | 127.6264 | 143.7286 | 132.0167 | 159.1524 | 100.4071 | 207.3383 | 215.5762 | 230.1524 | 274.8978 | 130.3903 | 113.3216 | 129.8476 | 125.7286 | 79.4238 | 69.4071 | 72.4238 | 117.6097 | 189.4573 | 185.5595 | 211.5929 | 157.2881 | 114.8308 | 117.9145 | 110.7454 | 147.3066 | 160.8643 | 154.1189 | 198.3048 | 174.5427 | 144.0167 | 123.4573 | 166.8643 | 139.1524 | 181.4758 | 160.2379 | 208.1357 | 267.4071 | 153.1022 | 130.1692 | 204.4573 | 215.2881 | 197.5242 | 193.7788 | 195.1692 | 374.1339 | 155.6952 | 139.3383 | 150.7956 | 272.2546 | 132.2379 | 134.4071 | 128.3216 | 139.526 | 176.5929 | 170.6449 | 208.0353 | 133.1189 | 98.2881 | 106.4405 | 100.6264 | 121.4573 | 207.8476 | 210.7454 | 213.1524 | 319.5242 | 207.1524 | 222.2714 | 251.7639 | 313.067 | 153.7788 | 73.6097 | 71.8476 | 153.7788 | 130.3216 | 94.7119 | 102.9833 | 141.7806 | 121.9498 | 99 | 100.8978 | 123.9498 | 77.4405 | 68.7119 | 75.3048 | 108.1022 | 110.7119 | 82.1692 | 114.2044 | 76.1022 | 245.3903 | 206.9145 | 206.0335 | 497.5093 | 143.9498 | 174 | 187.7286 | 421.5947 | 250.5093 | 168.5075 | 223.7974 | 376.3048
1734 | 21.0684 | 169.8933 | 192.0079 | 220.996 | 210.7036 | 142.5059 | 161.4901 | 141.7905 | 157.6048 | 153.8103 | 126.6008 | 120.9012 | 193.6877 | 207.1976 | 153.3004 | 161.6008 | 235.0159 | 167.7036 | 135.9052 | 147.6008 | 201.2016 | 148.1937 | 293.6996 | 297.8143 | 173.4032 | 224.502 | 212.3952 | 202.2016 | 339.3083 | 121.1107 | 114.6877 | 114.1067 | 141.8972 | 154.3083 | 139.1028 | 150.2845 | 97.3044 | 216.3913 | 218.502 | 227.8933 | 275.5849 | 122.3044 | 122.7984 | 129.1028 | 133.6048 | 78 | 66.004 | 71.9012 | 122.6917 | 189.5889 | 178.996 | 226.7115 | 154.3952 | 117.5139 | 132.8972 | 116.3952 | 114.004 | 161.7984 | 150.6048 | 207.004 | 165.3992 | 140.5889 | 135.004 | 166.8972 | 140.1976 | 167.419 | 156.8143 | 210.8063 | 267.3083 | 143.4032 | 139.502 | 210.2964 | 215.0988 | 221.7944 | 209.6956 | 204.2016 | 408.5929 | 150.6008 | 149.8933 | 168.0909 | 254.5889 | 122.5059 | 131.004 | 130.2885 | 135.6127 | 181.1976 | 146.3952 | 182.0079 | 131.7075 | 101.4032 | 111.502 | 107.9881 | 130.6008 | 197.1897 | 211.2885 | 210.8933 | 355.5099 | 211.502 | 211.6917 | 231.8103 | 326.0909 | 174.8024 | 82.8972 | 71.1028 | 174.8024 | 133.7905 | 96.1028 | 102.8063 | 130.1147 | 108.502 | 99.3004 | 114.502 | 112.9052 | 77.0948 | 67.6996 | 74.6917 | 96.9012 | 109.6996 | 96.0079 | 89.0988 | 67.9052 | 240.9091 | 213.7865 | 215.6996 | 480.8024 | 135.3083 | 175.502 | 184.1897 | 395.0119 | 234.1028 | 194.7036 | 205.3992 | 380.498
1735 | 21.0794 | 158.3127 | 200.0589 | 211.1858 | 210.1209 | 156.1269 | 140.5604 | 133.4426 | 162.2477 | 168.2477 | 123.9351 | 122.3746 | 174.1888 | 213.3187 | 155.5634 | 158.3097 | 266.4396 | 177.1269 | 147.8142 | 143.0589 | 201.1858 | 143.5664 | 297.0649 | 320.8112 | 182.4396 | 229.5634 | 213.0679 | 211.8791 | 338.991 | 131.8701 | 101.4426 | 118.8082 | 149.8202 | 164.3097 | 135.2447 | 124.2508 | 102.4985 | 206.9411 | 224.5015 | 213.4985 | 261.1948 | 127.8112 | 114.3716 | 144.6314 | 138.5604 | 70.4955 | 70.8112 | 69.6224 | 109.8761 | 177.8172 | 166.997 | 238.8701 | 144.4366 | 136.6843 | 128.6254 | 113.9411 | 119.4366 | 157.4366 | 159 | 203.0559 | 165.2508 | 121.9381 | 139.8112 | 162.9381 | 133.4985 | 192.8701 | 176.997 | 211.0559 | 271.6813 | 149 | 142.6873 | 210.2538 | 213.6254 | 209.8112 | 196.997 | 202.9351 | 398.1888 | 149.8112 | 135.4985 | 152.3127 | 227.1828 | 132.3746 | 142.0649 | 120.006 | 147.994 | 177.9381 | 146.4426 | 191.3097 | 136.5574 | 106.6873 | 111.2477 | 90.1269 | 125.7462 | 191.5075 | 195.6903 | 200.5634 | 358.6164 | 216.8761 | 201.0649 | 245.9351 | 310.6254 | 180.5634 | 73.6224 | 75.6873 | 180.5634 | 124.8172 | 110.0679 | 114.6254 | 145.994 | 105.4336 | 101.2508 | 110.8082 | 116.6843 | 74.2538 | 68.8761 | 60.3127 | 98.0619 | 114.3157 | 102.8082 | 92.6284 | 79.5015 | 248.2447 | 205.3217 | 214.0619 | 495.3187 | 145.3097 | 177.4366 | 173.8172 | 420.6903 | 248.0679 | 202.2508 | 207.4396 | 374.4985
1736 | 21.0904 | 153.1504 | 188.9499 | 198.8254 | 201.8245 | 161.0501 | 138.1997 | 154.9248 | 151.0752 | 162.2748 | 113.7001 | 132.3742 | 174.7753 | 230.8003 | 156.7252 | 147.675 | 266.2005 | 184.65 | 150.4252 | 133.2498 | 195.9749 | 160.2748 | 309.8997 | 310.4754 | 183.8254 | 234.6249 | 228.7753 | 222.4503 | 311.5246 | 116.2748 | 116.1003 | 102.0251 | 173.1253 | 149.1253 | 121.1997 | 131.2247 | 94.7252 | 216.7502 | 223.8254 | 209.6249 | 291.2498 | 123.65 | 112.0743 | 150.4512 | 132.2999 | 66.6742 | 65.35 | 64.2247 | 120.4996 | 195.9749 | 170.624 | 222.6249 | 147.1746 | 121.1003 | 123.5004 | 126.0251 | 116.325 | 149.7753 | 164.1997 | 179.325 | 165.0752 | 133.7494 | 137.9248 | 158.5004 | 137.0994 | 173.3751 | 165.675 | 193.4996 | 248.0251 | 143.4754 | 130.9507 | 215.8755 | 215 | 205.65 | 182.1003 | 189.4503 | 404.2999 | 148.8997 | 134.5497 | 155.5998 | 221.0493 | 134.2498 | 140.2757 | 135.1755 | 130.8747 | 177.0752 | 169.8747 | 183.2748 | 116.3751 | 107.6249 | 95.2005 | 100.2498 | 116.2247 | 211.5255 | 197.5506 | 206.2748 | 337.7243 | 213.5255 | 212.9248 | 239.5998 | 321.7494 | 178.4754 | 64.65 | 72.7252 | 178.4754 | 134.2005 | 127.4002 | 112.1003 | 127.5748 | 95.2247 | 110.1746 | 106.3742 | 112.1495 | 86.7001 | 74.9499 | 64.8997 | 96.65 | 120.8003 | 86.0251 | 103.0501 | 84.35 | 234.5246 | 230.2256 | 216.5497 | 512.8003 | 134.675 | 176.5998 | 183.5255 | 421.9006 | 266.0501 | 204.0251 | 202.0009 | 385.5739
1737 | 21.1014 | 143.6982 | 196.3727 | 193.0827 | 210.6037 | 148.5683 | 149.3373 | 163.0236 | 152.4199 | 164.3018 | 116.3963 | 140.8229 | 156.5565 | 211.231 | 154.0236 | 157.1181 | 246.29 | 190.3845 | 145.6982 | 143.71 | 196.3137 | 160.9528 | 327.4435 | 295.29 | 169.3137 | 218.7756 | 226.0709 | 201.5328 | 327.3845 | 119.9882 | 110.3609 | 103.3727 | 156.9291 | 142.3727 | 128.29 | 145.3963 | 91.3491 | 208.6509 | 209.9882 | 205.2428 | 299.0118 | 130.3963 | 126.9291 | 121.0827 | 127.9173 | 88.2782 | 63.3255 | 73.6745 | 127.2546 | 192.941 | 186.8347 | 230.0472 | 153.2546 | 115.6982 | 112.3609 | 134.7926 | 119.0236 | 144.71 | 166.5683 | 178.9882 | 164.3963 | 151.6391 | 136.9173 | 148.0354 | 155.6627 | 169.6627 | 160.2782 | 195.1956 | 250.3845 | 139.0827 | 124.1889 | 206.0827 | 207.2428 | 199.2428 | 186.479 | 185.7336 | 406.3255 | 153.9646 | 144.6745 | 159.6509 | 248.0472 | 141.6745 | 128.7926 | 119.6509 | 147.0709 | 173.6982 | 184.0472 | 190.0236 | 115.3609 | 110.3255 | 90.1299 | 101.941 | 128.3727 | 191.2782 | 186.0709 | 210.3255 | 349.8819 | 197.6627 | 215.9646 | 241.29 | 326.1483 | 167.6745 | 73.7572 | 69.3491 | 167.6745 | 117.6627 | 125.3727 | 108.7218 | 134.6627 | 102.3137 | 116.9291 | 119.8819 | 124.9764 | 73.5446 | 76.3018 | 75.6982 | 96.3137 | 109.3255 | 89.3963 | 101.3609 | 82.6627 | 250.3845 | 210.3137 | 222.6273 | 498.9646 | 128.6037 | 174.9173 | 176.769 | 403.3373 | 269.0827 | 197.9528 | 181.0709 | 407.8583
1738 | 21.1124 | 152.146 | 199.1548 | 212.2434 | 200.6991 | 142.1373 | 155.593 | 155.4645 | 169.7965 | 161.7965 | 133.2921 | 116.9425 | 150.2346 | 197.146 | 156.9513 | 174.2035 | 245.9912 | 194.1062 | 146.8053 | 146.8053 | 209.0798 | 152.3009 | 327.8228 | 291.146 | 166.3496 | 193.2346 | 232.0575 | 188.0886 | 334.6018 | 113.8053 | 115 | 111.3982 | 137.6592 | 147.9513 | 132.8851 | 147.261 | 95.7478 | 228.3719 | 216.0399 | 201.5355 | 277.929 | 136.8053 | 119.1947 | 136.0487 | 120.2921 | 78.6504 | 64.7965 | 76.3982 | 116.4469 | 195.2346 | 156.5531 | 227.1636 | 146.6417 | 127.2921 | 115.6018 | 144.4645 | 125.7965 | 155.4956 | 142.6592 | 177.6991 | 165.5619 | 140.7079 | 126.4956 | 152.9513 | 155.6991 | 170.0487 | 155.5443 | 177.593 | 250.9601 | 146.3583 | 147.261 | 222.7965 | 203.8851 | 183.3009 | 201.2123 | 190.7079 | 410.593 | 148.3496 | 148.0974 | 156.3009 | 251.1062 | 139.1548 | 140.2123 | 119.0974 | 149.2123 | 176.5531 | 178.7167 | 188.407 | 127.6903 | 102.7079 | 110.2522 | 97.593 | 128.3583 | 191.0886 | 197.6504 | 213.8938 | 336.2921 | 203.9912 | 210.3496 | 237.146 | 317.3719 | 164.8053 | 87.9513 | 79.3408 | 164.8053 | 122.593 | 123.6105 | 117.6504 | 132.6018 | 102.8452 | 105.6991 | 99.3009 | 124.4956 | 65.9337 | 75.1947 | 77.5044 | 96.4956 | 112.8938 | 101.3982 | 101.4557 | 88.6417 | 242.9202 | 213.9912 | 213.5531 | 493 | 120.4469 | 164.8452 | 196.4469 | 392.8141 | 268.3185 | 196.9912 | 201.3895 | 384.6991
1739 | 21.1234 | 156.5526 | 184.4011 | 220.5526 | 200.1907 | 155.5722 | 157.3156 | 150.2833 | 172.8289 | 149.993 | 133.0589 | 125.8289 | 165.0393 | 208.4278 | 147.763 | 169 | 254.6578 | 183.1907 | 141.533 | 137.5526 | 221.993 | 160.4081 | 297.6185 | 298.8093 | 173.5133 | 212.7433 | 215.2104 | 207.1711 | 336.7041 | 108.533 | 116.0856 | 117.6185 | 129.4278 | 157.9411 | 146.6774 | 129.5526 | 104.4278 | 247.927 | 222.3156 | 232.4081 | 244.0463 | 133.7041 | 128.0856 | 138 | 123.3156 | 75.8289 | 59.1444 | 75.7433 | 127.8752 | 206.7826 | 165.0267 | 209.1711 | 154.2959 | 134.2959 | 114.8093 | 121.1907 | 126.6578 | 149.3352 | 130.8093 | 182.2567 | 151.9804 | 128.2763 | 125.763 | 158.5989 | 151.2104 | 175.2567 | 160.3815 | 193.7896 | 249.7504 | 137.2567 | 131.3619 | 221.4867 | 225.6381 | 183.0856 | 197.1122 | 186.5989 | 416.2959 | 157.6844 | 152 | 152.467 | 244.1711 | 134.8948 | 123.8093 | 121.9144 | 137.5133 | 168.0196 | 161.5989 | 172.9341 | 138.0196 | 95.7041 | 107 | 108.3619 | 115.2763 | 208.7237 | 202.0659 | 220.8289 | 338.5919 | 212.6578 | 207.0196 | 238.6578 | 349.5919 | 159.533 | 74.7826 | 91.1907 | 159.533 | 123.23 | 113.7896 | 112.6578 | 136.1515 | 102.0393 | 105.9144 | 104.5133 | 122.3156 | 79.1444 | 78.6578 | 72.0856 | 96.8485 | 125.2567 | 109.0659 | 91.5526 | 93.763 | 231.0926 | 223.7433 | 215.1515 | 501.3226 | 118.8485 | 168.7433 | 194.8485 | 390.4544 | 248.237 | 196.9733 | 220.6185 | 372.2496
1740 | 21.1344 | 160.7346 | 183.3467 | 206.0274 | 195.8776 | 146 | 146.4897 | 175.2586 | 177.2311 | 137.9589 | 114.2449 | 120.5034 | 161.476 | 220.9931 | 142.1087 | 173.4897 | 252.6121 | 177.381 | 134.7757 | 138.7414 | 198.7346 | 168.3947 | 305.4966 | 304.619 | 180.8844 | 216.3536 | 209.4897 | 218.1087 | 340.5103 | 118.238 | 116.8776 | 114.6464 | 135.2586 | 174.8844 | 137.9863 | 124.0069 | 108.0137 | 226.5847 | 209.9931 | 247.5034 | 262.9931 | 140.8776 | 130 | 133.5103 | 114.7346 | 67.8844 | 50.476 | 70.3741 | 125.4084 | 180.2654 | 174.5309 | 211.5034 | 154.6053 | 128.2449 | 114.6327 | 118 | 121.2449 | 152.4417 | 138.8639 | 188 | 168.3536 | 132.6121 | 114.4966 | 166.3879 | 146.9863 | 177.6327 | 155.4966 | 191.2791 | 259.4084 | 141.8776 | 132 | 215.6121 | 215.6464 | 189.1156 | 202.9187 | 193.6396 | 393.0343 | 160.6533 | 144.8913 | 139.7689 | 234.5309 | 139.2517 | 128.8707 | 124.8639 | 149 | 166.238 | 159.286 | 175.5984 | 128.7551 | 104.3741 | 113.3604 | 107.1293 | 115.4966 | 206.6327 | 203.8913 | 211.3879 | 324.9931 | 212.4829 | 200 | 233.2449 | 320.2791 | 161.7551 | 58.3673 | 86.5034 | 161.7551 | 107.8707 | 121.381 | 111.3604 | 140.6396 | 95.4829 | 104.7483 | 109.6396 | 107.3741 | 67.1087 | 75.1156 | 70.6327 | 87.4966 | 116.7826 | 113.1361 | 94.238 | 88.857 | 246.1636 | 213.8844 | 220.3879 | 517.4966 | 118.476 | 169.7346 | 184 | 412.1499 | 256.8844 | 181.857 | 213.5309 | 373.0676
1741 | 21.1454 | 163.112 | 203.6136 | 194.5015 | 200.1829 | 146 | 155.1593 | 164.3628 | 187.9085 | 165.6814 | 120.705 | 120.705 | 168.5663 | 220.5899 | 145.112 | 169.407 | 262.3864 | 184.9794 | 129.1388 | 136.4985 | 202.6578 | 155.9322 | 304.5221 | 306.6814 | 166.6578 | 213.0647 | 208.112 | 229.6136 | 338.7965 | 129.3186 | 114.6136 | 105.9558 | 134.7965 | 159.112 | 139.112 | 125.5694 | 99.5457 | 250.9764 | 209.9764 | 231.4748 | 269.9322 | 145.9558 | 131.5457 | 126 | 136.4337 | 62.1593 | 68.7729 | 67.9085 | 105.6372 | 167.0678 | 167.4574 | 213.6372 | 163.9764 | 133.1593 | 107.0678 | 126.1151 | 122.6814 | 179.5426 | 136.9527 | 184.9085 | 175.4985 | 133.8849 | 122.7965 | 158.9322 | 158.1593 | 170.0678 | 155.295 | 176.5694 | 235.3864 | 135.3628 | 130.8407 | 216.4985 | 204.6372 | 195.6136 | 179.1151 | 172.1799 | 373.7965 | 141.5457 | 127.9764 | 134.1151 | 218.9558 | 135.295 | 128.9085 | 136.0915 | 150.9322 | 175.3864 | 135.3186 | 184.9527 | 126.1593 | 102.6814 | 111.6341 | 100.1357 | 123.0236 | 204.8643 | 201.2744 | 205.8643 | 333.4779 | 209.9291 | 198.4543 | 238.9322 | 297.4543 | 158 | 62.4543 | 76.6578 | 158 | 109.0678 | 113.9085 | 112.7256 | 128.0678 | 104.5221 | 112.5694 | 100.5457 | 105.6814 | 74.7492 | 81.6136 | 65.7729 | 84.2035 | 87.5899 | 98.4306 | 103 | 98.9085 | 239.6403 | 225.9353 | 204.8171 | 507.6341 | 131.7492 | 170.1799 | 187.0915 | 402.5015 | 253.4779 | 183.7934 | 196.0236 | 404.1799
1742 | 21.1564 | 154.1834 | 201.8038 | 206.8038 | 208.4051 | 138.823 | 152.2153 | 162.5821 | 172.6332 | 158.4115 | 119.0191 | 132.177 | 155.5885 | 206.8166 | 126.614 | 150.6013 | 268.9809 | 189.6459 | 147.2089 | 127.386 | 191 | 156.6077 | 297.8038 | 299.0128 | 165.7656 | 195.7592 | 193.2026 | 225.8102 | 353.9809 | 119.8421 | 117.9872 | 113.8102 | 148.7847 | 152.9745 | 117.823 | 136.3987 | 98.0128 | 241.0064 | 217.1515 | 217.9745 | 271.8038 | 143.8421 | 128.8166 | 127.5949 | 138.6459 | 68.386 | 70 | 64.9936 | 108.614 | 177.9554 | 183.2281 | 211.8293 | 158.7911 | 130.2153 | 110.7783 | 138.2026 | 116.6077 | 156.5949 | 150.1069 | 185.5821 | 173.563 | 131.3668 | 123.2281 | 165.1898 | 154.0191 | 172.1834 | 156.1834 | 192.1834 | 268.2967 | 134.3796 | 125.8102 | 210.177 | 214.3923 | 211.3477 | 178.8421 | 166.1643 | 395.3605 | 144.7974 | 127.9745 | 148.9936 | 222.4242 | 131.3987 | 116.0255 | 139.8038 | 154.3987 | 171.2153 | 140.9936 | 166.2089 | 137.9681 | 95.0128 | 93.9936 | 90.0128 | 131 | 209.4051 | 211.614 | 210.4051 | 342.9872 | 196.7656 | 190.4179 | 251.5694 | 296.9936 | 160.3923 | 62.3923 | 69.7847 | 160.3923 | 107.9936 | 111.3923 | 105.9745 | 122.2089 | 103.386 | 113.4306 | 103.7974 | 104.7911 | 67.2026 | 76.614 | 67.3987 | 81.7783 | 89.7656 | 89.1962 | 99.8102 | 92.0064 | 272.8038 | 242.4434 | 201.7783 | 490.7911 | 126.5949 | 159.3796 | 185.6204 | 423.9745 | 259.7974 | 171.9809 | 205.5949 | 392.5821
1743 | 21.1674 | 157.0678 | 189.3136 | 207.8771 | 207.233 | 146.9067 | 149.9322 | 159.4916 | 159.4661 | 158.7542 | 110.822 | 138.0678 | 161 | 210.5084 | 129.8644 | 154.9322 | 265.6695 | 165.9449 | 147.9322 | 137.6991 | 195.1102 | 148.0678 | 289.8348 | 287.6568 | 180.3559 | 210.0678 | 199.8093 | 225.9322 | 339.1611 | 106.5763 | 122.3559 | 106.9449 | 158.0551 | 155.9576 | 121.3856 | 137.822 | 82.9576 | 229.7797 | 232.6017 | 229.589 | 279.4534 | 124 | 131.6864 | 126.3009 | 117.822 | 78.2881 | 74.1102 | 72.9322 | 95.8348 | 192.0127 | 169.589 | 198.822 | 153.9576 | 126.6991 | 127.5763 | 129.6017 | 109.7119 | 158.178 | 174.2712 | 197.6991 | 175.3814 | 145.767 | 112.8771 | 166.3009 | 145.411 | 175.0678 | 156.6017 | 199.589 | 277.5425 | 146.0551 | 126.3432 | 215.2458 | 213.4788 | 216.6823 | 176.2627 | 192.4534 | 400.4365 | 153.8093 | 146.9873 | 146.2458 | 213.8771 | 139.3983 | 112.2203 | 138.411 | 156.6441 | 162.3559 | 144 | 164.8771 | 143.5466 | 92.2881 | 97.822 | 88.1102 | 119.0805 | 210.2881 | 206.233 | 203.4788 | 342.0127 | 208.4788 | 204.6059 | 253.2585 | 311.0974 | 168.9322 | 63.5339 | 75.3559 | 168.9322 | 116.3432 | 115.411 | 115.1229 | 115.9449 | 108.767 | 101.4661 | 103.767 | 113.1102 | 70.1102 | 69.178 | 70.8771 | 95.6991 | 100.2458 | 95.9322 | 103.6313 | 77.0805 | 251.2712 | 204.7373 | 207.4788 | 489.6568 | 124.4788 | 178.0424 | 184.2203 | 436.411 | 262.6441 | 179.767 | 204.3009 | 403.0551
1744 | 21.1784 | 150.8466 | 192.2022 | 210.3068 | 209.4233 | 157.0677 | 140.9143 | 147.6564 | 162.5767 | 160.2271 | 116.6564 | 134.3865 | 163.9631 | 219.3805 | 137.6504 | 155.6504 | 234.8775 | 168.0797 | 150.7669 | 137.9203 | 199.3068 | 145.6564 | 273.8038 | 291.4293 | 172.497 | 203.4233 | 211.4233 | 226.2271 | 324.0488 | 112.6135 | 114.0737 | 114.5827 | 155.9203 | 154.8526 | 132.6872 | 140.2699 | 87.7789 | 229.4293 | 223.6932 | 221.8038 | 286.2271 | 126.9631 | 138.9572 | 121.8466 | 118.5767 | 81.3068 | 66.4542 | 72.8038 | 97.5827 | 178.1902 | 168.1534 | 202.9631 | 152.4293 | 131.1534 | 129.8038 | 120.2699 | 120.2391 | 155.3068 | 149.8098 | 193.6872 | 162.3128 | 147.8098 | 125.0428 | 163.1165 | 140.497 | 171.8098 | 159.5458 | 196.4602 | 243.8586 | 143.497 | 138.2331 | 216.3128 | 205.7301 | 194.9263 | 205.2331 | 200.9203 | 395.3924 | 148.9143 | 147.2639 | 146.0428 | 216.3068 | 138.1474 | 129.0797 | 133.497 | 156.4602 | 166.3496 | 151.1962 | 180.006 | 127.8835 | 98.2699 | 93.497 | 99.0797 | 112.1594 | 233.6255 | 207.1534 | 207.5827 | 329.8835 | 203.6932 | 220.9143 | 243.8894 | 330.3865 | 151.4482 | 64.2331 | 74.2699 | 151.4482 | 134.5827 | 116 | 117.7729 | 111.3068 | 105.3068 | 105.4233 | 96.9203 | 121.1165 | 65.8406 | 71.3865 | 61.8775 | 98.0369 | 99.1962 | 95.3805 | 105.8406 | 71.4293 | 225.1165 | 208.7849 | 202.6932 | 487.0797 | 118 | 183.8038 | 183.3008 | 421.7609 | 245.1046 | 182.2331 | 196.8835 | 405.5767
1745 | 21.1894 | 145.466 | 210.466 | 203.2084 | 201.288 | 150.1171 | 126.8408 | 151.3864 | 168.0984 | 149.6932 | 125.6136 | 138.1288 | 170.178 | 203.3372 | 142.9391 | 148.7424 | 216.5644 | 172.3864 | 141.0796 | 136.2272 | 193.0796 | 153.7424 | 278.8103 | 307 | 168.9204 | 212.712 | 210.6932 | 216.1288 | 345.644 | 114.534 | 107.7424 | 117.2389 | 145.5152 | 165.7728 | 123.0492 | 136.3372 | 101.11 | 241.5152 | 220.3372 | 221.1476 | 269.5948 | 129.2576 | 131.9204 | 126.4848 | 114.5152 | 84.6628 | 62.8103 | 61.2576 | 110.6932 | 183.89 | 163.5152 | 199.1592 | 155.8033 | 128.2576 | 128.712 | 109.8033 | 133.1592 | 157.7916 | 157.6136 | 177.9508 | 164.5456 | 139.4965 | 132.5152 | 169.4848 | 139.9696 | 173.9508 | 178 | 189.5152 | 266.7236 | 137.7424 | 144.8712 | 221.1592 | 214.89 | 199.0796 | 208.3864 | 193.5644 | 419.7916 | 134.8408 | 146.6323 | 151.7728 | 235.7799 | 129.8408 | 138.178 | 129.0492 | 149.5152 | 164.9812 | 165.7916 | 188.0304 | 137.6323 | 100 | 88.6136 | 102.5152 | 119.0304 | 234.1709 | 202.0796 | 216.3372 | 327.4356 | 219.5035 | 205.534 | 239.8829 | 325.8524 | 147.1851 | 70.8712 | 75.1288 | 147.1851 | 138.9812 | 125.1476 | 116.5456 | 110.3068 | 103.4356 | 114.712 | 85.644 | 123.5644 | 62.0188 | 70.7728 | 56.6323 | 92.6932 | 102.466 | 99.8103 | 83.7236 | 82.644 | 244.1171 | 225.7845 | 195.4168 | 488.3372 | 123.2272 | 174.4356 | 182.1171 | 403.178 | 238.0375 | 177.11 | 197.0492 | 394.11
1746 | 21.2004 | 147.3014 | 205.5831 | 198.7916 | 186.4169 | 166.969 | 139.2394 | 156.5099 | 170.6253 | 146.6563 | 134.3747 | 129.3859 | 180.6141 | 213.0198 | 149.3549 | 143.8338 | 223.6141 | 173.4789 | 146.3014 | 141.6563 | 191.5831 | 162.7183 | 306.9578 | 314.1662 | 162.8028 | 230.7183 | 208.1042 | 223.5099 | 353.8846 | 114.0422 | 116.2704 | 106.031 | 147.2704 | 156.7606 | 135.5099 | 141.093 | 94.6141 | 236.5521 | 215.6873 | 205.6788 | 268.5099 | 128.3437 | 118.1887 | 118.4591 | 122.093 | 92.8958 | 82.8958 | 63.9268 | 110.7916 | 193.0732 | 159 | 197.062 | 146.7183 | 115.2507 | 126.1154 | 109.2282 | 116.2817 | 164.8958 | 151.5943 | 191.0198 | 163.9888 | 139.4677 | 115.907 | 170.8648 | 145.4169 | 165.969 | 183.8226 | 189.031 | 262.7183 | 148.062 | 143.3127 | 202.0422 | 228.5521 | 191.7606 | 211.2704 | 189.8648 | 419.7296 | 144.1042 | 156.7296 | 139.1774 | 242.0534 | 130.1464 | 141 | 132.1042 | 154.4057 | 148.8648 | 166.1774 | 177.2084 | 140.5633 | 100.8958 | 86.1774 | 93.0732 | 106.4169 | 206.031 | 213.124 | 214.8226 | 315.0112 | 228.4901 | 220.7183 | 224.8226 | 313.5521 | 182.7606 | 68.4169 | 78.9268 | 182.7606 | 113.907 | 119.0844 | 103.4479 | 116.9268 | 99.5211 | 117.0422 | 86.7183 | 123.8958 | 69.1774 | 73.8536 | 71.2084 | 97.2704 | 97.5831 | 114.0732 | 81.5099 | 75.2084 | 251.1154 | 197.4479 | 198.2282 | 488.6141 | 125.5211 | 177.2394 | 203.4479 | 390.7718 | 238.4703 | 168.3747 | 188.907 | 381.7916
1747 | 21.2114 | 158.6195 | 211.8407 | 198.2389 | 182 | 168.9027 | 148.9027 | 155.0353 | 164 | 143.1593 | 141.9204 | 113.5575 | 173.4335 | 217.4335 | 140.9823 | 149.2656 | 236.2213 | 179.4602 | 151.6371 | 140.4602 | 192.7787 | 163.1769 | 313.3982 | 321.6195 | 151.3009 | 221.9733 | 206.0796 | 224.7964 | 359.4955 | 105.7611 | 117.0973 | 115.6018 | 146.2566 | 147.0973 | 139.0973 | 143.115 | 98.938 | 239.2213 | 213.4778 | 189.0267 | 264.2742 | 122.1769 | 105.5841 | 115.885 | 125.0353 | 91.2389 | 76.6371 | 60.7964 | 115.7611 | 175.9557 | 170.9647 | 199.1769 | 155 | 120.8674 | 123.2656 | 130.3009 | 106.6195 | 162.3186 | 154.2656 | 207.3982 | 178.1593 | 153.1769 | 113.4249 | 146.2919 | 143.7611 | 160.0796 | 170.7521 | 195.8407 | 272.8407 | 156.6195 | 143.6814 | 186.3982 | 205.9114 | 191.7611 | 204.7344 | 179.5575 | 395.4335 | 148.9823 | 144.3982 | 139.3629 | 218.2213 | 120.7611 | 125.3539 | 137.4425 | 162.3186 | 159.2656 | 168.6638 | 179.1416 | 124 | 96.4778 | 91.4249 | 95.2832 | 107.062 | 199.0353 | 220.5751 | 217.8584 | 310.9647 | 206.4955 | 220.7168 | 226.0177 | 321.2832 | 175.3982 | 67.6814 | 74.876 | 175.3982 | 112.8054 | 105.7434 | 114.5841 | 128.5222 | 89.3982 | 102.7787 | 104.6638 | 118.0973 | 63.8407 | 82.3186 | 68.5398 | 105 | 92.3362 | 101.0973 | 94.3009 | 77.1416 | 235.8407 | 188.7964 | 225.2832 | 484.1946 | 124.1416 | 172.6371 | 191.5751 | 381.6638 | 222.6461 | 180.5222 | 170.3186 | 369.7344
1748 | 21.2224 | 154.1652 | 201.7157 | 195 | 185.3073 | 160.936 | 148.023 | 159.8988 | 160.2202 | 146.6697 | 127.4085 | 108.8348 | 166.2843 | 205.0871 | 136.6607 | 165.0551 | 238.1101 | 179.0551 | 142.5275 | 143.8348 | 189.4725 | 149.7478 | 306.5826 | 301.1012 | 164.8668 | 202.4725 | 209.7247 | 214 | 346.4725 | 111.8348 | 115.6146 | 121 | 137.8899 | 138.0551 | 122.4955 | 128.4725 | 91.1101 | 241.1101 | 211.7247 | 213.7247 | 255.032 | 113.9449 | 113.2753 | 135.2522 | 121.8668 | 80.4405 | 66.5826 | 62.7567 | 119.4725 | 172.9219 | 180.2753 | 202.7567 | 154.5275 | 138.2753 | 127.2433 | 124.968 | 108.7798 | 159.4174 | 155.8809 | 193.4955 | 174.1101 | 146.3624 | 119.4955 | 140.8988 | 135.6607 | 168.9219 | 152.1972 | 206.5045 | 266.968 | 140.8259 | 145.1652 | 190.9219 | 180 | 199.2522 | 175.2433 | 186.1741 | 383.0871 | 159.3073 | 129.5506 | 147.5275 | 226.7247 | 122.1101 | 121.6466 | 141.6927 | 154.2202 | 170.3303 | 174.8028 | 178.8028 | 121.1652 | 95.6697 | 100.8028 | 101.6927 | 104.968 | 200.1191 | 221.1191 | 216.7798 | 315.0781 | 199.6146 | 216.1972 | 219.4725 | 321.5506 | 163.3854 | 67.7478 | 74.8118 | 163.3854 | 124.2753 | 106.0781 | 122.2753 | 123.6607 | 97.7017 | 97.5826 | 109.8579 | 108.1101 | 66.4725 | 76.5826 | 74.1421 | 102.1652 | 95.1741 | 87.3303 | 102.1972 | 71.6056 | 240.3624 | 193.5915 | 225.0781 | 465.6376 | 127.5826 | 177.7017 | 183.1421 | 418.9859 | 249.032 | 187.4725 | 178.7567 | 368.1191
1749 | 21.2334 | 170.8753 | 188.9695 | 202.2714 | 199.6648 | 141.5457 | 143.3324 | 166.9419 | 149.6981 | 137.4876 | 133.2686 | 116.8476 | 182.4238 | 204.759 | 140.9667 | 163.5152 | 251.0276 | 176.5457 | 141.0305 | 150.3933 | 188.5457 | 143.5457 | 300.6371 | 298.421 | 169.7286 | 208.8171 | 209.1829 | 205.2743 | 339.2438 | 115.9695 | 123 | 114.2133 | 143.8781 | 146.2105 | 122.2105 | 117.3657 | 84.1524 | 240.4543 | 218.939 | 214.1524 | 271.7562 | 120.3324 | 117.6952 | 129.3352 | 129.1524 | 70.5457 | 65.9695 | 67.7895 | 114.1829 | 183.0305 | 170.6371 | 202.9419 | 150.6067 | 133.4848 | 119.3019 | 125.2105 | 115.4238 | 147.9114 | 153.0581 | 175.759 | 169.5762 | 155.7867 | 121.1495 | 149.8809 | 132.2105 | 157.7009 | 146.3657 | 198.5486 | 248.7895 | 133.241 | 144.9086 | 203.4543 | 189.2105 | 204 | 179.9057 | 194.7286 | 380.3352 | 156.2133 | 129.6648 | 151.3629 | 221.82 | 132.6038 | 128.3048 | 147.2105 | 156.3019 | 160.1219 | 162.6981 | 181.241 | 116.061 | 102.9695 | 101.7867 | 95.5762 | 107.1495 | 210.2133 | 214.2466 | 219.5457 | 320.9667 | 215.7257 | 210.3657 | 215.6371 | 310.0305 | 155.0305 | 71.241 | 78.7895 | 155.0305 | 110.759 | 102.2714 | 117.4848 | 113.9086 | 114.4543 | 97.939 | 108.8171 | 107.4543 | 74.2714 | 70.6371 | 78.5762 | 100.4543 | 106.1524 | 89.7257 | 94.4267 | 64.3629 | 236.2133 | 218.2714 | 224.6648 | 483.8448 | 126 | 183.3048 | 185.1524 | 451.8781 | 249.759 | 189.4543 | 193.4848 | 381.1219
1750 | 21.2444 | 167.6444 | 185.5265 | 205.0295 | 209.5088 | 154.9116 | 146.5147 | 150.0118 | 145.9823 | 137.4028 | 141.1002 | 128.4617 | 175.556 | 211.9057 | 144.0884 | 153.556 | 270.9764 | 167.5442 | 130.556 | 151.5147 | 192.4676 | 135.5383 | 291.5265 | 293.6444 | 162 | 212.5147 | 195.0471 | 196 | 339.947 | 117.9941 | 122.0059 | 103.5206 | 151.4794 | 138.6031 | 142.9352 | 116.9293 | 83.4735 | 231.0648 | 213.0471 | 228.3851 | 280.4971 | 124.5088 | 131.9764 | 115.0177 | 117.0413 | 69.9941 | 65.5088 | 68.9352 | 104.0236 | 178.5206 | 178.4263 | 185.5147 | 149.9823 | 131.0177 | 120.5324 | 121.5796 | 113.5265 | 154.8762 | 156.0707 | 176.444 | 167.4971 | 151.0707 | 124.5501 | 134.5147 | 143.4912 | 147.4204 | 147.9175 | 189.444 | 239.4971 | 147.9646 | 152.4735 | 197.0471 | 189.0589 | 200.0236 | 197.0177 | 170.1002 | 388.3851 | 154.9646 | 144.4794 | 151.0295 | 233.3674 | 123.6326 | 122.4735 | 149.5442 | 163.9941 | 156.9941 | 149.0413 | 170.1179 | 116.4853 | 103.5029 | 102.5383 | 93 | 116.0177 | 198.0295 | 189 | 220.9823 | 322.5972 | 223.5088 | 196.5088 | 213.4853 | 327.391 | 152.0118 | 85.4676 | 65.5206 | 152.0118 | 108.4617 | 110.4971 | 105.0766 | 123.9587 | 107.053 | 106.9587 | 114.5029 | 111.4853 | 77.5265 | 66 | 71.5265 | 97.0295 | 104.4794 | 91.5442 | 90.9352 | 75.4617 | 233.9705 | 228.9823 | 219.5972 | 498.5442 | 143.8939 | 179.4617 | 168.5678 | 414.7446 | 254.4204 | 179.0707 | 184.556 | 395.391
1751 | 21.2554 | 151.1495 | 191.1869 | 191.8505 | 194.2477 | 161.8505 | 149.6028 | 161.243 | 160.7149 | 150.9439 | 132.6589 | 125.8318 | 160.3972 | 217.3038 | 144.7897 | 152.1495 | 240.8739 | 173.243 | 124.5654 | 138.8131 | 204.1121 | 135.6215 | 304.3177 | 289.8831 | 181.8644 | 211.0187 | 186.4907 | 233.6915 | 358.1682 | 113.3972 | 123.5467 | 102.5467 | 145.8318 | 134.243 | 136.6823 | 122.3785 | 89.528 | 227.1308 | 215.6962 | 244.4346 | 278.4533 | 125.5467 | 127.3411 | 119.6402 | 121.7149 | 70.4907 | 65.0187 | 67.2664 | 110.6962 | 167.3598 | 191 | 183 | 156.5654 | 118.8318 | 118.5654 | 114.1121 | 122.7523 | 160.2103 | 147.0561 | 183.9626 | 169.0187 | 139 | 121.0934 | 153.9018 | 139.9066 | 173.2243 | 168.1121 | 209.1869 | 238.472 | 162.6589 | 141.2103 | 193.0748 | 187.1495 | 206.1869 | 183.8131 | 184.07 | 401.3785 | 148.7757 | 146.9813 | 134.2851 | 237.1543 | 114.243 | 127.5093 | 121.6262 | 155.3225 | 152.9066 | 150.6589 | 156.1121 | 119.5093 | 102.4907 | 97.528 | 101.6589 | 116.0561 | 192.4907 | 190.0187 | 207.701 | 317.7149 | 227.6028 | 213.8457 | 223.1308 | 337.3411 | 148.9813 | 86.9252 | 67.6028 | 148.9813 | 116.0187 | 117.1121 | 98.1121 | 132.0187 | 100.5467 | 110.4346 | 106.3598 | 115.0187 | 76.5654 | 67.528 | 66.4907 | 105.243 | 100.8692 | 88.5841 | 102.5093 | 77.4159 | 249.1869 | 223.3411 | 222.8644 | 502.2056 | 145.701 | 174.2851 | 169.2243 | 376.5654 | 264.9439 | 178.7149 | 170.9252 | 423.1682
1752 | 21.2664 | 166.8245 | 192.6538 | 181.3918 | 189.8678 | 147.7404 | 158.6514 | 172.4351 | 157.3942 | 152.1731 | 132.6538 | 120.6514 | 161.7813 | 198.1322 | 157.9135 | 157.3918 | 214.2596 | 186 | 132.6947 | 146.7356 | 202.6971 | 140.4351 | 308.4807 | 299.0457 | 190.5673 | 208.8702 | 196.9543 | 237.137 | 332.5721 | 109.0433 | 113.4807 | 108.1298 | 142.2164 | 142.3053 | 126.7813 | 109.2187 | 87.3486 | 238.6947 | 244.2572 | 220.1346 | 268.6971 | 116.0024 | 120.5649 | 133.2596 | 122.0457 | 75.738 | 67.0433 | 63.3462 | 122.0433 | 156.8702 | 189.9567 | 190.8245 | 150.0889 | 116.7813 | 115.7404 | 127.3029 | 125.0457 | 163.2572 | 151.0433 | 181.4784 | 162.6971 | 131.6971 | 123.3918 | 164.5673 | 135.5216 | 170.9158 | 163.5673 | 197.0913 | 243.7813 | 159.524 | 136.4327 | 204.8245 | 195.5216 | 207.1322 | 181.8245 | 187.3966 | 412.214 | 127.6106 | 144.9567 | 130.8245 | 225.2596 | 140.0409 | 125.9135 | 123.3871 | 153.8245 | 160.5193 | 154.3053 | 159.9135 | 116.8702 | 100.4351 | 101.6082 | 101.1322 | 120.0433 | 202.9543 | 198.8245 | 207.6491 | 312.8293 | 214.7428 | 218.9591 | 230 | 333.1731 | 149.0433 | 70.4807 | 61.0024 | 149.0433 | 123.7813 | 122.4784 | 102.9567 | 129.8702 | 98.8269 | 104.9135 | 103.6947 | 108.6971 | 70.6106 | 70.0433 | 64.4351 | 107.5673 | 100.2596 | 101.3462 | 97.7836 | 80.8245 | 257.4351 | 218.6514 | 233.1322 | 495.786 | 134.6947 | 160.3918 | 165.3509 | 422.2525 | 271.9111 | 186.3486 | 162.8269 | 395.4855
1753 | 21.2774 | 161.7197 | 185 | 193.9499 | 200.1357 | 135.059 | 152.3893 | 169.3982 | 146.6696 | 146.9232 | 123.3982 | 132.5428 | 163.7286 | 206.5517 | 145.3215 | 162.475 | 215.7286 | 189.2036 | 136.4661 | 152.3215 | 196 | 141.1357 | 300.2036 | 292.7375 | 171.9232 | 206.5339 | 208.0679 | 221.9499 | 328.7642 | 110 | 108.941 | 114.2036 | 140.5929 | 148.1446 | 134.6018 | 109.941 | 84.5339 | 243 | 225.7609 | 215.9499 | 283.3571 | 117.2803 | 126.8053 | 132.5929 | 121.6107 | 72.4572 | 75.475 | 69.3982 | 124.0679 | 163.0768 | 176.1858 | 207.0768 | 142.6018 | 116.5929 | 111.0679 | 120.1179 | 127.2803 | 169.8554 | 153.0679 | 192.2125 | 164.0089 | 128.7375 | 126.8732 | 150.1947 | 140.2714 | 153.1947 | 158.8053 | 191.416 | 262.2803 | 151.6696 | 139.059 | 201.8554 | 207.7464 | 200.0679 | 185.2625 | 170.475 | 408.7786 | 132.8821 | 128.5161 | 139.6018 | 230.4661 | 142.3893 | 124.5339 | 131.2536 | 151.3893 | 173.6018 | 166.5517 | 167.0768 | 109.7286 | 99.5339 | 93.3215 | 100.475 | 114.059 | 194.8465 | 199.059 | 213.4572 | 321.4928 | 196.9321 | 203.7964 | 216.1179 | 325.7875 | 156.4071 | 70.7464 | 61.7464 | 156.4071 | 119.8554 | 121.4661 | 102 | 122.7286 | 99.2714 | 109.4071 | 102.6607 | 119.6196 | 71.0768 | 64.059 | 67.8053 | 101.2036 | 95.3215 | 98.3215 | 103.6785 | 83.7286 | 244.2536 | 212.3893 | 233.5428 | 491.2125 | 129.3893 | 175.0856 | 164.8821 | 435.9055 | 275.1268 | 178.1947 | 167.5428 | 375.3482
1754 | 21.2884 | 144.4454 | 184.4538 | 191.7059 | 200.9076 | 131.1849 | 146.1849 | 154.3446 | 161.0168 | 140.6387 | 115.9916 | 135.6303 | 158.9076 | 202.9748 | 140.3697 | 162.9916 | 223.4706 | 185.4454 | 136.5462 | 140.8151 | 198.7311 | 151.1933 | 288.7983 | 303.1008 | 161.8151 | 205.3613 | 211.7311 | 222.437 | 323.5043 | 103.9916 | 106.2605 | 118.6387 | 133.3613 | 159.7311 | 129.9916 | 120.3697 | 89.916 | 237.5378 | 203.7395 | 224.6303 | 300.3613 | 129.6387 | 128.2689 | 129.1849 | 125.6303 | 71.0084 | 81.4538 | 68 | 114.0756 | 185.2017 | 152.437 | 208.9916 | 157.1092 | 117.5546 | 125.6554 | 112.916 | 126.5294 | 158.8151 | 155.0924 | 181.2437 | 178.647 | 125.9916 | 123.8067 | 159.1092 | 140.1765 | 171.3949 | 162.4538 | 192.5294 | 270.2689 | 135.9748 | 134.6387 | 205.563 | 204.8908 | 208.647 | 183.0924 | 190.1092 | 399.1008 | 131.3446 | 133.0252 | 136.084 | 228.3613 | 123.6219 | 137.0168 | 130.2857 | 136.4454 | 173.3613 | 174.9916 | 173.3613 | 105.4538 | 105.4622 | 94.9244 | 105.3613 | 108 | 189.9244 | 203.3781 | 196.1681 | 338.7227 | 199.2773 | 190.0756 | 217.6554 | 313.8151 | 169.647 | 66.7983 | 74.1849 | 169.647 | 118.647 | 118.8151 | 117.8403 | 121.1849 | 99.7227 | 110.084 | 101.8235 | 124.6219 | 78.4538 | 65.1008 | 77.4622 | 94.7143 | 99.1092 | 93.3697 | 103.7143 | 78.9076 | 236.1849 | 208.916 | 217.5126 | 490.0756 | 125.916 | 181.2605 | 168.2605 | 402.5294 | 241.0505 | 164.7143 | 185.3781 | 400.6554
1755 | 21.2994 | 153.5212 | 195.7287 | 179.5585 | 183.8033 | 146.9627 | 150.7925 | 144.6755 | 159.2713 | 139.766 | 108.766 | 126.9734 | 146.2713 | 208.1063 | 142.3245 | 156.883 | 231.883 | 181.6755 | 138.6755 | 139 | 212.1702 | 148.5054 | 293.1967 | 295.0373 | 161.6755 | 201.2075 | 197.8033 | 212 | 316.3138 | 117.4308 | 105.7021 | 125.5851 | 130.3245 | 144.6862 | 125.5585 | 124 | 93.4415 | 236.3511 | 211.5585 | 226.0266 | 292.2979 | 123.1809 | 128.7925 | 127.0904 | 115.2979 | 80.4681 | 72.0639 | 65.2075 | 103.3245 | 178.5692 | 183.8881 | 200.0904 | 175.2606 | 120.766 | 131.4149 | 119.234 | 123.7021 | 176.5478 | 156 | 167.9096 | 168.8033 | 136.0797 | 115.3245 | 162.7394 | 138.117 | 183.2979 | 151.9468 | 190.8191 | 231.6969 | 126.0266 | 132.6489 | 225.4946 | 190.0904 | 196.0107 | 181.766 | 182.0107 | 415.6117 | 127.8191 | 141.8564 | 156.5744 | 220.8564 | 133.9893 | 149.234 | 137.4415 | 133.234 | 171.4415 | 171.117 | 168.6489 | 116.7287 | 95.4788 | 102.3245 | 109.0266 | 114.1436 | 200.6755 | 196.9203 | 190.7925 | 332.0904 | 188.5958 | 187.7021 | 212.2447 | 289.1011 | 169.2979 | 61.7021 | 76.5585 | 169.2979 | 119.9734 | 105.8298 | 132.6755 | 115.1809 | 94.766 | 105.4415 | 92.7128 | 118.7925 | 70.7394 | 74.9096 | 74.7394 | 83.0904 | 106.0904 | 81.9203 | 99.3511 | 74.6489 | 248.0532 | 216.3511 | 217.5478 | 477.6489 | 122.7394 | 163.9468 | 168.2606 | 398.0266 | 242.6914 | 160.3511 | 186.1809 | 396.9203
1756 | 21.3104 | 168.9956 | 189.0176 | 187.9912 | 192.1196 | 148.2964 | 151.2876 | 156.2744 | 146.5752 | 143.1372 | 107.5708 | 119.0044 | 160.9736 | 210.8716 | 139.2876 | 142.8716 | 241.8452 | 175.0088 | 134.292 | 144.1372 | 198.168 | 147.2788 | 295.7256 | 297.6992 | 169.2788 | 206.9912 | 196.9868 | 214.2832 | 315.73 | 125.7212 | 107.0044 | 117.4424 | 142.1284 | 141.2744 | 115.7256 | 129.708 | 93 | 241.2832 | 213.7124 | 223.7212 | 280.7212 | 122.1372 | 125.292 | 114.2964 | 110 | 80.0044 | 70.1372 | 76.1284 | 111.7036 | 167.9956 | 191.8848 | 197.5708 | 167.8716 | 122.9956 | 121.8628 | 125.5664 | 123.292 | 168.0264 | 150.292 | 172.7124 | 170.27 | 139.438 | 115.1416 | 163.2788 | 149.2744 | 165.4424 | 147.4248 | 194.7168 | 244.6684 | 136.2788 | 130.5708 | 199.8936 | 191.5664 | 196.9824 | 179.4292 | 185.8364 | 398.8848 | 132.8584 | 130.1504 | 159.4468 | 220.5664 | 129.5928 | 139.584 | 130.7212 | 144.7036 | 163.5796 | 159.4424 | 167.1416 | 128.854 | 84.5708 | 105.5664 | 109.0044 | 120.7124 | 195.7212 | 188.9956 | 200.9912 | 338.1328 | 189.416 | 200.9912 | 201.854 | 294.4028 | 167.4248 | 63.5752 | 71.292 | 167.4248 | 119.9956 | 94.7168 | 136.854 | 111.2832 | 104.416 | 97.0088 | 102.978 | 113.0088 | 68.9956 | 73.4336 | 71.854 | 87.9912 | 100.7168 | 88.2656 | 98.0044 | 84.5576 | 256 | 212.7212 | 218.1592 | 478.4248 | 130.6992 | 157.1416 | 173.4292 | 420.8364 | 246.4468 | 170.9912 | 183.4248 | 389.5664
1757 | 21.3214 | 174.3324 | 182.2478 | 206.2449 | 190.5073 | 141.5831 | 141.2536 | 161.0845 | 153.3295 | 145.2507 | 114.9971 | 118.9155 | 175.5015 | 212.0787 | 143.8309 | 149.9096 | 234.8396 | 156.172 | 133.4985 | 131.0904 | 182.7493 | 144.4198 | 288.5831 | 298.0875 | 166.4198 | 200.172 | 191.4227 | 207.8367 | 319.5773 | 113.4198 | 111.5802 | 116.1633 | 155.4985 | 138.5044 | 117.3295 | 127.586 | 102.9125 | 233.0875 | 204.5044 | 225.414 | 277.7493 | 120.7522 | 120.4169 | 114.5802 | 105.3353 | 78.1662 | 79.2478 | 77.8367 | 113.7522 | 169.8338 | 173 | 200.3324 | 147.5044 | 116.6705 | 126.1633 | 128.4169 | 114.9184 | 156.414 | 150.6647 | 175.1662 | 177.5015 | 139.414 | 116 | 156.9213 | 142.4227 | 168.2449 | 161.6618 | 193 | 266.6705 | 129.9213 | 134.4985 | 186.0787 | 192.6676 | 197.3382 | 183.6647 | 181.3411 | 395.7434 | 130.2507 | 137.828 | 126.0087 | 233.9125 | 124.9125 | 129.8338 | 131.2478 | 148.5015 | 163.8309 | 152.9155 | 171.4985 | 119.3382 | 86.7493 | 110.1662 | 113.5802 | 116.7522 | 182.2536 | 198.9971 | 195.3382 | 356.6618 | 208.4956 | 198.2536 | 199.3353 | 309.6676 | 169.4169 | 70.3295 | 69.9155 | 169.4169 | 111.9213 | 102.9125 | 135.5015 | 121.1633 | 114.1662 | 96.8309 | 94.0933 | 104.0845 | 69.0845 | 65.3353 | 79.8309 | 92.2507 | 97.2507 | 98.5015 | 95 | 75.9242 | 247.8367 | 202.1691 | 205.2507 | 471.6705 | 135.1691 | 165.5802 | 186.4111 | 426.2536 | 240.9971 | 177.5831 | 195.3295 | 367.344
1758 | 21.3324 | 179.5723 | 203.2661 | 192.3755 | 191.0984 | 148.5492 | 154.0522 | 163.7631 | 162.3815 | 132.0924 | 120 | 134.0984 | 168.2369 | 211.6647 | 147.4046 | 145.9247 | 227.5723 | 161.2892 | 121.1155 | 134.4799 | 186.3815 | 136.6185 | 291.9769 | 274.9247 | 161 | 217.1968 | 195.2892 | 201.4046 | 320.4739 | 115.74 | 115.8092 | 113.0693 | 142.5201 | 139.9307 | 131.74 | 125.3815 | 93.9247 | 221.8323 | 197 | 211.5432 | 285.5492 | 131.8845 | 114.6416 | 120 | 118.0753 | 74.2369 | 74.6647 | 72 | 111.1908 | 172.5723 | 174.7861 | 184.7339 | 144.7631 | 121.9076 | 139.74 | 125.0231 | 122.7169 | 181.243 | 149.2369 | 183.1446 | 176.7861 | 127.9247 | 116 | 142.4508 | 147.4799 | 184.1446 | 163.4508 | 194.7861 | 246.9016 | 138.0753 | 131.0462 | 206.503 | 201.1215 | 182.4508 | 186.4046 | 184.6707 | 397.4739 | 135.5492 | 144.0231 | 125.8845 | 239.8092 | 142.1215 | 118.2831 | 148.6938 | 139.4508 | 165.0231 | 153.2139 | 173.4046 | 118.1446 | 97.5261 | 113.9769 | 110.6647 | 107.6416 | 193.2661 | 190.3062 | 192.3584 | 365 | 212.4277 | 197.3584 | 198.9769 | 300.26 | 174.3584 | 74.6185 | 74.9769 | 174.3584 | 107.5723 | 121.3123 | 125.26 | 127.5954 | 89.3986 | 104.5723 | 90.6938 | 107.3584 | 65.2139 | 48.9016 | 77.4508 | 97.5492 | 107.3123 | 97.1908 | 87.8554 | 68.9538 | 246.1677 | 188.4739 | 205.7861 | 478.6938 | 128.6185 | 172.7861 | 178.7339 | 412.26 | 256.1215 | 173.2369 | 179.3755 | 357.9307
1759 | 21.3434 | 167.4159 | 199.8082 | 200.6992 | 192.1003 | 140.8465 | 156.6696 | 153.0235 | 165.2153 | 159.8527 | 127.8997 | 128.4159 | 168.646 | 199.3156 | 153.6844 | 130.7463 | 214.2005 | 177.6844 | 120.1151 | 132.2389 | 180.7079 | 145.3304 | 309.1918 | 270.0767 | 154.3156 | 197.3243 | 217.1535 | 216.1918 | 324.9381 | 119.7847 | 116.2153 | 113.6844 | 144.1535 | 125.3391 | 149.1535 | 134.2921 | 91.5074 | 221.4307 | 196.3923 | 217.8379 | 284.5309 | 117.6931 | 111.6077 | 121.823 | 116.2389 | 75.8614 | 71.4307 | 62.8849 | 126.5841 | 164.6696 | 180.8614 | 197.3069 | 150.4307 | 105.6931 | 134.0619 | 134.5458 | 131.2153 | 160.9703 | 148.4307 | 197.7228 | 155.5161 | 120.646 | 123.8997 | 151.9765 | 140.9852 | 167.7315 | 156.5693 | 192.9616 | 263.5223 | 140.4926 | 137.9381 | 216.823 | 206.7847 | 178.6077 | 175.0619 | 191.1386 | 381.2772 | 143.0384 | 138.9616 | 139.0384 | 216.5161 | 138.0619 | 109.7847 | 143.4159 | 135 | 157.5309 | 164.1535 | 183.3304 | 125.4307 | 96.7079 | 99.5928 | 112.2921 | 110.0767 | 197.1003 | 186.4691 | 200.2537 | 351.0235 | 216.6844 | 204.646 | 221.0532 | 301.0767 | 182.8614 | 72.3923 | 69.7079 | 182.8614 | 106.354 | 121.7079 | 116.9616 | 133.4691 | 87.7995 | 91.8082 | 107.8997 | 119.5074 | 63.3923 | 58.2302 | 61.4542 | 97.7463 | 100.4159 | 118.6609 | 104.8762 | 73.6077 | 233.2005 | 189.8997 | 216.1151 | 498.3304 | 113.6312 | 164.2772 | 182.7995 | 416.7228 | 244.1621 | 187.6225 | 187.6992 | 394.3836
1760 | 21.3544 | 165.4396 | 195.5797 | 209.8004 | 190.7198 | 124.9405 | 143.1804 | 145.8599 | 151.1209 | 158.8214 | 133.62 | 128.2995 | 158.6007 | 186.3205 | 143.7408 | 144.739 | 232.5183 | 168.3608 | 137.1593 | 135.7793 | 174.1401 | 161.9194 | 296.6813 | 272.1401 | 162.3993 | 208.3379 | 205.1612 | 208.021 | 332.62 | 120.2399 | 117.62 | 124.8196 | 147.6603 | 129.359 | 140.261 | 140.8599 | 96.38 | 213.0806 | 207.1593 | 231.2399 | 247.522 | 119.359 | 121.9194 | 126.7198 | 103.0403 | 77.1401 | 75.4799 | 59.4799 | 111.2014 | 166.6795 | 172.2207 | 197.5009 | 152.62 | 119.1383 | 120.1804 | 133.3205 | 128.2802 | 143.9597 | 158.0595 | 204 | 159.5989 | 121.1401 | 137.0595 | 156.0403 | 134.9597 | 167.6392 | 159.9597 | 212.0788 | 244.6621 | 139.3397 | 140.0403 | 192.5815 | 192.9808 | 172.1804 | 155.6007 | 184.9002 | 368.1804 | 133.8407 | 135.7601 | 139.1401 | 218.1191 | 119.2207 | 116.4396 | 153.8388 | 142.4396 | 149.6603 | 161.4606 | 196.8196 | 113.3608 | 88.9002 | 111.9387 | 109.5604 | 110.9002 | 185.1804 | 186.2802 | 206.0998 | 341.38 | 206.7408 | 201.4203 | 206.1016 | 300.6603 | 177.3205 | 77.5797 | 73.0595 | 177.3205 | 104.62 | 112.6603 | 113.7601 | 128.3205 | 99.2399 | 91.9194 | 101.2207 | 128.0998 | 69.1996 | 66.2802 | 57.0998 | 91.2802 | 104.0192 | 100.3818 | 104.7408 | 78.3397 | 223 | 209.2592 | 223.8599 | 495.7005 | 122.979 | 170.3993 | 196.0998 | 429.8196 | 222.7005 | 187.8407 | 194.3205 | 392.3013
1761 | 21.3654 | 170.6323 | 183.8259 | 182.6066 | 193 | 139.5999 | 144.6903 | 150.7935 | 162.8644 | 143.3677 | 127.6775 | 130.8387 | 169.9676 | 188.5871 | 140.058 | 154.3677 | 244.3097 | 162.529 | 135.1484 | 137.3097 | 176.1613 | 135.1227 | 280.471 | 292.4967 | 177.5871 | 195.226 | 201.2193 | 212.8064 | 330.471 | 117.2065 | 113.5742 | 123.942 | 137.4129 | 139.3677 | 127.471 | 132.5162 | 89.6775 | 218.3806 | 215.8968 | 216.8259 | 252.9225 | 116.0904 | 121.0452 | 127.1032 | 113.2774 | 72.2065 | 73.8387 | 61 | 93.471 | 183.3806 | 178.2774 | 202.5419 | 143.5162 | 123.7226 | 115.3677 | 124.8387 | 119.6775 | 149.529 | 167.4258 | 192.6194 | 157.0904 | 122.529 | 123.0324 | 153.6323 | 127.8839 | 180.4258 | 164.8968 | 207.2969 | 250.187 | 133.1484 | 128.7807 | 207.9805 | 175.5162 | 185.7031 | 178.9805 | 181.1032 | 381.0709 | 142.1741 | 129.942 | 143.058 | 222.942 | 126.5419 | 126.6903 | 147.6646 | 147.6323 | 150.1613 | 172.7031 | 175.7098 | 134.716 | 97.7483 | 106.4001 | 98.0452 | 117.8516 | 195.5419 | 185.8968 | 211.1613 | 322.0324 | 185.3549 | 199.2645 | 194.058 | 308.1161 | 168.8387 | 77.2065 | 70.4129 | 168.8387 | 106.2645 | 120.1161 | 115.529 | 107.8259 | 107.5871 | 101.1613 | 106.0128 | 134.4258 | 78.058 | 67.1613 | 73.5419 | 87.7355 | 101.2517 | 84.1613 | 91.5742 | 84.7935 | 225.529 | 190.813 | 219.3097 | 496.9548 | 123.2517 | 181.7935 | 166.3873 | 442.8516 | 225.8516 | 186.6903 | 182.0452 | 391.6451
1762 | 21.3764 | 156.82 | 193.0473 | 180.6018 | 182.0427 | 145.8436 | 137.0427 | 145.2654 | 164.0427 | 132.0427 | 115.6209 | 121.2654 | 159.73 | 204.6018 | 151.3791 | 150.1327 | 233.91 | 154.9764 | 124.8436 | 135.2891 | 167.6873 | 142.4264 | 280.9336 | 285.0191 | 158.1518 | 187.8909 | 200.8436 | 211.6873 | 323.1991 | 104.0427 | 103.91 | 117.1327 | 135.5782 | 143.5118 | 127.2891 | 141.18 | 91.1564 | 231.4455 | 206.6873 | 224.1136 | 268 | 109.2891 | 114.4218 | 119.5545 | 115.1991 | 67.4218 | 66.8436 | 64.8673 | 106.8246 | 168.0854 | 164.73 | 202.6209 | 154.1136 | 126.9573 | 130.4691 | 128.8009 | 123.7346 | 153.5782 | 144.8627 | 184.0664 | 153.5118 | 133.0236 | 121.0236 | 157.8673 | 118.1327 | 173.3318 | 160.1991 | 206.0236 | 242.7963 | 128.6445 | 128.5118 | 202.1518 | 180.9573 | 183.1091 | 180.2418 | 193.5355 | 384.5545 | 165.8246 | 132.1564 | 147.9336 | 216.7773 | 132.4218 | 130.6445 | 135.0664 | 149.2891 | 168.1136 | 173.9764 | 168.7346 | 126.8627 | 80.7963 | 94.3555 | 89.4882 | 120.4218 | 183.3745 | 192.1564 | 218.1564 | 318.7346 | 181.8673 | 186.4645 | 209.2464 | 303.0427 | 170.8673 | 78.2227 | 66.6445 | 170.8673 | 106.3555 | 126 | 115.0664 | 110.6018 | 105.5545 | 115.2464 | 118.1564 | 118.9527 | 72.6209 | 65.1327 | 66.5309 | 80.5545 | 92.4218 | 86 | 90.9336 | 89.5782 | 226.3555 | 185.3127 | 224.3791 | 490.0427 | 120.8673 | 170.4645 | 166.0473 | 434.4645 | 234.2227 | 180.3318 | 190.2464 | 371.2845
1763 | 21.3874 | 168.7051 | 208.2547 | 190.2842 | 191.1073 | 143.6568 | 145.4504 | 150.1957 | 159.3137 | 130.5979 | 132.6756 | 114.3727 | 156.5389 | 196.5496 | 136.2949 | 148.6568 | 223.4316 | 167.7346 | 120.0295 | 128.7748 | 168.5684 | 147.8043 | 280.0295 | 297.6461 | 149.5979 | 203.5389 | 214.4209 | 206.6568 | 326.5684 | 99.3137 | 109.8526 | 124.8526 | 135.6863 | 137.4611 | 126.0295 | 143.0885 | 90.059 | 214.6381 | 189.8338 | 222.4906 | 270.6273 | 119.1957 | 114.3137 | 123.882 | 115.941 | 69.941 | 72.5389 | 62.4021 | 109.0885 | 174.3914 | 177.303 | 215.0778 | 164.3137 | 113.9517 | 117.3244 | 119.5201 | 114.2064 | 157.6273 | 159.244 | 178.4021 | 158.6273 | 120.9222 | 129.2842 | 158.6863 | 146.8713 | 151.5791 | 162.9115 | 191.2949 | 226.059 | 140.1662 | 134.941 | 210.0188 | 196.1957 | 159.5791 | 177.5684 | 192.4611 | 361.9517 | 149.697 | 135 | 149 | 218.2842 | 129.0295 | 127.059 | 120.8633 | 142.7748 | 163.8633 | 146.0107 | 165.7748 | 122.1957 | 77.8526 | 96.6273 | 100.7936 | 115.7158 | 173.9705 | 199.5979 | 204.5791 | 313.1474 | 187.941 | 207.9008 | 217.3137 | 316.0483 | 167.7453 | 68.1769 | 62.4021 | 167.7453 | 117.1662 | 107.6086 | 108.0885 | 122.9115 | 97.4021 | 104.2654 | 113.7748 | 104.4021 | 68.6568 | 65.6273 | 75.0778 | 88.1662 | 113.9893 | 86 | 88.059 | 92.9705 | 210.2359 | 206.7641 | 224.4021 | 470.8633 | 110.5201 | 169.5684 | 171.4021 | 410.5791 | 238.6273 | 188.7936 | 189.7748 | 357.3137
1764 | 21.3984 | 159.5953 | 198.9557 | 211.4047 | 211.0443 | 146.0074 | 140.2866 | 148.9778 | 154.647 | 145.0443 | 127.2718 | 108.3161 | 165.6839 | 182.3087 | 127.3383 | 142.3087 | 203.9409 | 166.294 | 129.0369 | 132.3604 | 188.0591 | 138.3161 | 281.6765 | 307.9926 | 147.9852 | 191.9409 | 218.3161 | 208.3383 | 321.9705 | 97.3235 | 110.9852 | 125.3161 | 135.6691 | 121.6248 | 129.0148 | 133.3087 | 85.9926 | 208.6839 | 198.0517 | 225.0295 | 270.6617 | 109.9483 | 120.353 | 128.6691 | 113.9852 | 72 | 74.3235 | 68.6987 | 118.0443 | 181.3235 | 187.9926 | 199.9114 | 140.2423 | 121.3973 | 118.7134 | 123.706 | 108.3383 | 149.6322 | 165.3013 | 187.3752 | 147.2866 | 148.802 | 132.3383 | 153.9852 | 139.5805 | 148.353 | 157.9705 | 200.4047 | 227.3456 | 135.294 | 145.0295 | 189.551 | 201.6691 | 168.3973 | 180.3309 | 187.3309 | 362.706 | 124.2866 | 124.9631 | 140.3013 | 220.6691 | 129.3383 | 122.9926 | 131.3973 | 149.0369 | 157.6691 | 159.4269 | 157.9852 | 129.6765 | 97.7208 | 96.6617 | 99.3013 | 110.6544 | 181.0222 | 200.6617 | 226.1108 | 302.647 | 192.0074 | 214.9705 | 224.6913 | 308.6027 | 161.6544 | 80.0665 | 73.3826 | 161.6544 | 116.3087 | 105.3604 | 96.9705 | 116.6322 | 88.9778 | 91.6544 | 94.6101 | 126.7577 | 70.3383 | 69.0074 | 78.647 | 101.3604 | 115.9631 | 92.0222 | 95.3678 | 87.3087 | 220.7356 | 208.9778 | 240.7356 | 444.5953 | 110.6913 | 172.3309 | 181.0443 | 392.6322 | 253.3826 | 181.9483 | 190.0148 | 381.4195
1765 | 21.4094 | 172.4868 | 186.8672 | 202.6017 | 211.5486 | 142.9115 | 142.8584 | 157.5841 | 152 | 155.0885 | 125.8142 | 123.0355 | 155.7345 | 194.6726 | 138.2213 | 148.5399 | 213.0355 | 167.1328 | 127.1858 | 136 | 198.0443 | 121.6902 | 277.5486 | 304.2743 | 138.5044 | 184.6814 | 187.3804 | 202.1858 | 308.4601 | 100.7699 | 101.5044 | 124.3628 | 141.4514 | 143.9382 | 122.823 | 131.3628 | 78.1858 | 206.2301 | 235.6639 | 207.1946 | 243.4247 | 108.4514 | 117.5486 | 107.876 | 119.4956 | 64.5044 | 77.0885 | 73.6814 | 116.5044 | 184.7699 | 174.7345 | 195.4956 | 159.3453 | 125.2301 | 128.4071 | 111.2831 | 120.5841 | 166.124 | 150.0973 | 205.2655 | 160.7612 | 127.4777 | 141.8584 | 136.3274 | 125.2743 | 168.7169 | 161.4956 | 217.4956 | 239.9027 | 140.9027 | 146.2743 | 196.4868 | 189.053 | 174.2743 | 154.7875 | 199.2655 | 353.6902 | 145.9382 | 135.6726 | 128.5044 | 216.2301 | 132.7257 | 128.8142 | 137.2743 | 138.3274 | 162.7699 | 158.3716 | 171.6726 | 128.9557 | 117.947 | 94.6372 | 103.8584 | 106.2743 | 193.5399 | 204.0885 | 217.1503 | 311.5841 | 195.7257 | 207.5929 | 218.4601 | 311.5841 | 173.6284 | 70.6017 | 80.6814 | 173.6284 | 103.4601 | 113.0885 | 93 | 114.7257 | 96.9027 | 92.7257 | 89.0443 | 119.9203 | 62.1416 | 70.6814 | 77.3628 | 98.8672 | 100.0973 | 90.2301 | 98.6372 | 86.0443 | 230.6814 | 194.4159 | 241.823 | 449.3098 | 122.177 | 156.3274 | 177.4601 | 411.1683 | 259.3186 | 193.3983 | 176.3274 | 387.5486
1766 | 21.4204 | 151.783 | 162.495 | 202.3245 | 186.8013 | 127.8195 | 144.2252 | 157.4503 | 155.4686 | 159.0811 | 108.8824 | 124.7566 | 161.0994 | 206.856 | 134.6755 | 141.2069 | 208.5132 | 172.7748 | 139.9553 | 145.7119 | 190.6755 | 125.4057 | 286.7931 | 299.5314 | 157.8924 | 183.6126 | 173.3063 | 208.0182 | 294.2881 | 102.3063 | 104.2434 | 113.2069 | 134.2881 | 138.2698 | 128.7119 | 134.0811 | 75.6937 | 199.8377 | 218.783 | 201.856 | 265.6855 | 122.0994 | 111.5314 | 110.4868 | 116.0629 | 83.1987 | 73.4503 | 66.3692 | 105.3692 | 177.2881 | 162.0629 | 180.9635 | 170.5314 | 131.3245 | 120.9818 | 111.3245 | 112.1258 | 145.4768 | 160.9553 | 195.0447 | 160.9818 | 127.7302 | 122.4139 | 147.7302 | 121.2252 | 152.0264 | 164.3063 | 184.9271 | 242.2252 | 130.0447 | 131.8195 | 195.9006 | 183 | 184.0994 | 176.2981 | 190.4321 | 361.5679 | 152.7566 | 152.7119 | 129.856 | 213.3063 | 119.4321 | 143.0994 | 131.144 | 125.4503 | 173.3245 | 164.4868 | 179 | 120.3692 | 100.4139 | 112.0365 | 92.0447 | 112.6308 | 201.4686 | 186.5761 | 216.4057 | 321.1623 | 202.5497 | 193.5132 | 218.856 | 319.0811 | 152.9453 | 68.9371 | 71.9818 | 152.9453 | 98.3063 | 103.9006 | 113.8113 | 115.3063 | 112.4057 | 101.6308 | 96.9371 | 116.5497 | 58.6937 | 66.8377 | 68.9818 | 105.0182 | 82.5132 | 81.7566 | 104.2434 | 83.5314 | 214.351 | 199.4057 | 226.9006 | 470.5679 | 125.3063 | 158.7119 | 181.3245 | 412.2881 | 258.3063 | 183.9635 | 169 | 377.3692
1767 | 21.4314 | 147.884 | 174.18 | 204.588 | 178.412 | 135.414 | 149.354 | 155 | 148.528 | 138.82 | 119.062 | 127.648 | 156.116 | 203.352 | 155.004 | 154.356 | 218.532 | 168.352 | 146.294 | 127.408 | 175.704 | 129.294 | 274.35 | 306.472 | 151.762 | 204.18 | 193.474 | 184.466 | 301.296 | 114.002 | 109.824 | 96.704 | 132.824 | 116.586 | 130.882 | 118.762 | 81.648 | 205.766 | 201.47 | 213.178 | 273.234 | 110.056 | 117.06 | 114.588 | 114.412 | 69.702 | 58.998 | 62.294 | 104.824 | 208.3001 | 165.354 | 187.12 | 166.176 | 108.172 | 126.884 | 116.412 | 118.708 | 159.534 | 172.942 | 177.41 | 154.882 | 120.468 | 131.768 | 148.94 | 123.53 | 173.4761 | 177.414 | 188.768 | 231.822 | 125.118 | 134.472 | 205.826 | 179.47 | 189 | 188.882 | 177.646 | 342.584 | 142.234 | 131.584 | 138.354 | 218.648 | 122.178 | 128.938 | 150.18 | 139.944 | 169.94 | 149.526 | 175.47 | 104.998 | 99.178 | 95.996 | 100.532 | 118.118 | 217.826 | 186.472 | 204.056 | 323 | 188.056 | 183.058 | 230.884 | 306.586 | 165.004 | 69.176 | 77.178 | 165.004 | 109.296 | 103.236 | 112.41 | 122.06 | 95.82 | 79.584 | 98.588 | 102.056 | 74.532 | 68.53 | 67.118 | 102.646 | 88.296 | 88.178 | 109.824 | 73.528 | 216.178 | 208.236 | 200.82 | 505.9461 | 127.824 | 162.294 | 175.822 | 424.238 | 259.412 | 186.59 | 194.416 | 406.4761
1768 | 21.4424 | 159.1829 | 170.0707 | 196.0988 | 176.1268 | 141.7183 | 150.5634 | 152.1268 | 152.9012 | 117.789 | 116.9439 | 119.9439 | 160.6195 | 205.3098 | 154.9439 | 157.6902 | 204.8878 | 154.5073 | 133.0707 | 129.4927 | 163.0988 | 135.4646 | 266.2817 | 280.55 | 155.7744 | 168.4659 | 205.5915 | 190.9573 | 311.7463 | 114.6902 | 112.4366 | 110.6756 | 124.6622 | 126.0841 | 117.0707 | 125.6476 | 89.028 | 205.4085 | 214.3658 | 216.2817 | 262.0988 | 106.5915 | 116.4085 | 118.3098 | 119.3098 | 64.3098 | 63.3378 | 58.4085 | 97.3805 | 183.4939 | 168.7183 | 181.5073 | 152.789 | 118.5488 | 119.5073 | 122.7463 | 120.4085 | 156.6342 | 152.0146 | 171.5634 | 161.9012 | 131.239 | 134.2537 | 139.5354 | 122.8451 | 173.3524 | 159.2964 | 212.0841 | 277.5622 | 136.0561 | 137.2817 | 196.9159 | 198.8305 | 192.5915 | 177.2256 | 193.6756 | 361.45 | 139 | 139.6756 | 138.8451 | 210.9439 | 104.4512 | 115.972 | 148.9439 | 139.8171 | 165.5634 | 153.211 | 181.1829 | 119.3939 | 95.0988 | 106.1122 | 109.1549 | 114.6902 | 208.9159 | 186.4085 | 230.7597 | 335.9293 | 196.0841 | 182.4366 | 217.761 | 293.0988 | 172.8451 | 57.2256 | 83.8732 | 172.8451 | 125.4927 | 99.2537 | 108 | 115.6622 | 97.7744 | 86.239 | 96.5634 | 92.8451 | 71.6622 | 66.4085 | 57.2256 | 99.2817 | 87.972 | 87.6902 | 88.0146 | 64.2537 | 219.2817 | 192.0427 | 207.8024 | 484.2403 | 121.8171 | 155.5354 | 164.0988 | 400.8317 | 251.3805 | 168.0146 | 191.3524 | 391.9866
1769 | 21.4534 | 151.0412 | 173.0364 | 206.8811 | 180.8447 | 132.5024 | 141.9636 | 144.4247 | 149.4247 | 128.3423 | 128.3423 | 117.4612 | 153.7718 | 190.9272 | 131.2742 | 148.6942 | 206.4976 | 161.6893 | 132.3835 | 131.8082 | 179.7258 | 143.1141 | 287.1869 | 279.2282 | 165.8447 | 177.301 | 215.767 | 200.1918 | 316.1918 | 102.7718 | 112.2694 | 112.1553 | 115.8859 | 144.4199 | 107.6165 | 121.5024 | 88.8082 | 206.9223 | 214.8859 | 232.0728 | 266.3058 | 126.9952 | 132.534 | 134.6117 | 124.6529 | 74.0364 | 64.8082 | 73.0728 | 105.6893 | 183.3423 | 182.8811 | 190.8811 | 143.6165 | 114.5801 | 118.6529 | 114.0412 | 110.233 | 152.4612 | 145.1918 | 173.9223 | 156.9636 | 142.3835 | 123.9636 | 142.1141 | 131.4976 | 162.1553 | 150.7306 | 207.7718 | 255.3567 | 136.3471 | 129.6942 | 192.4612 | 192.3883 | 187.4247 | 162.7718 | 189.3107 | 373 | 122.1966 | 148.4612 | 154.8034 | 190.9272 | 121.5704 | 116.9223 | 113.5849 | 134.0777 | 152.5801 | 153.6165 | 172.3107 | 116.7718 | 102.2282 | 103.7718 | 112.9223 | 123.9588 | 188.3883 | 185 | 233.0412 | 327.1189 | 203.4612 | 176.4247 | 213.9223 | 295.8447 | 164.6942 | 50.8082 | 71.8495 | 164.6942 | 125.6165 | 105.0364 | 115.3058 | 127.3423 | 102.7306 | 93 | 91.6165 | 97.8447 | 75.3058 | 62.0777 | 69.8034 | 95.3471 | 97.6893 | 85.2694 | 93.6117 | 80.2646 | 221.9223 | 185 | 222.0364 | 471 | 141.6481 | 159.5753 | 161 | 415.301 | 246.5388 | 176.534 | 198.4199 | 397.0728
1770 | 21.4644 | 154.4287 | 181.943 | 194.9139 | 168.1425 | 138.6574 | 142.7144 | 143.4857 | 157.4002 | 141.4287 | 128.057 | 111.3141 | 161.1431 | 197.6289 | 119.5428 | 144.5143 | 224.1146 | 161.5428 | 126.5713 | 131 | 169.171 | 137.5713 | 259.342 | 286.7144 | 164.2856 | 198.1431 | 192.9994 | 198.0285 | 340.0291 | 96.7714 | 112 | 97.3711 | 120.6859 | 135.6283 | 129.0291 | 141.772 | 88 | 218.4002 | 202.5998 | 228.3426 | 254.8854 | 116.9139 | 124.8854 | 117.7138 | 115.5998 | 77 | 77.3717 | 61.9139 | 111.4857 | 183.057 | 171.6568 | 222.7435 | 161.3147 | 121.1431 | 117.7714 | 124.1146 | 126.3147 | 142.5998 | 157.1431 | 185.4002 | 150.2856 | 119.4851 | 133.6289 | 162.5719 | 130.5428 | 151.0855 | 151.7429 | 191.3711 | 237.7714 | 143.9145 | 138.886 | 198.943 | 192.2001 | 193.1716 | 178.3147 | 178.3141 | 378.2001 | 142.0006 | 163.8575 | 129.7993 | 205.8005 | 116.1425 | 144.7435 | 124.2862 | 155.2862 | 130.9139 | 149.7714 | 167.2571 | 115.2286 | 107.4857 | 91.0855 | 99.1425 | 110.171 | 183.7429 | 187.2286 | 204.4851 | 301.9424 | 188.3996 | 188.1146 | 209.7999 | 332.9151 | 166.4572 | 58.1716 | 91.5149 | 166.4572 | 113.5998 | 100.5713 | 123.943 | 121.114 | 110.4287 | 100.4287 | 107.829 | 93.3141 | 86.9145 | 66.943 | 67.8284 | 88.7999 | 102.7429 | 97.6289 | 97.5143 | 75.8569 | 216.3141 | 213.9721 | 215.3426 | 480.6574 | 126.228 | 161.2571 | 172.886 | 391.5707 | 253.4287 | 168.1425 | 176.9994 | 355.456
1771 | 21.4754 | 146.4277 | 188.531 | 190.5103 | 163.7552 | 138.2242 | 154.5723 | 144 | 151.1828 | 128.8967 | 97.3037 | 104.469 | 158.9587 | 188.407 | 135.3688 | 140.2242 | 217.6725 | 155.469 | 137.593 | 127.2242 | 182.6343 | 128.9587 | 277.2066 | 300.0826 | 149.407 | 190.6725 | 182.4897 | 203.7965 | 301.9349 | 113.3688 | 96.1415 | 108.8585 | 128.2862 | 118.1622 | 124.1622 | 131.1209 | 75.1622 | 197.8347 | 199 | 224.2448 | 275.6756 | 116.2862 | 113.9587 | 109.2448 | 126.3481 | 72.469 | 78.9793 | 74.8791 | 105.9587 | 174.9587 | 173.5517 | 209.345 | 150.6312 | 122.7345 | 114.7345 | 113.1415 | 121.6725 | 141.2655 | 154.9587 | 176.1622 | 145.2242 | 132.8998 | 147.062 | 160.6932 | 132.0207 | 171.4101 | 157.2862 | 205.8791 | 243.7965 | 126.6105 | 146.7758 | 201 | 215.8998 | 196 | 171.407 | 169.9587 | 363.3864 | 148.7345 | 126.7107 | 129.5723 | 205.4483 | 123.8378 | 129.0795 | 126.7138 | 147.1415 | 143.124 | 142.2035 | 156.4277 | 117.5103 | 101.9587 | 86.4897 | 100.0413 | 121.3688 | 191.5517 | 210.655 | 205.062 | 321.4308 | 184.5103 | 188.469 | 201.9587 | 312.5279 | 163.469 | 68.5517 | 88.6725 | 163.469 | 110.7552 | 108.5723 | 107.1209 | 107.938 | 95.6312 | 89.407 | 114.7552 | 105.3481 | 78.6725 | 64.469 | 72.5517 | 88.5103 | 104.5103 | 92.1828 | 90.9587 | 79.5517 | 222.3068 | 196.814 | 229.3688 | 503.6343 | 121.0207 | 167.7965 | 177 | 390.5723 | 251.469 | 177.3481 | 170.2655 | 361.655
1772 | 21.4864 | 143 | 170.8136 | 208.6515 | 164.7675 | 126.2556 | 140.3485 | 170.0935 | 148.7675 | 113.2556 | 111.7912 | 106.8373 | 168.5118 | 171.7207 | 142.5349 | 142.0698 | 204.0231 | 169.3491 | 126.6509 | 135.2095 | 188.2325 | 133.1397 | 278.558 | 300.9302 | 146.5349 | 178.558 | 224.9776 | 210.6047 | 309.2562 | 109.0231 | 95.6047 | 111.6976 | 134.6047 | 131.6515 | 119.2325 | 118.8603 | 80.9769 | 219.9308 | 205.9071 | 236.2793 | 257.9065 | 110.3254 | 113.5349 | 117.442 | 124.0929 | 83.2793 | 76.4651 | 75.6278 | 118.5817 | 153.8136 | 169.8603 | 188.1858 | 144.2325 | 108.1858 | 97.8834 | 117.9769 | 128.7444 | 146.6047 | 149.9302 | 194.2562 | 159.3491 | 110.8367 | 136.1858 | 146.4882 | 153.7213 | 157.5112 | 142.8834 | 181.3017 | 231.4183 | 127.6746 | 129.581 | 202.5349 | 207.6509 | 186.7905 | 161.6278 | 195.6284 | 361.8373 | 144.1627 | 155.9776 | 118.4183 | 219.1166 | 134.1397 | 138.6515 | 133.442 | 124.3485 | 136.7207 | 157.6515 | 153.7675 | 122.6047 | 109.2095 | 89.0698 | 106.6047 | 95.5343 | 204.7444 | 191.139 | 196.4882 | 348.419 | 182.6976 | 187.7675 | 200.7675 | 305.0698 | 146.6509 | 54.8834 | 71.9532 | 146.6509 | 98.7207 | 109.6976 | 98.6976 | 98.8603 | 92.3024 | 97.2793 | 109.6278 | 114.6047 | 61.1858 | 63.7675 | 86.5118 | 99.7444 | 93.4882 | 92.8373 | 76.7207 | 74.3254 | 238.0468 | 158.8367 | 215.8136 | 502.3254 | 125.8373 | 170.7675 | 156.2787 | 418.5586 | 236.1858 | 169.7207 | 187.884 | 362.0929
1773 | 21.4974 | 149.238 | 167.3392 | 201.524 | 162.6608 | 117.5418 | 125.643 | 157.7266 | 151.3392 | 128.7139 | 115.881 | 125.9342 | 160.3038 | 174.238 | 135.9823 | 149.238 | 221.2734 | 149.8278 | 108.7443 | 127.8633 | 181.762 | 129.5418 | 278.3392 | 314.8152 | 164.1544 | 182.238 | 197.4531 | 218.238 | 324.5772 | 112.238 | 92.3215 | 121.9165 | 132.881 | 149.476 | 123.6785 | 127.9165 | 92.5772 | 206.3873 | 217.357 | 220.5063 | 257.0177 | 111.119 | 107.762 | 121.5595 | 101.7266 | 59.7088 | 66.643 | 70.1013 | 108.1848 | 155.7975 | 191.3924 | 179.3215 | 143.2203 | 110.238 | 114.8329 | 110.0835 | 108.6076 | 151.8987 | 141.2025 | 196.3215 | 160.1013 | 126.9519 | 120.3038 | 146.119 | 145.9645 | 143.9823 | 152.0355 | 186.8152 | 251.9519 | 134.6785 | 142.7139 | 195.9823 | 192.0835 | 175.4228 | 174.0355 | 192.3038 | 367.6785 | 131.3038 | 147.9468 | 157.6659 | 204.5063 | 118.0658 | 124.5063 | 132.1013 | 139.2734 | 163.4102 | 147.405 | 154.7797 | 121.6608 | 87.0481 | 95.4582 | 101.762 | 117.1899 | 202.5418 | 175.2025 | 192.2203 | 326.7088 | 198.3747 | 202.0355 | 209.5772 | 304.4405 | 149.7975 | 68.7139 | 78.9165 | 149.7975 | 128.5291 | 84.8278 | 91.762 | 103.238 | 120.2912 | 101 | 103.3215 | 114.4405 | 64.0177 | 60.881 | 77.524 | 71.0304 | 113.3924 | 85.4228 | 76.8987 | 80.5772 | 213.1493 | 171.0532 | 204.5418 | 471.9291 | 132.4582 | 168.6608 | 153.119 | 418.9823 | 231.2203 | 155.8633 | 158.6911 | 377.9342
1774 | 21.5084 | 150.208 | 173.5443 | 204.3363 | 176.2567 | 132.6328 | 124.5841 | 155.1682 | 167.0487 | 118.1592 | 129.2567 | 122.2876 | 151.4557 | 177.5841 | 136.3761 | 147.0398 | 208.7831 | 154.8806 | 120.0487 | 118.6637 | 183.1682 | 136.7124 | 295.6328 | 286.5264 | 146.8229 | 180.0398 | 189.1682 | 225.5443 | 308.7831 | 106.8717 | 98.9204 | 102.8229 | 145.4647 | 138.7433 | 113.1194 | 147.6328 | 98.9602 | 197.6239 | 227.1283 | 207.8717 | 251.0796 | 121.5045 | 116.2965 | 144.9691 | 123.7214 | 71.801 | 67.1682 | 70.5841 | 123.0089 | 179.385 | 172.6548 | 199.385 | 157.2567 | 100.1194 | 114.6637 | 126.801 | 109.1283 | 160.1283 | 128.7035 | 191.8318 | 171.6726 | 138.7522 | 126.5045 | 142.2478 | 138.8318 | 145.9602 | 152.8318 | 191 | 236.8229 | 139.9602 | 123.4468 | 213.801 | 207.2169 | 177.7522 | 171.6637 | 180.2876 | 360.2876 | 132.7522 | 141.208 | 155.7433 | 211.6726 | 121.7124 | 124.5443 | 119.9115 | 146.5841 | 146.6548 | 132.7035 | 158.9602 | 131.2965 | 83.1682 | 89.0796 | 95.2478 | 114.1194 | 201.792 | 165.8717 | 245.0666 | 302.3672 | 193.4955 | 191.7433 | 207.2478 | 299.2478 | 150.4159 | 71.6239 | 88.3363 | 150.4159 | 114.2388 | 100.1771 | 88.4159 | 113.7124 | 96.3185 | 103.3761 | 99.6239 | 115.5841 | 69.9602 | 62.3761 | 66.8717 | 88.1373 | 95.4468 | 116.2656 | 85.9204 | 88.5443 | 227.9691 | 199.1771 | 192.7035 | 502.8099 | 107.0707 | 161.6637 | 179.3452 | 402.7433 | 234.9602 | 175.9691 | 175.1373 | 362.407
1775 | 21.5194 | 158.0433 | 170.9784 | 191.5221 | 171.1524 | 142.6303 | 129.0216 | 160.0216 | 179.8476 | 133.3039 | 101.6311 | 120 | 158.0433 | 174.7827 | 149.8693 | 134.7394 | 205.6087 | 154.7827 | 130.4346 | 149.9775 | 175.9567 | 122.9134 | 273.4571 | 317.4122 | 150.652 | 175.7827 | 178.7394 | 206.8918 | 329.7385 | 109.0216 | 122.7169 | 83.3264 | 132.9134 | 118.9134 | 118.8476 | 130.2831 | 93.5654 | 209.8693 | 230.6087 | 234.1515 | 267.4996 | 115.1524 | 119.8043 | 131.6961 | 116.5221 | 77.8043 | 74.4346 | 69.3913 | 110.3047 | 189.0216 | 187.7169 | 199.3697 | 149.7394 | 117.9126 | 124.2606 | 110.2831 | 111 | 157.174 | 148.5212 | 180.5437 | 175.8043 | 122.3047 | 117.7394 | 132.1524 | 141.2173 | 138.9567 | 168.0866 | 195.826 | 254.3256 | 142.6087 | 126.652 | 194.0658 | 209.587 | 179 | 175.6303 | 197.3039 | 378.1082 | 153.3039 | 127.3264 | 146.3697 | 207.761 | 124.8043 | 121.9784 | 112.174 | 136.5437 | 159.3039 | 133.2173 | 171.2606 | 110.6744 | 91.239 | 91.0216 | 94 | 102.1524 | 190.7394 | 175.2606 | 222.0008 | 348.6728 | 193.413 | 204.0866 | 215.652 | 299.6087 | 150.8043 | 65.3697 | 67.4788 | 150.8043 | 108.8043 | 97.1524 | 115.3472 | 115.1957 | 106.4996 | 94.348 | 105.4346 | 103.9351 | 82.2606 | 76.6736 | 85.9126 | 96.6087 | 102.6736 | 96.0441 | 92.0216 | 75.5221 | 217.9134 | 198.5654 | 212.5212 | 496.9134 | 124.1299 | 166.4346 | 195.652 | 403.826 | 237.6087 | 181.1957 | 182 | 390.7818
1776 | 21.5304 | 149.3839 | 180.6161 | 209.2321 | 175.533 | 141.5501 | 143.0659 | 136.5014 | 146.8853 | 155.9657 | 108.6992 | 129.7994 | 161.6332 | 176.4499 | 135.851 | 148.3324 | 206.8166 | 164.6161 | 149.9657 | 133.5014 | 169.1003 | 144.3152 | 282.5158 | 291.8853 | 155.4499 | 194.5989 | 180.0831 | 222.4155 | 327.8338 | 118.9828 | 113.3008 | 84.8997 | 115.1174 | 118.2665 | 132.4327 | 132.3496 | 78.9341 | 211.3668 | 214.6676 | 257.3324 | 280.9828 | 102.3839 | 122.4499 | 137.6161 | 116.2665 | 74.7335 | 67.0172 | 68.1834 | 103.5501 | 172.0343 | 188.1003 | 187.3839 | 154.3496 | 114.8338 | 118.8338 | 131.9485 | 112.6332 | 138.851 | 142.5673 | 172.2837 | 154.7679 | 144.9485 | 132.149 | 152.8654 | 128.1174 | 145.9828 | 162.2006 | 200.2665 | 264.0831 | 135.6504 | 124.9169 | 190.4499 | 201.6504 | 190.4327 | 169.6504 | 209.3496 | 356.8681 | 117.169 | 133.7994 | 147.4499 | 210.0831 | 103.7679 | 103.851 | 131.4155 | 136.4499 | 164.8166 | 166.6648 | 157.6676 | 120.5158 | 92.1834 | 83.8338 | 94.8166 | 111.4327 | 189.6332 | 188.6161 | 204.8338 | 333.2349 | 191.5501 | 190.851 | 214.7335 | 295.1003 | 155.0831 | 54.2006 | 68.533 | 155.0831 | 107.3668 | 93.3668 | 115.467 | 124.7994 | 102.8338 | 91.1834 | 115.1662 | 93.6504 | 62.9512 | 63.6676 | 76.3008 | 83.1174 | 86.4011 | 85.7335 | 89.7335 | 84.2493 | 209.9169 | 186.3839 | 221.2665 | 448.0859 | 130.8166 | 183.5158 | 153.9025 | 403.3668 | 243.7163 | 202.2321 | 198.3324 | 388.3839
1777 | 21.5414 | 145.3422 | 183.8289 | 195.7639 | 164.5663 | 144.3156 | 137.7109 | 142.6048 | 161.3806 | 151.7109 | 112 | 118.7374 | 155.3687 | 172.0265 | 150.2361 | 152.8289 | 212.8024 | 144.6194 | 132.4483 | 113.9085 | 163.0265 | 133.4218 | 284.3422 | 269.9085 | 137.7639 | 204.8024 | 200.8939 | 224.5133 | 320.1976 | 98.6194 | 123.2626 | 105.065 | 131.065 | 132.2626 | 130.0265 | 132.3422 | 85.118 | 200.2241 | 212.6578 | 227.8435 | 267.2507 | 114.0915 | 109.7374 | 133.3687 | 107.0531 | 65.7109 | 69.1446 | 71.3156 | 97.1976 | 164.6844 | 177.882 | 189.9735 | 161.8024 | 121.2891 | 110.3687 | 128.882 | 122.9469 | 152.4071 | 128.3952 | 190.065 | 161.6048 | 136.0796 | 133.5133 | 129.817 | 129.1446 | 120.6459 | 155.0265 | 174.4748 | 239.3037 | 131.5133 | 134.7759 | 191 | 186.7374 | 165.6459 | 168 | 181.9881 | 332.7639 | 150.2745 | 120.2507 | 150.4867 | 235.8674 | 125.5252 | 101.6578 | 127.7109 | 170.1565 | 147.5929 | 140.8435 | 174.7228 | 127.3156 | 103.6048 | 93.6048 | 102.4602 | 102.3952 | 220.6697 | 191.8289 | 188.9085 | 295.5014 | 215.8674 | 180.3687 | 220.6313 | 292.3422 | 161.8024 | 61.9469 | 84.9204 | 161.8024 | 104.5133 | 101.2891 | 100.7374 | 130.3156 | 100.1711 | 89.3422 | 93.7905 | 110.2361 | 65.4602 | 72.4337 | 70.5133 | 99.065 | 99.4071 | 93.2891 | 92.3156 | 76.2241 | 210.6578 | 188.1446 | 213.7109 | 496.0238 | 126.0265 | 158.817 | 153.9469 | 394.7109 | 245 | 153.1207 | 173.817 | 359.4748
1778 | 21.5524 | 162.6653 | 189.0472 | 207.1417 | 171.2532 | 155.9356 | 140.206 | 150.8884 | 171.8884 | 145.794 | 127.9829 | 116.8412 | 154.8412 | 216.425 | 145.588 | 152.1588 | 193.8111 | 155.1417 | 128.8412 | 116.8884 | 177.1417 | 152.7125 | 252.0343 | 295.6009 | 146.618 | 195.0644 | 203.3176 | 202.1287 | 306.382 | 132.6953 | 96.5579 | 110.6824 | 122.382 | 129.1116 | 126.4764 | 137.0472 | 77.7468 | 208.9356 | 223.9356 | 243.7125 | 248.0171 | 111.9528 | 112.0472 | 115.1759 | 124.3477 | 69.0472 | 70.8412 | 63.588 | 103.5708 | 176.618 | 167.588 | 203.618 | 155.4292 | 119.6352 | 109.8412 | 105.1287 | 104.8111 | 148.4292 | 134.412 | 194.8412 | 156.4292 | 122.0644 | 131.3176 | 146.7125 | 119.9056 | 141.0773 | 164.0944 | 172.3648 | 221.382 | 124.2704 | 136.1588 | 222.9657 | 186.5236 | 172.618 | 176.412 | 198.7125 | 381.1546 | 148.9056 | 138.8713 | 129.1287 | 215.7639 | 114.1759 | 113.7768 | 135.2532 | 161.8583 | 166.7125 | 144.0944 | 168.9056 | 123.794 | 85.8111 | 86.7468 | 89.6996 | 100.8412 | 223.6352 | 197.0472 | 201.9829 | 316.7597 | 181.4635 | 181.5236 | 206.8583 | 303.7768 | 145.3347 | 73.2532 | 73.6996 | 145.3347 | 99.794 | 108.0472 | 111.4592 | 125.9528 | 97.4764 | 91.5236 | 101.618 | 114.8412 | 65.3176 | 73.3176 | 60.7468 | 97.9528 | 89.5408 | 87.4292 | 103.0944 | 82.412 | 208.4764 | 182.2704 | 194.3347 | 497.0644 | 127.5236 | 172.3477 | 172.8241 | 401.412 | 257.618 | 171.4421 | 169.6824 | 372.5065
1779 | 21.5634 | 157.4651 | 160.1277 | 189.5161 | 186.6559 | 146.0511 | 146.9745 | 145.172 | 143.1277 | 127.9301 | 123.3186 | 110.172 | 144.7581 | 180.6183 | 134.6116 | 146.879 | 222.4328 | 150.3186 | 155.4583 | 118.8535 | 192.8024 | 144.1976 | 279.2863 | 301.8535 | 153.2675 | 204.9489 | 192.7581 | 219.3373 | 304 | 119.5161 | 112.3373 | 91.3696 | 133.6559 | 110.0766 | 131.121 | 120.0766 | 90.5094 | 201.6116 | 194.4207 | 218.1277 | 277.4328 | 98.1976 | 116.414 | 124.8024 | 114.3441 | 68.293 | 56.4906 | 69.6814 | 99.879 | 160.2231 | 161.7325 | 198.3186 | 157.414 | 144.6048 | 132.1908 | 107.828 | 107.828 | 144.4395 | 147.0954 | 182.1976 | 138.7836 | 125.9745 | 138.6814 | 158.6814 | 127.3884 | 147.707 | 167.707 | 180.6814 | 224.9745 | 134.0954 | 136.8535 | 199.1277 | 189.5605 | 169.879 | 185.6814 | 192.7581 | 355.1533 | 137.6116 | 131.9046 | 151.4583 | 243.4328 | 133.1908 | 108.3186 | 124.1976 | 157.293 | 142.8347 | 174.1653 | 171.2675 | 128.121 | 94.8024 | 107.1908 | 82.7325 | 109.5349 | 193.9812 | 181.7836 | 198.172 | 349.3118 | 202.1653 | 198.2164 | 215.0954 | 291.4906 | 154.8024 | 54.5161 | 76.121 | 154.8024 | 122.8978 | 114.9745 | 107.172 | 103.6627 | 93.586 | 97.9745 | 93.7581 | 110.7325 | 71.828 | 70.4395 | 68.3884 | 97 | 99.8024 | 80.879 | 93.0511 | 84.8535 | 218.2419 | 181 | 208.9234 | 481.3441 | 128 | 174.293 | 180.2675 | 400.4395 | 262.5605 | 159.0766 | 174.2675 | 358.9301
1780 | 21.5744 | 123.966 | 170.5841 | 195.5236 | 173.4159 | 137.0737 | 132.4159 | 153.5236 | 142.3289 | 133.6578 | 131.5236 | 109 | 159.4499 | 190.3157 | 139.0605 | 139.9395 | 240.121 | 151.5974 | 121.0397 | 136.3157 | 170.758 | 123.8185 | 291.0605 | 310.6578 | 139.2817 | 209.5974 | 184.0737 | 209.1474 | 333.4367 | 95.2212 | 111.6711 | 94.0605 | 110.8922 | 113.0605 | 120.7448 | 127.3894 | 99.9263 | 220.7788 | 225.363 | 224.2552 | 276.0737 | 122.8393 | 109.2079 | 120.9395 | 130.1815 | 70.5974 | 57.4631 | 61.4764 | 100.7316 | 174.3157 | 168.7921 | 179.6843 | 160.5974 | 116.1002 | 99.637 | 112.4631 | 109.8658 | 137.0737 | 140.3422 | 182.5974 | 131.6711 | 134.7921 | 121.8185 | 156.5369 | 124.6711 | 135.0132 | 155.879 | 172.4764 | 240.7183 | 131.6711 | 130.9395 | 196.5974 | 202.9868 | 180.2552 | 188.7316 | 189.2684 | 357.6578 | 129.9395 | 131.7316 | 148.2079 | 216.966 | 120.5501 | 135.5709 | 132.3894 | 131.8922 | 119.8185 | 152.1607 | 172.8658 | 129.8658 | 96.1342 | 104.9395 | 99.3157 | 97.1474 | 182.0737 | 196.3157 | 182.2817 | 293.3951 | 189.6843 | 211.3894 | 212.6711 | 289.8658 | 149.2079 | 57.0605 | 69.2079 | 149.2079 | 121.8053 | 99.5501 | 119.8526 | 133.7656 | 99.0605 | 91.2079 | 98.9263 | 107.4026 | 71.2684 | 72.5974 | 81.2552 | 84.0132 | 92.4764 | 90.3894 | 106.5841 | 82.4026 | 194.8922 | 184.4631 | 205.0737 | 483.3289 | 128 | 166.2079 | 179.2684 | 414.7183 | 228.3686 | 170.7183 | 185.3894 | 370.7183
1781 | 21.5854 | 133.0492 | 190.5615 | 197 | 192.9519 | 137.7562 | 137.9027 | 148.8535 | 146.5123 | 157.83 | 112.8043 | 106.3658 | 167.2685 | 190.3658 | 115.4139 | 139.8781 | 201.6085 | 131.8043 | 136.9519 | 166.2204 | 183.6834 | 128.9027 | 277.9508 | 297.0727 | 144.0246 | 172.2427 | 197.9273 | 208.7562 | 320.4385 | 98.1465 | 100.4631 | 91.4877 | 122.8054 | 128.0492 | 125.1465 | 127.2438 | 99.2438 | 222.2438 | 194.9989 | 240.9273 | 282.0246 | 114.7069 | 107.1219 | 112.0973 | 138.2685 | 77.1465 | 65.9027 | 66.1465 | 112.415 | 172.6096 | 173.5123 | 214.7573 | 146.0727 | 118.0246 | 133.5134 | 104.2192 | 92.4385 | 139.5123 | 126.7069 | 191.7808 | 145.9273 | 135.1219 | 133.9273 | 165.6588 | 138.9273 | 158.4642 | 148.7315 | 187.6834 | 228.9508 | 138.9027 | 140.5369 | 193.4877 | 189.1946 | 177.6096 | 178.4631 | 181.9754 | 350.2192 | 137.7808 | 121.4631 | 119.7796 | 218.1465 | 118 | 126.8289 | 127.8535 | 142.9273 | 149.4888 | 170.83 | 158.9508 | 114.1946 | 89.8535 | 79.4139 | 100.2438 | 115.1957 | 195.9273 | 183.1946 | 212.4888 | 316.4888 | 215.0984 | 202.4631 | 210.2438 | 298.7808 | 159.415 | 72.9273 | 69.7562 | 159.415 | 104.3166 | 89.0973 | 107.0727 | 100.3647 | 87.7069 | 97.9027 | 94.7315 | 124.5615 | 64.8535 | 81.7808 | 75.9754 | 87.2685 | 81.3412 | 84.9754 | 89.6823 | 94.2931 | 212.0738 | 177.0973 | 224.1957 | 484 | 113.9508 | 194.8546 | 177.2438 | 403.8289 | 237.9273 | 168.6096 | 156.2673 | 354.5604
1782 | 21.5964 | 132.3289 | 176.9734 | 195.2193 | 162.1661 | 139.7807 | 136.3289 | 150.6711 | 155.0133 | 127.1661 | 118.9037 | 119.3555 | 149.3023 | 189.1096 | 129.8073 | 137.3289 | 188.8771 | 157.4917 | 135.5482 | 154.8638 | 189.5615 | 135.3422 | 283.1229 | 290.5482 | 153.0133 | 183.917 | 191.9867 | 221.4651 | 307.3156 | 106.1229 | 109.6844 | 119.4917 | 116.0963 | 112.1927 | 117.9867 | 116.3156 | 81.1927 | 215.7674 | 195.3422 | 224.3023 | 250.0564 | 123.6844 | 108.7807 | 132.3688 | 102.495 | 63.7541 | 62.5482 | 73.2326 | 97.9734 | 181.794 | 165.9867 | 178.1527 | 168.0399 | 120.7807 | 125.6445 | 107.4518 | 116.711 | 154.2459 | 141.917 | 174.3023 | 160.4651 | 118.083 | 126.206 | 150.9734 | 108.9468 | 142.4119 | 150.6711 | 184.6578 | 218.9867 | 133.7674 | 127.7541 | 205.4651 | 209.2592 | 187.6844 | 173.4385 | 180.1096 | 378.3821 | 131.8771 | 155.6147 | 145.3821 | 207.4252 | 136.6977 | 125 | 127 | 141.4385 | 149.5482 | 163.3156 | 172.1362 | 136.0399 | 81.8771 | 108.9436 | 101.7807 | 116.2193 | 193.5482 | 180.1096 | 219.6711 | 306.7541 | 197.6312 | 194.7674 | 183.289 | 300 | 153.8771 | 55.4119 | 64.6578 | 153.8771 | 103.7807 | 92.4518 | 121.0266 | 112.8073 | 90.4518 | 95.4385 | 82.4252 | 110.083 | 65.7807 | 60.7408 | 76.7807 | 107.5881 | 104.9303 | 87.5615 | 96.794 | 75.5216 | 214.1096 | 204.4917 | 228.7807 | 469.7541 | 120.0133 | 182.9734 | 154.7408 | 433.1628 | 240 | 168.8904 | 174.2592 | 367.1362
1783 | 21.6074 | 154.5665 | 191.2478 | 199.5133 | 177.8585 | 144.5133 | 129.6814 | 159.1239 | 160.5133 | 135.6372 | 112.7787 | 126.4159 | 153.3186 | 184.4867 | 137.4159 | 157.7611 | 234.9382 | 148.3628 | 123.2654 | 136.7522 | 178.2654 | 148.6372 | 264.1415 | 282.7787 | 144.0708 | 188.708 | 194.6106 | 228.4159 | 298.7787 | 95.2654 | 97.4601 | 107.6548 | 115.9027 | 127.1505 | 117.9027 | 126.7346 | 92.5399 | 210.4867 | 193.292 | 238.2478 | 271.2744 | 113.2654 | 114.4159 | 107.9203 | 117.8585 | 75.5399 | 54.7787 | 62.2654 | 86.0708 | 176.6814 | 171.3186 | 196.4691 | 168.292 | 121.9027 | 102.3362 | 112.5133 | 120.9027 | 142.4601 | 145.8053 | 180.1239 | 134.9203 | 125.9292 | 115.0708 | 152.6106 | 146.5223 | 157.1505 | 160.0266 | 171.3628 | 217.0973 | 133.9027 | 133.2213 | 175.407 | 191.2389 | 189 | 152.2389 | 166.4601 | 358.5309 | 131.9027 | 144.6548 | 157.1239 | 223.1505 | 127.2654 | 138.5399 | 115.2654 | 138.292 | 151.708 | 137.6282 | 157.7522 | 129.0708 | 90.0266 | 90.4335 | 102 | 123.2213 | 198.4159 | 200.7611 | 186.6017 | 352.8409 | 185.9734 | 194 | 181.8053 | 290.0708 | 171.9558 | 62.9292 | 73.0266 | 171.9558 | 109.4159 | 110.1505 | 102.2389 | 105.0708 | 92.8053 | 91.3894 | 90.0266 | 114.3186 | 52.4601 | 63.4159 | 62.5575 | 94.6548 | 101.6814 | 90.708 | 89.8761 | 73.9027 | 193.2389 | 210.708 | 241.6372 | 474.3186 | 134.5399 | 204.4691 | 162.8319 | 412.6282 | 246.3186 | 179.8319 | 165.2654 | 366.292
1784 | 21.6184 | 139.616 | 203.0644 | 196.3402 | 177.2552 | 136.7655 | 150.0438 | 140.7861 | 152.7655 | 124.1907 | 108.3402 | 122.4253 | 159.4897 | 201.384 | 127.0206 | 142.616 | 207.9768 | 136.0206 | 119.2552 | 116.701 | 169.6804 | 135.3608 | 289.4485 | 307.6186 | 145.7448 | 214.6186 | 186.7655 | 218.0206 | 298.9149 | 94.9149 | 94.1701 | 106 | 133.384 | 142.7242 | 127.1495 | 147.2139 | 94.9149 | 183.4665 | 216.7887 | 240 | 199.8891 | 108.3402 | 117.7448 | 124.2139 | 117.2552 | 77 | 64.9794 | 73.8093 | 102.384 | 164.1057 | 177.4897 | 188.9356 | 170.7448 | 115.5954 | 110.0644 | 109.3402 | 122.8299 | 154.7242 | 151.4897 | 156.2964 | 139.3196 | 126.0851 | 140.5335 | 139.2758 | 136.3608 | 153.5103 | 171.9794 | 200.1933 | 225.2345 | 129.4253 | 130.3402 | 190.299 | 178.9356 | 192.6598 | 171.0438 | 176.8943 | 371.5541 | 139.3196 | 149.4046 | 120.4871 | 226.8299 | 135.1495 | 121.701 | 123.1495 | 148.0644 | 169.384 | 160.6186 | 139.5309 | 136.2345 | 97.4046 | 99.8943 | 81.8711 | 114.8505 | 204.4897 | 192.0206 | 192.1495 | 323.2319 | 199.6392 | 186.6804 | 221.3428 | 298.1495 | 142.8917 | 65.8299 | 76.7448 | 142.8917 | 90.7861 | 105.5954 | 102.7448 | 102.1701 | 97.5747 | 102.8943 | 99.2345 | 97.616 | 74.7887 | 45.701 | 58.2552 | 107.6392 | 105.5747 | 102.8943 | 78.9356 | 74.9149 | 214.7887 | 196.3608 | 214.6366 | 467.6804 | 127.7655 | 178.6366 | 198.7681 | 426.4691 | 236.0206 | 168.1907 | 164.9149 | 405.3428
1785 | 21.6294 | 138 | 162.2742 | 170.9645 | 178.8545 | 129.5093 | 147.3638 | 151.0541 | 157.5634 | 137.8358 | 114.4907 | 114.5821 | 162.7817 | 152.0018 | 152.89 | 147.4907 | 213.3452 | 150.7631 | 113.4366 | 135.3993 | 163.4366 | 145.1269 | 288.291 | 290.528 | 162.6903 | 191.9645 | 179.5093 | 203.0914 | 301.7817 | 104.2724 | 102.3452 | 116.1996 | 122.9459 | 117.11 | 118.7276 | 134.1642 | 91.291 | 201.3993 | 206.9459 | 222.3824 | 222.6717 | 127.472 | 111.5093 | 117.6548 | 117.9272 | 79.7817 | 53.0186 | 72.2183 | 94.7276 | 166.709 | 181.709 | 215.8172 | 160.8004 | 122.4179 | 125.8358 | 116.4179 | 114.6548 | 132.819 | 133.4552 | 167.9086 | 165.9627 | 100.9645 | 127.2369 | 143.5634 | 150.7631 | 154.8545 | 168.3638 | 178.8918 | 226 | 121.5821 | 125.3638 | 204.0541 | 203.9627 | 187.4366 | 176.709 | 173.3638 | 337.7649 | 128.8731 | 142.5821 | 140.181 | 226.0728 | 116.528 | 140.3993 | 144.3993 | 152.709 | 168.2183 | 161.1455 | 158.3993 | 125.8731 | 103.5634 | 88.0186 | 104.1082 | 110.291 | 191.0914 | 205.8358 | 189.291 | 307.0186 | 213.0541 | 202.6903 | 222.2183 | 276.7462 | 136.291 | 66 | 52.8918 | 136.291 | 109.3993 | 80.8918 | 116.9086 | 125.181 | 93.3638 | 86.3824 | 87.0186 | 108.9814 | 70.5093 | 59.7631 | 59.8545 | 91.3824 | 97.6548 | 91.9459 | 119.7258 | 74.0728 | 220.709 | 204.2724 | 209.2183 | 467 | 101.9645 | 177.8545 | 175.11 | 420.5821 | 255.3993 | 162.3638 | 177.9814 | 381.1828
1786 | 21.6404 | 145.5163 | 170.2744 | 181.2139 | 148.935 | 131.8186 | 139.4837 | 159.5163 | 174.9116 | 113.6327 | 121.5767 | 133.7302 | 155.4837 | 203.4323 | 140.907 | 167.7302 | 218.6977 | 153.8791 | 129.9116 | 139.8186 | 193.065 | 152.5767 | 248.5397 | 309.6697 | 164 | 205.972 | 181.8186 | 202 | 307.6372 | 115.3348 | 111.4558 | 101.028 | 117.3023 | 133.7906 | 129.2744 | 129.2419 | 97.5767 | 218.972 | 184.3908 | 261.3998 | 211.8466 | 112.0884 | 126.0325 | 117.9395 | 114.2419 | 83.7581 | 83.9441 | 66.3628 | 90.2419 | 154.7861 | 181.0605 | 192.6327 | 133.6931 | 114.5442 | 120.4233 | 112.3023 | 111.1814 | 149.7906 | 135.7581 | 181.2139 | 158.6047 | 133.7627 | 112.8466 | 159.0325 | 123.8141 | 141.8466 | 176.4558 | 186.3953 | 241.0325 | 115.3628 | 126.8791 | 180.5722 | 234.1859 | 178.5442 | 191.093 | 170.1814 | 359.4278 | 155.2464 | 152.3348 | 139.1814 | 226.9395 | 124.3953 | 110.0559 | 130.9675 | 135.1489 | 157.6652 | 132.8141 | 145.907 | 123.1209 | 85.2094 | 87 | 82.5117 | 94.9675 | 185.3023 | 182.5722 | 210.6092 | 358.6138 | 200.8466 | 208.6977 | 208.8466 | 284.3953 | 151.0325 | 76.3348 | 73.5488 | 151.0325 | 103.4837 | 104.3673 | 109.5442 | 134.5163 | 103.3348 | 117.8836 | 92.5767 | 110.9395 | 77.5163 | 71.3348 | 74.093 | 103.1534 | 113.9116 | 73.1489 | 92.935 | 72.1209 | 217.2419 | 179.6327 | 222.1534 | 477.3348 | 134.7627 | 162.9675 | 163.6047 | 436.9116 | 263.5767 | 175.1534 | 193.093 | 379
1787 | 21.6514 | 134.5781 | 165.2891 | 216.2656 | 183.1692 | 131.0482 | 126.6263 | 174.2773 | 150.3008 | 130.0846 | 147.6992 | 105.4935 | 155 | 184.1563 | 122.8672 | 149.9636 | 207.5781 | 126.3972 | 141.47 | 135.2409 | 152.168 | 130.1563 | 304.0612 | 294.819 | 140.2044 | 184.1563 | 194.3737 | 218.181 | 290.8672 | 107.4336 | 99.6263 | 113.3255 | 118.9036 | 112.1563 | 123.3372 | 118.53 | 86.5781 | 200.9636 | 201.0846 | 228.7826 | 261.4466 | 127.181 | 92.7344 | 119.9036 | 114.9518 | 71.6263 | 54.5899 | 75.5182 | 108.0846 | 193.9765 | 176.2409 | 222.4101 | 149.1328 | 95.9154 | 124.7591 | 116.7591 | 111.9518 | 145.2891 | 139.8073 | 184.8555 | 168.47 | 116.9636 | 100.5781 | 139.0599 | 144.8437 | 166.6992 | 164.6263 | 179.3854 | 238.1927 | 124.5182 | 146.9883 | 204.6992 | 194.1198 | 189.4219 | 176.7708 | 189.0364 | 342.9154 | 134.1563 | 158.7109 | 143.7591 | 225.0964 | 112.6263 | 93.7227 | 125.2409 | 161.6028 | 155.0964 | 150.0364 | 171.651 | 129.6628 | 89.7109 | 95.5664 | 96.2292 | 113.9883 | 182.1445 | 197.181 | 185.349 | 328.6862 | 215.2292 | 190.9154 | 175.638 | 313.5547 | 150.0964 | 53.2044 | 60.7227 | 150.0964 | 103.9518 | 98.3854 | 114.7109 | 105.4935 | 105.9036 | 92.3972 | 94.9036 | 115.7591 | 58.9636 | 75.8073 | 70.2409 | 89.7227 | 102.6263 | 100.5547 | 105.2773 | 62.4818 | 207.4818 | 182.7591 | 197.3008 | 475.1445 | 137.9518 | 171.5182 | 195.362 | 434.1927 | 256.3854 | 180.7591 | 153.0716 | 366.6263
1788 | 21.6624 | 161.9593 | 183.3182 | 212.2153 | 167.646 | 125.2153 | 141.4258 | 162.4665 | 161.5335 | 141.6052 | 114.2919 | 121.354 | 147.2871 | 203.2464 | 154.7799 | 154.7847 | 192.5383 | 178.9903 | 130.4306 | 130.1794 | 164.4617 | 135.7488 | 282.8972 | 324.8516 | 145.7488 | 181.0718 | 193.0718 | 202.6101 | 288.0718 | 110.8565 | 97.0718 | 114.9641 | 131.5335 | 120.6411 | 146.1387 | 116.0718 | 93.7129 | 200.9641 | 201.0359 | 242.4258 | 296.7799 | 133.7847 | 109.3182 | 119.0359 | 113.0718 | 70.0359 | 69.3899 | 87.5694 | 122.4976 | 183.4665 | 171.1794 | 197.0048 | 128.7895 | 117.1746 | 112.4665 | 120.8565 | 124.5335 | 149.8206 | 140 | 187.8923 | 169 | 113.1077 | 127.9593 | 147.6411 | 125.7536 | 154.5024 | 156.2871 | 194.4258 | 242.8206 | 118.2512 | 133.5383 | 181.8972 | 184.2871 | 170.7177 | 199.1387 | 198.677 | 360.3182 | 149.3899 | 119.4714 | 142.0718 | 200.8972 | 127.4258 | 126.744 | 131.7488 | 127.3278 | 166.5694 | 131.7177 | 143.1125 | 125.1794 | 85.1794 | 105.6411 | 113.3899 | 117.8923 | 213.8157 | 202.8206 | 199.4258 | 337.6052 | 202.5024 | 190 | 200.9952 | 287.0407 | 163.4976 | 71.2823 | 60 | 163.4976 | 93.3948 | 122.1028 | 96.6818 | 113.6411 | 97.323 | 101.6052 | 105.6052 | 115.0359 | 44.5024 | 68.2871 | 66.1435 | 74.5383 | 104.8923 | 97.1794 | 92.5024 | 71.6411 | 204.1077 | 199.3899 | 220.1028 | 489.4617 | 135.1077 | 168.1435 | 168.0766 | 372.2967 | 229.9689 | 161.7177 | 171.2464 | 367.9282
1789 | 21.6734 | 147.3775 | 160.5663 | 185.6371 | 167.9764 | 139.6461 | 135.1652 | 157.118 | 165.9056 | 147.8584 | 97.3775 | 104.4247 | 130.4011 | 152.2506 | 131.5899 | 158.9056 | 205.6697 | 171.236 | 115.3539 | 146.5989 | 172.8112 | 115.4955 | 320.0798 | 290.8494 | 136.236 | 201.5045 | 192.0236 | 181.4955 | 316.3157 | 123.6933 | 101.882 | 103.2831 | 117.3539 | 124.9056 | 116.7315 | 125.764 | 91.0708 | 201 | 199.0472 | 221.5191 | 248.2034 | 129.118 | 134.4101 | 119.9764 | 108.118 | 80.7404 | 51.4483 | 62.6135 | 112.2596 | 174.2124 | 168.0708 | 196 | 155.3393 | 119.9528 | 108.0944 | 106.3539 | 109.3775 | 150 | 132.1888 | 176.2831 | 169.9764 | 127.6461 | 135.8348 | 152.882 | 114.2596 | 154.9764 | 138.4247 | 214.5281 | 221.5191 | 141.4337 | 131.0472 | 204.4337 | 213.2921 | 161.2124 | 162.8966 | 195.0944 | 367.8348 | 139.2596 | 144.3629 | 156.6461 | 229.2921 | 147.5281 | 131.9056 | 132.9764 | 152.3629 | 171.882 | 127.0944 | 169.3393 | 117.1888 | 99.6461 | 97.2124 | 105.2124 | 106.2831 | 208.1652 | 204.9528 | 186.3303 | 311.6371 | 191.2596 | 185.118 | 180.5191 | 287.9528 | 147.4011 | 81.764 | 60 | 147.4011 | 104.7169 | 91.7551 | 101.8584 | 109.118 | 105.7876 | 109.8112 | 104.0472 | 102.3067 | 73.2921 | 61.1652 | 67.9528 | 82.7876 | 102.0708 | 86.2596 | 90.0472 | 82.7404 | 235.2449 | 188.2831 | 206.3539 | 500.7404 | 125.236 | 175.8112 | 167 | 424.6786 | 212.4011 | 174.6697 | 217.891 | 399.2449
1790 | 21.6844 | 141.0678 | 180.7626 | 181.0452 | 198.6496 | 108.3617 | 147.853 | 138.2148 | 164.0226 | 142.0678 | 106.887 | 123.7739 | 175.48 | 172.7513 | 150.7739 | 137.2487 | 198.0904 | 187.8078 | 133.7852 | 120.3052 | 198.7061 | 146.6383 | 273.5426 | 306.8078 | 155.7739 | 185.1922 | 192 | 190.887 | 299.2035 | 108.1809 | 102 | 103.9887 | 113.0452 | 137.853 | 128.853 | 130.9435 | 82.1017 | 168.373 | 192.0791 | 218.0339 | 237.113 | 133.9435 | 105.3391 | 119.0113 | 103.0565 | 78.0339 | 58.9096 | 60.0226 | 90.2487 | 174 | 165.0339 | 193.0339 | 132.2713 | 138.7852 | 105.0339 | 113.9096 | 119.8757 | 142.0904 | 161.6609 | 209.6156 | 165.0565 | 146.7852 | 109.3052 | 147.0678 | 136.74 | 159.9435 | 160.74 | 194.2374 | 225.9435 | 118.2713 | 121.113 | 224.7739 | 225.8643 | 176.8191 | 166.9435 | 187.0904 | 360.0904 | 117.2487 | 130.1696 | 147.113 | 227.0339 | 132.1809 | 119.147 | 118.1696 | 168.8191 | 159.147 | 131.9435 | 139.3504 | 119.9661 | 82.2035 | 90.0791 | 96.1017 | 94.1357 | 221.8417 | 176.3278 | 189.9548 | 312.9774 | 192.9774 | 201.8078 | 239.3217 | 282.0678 | 141.0678 | 62.2261 | 69.887 | 141.0678 | 99.0678 | 119.6722 | 96.0678 | 114.9322 | 111.9322 | 86.2713 | 94.113 | 118.8078 | 51.26 | 70.887 | 66.0226 | 83 | 95.0791 | 99.8417 | 106.8078 | 78.0565 | 227.1017 | 203.8191 | 209.9548 | 502.9774 | 125.9887 | 167.1017 | 188.7513 | 410.1809 | 230.7852 | 181.9209 | 187.3617 | 400.9887
1791 | 21.6954 | 138.0247 | 187.9921 | 164.0345 | 187.9921 | 138.9882 | 158.9704 | 154.9822 | 153.0138 | 142.9852 | 121.9891 | 124.999 | 168.9793 | 184.0256 | 154.9901 | 145.0099 | 194.0286 | 156.9961 | 127.0059 | 127.0128 | 165.9961 | 123.002 | 299.9763 | 277.0217 | 155.0158 | 185.0069 | 185.9832 | 231.9586 | 310.9901 | 113.9832 | 115 | 107.0178 | 99.0523 | 131.9901 | 124.997 | 135.9783 | 92.0158 | 238.9457 | 188.9872 | 200.0148 | 262.0079 | 108.0138 | 118.0069 | 97.0138 | 124.9793 | 76.9842 | 69.9921 | 65.002 | 104.0118 | 150.0316 | 163.0118 | 196.9941 | 173.9645 | 100.0099 | 133.9822 | 132.997 | 123.9891 | 143.9744 | 138.9694 | 186.9921 | 149.0128 | 103.0187 | 118.0039 | 135.9961 | 130.9813 | 148.0256 | 171.9882 | 200.9891 | 240.9714 | 117.0276 | 127.0099 | 184.0316 | 191.003 | 185.9931 | 184.0148 | 191.9911 | 356.0602 | 139.9931 | 140.9941 | 152.9605 | 226.9813 | 104.0326 | 119.0099 | 120.997 | 141.9931 | 175.9882 | 150.0059 | 173.9822 | 103.0355 | 90.0158 | 101.0059 | 96.997 | 86.0286 | 200.9773 | 172.0158 | 201.9931 | 353.9201 | 223.9832 | 190.0069 | 209.9862 | 316.9931 | 165.0049 | 61.002 | 77.9862 | 165.0049 | 111.001 | 99.0148 | 92.0079 | 123.9931 | 85.0138 | 103.0178 | 103.9852 | 109.9813 | 56.0089 | 59.0306 | 75.9862 | 95.9882 | 101.998 | 106.9852 | 99.003 | 81.0089 | 217.9813 | 194.9882 | 186.001 | 463.0306 | 124.997 | 143.0256 | 176.0118 | 415.0089 | 226.9813 | 184.9783 | 154.0335 | 383
1792 | 21.7064 | 162.4689 | 179.9602 | 198.7876 | 180.0531 | 127.0266 | 129.1461 | 136.8539 | 166.9469 | 127.9867 | 111.0531 | 123.8672 | 147.9469 | 209.6282 | 144.9734 | 154.8274 | 223 | 153.1593 | 132.8805 | 139.8539 | 162.0664 | 125.385 | 275.9071 | 299.3187 | 170.5486 | 191.8672 | 169.3718 | 190.2257 | 301.1726 | 97.0398 | 114.8805 | 124.7212 | 151.4955 | 122.239 | 122.0133 | 114.0797 | 107.8938 | 184.3452 | 176.5709 | 215.0133 | 270.0133 | 121.8141 | 124.7743 | 111.1062 | 104.1726 | 61.0266 | 61.9867 | 67.0133 | 115.8274 | 181.8407 | 175.0266 | 191.2124 | 138.1062 | 110.0266 | 116.1593 | 129.8938 | 113.0929 | 118.478 | 108.4249 | 179.0398 | 161.615 | 121.9602 | 121.6548 | 132.2124 | 112.3718 | 173.615 | 160.1992 | 190.2523 | 212.2788 | 144.8407 | 136.9867 | 215.6946 | 193.9734 | 178.9734 | 198.5884 | 182.9734 | 416.3494 | 133.0531 | 134.8008 | 113.5577 | 207.8141 | 136.6681 | 128.8141 | 118.3054 | 134.8407 | 164.1461 | 155.9867 | 156.0797 | 138.5618 | 105.7079 | 106.6548 | 94.1062 | 114.9602 | 178.5975 | 187.6415 | 195.2257 | 273.8896 | 206.9336 | 196.8938 | 195.8938 | 309.4291 | 170.1461 | 62.9469 | 64.2788 | 170.1461 | 111.7876 | 113.9867 | 100.0664 | 117.0929 | 98.9469 | 120.5486 | 89.3585 | 91.4514 | 65.1062 | 89.7477 | 62.0797 | 84.2788 | 99.9734 | 91.9203 | 101.9469 | 89.8274 | 199.4382 | 183.2257 | 187.3452 | 494.0133 | 122.0398 | 168.6415 | 187.7344 | 423.6946 | 208.4382 | 163.1195 | 187.9071 | 383.1195
1793 | 21.7174 | 123.179 | 177.4347 | 184.2529 | 184 | 128.7187 | 139.8977 | 126.1534 | 162.5653 | 126.8722 | 114.821 | 114.1023 | 144.2813 | 182.4347 | 142.3864 | 142.2301 | 222.5398 | 164.8977 | 124.0767 | 129.1278 | 167.3835 | 153.7187 | 269.4858 | 322.2585 | 137.6136 | 182.1023 | 196.5653 | 206.8722 | 314.2301 | 100.358 | 105.821 | 104.2813 | 114.7671 | 139.4886 | 122.821 | 120.0767 | 99.6932 | 209.4375 | 218.2585 | 215.821 | 270.4119 | 108.5625 | 108.0256 | 119.358 | 117.5114 | 63.2301 | 60.9489 | 67.821 | 102.821 | 169.7699 | 177.2301 | 207.4858 | 146.2045 | 112.5369 | 128.2557 | 122.179 | 119.642 | 153.7699 | 139.8977 | 181.3864 | 134.1506 | 118.9489 | 96.4858 | 147.6932 | 139.3608 | 144.9744 | 174.8466 | 208.0795 | 233.6648 | 132.7443 | 136.1023 | 192.7443 | 191.8466 | 177.1023 | 167.4886 | 180.7955 | 368.6648 | 136.821 | 120.1534 | 154.5142 | 194.9716 | 112.358 | 115.1534 | 140.3608 | 123.4858 | 174.6932 | 155.1534 | 161.8466 | 106.7159 | 83.9489 | 81.7415 | 102.2813 | 111.6676 | 222.2841 | 162.1506 | 211.2074 | 339.821 | 201.2841 | 189.1278 | 188.4858 | 268.5341 | 180.4119 | 58.9233 | 84.7187 | 180.4119 | 95.8977 | 112.9489 | 105.2557 | 123.7955 | 95.4347 | 87.5114 | 115.4886 | 124.4631 | 72.9233 | 71.0511 | 67.8977 | 104.8722 | 98.1534 | 86.1534 | 98.1534 | 76.642 | 231.4631 | 199.4631 | 212.5909 | 494.6676 | 124.8722 | 142.6392 | 167.6676 | 400.6165 | 241.6392 | 172.3068 | 180.9489 | 392.1534
1794 | 21.7284 | 130.1893 | 193.5078 | 230.5613 | 183.8486 | 117.2807 | 136.53 | 132.7572 | 146.6815 | 122.6058 | 108.4165 | 117.8486 | 155.53 | 199.1514 | 120.4387 | 150.6214 | 205.4165 | 160.6971 | 127.265 | 134.3029 | 181.2807 | 142.3564 | 288.265 | 294.1136 | 160.1671 | 186.3407 | 180.4922 | 202.1514 | 321.5992 | 113.3185 | 99.7951 | 114.6214 | 143.5078 | 119.9243 | 116.4543 | 122.9621 | 87.5457 | 189.4387 | 190.1136 | 209.3407 | 249.7415 | 129.8107 | 109.3407 | 132.9621 | 135.9399 | 71.3564 | 59.6815 | 61.9844 | 96.6058 | 161.5679 | 185.4321 | 224.9021 | 153.7728 | 133.3407 | 136.7506 | 128.1292 | 106.9086 | 144.9243 | 137.0222 | 158.53 | 176.9399 | 117.3786 | 114.2049 | 136.0757 | 114.735 | 143.9243 | 168.3564 | 174.2872 | 257.7885 | 123.8329 | 140.5679 | 183.4543 | 186.6436 | 180.6593 | 148.1893 | 174.8173 | 392.4856 | 130.1893 | 126.6815 | 136.0379 | 231.3942 | 125.7728 | 120.3942 | 116.8708 | 142.5679 | 162.9243 | 160.4321 | 156.7572 | 132.9778 | 81.5078 | 109.2049 | 113.2272 | 99.4165 | 194.8864 | 206.7951 | 182.0222 | 332.3185 | 175.7415 | 193.6214 | 208.0979 | 325.4478 | 158.53 | 55.9621 | 73.735 | 158.53 | 92.8329 | 110.6214 | 114.6593 | 116.1514 | 111.7728 | 106.9243 | 95.8107 | 104.7951 | 69.9621 | 73 | 64.2272 | 99.5835 | 103.735 | 91.735 | 104.4165 | 63.6058 | 210.735 | 179.2272 | 197.7193 | 480.0692 | 119.47 | 166.3942 | 155.9844 | 387.628 | 265.1671 | 183.3564 | 178.1671 | 397.0535
1795 | 21.7394 | 133.997 | 180.0973 | 193.4985 | 180 | 100.655 | 149.8495 | 151.5486 | 162.1946 | 137.9498 | 120.7051 | 114 | 169.1505 | 202.1976 | 154.6931 | 140.6489 | 215.6489 | 154.1535 | 133.1976 | 141.2979 | 162.8495 | 127.0532 | 294.5486 | 295.5958 | 140.5547 | 193.4955 | 193.7021 | 204.2949 | 287.655 | 97.1535 | 120 | 104.8495 | 131.3511 | 118.3009 | 127.6489 | 121.8495 | 76.5015 | 225.4483 | 194.3039 | 219.1535 | 292.2447 | 125.4012 | 117.348 | 131.1976 | 109.1505 | 54.9498 | 76.4483 | 86.4483 | 111.2979 | 175.1976 | 171.7021 | 197.5015 | 148.2006 | 141.5988 | 105.7523 | 106.8024 | 129.348 | 141.7462 | 161.4955 | 171.4483 | 149.5988 | 126.348 | 95.2036 | 138.8526 | 108.7523 | 142.3511 | 151.7492 | 206.5988 | 227.1003 | 143.5958 | 154.2477 | 194.0973 | 201.2949 | 172.5517 | 153.1505 | 218.6931 | 355.4544 | 135.4012 | 143.1474 | 137.6018 | 218.0562 | 120.0502 | 105.4012 | 137.4453 | 156.1976 | 160.0973 | 146.8024 | 174.6961 | 107.8024 | 70.3541 | 89.2006 | 117.4955 | 109.9498 | 191.3982 | 225.7462 | 208.2006 | 316.3039 | 217.8937 | 184.0502 | 234.8967 | 288.8084 | 171.8495 | 55.9529 | 66.3982 | 171.8495 | 114.7523 | 100.6489 | 106.7021 | 119.5486 | 105.2979 | 104.2979 | 91.8024 | 123.4955 | 68.7994 | 72.8997 | 69.7994 | 89 | 97.1505 | 84.8495 | 114.2477 | 78.6991 | 203.7492 | 186.0532 | 215.9498 | 432.9559 | 106.7523 | 151.652 | 179.2447 | 429.1505 | 244.9529 | 167.4012 | 157.9027 | 372.0471
1796 | 21.7504 | 116.9357 | 162.6878 | 185.6235 | 180.6869 | 132.0624 | 146.1258 | 144.124 | 127.7502 | 136.0634 | 151.5014 | 114.6869 | 171.3131 | 186.5005 | 113.5601 | 134.9366 | 209.8117 | 174.439 | 120.6226 | 129.3113 | 161.561 | 147 | 285.3756 | 272.0596 | 168.5014 | 167.9348 | 206.0009 | 174.2479 | 317.8136 | 118.2507 | 117.8145 | 102.3747 | 137.8751 | 122.439 | 121.9991 | 120.0615 | 86.0624 | 214.8127 | 218.0009 | 239.1267 | 257.3765 | 133.3122 | 105.6869 | 116.4995 | 111.7502 | 55.2488 | 66.7493 | 76.2498 | 98.562 | 161.9981 | 185 | 205.439 | 153.124 | 132.3141 | 119.438 | 122.2498 | 115.876 | 119.6235 | 134.3113 | 161.9366 | 143.2488 | 114 | 117.9376 | 155.4995 | 123.6869 | 150.0615 | 148.1864 | 197.8136 | 230.1864 | 117.4995 | 140.4371 | 177.3747 | 168.8751 | 182.8751 | 156 | 176.2488 | 379.1911 | 141.0643 | 127.1873 | 149.4371 | 254.3774 | 120.1258 | 113 | 109.3737 | 140.3756 | 143.8117 | 160.8136 | 150.4371 | 122.5629 | 95.8751 | 93.6244 | 89.1249 | 107.6887 | 181.7484 | 200.6887 | 211.3756 | 338.5648 | 179.3113 | 184.5629 | 212.9385 | 341.3756 | 167.8136 | 73.3131 | 56.1864 | 167.8136 | 126.2526 | 95.0615 | 119.3131 | 110.6253 | 93.0615 | 91.6869 | 106.6878 | 96.3113 | 65.1249 | 70.2507 | 65.5005 | 88.9376 | 100.6244 | 82.4371 | 99.5005 | 73.8742 | 199.562 | 203.7521 | 214.8751 | 469.438 | 121.6869 | 165.4986 | 148.3103 | 429.315 | 260.2526 | 174.9376 | 175.9366 | 355.124
1797 | 21.7614 | 145.2527 | 159.7936 | 208.1317 | 191.0747 | 132.3274 | 132.5516 | 161.9715 | 126.242 | 122.6726 | 127.3559 | 123.879 | 160.4021 | 179.1495 | 149.0854 | 149.5231 | 219.9822 | 150.7758 | 157.6085 | 147.8043 | 183.9537 | 146.3274 | 275.7758 | 315.5338 | 147.7189 | 210.5053 | 190.1779 | 207.9253 | 298.4306 | 108.7936 | 86.7189 | 107.4769 | 135.5516 | 102.4377 | 135.8043 | 135.2527 | 87.5979 | 211.6263 | 203.0747 | 209.4306 | 237.484 | 136.2064 | 117.121 | 123.7758 | 107.1779 | 75.4947 | 77.7011 | 77.6833 | 107.8968 | 191.4021 | 184.9253 | 183.4199 | 166.4128 | 108.2705 | 111.3737 | 124.9537 | 100.4947 | 144 | 153.7011 | 174.5053 | 161.4021 | 114.5979 | 116.7758 | 162.0285 | 135.121 | 164.5801 | 165.7758 | 180.3737 | 249.8683 | 127.6156 | 145.2064 | 184.5694 | 168.8683 | 180.4021 | 156.9715 | 195.0747 | 324.6904 | 114.0925 | 129.8505 | 153.3096 | 214.9253 | 107.8968 | 113.7473 | 129.4769 | 130.8505 | 156.5231 | 145.1673 | 158.3452 | 115.1032 | 93.6263 | 102.7011 | 92.7189 | 90.5409 | 207.4769 | 181.8221 | 202.6726 | 303.9431 | 199.5231 | 178.0747 | 196.6263 | 328.3381 | 150.6726 | 62.1032 | 73.7011 | 150.6726 | 87.8398 | 109.5801 | 108.6263 | 104.6263 | 107.879 | 101.8505 | 101.4769 | 114.7295 | 67.3737 | 60.4947 | 59.1957 | 88.0747 | 109.2527 | 88.8505 | 93.121 | 86.4021 | 207.1779 | 171.6441 | 213.4484 | 461.8968 | 132.5231 | 187.1032 | 183 | 391.9146 | 219.7473 | 174 | 189.2527 | 371.6263
1798 | 21.7724 | 136.6092 | 182.174 | 184.826 | 190.8687 | 124 | 128.2183 | 173.1298 | 152.6076 | 130.4779 | 107.174 | 111.0443 | 152.913 | 180.913 | 115.8718 | 155.1298 | 198.742 | 148.6092 | 129.0443 | 134.5664 | 172.6535 | 137.6519 | 270.3893 | 285.174 | 166.6504 | 192.6092 | 179.4779 | 207.2611 | 280.3054 | 130.2168 | 106.8687 | 101 | 130.5221 | 142.1282 | 119.8687 | 126.087 | 95.174 | 204.8244 | 204.2611 | 193.9161 | 289.0839 | 114.9573 | 130.3038 | 121.4351 | 98.0443 | 93.0427 | 73.826 | 49.7832 | 117.3465 | 182.7817 | 183.6519 | 200.4779 | 124.7848 | 125.8275 | 116.4351 | 112.7832 | 118.5649 | 144.6962 | 149.0427 | 160.0031 | 154.3481 | 122.2611 | 113.6519 | 148.5206 | 144.9969 | 147 | 164.6535 | 184.913 | 269.345 | 158.2595 | 123.5221 | 202.6946 | 190.5206 | 173.3481 | 169.6962 | 192.6061 | 358.3481 | 141.2183 | 129.3054 | 121.6535 | 216.7848 | 118.3038 | 122.826 | 123.6092 | 129.5221 | 161.5206 | 170.7817 | 172.6504 | 107.568 | 90.3924 | 97.9557 | 112.6076 | 121.7389 | 199.6076 | 191.3038 | 211.3481 | 329.3924 | 205.3038 | 179.9573 | 192.2611 | 286.1329 | 160.3924 | 52.3924 | 69.3908 | 160.3924 | 119.7801 | 93.3054 | 104.9573 | 101.8275 | 92.7817 | 100.6092 | 97.6977 | 99.5221 | 72.174 | 77.6076 | 74.5221 | 89.174 | 98.6076 | 86.826 | 104.4763 | 80.3054 | 197.2611 | 191.3908 | 217.6946 | 472.7389 | 137.5206 | 174.6076 | 182.5649 | 428.6092 | 229.2611 | 174 | 180 | 367.6092
1799 | 21.7834 | 141.7089 | 180.8219 | 185.2842 | 178.9007 | 123.6027 | 139.7089 | 161.5171 | 136.411 | 125 | 109.0993 | 121.3048 | 150.5103 | 179.5034 | 145.2055 | 145.1062 | 234.226 | 153.411 | 140.6952 | 150.5034 | 187.4178 | 132.3116 | 244.7877 | 285.6096 | 143.4897 | 198.0069 | 176.4829 | 209.2055 | 293.8014 | 118.0274 | 95 | 101.1986 | 134.0137 | 112.9931 | 107.2055 | 127.6952 | 95.2123 | 187.0651 | 204.8151 | 235.1986 | 247.0069 | 123.9075 | 121.6096 | 126.7945 | 110.4897 | 81.7089 | 70.9075 | 60.1849 | 99.5034 | 171.9863 | 181.1918 | 196.0925 | 155.3904 | 142.5171 | 120.7021 | 119.5103 | 113.6027 | 151.0069 | 140.7877 | 198.6233 | 157.6027 | 126.6884 | 108.9075 | 134.6884 | 106.0788 | 147.9931 | 180.6096 | 186.7808 | 230.7021 | 140.6952 | 129.1986 | 189.8938 | 177.4829 | 176.9007 | 177.2979 | 160.1781 | 361.2055 | 151.4178 | 141.7089 | 136.1199 | 239.8425 | 111.2979 | 120.2055 | 131.4897 | 136.6884 | 147.0925 | 158.4966 | 148.9931 | 141.3253 | 103.113 | 88.3904 | 97.8014 | 118.8014 | 186.887 | 187.1781 | 212.7158 | 340.9212 | 194.7226 | 193.0788 | 196.7877 | 314.8288 | 171.9212 | 67.9931 | 63.4966 | 171.9212 | 86.7945 | 106.8014 | 114.0069 | 118.7158 | 83.0788 | 104.9144 | 123.0206 | 105.7945 | 73.3048 | 64.0925 | 77.6164 | 91.1986 | 85.7877 | 84.8014 | 81.0788 | 91.3185 | 202.5822 | 186.2911 | 203.8014 | 471.2911 | 121.9007 | 160.6952 | 177.3048 | 431.0274 | 233.6884 | 174.0993 | 176.6233 | 373.1062
1800 | 21.7944 | 128.2143 | 156.3526 | 204.2143 | 175.6563 | 121.4509 | 130.1205 | 139.7812 | 123.8973 | 124.1072 | 108.8839 | 114.5536 | 137.8928 | 173.4375 | 136.7723 | 138.2277 | 198.6607 | 141.3437 | 145.7723 | 144.9955 | 168.3526 | 116.8884 | 263.3437 | 272.3304 | 156.1072 | 188.1072 | 196.9911 | 202.2232 | 292.7812 | 83.7857 | 95.3348 | 103.8928 | 117.6741 | 120.3259 | 103.2366 | 133.2188 | 80.7857 | 228.759 | 188.3482 | 236.2188 | 243.8035 | 116.9018 | 108.8884 | 132.2143 | 123.442 | 69.2188 | 60.8884 | 78.6607 | 97.2321 | 190 | 187.9822 | 204.2143 | 166.7723 | 118.7723 | 115.5447 | 103.8795 | 109.6652 | 142.7812 | 155.2143 | 171.3482 | 153.1072 | 138.9866 | 100.1161 | 147.9911 | 134.8928 | 155.4375 | 164.875 | 210.8839 | 227.3304 | 146.3304 | 133.4553 | 197.8884 | 198.7723 | 174.6607 | 179.7768 | 189 | 354.2232 | 127.7768 | 130.558 | 114.0178 | 188.0178 | 116.3437 | 116.125 | 142.2098 | 149.7679 | 154.7679 | 160.7679 | 159.1161 | 107.7768 | 85.7768 | 102.3393 | 97.0045 | 117.1116 | 205.2277 | 209.9733 | 193.7857 | 311.7723 | 165.8928 | 217.875 | 211.2143 | 279.7902 | 144.7812 | 74.9911 | 68 | 144.7812 | 92.6607 | 104.5536 | 104.8884 | 100.5669 | 107.7634 | 87.6741 | 96.6696 | 112.2188 | 67.558 | 75.2277 | 57.8973 | 93.558 | 103.8973 | 83.4464 | 105.6518 | 67.1116 | 224.3259 | 194.6518 | 203.6741 | 484.7857 | 124.3482 | 166.3304 | 171.558 | 397.9018 | 252.4598 | 174.2188 | 151.8035 | 365.1116
1801 | 21.8054 | 115.3629 | 191.4956 | 187.7699 | 161.0886 | 132.6283 | 148.239 | 143.2743 | 140.8585 | 119.2301 | 102.1063 | 112.1151 | 150.3276 | 166.0798 | 126.2566 | 148.9912 | 189.6107 | 157.8937 | 133.8937 | 136.5044 | 198.7875 | 116.3717 | 282.9912 | 270.3452 | 146.8937 | 183.4779 | 181.1239 | 203.7522 | 299.9912 | 99.1151 | 100.1063 | 111.2478 | 130.2566 | 108.9824 | 129.2478 | 129.2301 | 94.0088 | 197.8496 | 209.0529 | 228.2654 | 281.575 | 135.4158 | 109.1151 | 117.5044 | 118.0088 | 64.239 | 61.1151 | 66.5132 | 115.8673 | 189.0088 | 158.1063 | 189.2566 | 154.3982 | 111.2301 | 96.9912 | 91.0886 | 107.3717 | 150.239 | 139.6371 | 193.4158 | 147.4868 | 118.9647 | 110.7346 | 133.8585 | 140.3893 | 143.1239 | 143.3452 | 199.1504 | 218.407 | 139.2654 | 148.5397 | 199.1063 | 190.3629 | 165.3629 | 176.6371 | 188.0088 | 357.6107 | 127.7346 | 134.3805 | 143.7699 | 219.2566 | 132.3982 | 139.8849 | 118.0176 | 132.8673 | 140.0974 | 144.4868 | 169.1151 | 108.6018 | 84.9912 | 110.8937 | 105.2478 | 117.6283 | 215.7434 | 166.0886 | 209.4868 | 309.3099 | 175.1063 | 193.8673 | 198.9824 | 302 | 151.6195 | 59.8673 | 67.5044 | 151.6195 | 84.6018 | 102.239 | 103.2566 | 121 | 83.239 | 98.0265 | 102.1239 | 109.9647 | 72.1239 | 85.7434 | 70.8937 | 97.3805 | 117.5132 | 88.4868 | 81.4779 | 69.6107 | 205.1768 | 171.2213 | 214.646 | 496.646 | 148.5221 | 163.1063 | 179.3452 | 425.2125 | 277.3982 | 167.5044 | 196.761 | 369.2213
1802 | 21.8164 | 127.0428 | 193.5019 | 175.2685 | 188.5019 | 130.5448 | 154.2762 | 129.1751 | 150.7315 | 135.5448 | 116.3191 | 118.3657 | 182.4163 | 197.0972 | 121.6809 | 156.9533 | 197.8677 | 146.5409 | 119.1362 | 132.7276 | 172.5409 | 119.4086 | 290.4086 | 296.7238 | 136.358 | 198.0039 | 185.1323 | 206.358 | 309.4514 | 105.2296 | 114.1829 | 108.642 | 123.6381 | 122.7276 | 129.5019 | 144.7276 | 85.9105 | 218.5448 | 170.9495 | 215.5914 | 225.6771 | 103.2724 | 117.6809 | 114.6809 | 111.6809 | 72.5914 | 69.6809 | 59.6771 | 118.4591 | 180.0934 | 173.8171 | 185.2257 | 139.1323 | 125.7743 | 104.1362 | 117.821 | 111.9066 | 159.6809 | 132.179 | 164.5409 | 156.2296 | 139.782 | 120.4124 | 147 | 128.7276 | 146.7238 | 166.5914 | 180.0895 | 202.9417 | 130.4047 | 120.5409 | 211.821 | 196.3229 | 177.0428 | 169.86 | 182.4981 | 363.2801 | 137.2762 | 131.6343 | 128.6809 | 216.0428 | 114.8171 | 130.9105 | 108.2218 | 140.6343 | 161.8638 | 154.7276 | 175.7743 | 126.8638 | 92.6809 | 95.8638 | 107 | 114.4552 | 217.8677 | 196.9066 | 220.2724 | 354.463 | 194.6304 | 200.2724 | 211.5019 | 298.4591 | 156.5448 | 66.9533 | 65.4981 | 156.5448 | 103.1362 | 111.8171 | 100.9961 | 119.0934 | 92.8171 | 81.9028 | 104.4981 | 135.0039 | 72.5914 | 87.7315 | 58.0428 | 93.5448 | 107.2724 | 99 | 98.3191 | 80.0467 | 165.9066 | 198.358 | 199.6343 | 480.2724 | 130.3191 | 177.0467 | 201.3657 | 447.3735 | 255.3229 | 167.9495 | 185.14 | 389.5486
1803 | 21.8274 | 142.6726 | 181.1789 | 197.3363 | 180.7818 | 137.415 | 138.3363 | 158.7335 | 131.9302 | 137.3667 | 111.4061 | 106.6637 | 146.2576 | 172.83 | 126.8909 | 158.9911 | 181.415 | 167.3667 | 120.1485 | 131.4454 | 195.5939 | 122.7424 | 292.703 | 310.8819 | 162.6637 | 178.703 | 202.3274 | 224.4973 | 329.0089 | 96.0787 | 105.8819 | 87.5331 | 116.4848 | 122.4848 | 120.4454 | 140.0304 | 81.0787 | 219.1789 | 198.6726 | 213.1485 | 251.8211 | 106.3363 | 120.9606 | 117.8122 | 116.5939 | 69.4061 | 73.1091 | 78.9911 | 94.9606 | 170.2272 | 180.6333 | 194.0394 | 153.8032 | 117.4061 | 105.4454 | 107.5635 | 121.9213 | 164.8909 | 146.297 | 184.4759 | 144.8515 | 97.5635 | 108.7513 | 145.8122 | 129.2272 | 164.1485 | 164 | 189.2272 | 260.3578 | 151.5546 | 142.8515 | 198.8909 | 178.1968 | 189.703 | 184.8819 | 189.4759 | 329.9696 | 123.415 | 139.3274 | 137.3059 | 227.0698 | 121.7818 | 122.6637 | 133.9606 | 148.3274 | 158.9213 | 152.2576 | 165.0304 | 125.4061 | 96.703 | 96.0394 | 105.5152 | 109.6637 | 197.5546 | 208.5546 | 222.297 | 317.7513 | 220.8819 | 199.1789 | 203.6333 | 282.2361 | 159.4061 | 71.3667 | 74.8515 | 159.4061 | 103.5546 | 114.9213 | 118.9606 | 109.6726 | 100.5241 | 111.8122 | 111.4759 | 113.7728 | 67.9213 | 69.0787 | 72.4848 | 95.0698 | 106.1789 | 96.4759 | 95.1878 | 74.297 | 183.3453 | 221.6941 | 208.6637 | 479.7728 | 125.2576 | 172.3363 | 189.6637 | 400.1878 | 238.6816 | 189.1395 | 168.6333 | 374.297
1804 | 21.8384 | 152.0515 | 165.1608 | 200.4984 | 186.6592 | 152.0161 | 146.3215 | 144.0193 | 145.4116 | 128 | 105.2669 | 101.09 | 146.5016 | 208.3023 | 136.18 | 135.6785 | 203.8939 | 159.4469 | 120.0354 | 133.5177 | 197.7138 | 127 | 295.5016 | 287.2123 | 153.3923 | 174.4277 | 189.7331 | 175.7877 | 347.4984 | 107.1961 | 92.0193 | 124.0354 | 120.3377 | 131.4823 | 121.7138 | 126.6977 | 89.4631 | 202.0354 | 211.4277 | 212.3923 | 264.6238 | 113.5177 | 114.3569 | 113.2508 | 120.3215 | 64.8746 | 65.5884 | 59.2154 | 88.6785 | 183.9646 | 185.4984 | 200.3215 | 128.9839 | 112.8746 | 109.4469 | 125.91 | 111.2862 | 164.6946 | 151.2154 | 171.2862 | 147.3762 | 117.3569 | 133.4277 | 139.4823 | 141.5177 | 162.91 | 162.8746 | 196.373 | 231.5369 | 148.6785 | 140.5531 | 207.3762 | 205.1061 | 187.6785 | 168.7685 | 174.1961 | 352.9485 | 142.0354 | 126.8939 | 157.82 | 213.2669 | 131.0354 | 117.09 | 127.0354 | 134.2862 | 147.9646 | 149.6077 | 152.18 | 119.4277 | 95.6431 | 99.7492 | 96.3569 | 102.9646 | 195 | 205.3569 | 222.3923 | 340.4984 | 204.1254 | 183.4823 | 207.8554 | 317.6785 | 154.8746 | 62.9646 | 72.7138 | 154.8746 | 99.7138 | 102.6785 | 112.0354 | 120.4984 | 107.9261 | 103.3923 | 96.3569 | 101.6431 | 57.9292 | 77.1415 | 82.6077 | 85.1254 | 93.0546 | 80.3923 | 102.6431 | 77.2862 | 210.6238 | 198.9292 | 200.1961 | 471.5016 | 123.8939 | 178.5531 | 178.4631 | 407.3215 | 267.193 | 168.1254 | 178.4823 | 372.4631
1805 | 21.8494 | 124.4612 | 164.0963 | 178.2114 | 170.5763 | 129.6343 | 146.4424 | 161.0196 | 155.4041 | 129.0384 | 93.5959 | 113.0384 | 163.7698 | 187.7886 | 153.6735 | 131.9233 | 214.1543 | 164.9233 | 115.8653 | 126.1543 | 189.4424 | 128.9037 | 306.7126 | 264.2686 | 145.6922 | 165.1535 | 204.6922 | 208.769 | 322.6155 | 105.942 | 112.8269 | 116.5771 | 101.1918 | 132.9616 | 114.6539 | 143.231 | 75.4996 | 198.9037 | 199.2114 | 206.2498 | 243.0196 | 111.5192 | 113.0767 | 124.4808 | 119.5771 | 62.2882 | 55.5959 | 73.058 | 87.8653 | 189.6922 | 163.5576 | 203.2114 | 155.8653 | 108.9037 | 118.7306 | 115 | 118.1731 | 144.4416 | 162.9812 | 178.1731 | 164.4808 | 118.3265 | 121.5576 | 144.4229 | 139.8653 | 152.1731 | 157.3461 | 168.0384 | 246.5388 | 145.7886 | 133 | 219.1159 | 186.5388 | 182.8849 | 177.8269 | 174.8457 | 377.5967 | 138.2114 | 146.4996 | 133.4041 | 202.9804 | 126.3461 | 124.5388 | 124.942 | 136.8465 | 149.0196 | 156.0963 | 173.6539 | 107.9037 | 99.5192 | 87.3078 | 91.9616 | 108.8653 | 198.2882 | 201.6539 | 212.7886 | 318.2114 | 202.7314 | 186.8653 | 186.0188 | 313.231 | 149.5192 | 68 | 64.7886 | 149.5192 | 95.4229 | 99.2694 | 109.4229 | 101.4996 | 74.9804 | 94.3078 | 91.7886 | 106.7306 | 70.0767 | 57 | 89.4424 | 90.1347 | 107.0963 | 73.2114 | 103.058 | 82.7886 | 196.1151 | 210.3845 | 194.9616 | 492.231 | 137.2694 | 169.0963 | 161.2114 | 405.8849 | 219.6343 | 171.5771 | 180.1347 | 356.4229
1806 | 21.8604 | 141 | 158.5217 | 183.6293 | 195.3273 | 141.9955 | 139.1854 | 141.2586 | 147.6338 | 133.2586 | 105.8146 | 116.3318 | 139.5606 | 182.5561 | 132.6293 | 123.2975 | 191.5561 | 155.5561 | 122.4096 | 108.5606 | 179.0344 | 142.8192 | 253.3752 | 299.7414 | 147.1465 | 188.151 | 196.5082 | 216.4394 | 317.8535 | 112.9566 | 113.6682 | 108.8924 | 127.2975 | 127.0732 | 113.9268 | 130.5561 | 94.0344 | 206.8879 | 204.8146 | 218.6682 | 231.9702 | 113.4439 | 123.3707 | 120.8879 | 111.5217 | 77.4439 | 67.4439 | 58.0732 | 98.1166 | 187.6248 | 169.3318 | 207.8879 | 158.8879 | 118.3707 | 124.7758 | 115.9268 | 118.0732 | 174.4394 | 139.3707 | 179.9268 | 163.1121 | 138.6293 | 125.1076 | 154.3318 | 141.2197 | 153.1854 | 161.2242 | 181.3409 | 238.1555 | 135.7369 | 131.8879 | 192.6682 | 178.3409 | 171.3363 | 178.6682 | 201.707 | 349.4096 | 141.4051 | 159.8445 | 122.2975 | 221.2586 | 126.5172 | 114.858 | 136.9611 | 120.707 | 138.5262 | 147.7414 | 172.3707 | 119.5949 | 98.6637 | 88.2631 | 86.0732 | 111.7025 | 208.2541 | 194.6248 | 203.849 | 323.2586 | 173.9313 | 194.3363 | 207.9611 | 302.5949 | 152.1854 | 69.2975 | 61.0389 | 152.1854 | 106.8146 | 93.5949 | 119.3318 | 124.1121 | 91.5904 | 86.1808 | 88.0389 | 112.7758 | 79.8146 | 59.9656 | 79.7758 | 88.9656 | 98 | 77.5172 | 91.0389 | 74.9611 | 210.4439 | 214.4051 | 195.0045 | 477.7025 | 130.4828 | 165.5606 | 171.4485 | 397.4873 | 235.5172 | 163.151 | 178.9656 | 378.009
1807 | 21.8714 | 142.7788 | 174.7906 | 187.9293 | 166.9882 | 122.9647 | 135.6518 | 140.5694 | 163.0589 | 128.2212 | 105.7906 | 117.6963 | 161.233 | 181.2448 | 129.2212 | 130.7788 | 193.4071 | 156.2212 | 131.6165 | 137.3482 | 166.3953 | 161.8259 | 278.9764 | 295.2919 | 144.9293 | 197.4542 | 156.6727 | 196.9411 | 317.5013 | 87.9293 | 116.4542 | 124.2212 | 130.233 | 119.4424 | 118 | 127.466 | 80.4071 | 202.3953 | 206.767 | 219.4778 | 266.4071 | 110.4071 | 125.1976 | 118.1741 | 124.8259 | 76.1859 | 67.3717 | 57.5576 | 117.6872 | 170.9175 | 164.767 | 204.7788 | 155.7788 | 123.1623 | 113.8141 | 120.1976 | 116.9647 | 150.3953 | 142.9764 | 181.0353 | 161.0825 | 133.4424 | 111.3599 | 146.2094 | 129.9882 | 148.2683 | 168.8259 | 209.8966 | 272.2212 | 117.9882 | 132.7317 | 197.6283 | 209.6636 | 188.767 | 184.0236 | 207.3273 | 354.2683 | 134.5458 | 124.1976 | 129.9764 | 213.0589 | 118.8141 | 136.8966 | 128.9882 | 138.9764 | 168.8848 | 151.3953 | 175.9764 | 131.1976 | 85.3835 | 105.6165 | 89.3128 | 107.7788 | 185.767 | 176.534 | 194.5458 | 321.3835 | 202.1741 | 204.8495 | 194.8495 | 313.0118 | 154.3835 | 72.6283 | 68.8141 | 154.3835 | 104.2212 | 107.3717 | 110.0236 | 128.8024 | 81.9882 | 69.1505 | 96.0118 | 105.9647 | 78.6047 | 72.4071 | 69.8024 | 101.0118 | 97.4071 | 72.1859 | 99.6047 | 66.5929 | 207.6047 | 199.2448 | 217.1976 | 477.7317 | 132.0589 | 185.4542 | 190.2094 | 421.2566 | 231.9647 | 178.9764 | 194.767 | 421.0118
1808 | 21.8824 | 144.1218 | 176.3205 | 206.3318 | 177.9278 | 135.2099 | 125.7676 | 151.8397 | 139.4311 | 117.851 | 106.9007 | 147.7516 | 151.8894 | 167.6795 | 126.4086 | 143.2484 | 199.8172 | 152.9888 | 126.6298 | 152.4808 | 166.7404 | 149.4311 | 282.3814 | 274.2868 | 169.8397 | 180.2596 | 193.492 | 217.8397 | 284.0497 | 111.7901 | 97.1603 | 116.1603 | 120.6795 | 113.1874 | 121.7788 | 108.5689 | 77.5801 | 204.2099 | 214.4311 | 194.6795 | 263.5801 | 108.6298 | 121.1715 | 124.6907 | 118.0993 | 83.5192 | 73.2212 | 68.641 | 92.6186 | 192.492 | 179.7788 | 206.9616 | 157.5417 | 135.3702 | 108.5801 | 126.0384 | 127.9503 | 147.3814 | 150.1603 | 165.4311 | 138.4583 | 120.6795 | 123.3205 | 146.149 | 135.4695 | 130.4583 | 157.4808 | 179.0993 | 257.4424 | 125.7788 | 148.8622 | 191.149 | 188.4199 | 197.6907 | 176 | 162.9503 | 336.4583 | 151.9503 | 126.4695 | 138 | 197.6186 | 116.0993 | 107.5689 | 135.7291 | 144.0609 | 143.7788 | 154.2596 | 178.3318 | 126.3318 | 85.5417 | 95.5914 | 113.7516 | 116.4695 | 193.851 | 197.1106 | 211.1106 | 326.5801 | 210.1603 | 186.9503 | 186.6073 | 311.7291 | 162.3093 | 64.8894 | 64.4199 | 162.3093 | 97.4199 | 115.5305 | 110.1874 | 124.851 | 86.2099 | 88.3093 | 94.0993 | 112.9616 | 75.7404 | 68.9503 | 66.9007 | 94.9007 | 95.4199 | 78.6795 | 97.1603 | 72.7788 | 201.5914 | 187.5689 | 218.8397 | 502.2596 | 116.8285 | 172.8782 | 190.7788 | 410.6073 | 241.0609 | 183.2212 | 202.0112 | 418.8894
1809 | 21.8934 | 122.2222 | 166.6666 | 187 | 194.8888 | 137.3334 | 148.9999 | 152.1111 | 130.0001 | 116.889 | 101 | 123.7777 | 152.5555 | 179.3334 | 140.5555 | 158.9999 | 225.4443 | 163.111 | 127 | 143.2222 | 164.8889 | 134.6666 | 261.4444 | 306.111 | 170.1111 | 183.8889 | 174.3334 | 218.3333 | 299.5557 | 106.3333 | 92.2222 | 109.4444 | 125.8889 | 140.2222 | 128.4443 | 123.1111 | 79.5556 | 199.4444 | 199.2222 | 205.889 | 255.9999 | 108.1111 | 108.5556 | 111.5555 | 122.2222 | 90.5555 | 59.6667 | 59.5556 | 109.7778 | 171.5556 | 189.5555 | 188.2222 | 145.4446 | 129.2222 | 111.4445 | 139.2221 | 124.2222 | 132.4445 | 147.2222 | 160.0001 | 161.4445 | 125.2222 | 113.8888 | 155.3333 | 139.8889 | 148.7777 | 153.3334 | 185.2222 | 229.4445 | 141.1111 | 121.7778 | 200.3333 | 192 | 181.6666 | 177.1111 | 159.2222 | 354.7777 | 146.8889 | 133.3334 | 135.7777 | 219.2223 | 125.7778 | 125.2223 | 145.1111 | 135.2223 | 161.3334 | 159.4445 | 165.0001 | 108.1112 | 73.2223 | 83.0001 | 98.6667 | 119.5555 | 186.6667 | 194.0001 | 212.4446 | 333.2223 | 208.7778 | 188.1112 | 220.5556 | 324 | 168.7777 | 70 | 69.1112 | 168.7777 | 99.8889 | 110.8889 | 139 | 112.3333 | 92.5556 | 95.8889 | 101.5555 | 101.1112 | 72.7778 | 65.8889 | 60.3334 | 85.8889 | 98.7778 | 83.6666 | 90.4444 | 85.2222 | 189.6667 | 197.8888 | 220.6666 | 496.3332 | 137.2223 | 191.8888 | 191.4442 | 435.8888 | 234.2223 | 171.6667 | 180.3332 | 409.7775
1810 | 21.9044 | 127.2213 | 158.0487 | 192.8274 | 167.4558 | 142.469 | 137.3807 | 144.8142 | 149.8629 | 148.4335 | 111.0487 | 111.3452 | 134.6903 | 193.7787 | 130.3584 | 142.8894 | 203.9248 | 150.6903 | 120 | 141.8894 | 174.2965 | 128.3584 | 254.8629 | 278.0132 | 164.4558 | 180.2345 | 188.0752 | 206.6548 | 327.5706 | 91.1726 | 88.1106 | 104.0355 | 125.9868 | 128.8761 | 112.0842 | 130.2832 | 92.7035 | 188.9116 | 194.1726 | 227.2213 | 238.5177 | 104.3319 | 124.0132 | 99.5177 | 106.8894 | 82.9381 | 59.6548 | 69.4203 | 109.2345 | 181.4203 | 175.1726 | 192.2832 | 171.0132 | 117.6416 | 120.9026 | 117.3584 | 120.5442 | 144.0487 | 143.5442 | 181.9026 | 169.0619 | 119.9381 | 108.1594 | 148.8142 | 135.2965 | 140.2965 | 171.7655 | 181.3452 | 229.5929 | 142.6548 | 124.5177 | 190.7655 | 197.3584 | 158.8142 | 184.0487 | 163.9868 | 347 | 141.3584 | 138.469 | 138.7878 | 228.8406 | 131.531 | 137.6548 | 142.7035 | 143.469 | 165.4955 | 157.0132 | 189.9513 | 122.7523 | 102.2345 | 95.5929 | 112.6548 | 109.8274 | 199.0132 | 207.2965 | 233.3229 | 359.4203 | 216.4071 | 204.8274 | 222.7168 | 323.73 | 157.7168 | 64.6416 | 81.1726 | 157.7168 | 104.8761 | 115.8761 | 129.6548 | 109.0619 | 106.1371 | 94.5797 | 97.8894 | 122.9513 | 77.1239 | 73.6903 | 66.6416 | 85.2832 | 103.1239 | 81.3452 | 79.4071 | 78.531 | 206.9381 | 182.3097 | 214.5929 | 464.8629 | 146.8761 | 182.6416 | 151.0355 | 414.9868 | 256.5045 | 183.5797 | 160.9868 | 360.5045
1811 | 21.9154 | 140.7532 | 164.7915 | 193.9362 | 186.6893 | 141.7787 | 165.7787 | 154.2468 | 169.2468 | 125.2723 | 117.5447 | 116.2851 | 156.1702 | 182.7149 | 124.5064 | 141.9234 | 222.2213 | 160.5702 | 123.4553 | 131.5447 | 171.0128 | 126.2085 | 274.2468 | 273.1447 | 174.0638 | 184.9489 | 179.651 | 199.2468 | 286.7532 | 94.5064 | 96.9745 | 120.0894 | 138.0128 | 135.9872 | 137.349 | 142.234 | 93.2723 | 224.417 | 197.0128 | 239.2723 | 251.4808 | 126.234 | 111.9872 | 113.7149 | 104.1957 | 81.3107 | 53.9745 | 61.4936 | 105.0638 | 166.3362 | 179.7404 | 205.2213 | 158.4936 | 124.7277 | 102.7277 | 115.2085 | 110.1957 | 154.9872 | 131.9617 | 172.8595 | 161.3107 | 125.4681 | 126.234 | 153.0638 | 136.9489 | 138.9872 | 165.817 | 188.0128 | 228.7021 | 137.2213 | 136.7404 | 195.4298 | 192.7404 | 174.1702 | 197.4553 | 180.7021 | 347.2468 | 135.2596 | 140.4936 | 167.8298 | 217.9234 | 128.5192 | 129.0128 | 143.5192 | 140.5575 | 140.0128 | 150.417 | 180.9872 | 135.7404 | 105.9617 | 93.2213 | 105.7404 | 107.7277 | 190.4425 | 202.7787 | 204.5064 | 354.9489 | 218.5319 | 199.0255 | 211.0128 | 302.9617 | 154.8978 | 71.9745 | 76.8553 | 154.8978 | 110.5064 | 112.1277 | 120.5192 | 105.0128 | 91.6893 | 103.2468 | 93.7149 | 112.7532 | 69.5192 | 87.0255 | 70.766 | 94.2723 | 97.2468 | 86.7787 | 84.9872 | 80.2085 | 205.5575 | 175.3915 | 209.0128 | 468.2043 | 147.0766 | 182.817 | 174.2468 | 420.1022 | 274.8298 | 189.7787 | 170.5447 | 381.5447
1812 | 21.9264 | 139.4818 | 166.1412 | 205.6317 | 194.7816 | 136.0364 | 153.5589 | 155 | 165.8544 | 120.5182 | 108.0364 | 115.5139 | 170.8544 | 191.7773 | 124.8137 | 152.227 | 223.2998 | 147.8865 | 134 | 122.5546 | 171.6274 | 131.8544 | 265.9315 | 295.1499 | 159 | 198.2955 | 196.7816 | 201.8137 | 278.4861 | 97.4047 | 103 | 98.2227 | 139.1456 | 140.2955 | 119.4411 | 139.2634 | 85.6681 | 232.3406 | 190.6317 | 232.4454 | 258.591 | 127.1499 | 114.7409 | 124.0728 | 117 | 68.2955 | 61.0364 | 63.5182 | 92.0728 | 152.9957 | 179.1499 | 202.1543 | 155.8544 | 129.2227 | 103.3362 | 119.5589 | 121.9636 | 158.5182 | 139.4454 | 201.8544 | 147.5182 | 129.6317 | 126.6317 | 142.1456 | 148.7409 | 147.9593 | 156.4775 | 191.9957 | 236.8908 | 145.5546 | 143.6638 | 200.8588 | 192.9272 | 183.6724 | 202.2998 | 193.8137 | 349.8137 | 129.8908 | 145.6274 | 141.227 | 236.2591 | 125.2955 | 123.9272 | 137.4454 | 132.182 | 137.5182 | 166.6317 | 177.0043 | 130.4861 | 112.409 | 98.9636 | 104.1135 | 111.7045 | 198.818 | 195.4818 | 201.9636 | 355.3406 | 204.5225 | 197.9229 | 206.7045 | 313.2955 | 171.4861 | 80.8501 | 55.6274 | 171.4861 | 109.7773 | 90.0321 | 120.6638 | 100.4454 | 101.5953 | 101.6681 | 106.4047 | 110.4454 | 66.2955 | 76.5953 | 67.2591 | 86.9272 | 99.2955 | 83.6274 | 93.4411 | 88.1863 | 195.8865 | 201 | 200.5589 | 439.1733 | 127.4454 | 168.0364 | 180.4411 | 394.9636 | 256.2591 | 179.1135 | 164.4047 | 371.7773
1813 | 21.9374 | 143.9706 | 187.1444 | 192.4719 | 185.4131 | 135.7433 | 140.9853 | 151.4719 | 155.6283 | 129.3275 | 111 | 132.6578 | 163.885 | 196.4425 | 132.1711 | 139.2273 | 208.3569 | 156.5575 | 135.6283 | 131.3422 | 183.8997 | 123.2567 | 258.4545 | 286.1858 | 161.7139 | 199.0147 | 178.4572 | 207.2714 | 271.1123 | 108.3717 | 102.1858 | 95.7286 | 143.4011 | 141.2861 | 132.0147 | 122.4425 | 85.242 | 201.4425 | 184.7995 | 219.0441 | 260.3008 | 117.6431 | 113.1858 | 123.4866 | 117 | 70.6431 | 63.1858 | 67.4425 | 101.7995 | 178.4278 | 173.7139 | 186.2125 | 151.8703 | 127 | 95.1564 | 109.9706 | 116.8289 | 160.5428 | 139.3422 | 193.2567 | 155.7847 | 124.0708 | 122.9706 | 152.9144 | 143.9292 | 156.8583 | 171.2005 | 201.4158 | 229.3569 | 149.1858 | 151.5722 | 180.9411 | 185.1003 | 159.9853 | 182.4719 | 195.4719 | 362.3275 | 132.6698 | 156 | 123.0708 | 234.5575 | 126.2861 | 119.8997 | 134.8997 | 145.9144 | 144.9706 | 152.9292 | 161.3422 | 110.0856 | 101.2005 | 97.8997 | 94.0856 | 109.8997 | 182.1003 | 193.7286 | 203.6136 | 322.5428 | 196.5842 | 212.5428 | 214.9411 | 324.8703 | 151.3569 | 88.1858 | 62.4719 | 151.3569 | 100.873 | 107.3008 | 129.3864 | 100.885 | 92.5281 | 101.242 | 114.115 | 108.1711 | 72.4425 | 66.4425 | 68.5428 | 83.9853 | 106.2567 | 92.3717 | 100.5869 | 82.7286 | 205.6431 | 195.3008 | 190.1564 | 479.6163 | 121.1003 | 172.8997 | 190.5722 | 395.7995 | 254.2861 | 164.4719 | 178.8997 | 374.2714
1814 | 21.9484 | 150.6385 | 159.4184 | 182.7163 | 194.667 | 132.3899 | 151.2694 | 137.1774 | 152.908 | 141.908 | 116.9573 | 123.2694 | 171.3264 | 188.816 | 136.2979 | 155.4896 | 203.7733 | 149.4326 | 136.5958 | 143.2837 | 188.4326 | 127.4611 | 298.6385 | 278.3264 | 167.5816 | 191.2837 | 184.9431 | 206.8653 | 306.4469 | 104.2837 | 100 | 97.8368 | 118.1774 | 132.2979 | 125.7021 | 134.9573 | 98.5958 | 208.2837 | 194.149 | 199.2552 | 242.7306 | 112.0142 | 110.7163 | 114.5104 | 120.9715 | 66.149 | 67.8083 | 71.4469 | 110.5816 | 183.9222 | 177.8795 | 215.3122 | 170.1632 | 125.8653 | 114.6736 | 127.1347 | 109.8653 | 153.4896 | 154.6878 | 198.3122 | 168.0427 | 131.3122 | 137.0142 | 146.3122 | 131.0142 | 138.6878 | 159.5816 | 178.9573 | 233.5674 | 141.0427 | 133.8795 | 208.1774 | 183.4042 | 172.8226 | 176.6878 | 181.7448 | 376.8937 | 166.2694 | 147.7733 | 133.7163 | 224.3122 | 118.4326 | 121.5958 | 136.5958 | 143.2837 | 157.5958 | 142.2837 | 167.3264 | 114.9858 | 94.1205 | 93.9222 | 100.4042 | 114.149 | 181.8226 | 194.4184 | 218.5531 | 326.5531 | 217.0569 | 210.8795 | 237.0712 | 338.0569 | 153.0142 | 77.2059 | 53.851 | 153.0142 | 81.2125 | 100.5104 | 115.2979 | 116.5531 | 97.1774 | 112.8937 | 97.5958 | 110.0285 | 73.3264 | 69.0285 | 71.4184 | 94.2694 | 111.3122 | 86.2979 | 81.9715 | 77.7448 | 206.8226 | 188.2267 | 204.2837 | 453.482 | 128.482 | 176.0142 | 180.8226 | 409.6878 | 240.1917 | 162.092 | 187.9715 | 379.8226
1815 | 21.9594 | 126.4081 | 168.9194 | 192.951 | 168.1033 | 136.8249 | 159.9597 | 134.1751 | 137.3678 | 124.8879 | 126.6725 | 120.1208 | 151.0806 | 159.5114 | 129.6322 | 132.8161 | 192.1751 | 153.9194 | 126.5201 | 149.9194 | 187.296 | 121.9194 | 280.6235 | 269.6235 | 157.7846 | 194.3678 | 199.6725 | 207.2557 | 311.5429 | 108.5517 | 98.8161 | 107.6637 | 112.5114 | 126.2242 | 128.8161 | 136.0893 | 93.5517 | 212.8476 | 198.5114 | 215.9194 | 246.1751 | 117.655 | 112.3678 | 133.6322 | 121.5289 | 68.4396 | 80.5604 | 65 | 106.4081 | 167.2154 | 170 | 201.8564 | 160.3363 | 122.4081 | 121.6007 | 127.6322 | 106.4081 | 133.7758 | 161.0403 | 176.8652 | 150.6322 | 124.6637 | 129.0403 | 139.0718 | 125.1121 | 154.8074 | 162.8074 | 184.8249 | 234.704 | 138.4308 | 125.1121 | 192.4483 | 195.8476 | 179.0806 | 184.8161 | 170.1121 | 360.1524 | 154.8336 | 130.8476 | 134.4799 | 220.0315 | 117.8879 | 110.6322 | 132.4396 | 144 | 155.5114 | 142.4081 | 157.7355 | 118.5201 | 101.704 | 78.9912 | 108.704 | 112 | 189.8564 | 199.7758 | 217.8967 | 321.1611 | 192.8564 | 205.9597 | 209.0403 | 303.2014 | 149.4799 | 56.4799 | 61.3275 | 149.4799 | 101.6725 | 121.4081 | 110.704 | 122.704 | 89.1436 | 95.8564 | 94.9194 | 104.7355 | 61.6637 | 59.5919 | 73.2242 | 98.2242 | 97.2645 | 83.7758 | 94.9597 | 67 | 213.3765 | 204.8564 | 200.8564 | 476.4168 | 154.3765 | 168.0403 | 185.8967 | 425.8074 | 235.7268 | 178.5201 | 192.3765 | 397.6235
1816 | 21.9704 | 127.1578 | 185.1578 | 206.3612 | 186.8496 | 119.6239 | 164.5339 | 157.2331 | 140.2256 | 128.5413 | 115.8496 | 99.9248 | 139.1578 | 171.3008 | 125.2331 | 135.5487 | 210.3008 | 159.6769 | 123.6165 | 153.2108 | 189.8496 | 139.0826 | 288.4513 | 284.2331 | 140.5339 | 194.7595 | 196.2479 | 207.6017 | 317.4216 | 113.9174 | 101.2405 | 112.4587 | 129.2331 | 118.9174 | 124.7669 | 109.7818 | 94.91 | 208.7447 | 205.3686 | 212.1208 | 257.2108 | 134.9852 | 119.2331 | 123.3761 | 94.3761 | 79.3083 | 68.7669 | 65.3083 | 107.4661 | 190.9396 | 171.8496 | 194.1578 | 155.2331 | 121.9248 | 104.6165 | 119.5339 | 105 | 138.9248 | 158.9322 | 155.09 | 146.8496 | 131.3083 | 123.8496 | 140.6091 | 120.5339 | 166.0678 | 175.3008 | 168.857 | 237.3909 | 158.4439 | 123 | 187.6992 | 204.0752 | 174.2479 | 176.143 | 179.714 | 360.5561 | 123.5561 | 139.0752 | 139.2331 | 227.0678 | 122.4661 | 113.9396 | 136.8347 | 146.7744 | 174.0826 | 141.6165 | 165.3008 | 121.7818 | 97.6091 | 96.4587 | 105.5339 | 108.3008 | 180.6165 | 195.3686 | 196.3761 | 291.1504 | 193.7892 | 205.2934 | 204.7744 | 277.8644 | 159.7818 | 60 | 71.8422 | 159.7818 | 96.3983 | 114.4513 | 106.9174 | 125.0826 | 93.7595 | 90.6239 | 110.5413 | 118.4661 | 67.0752 | 65.6239 | 69.9248 | 93.6917 | 92.1653 | 86.1504 | 99.8422 | 71.9322 | 207.09 | 188.2182 | 193.1578 | 458.7744 | 136.9926 | 165.0074 | 173.6239 | 433.0604 | 258.2934 | 179.3157 | 180.5413 | 409.1504
1817 | 21.9814 | 132.9617 | 189.3589 | 183.1289 | 195.8083 | 130 | 151.6273 | 153.9095 | 138.8083 | 130.7178 | 116.7945 | 105.2055 | 150.7316 | 169.7316 | 131.5261 | 146.0767 | 202 | 140.3589 | 119.23 | 129.6411 | 186.3067 | 143.115 | 279.2055 | 287.6411 | 138.2055 | 184.6028 | 209.9478 | 204.1427 | 285.0277 | 101.8711 | 107.8711 | 109 | 139.3727 | 117.4494 | 118.7945 | 120.8328 | 80.2822 | 185.7316 | 189.8466 | 188.9894 | 241.655 | 115.7178 | 124.885 | 117.8083 | 94.2577 | 75.8328 | 65.6794 | 70.4878 | 114.6028 | 204.5889 | 173.7561 | 191.9478 | 153.1917 | 120.1534 | 105.6794 | 109.5122 | 105 | 141.9617 | 162.9478 | 177.0905 | 152.9233 | 129.115 | 131.2055 | 138.4494 | 121.7316 | 159.8083 | 164.4356 | 175.5889 | 240.1917 | 136.6794 | 130.0522 | 188.6273 | 199.7561 | 184.7423 | 172.6166 | 192.8573 | 372.3834 | 143.7178 | 133.1534 | 138.7945 | 211.5122 | 134.0905 | 129.5122 | 127.1672 | 143.7039 | 169.7806 | 144.2822 | 166.9372 | 136.6794 | 95.4494 | 94.9233 | 102.5644 | 98.3972 | 193.8605 | 181.1289 | 185.3589 | 303.0277 | 214.1534 | 187.3589 | 203.3067 | 301.6028 | 168.9095 | 62.885 | 63.7945 | 168.9095 | 107.8711 | 98.1534 | 100.6411 | 123.9862 | 93.8605 | 94.2684 | 114.3206 | 117.9095 | 64.3589 | 78.2439 | 68.1534 | 91.7178 | 101.1917 | 85.8466 | 98.2055 | 79.4739 | 220.4356 | 173.9617 | 199.9233 | 467.2439 | 130.5644 | 171.115 | 176.6273 | 406.9095 | 236.8328 | 189.6411 | 181.4356 | 390.5751
1818 | 21.9924 | 138.6613 | 188.6657 | 190.0059 | 199.6759 | 134.9927 | 137.9971 | 148.6554 | 149.9985 | 125.6701 | 114.6598 | 109.6701 | 150.022 | 186.9956 | 138.3343 | 145.3284 | 203.3314 | 141.9956 | 110.6613 | 137.9897 | 175.9912 | 144.6554 | 280.3416 | 289.6657 | 152.3226 | 190.6628 | 189.6745 | 225.3431 | 317.6686 | 101.9839 | 105.654 | 107.0029 | 151.3387 | 130.3314 | 116.327 | 120.9868 | 80.6701 | 201.6642 | 202.6598 | 220.3402 | 249.0161 | 123.651 | 120.6818 | 122.6745 | 121.3255 | 70.3284 | 62.0044 | 77.3372 | 117.6672 | 186.673 | 171.6628 | 191.978 | 138.673 | 111.3343 | 110.3255 | 106.3241 | 105 | 148.9927 | 150.3299 | 189.3343 | 157.3328 | 123.6657 | 138.3328 | 147.6701 | 136.6657 | 159.349 | 154.3402 | 163.6642 | 243.6466 | 130.3416 | 146.3314 | 181.9868 | 194.0015 | 167.327 | 182.0161 | 180.9766 | 350.6686 | 142.3314 | 124.3343 | 132 | 211.3197 | 137.6803 | 121.6642 | 126.3475 | 124 | 148.3299 | 147 | 177.0161 | 129.3431 | 106.6672 | 94.3402 | 101.0103 | 100.9912 | 213.6745 | 187.673 | 187.3284 | 329.3446 | 202.6716 | 195.9853 | 195.654 | 305.9985 | 159.9941 | 66.3372 | 58.9971 | 159.9941 | 97.6657 | 96.3241 | 102.3328 | 108.9971 | 110.346 | 89.3197 | 108.3431 | 109.9927 | 67.6598 | 80.0044 | 63.6613 | 92.3284 | 89.0029 | 91.6701 | 97.3446 | 71.3343 | 213.6686 | 178.3299 | 204.6657 | 476.6598 | 131.0073 | 162.6701 | 167.6569 | 386.6774 | 233.6584 | 188.6701 | 176.6657 | 371.6496
1819 | 22.0034 | 143.9292 | 182.7521 | 195.1152 | 181.4778 | 141.8938 | 140.2743 | 156.7521 | 150.6195 | 132.7346 | 122.2743 | 113.9735 | 134.8408 | 193.3451 | 141.4868 | 146.823 | 208.7611 | 156.2833 | 122.4868 | 144.7521 | 178.6813 | 152.407 | 281.0798 | 294.1062 | 164.584 | 197.0354 | 175.8938 | 213.7257 | 309.1327 | 116.0619 | 115.7521 | 103.3451 | 137.7876 | 130.9292 | 123.9646 | 135.5486 | 83.2479 | 208.4514 | 204.7521 | 212.0354 | 222.5132 | 137.0619 | 108.9735 | 112 | 127.1681 | 77.3451 | 59.4514 | 73.3805 | 115.6195 | 182.1416 | 182.5222 | 205.0883 | 138.2833 | 109.3097 | 124.4514 | 114.5132 | 101.5486 | 154.1681 | 154.6549 | 183.1681 | 155.2389 | 121.5486 | 133.823 | 148.177 | 136.2743 | 135.2389 | 147.416 | 169.0708 | 259.6105 | 118.3097 | 146.584 | 202.416 | 195.4514 | 180.8319 | 166.2124 | 206.4424 | 349.0354 | 139.823 | 122.3097 | 135.4514 | 219.9911 | 116.9292 | 126.3805 | 106.3097 | 136.77 | 157.8319 | 151.8319 | 158.7965 | 117.0354 | 104.2743 | 91.5576 | 93.9027 | 109.8938 | 196.0973 | 195.2214 | 191.584 | 316.7611 | 198.7611 | 206.6813 | 206.7876 | 305.2389 | 161.4424 | 62.7257 | 64.0354 | 161.4424 | 100.0354 | 109.6903 | 104.7257 | 114.7257 | 90.584 | 96.6105 | 95.3451 | 116.5486 | 71.4778 | 77.1062 | 69.9646 | 88.9911 | 92.593 | 88.0354 | 78.5486 | 68.2743 | 208.584 | 186.1062 | 210.4868 | 481.8584 | 127.9027 | 160.416 | 182.3805 | 386.1506 | 251.5576 | 190.2124 | 185.9381 | 373.6016
1820 | 22.0144 | 129.6346 | 167.9277 | 203.5584 | 169.5703 | 131.3534 | 133.0683 | 152.2931 | 143.3534 | 143.5623 | 110.0643 | 126.7109 | 162.6426 | 191.1406 | 144.2812 | 142.3614 | 210.4257 | 166.2812 | 131.7149 | 141.008 | 172.0802 | 138 | 305.0723 | 292.1366 | 158.2129 | 189.7069 | 182.5099 | 212.3534 | 316.9436 | 105.2891 | 101.6426 | 104 | 129.5743 | 123.4257 | 119.1406 | 132.2168 | 99.502 | 210.3534 | 194.5743 | 213.2851 | 212.5703 | 119.1406 | 124.5703 | 111.2851 | 118 | 76.2129 | 62.0683 | 77.7871 | 118.004 | 181.0643 | 187.6386 | 188.0119 | 149.3534 | 106.9277 | 122.0643 | 111.004 | 96.4297 | 144.2851 | 148.2812 | 167.9237 | 149.6426 | 126.4377 | 132.9357 | 152.2812 | 134.7871 | 137.506 | 149.1406 | 178.004 | 237.6426 | 128.7951 | 135.9237 | 197.7069 | 203.7871 | 179.6346 | 179.4297 | 188.996 | 355.2891 | 128.9277 | 130.2931 | 144.1446 | 206.7188 | 118.004 | 125.0683 | 111.4337 | 153.8515 | 159.8515 | 160.2851 | 167.0723 | 119.7149 | 95.996 | 97.2089 | 105.2129 | 107.2168 | 189.4337 | 216.7069 | 186.6426 | 325.2168 | 193.2772 | 195.7911 | 206.0802 | 306.7911 | 150.7871 | 63.1406 | 63.8554 | 150.7871 | 99.1406 | 103.1366 | 104.4257 | 116.9277 | 87.4297 | 81.7911 | 94.9277 | 112.502 | 59.2129 | 73.3495 | 70.1446 | 77.8634 | 105.2129 | 90 | 82.7228 | 67.8594 | 205.4297 | 184.8515 | 223.2891 | 469.7109 | 139.5703 | 159.6386 | 189.2891 | 421.5783 | 262.2129 | 192.3495 | 185.6306 | 346.2248
1821 | 22.0254 | 125.419 | 173.7641 | 181.3944 | 173.2852 | 122.2606 | 132.2852 | 163.0845 | 144.243 | 118.9331 | 95.4789 | 114.9331 | 160.5211 | 183.412 | 132.8908 | 149.0423 | 206.588 | 156.3697 | 131.5211 | 152.3451 | 180.5387 | 143.9155 | 291.1021 | 280.2183 | 154.2606 | 185.2007 | 197.6056 | 195.4965 | 336.9507 | 111.5211 | 100.2606 | 98.4542 | 129.7148 | 114.7817 | 112.8908 | 139.1092 | 96.2359 | 203.4789 | 209.8732 | 208.3028 | 222.5704 | 114 | 111.2782 | 117.7641 | 115.7817 | 78.5458 | 60.1761 | 79.8908 | 119.2359 | 163.5211 | 174.3028 | 211.6725 | 145.0669 | 111.6549 | 113.3944 | 118.5211 | 107.8732 | 143.7394 | 145.7641 | 162.1761 | 153.4366 | 137.0176 | 136.0599 | 143.8486 | 133.9331 | 153.9577 | 144.7394 | 182.1937 | 224.7993 | 146.6725 | 122.0423 | 187.2852 | 194.0599 | 158.7817 | 179.0423 | 182.588 | 374.4613 | 135.1338 | 144.0423 | 140.6056 | 220.3275 | 127.3697 | 127.243 | 119.6725 | 141 | 145.1514 | 151.6056 | 163.8239 | 121.3697 | 96.6127 | 89.8063 | 107.5458 | 108.9331 | 204.6972 | 201.8486 | 196.7218 | 336.5458 | 176.9577 | 209.4789 | 221.9331 | 309.0176 | 152.5211 | 61.3275 | 62.588 | 152.5211 | 93.2606 | 90.4789 | 108.3521 | 120.1761 | 91.1092 | 100.2852 | 100.3944 | 116.6303 | 56.3697 | 62.2852 | 72.1514 | 89.412 | 103.1092 | 93.3275 | 102 | 78.1761 | 214.2852 | 178.2852 | 231.3697 | 462 | 133.3028 | 153.3275 | 193.3028 | 429.9754 | 252.3451 | 174.2606 | 167.9824 | 378.5035
1822 | 22.0364 | 143.8879 | 192.7301 | 197.8201 | 175.9778 | 128.2582 | 132.3037 | 144.5958 | 155.4159 | 126.1343 | 101.8201 | 114.4942 | 149.2138 | 179 | 129.854 | 138.6519 | 216.3481 | 160.4381 | 130.146 | 137.9439 | 162.91 | 150.5619 | 264 | 284.382 | 157.9661 | 205.618 | 183.854 | 186.3481 | 305.3598 | 105.9439 | 97.09 | 93.5841 | 117.854 | 107.9439 | 112.91 | 138.3259 | 83.764 | 199.1238 | 216.2138 | 209.6402 | 255.0783 | 118.2021 | 98.3481 | 126.0339 | 108.9439 | 81.8879 | 69 | 73.0339 | 102.1799 | 168.2582 | 177.5502 | 204.472 | 148.1799 | 112.6858 | 122.1799 | 111.7979 | 121.472 | 143.09 | 162.4381 | 170.236 | 162.146 | 118.854 | 119.3481 | 138.9778 | 131.9661 | 159 | 150.2021 | 172.09 | 228.8306 | 141 | 123.1121 | 192.2699 | 183.4603 | 167.6063 | 173.618 | 182.0339 | 384.0678 | 142.8879 | 136.7079 | 131.8201 | 218.7418 | 124.5619 | 134.9778 | 117.8201 | 138.7079 | 146.2021 | 150.4603 | 158.4381 | 128.8762 | 103.4498 | 101.4381 | 107.8318 | 102.382 | 195.3376 | 201.1799 | 203.1577 | 344.09 | 185.8201 | 219.6402 | 215.764 | 285.8879 | 155.3481 | 71.5841 | 64.3259 | 155.3481 | 93.146 | 93.764 | 118.7418 | 133.2021 | 101.7862 | 105.6519 | 106.5058 | 121.3481 | 57.764 | 75.2921 | 70.91 | 88.8201 | 90.8657 | 93.2699 | 99.7079 | 79.7418 | 218.1238 | 198.9322 | 216.3376 | 461.618 | 126.618 | 162.4381 | 176.3037 | 401.1916 | 236.7979 | 170.3259 | 191.1121 | 392.236
1823 | 22.0474 | 134.7886 | 194.5083 | 207.943 | 172.0689 | 137.6342 | 120.9145 | 139.6746 | 144.057 | 140.7197 | 111.943 | 121.057 | 138.5487 | 173.0855 | 139.829 | 138.6746 | 224.6057 | 160.8741 | 127.2684 | 131.4228 | 167.9715 | 155.2518 | 256.5083 | 287.3658 | 154.171 | 210.1259 | 187.1259 | 193.0285 | 305.6176 | 99.8171 | 96.3658 | 99.4228 | 122.7031 | 111.2803 | 115.2114 | 132.8171 | 86.5772 | 195.886 | 207.1259 | 226.3373 | 261.5653 | 115.9311 | 110.5487 | 116.0285 | 108.3373 | 72 | 65.057 | 66.1829 | 100.7316 | 170.9311 | 188.6627 | 198.057 | 145.5487 | 104.6746 | 108.5083 | 109.7316 | 119 | 141.9715 | 157.7482 | 172.1544 | 161.6342 | 113.8456 | 120.114 | 132.3088 | 142.3658 | 153.8741 | 164.4917 | 176.886 | 265.057 | 134.6912 | 130.2399 | 191.2803 | 198.057 | 178.5368 | 171.0285 | 172.0285 | 359.3254 | 145.2233 | 128.2684 | 139.5772 | 209.7601 | 120.9715 | 112.9311 | 125.9715 | 139.3373 | 158.1259 | 167.0285 | 162.0285 | 128.171 | 94.7031 | 105.4228 | 89.8456 | 112.4632 | 179.4632 | 197.7601 | 187.1544 | 331.1425 | 195.943 | 221.7482 | 218.9715 | 281.1259 | 160.057 | 75.057 | 69.8575 | 160.057 | 104.4632 | 97.3658 | 104.2399 | 124.6223 | 113.6342 | 97 | 101.9145 | 124.8741 | 62.943 | 67.9596 | 72.0285 | 94.2114 | 81.0119 | 81.6342 | 98.7601 | 82.9834 | 208.9715 | 213.4228 | 200.4632 | 460.6057 | 126 | 165.6342 | 177.1378 | 373.6057 | 237.0974 | 164.8171 | 185.6223 | 402.829
1824 | 22.0584 | 135.5934 | 172.8352 | 201.3957 | 192.0989 | 139.2857 | 128.7802 | 144.4616 | 138.4066 | 133.6923 | 122.0659 | 112.1539 | 147.2527 | 177.8241 | 143.7693 | 135.3298 | 224 | 152.1868 | 120.8132 | 118.022 | 181.5714 | 155.1429 | 244.1868 | 293.4395 | 142.0989 | 204.5825 | 188.4945 | 204.6373 | 303.8573 | 104.5055 | 99.5934 | 99.033 | 134.2198 | 123.5934 | 123.3516 | 125.3077 | 84.1209 | 208 | 205.6484 | 233.4066 | 244.4066 | 102.8132 | 109.9011 | 118.6923 | 110.1209 | 75.2527 | 58.1868 | 67.5934 | 108.4066 | 165.945 | 177.1209 | 187.9341 | 152.6264 | 122.8791 | 104.3186 | 119.4395 | 115.7473 | 136.055 | 150.945 | 178.2198 | 157.1868 | 116.7252 | 115.7252 | 150.5384 | 143.967 | 150.4725 | 171.5275 | 182.9011 | 266.7252 | 125 | 127.6264 | 200.7473 | 193.6264 | 169.6923 | 177.3516 | 169.8132 | 353.7252 | 150.5825 | 130.3516 | 146.4725 | 219.6923 | 126.033 | 112.4175 | 126.967 | 151.6923 | 150.5495 | 160.3407 | 160.2198 | 118.1318 | 105 | 111.1318 | 90.2857 | 120.0879 | 195.2198 | 197.9341 | 198.0989 | 324.5384 | 203.6264 | 201.9341 | 215.4945 | 300.3846 | 152.7802 | 71.033 | 81.3407 | 152.7802 | 115.3407 | 96.1209 | 103.6593 | 108.7252 | 99.022 | 97 | 102.4616 | 114.1539 | 70.6264 | 58.3186 | 70.6264 | 91.3736 | 98.5605 | 82.4725 | 94.4616 | 94.2089 | 209.5605 | 193.5166 | 216.2198 | 475.0439 | 131.2857 | 168.5055 | 198.0879 | 388.4505 | 247.1868 | 169.9121 | 158.3407 | 418.967
1825 | 22.0694 | 139.1888 | 168.8909 | 187.6077 | 185.9202 | 149.5133 | 124.9056 | 139.5399 | 147.3776 | 148.2832 | 118.3657 | 116.3776 | 149.9056 | 185.8524 | 129.7021 | 129.4308 | 218.9734 | 151.8378 | 127.4455 | 108.7021 | 192.3923 | 138.7021 | 271.0651 | 282.3391 | 149.3245 | 189.6077 | 186.9588 | 215.5279 | 284.7287 | 108.9734 | 101.0944 | 100.7434 | 133.4867 | 125.9322 | 133.2301 | 123.2832 | 83.2832 | 210.0944 | 197.0266 | 224.7846 | 246.6755 | 109.4455 | 105.1888 | 117.7846 | 105.5133 | 77.0678 | 57.4189 | 70.7699 | 107.3245 | 175.2301 | 181.3098 | 183.6755 | 155.8378 | 118.2713 | 114.5811 | 112.5279 | 114.3511 | 133.4721 | 149.758 | 176.6489 | 162.2832 | 129.6755 | 126.5811 | 156.7168 | 132.6224 | 148.2035 | 160.8112 | 186.5665 | 262.5014 | 138.4042 | 130.4189 | 196 | 186.3657 | 174.6489 | 179.2713 | 185.242 | 356.6224 | 131.8378 | 143.5811 | 153.4189 | 220.4601 | 121.0412 | 129.0678 | 124 | 146.1769 | 133.8644 | 155.8378 | 159.0678 | 117.4335 | 102.9056 | 110.0146 | 103.0266 | 101.2035 | 199.9322 | 202.8909 | 209.9322 | 342.4455 | 216.4721 | 196 | 199.7168 | 315.7434 | 154.3511 | 72.7434 | 69.7168 | 154.3511 | 111.2566 | 91.0944 | 120.7287 | 110.7846 | 93.4721 | 101.1888 | 95.8644 | 105.8112 | 67.5545 | 65.6489 | 66.7168 | 97.7965 | 95.4189 | 90.6755 | 83.6755 | 81.1888 | 209.7699 | 173.0266 | 218 | 502.4455 | 144.4455 | 170.879 | 179.6224 | 405.1356 | 246 | 180 | 157.1888 | 397.1503
1826 | 22.0804 | 134.6518 | 181.8441 | 185.8077 | 179.2287 | 140.9271 | 134.504 | 147.413 | 148.2207 | 143.6336 | 113.1923 | 114.2025 | 154.7611 | 169 | 137.2753 | 144.6518 | 216.7429 | 154.2935 | 126.8077 | 115.413 | 186 | 147.1377 | 306.9818 | 263.2935 | 149.587 | 190.3947 | 192.8259 | 210.0546 | 297.9818 | 103.2935 | 107.0182 | 100.7247 | 131.2935 | 137.7611 | 122.3947 | 123.3765 | 95.4494 | 203.5142 | 210.0364 | 211.1377 | 252.0182 | 116.1377 | 110.1377 | 108.0182 | 110.7247 | 71.7065 | 64.0364 | 64.7895 | 100.2571 | 176.4676 | 187.9454 | 199.7975 | 158.7247 | 106.7429 | 118.7429 | 104.4676 | 124.6053 | 128.5972 | 136.8988 | 179.7611 | 164.5324 | 128.9818 | 118.6701 | 146.7065 | 127.0364 | 139.4494 | 148.1012 | 195.8077 | 245.3117 | 138.8907 | 137.4676 | 199.0182 | 185.504 | 176.4676 | 166.8806 | 186.8907 | 367.8522 | 141.1923 | 136.5324 | 149.2571 | 221.6417 | 114.3117 | 123.7065 | 129.1741 | 130.4494 | 160.1093 | 162.1741 | 164.917 | 109.917 | 98.2753 | 99.3299 | 113.4494 | 98.0546 | 203.1377 | 205.9271 | 223.4858 | 345.6883 | 224.5324 | 195.1377 | 200.4858 | 319.1741 | 161.1559 | 73.1559 | 68.7611 | 161.1559 | 108.6883 | 95.7247 | 123.6336 | 102.7429 | 100.2389 | 100.9636 | 102.7065 | 104.3117 | 62.587 | 68.3299 | 68.3482 | 111.2935 | 101.1741 | 92.1377 | 80.3947 | 87.8623 | 207.6701 | 187.3299 | 214.9818 | 491.0284 | 147.2571 | 158.413 | 163.2571 | 411.6599 | 256.7793 | 171.3765 | 172.917 | 367.5506
1827 | 22.0914 | 133.8605 | 171.9046 | 184.3128 | 178.0441 | 124.5565 | 157.652 | 144 | 145.0866 | 127.7739 | 119.5654 | 99.9912 | 166.3304 | 178.3128 | 131.0088 | 138.9824 | 213.2438 | 150.6784 | 119.5477 | 124.8605 | 186 | 142.0088 | 290.1395 | 288.5036 | 153.4435 | 189.652 | 198.8605 | 213.5742 | 306.8605 | 105.4435 | 110.1131 | 95.4611 | 127.6784 | 133.3657 | 114.1131 | 115.9912 | 96.8957 | 191.4435 | 215.3392 | 206.4523 | 252.8957 | 113.2261 | 110.3304 | 115.5477 | 109.8957 | 70.4435 | 66.235 | 58.8693 | 95.3304 | 181.0088 | 170.6784 | 220.2173 | 167.2085 | 121.8693 | 121.8957 | 120.0954 | 134.2173 | 118.5213 | 154.4258 | 201.5301 | 167.5301 | 118.348 | 104.1219 | 145 | 138.5477 | 141.3392 | 146.5389 | 182.3392 | 248.7827 | 123.8693 | 145.5565 | 203.8869 | 197.1219 | 188.9912 | 174.6608 | 172.7562 | 395.3392 | 155.1043 | 133.765 | 138.5654 | 231.235 | 117.3392 | 119.3392 | 139.5477 | 131.4523 | 161.8268 | 164.5654 | 174.0088 | 136.3039 | 99.5477 | 119.1997 | 110.9046 | 105.1219 | 200.6696 | 200.1997 | 211.166 | 331.5742 | 208.9046 | 197.5477 | 203.6872 | 314.0265 | 166.2173 | 76.8869 | 75.8957 | 166.2173 | 99.8957 | 91.7915 | 110.4346 | 111.2173 | 92.2173 | 96.1043 | 101 | 113.5477 | 69.1043 | 76.6696 | 70.9134 | 105.9046 | 111.5477 | 94.5477 | 78.765 | 83.6784 | 184.6961 | 204.0954 | 215.8781 | 463.3392 | 141.8869 | 159.4346 | 165.4258 | 421.2438 | 274.1043 | 173.7474 | 180.6784 | 388.2775
1828 | 22.1024 | 143.6194 | 164.823 | 206.4824 | 170.7745 | 119.8982 | 153.6946 | 145.823 | 156.7212 | 129.5442 | 113.5442 | 111.3806 | 169.3673 | 188.6327 | 132.3806 | 142.1637 | 191.8849 | 144.4558 | 113.9734 | 137.8097 | 186.9115 | 126.9734 | 273.0885 | 303.4158 | 151.7212 | 192.0752 | 202.2388 | 207.5576 | 308.8716 | 105.0885 | 108.0885 | 87.823 | 131.9381 | 121.3806 | 110.7212 | 118.7212 | 87.5309 | 200.6594 | 213.823 | 217.9514 | 255.8363 | 113.2788 | 117.0133 | 121.3673 | 109.1903 | 71 | 59.4558 | 69.0885 | 98.3673 | 175.0885 | 170.8363 | 220.2655 | 166.7964 | 125.2522 | 116.177 | 117.6061 | 120.5928 | 143.0133 | 153.6061 | 211.8982 | 173.3406 | 115.9248 | 102.1018 | 134.9734 | 147.1018 | 137.5442 | 152.354 | 181.2788 | 247.823 | 137.7345 | 150.4691 | 201.354 | 193.7345 | 188.9867 | 181.646 | 185.4558 | 373.314 | 149.4291 | 143.2788 | 142.1151 | 222.6327 | 121.7478 | 115.0885 | 135.7964 | 135.2036 | 138.1018 | 155.354 | 169 | 140.2388 | 104.4558 | 108.7565 | 102 | 91.7079 | 188.9734 | 202.0619 | 183.646 | 331.9381 | 192.2522 | 197.4424 | 213.4205 | 303.1018 | 175.8363 | 79.3673 | 79.2921 | 175.8363 | 89.6194 | 93.5709 | 106.4291 | 113.5442 | 100.9248 | 100.4558 | 93.2522 | 121.646 | 72.0885 | 72.7212 | 62.0133 | 92.4424 | 110.5309 | 102.646 | 87.3673 | 77 | 192.6108 | 204.3406 | 202.8583 | 454.9867 | 147.5576 | 173.2036 | 179.646 | 403.9867 | 268.8849 | 184.6194 | 182.6594 | 415.8982
1829 | 22.1134 | 137.4041 | 174.4887 | 195.1296 | 177.4906 | 132.3177 | 138.532 | 149.8722 | 153.0639 | 135.5527 | 112.532 | 113.8515 | 162.5752 | 181.8515 | 134.8515 | 137.2763 | 197.8497 | 146.4041 | 111.829 | 130.2556 | 196.8929 | 135.126 | 271.0639 | 280.3834 | 154.6166 | 184.0639 | 187.468 | 214.4041 | 301.829 | 111.9568 | 100.9154 | 101.7012 | 146.4041 | 132.2763 | 109.4041 | 127.7012 | 79.5959 | 214.7444 | 197.7462 | 234.1917 | 266.8083 | 120.2124 | 126.2763 | 123 | 111.5959 | 66.7876 | 60.9361 | 61.4473 | 105.1485 | 173.532 | 179.9361 | 199.6823 | 154.6598 | 123.0207 | 115.4041 | 105.8722 | 118.3177 | 139.171 | 136.7237 | 198.5752 | 163.468 | 130.9568 | 115.8929 | 119.7237 | 142.171 | 134.1917 | 146.5959 | 184 | 252.8083 | 137.2556 | 139.2142 | 209.1692 | 192.3195 | 169.8947 | 182.7237 | 189.7444 | 367.3816 | 151.1053 | 146.468 | 153.468 | 208.3627 | 124.4473 | 123.3609 | 114.7669 | 145 | 153.297 | 146.3195 | 161.0432 | 123.1278 | 99.8515 | 86.6166 | 100.5959 | 94.2331 | 182.1485 | 189 | 207.658 | 352.9568 | 204.9982 | 199.0207 | 231.0432 | 304.7237 | 171.8308 | 74.4473 | 80.0432 | 171.8308 | 90.8929 | 108.2763 | 104.829 | 118.6166 | 110.3402 | 98.1917 | 87.7444 | 133.4887 | 71 | 68.5959 | 73.6166 | 85.5959 | 100.2556 | 100.9154 | 94.6166 | 74.1917 | 211.703 | 187.4473 | 178.1278 | 453.2124 | 147.3834 | 191.4248 | 182.5959 | 410.6373 | 265.4887 | 179.3402 | 177.0864 | 409.5959
1830 | 22.1244 | 145.7247 | 174.8343 | 182.4101 | 170.6068 | 148.6376 | 143.2837 | 144.3146 | 141.4326 | 142.0393 | 117.2837 | 118.5674 | 155.882 | 172.1573 | 126.5983 | 142.9214 | 218.441 | 152.8034 | 126.8427 | 120.7163 | 202.1966 | 146.9522 | 278.1657 | 279.7641 | 161 | 174.8343 | 190.882 | 207.3539 | 314.441 | 113.3146 | 94.9607 | 112.5983 | 142.7163 | 127.8343 | 116.2837 | 139.559 | 81.3624 | 215.1573 | 186.1264 | 223.7949 | 274.323 | 116.8343 | 118.4719 | 125.4017 | 111.9214 | 69.2051 | 60.0786 | 62.6461 | 97.5505 | 183.0871 | 168.5112 | 179.5197 | 164.4888 | 127.1573 | 121.323 | 115.2051 | 119.2668 | 137.1264 | 137.8034 | 187.559 | 164.9607 | 136.6376 | 127.441 | 132.8118 | 138.2051 | 140.6068 | 152.6854 | 192.6461 | 246.3932 | 137.323 | 118.0393 | 218.1966 | 185.559 | 172.7725 | 173.2359 | 192.0786 | 376.8736 | 155.4326 | 134.9916 | 147.2753 | 197.3624 | 118.441 | 135.441 | 115.4495 | 149.8034 | 160.1966 | 141.4803 | 159.6854 | 118.5983 | 93.5197 | 95.882 | 105.7247 | 117.0871 | 187.5197 | 195.2444 | 215.5899 | 342.7865 | 225.1966 | 193.0702 | 215.7556 | 304.8427 | 163.2837 | 67.9607 | 71 | 163.2837 | 110.6068 | 111.0393 | 117.441 | 120.677 | 97.6292 | 95.4803 | 95.3624 | 127.5899 | 69.0786 | 68.9214 | 76.1573 | 88.323 | 101.7641 | 95.441 | 101 | 71 | 212.9691 | 199.6938 | 172.6376 | 449.3539 | 158.2922 | 191.3932 | 180.0393 | 411.5505 | 277.8427 | 178.6376 | 176.7725 | 394.0702
1831 | 22.1354 | 143.6401 | 173.3894 | 188.6401 | 183.4631 | 148.4484 | 143.5811 | 141.4189 | 142.3157 | 132.1327 | 127.9263 | 121.6253 | 161.9558 | 188.1978 | 135.8231 | 145.4926 | 221.4779 | 163.9115 | 128.0442 | 123.8673 | 187.6401 | 134.5074 | 284.0442 | 291.4189 | 156.5663 | 179.301 | 190.5516 | 204.8968 | 301.2211 | 111.4042 | 101.9115 | 106.059 | 129.1179 | 123.9263 | 115.1032 | 136.0295 | 85.9853 | 214.4484 | 196.4926 | 210.5811 | 272.1032 | 118.8378 | 104.0295 | 123.5663 | 119.4189 | 79.4631 | 68.3452 | 74.9558 | 95.8083 | 192.5221 | 157.9558 | 197.2273 | 175.5369 | 121.0295 | 117.6253 | 123.4926 | 96.5369 | 148.9705 | 143.9853 | 186.9853 | 164.5221 | 129.5516 | 129.9853 | 151.4926 | 139.5958 | 141.1474 | 157.5516 | 199.5369 | 264.0798 | 150.3747 | 124.8821 | 194.7727 | 191.8821 | 183.6253 | 163.5516 | 187.0442 | 384.7346 | 142.5221 | 136.2862 | 142.4631 | 212.3305 | 125.9115 | 137 | 121.6253 | 145.6401 | 147.1179 | 142 | 159.6253 | 112.059 | 98.4189 | 96.5369 | 111.5221 | 125.5369 | 192.9115 | 195.1032 | 193.5958 | 346.1388 | 212.1179 | 190.3157 | 224.7641 | 308.0147 | 166.5369 | 69 | 69.0295 | 166.5369 | 101.2949 | 104.0885 | 109.1474 | 119.4484 | 90.3599 | 99.4484 | 102.941 | 111.5074 | 66.0147 | 70.5074 | 78.8968 | 92.5074 | 94.699 | 102.4189 | 102.4779 | 72.4779 | 226.0147 | 208.6843 | 194.1241 | 456.7494 | 162.2211 | 174.5958 | 172.5958 | 410.301 | 274.6106 | 178.941 | 197.9705 | 378.5074
1832 | 22.1464 | 141.5738 | 188.0689 | 182.5344 | 173.377 | 135.3376 | 145.5738 | 149.5246 | 148.941 | 128.5541 | 117.3475 | 121.0787 | 160.9606 | 202.941 | 141.436 | 149.0295 | 221.4852 | 178.0787 | 125 | 134.5148 | 181.0295 | 131.9803 | 273.9311 | 298.5148 | 148.9705 | 202.0886 | 187 | 206.4459 | 309.7312 | 119.0098 | 106.4852 | 104.0197 | 125.0394 | 128.4951 | 110.0197 | 119.8622 | 85.4852 | 218.5049 | 203.564 | 208.0295 | 242.7835 | 136.059 | 104.5246 | 117.4852 | 120.4557 | 90.0787 | 61.8327 | 71.941 | 111.0197 | 186.4557 | 165.0394 | 209.436 | 172.4951 | 114.4557 | 111.0197 | 118.4459 | 98.5443 | 152.5148 | 146.5148 | 182.9508 | 164.0098 | 118.9311 | 128.4754 | 140.8918 | 139.059 | 135.5443 | 149.4557 | 191.9508 | 268.7639 | 150.4164 | 138.0492 | 190.6427 | 203.5344 | 170.9606 | 174.1378 | 187.5344 | 385.8327 | 137.4656 | 140.3967 | 145.5049 | 229.0492 | 125.941 | 139.0197 | 121.0787 | 144.5836 | 150.1082 | 147.5541 | 157.564 | 114.059 | 103.4951 | 97.5344 | 96.3672 | 121.4852 | 188.9016 | 195.5738 | 194.5738 | 360.9016 | 202.9902 | 210.5738 | 225.8524 | 303.9705 | 170.059 | 67.9902 | 68.0098 | 170.059 | 91.6033 | 97.4951 | 107.0787 | 123.5049 | 105.0492 | 102.4951 | 104.4754 | 107.9705 | 60.9606 | 66.9705 | 74.3868 | 88.9705 | 83.0197 | 100.4262 | 96.4262 | 85.1082 | 232.0689 | 199.5344 | 213.941 | 465.9213 | 143.9705 | 171.0295 | 174.5836 | 397.7442 | 265.4852 | 185.0197 | 210.0984 | 397.1869
1833 | 22.1574 | 144.8623 | 185.173 | 184.9656 | 168.241 | 121.0688 | 159.2065 | 150.9656 | 138.3451 | 140.2065 | 103.5516 | 123.8279 | 151.8279 | 205.2754 | 138.1033 | 148.3795 | 213.2763 | 178.2418 | 124.4828 | 139.6205 | 183.4828 | 137.241 | 264.9312 | 306.7237 | 150.6549 | 185.6567 | 195.7926 | 217.0335 | 319.0353 | 114.3107 | 101.3795 | 108.0688 | 131.5861 | 119.7246 | 114.0688 | 125.654 | 89.1721 | 217.9656 | 190.8632 | 191.346 | 233.4131 | 117.6911 | 110.6205 | 122.7237 | 111.3451 | 88.173 | 56.8958 | 72.7237 | 109.8967 | 186.6549 | 166.9312 | 201.9656 | 173.0344 | 115.6893 | 107.3107 | 108.8623 | 106.6205 | 162.2754 | 134.5525 | 181.6205 | 163.9656 | 126.9991 | 128.0688 | 135.1721 | 141.8967 | 167.9294 | 151.2065 | 195.7926 | 243.9656 | 131.1386 | 136.7935 | 197.759 | 198.7246 | 179.9303 | 184.3795 | 176.5181 | 362.7935 | 148.4819 | 131.0344 | 139.7935 | 225.2074 | 113.7935 | 133.2418 | 126.9312 | 150.4139 | 147.5525 | 157.6549 | 163.4828 | 124.1377 | 96.7935 | 97.3795 | 94.8958 | 116.8967 | 177.9656 | 202.4828 | 226.8261 | 336.0009 | 199.9312 | 205.0697 | 214.6205 | 309.7926 | 176 | 75.2754 | 59.6902 | 176 | 102 | 101.6549 | 122.7582 | 114.6902 | 99.6558 | 93.7246 | 105.6205 | 107.5861 | 61.1377 | 67.6205 | 68.1721 | 82.3795 | 109.3089 | 82.6558 | 94.6893 | 85.6558 | 224.0009 | 198.8623 | 210.5861 | 495.7564 | 139.9656 | 175.0344 | 178.3451 | 400.4466 | 265.0344 | 189.5861 | 193.6222 | 411.3451
1834 | 22.1684 | 136.2345 | 169.646 | 199.3556 | 193.0031 | 127.236 | 157.0575 | 154.236 | 144.59 | 148.118 | 111.8835 | 122.177 | 148.059 | 197.1149 | 130.9395 | 140.2345 | 209.6475 | 171.5295 | 130.354 | 138.764 | 190.413 | 141.882 | 275.1785 | 297.1149 | 163.472 | 189.0046 | 200.823 | 227.2345 | 304.4115 | 123.2965 | 94.823 | 106.2935 | 132.4115 | 133.1801 | 120.7655 | 142.4115 | 89.2345 | 209.587 | 190.4736 | 210.5946 | 263.0031 | 125.7111 | 115.059 | 122.1165 | 111.7655 | 72.646 | 62.1755 | 74.2345 | 115.472 | 189.941 | 162.3525 | 188.8215 | 171.882 | 124.177 | 104.118 | 100.764 | 122.1785 | 165.2345 | 143.1801 | 181.2935 | 155.587 | 139.4115 | 129.4705 | 127.292 | 147.472 | 177.0559 | 164.413 | 190.233 | 252.0015 | 119.941 | 127.823 | 195.7655 | 197.8835 | 181.9395 | 184.177 | 163.5295 | 351.705 | 140.2904 | 134.118 | 135.5885 | 215.4115 | 109.059 | 126.5295 | 134.531 | 146.4115 | 142.9425 | 169.9425 | 175.708 | 133.295 | 98.9425 | 100.354 | 113.9425 | 124.0605 | 182.295 | 192.9985 | 229.8789 | 326.236 | 212.8261 | 194.059 | 215.3525 | 313.764 | 167.528 | 81.4115 | 66.8851 | 167.528 | 106.236 | 104.941 | 117.292 | 117.649 | 91.5885 | 87.059 | 99.469 | 106.823 | 71.8835 | 74.7065 | 68.2345 | 89.59 | 115.0559 | 85.708 | 93.1165 | 87.649 | 226.9441 | 189.1755 | 214.059 | 497.1119 | 149.0605 | 180.7655 | 175.5885 | 408.9969 | 284.0031 | 197.295 | 186.4736 | 429.7686
1835 | 22.1794 | 136.3343 | 167.2508 | 205.9567 | 200.3313 | 136.9597 | 159.2105 | 153.1239 | 167.5448 | 127.2447 | 112.5821 | 119.0403 | 150.0836 | 183.5418 | 124.709 | 134.9164 | 216.8761 | 141.6596 | 132.7492 | 138.7926 | 183.997 | 134.5821 | 306.0463 | 283.5418 | 160.7059 | 205.4149 | 189.3313 | 225.709 | 316.003 | 119.2044 | 110.421 | 104.0836 | 137.5015 | 128.1611 | 120.1239 | 131.7895 | 92.0433 | 216.003 | 188.6626 | 215.8268 | 269.7895 | 137.288 | 116.5418 | 116 | 112.2074 | 66.4582 | 59.709 | 76.5015 | 130.0433 | 190.0836 | 178.421 | 209.424 | 167.291 | 121.0403 | 111.4179 | 107.2941 | 119.997 | 154.4985 | 144.1208 | 180.1672 | 166.8792 | 136.9164 | 140.9194 | 131.1702 | 152.291 | 157.1239 | 172.0836 | 175.291 | 249.1641 | 132.003 | 129.8761 | 208.6687 | 195.8731 | 178.4179 | 163.1611 | 179.1702 | 365.9628 | 134.5448 | 138.709 | 139.1672 | 217.2508 | 120.8359 | 134.5851 | 139.7492 | 140.6657 | 152.1672 | 177.5418 | 172.3716 | 128.2477 | 106.5418 | 101.6657 | 113.9567 | 119.997 | 200.003 | 186.6254 | 210.9164 | 344.0866 | 221.1239 | 191.7492 | 215.709 | 308.3746 | 160 | 76.7492 | 71.7895 | 160 | 105.1239 | 104.5418 | 106 | 125.8328 | 101.6687 | 99.9194 | 96.4179 | 105.6254 | 69.8731 | 68.2477 | 67.7926 | 85.4552 | 109.2105 | 85.0806 | 82.6657 | 92.5821 | 234.4149 | 197.003 | 211.7492 | 462.1239 | 156.9164 | 178.4985 | 184.0433 | 384.8731 | 292.4149 | 188.997 | 193.3313 | 417.4924
1836 | 22.1904 | 141.6623 | 158.9183 | 199.4459 | 191.3377 | 144.7704 | 142.0662 | 161.1898 | 157.9448 | 122.9603 | 114.649 | 120.649 | 153.7704 | 183.4459 | 124.2296 | 131.2296 | 221 | 138.1634 | 127.2296 | 137.5673 | 174.6623 | 134.9868 | 318.0132 | 288.4327 | 144.7969 | 202.3245 | 189.2031 | 220.7969 | 323.1214 | 106.6623 | 119.351 | 110.5408 | 132.4724 | 115.7572 | 110.4592 | 127.8786 | 86.9183 | 218.1347 | 188.9603 | 208.6225 | 250.9183 | 131.2031 | 121.4327 | 124.8653 | 118.4194 | 62.1214 | 65.8786 | 70.9183 | 123.2561 | 184.9051 | 182.9183 | 213.2826 | 175.5276 | 118.7704 | 112.6755 | 108.2428 | 112.8786 | 154.5408 | 133.4327 | 195.8521 | 174.7969 | 137.1082 | 134.3775 | 139.0132 | 148.8918 | 146.9051 | 156.9316 | 173 | 262.7174 | 140.7837 | 142.8653 | 204.3642 | 192.8786 | 173.5806 | 160.7307 | 179.8102 | 384.7704 | 142.1214 | 129.3642 | 145.4327 | 211.6887 | 133.8786 | 132.1347 | 137.5541 | 149.7439 | 154.5541 | 171.9051 | 155.0132 | 120.5541 | 102.0132 | 84.702 | 109.6623 | 115.649 | 220.4194 | 200.7439 | 218.3113 | 346.5806 | 202.5938 | 202.8521 | 221.3245 | 323.0684 | 157.2296 | 75.1082 | 71.2031 | 157.2296 | 97.1214 | 103.3377 | 107.1082 | 114.5806 | 106.7837 | 106.6755 | 108.2031 | 111.9868 | 67.4327 | 71.6358 | 56.0397 | 81.2031 | 102.0397 | 79.9868 | 85.649 | 88 | 219.6887 | 197.4724 | 217.3113 | 460.2163 | 149.9051 | 176.3245 | 194.4327 | 410.691 | 281.5673 | 175.7837 | 179.351 | 427.1236
1837 | 22.2014 | 136.2035 | 169.2568 | 191.6371 | 175.8406 | 144.7345 | 140.124 | 174.3983 | 145.531 | 143.4956 | 109.8052 | 122.0354 | 161.6638 | 181.8672 | 119.1681 | 132.3983 | 222.6991 | 153.1681 | 128.9646 | 147.593 | 168.6371 | 154.8585 | 314 | 314.655 | 149.1948 | 205 | 198.9646 | 199.1415 | 313.2035 | 102.9026 | 106.6371 | 111.0354 | 122.3009 | 125.3983 | 111.6638 | 126.469 | 82.531 | 203.6371 | 197.0354 | 213.5044 | 252.7611 | 126.2389 | 115.3716 | 129.7345 | 123.7345 | 70.8939 | 76.3629 | 67.6638 | 119.7965 | 194.1594 | 189.8585 | 202.9293 | 174.938 | 121 | 115.0974 | 129.186 | 116 | 149.938 | 147.7611 | 189.4424 | 170.6991 | 135.1681 | 134.593 | 154.2568 | 152.531 | 136.3362 | 154.1948 | 187.1594 | 259.4777 | 141.8319 | 142.6371 | 208.0266 | 200.531 | 174.4956 | 177.8319 | 193.2214 | 377.938 | 132.7699 | 128.4956 | 151.8319 | 216.8939 | 138.1328 | 137.4603 | 136.8672 | 144.7079 | 156.1328 | 169.2655 | 155.531 | 126.0974 | 101.3983 | 92.124 | 103.0707 | 113.6371 | 220.6371 | 204.7699 | 202.345 | 333.3362 | 236.8764 | 198.1415 | 217.2035 | 334.8672 | 160.6638 | 70.3362 | 70.7699 | 160.6638 | 99.0974 | 102.5664 | 119.3275 | 109.2655 | 111.7965 | 106.2655 | 97.0087 | 113.7345 | 68.1681 | 68.5397 | 54.7611 | 91.531 | 93.9646 | 82.8672 | 96.0974 | 77.8052 | 204.5044 | 201.4603 | 209.8406 | 465.9646 | 161.4249 | 196.5576 | 186.6725 | 413.7786 | 288.1948 | 188.7258 | 176.2655 | 433.2122
1838 | 22.2124 | 147.2026 | 176.4773 | 198.152 | 186.4106 | 148.208 | 151.1067 | 162.5333 | 149 | 171.4106 | 114.152 | 117.8427 | 169.472 | 179.8427 | 126.2587 | 132.1067 | 221.1067 | 170.6746 | 128.528 | 149.32 | 181.5173 | 160.056 | 308.2133 | 302.488 | 162.7867 | 195.7413 | 180.0107 | 198.5733 | 303.3707 | 96.1067 | 109.68 | 106.264 | 135.4667 | 135.5227 | 126.9947 | 129.3653 | 87.736 | 206.1013 | 202.68 | 230.7253 | 250.584 | 119.736 | 113.7867 | 119.8987 | 120.8427 | 79.4213 | 72.1627 | 74.3147 | 123.8427 | 196.32 | 193.32 | 202.6347 | 157.584 | 130.8373 | 127.1013 | 126.8534 | 119.472 | 154.5733 | 151.9493 | 190.4667 | 152.904 | 133.5787 | 133.4267 | 169 | 150.792 | 145.3093 | 160.8427 | 192.792 | 253.9947 | 139.3707 | 147.416 | 216.264 | 201.6853 | 173.4773 | 188.1013 | 224.888 | 362.8987 | 143.624 | 140.208 | 156.8933 | 217.32 | 140.736 | 152.208 | 129.056 | 139.3653 | 160.472 | 165.2133 | 167.1013 | 120.1627 | 90.6907 | 113.5227 | 95.8427 | 111.472 | 211.528 | 202.3147 | 199.4667 | 345.2026 | 252.904 | 195.8373 | 219.5227 | 326.4773 | 171.944 | 72.944 | 66 | 171.944 | 105.8933 | 116.888 | 132.6293 | 104.056 | 110.6347 | 97.0053 | 80.528 | 112.5787 | 74.1013 | 60.736 | 74.5733 | 95.5787 | 95.8427 | 81.4213 | 99.4213 | 69.4213 | 203.7867 | 205.792 | 199.5787 | 474.7867 | 173.4213 | 182.8027 | 188.416 | 404.1013 | 289.6347 | 184.9547 | 187.8373 | 438.1466
1839 | 22.2234 | 145.408 | 169.8181 | 204.0452 | 194.9085 | 144.9085 | 149.591 | 153.1819 | 169.0925 | 156.2702 | 123.591 | 125.8644 | 173.1819 | 179.591 | 146.0011 | 150.0925 | 221.3638 | 162.5436 | 132.5005 | 140.0452 | 196.7729 | 157.9548 | 289.2261 | 299.9106 | 156.9537 | 180.7266 | 181.1378 | 199.3176 | 310.0463 | 108.7739 | 97.6803 | 107.3638 | 148.3638 | 123.2713 | 130.8633 | 142.2734 | 91.9548 | 213.7729 | 194.817 | 239.1356 | 271.3659 | 115.6814 | 110.3176 | 112.2271 | 117.6362 | 80.1819 | 63.2271 | 77.7729 | 117.6814 | 197.6824 | 172.817 | 204.5005 | 163.0011 | 132.6814 | 122.3628 | 118.8644 | 119.6362 | 156.4995 | 141.9085 | 208.6824 | 153.7739 | 132.8181 | 127.3638 | 170.1819 | 142.8633 | 152.0452 | 181.8654 | 193.7277 | 262.591 | 132.4543 | 144.3628 | 210.2723 | 205.9096 | 178.6372 | 182.7718 | 201.9064 | 364.0915 | 153.8181 | 146.9548 | 143.0442 | 210.409 | 139.6362 | 144.7718 | 135.8191 | 138.0904 | 161.8181 | 167.5005 | 175.3638 | 131.3197 | 96.3649 | 107.1809 | 89.6814 | 121.0915 | 212.5457 | 179.7708 | 208.2271 | 353.4543 | 234.2723 | 188.817 | 221.4543 | 332.2282 | 165.7718 | 70.3176 | 63.0452 | 165.7718 | 101.4995 | 113.9989 | 127.7266 | 109.0463 | 98.3176 | 94.3186 | 93.9558 | 114.7729 | 80 | 67.3186 | 72.3628 | 105.4553 | 97.3638 | 85.7277 | 94.2723 | 74.3186 | 200.9085 | 198.4543 | 210.6372 | 468.9537 | 161.1809 | 166.5457 | 180.6351 | 416.5005 | 315.1388 | 189.9558 | 197.3638 | 444.9537
1840 | 22.2344 | 139.6195 | 160.5545 | 198.9838 | 196.0325 | 133.9675 | 154.826 | 164.8585 | 154.0441 | 139.2227 | 119.174 | 114.5058 | 159.522 | 178.7935 | 152.587 | 152.5383 | 219.3805 | 160.2715 | 137 | 153.0812 | 196.1903 | 143.7123 | 312.1786 | 313.6032 | 161.4617 | 184.652 | 197.826 | 207.2715 | 305.5383 | 110.5545 | 106.1137 | 112.0162 | 143.9675 | 117.6357 | 126.1903 | 140.1578 | 94 | 206.5545 | 188.0162 | 224.5545 | 271.0928 | 132.5104 | 111.6357 | 111 | 117.8097 | 84.0162 | 66.826 | 69.9513 | 113.3968 | 196.3643 | 166.0162 | 208.3968 | 172 | 118.1415 | 119.826 | 111.7285 | 130.6682 | 166.478 | 156.9072 | 196.5058 | 165.8097 | 133.2065 | 126.587 | 164.9675 | 153.2715 | 140.348 | 178.9026 | 192.174 | 278.0812 | 123.3643 | 132.7773 | 203.9838 | 187.0766 | 178.761 | 197.9234 | 182.0162 | 355.7285 | 144.5545 | 142.9675 | 148.2877 | 219.652 | 128.9513 | 127.9513 | 130.1253 | 146.6682 | 161.6032 | 177.4292 | 198.7169 | 125.9026 | 94.3318 | 95.3805 | 100.478 | 134.4455 | 206.5545 | 193.1624 | 205.7935 | 331.0928 | 238.239 | 189.2552 | 222.0162 | 324.9188 | 161.0162 | 67.413 | 69.4455 | 161.0162 | 103.0325 | 108.0162 | 116.5708 | 130.478 | 89.3805 | 98.6032 | 98.3643 | 116 | 77.587 | 72.2065 | 74.0487 | 112 | 100.2065 | 96.239 | 95.826 | 82.826 | 202.0325 | 203.8422 | 213.174 | 483.7169 | 143.348 | 172.6195 | 169.3968 | 411.9513 | 340.2227 | 190.1415 | 194.7773 | 447.652
1841 | 22.2454 | 148.7709 | 172.2351 | 197.6155 | 185.2271 | 144.1574 | 145.6892 | 153.5338 | 147.3108 | 140.1534 | 109.8446 | 110.3865 | 159.2331 | 192.1574 | 157.7709 | 143.7689 | 211.4601 | 171.4622 | 122.2271 | 144.5338 | 201.1554 | 144.0797 | 306.9163 | 301.6892 | 168.3845 | 175.6115 | 189.3048 | 228.3108 | 301.6932 | 108.6932 | 116.3068 | 115.2311 | 138.7689 | 122.6155 | 127.4622 | 144.8486 | 84.1514 | 189.9203 | 186.3068 | 208.5358 | 258.6155 | 126.3027 | 127.6952 | 112.8466 | 117.7689 | 77.3825 | 63.2291 | 65.2311 | 104.3825 | 190.7689 | 179.6952 | 207.3845 | 167.0757 | 108.5378 | 107.6115 | 107.9243 | 127.9203 | 150.1474 | 152.7649 | 180.0757 | 165.1534 | 116.1494 | 126.2311 | 158.5378 | 150.9203 | 153.0817 | 162.0757 | 185.3068 | 255.3764 | 141.1594 | 134.3088 | 202 | 189.2351 | 206.008 | 189.6095 | 186.4622 | 360.5419 | 142.6932 | 144.5399 | 151.6135 | 219.2291 | 117.4601 | 124.4622 | 137.6196 | 155.6155 | 163.2311 | 168.6892 | 190.225 | 128.7729 | 105.0817 | 89.9223 | 109.3845 | 136.9223 | 209.002 | 218.5399 | 203.7689 | 362.3188 | 226.3804 | 221.2371 | 230.1554 | 306.998 | 165.4622 | 66.5378 | 74.3845 | 165.4622 | 110.6932 | 111.8466 | 109.3068 | 124.6115 | 101.1574 | 97.1534 | 106.9263 | 122.7709 | 74.7689 | 76.6932 | 70.1514 | 94.7649 | 109.6175 | 101 | 102.6932 | 91.5399 | 194.9203 | 209 | 221.6952 | 468.4541 | 145.6175 | 178.0777 | 176.3865 | 402.3068 | 317.7629 | 203.9283 | 196.3088 | 482.9303
1842 | 22.2564 | 161.7896 | 174.2104 | 198.6278 | 180.3948 | 134.1651 | 140.5113 | 150.1618 | 166.9288 | 128.9547 | 99.0939 | 126.3009 | 168.1392 | 167.725 | 150.6991 | 144.8835 | 199.8382 | 151.6538 | 133.7184 | 145.5566 | 191.1877 | 135.9321 | 313.8576 | 278.9321 | 152.3042 | 191.1132 | 180.1165 | 230.9774 | 290.1877 | 108.4887 | 110.8608 | 119.767 | 126.0712 | 109.8156 | 125.233 | 125.259 | 94.3236 | 198.0679 | 195.3009 | 211.4175 | 267.7896 | 119.6505 | 124.9547 | 129.6958 | 114.4887 | 74.5113 | 61.5113 | 73.534 | 104.6505 | 190.6278 | 173.1877 | 195.6991 | 172.1618 | 130.8576 | 119.3462 | 117.2783 | 137.3236 | 163.508 | 153.3009 | 178.2557 | 169.6505 | 142.042 | 132.0226 | 170.1844 | 154.6731 | 157.466 | 160.8835 | 192.4175 | 250.0679 | 151.8608 | 145.1618 | 193.8382 | 189.3269 | 195.8641 | 168.0939 | 176.0712 | 380.9288 | 160.0679 | 136.6991 | 141.9774 | 236.3462 | 124.0453 | 117.2104 | 146.1165 | 147.2104 | 184.7184 | 177.9514 | 174.1165 | 135.4887 | 110.0939 | 103.0679 | 111.5113 | 126.8382 | 210.233 | 203.1651 | 189.1877 | 366.1424 | 246.1359 | 215.6538 | 215.1651 | 298.233 | 162.6052 | 78.8123 | 78.3948 | 162.6052 | 110.4887 | 106.7217 | 112.6505 | 106.2104 | 91.7929 | 109.8123 | 101.8156 | 130.1392 | 76.5113 | 75.233 | 70.9061 | 105.3462 | 99.9321 | 95.9774 | 97.466 | 81.1877 | 208.0906 | 199.5825 | 207.6538 | 481.3915 | 151.6278 | 208.2524 | 187.2783 | 422.6019 | 347.8316 | 188.1198 | 213.4401 | 498.4887
1843 | 22.2674 | 143.9556 | 172.8407 | 193.879 | 190.0413 | 132.6018 | 138.7994 | 149.2389 | 166.3185 | 117.879 | 109.7228 | 118.9173 | 168.0796 | 183.8468 | 115.8347 | 147.9204 | 217.4042 | 163.3246 | 147.2802 | 138.2772 | 186.8407 | 136.6018 | 299.8347 | 279.6018 | 167.3246 | 184.2359 | 205.8851 | 223.5192 | 304.4042 | 99.9587 | 114.121 | 121.3599 | 131.8024 | 116.0827 | 134.5222 | 139.4456 | 102.0796 | 196.7581 | 219.2833 | 207.3982 | 251.2359 | 117.2389 | 111.9587 | 128.7581 | 107.2389 | 83.6815 | 61.0796 | 77.3599 | 113.7611 | 183.3185 | 177.8024 | 217.8055 | 186.6018 | 128.3568 | 114.0766 | 117.1593 | 125.876 | 170.1179 | 156.1593 | 191.1623 | 158.9173 | 139.0353 | 140.121 | 161.3568 | 157.1593 | 160.6815 | 155.5988 | 210.8024 | 235.3155 | 143.5988 | 157.6815 | 197.9617 | 193.2419 | 169.6784 | 168.4808 | 168.3599 | 389.2802 | 151.7167 | 131.2802 | 143.4808 | 253.4808 | 135.121 | 128.6432 | 134.7581 | 148.4012 | 190.5988 | 191.2006 | 164.0383 | 148.7228 | 104.7198 | 105.5988 | 96.3568 | 118.7994 | 221.4425 | 209.9234 | 189.3216 | 337.2772 | 264 | 215.1623 | 208.4808 | 306.2419 | 152.3185 | 81.2389 | 88.0413 | 152.3185 | 119.2419 | 97.879 | 112.1593 | 98.4395 | 89.3216 | 102.6371 | 106.1623 | 114.0766 | 91.4425 | 74.2802 | 68.5988 | 98.1562 | 92.9204 | 96.8407 | 97.4808 | 81.9617 | 213.5988 | 205.5222 | 227.6462 | 453.2711 | 159.6815 | 210.2772 | 201.8821 | 432.1593 | 354.5161 | 178.0413 | 199.5958 | 467.5544
1844 | 22.2784 | 148.6151 | 173.0428 | 204.0482 | 184.2138 | 145.1765 | 154.6151 | 151.8717 | 152.8662 | 135.8772 | 107.5614 | 114.9145 | 162.4331 | 186.3849 | 128.0537 | 138.5614 | 214.9945 | 160.6897 | 150.6096 | 140.4386 | 185.7379 | 138.7379 | 285.3048 | 257.5987 | 156.8607 | 192.8772 | 217.0855 | 219.6952 | 315.6952 | 93.3904 | 98.7325 | 109.909 | 140.9572 | 143.5724 | 126.6469 | 146.2138 | 108.8289 | 193.6096 | 237.4814 | 208.5241 | 265.7489 | 125.091 | 100.8235 | 130.1765 | 108.5669 | 75.6042 | 67.1765 | 69.8235 | 107.2138 | 192.7007 | 186.9572 | 217.0373 | 180.909 | 113.7807 | 115.091 | 113.6952 | 114 | 162.0428 | 159.2193 | 190.8235 | 166.0055 | 136.091 | 127.3421 | 158.5241 | 145.2138 | 154.5614 | 162.4386 | 231.0482 | 246.1393 | 134.1283 | 143.7325 | 198.4331 | 203.5669 | 155.2138 | 181.4386 | 171.2621 | 372.3849 | 153.7434 | 125.4759 | 155.7862 | 246.8662 | 122.3421 | 134.7379 | 136.8289 | 152.6524 | 181.1283 | 201.4814 | 196.1875 | 143.2993 | 102.0428 | 111.7862 | 103.9627 | 126.1338 | 239.4386 | 221.3048 | 209.0055 | 338.1338 | 245.7325 | 198.5132 | 212.3048 | 305.4759 | 155.1765 | 76.0428 | 69.1656 | 155.1765 | 119.1283 | 89.7807 | 110.6524 | 105.4814 | 105.091 | 106.3958 | 107.1283 | 116.3958 | 76.818 | 76.9572 | 66.9572 | 102.6579 | 95.3048 | 102.9145 | 98.0428 | 88.3048 | 213.9145 | 200.2566 | 212.0746 | 470.0592 | 163.3476 | 207.2193 | 206.6952 | 434.5669 | 356.7007 | 194.3958 | 184.3904 | 478.4013
1845 | 22.2894 | 146.0353 | 172 | 205.0177 | 181.6588 | 135.0412 | 141.4001 | 163.6411 | 154.647 | 147 | 104.6588 | 109.6883 | 152.0236 | 183.647 | 150.9764 | 142.9706 | 202.3471 | 141.3647 | 138.0412 | 132.7059 | 179.3471 | 144.9764 | 280.6824 | 295.5116 | 149.9823 | 213.9706 | 201.0412 | 211.6883 | 336.2705 | 115.2646 | 106.2823 | 113.3058 | 136.0177 | 140.7059 | 121.6588 | 138.3412 | 98.3765 | 209.6235 | 227.3765 | 229.947 | 269.3647 | 115.7118 | 103.647 | 120.7059 | 122.9647 | 69 | 59.7 | 72.647 | 107.9823 | 180.7236 | 184.0118 | 218.6353 | 147.0825 | 130.2764 | 123.6588 | 115.6588 | 123.9706 | 157.6765 | 153.3589 | 176.3647 | 153.3942 | 155.9588 | 126.9764 | 147.3765 | 160.6058 | 141.0236 | 185.947 | 220.7177 | 262.9882 | 142.9647 | 126.0236 | 212.6176 | 204.0059 | 169.2764 | 179.6883 | 187.6235 | 380.2823 | 138.0648 | 139.2823 | 159.6706 | 266.594 | 132.6117 | 148.2882 | 129.0353 | 154.3294 | 191.2941 | 200.0177 | 191.406 | 145.3058 | 102.3294 | 113.6765 | 107.6824 | 121.0294 | 226.3883 | 232.6353 | 220.3235 | 362.9411 | 278.541 | 206.6058 | 218.3176 | 328.5881 | 160.9941 | 71.6765 | 74.947 | 160.9941 | 108.0236 | 103.6176 | 112.9941 | 121.9647 | 103.0236 | 108.3471 | 97.353 | 115.0236 | 69.6529 | 70.6883 | 61.353 | 100.3647 | 120.594 | 101.6765 | 101.6529 | 89 | 217.3294 | 203.9706 | 208.6235 | 459.1001 | 176.9588 | 206.0118 | 203.3412 | 446.9706 | 374.6353 | 232.2411 | 192.9706 | 496.9882
1846 | 22.3004 | 147.4469 | 171.323 | 198.615 | 180.292 | 117.8451 | 131.354 | 161.907 | 169.5089 | 149.031 | 110.062 | 108.031 | 154.7699 | 185.646 | 134.6902 | 156.1239 | 204.739 | 163.0798 | 145.8939 | 129.385 | 194.6019 | 142.2301 | 292.8939 | 313.8761 | 158.416 | 202.7521 | 200.093 | 228.3098 | 333.4601 | 115.4911 | 102.1371 | 109.1991 | 136.385 | 136.031 | 116.907 | 126.8451 | 96.416 | 206.1681 | 225.416 | 212.5972 | 249.1061 | 131.6948 | 100.907 | 134.9868 | 109.3672 | 73.062 | 67.5399 | 72.615 | 100.8451 | 192.6638 | 192.8319 | 229.416 | 166.1727 | 117.6592 | 131.7699 | 107.5221 | 122.2301 | 153.969 | 145.615 | 197.4028 | 163.3098 | 148.1061 | 141.8319 | 161.6638 | 144.5972 | 148.5089 | 176.0442 | 217.416 | 253.4911 | 149 | 136.8939 | 204.7521 | 198.261 | 176.969 | 203.7568 | 188.2301 | 390.354 | 140.5399 | 135.8141 | 165.093 | 263.4291 | 144.031 | 141.7831 | 134.5089 | 146.1991 | 199.354 | 193.615 | 184.416 | 139.1681 | 96.2301 | 107.938 | 99.584 | 120.062 | 214.292 | 224.4601 | 207.7831 | 381.1239 | 274.2742 | 219.031 | 214.907 | 335.2301 | 169.4469 | 70 | 75.8761 | 169.4469 | 118.2169 | 102.5221 | 109.938 | 130.708 | 106.4469 | 103.292 | 107.5399 | 129.2789 | 71.323 | 63.615 | 73.5709 | 92.969 | 121.4911 | 92.5531 | 91.1371 | 81.5531 | 220.708 | 217.8009 | 210.584 | 475.1727 | 160.3052 | 224.9868 | 215.5399 | 432.3672 | 391.5089 | 228.3672 | 208.1549 | 528.7878
1847 | 22.3114 | 158.5821 | 188.2321 | 206.65 | 179 | 124.7179 | 140.9607 | 136.2536 | 166.0393 | 140.35 | 123.0286 | 122.0964 | 158 | 167.7679 | 131.2036 | 155.8643 | 202.8643 | 157.3893 | 145.4179 | 139.9607 | 178.875 | 137.6214 | 263.4679 | 318.2714 | 149.2821 | 199.825 | 193.35 | 240.7571 | 325.6214 | 123.0964 | 99.4464 | 97.4179 | 128.35 | 128.0393 | 139.5036 | 118.5536 | 94.8643 | 211.3393 | 221.7964 | 226.1929 | 256.475 | 134.0393 | 109.7179 | 141.5536 | 113.7857 | 75.6893 | 73.3107 | 71.6893 | 103.5821 | 212.65 | 189.7286 | 218.9036 | 163.3893 | 117.3393 | 130.175 | 122.9893 | 121.7571 | 158.5143 | 152.9607 | 184.4964 | 173 | 139.6214 | 140.7964 | 155.4571 | 165.775 | 147.7964 | 173.2036 | 193.1179 | 255.5821 | 137.2821 | 144.6893 | 203.8929 | 208.4071 | 189.0964 | 201.8357 | 167.0786 | 378.5929 | 139.4179 | 141.0286 | 154.2143 | 227.05 | 130.525 | 139.825 | 144.1357 | 138.5536 | 208.9607 | 183.0393 | 173.2143 | 152.6107 | 93.6893 | 116.3393 | 95.6214 | 127.5143 | 213.6893 | 211.7964 | 207.8929 | 373.9714 | 232.3607 | 227.5821 | 226.475 | 323.7286 | 170.9321 | 73.4464 | 66.4857 | 170.9321 | 127.0679 | 95.9321 | 114.2036 | 138.8929 | 112.7571 | 99.9321 | 114 | 129.0393 | 71 | 68.2036 | 79.6214 | 115.4357 | 109.7964 | 98.65 | 98.7857 | 72.4857 | 254.3964 | 224.0679 | 227.9893 | 489.6214 | 147.6214 | 206.05 | 223.3786 | 454.7071 | 410.0964 | 199.7 | 203.35 | 547.1357
1848 | 22.3224 | 145.8637 | 174.9527 | 194.8637 | 190.2252 | 121.5811 | 152.0158 | 139.33 | 158.4921 | 154.241 | 124.4921 | 113.2669 | 170.6284 | 171.9268 | 150.1363 | 146.9842 | 217.1363 | 152.5079 | 132.8795 | 147.5079 | 189.7489 | 143.3142 | 297.5135 | 338.1363 | 162.241 | 183.0574 | 184.089 | 229.3716 | 348.152 | 116.3716 | 110.8221 | 101.7173 | 138.0315 | 133.8221 | 143.9842 | 113.4921 | 90.8953 | 206.8795 | 217.5968 | 220.4606 | 248.2669 | 132.1047 | 115.7016 | 130.8795 | 134.7331 | 73.1937 | 71.5968 | 69.8953 | 120.33 | 221.911 | 192.3142 | 211.5968 | 173.241 | 123.4032 | 131.5079 | 120.7748 | 119.4921 | 172.9268 | 143.67 | 203.1678 | 182.8221 | 148.8221 | 145.7173 | 157.8221 | 159.9527 | 142.8953 | 174.5968 | 197.8378 | 281.4505 | 152.3457 | 133.7748 | 207.7016 | 246.2725 | 201.3142 | 188.6858 | 172.33 | 398.9583 | 140.2095 | 159.33 | 160.6284 | 238.5552 | 145.0473 | 143.4032 | 138.9842 | 152.4347 | 210.8953 | 202.8536 | 171.2095 | 155.3874 | 103.1205 | 110.4764 | 102.0158 | 115.2669 | 217.5079 | 216.7173 | 215.2095 | 379.5236 | 275.8277 | 220.4764 | 223.1779 | 349.4662 | 172.1047 | 60.9685 | 84.3457 | 172.1047 | 126.5968 | 101.3142 | 109.2827 | 137.089 | 99.2669 | 107.4189 | 110.4921 | 136.9268 | 67.4921 | 75.911 | 74.089 | 111.2669 | 101.3874 | 98.089 | 114.8221 | 84.7331 | 278.1205 | 210.2669 | 234.8953 | 451.8164 | 184.1836 | 220.3615 | 224 | 480.9268 | 425.1205 | 211.3457 | 221.4505 | 555.3142
1849 | 22.3334 | 134.7168 | 168.8555 | 205.1328 | 198.5693 | 122.9971 | 139.2949 | 159.2774 | 159.8555 | 157.7168 | 123.7139 | 116.2803 | 171.7168 | 183.4248 | 150.5752 | 148.2832 | 224 | 167.5635 | 134.7109 | 152.5693 | 202.5693 | 140.2891 | 300.4394 | 339.2891 | 165.7168 | 180.7109 | 174.1475 | 198.3008 | 333.0146 | 108.1445 | 111.4307 | 114.9941 | 136.1475 | 150.1358 | 137.4307 | 125.5635 | 83.5752 | 201.5723 | 207.7197 | 247.2656 | 257.7051 | 117.2949 | 108.8613 | 115.5781 | 121.7256 | 70.5723 | 78.8525 | 73.9971 | 103.1562 | 216.8613 | 196.4277 | 246.6934 | 178.8584 | 130.4248 | 109.4424 | 116 | 123.7109 | 161.581 | 140.8555 | 193.8701 | 182.0029 | 150.1445 | 167.5606 | 149.1504 | 151.7139 | 146.2832 | 166.1475 | 193.5781 | 256.0205 | 168.8525 | 127.5723 | 222.2774 | 229.3037 | 203.2861 | 185.2861 | 191.5635 | 390.7256 | 157.7051 | 149.2949 | 143.1562 | 261.1358 | 144.7197 | 144.7139 | 155.9883 | 155.4307 | 219.9941 | 193.7256 | 185.1358 | 169.419 | 104.8584 | 110.9971 | 106.4277 | 139.6963 | 228.2803 | 236.419 | 212.7168 | 370.4365 | 286.4365 | 225.9941 | 231.8496 | 334.1592 | 195.8438 | 65.708 | 91.5723 | 195.8438 | 131.7109 | 108.9971 | 129.5576 | 147.1358 | 102.2803 | 104.5752 | 106.8584 | 134.1475 | 73.1387 | 68.0059 | 64.1475 | 112.1387 | 103.2832 | 107.4219 | 104.0088 | 85.2891 | 278.4307 | 204.7139 | 217.581 | 462.4131 | 170.0176 | 229.1445 | 201.1562 | 468.1533 | 448.2744 | 222.8555 | 218.8642 | 600.1182
1850 | 22.3444 | 152.6063 | 179.4401 | 193.1199 | 189.1076 | 136.6185 | 145.3447 | 171.5354 | 171.1662 | 154.5477 | 117.4646 | 133.797 | 157.6553 | 189.6308 | 133.4768 | 155.8093 | 247.9632 | 175.1785 | 152.9755 | 145.2861 | 208.357 | 153.2493 | 305.3447 | 321.0245 | 170.5354 | 201.154 | 194.2371 | 207.6063 | 306.2983 | 101.9169 | 107.0954 | 99.3937 | 159.1417 | 148.4646 | 125.1076 | 134.6308 | 91.8924 | 190.1076 | 217.797 | 225.4182 | 261.0954 | 134.2371 | 120.5232 | 129.154 | 129.9755 | 62.7384 | 75.4646 | 76 | 105.6185 | 229.9755 | 203.5354 | 228.3229 | 198.3324 | 138.0831 | 122.5109 | 121.0831 | 116.5599 | 162.2616 | 152.8924 | 200.2493 | 190.1662 | 145.3692 | 158.2983 | 169.4156 | 140.3815 | 134.9291 | 170.9877 | 202.797 | 278.3079 | 156.7507 | 141.5232 | 220.0123 | 246.0463 | 190.6553 | 177.0123 | 171.5844 | 382.2738 | 169.8093 | 151.7139 | 152.154 | 252.9291 | 156.2493 | 164.6063 | 135.6798 | 171.4278 | 223.2738 | 210.6894 | 180.5599 | 160.0245 | 104.7262 | 104.2861 | 104.8215 | 122.9537 | 232.7262 | 227.7507 | 229.154 | 400.5818 | 271.5599 | 228.5477 | 201.4182 | 345.5109 | 178.1321 | 75.8093 | 75.0245 | 178.1321 | 139.0831 | 118.9877 | 117.9414 | 134.5722 | 108.1785 | 92.5599 | 99.4646 | 125.1907 | 70.9169 | 71.2616 | 69.7139 | 109.1907 | 121.7017 | 104.7384 | 91.4646 | 72.8338 | 253.0368 | 204.2738 | 216.0831 | 491.8801 | 214.6404 | 206.2152 | 202.8924 | 458.0954 | 498.1172 | 226.9046 | 212.5477 | 606.8338
1851 | 22.3554 | 148.9233 | 180.0462 | 221.4456 | 180.5693 | 145.4307 | 152.9538 | 173.2616 | 171.3078 | 148.0924 | 117.9538 | 131.6155 | 167.0305 | 212.415 | 133.1691 | 146.1848 | 248.8771 | 174.5231 | 150.1386 | 141.2616 | 219.5998 | 156.0462 | 316.646 | 302.4464 | 164.8771 | 211.6922 | 206.6922 | 232.7228 | 321.415 | 124.3688 | 107.4769 | 107.5074 | 129.8622 | 149.6922 | 112.8771 | 133.0462 | 101.1691 | 204.4612 | 222.2616 | 211.2616 | 265.1691 | 143 | 111.9695 | 133.7847 | 146.3383 | 71.8152 | 69.3078 | 64.9233 | 91.5388 | 240.4307 | 203.7847 | 227.8152 | 192.7079 | 128.9233 | 123.6155 | 126.6922 | 110.0462 | 157.6155 | 150.354 | 198.6155 | 193.2616 | 149.1691 | 142.4002 | 167.9233 | 171.4456 | 158.0611 | 201.3226 | 210.9538 | 283.8771 | 140.6617 | 145.5231 | 210.354 | 212.693 | 197.8152 | 181.3845 | 173.0767 | 399.7228 | 160.9233 | 152.0462 | 149.4002 | 264.9843 | 159.7847 | 143.2005 | 139.0305 | 164.7079 | 214.8771 | 213.354 | 177.7384 | 143.6617 | 119.769 | 100.2616 | 89.9695 | 122.3383 | 212.3234 | 234.5536 | 220.4926 | 381.5082 | 317.4761 | 218.4002 | 230.0916 | 360.646 | 193.8457 | 70.6155 | 67.5231 | 193.8457 | 126.9695 | 107.231 | 116.646 | 142.0305 | 108.2616 | 103.0305 | 95.5231 | 140.7228 | 74.1691 | 66.6155 | 70.7847 | 115.8614 | 115.4464 | 107.9076 | 93.4307 | 88.9381 | 227.0157 | 211.3845 | 214.3078 | 468.7236 | 208.6774 | 213.5074 | 220.2921 | 446.6617 | 507.354 | 242.769 | 212 | 600.0462
1852 | 22.3664 | 147.2522 | 211.2819 | 207.7241 | 212.7833 | 145.4985 | 164.5104 | 159.4867 | 170 | 158.7626 | 129.5104 | 114.7359 | 181.7596 | 183.2137 | 138.7537 | 135.9941 | 225.73 | 162.7389 | 140.9941 | 131.2404 | 177.9556 | 151.997 | 304.7359 | 289.7419 | 155.9941 | 210.7478 | 214.7567 | 241.2522 | 333.5044 | 107.4748 | 117.0089 | 112.2493 | 139.273 | 134.4837 | 119.7596 | 135.003 | 87.9852 | 203.4926 | 200.9793 | 218.5074 | 282.7656 | 128.7359 | 137.0296 | 125.4926 | 143.9941 | 76 | 69.5015 | 68.5074 | 94.0089 | 210.4689 | 222.5193 | 228.997 | 206.7685 | 128.7537 | 124.7537 | 127.2493 | 106.7478 | 164.7596 | 145.7478 | 202.0059 | 180.9881 | 181.7804 | 142.003 | 171.5074 | 190.0059 | 164.997 | 200.4896 | 193.9822 | 263.733 | 158.7715 | 150.2552 | 204.2463 | 212.7656 | 202 | 196.7626 | 173.2463 | 402.997 | 153.997 | 166.7656 | 144.4985 | 240.2196 | 137.9793 | 145.0119 | 136.4926 | 175.7656 | 232.27 | 210.2493 | 181.7537 | 151.2611 | 97.2226 | 100 | 97.7626 | 140.2641 | 254.5489 | 235.2463 | 221.0059 | 394.0237 | 339.5044 | 236.7715 | 220.9763 | 351.7389 | 198.4956 | 74.7567 | 66.2493 | 198.4956 | 116.7448 | 110.2581 | 133.2641 | 134.9881 | 111.7537 | 145.537 | 117.5222 | 138.7419 | 79.003 | 70.7567 | 74.5044 | 112.2433 | 122.2611 | 110 | 108.5133 | 96.7507 | 247.2759 | 205.7419 | 224.2611 | 497.7893 | 205.0059 | 255.0356 | 230.2552 | 446.7537 | 530.5252 | 263.7656 | 220.2581 | 618.5193
1853 | 22.3774 | 153.3412 | 200.6351 | 216.5497 | 197.2939 | 138.8957 | 143.583 | 173.3128 | 172.2891 | 145.4503 | 133 | 121.4455 | 201.0237 | 182.4455 | 137.7109 | 136.2891 | 247.2322 | 184.9431 | 133.6588 | 134.1043 | 213.3601 | 157.1043 | 299.237 | 313.7061 | 157.8152 | 214.5782 | 199.4503 | 221.3981 | 352.5497 | 112.7346 | 110.8436 | 93.6872 | 142.4218 | 134.3412 | 136.7346 | 139.8152 | 100.5497 | 204.0521 | 188.6588 | 224.8152 | 268.1612 | 113.3176 | 134.0285 | 110.0285 | 146.5782 | 73.7109 | 71.5261 | 62.6066 | 113.0237 | 217.5497 | 232.0521 | 233.3412 | 185.5309 | 123.8957 | 136.6824 | 123.1848 | 118.9715 | 164.9479 | 165.6019 | 191.0285 | 180.8152 | 175.2133 | 156.7346 | 180.1043 | 175.2133 | 169.3412 | 210.7346 | 194.1043 | 287.7582 | 156.8436 | 152 | 228.1801 | 227.1564 | 195.1327 | 201 | 203.2843 | 411.9194 | 165.9715 | 169.7109 | 148.5782 | 258.2322 | 131 | 138.3176 | 137.8152 | 177.1848 | 199.3223 | 206.9479 | 203.6019 | 155 | 101.7346 | 91.6066 | 105.8152 | 135.8436 | 245.0569 | 238.5782 | 230.6303 | 418.0237 | 344.0521 | 228.7394 | 208.4218 | 364.0237 | 201.5782 | 84.6303 | 78.2085 | 201.5782 | 123.3934 | 111.4739 | 125.7915 | 127.9479 | 108.4218 | 112.2181 | 99.8199 | 119.2133 | 65.5024 | 63.8436 | 73.7109 | 103.1327 | 119.1327 | 90.1612 | 111.4739 | 85.5545 | 225.4787 | 199.1848 | 253.9431 | 468.2702 | 204.7109 | 231.1375 | 196.9005 | 453.3412 | 517.6351 | 277.3934 | 223 | 650.9431
1854 | 22.3884 | 158.8766 | 176.1676 | 203.3922 | 188.2247 | 146.3039 | 135.2247 | 151.0884 | 178.4273 | 153.1805 | 118.2689 | 111.8195 | 168.0092 | 165.0663 | 127.6961 | 127.6961 | 250.5727 | 179.0442 | 143.6298 | 145.3039 | 219.6961 | 149.6961 | 311.4052 | 335.1805 | 155.1234 | 208.2468 | 211.8324 | 227.4052 | 358.7753 | 97.6169 | 125.0571 | 110.4844 | 153.4052 | 131.3481 | 116.965 | 137.8987 | 101.3481 | 227.4844 | 203.2818 | 204.291 | 286.8103 | 138.6869 | 127.674 | 116.0792 | 134.8195 | 67.5727 | 77.4273 | 63.8766 | 104.0442 | 218.3481 | 223.6961 | 239.6519 | 190.9558 | 123.5506 | 115.965 | 139.8324 | 126.8766 | 202.7661 | 179.7532 | 178.4714 | 165.7182 | 158.3702 | 154.7974 | 150.9871 | 164.5727 | 153.9429 | 193.291 | 203.7532 | 287.6961 | 158.6519 | 158.2026 | 253.0571 | 203.639 | 205.4052 | 215.7311 | 200.5948 | 397.9429 | 166.8987 | 177.5286 | 159.3039 | 245.291 | 131.7753 | 144.3039 | 130.4714 | 175.2247 | 206.3831 | 216.8545 | 210.7753 | 133.291 | 88.1676 | 113.8103 | 108.5506 | 126.7974 | 215.291 | 247.7532 | 239.2026 | 388.1105 | 367.4844 | 251.1363 | 224.0571 | 347.291 | 190.5948 | 79.2468 | 77.3481 | 190.5948 | 126.7753 | 114.1013 | 130.5286 | 148.709 | 103.8987 | 98 | 88.8987 | 134.9337 | 81.1584 | 75.7311 | 70.674 | 102.5506 | 116.2247 | 115.0129 | 115.6519 | 76.5727 | 225.3039 | 211.9558 | 234.0884 | 501.5194 | 210.2026 | 228.5286 | 192.2026 | 413.1326 | 506.3481 | 283.8766 | 250.1363 | 580.6924
1855 | 22.3994 | 137.9468 | 182.0266 | 184.6106 | 202.9647 | 151.3628 | 140.5133 | 165.8408 | 164.2477 | 149.1239 | 133.6903 | 111.9381 | 157 | 196.8055 | 147.0532 | 136.8142 | 239.5486 | 159.2477 | 145.4248 | 138.5486 | 192.584 | 150.9381 | 322.8761 | 331.9115 | 180.7789 | 234.3541 | 215.4248 | 220.761 | 325.9202 | 107.7523 | 107.1592 | 108.3362 | 145.9734 | 132.3628 | 139.9293 | 139.3628 | 94.4867 | 222.1858 | 225.3275 | 220.0532 | 297.1505 | 131.2477 | 122.2743 | 113.4867 | 118.3982 | 72.3009 | 72.6991 | 80.7523 | 112.6018 | 255.5931 | 238.3275 | 237.0619 | 189.4867 | 136.6018 | 118.4514 | 123.7344 | 139.0266 | 170.6812 | 180.4248 | 186.239 | 189.3541 | 151.8495 | 156.1505 | 167.2036 | 162.2124 | 149 | 225.5931 | 194.9734 | 273.9734 | 176.5399 | 147.3982 | 228.8583 | 226.7169 | 215.3009 | 183.7698 | 181.2477 | 384.3362 | 166 | 152.4335 | 161.2124 | 258.6903 | 145.3894 | 137.5486 | 139.0266 | 158.4601 | 213.5752 | 219.2124 | 215.7257 | 134.8761 | 95.6018 | 125.7257 | 124.7523 | 120.2743 | 272.7967 | 239.761 | 240.2124 | 417.3807 | 392.1151 | 260.5752 | 239.239 | 379.5931 | 191.7257 | 80.1505 | 69.6991 | 191.7257 | 121.4867 | 111.8495 | 121.1858 | 121.1326 | 108.5133 | 108.239 | 108.478 | 116.584 | 75.9734 | 64.2477 | 75.5133 | 101.4248 | 117.5752 | 107.4601 | 105.9734 | 82.8761 | 224.0619 | 210.4867 | 230.7257 | 524.4514 | 177.345 | 228.6372 | 231.0179 | 442.7435 | 499.4867 | 279.4867 | 227.2831 | 609.9825
1856 | 22.4104 | 153.5974 | 185.7999 | 200.1977 | 203.8004 | 136.8018 | 157.9981 | 172.7999 | 172.7985 | 140.6008 | 127.0014 | 123.3988 | 152.2006 | 208.5998 | 149.8004 | 163.1972 | 248.1987 | 171.798 | 129.0019 | 138.3997 | 182.8004 | 158.3992 | 323.4002 | 324.4007 | 172.8018 | 231.6012 | 206.201 | 225.9991 | 340.1972 | 108.0005 | 113.7985 | 116.3988 | 120.6027 | 134.5998 | 149.5998 | 127.2015 | 93.7999 | 227.799 | 229.2001 | 198.8032 | 267.6036 | 137.3988 | 129.799 | 108.8004 | 140.597 | 63.6012 | 72.5998 | 90.5993 | 104.8013 | 232.404 | 227.0019 | 261.597 | 182.4007 | 132.8009 | 126.5993 | 126.799 | 141.2001 | 159.7999 | 170.4011 | 196.9991 | 176.2025 | 158.1991 | 138.6022 | 170.8004 | 160.4002 | 153.7994 | 208.8032 | 236.7947 | 290.1977 | 173.0009 | 133.6012 | 219.4002 | 211.8028 | 196.2025 | 185.1987 | 208.9962 | 385.9995 | 156.4011 | 153.799 | 153.8009 | 243.2025 | 149 | 146.9986 | 142 | 165.1987 | 204.4011 | 200.6022 | 214.6003 | 148.9986 | 104.5993 | 114.2015 | 132.1996 | 119 | 242.8056 | 239.3997 | 263.9972 | 415.2015 | 373.0028 | 252.201 | 247.5993 | 378.0014 | 175.4021 | 69.8013 | 76.799 | 175.4021 | 96.0028 | 110.2001 | 125.1991 | 109.6003 | 94.8018 | 110.2001 | 110.8004 | 122.7985 | 75.3997 | 71.1987 | 72.2006 | 97.8004 | 122.7994 | 127.7971 | 110.9991 | 73.8013 | 238.9981 | 212.9995 | 228.0005 | 481.4054 | 194.3969 | 228.7999 | 217.8028 | 423.6036 | 490.8009 | 286.799 | 281.3926 | 569.6065
1857 | 22.4214 | 128.9487 | 194.9342 | 180.6341 | 202.1878 | 142.7464 | 176.6195 | 158.3805 | 163.8171 | 134.939 | 119.1268 | 111.3805 | 151 | 209 | 151.4366 | 169 | 225.0097 | 187.3707 | 146.9293 | 144.6854 | 177.939 | 139.6951 | 322.1878 | 334.3707 | 160.0658 | 237.122 | 196.6902 | 228.8122 | 342.7512 | 98.878 | 103.1927 | 111.6902 | 133.6805 | 126.878 | 145.939 | 128.8732 | 87.5024 | 232.4366 | 241.1829 | 222.8635 | 267.3098 | 145.6854 | 132 | 122.6195 | 130.7561 | 69.9342 | 58.3805 | 80.6293 | 106.061 | 227.2488 | 244.1171 | 275.3098 | 192.3707 | 132.6244 | 130.4366 | 129.8122 | 119.8829 | 203.0464 | 173.6854 | 178.6951 | 176.6854 | 153.5024 | 153.4927 | 186.2439 | 142.9439 | 138.7561 | 195.5024 | 230.1317 | 284.4415 | 144.1975 | 161.8635 | 218.1878 | 221.4317 | 195.8732 | 186.3756 | 199.9439 | 404.8683 | 157.2488 | 165.7464 | 146.3146 | 246.122 | 157.9342 | 140.2536 | 139.5634 | 180.9951 | 202 | 192.7512 | 201.8171 | 171.4927 | 107.6244 | 98.0049 | 110.2585 | 123.8732 | 232.6244 | 224.5683 | 248.8829 | 423.3707 | 362.1268 | 289.7976 | 257.122 | 393.6805 | 173.4366 | 70.2488 | 70.878 | 173.4366 | 120.0513 | 119.7464 | 112.3805 | 128.4927 | 111.3049 | 114.061 | 120.5585 | 132.4976 | 71.1268 | 70.7512 | 75.061 | 105.122 | 111.8171 | 114.5073 | 108.1268 | 77.4976 | 237.3146 | 201.8171 | 198.5731 | 500.0513 | 202.6244 | 216.8171 | 201.4415 | 428.9951 | 445.1414 | 319.0513 | 283.1927 | 549.5024
1858 | 22.4324 | 127.7714 | 187.1062 | 207.155 | 225.9101 | 140.8776 | 154.4409 | 160.7714 | 155.2286 | 171.9264 | 115.5265 | 112.1224 | 155.1224 | 198.2817 | 135.5103 | 144.2654 | 229.7183 | 167.7388 | 158.5959 | 160.0163 | 191.0163 | 139.1224 | 340.1387 | 355.1387 | 172.0163 | 214.2654 | 223.857 | 238.8938 | 326.3347 | 121.7346 | 111.5428 | 124.0163 | 133.8776 | 150.5591 | 121.9144 | 115.1592 | 95.8938 | 257.7346 | 228.3347 | 238.2449 | 277.2449 | 139.5796 | 132 | 115.2817 | 122.8776 | 74.4735 | 76.4367 | 82.9469 | 106.1755 | 225.5265 | 233.3347 | 266.2817 | 175.2123 | 109.0899 | 136.7714 | 126.7021 | 94.3878 | 195.6858 | 169.2286 | 213.5754 | 194.4897 | 139.6327 | 148.1062 | 157.0205 | 147.2449 | 145.7183 | 208.0163 | 215.2817 | 299.3142 | 131.4041 | 157.4572 | 227.8938 | 206.8613 | 192.0531 | 216.506 | 195.1755 | 385.9144 | 163.7714 | 171.2979 | 163.9632 | 244.7021 | 155.0531 | 141.2979 | 150.5428 | 194.7183 | 202 | 206.8408 | 223.7346 | 138.8981 | 89.8613 | 83.4572 | 99.2286 | 126.649 | 220.6327 | 221.8245 | 254.7183 | 416.1062 | 353.5796 | 284.1592 | 266.4204 | 426.857 | 190.4897 | 75.9469 | 65.7021 | 190.4897 | 120.4041 | 125.2979 | 120.5428 | 135.4735 | 114.351 | 111.7021 | 111.4572 | 129.0531 | 65.8776 | 64.2286 | 68.5796 | 102.0531 | 90.0368 | 117.4204 | 105.351 | 85.5959 | 225.2817 | 194.0531 | 216.7346 | 488.8613 | 194.7551 | 227.1918 | 222.9101 | 431.1755 | 444.0694 | 315.2817 | 319.5754 | 523.2654
1859 | 22.4434 | 131.5103 | 196.7149 | 214 | 189.1611 | 133.3058 | 154.0207 | 154.469 | 174.0827 | 161.5909 | 127.5517 | 123.0413 | 164.3678 | 200.1839 | 137.8574 | 141.5103 | 212.7541 | 151.2851 | 155.8161 | 155.469 | 183.1219 | 155.8988 | 333.9587 | 352.3058 | 145.7128 | 228.2459 | 218.2851 | 225.938 | 317.1426 | 115.2851 | 103.9587 | 116.9587 | 107.0599 | 130.8967 | 146.2872 | 127.062 | 92.1426 | 216.9772 | 228.3471 | 233.3058 | 269.7954 | 131.3058 | 123.6322 | 113 | 108.6115 | 85.0413 | 81 | 84 | 111.0207 | 215.7954 | 241.7149 | 262.3264 | 196.1033 | 122.4091 | 131.3058 | 111.7748 | 116.7769 | 169.407 | 175.531 | 232.0413 | 194.6529 | 143.6942 | 156.8781 | 173.4298 | 156.531 | 174.7769 | 221.8781 | 213.8368 | 268.6921 | 140.0413 | 153.3264 | 228.3264 | 196.3058 | 201.0413 | 197.8967 | 176.5909 | 407.7769 | 145.7541 | 151.0805 | 137.0392 | 256.5517 | 124.7128 | 170.4505 | 155.5103 | 165.2024 | 173.5495 | 182.3863 | 214.7748 | 136.0207 | 101.8988 | 97.7356 | 103.0207 | 114.4483 | 207.1219 | 227.0207 | 243.6115 | 393.0805 | 335.2644 | 313.6343 | 262.1426 | 417.1012 | 195.6736 | 83.6942 | 80.062 | 195.6736 | 123.1839 | 117.6322 | 123 | 130.1426 | 119.8574 | 121.8781 | 113.1839 | 131.3471 | 70.0207 | 61.3264 | 82.062 | 101 | 113.6137 | 117.3264 | 140.1447 | 97.8781 | 216.3058 | 185.469 | 237.8988 | 466.5909 | 146.9772 | 195.6921 | 226.3264 | 413.4277 | 395.7933 | 330.5723 | 281.9772 | 513.8161
1860 | 22.4544 | 135.3963 | 166.7354 | 254.7553 | 198.8305 | 160.0193 | 160.9435 | 157.2453 | 178 | 149.5093 | 109.6223 | 123.3019 | 171.0944 | 184.0186 | 137.3019 | 163.2267 | 233.623 | 178.7174 | 138.0186 | 131.0751 | 186.0944 | 126.7354 | 326.9056 | 363.736 | 164.623 | 229.4528 | 210.0565 | 244.2267 | 352.51 | 115.5472 | 90.113 | 130.2832 | 122.3777 | 135.3398 | 128.2261 | 134.2453 | 98.6416 | 235.1702 | 219.6602 | 202.2826 | 289.2267 | 139.3398 | 129.6416 | 132.5286 | 129.7739 | 81.0565 | 76.7547 | 77.2074 | 109.4528 | 188.5279 | 231.264 | 268.7926 | 180.6223 | 121.7547 | 113.0186 | 132.7739 | 114.3584 | 167.5472 | 160.8677 | 205.1316 | 165.9807 | 139.9056 | 149.6602 | 172.9056 | 158 | 168.113 | 193.4335 | 225.0379 | 287.4721 | 153.0379 | 159.7926 | 228 | 214.5286 | 204.6981 | 204.887 | 201.0193 | 408.7547 | 142.8491 | 147 | 143.736 | 264.9435 | 128.3398 | 163.264 | 145.8112 | 159 | 178.1888 | 173.6037 | 188.2261 | 128.5093 | 91.4149 | 76.377 | 105.6981 | 114.5472 | 212.6416 | 235.6416 | 238.4528 | 372.0186 | 352.3777 | 347.1702 | 267.7926 | 449.6609 | 195.1509 | 90.9435 | 76.2074 | 195.1509 | 108.7168 | 121.9435 | 123 | 115.4149 | 109.9621 | 112.113 | 103.8112 | 125.2074 | 70.1509 | 62.6981 | 81.6037 | 105.2453 | 79.0937 | 103.4149 | 114.7354 | 93.2074 | 212.4528 | 192.4907 | 215.5279 | 510.5479 | 143.9435 | 204.2832 | 214.113 | 387.9242 | 383.4528 | 317.0186 | 303.5665 | 487.5279
1861 | 22.4654 | 148.9204 | 186.8408 | 210.3184 | 198.5546 | 140.0206 | 159.4159 | 171.7818 | 186.6136 | 129.0501 | 119.7818 | 126.4454 | 169.4159 | 196.5045 | 154.2272 | 145.466 | 231.1091 | 166.7728 | 136.7227 | 157.1476 | 174.941 | 110.6637 | 328.5841 | 336.7137 | 151.7728 | 220.3864 | 215.0295 | 222.1592 | 322.8229 | 107.3864 | 103.5045 | 125.2478 | 133.7522 | 149.9204 | 131.7522 | 128.9705 | 90.525 | 210.7137 | 235.2272 | 227.1476 | 285.2478 | 113.4365 | 131 | 126.525 | 116.7728 | 60.1887 | 76 | 70.8318 | 118.475 | 189.1682 | 216.941 | 240.7137 | 163.2182 | 133.059 | 134.1181 | 129.2478 | 123.3363 | 168 | 193.3158 | 209.475 | 168.7522 | 131.2478 | 161.7818 | 174.5841 | 156.2773 | 154.8023 | 178.525 | 230.4454 | 250.6547 | 148.9705 | 153.2478 | 227.1386 | 210.2478 | 208.4454 | 188.9114 | 197.3864 | 392.4955 | 178.3158 | 144.4159 | 140.8318 | 232.4069 | 130 | 129.1297 | 146.5841 | 148.6637 | 178.2773 | 191.9499 | 190.0295 | 139.9204 | 108.8113 | 86.6431 | 104.2773 | 108.1091 | 215.7227 | 231.8318 | 230.2478 | 381.9204 | 325.8524 | 339.0796 | 309.4839 | 440.4955 | 191.5546 | 80.8023 | 73.2773 | 191.5546 | 104.2773 | 123.8614 | 107.4955 | 113 | 95.941 | 93.6342 | 114.059 | 111.0796 | 63.1091 | 71.6136 | 56.0206 | 90.4955 | 109.8999 | 85.4955 | 115.8909 | 101.475 | 232.6726 | 217.2567 | 223.9204 | 471.6252 | 165.6726 | 225.0886 | 235.2567 | 415.8703 | 356.2978 | 325.1977 | 270.2388 | 455.4365
1862 | 22.4764 | 150.1263 | 206.7257 | 230.8305 | 189.2635 | 122.8952 | 145.8952 | 164.3898 | 136.4546 | 131.2419 | 150.8305 | 129.621 | 155.0216 | 189.3898 | 175.3467 | 158.5994 | 227.379 | 178.8521 | 112.5378 | 136.6641 | 189.5994 | 132.5886 | 308.906 | 318.0216 | 166.473 | 203.2743 | 214.2527 | 241.5886 | 353.6934 | 130.4622 | 115.6102 | 116.1371 | 111.4114 | 149.379 | 126.0108 | 128.8737 | 103.8521 | 200.7581 | 224.8952 | 226.7581 | 269.1479 | 123.8521 | 124.0108 | 133.7365 | 140.2095 | 70.1048 | 61.1479 | 66.5054 | 134.8521 | 183.0108 | 197.527 | 240.3683 | 161.8737 | 125.3898 | 140.621 | 133.2419 | 126.7473 | 158.3898 | 157.0647 | 225.8521 | 144.6641 | 143.1048 | 159.6317 | 188.9784 | 171.7257 | 160.8629 | 163.0216 | 203.0432 | 289.4299 | 142.7581 | 134.527 | 192.9276 | 195.0216 | 172.3066 | 208.7149 | 203.8629 | 382.1371 | 139.4438 | 154.4838 | 170.5778 | 240.9784 | 149.2203 | 133.6102 | 136.5162 | 151.3683 | 184.9892 | 173.1587 | 196.2419 | 130.6425 | 86.6641 | 86.379 | 89.1479 | 127.094 | 184.5486 | 214.4006 | 236.8629 | 391.8629 | 295.6641 | 372.8197 | 307.2635 | 527.1124 | 171.7797 | 78.1263 | 93.9676 | 171.7797 | 127.5886 | 117.8844 | 113.7365 | 113 | 101.8629 | 113.7149 | 96.7797 | 109.8737 | 72.4838 | 68.6317 | 73.8413 | 83.6317 | 97.6533 | 93.4838 | 103.0216 | 77.6641 | 227.2635 | 209.6425 | 241.7257 | 472.7365 | 156.7689 | 203.5378 | 230.2635 | 427.9892 | 352 | 327.8737 | 322.5346 | 408.1911
1863 | 22.4874 | 130.5093 | 209.8859 | 214.6233 | 180.9125 | 144.0345 | 167.9204 | 162.1141 | 167.9815 | 121.3687 | 123.992 | 118.4827 | 144.1406 | 195.9735 | 155.8514 | 160.1141 | 214.5968 | 199.6048 | 132.0345 | 144.5173 | 198.2016 | 139.5438 | 347.6393 | 323.9735 | 169.8859 | 189.4827 | 208.6843 | 215.7639 | 318.2466 | 126.9125 | 101.939 | 138.0345 | 122.1751 | 138.3687 | 135.6313 | 125.4562 | 75.878 | 214.1751 | 227.4297 | 231.3157 | 266.1141 | 118.0265 | 130.0875 | 108.4217 | 117.4217 | 72 | 74.061 | 66 | 118.3952 | 186.4297 | 206.5173 | 225.939 | 177.061 | 116.9125 | 122.3952 | 130.4562 | 111.053 | 164.0875 | 168.7188 | 214.7109 | 156.947 | 121.9655 | 150.1406 | 167.0796 | 151.8514 | 169.9735 | 189.3502 | 220.2626 | 262.3342 | 147.3157 | 163.8939 | 194.2016 | 193.8859 | 189.1486 | 189.8514 | 208.5438 | 356.1936 | 139.2016 | 145.3687 | 152.8514 | 243.8859 | 144.0265 | 139.4297 | 156.2626 | 159.0875 | 154.992 | 177.0875 | 198.7719 | 147.6048 | 92.7454 | 96.6313 | 88.7719 | 117.5968 | 198.6048 | 204.0265 | 214.9655 | 340.7294 | 300.8594 | 388.6313 | 325.4907 | 502.7904 | 179.6313 | 66.4827 | 83.7109 | 179.6313 | 99.1061 | 113.4562 | 124.7454 | 118.3157 | 95.0265 | 121.4297 | 92.2281 | 115.3157 | 79.3157 | 77.7454 | 70.7984 | 92.7454 | 97.6578 | 79.053 | 96.5703 | 81.0875 | 215.3687 | 212.4297 | 220.9655 | 463.3687 | 158.5438 | 164.5623 | 211.2812 | 414.8249 | 337.8249 | 310.2812 | 304.4217 | 444.5252
1864 | 22.4984 | 164.8276 | 178.5618 | 197.6283 | 171.9159 | 161.3717 | 150.3406 | 159.3053 | 144.2565 | 140.6594 | 128.0841 | 134.0664 | 150.1859 | 173.6459 | 181.7435 | 159.1018 | 235.4558 | 163.376 | 150.27 | 151.3894 | 190.0177 | 140.8982 | 310.783 | 322.3053 | 156.5265 | 210.4558 | 199.9159 | 200.323 | 335.4558 | 91.8671 | 117.9647 | 99.6813 | 126.6947 | 126.2212 | 122.6283 | 127.6947 | 91.7611 | 222.2876 | 224.4071 | 246.3717 | 282.1682 | 131.3717 | 124.7124 | 120.27 | 114 | 86.3717 | 59.8318 | 67.7965 | 107.0177 | 171.73 | 205.3053 | 238.3717 | 150.2565 | 119.5929 | 135.27 | 125.5088 | 119.7788 | 159.6106 | 174.5929 | 188.7477 | 150.9159 | 138.7611 | 136.4247 | 146.9336 | 147.2035 | 171 | 171.4424 | 222.1018 | 256.2035 | 146.2035 | 160.8141 | 214.7611 | 201.1859 | 198.2876 | 188.7965 | 193.73 | 379.0488 | 128.323 | 149.3894 | 152.6947 | 214.3583 | 138.5088 | 143.5929 | 163.4912 | 147.4247 | 166.27 | 179.7965 | 198.1018 | 131.1371 | 102.0841 | 108.7788 | 96.1859 | 116 | 211.7788 | 220.0664 | 215.5929 | 347.4735 | 286.73 | 358.5618 | 334.2876 | 506.9823 | 188.1859 | 82.9647 | 83.7965 | 188.1859 | 97.6947 | 108.5088 | 139.4735 | 135.1682 | 97.5929 | 99.5442 | 106.3717 | 113.3053 | 71.9159 | 70.9159 | 79.8806 | 91.3053 | 91.7124 | 95.8629 | 112.1682 | 80.2035 | 239.1505 | 195.9336 | 226.9823 | 471.8806 | 129.3583 | 180.6594 | 214.3894 | 402.2212 | 301.8671 | 296.323 | 292.9159 | 431.1371
1865 | 22.5094 | 151.7 | 184.1052 | 181.4316 | 187.3894 | 149.3421 | 178.9578 | 165.3737 | 146.4631 | 115.6843 | 111.7 | 121.4316 | 172.8526 | 191.9421 | 173.1632 | 161.7316 | 209.7737 | 174.4789 | 138.3421 | 134.7 | 175.3421 | 122.7895 | 306 | 329.2842 | 162.2842 | 192.0579 | 225.4052 | 222.6736 | 337.0895 | 113.4947 | 107.2527 | 115.0315 | 139.7473 | 145.9421 | 126.4631 | 150.7631 | 98.5526 | 203.879 | 224 | 245.2684 | 299.4789 | 112.9685 | 116.7158 | 112.8948 | 97.6106 | 77.0737 | 67.1052 | 73.4631 | 118.7473 | 190.9421 | 193.1632 | 226.3421 | 153.3737 | 119.0895 | 126.0737 | 137.7473 | 130.1052 | 150.8053 | 177.7316 | 198.7473 | 146.3579 | 130.9842 | 144.1052 | 148.6421 | 140.6263 | 205.5999 | 188.121 | 205.6106 | 246.8948 | 125.0579 | 136.3264 | 216.0895 | 215.6579 | 183.5211 | 206.3 | 180.1632 | 362.879 | 126.0895 | 124.5053 | 146.6263 | 259.2578 | 137.0895 | 131.2527 | 143.0579 | 176.0473 | 176.1947 | 158.1474 | 198 | 138.1052 | 92.9842 | 99.9842 | 107.9263 | 105.9842 | 198.4316 | 224.7316 | 228.7473 | 333.5211 | 247.6685 | 345.8948 | 307.6843 | 499.8053 | 158.0422 | 65.879 | 83.0895 | 158.0422 | 106.1947 | 103.4474 | 127.3421 | 117.879 | 100.7316 | 94.2684 | 110.7316 | 123.9263 | 73.7316 | 79.1052 | 75.5369 | 105.5684 | 110.121 | 91.6263 | 106.7158 | 64.5211 | 213.7737 | 189.4474 | 205.2369 | 462.0737 | 128.7316 | 191.1947 | 202.2527 | 374.5948 | 298.9105 | 295 | 279.2527 | 411.7
1866 | 22.5204 | 158.3053 | 186.8456 | 215.067 | 189 | 151.6913 | 165.3893 | 166 | 124.8524 | 125.9194 | 118.3053 | 112.6175 | 153.7752 | 175.5437 | 145.2383 | 173.9966 | 215.3053 | 172.3087 | 129.6175 | 148.6879 | 187.8422 | 123.7684 | 267.2417 | 341.0738 | 182.3791 | 214.916 | 186.4733 | 236.0738 | 342.5369 | 106.7718 | 100.4631 | 103.1578 | 131.7718 | 150.7684 | 133.4597 | 141.9262 | 82.3893 | 206.6141 | 192.6243 | 230.235 | 301.9228 | 116.5369 | 135.3791 | 127.6879 | 117.2248 | 76 | 51.3893 | 84.151 | 100.6209 | 198.5369 | 192 | 198.2383 | 179.8388 | 109.7718 | 118.5403 | 127.9262 | 119.0034 | 148.1544 | 162.3121 | 218.4563 | 179.2214 | 119.849 | 131.1578 | 165.6107 | 119.6981 | 180.3927 | 173.3893 | 225.2248 | 248.7684 | 150.6845 | 154.3019 | 228.9194 | 225.3053 | 146.933 | 176.6243 | 214.067 | 348.0806 | 135.2282 | 105.3893 | 161.6879 | 237.2383 | 137 | 115.235 | 142.8456 | 149.4699 | 145.6243 | 161.5369 | 204.4597 | 134.3859 | 90.1544 | 97.1544 | 91.4665 | 101.3087 | 185.0034 | 213.9262 | 211.5437 | 340.3053 | 226.4665 | 359.765 | 323.4563 | 468.5471 | 147.6175 | 65.8456 | 69.1578 | 147.6175 | 118.0738 | 103.9228 | 107.5437 | 126.151 | 80.6981 | 105.9966 | 111 | 108.3893 | 73.0772 | 62.4665 | 78.6913 | 89.4665 | 99.0806 | 90.0772 | 104.1544 | 72.2282 | 201.7718 | 219.4529 | 254.6777 | 469.3053 | 149.3019 | 175.3893 | 188.0806 | 377.5369 | 315.6107 | 266.3927 | 253.084 | 441.3757
1867 | 22.5314 | 126.2713 | 202.8968 | 198.3628 | 184.3245 | 147.3245 | 169.6106 | 167.8702 | 183.7819 | 139.1564 | 123.6755 | 128.8319 | 143.584 | 191.767 | 168.2479 | 152.5574 | 217.8702 | 181.3511 | 147.7021 | 154.6755 | 203.9617 | 150.183 | 306.0798 | 331.7138 | 143.7904 | 203.9085 | 178.3245 | 261.3128 | 331.7787 | 106.9351 | 114.9617 | 123.5075 | 125.3894 | 139.7787 | 125.584 | 144.7404 | 89.416 | 181.7521 | 192.8053 | 254.2479 | 262.7255 | 139.4426 | 146.3511 | 133.6755 | 99.3628 | 75.0649 | 60.2862 | 77.5191 | 108.3511 | 167.2064 | 184.5191 | 231.5341 | 190.416 | 115.5457 | 123.6106 | 127.9351 | 94.6223 | 148 | 161 | 189.1415 | 168.9085 | 122.7404 | 136.5457 | 165.1298 | 132.0266 | 216.3394 | 179.4809 | 195.2064 | 234.0383 | 144.584 | 134.4925 | 203.817 | 231.6106 | 186.0798 | 168.3894 | 199.233 | 356.3511 | 157.5075 | 126.4426 | 169.5457 | 208.817 | 141.6755 | 112.1298 | 135.5191 | 137.6489 | 143.9351 | 126.4659 | 191.9085 | 145.2213 | 93.7404 | 90.4543 | 100.2862 | 118.767 | 205.5075 | 216.7404 | 217.4809 | 366.2479 | 241.8319 | 369.416 | 341.8319 | 469.7404 | 160.0915 | 65.0649 | 76.416 | 160.0915 | 110.584 | 95.584 | 104.1298 | 125.1298 | 107.0532 | 103.2596 | 102.584 | 109.8053 | 82.3511 | 63.8053 | 92.0915 | 101.0915 | 94.2596 | 89.0649 | 104.9351 | 73.9351 | 220.6372 | 177.1149 | 229.0766 | 468.1298 | 156.6106 | 162.7787 | 201.9617 | 429.4309 | 272.1149 | 246.233 | 282.7936 | 396.3096
1868 | 22.5424 | 140.1058 | 219.1584 | 209.3162 | 169.789 | 144.1578 | 164.3156 | 179.3688 | 145.3671 | 126.7364 | 103.1573 | 122.4208 | 165.7375 | 174.9994 | 166.2104 | 177.5271 | 204.7364 | 180.1052 | 117.7359 | 158.7896 | 175.6307 | 144.4208 | 304.263 | 299.7359 | 144.7896 | 188.789 | 201.6849 | 241.2099 | 302.5781 | 127.8427 | 94.2099 | 106.9994 | 133.5266 | 134.263 | 126.8948 | 158.2636 | 100.4214 | 213.1589 | 196.7896 | 256 | 258.1052 | 132.4734 | 115.7359 | 125.4734 | 124.5271 | 74.0526 | 61 | 80.7896 | 94.789 | 194.3693 | 184.9474 | 198.9463 | 153.1041 | 137.7901 | 139.1584 | 126.1052 | 129.0011 | 158.4214 | 163.8422 | 204.0532 | 149.9994 | 110.6838 | 137 | 163.1052 | 131.1052 | 174.4723 | 166.7364 | 200.5792 | 242.474 | 135.4734 | 143.4214 | 203.8948 | 198.8411 | 188.0526 | 182.211 | 197.0526 | 371.211 | 138.1573 | 149.7901 | 162.4208 | 221.211 | 121.1573 | 119.5792 | 131.2104 | 137.9474 | 168.6323 | 168.5277 | 199.5266 | 135.5786 | 91.1578 | 101.3688 | 103.8422 | 124.737 | 222.1584 | 202.789 | 240.7375 | 357.5786 | 208.8937 | 354.8416 | 318.3677 | 472.8422 | 171.4214 | 57.4208 | 65.6312 | 171.4214 | 104.3156 | 97.8422 | 104.9474 | 113.6312 | 114.6844 | 109.6318 | 120.0006 | 109.0526 | 72.5786 | 69.6844 | 73.1046 | 108.6318 | 101.5792 | 77.6312 | 113.5266 | 72.1052 | 190.7359 | 185.3688 | 227 | 480.3162 | 142.789 | 175.2636 | 192.5786 | 414.052 | 269.9474 | 224.1573 | 289.737 | 411.948
1869 | 22.5534 | 146.7581 | 230.5566 | 203.2822 | 209.3069 | 120.9675 | 171.6775 | 156.9675 | 151.6372 | 116.4031 | 113.5163 | 110.4837 | 173.7178 | 162.4837 | 194.7907 | 166.524 | 204.9597 | 137.7737 | 159.186 | 140.7659 | 175.9194 | 147.8388 | 283.8465 | 287.4434 | 156.5163 | 191.8388 | 179.0078 | 221.7659 | 319.2341 | 141.476 | 99.7178 | 101.2016 | 128.2419 | 123.4434 | 131.7984 | 139.8062 | 77.0078 | 206.3628 | 212.355 | 233.9272 | 267.5969 | 120.4837 | 116.8791 | 132.6775 | 114.4837 | 69.2016 | 64.8388 | 68.524 | 114.1535 | 191.2016 | 187.8791 | 221.9519 | 151.9597 | 134.2016 | 120.8062 | 112.5644 | 120.4434 | 169.5566 | 161.1209 | 204.0403 | 165.3147 | 132.0728 | 137.9597 | 140.9272 | 151.1535 | 178.7178 | 155.4434 | 201.9597 | 263.1535 | 135 | 144 | 223.1938 | 196.0403 | 173.6047 | 185.8791 | 182.6047 | 332.6528 | 131.2419 | 150.0403 | 145.6853 | 217.2016 | 111.3628 | 113.2822 | 143.476 | 126.4837 | 165.2016 | 156.6047 | 191.3628 | 152.2744 | 92.9194 | 99.1209 | 101.1209 | 101.9675 | 214.3628 | 174.169 | 239.1209 | 334.9272 | 229.0728 | 318.4915 | 340.9922 | 411.5799 | 172 | 64.6775 | 71.7178 | 172 | 100.1612 | 97.0403 | 114.5969 | 120.6775 | 86.2093 | 105.2016 | 121 | 112.8388 | 72 | 91.1132 | 78.7178 | 85.0078 | 96.2419 | 104.831 | 104.4031 | 69.1209 | 213.9519 | 196.5566 | 217.4031 | 487.7178 | 133.3628 | 188.476 | 194.8791 | 431.2341 | 274.7984 | 214.3628 | 266.0078 | 363.0962
1870 | 22.5644 | 136.3082 | 200.8686 | 217.5797 | 165.317 | 147.2154 | 137.9807 | 174.4676 | 148.1121 | 148.0753 | 102.3362 | 122.6357 | 147.7565 | 194.0753 | 161.9807 | 142.6725 | 205.972 | 166.1314 | 145.4483 | 163.3275 | 169.1961 | 131.4763 | 301.4676 | 321.9913 | 155.056 | 210.4676 | 187.7198 | 226.8319 | 356.9352 | 115.7565 | 102.9159 | 116.5517 | 138.6918 | 114.2522 | 101.8686 | 148.7198 | 89.6077 | 215.7198 | 196.4763 | 228.1401 | 274.8039 | 119.028 | 117 | 120.3643 | 130.5237 | 90.3836 | 77.6357 | 77.7198 | 91.6725 | 203.6357 | 183.1401 | 214.2522 | 162.6918 | 129.1401 | 123.8879 | 105.1961 | 114.1681 | 148.6164 | 170.7198 | 215.6638 | 152.3923 | 148.5517 | 120.5044 | 167.2154 | 141.3082 | 200.3836 | 152.0841 | 199.0841 | 242.6164 | 132.0841 | 142.056 | 211.3643 | 195.028 | 200.2154 | 187.944 | 190.7478 | 387.3748 | 133.9159 | 126.6725 | 155.6918 | 233.5237 | 123.6357 | 133.4116 | 156.6357 | 130.8599 | 157.2242 | 168.6357 | 186.1401 | 131.6164 | 104.6638 | 92.1961 | 95.1681 | 99.056 | 232.4676 | 225.4869 | 226.3643 | 355.3836 | 206.6725 | 316.028 | 344.9159 | 424.5517 | 188.5237 | 67.9159 | 82.6918 | 188.5237 | 89.3082 | 119.3555 | 109.1681 | 121 | 106.3836 | 93.3362 | 119.056 | 108.1401 | 72 | 64.7846 | 83.8599 | 82.056 | 105.7198 | 62.2609 | 102.056 | 70.944 | 223.7478 | 223.2435 | 216.028 | 451.0648 | 157.2995 | 179.2802 | 190.1401 | 401.8686 | 247.7846 | 223.7198 | 249.4483 | 359.056
1871 | 22.5754 | 153.7169 | 185.2359 | 209.1416 | 186.6382 | 148 | 156.6854 | 156.2989 | 143.0786 | 113.5662 | 110.8584 | 104.2989 | 147.9843 | 156.6134 | 147.2202 | 145.9371 | 236.5124 | 162.0786 | 149.9214 | 134.4719 | 169.9843 | 112.2989 | 279.3618 | 315.1258 | 150.0786 | 194.2674 | 185.0472 | 224.0472 | 301.8966 | 105.1573 | 97.0944 | 111.0944 | 129.1573 | 126.7955 | 119.7011 | 111.5977 | 84.0944 | 199.2674 | 217.654 | 200.4404 | 269.0944 | 118.0157 | 127.827 | 120.9843 | 115.2517 | 71.3146 | 71.1101 | 89.8113 | 105.7641 | 161.6764 | 174.1416 | 232.7011 | 172.8427 | 98.4876 | 117.1101 | 130.591 | 120.8899 | 140.1258 | 171 | 188.4404 | 137.2359 | 130.2989 | 150.5124 | 163.0786 | 148.8742 | 201.9843 | 194.3236 | 204.9056 | 223.2989 | 127.0786 | 147.9056 | 204.1101 | 169.409 | 190.173 | 182.0944 | 189.0315 | 383.0944 | 138.9214 | 150.6068 | 143.2045 | 256.6382 | 119.0786 | 134 | 153.0629 | 134.9371 | 162.9056 | 142.4247 | 198.7955 | 131.9843 | 101.0629 | 76.2517 | 95.9843 | 112.7798 | 200.5191 | 209.2831 | 234.8584 | 363.8742 | 206 | 276.6292 | 332.2045 | 409.2517 | 155.5348 | 59.1416 | 62.3303 | 155.5348 | 110.654 | 105.2359 | 117.8584 | 121.9843 | 97.1573 | 108.7483 | 99.3146 | 124.7326 | 85.7798 | 70.8899 | 58.409 | 92.827 | 117.8113 | 99.3866 | 97.0786 | 77.8899 | 199.3932 | 183.6449 | 214.0315 | 439.173 | 156.0315 | 163.2517 | 187.0472 | 429.5439 | 233.2202 | 187.582 | 246.0472 | 408.2135
1872 | 22.5864 | 158.9828 | 199.9484 | 178.1066 | 176.0378 | 133.0516 | 145.0413 | 156.9966 | 132.0378 | 113.9966 | 110.0034 | 147.8487 | 162.9484 | 189.8831 | 129.0619 | 175.8969 | 204.1135 | 138.0825 | 149.0034 | 150.9416 | 199.8969 | 137.9106 | 315.8728 | 312.0103 | 140.0344 | 211.9381 | 200.945 | 196.0963 | 300.0034 | 98.0241 | 113.9416 | 122.9587 | 110.0653 | 131.9828 | 125.9794 | 116.9794 | 115.89 | 205.9759 | 191.0928 | 218.9347 | 241.0963 | 111.0241 | 125.0103 | 134.9519 | 114.0034 | 88.9381 | 62.0309 | 76.0481 | 116.9622 | 196.8762 | 209.8762 | 207.0894 | 158.0516 | 110.9553 | 117.9966 | 126.0172 | 110.0378 | 162.9209 | 155.055 | 196.9691 | 149.9553 | 154.914 | 134.0584 | 157.0206 | 146.0103 | 159.1478 | 168.0928 | 211.9759 | 276.8143 | 121.0206 | 140.0275 | 228.914 | 192.9175 | 199.9656 | 181.0034 | 212.9175 | 371.0413 | 131.0275 | 151.9966 | 163.9278 | 216.141 | 119 | 126.0275 | 160.9725 | 136.9931 | 165.9897 | 166.914 | 214.945 | 138.9759 | 95.0206 | 90.9484 | 90.0206 | 126.9519 | 201.9931 | 209 | 235.9966 | 333.1066 | 203.0103 | 267.0309 | 311.0722 | 370.1341 | 153.0069 | 62.9862 | 75.9519 | 153.0069 | 98.0447 | 100.0172 | 127.9656 | 120.0069 | 110.9519 | 94.0516 | 101.9897 | 116.0309 | 74.0413 | 63.0275 | 66.9691 | 104.9587 | 85.1135 | 96.0138 | 97.9966 | 82.9828 | 222.9175 | 197.9484 | 233.9312 | 468.8969 | 140.055 | 180.9381 | 182.0172 | 392.1306 | 264.89 | 201.9484 | 229.0584 | 422.9519
1873 | 22.5974 | 153.9292 | 175.301 | 179.1948 | 200.6105 | 140.8229 | 137.0531 | 177.7344 | 167.7875 | 135.2036 | 121 | 122.9734 | 137.2036 | 193 | 154.7875 | 142.0797 | 192.1416 | 155.7521 | 153.6901 | 132.062 | 194.7079 | 133.938 | 301.7964 | 309.9646 | 139.1948 | 191.9557 | 187.8584 | 244.8406 | 315.9115 | 122.8672 | 106.8052 | 115.9026 | 130.8141 | 130.7521 | 120.0443 | 150.9469 | 94.9734 | 189.3541 | 219.8672 | 256.0177 | 251.2036 | 127.9469 | 120.0443 | 111.1062 | 111.0177 | 69.0089 | 67.0885 | 74.2302 | 89.2036 | 162.1151 | 187.938 | 203.1594 | 129.0797 | 118.9911 | 122.0354 | 118.9026 | 115.0797 | 158.0708 | 177.6901 | 210.0177 | 145.2213 | 150.699 | 131.9469 | 170.8495 | 129.0885 | 186.7875 | 174.938 | 227.7167 | 260.7698 | 123.0266 | 121.2479 | 212.8052 | 196.9292 | 176.9911 | 184.8672 | 181.1948 | 386.3891 | 149.8584 | 145.8938 | 144.0885 | 217.1505 | 109.1062 | 121.0797 | 148.0443 | 147.0797 | 161.9115 | 160.7521 | 192.9469 | 132.9646 | 106.9646 | 112.8938 | 100.9469 | 104.0885 | 179.2125 | 209 | 218.1328 | 366.8406 | 214.0354 | 224.2036 | 329.9203 | 356.8672 | 150.239 | 69.1771 | 90.9203 | 150.239 | 113.8938 | 118.7964 | 145.7521 | 126.1151 | 103.0354 | 114.8406 | 119.8761 | 106.0708 | 74.0443 | 73.0177 | 77 | 94.1062 | 97.9911 | 105.8141 | 119.8141 | 76.1859 | 218 | 184.2479 | 228.8584 | 446.0177 | 141.1771 | 184.8141 | 197.8052 | 439.5574 | 248.8406 | 200.8495 | 260.6547 | 358.1859
1874 | 22.6084 | 146.2537 | 208.3022 | 201.1057 | 157.5286 | 121.148 | 142.8097 | 148.4652 | 143.8731 | 157.6194 | 120.7886 | 120.3383 | 160.1903 | 192.556 | 131 | 151.3383 | 208.4863 | 128.3594 | 119.1269 | 139.2537 | 161.6194 | 127.7612 | 279.5075 | 306.3172 | 160.7251 | 186.8943 | 172.5075 | 227 | 305.6828 | 107.6828 | 85.6132 | 105.2749 | 110.4652 | 103.6132 | 124.8731 | 144.6617 | 91.7674 | 228.8308 | 206.2898 | 257.4925 | 273.2388 | 122.2326 | 125.2114 | 123.148 | 113 | 70.0211 | 76.704 | 99.1331 | 111.7463 | 175.3594 | 181 | 221.4863 | 138.6766 | 117.9577 | 125.6828 | 108.1903 | 124.3594 | 165.4714 | 143.2326 | 211.4714 | 169.9154 | 117.2326 | 126.3383 | 154.0423 | 139.0846 | 163.2537 | 168.4229 | 196.0423 | 235.6766 | 126.5286 | 148.4925 | 191.6766 | 189.6132 | 176.4863 | 170.1692 | 202.8943 | 318.4017 | 134.1269 | 133.9577 | 153.5771 | 233.8308 | 121.1057 | 129.704 | 152.3445 | 155.8943 | 151.9366 | 133.8669 | 187.3172 | 128.7251 | 103.4652 | 100.8731 | 95.2326 | 114.0211 | 203.1903 | 208.4503 | 232.7463 | 349.0634 | 218.2326 | 247.1903 | 321 | 342.5075 | 176.6194 | 88.5771 | 81.4714 | 176.6194 | 102.1269 | 96.1269 | 117.9789 | 138.556 | 106.9366 | 96.7886 | 106.2537 | 113.4714 | 78.4291 | 74.7886 | 77.148 | 105.556 | 96.7886 | 85.3806 | 99.1903 | 96.4291 | 218.2326 | 211.6194 | 213.444 | 448.592 | 160.5348 | 164.4229 | 176.6555 | 390.2326 | 231.2326 | 183.8308 | 221.6406 | 378.6617
1875 | 22.6194 | 157.9331 | 176.8359 | 205.6656 | 182.0669 | 128.535 | 133.6991 | 169.2644 | 139.1034 | 139.0973 | 111.1003 | 135.8663 | 168.1976 | 172.7021 | 131.234 | 167.4681 | 229.4286 | 145.7021 | 125.234 | 150.231 | 144.6687 | 162.1641 | 302.4985 | 319.7963 | 148.0669 | 181.5319 | 196.6353 | 226.1976 | 291.6353 | 93.5015 | 113.0973 | 117.5653 | 131.6656 | 131.4316 | 119.0334 | 129.0334 | 81.4347 | 220.6656 | 264.1611 | 233.3647 | 239.4711 | 132.5653 | 134.3982 | 130.5015 | 112.8663 | 70.8328 | 63.6353 | 59.7021 | 100.3344 | 192.6353 | 181.0669 | 242.8966 | 170.0669 | 115.8663 | 111.1337 | 117.6018 | 140.0304 | 140.9666 | 153.231 | 187.535 | 166.0669 | 127.7325 | 141.5988 | 156.535 | 143.4681 | 174.1976 | 187.231 | 199.0031 | 265.462 | 150.9331 | 125.8024 | 222.3647 | 217.6656 | 198.3647 | 177.2979 | 197.9666 | 336.7994 | 139.7994 | 132.0334 | 134.6353 | 226.535 | 125.5653 | 116.9027 | 123.0031 | 150.7994 | 149 | 173.3982 | 200.6292 | 116.5684 | 124.1976 | 95.2675 | 106 | 114.8663 | 212.3344 | 183.7021 | 220.5653 | 351.2979 | 228.3313 | 254.6292 | 321.8359 | 363.5592 | 159.1337 | 68.7325 | 57.234 | 159.1337 | 108.1672 | 101.8997 | 116.8328 | 117.8663 | 103.9666 | 87.3344 | 117.3647 | 89.7356 | 53.0031 | 64.8663 | 83.6322 | 84.9331 | 86.6991 | 102.7994 | 107.5319 | 70.3009 | 229.3009 | 194.2675 | 233.3313 | 474.696 | 138.465 | 183.5319 | 205.8632 | 400.9331 | 242.1337 | 176.9027 | 205.4012 | 363.6353
1876 | 22.6304 | 155.8628 | 198.8966 | 196.6859 | 184.2286 | 142.5825 | 127.1949 | 148.1372 | 169.7197 | 114.8748 | 114.0457 | 131.7256 | 146.3261 | 192.6342 | 138.0915 | 180.1769 | 184.8688 | 165.5427 | 131.7714 | 128.0915 | 164.8688 | 137.2227 | 288.2286 | 286.6461 | 150.7773 | 168.2286 | 214.1769 | 203.7773 | 311.006 | 107.497 | 87.3658 | 105.7773 | 122.7316 | 115.503 | 120.3658 | 128.4453 | 93.9085 | 211.2286 | 210.9085 | 215.6402 | 281.5427 | 120.2744 | 116.6342 | 144.4513 | 109.503 | 65.3598 | 80.6739 | 79.0855 | 110.7316 | 209.5825 | 182.0855 | 211.5944 | 171.3598 | 112.3658 | 115.4115 | 134.7714 | 112.5944 | 140.0915 | 132.3718 | 202.3141 | 168 | 120.6859 | 129.5427 | 171.4056 | 156.1769 | 151 | 164.6342 | 226.5825 | 221.3658 | 148.8628 | 149 | 205.6461 | 206.994 | 180.8688 | 157.7773 | 196.5427 | 331.3201 | 134.4115 | 133.9145 | 152.2227 | 241.497 | 114.4632 | 141.0795 | 150.9026 | 144.9543 | 149.0915 | 156.1372 | 161.8231 | 132.9543 | 101.5487 | 102.1312 | 106.2286 | 111.2744 | 220.9026 | 203.5885 | 208.5487 | 331.7316 | 208.8171 | 213.9026 | 343.0735 | 294.2803 | 163.503 | 61.0457 | 64.8231 | 163.503 | 112.497 | 99.3658 | 111.8171 | 114.5944 | 102.3598 | 96.3598 | 99.7316 | 110.2227 | 81.7714 | 61.6859 | 72.3141 | 84.0974 | 78.4115 | 95.8568 | 93.9543 | 78.7256 | 236.6739 | 201.9085 | 214.5487 | 439.1949 | 123.503 | 169.7714 | 174.3718 | 399.8231 | 248.4235 | 202.4056 | 216.2684 | 383.0974
1877 | 22.6414 | 153.4061 | 157.3207 | 210.942 | 188.6519 | 112.884 | 170.2732 | 151 | 143.6962 | 153.1878 | 115.3481 | 126.4061 | 173.7679 | 184.7099 | 140.6382 | 164.6245 | 203.4641 | 156.058 | 127.4061 | 131.1603 | 183.7542 | 121.9283 | 292.942 | 320.6519 | 167.1741 | 172.2458 | 196.5359 | 218.6076 | 329.5053 | 97 | 95.9283 | 121.1297 | 138.8122 | 125.942 | 127.3618 | 97.0443 | 92.9283 | 216 | 209.1741 | 229.9863 | 272.2901 | 126.3481 | 109.9283 | 133.3481 | 119.7099 | 53.3344 | 53.6962 | 61.5084 | 124.2595 | 178.8977 | 164.5084 | 222.3175 | 157.5359 | 120.4641 | 123.8259 | 130.1741 | 124.6519 | 143.1603 | 159.7236 | 189.5084 | 168.116 | 133.3175 | 120.6962 | 158.1878 | 138.5939 | 151.8703 | 155.3175 | 198.9726 | 229.5802 | 145.1878 | 148.3038 | 238.0274 | 187.0443 | 198.5359 | 174.4061 | 185.8977 | 338.1741 | 142.7099 | 152.4198 | 136.116 | 229.4916 | 143.4335 | 102.2184 | 128.4641 | 144 | 152.1603 | 159.6382 | 177.2595 | 131.942 | 112.3038 | 83.5939 | 110.0137 | 115.8397 | 197.4778 | 195.116 | 219.1297 | 347.5222 | 206.9146 | 193.4367 | 283.3344 | 320.5359 | 172.2015 | 61.7099 | 81.0717 | 172.2015 | 102.4061 | 106.0137 | 108 | 125.7236 | 89.058 | 85.3207 | 114.884 | 93.942 | 76.8259 | 76.4061 | 59.0443 | 104.6793 | 87.8122 | 72.6382 | 92.4778 | 73.5222 | 208.884 | 199.5939 | 225.7679 | 483.7236 | 135.4504 | 165.9863 | 202.1297 | 417.2901 | 298.5939 | 190 | 201.5222 | 405.1878
1878 | 22.6524 | 159.8594 | 192.9611 | 208.5235 | 184.1952 | 111.9843 | 127.7577 | 152.828 | 154.8594 | 140 | 120.0157 | 130.8908 | 169.086 | 181.3359 | 149.3126 | 187.6955 | 212.4765 | 157.5625 | 131.6798 | 148.1017 | 193.0471 | 136.086 | 292.4218 | 315.2812 | 167.0471 | 163.3045 | 190.7577 | 198.0389 | 291.6717 | 98.6874 | 111.6328 | 107.0703 | 150.3126 | 125.6328 | 116.8594 | 115.1952 | 106.0314 | 215.2266 | 212.4218 | 244.5235 | 277.2812 | 130.9454 | 123.8048 | 138.2266 | 114.9297 | 74.4375 | 66.2655 | 84.3829 | 96.4218 | 162.1249 | 186.6798 | 195.7031 | 148.8048 | 127.2969 | 121.2109 | 132.2969 | 119 | 160.0314 | 139.3515 | 213.7188 | 169.1563 | 106.914 | 132 | 145.2812 | 131.7891 | 164.8751 | 130.1796 | 227.9767 | 240.3359 | 131.5782 | 137.5625 | 206.7734 | 203.5078 | 194.0233 | 181.0703 | 169.828 | 342.828 | 138.2109 | 140.6095 | 137.2266 | 205.4218 | 118.5625 | 121.7188 | 135.6485 | 144.4922 | 169.5235 | 169.9297 | 151.2499 | 130.2266 | 102.1952 | 78.6171 | 94.3515 | 98.4062 | 189.9843 | 195.5938 | 204.7891 | 352.9064 | 236.4532 | 230.8908 | 303.6641 | 313.7031 | 144.7577 | 57.7031 | 67.0546 | 144.7577 | 110.1952 | 90.2812 | 109.6171 | 105.2109 | 90.914 | 120.9611 | 112.5782 | 93.4218 | 74.2812 | 82.1563 | 75.7188 | 68.9686 | 100.9297 | 83.2109 | 84.4922 | 81.2266 | 208.4765 | 192.5078 | 220.5235 | 462.4375 | 158.0157 | 181.0157 | 188.7734 | 422 | 292.2812 | 189.7891 | 210.7734 | 392.9843
1879 | 22.6634 | 156.1828 | 167.4868 | 189.4956 | 196.696 | 125.0826 | 149.348 | 175.5132 | 152.2566 | 139.8348 | 106.0088 | 104.239 | 156.7522 | 196.522 | 128.1564 | 146.8084 | 228.6134 | 164.7522 | 103.7258 | 123.1652 | 156.1476 | 125.3216 | 297.6696 | 321.3954 | 132.13 | 206.174 | 210.5308 | 225.1652 | 326.5044 | 120.3392 | 118.0176 | 107.3392 | 129.8998 | 135.239 | 115.2478 | 130.3392 | 81.8172 | 204.7522 | 215.3568 | 226.239 | 277.6134 | 116.8348 | 108 | 127.5044 | 113.8348 | 67.826 | 81.9262 | 64.9824 | 102.5782 | 178.5694 | 158.065 | 205.5782 | 135.9736 | 118.2478 | 125.0738 | 122.3392 | 119.7434 | 134.5782 | 146.7258 | 209.2566 | 157.174 | 120.8172 | 130.761 | 148.587 | 129.4956 | 150.6608 | 159.413 | 187.3392 | 255.7698 | 126.7434 | 143.348 | 215.0176 | 196.3392 | 229.5396 | 180.0176 | 192.5958 | 366.8348 | 143.065 | 110.652 | 128.652 | 211.3304 | 125.587 | 116.9262 | 132.1564 | 149.9262 | 151.8998 | 167.1828 | 182.7346 | 121.4042 | 116.1828 | 100.3304 | 99.5782 | 115.5958 | 201.8436 | 177.826 | 202.9912 | 316.2126 | 217.9824 | 204.8172 | 287.3216 | 321.348 | 168.0826 | 66.4218 | 80.174 | 168.0826 | 124.0176 | 94.5782 | 130.5044 | 109.0738 | 102.2566 | 94.5782 | 106.7522 | 100.4868 | 76.1828 | 70.174 | 71.587 | 94.9174 | 99.7522 | 86.1652 | 91.9086 | 72.0738 | 226.9262 | 185.6696 | 202.1564 | 455.1652 | 144.8348 | 168.826 | 198.9174 | 420.0176 | 292.8612 | 188.0738 | 221.2478 | 404.1002
1880 | 22.6744 | 137.3285 | 186.3723 | 196.9927 | 160.7591 | 126.9489 | 134.4161 | 159.3796 | 144.3796 | 137.0511 | 96.0438 | 117.0511 | 153.6204 | 169.9489 | 140.5255 | 175.1533 | 193.1825 | 160.3869 | 134.9489 | 124.5255 | 178.1971 | 137.8175 | 296.3723 | 343.3504 | 177.2409 | 196.0511 | 173.708 | 226.6204 | 317.5839 | 113 | 95.4307 | 99.7153 | 148.438 | 146.2482 | 117.7153 | 123.7591 | 100.3869 | 202.854 | 189.5109 | 237.3431 | 242.2774 | 117.4671 | 107.8102 | 122.0949 | 112.7591 | 76.0511 | 69.2409 | 84.6277 | 107.5766 | 194.1022 | 180.9051 | 212.9489 | 169 | 119.8613 | 138.1387 | 114.9051 | 127.4307 | 141.4745 | 176.6204 | 200.5255 | 149.8978 | 142.2336 | 117.5693 | 144.9489 | 136.5182 | 157.146 | 164 | 180.1898 | 230.4307 | 137.6569 | 126.8978 | 194.2336 | 190.2336 | 181.2336 | 158.2847 | 176.5255 | 361.7737 | 163.1533 | 129.3796 | 149.1825 | 217.8467 | 119.7664 | 107.2774 | 142.6277 | 139.3285 | 173.2847 | 150.7956 | 201.4307 | 136.5255 | 97.8978 | 104.3796 | 106.6642 | 102.3723 | 190.9927 | 189.7518 | 213.0511 | 359.1095 | 239.4307 | 222.438 | 300.1971 | 306.1314 | 167.3869 | 61.4234 | 70.4307 | 167.3869 | 102.5693 | 99.8613 | 124.4307 | 121.4745 | 94.5693 | 102.4234 | 104.3796 | 115.7664 | 57.4234 | 60.9051 | 66.8102 | 94.0949 | 96.8102 | 87.146 | 103.5182 | 84.4745 | 214.8102 | 183.708 | 215.0949 | 458.9927 | 141.3869 | 177.8102 | 196.7664 | 398 | 258 | 198.7664 | 224.0949 | 384.0438
1881 | 22.6854 | 152.2508 | 203.2122 | 213.3205 | 168.2144 | 136.5359 | 164.3205 | 162.2497 | 147.5713 | 131.4298 | 107.179 | 111.2154 | 152.4652 | 178.8928 | 133.8564 | 148.8575 | 214.1072 | 145 | 145.0718 | 122.1436 | 160.5713 | 120.3226 | 317.7138 | 300.8221 | 169.0354 | 185.4995 | 189.8928 | 221.1779 | 285.8575 | 110.9636 | 91.3934 | 97.4287 | 125.5005 | 119.4641 | 116.3934 | 131.4287 | 86.3934 | 208.8564 | 222.8928 | 212.4287 | 265.179 | 140.0354 | 108.5724 | 124.9293 | 121.2862 | 68.8221 | 62.7503 | 64.1436 | 95.179 | 175.0354 | 178.9282 | 219.8564 | 167.0707 | 110.2862 | 144.6056 | 116.358 | 121.3569 | 146.4287 | 168.9271 | 192.8917 | 167.8928 | 150.463 | 123.1072 | 153.4641 | 148.9636 | 149.4287 | 164.3215 | 184.2508 | 226.2862 | 159.1415 | 144.6785 | 205.7856 | 202.3215 | 190.7492 | 165.6088 | 170.7138 | 332.7149 | 136.2144 | 136.1083 | 165.6056 | 242.3205 | 111.1083 | 127.0707 | 121.0718 | 149.7492 | 176.4287 | 183.8917 | 197.9293 | 133.179 | 115.2497 | 107.5713 | 113.1072 | 123.0707 | 208.4995 | 216.5359 | 205.3934 | 326.9303 | 232.0707 | 198.4287 | 281.3923 | 335.2497 | 155.537 | 74.7856 | 66.179 | 155.537 | 108.1072 | 88.5713 | 118.8928 | 126.1072 | 98.821 | 114.1779 | 106.6066 | 104.7503 | 69.9282 | 60.4287 | 64.4641 | 95.5359 | 93.6066 | 79.4287 | 116.2851 | 87.7138 | 216.1083 | 198.6066 | 214.6066 | 473.284 | 130.5016 | 175.7856 | 185.9646 | 396.0707 | 259.9293 | 186.2497 | 223.6066 | 395.0718
1882 | 22.6964 | 166.46 | 160.9736 | 193.8673 | 170.9558 | 138.8495 | 145.2258 | 153.8495 | 145.3142 | 156.7742 | 114.969 | 135.3716 | 171.6858 | 177.5221 | 117.5531 | 157.3142 | 216.7921 | 145.8363 | 151.3716 | 141.9116 | 159.9868 | 150.9868 | 303.4158 | 311.46 | 160.0442 | 176.9426 | 189.4779 | 207.5531 | 293.8363 | 97.2258 | 104.3142 | 101.1195 | 137.7611 | 116.4337 | 126.0884 | 137.6284 | 95.73 | 192.6726 | 219.3716 | 217.1947 | 274.4026 | 129.0132 | 127.73 | 138.1327 | 131.5221 | 83.6416 | 69.9558 | 81.5221 | 104.4026 | 170.9426 | 170.9558 | 200.8053 | 151.8363 | 119.8053 | 108.9558 | 134.6858 | 118.031 | 150.9558 | 133.8053 | 171.301 | 164.0132 | 120.6416 | 125.9116 | 150.1947 | 133.5531 | 153.5974 | 168.9116 | 202.2832 | 240.1063 | 127.9558 | 141.7611 | 203.8805 | 203.6858 | 172.8363 | 202.0884 | 160.9558 | 343.8937 | 140.031 | 156.5043 | 134.0178 | 220.4779 | 137.1195 | 113.27 | 129.2832 | 131.9558 | 178.5663 | 159.4337 | 211.0132 | 141.8053 | 109.5974 | 104.9558 | 110.6548 | 107.5974 | 195.8805 | 217.0574 | 217.5974 | 364.1769 | 214.6858 | 202.4779 | 268.3584 | 327.6858 | 183.5663 | 73.3584 | 75.6416 | 183.5663 | 107.4469 | 88.2258 | 116.4469 | 123.1769 | 89.7168 | 98.8805 | 95.1195 | 112.7921 | 61.9558 | 64.9558 | 61.9116 | 98.5663 | 82.9558 | 85.8673 | 101.6416 | 77 | 237.699 | 187.8363 | 203.1195 | 436.5089 | 166.4469 | 173.1637 | 192.3274 | 381.9116 | 275.2832 | 178.8053 | 213.1947 | 406.6284
1883 | 22.7074 | 134.9765 | 202.233 | 211.8377 | 180.24 | 126.5576 | 165.3106 | 141.2941 | 154.0777 | 132.7365 | 105.4023 | 116.7906 | 161.4729 | 177.9529 | 130.9765 | 167.2635 | 228.1553 | 151.3412 | 134.1082 | 151.52 | 179.8918 | 171.2871 | 280.2941 | 285.6423 | 152.3412 | 200.3342 | 194.1859 | 217.2871 | 302.6353 | 119.4188 | 114.2635 | 104.1082 | 133.76 | 125.8141 | 116.7435 | 152.52 | 80.8447 | 202.8918 | 203.6894 | 224.5506 | 268.6824 | 108.7129 | 110.0777 | 118.5812 | 126.6824 | 79.6824 | 75.9459 | 75.1012 | 99.6047 | 196.8377 | 177.2094 | 192.1318 | 157.0777 | 112.5812 | 113.6283 | 123.4188 | 133 | 157.3412 | 128.2941 | 200.4965 | 146.8752 | 121.0306 | 137.76 | 157.4188 | 143.9459 | 154.7059 | 180.233 | 195.1553 | 259.4424 | 135.1318 | 140.5271 | 199.8377 | 196.767 | 178.2094 | 184.3106 | 170.1082 | 331.9364 | 153.1553 | 119.4259 | 167.0471 | 223.6047 | 134.3106 | 128.9459 | 123.6047 | 142.6894 | 167.0777 | 168.8141 | 186.1012 | 133.9223 | 114.7906 | 112 | 87.0541 | 109.233 | 194.3412 | 189.5812 | 222.1318 | 333.76 | 203.0235 | 204.8141 | 273.5035 | 318 | 171.5506 | 75.6047 | 73.6588 | 171.5506 | 97.4494 | 113.5812 | 109.7435 | 100.6658 | 96.1859 | 99.4494 | 97.0541 | 122.9694 | 68.3412 | 70.9459 | 74.8141 | 89.5812 | 88.8141 | 105.2871 | 100.0541 | 73.9694 | 207.0541 | 195.5812 | 204.6588 | 453.8141 | 155.1318 | 169.3717 | 177.1012 | 397.1859 | 266.9694 | 175.1388 | 220.5506 | 423.1012
1884 | 22.7184 | 148.5762 | 181.407 | 193.0168 | 192.1189 | 146.5427 | 144.4573 | 160.8392 | 148.5762 | 129.9916 | 139.4154 | 119.407 | 156.4154 | 200.5427 | 144 | 171.4489 | 218.1608 | 145.4154 | 144.9749 | 122.1441 | 166.2881 | 157.6951 | 302.1441 | 336.1441 | 145.407 | 168.33 | 204.1608 | 204.9916 | 316.1105 | 110.1524 | 116.2881 | 116.5595 | 117.9916 | 118 | 141.9497 | 123.5762 | 91.8476 | 194.7622 | 223.1021 | 214.8559 | 257.6951 | 121.7287 | 105.0084 | 127.2713 | 119.2965 | 70.4238 | 68.8559 | 57.8727 | 98.4405 | 175.7119 | 171.8559 | 193.5762 | 153.0168 | 121.8476 | 100.8811 | 111.9916 | 132.2797 | 156.0251 | 148.9916 | 181.5511 | 173.0938 | 140.2797 | 122.5678 | 149.3049 | 134.5511 | 139.9146 | 160.9916 | 185.3049 | 240.7538 | 135.8559 | 144.2881 | 181.4489 | 209.9581 | 177.0335 | 166.3384 | 184.8643 | 402.4154 | 141.8643 | 138.9749 | 142.4405 | 220.8559 | 113.7454 | 122.7203 | 121.7203 | 160.9497 | 163.0168 | 166.186 | 168.1524 | 137.263 | 116.8308 | 109.6951 | 92.9916 | 96.0419 | 191.1524 | 200.5762 | 221.7035 | 324.9062 | 191.7287 | 195.1273 | 288.1273 | 317.2797 | 162.2881 | 73.2881 | 80.3049 | 162.2881 | 106.4238 | 124.2881 | 140.2797 | 134.3986 | 103.2797 | 107.9916 | 102.8476 | 103.8643 | 64.7203 | 64.1441 | 66.1357 | 98.8476 | 84.0251 | 92.9916 | 106.2797 | 56.3049 | 214.4322 | 207.5846 | 211.737 | 443.6951 | 151.9665 | 184.4238 | 160.1608 | 400.9665 | 249.593 | 208.5678 | 205.3819 | 403.1357
1885 | 22.7294 | 149.6548 | 170.534 | 204.0619 | 176.7198 | 126.9381 | 161.0914 | 150.3452 | 147.935 | 125.7198 | 128.9056 | 106.1888 | 153.0975 | 183.5959 | 143.6873 | 179.2771 | 228.8112 | 138.0325 | 128.5635 | 126.9086 | 169.2508 | 138.6842 | 300.0294 | 328.7137 | 137.6609 | 197.4985 | 214.0294 | 205.0975 | 298.3158 | 116.0619 | 114.8731 | 109.0944 | 113.5635 | 119.8761 | 111.1594 | 124.0294 | 84.8437 | 224.7787 | 193.531 | 216.6579 | 248.8467 | 129.4985 | 111.9381 | 117 | 127.3127 | 66.6873 | 68.9381 | 69.9381 | 107.3127 | 174.9381 | 173.1888 | 202.1594 | 164.2183 | 116.5635 | 117.0619 | 107.5635 | 125.9675 | 170.0914 | 145.6579 | 171.2213 | 144.0975 | 133.8112 | 121.0944 | 161.7492 | 122.9706 | 178.4985 | 155.7817 | 197.4365 | 261.1827 | 135 | 142.7167 | 197.0944 | 184.2213 | 198.3421 | 196.9644 | 192.9706 | 384.0883 | 145.5929 | 126.4721 | 149.1239 | 217.4985 | 136 | 124.8112 | 125.2183 | 130.0031 | 174.0619 | 190.5898 | 172.6548 | 123.6579 | 97.3746 | 96.7817 | 88.7198 | 122.7167 | 196.4365 | 202.4365 | 210.8731 | 358.7492 | 204.5015 | 183.531 | 275.2802 | 313.7817 | 160.0914 | 73.9056 | 91.4985 | 160.0914 | 104.5635 | 123.9675 | 132.2477 | 114.2508 | 99 | 103.0944 | 98.3452 | 108.2183 | 69 | 66.0944 | 62.2508 | 90.5929 | 102.6254 | 85.7492 | 101.6873 | 72.8142 | 217.7817 | 214.1858 | 220.9644 | 434.8467 | 134.8792 | 181.6254 | 174.8761 | 386.8498 | 267.5015 | 204.2802 | 174.8142 | 391.7462
1886 | 22.7404 | 136.4941 | 193.6509 | 197.3136 | 182.4586 | 133.6864 | 140.8432 | 162.8314 | 129.2041 | 134.4941 | 120.1332 | 120.6982 | 183.1332 | 197.435 | 142.1686 | 148.0591 | 216.6627 | 148.4941 | 135.145 | 146.4823 | 174.9882 | 117.9259 | 284.3373 | 287.8786 | 174.9527 | 185.6864 | 203.9023 | 231.4232 | 327.1686 | 111.8432 | 98.8432 | 115.3373 | 119.8077 | 130.5059 | 136.8077 | 107.3255 | 85.0118 | 202.8432 | 193.8668 | 233.0236 | 277.6982 | 117.3491 | 117.5059 | 118.6864 | 126.4705 | 66.5177 | 73.4941 | 76.6864 | 107.9882 | 182.3609 | 185.1686 | 220.2191 | 163.0355 | 121.1214 | 113.3491 | 118.5295 | 112.9764 | 154.565 | 159.6627 | 190.1214 | 173.6273 | 118.29 | 126.1568 | 154.4941 | 142.0355 | 164.6627 | 160.5059 | 188.6627 | 228.4437 | 131.7959 | 129.0473 | 199.7959 | 204.9764 | 180.9645 | 159.1686 | 204.7841 | 351.1773 | 137.4941 | 152.9645 | 138.6627 | 206.3609 | 135.3255 | 115.5295 | 121.6745 | 157.1686 | 168.1568 | 158.1923 | 163.0355 | 134.4586 | 93.8314 | 98.1332 | 97.4941 | 104.4941 | 195.2514 | 195.0118 | 197.2041 | 353.6864 | 218.3373 | 186.7336 | 266.8432 | 320.3609 | 142.2041 | 68.5059 | 78.8432 | 142.2041 | 110.3018 | 109.29 | 112.1686 | 119.8195 | 96.9764 | 104.1095 | 108.3018 | 107.0355 | 68.6627 | 69.9764 | 70.855 | 82.3255 | 104.8195 | 83.7218 | 99.1568 | 88 | 215.9291 | 197.3136 | 186.71 | 459.3136 | 163.9054 | 186.5177 | 185.6745 | 432.58 | 279.8195 | 192.9764 | 190.1686 | 363.2277
1887 | 22.7514 | 134.1809 | 185.3717 | 190.8093 | 171.8753 | 142.1809 | 144.4572 | 164.8948 | 148.9804 | 135.6186 | 107.2665 | 136.7139 | 169.5428 | 176.1515 | 145.1711 | 166.4866 | 213.3717 | 146 | 126.5428 | 136.8289 | 169.4572 | 154.6479 | 287.2665 | 306.2763 | 157.533 | 194.9046 | 225.0392 | 200.6283 | 322.5722 | 114.9144 | 102.2763 | 112.4768 | 112.7995 | 130.1052 | 125.0954 | 106.3521 | 88.7335 | 208.6283 | 205.467 | 238.4768 | 279.2397 | 121.2861 | 119.2763 | 123.7433 | 114 | 75.4474 | 69.3717 | 83.7335 | 102.6381 | 188.2103 | 185.4572 | 180.3227 | 170.0392 | 105.9804 | 120.3619 | 130.2763 | 106.6186 | 178.0196 | 159.9902 | 174.4376 | 158.4474 | 105.6088 | 123.0856 | 152.9046 | 151.7531 | 163.1809 | 162.8191 | 188.4474 | 267.582 | 119.0758 | 147.5526 | 189.4278 | 192.6479 | 166 | 158.9144 | 181.2861 | 410.6211 | 138.0758 | 136.0098 | 138.2665 | 216.5526 | 131.8191 | 128.7237 | 126.0856 | 154.5624 | 162.9144 | 169.9046 | 181.0758 | 120.9804 | 92.2763 | 85.8093 | 97.3521 | 103.9902 | 228.6479 | 199.2763 | 210.467 | 354.401 | 212.7629 | 214.8191 | 269.0098 | 327.8387 | 157.6381 | 69.1907 | 78.6577 | 157.6381 | 97.3619 | 91 | 114.8093 | 113.8191 | 88.9902 | 81.8093 | 99.885 | 118.0196 | 62.8387 | 61.0856 | 78.0098 | 80.0856 | 101.1711 | 108.0856 | 97.7139 | 85.2861 | 191.6088 | 190.9902 | 202.6479 | 452.2665 | 129.2469 | 191.6479 | 190.6283 | 391.4768 | 272.3717 | 181.9144 | 194.4376 | 391.1711
1888 | 22.7624 | 137.7051 | 177.2271 | 200.1593 | 200.4882 | 138.556 | 140.261 | 172.7831 | 163.1356 | 150.261 | 112.2271 | 141.6577 | 165.6238 | 201.8746 | 150.556 | 137.3864 | 209.4779 | 150.444 | 125.1356 | 129.8983 | 169.8983 | 143.6136 | 297.2508 | 302.8068 | 166.8746 | 199.1932 | 185.1356 | 208 | 308.0474 | 108.4542 | 99.5797 | 96.0576 | 130.9661 | 120.0915 | 118.8746 | 117.3864 | 82.2271 | 214.261 | 194.3864 | 222.0576 | 228.5118 | 111.7051 | 117.1593 | 112.6712 | 115.5458 | 81.4203 | 64.1254 | 77.4203 | 102.3525 | 168.3762 | 184.739 | 217.2271 | 135.739 | 125.3525 | 126.444 | 126.6136 | 119.9085 | 160.5797 | 163.9763 | 190.1932 | 166.1593 | 134.8644 | 127.4203 | 150.6238 | 136.3084 | 160.9085 | 164.1254 | 194.0339 | 239.7966 | 133.5797 | 139.6475 | 212.4542 | 183.3525 | 171.2169 | 160.5694 | 175.3762 | 338.2949 | 151.8068 | 128.739 | 148.444 | 216.7627 | 126.9424 | 131.0339 | 135.8305 | 135.7153 | 159.3525 | 169.7492 | 189.7831 | 142.0915 | 88.8068 | 93.4203 | 106.0678 | 116.8644 | 219.9322 | 200.6373 | 203.444 | 329.4677 | 180.3864 | 210.1356 | 260.3864 | 314.4101 | 157.3525 | 63.7051 | 62.4779 | 157.3525 | 98.2271 | 93.1254 | 121.0678 | 112.6136 | 96.4542 | 89.8068 | 116.5899 | 102.8983 | 51.922 | 66.7729 | 69.5797 | 82.2949 | 109.4542 | 110.6814 | 105.7492 | 72.4203 | 212.1695 | 197.8746 | 195.2847 | 454.7153 | 150.2949 | 180.8068 | 197.0339 | 374.6712 | 260.1695 | 178.739 | 208.261 | 402.739
1889 | 22.7734 | 149.411 | 174.411 | 208.9047 | 181.9851 | 118.3559 | 141.6314 | 147.3008 | 152.1377 | 140.1229 | 110.4386 | 110.7945 | 138.589 | 193.6165 | 131.589 | 143.1102 | 221.9449 | 162.6292 | 114.7543 | 137.2733 | 181.7945 | 140.9725 | 306.3686 | 300.3559 | 159.6441 | 198.5614 | 176.1928 | 206.9725 | 341.7796 | 102.411 | 108.3708 | 112.178 | 135.411 | 131.1504 | 113.0826 | 118.589 | 78.3835 | 208.0275 | 199.4937 | 233.4512 | 243.9322 | 124.2331 | 118.9174 | 129.2331 | 123.8496 | 79.411 | 70.3284 | 71.5063 | 109.4386 | 190.4788 | 196.1102 | 212.5614 | 146.9047 | 127.9174 | 142.7394 | 124.178 | 108.2331 | 164.6441 | 145.2606 | 187.9174 | 165.4512 | 149.1504 | 122.5339 | 135.5085 | 162.7796 | 152.1102 | 174.8496 | 189.178 | 253.93 | 131.8898 | 137.0826 | 207.4386 | 189 | 194.4386 | 183.2733 | 197.6843 | 336.659 | 143.3961 | 136.2055 | 164.1229 | 243.1504 | 112.6716 | 129.6716 | 153.3284 | 126.1102 | 165 | 154.8771 | 165.1229 | 147.3961 | 87.589 | 95.7267 | 103.3433 | 130.9449 | 219.9851 | 219.3835 | 220.1504 | 370.2331 | 185.2882 | 193.3835 | 261.7945 | 324.8749 | 164.4386 | 72.7394 | 77 | 164.4386 | 97.6716 | 99.5339 | 111.7669 | 117.1653 | 96.9873 | 90.6441 | 96.0551 | 116.0275 | 81.7118 | 67.7394 | 69.7394 | 75.3157 | 101.9725 | 100.7945 | 89.2331 | 73.9047 | 199.3984 | 188.178 | 205.5063 | 441.6165 | 140.4386 | 180 | 188.4788 | 391.0275 | 247.6039 | 187.6441 | 202.2331 | 404.8623
1890 | 22.7844 | 153.178 | 171.644 | 218.7728 | 156.6628 | 113.7424 | 163.6744 | 133.8712 | 172 | 129.6534 | 111.8618 | 110.8712 | 135.0398 | 193.6042 | 135.2272 | 154.2084 | 210.2974 | 142.5738 | 138.7424 | 120.7424 | 184.9204 | 137 | 276.7026 | 301.623 | 161.8618 | 193.6534 | 197.0304 | 207.1382 | 306.9508 | 101.1686 | 127.1194 | 107.0398 | 134.6042 | 121.1686 | 123.4754 | 115.6932 | 77.5246 | 212.2178 | 211.911 | 213.9602 | 267.5738 | 127.0398 | 109.6136 | 129.644 | 127.0796 | 84.267 | 62.6136 | 59.5246 | 111.5152 | 175.3958 | 203.7236 | 204.1686 | 159.1686 | 120.356 | 133.1288 | 124.9204 | 116.267 | 164.6838 | 156.6136 | 175.3466 | 159.0283 | 138.5152 | 112.5644 | 171.5058 | 136.8806 | 161.466 | 179.6042 | 187.089 | 222.4052 | 117.089 | 144.2084 | 212.7822 | 187.911 | 193.9016 | 166.9602 | 178.822 | 376.3958 | 120.534 | 136.7822 | 152.5644 | 225.5456 | 124.7424 | 136.2974 | 140.6042 | 145.9204 | 165.2178 | 166.4356 | 158.1382 | 121.0492 | 87.5246 | 109.2084 | 117.6838 | 120.1686 | 192.4848 | 210.466 | 212.3466 | 369.3372 | 205.8408 | 189.2576 | 261.8712 | 291.5037 | 166.733 | 64.2178 | 74.822 | 166.733 | 111.0492 | 92.8314 | 108.3068 | 140.7822 | 117.9016 | 97 | 103.5644 | 119.7822 | 70.089 | 59.8712 | 66.2272 | 93.7424 | 103.8806 | 99.1288 | 93.3466 | 85.2974 | 227.3466 | 182.1686 | 191.3466 | 445.089 | 143.822 | 183.0492 | 175.1382 | 397.3958 | 276.7822 | 195.3068 | 201.555 | 385.2178
1891 | 22.7954 | 153.8672 | 162.9028 | 203.4514 | 177.2301 | 121.8407 | 128.5222 | 143.9028 | 172.4602 | 132.6903 | 108.0444 | 115.6106 | 138.8143 | 182.2389 | 152.8495 | 137 | 196.5398 | 160.1505 | 141.0884 | 127.6903 | 190.7167 | 141.8319 | 290.2301 | 325.7787 | 159.425 | 190.0176 | 174.3009 | 222.6903 | 292.5486 | 94.2213 | 105.5398 | 105.062 | 130.1417 | 114.6815 | 120.0708 | 116.5222 | 82.3097 | 211.8495 | 208.9204 | 221.4602 | 275.4866 | 121.1505 | 121.7611 | 115.6106 | 113.2301 | 89.3273 | 69.6991 | 62.469 | 101.7611 | 180.0884 | 172.8495 | 193.7699 | 154.5222 | 130.4778 | 132.0708 | 132.5574 | 125.2389 | 138.5486 | 161.1681 | 171.3894 | 200.1149 | 125.0796 | 118.3629 | 147.9028 | 152.4778 | 144.4426 | 167.7787 | 192.3806 | 250.6106 | 124.4514 | 133.2125 | 208.3185 | 186.3009 | 169.6991 | 178.6018 | 171 | 376.7167 | 143.0796 | 134.1593 | 157.2125 | 201.8055 | 131.6903 | 129.2125 | 128.5486 | 152.407 | 164.3894 | 167.7699 | 166.3273 | 130.6194 | 94.3806 | 91.3097 | 94.5398 | 110.4602 | 202.9292 | 194.1328 | 214.1417 | 354.9824 | 226.6371 | 192.894 | 258.3273 | 334.938 | 158.6815 | 66.3806 | 67.2301 | 158.6815 | 119.9292 | 98.3982 | 115.531 | 137.6991 | 98.9912 | 98.3806 | 103.6106 | 118.5398 | 76.531 | 65.0708 | 80.1681 | 96.5486 | 118.5574 | 99.4514 | 94.4514 | 75.9116 | 223.6194 | 174.9912 | 186.0088 | 452.4514 | 137.3806 | 194.4602 | 188.3894 | 404.1593 | 277.8407 | 197.0088 | 186.3717 | 394.5134
1892 | 22.8064 | 134.8477 | 184.3033 | 210.6371 | 174.849 | 128.2424 | 147.4848 | 162.8795 | 163.5775 | 134.7576 | 140.6371 | 120.7576 | 166.759 | 173.9695 | 145.1219 | 140.3934 | 195 | 160.6066 | 128.9695 | 132.1814 | 174.1496 | 151.2133 | 289.0609 | 316.4543 | 188.8809 | 174.3629 | 192.1801 | 228.151 | 292.392 | 104.2729 | 104 | 119.3338 | 141.3338 | 122.8795 | 125.5457 | 135.2424 | 80.7271 | 210.9086 | 208.3643 | 225.1814 | 249.1814 | 126.9391 | 131.7271 | 118.3338 | 112.3033 | 67.9695 | 66.1205 | 56.4238 | 111.7271 | 169.9086 | 171.9086 | 188.3947 | 173.2424 | 116.1205 | 132.2133 | 117.5443 | 120.8477 | 129.9086 | 145.2424 | 177.1496 | 162.09 | 126.6053 | 141.5457 | 165.1828 | 134.9695 | 162.1219 | 157.7271 | 198.2119 | 256.4848 | 138.6662 | 145.2742 | 199.6357 | 194.9695 | 162.2424 | 194.2424 | 168.8186 | 358.4529 | 138.5457 | 128.2424 | 168.3047 | 244.6066 | 128.91 | 140.5471 | 120.151 | 132.8172 | 158.0305 | 169.9086 | 157.8158 | 128.9086 | 94.8795 | 90.4543 | 94.2119 | 114.9086 | 198.6662 | 209.2742 | 221.2133 | 374.2729 | 205.151 | 227.6967 | 236.7271 | 317.2729 | 167.849 | 70.5152 | 70.1814 | 167.849 | 113 | 87.6053 | 124.4848 | 130.7271 | 102.9404 | 104.2119 | 106.3338 | 118.6967 | 82.5762 | 74.4238 | 65.9391 | 92.5139 | 100.151 | 110.2729 | 105.0305 | 87.3033 | 219.4848 | 185.9695 | 183.7563 | 461.0914 | 142.2424 | 192.6066 | 186.8781 | 398.7271 | 281.0914 | 188.4543 | 199.5457 | 421.5457
1893 | 22.8174 | 148.7266 | 177.472 | 191.9065 | 173.4042 | 124.1612 | 150.528 | 145.4159 | 137.6215 | 134 | 122.9252 | 118.2173 | 143.0374 | 181.3295 | 128.9252 | 153.0374 | 196.0187 | 159.6776 | 136.3295 | 136.1986 | 203.9065 | 132.5467 | 296.4977 | 320.7453 | 156.0935 | 189.0187 | 216.8691 | 230.6145 | 319.0864 | 104.0374 | 99.4159 | 112.2734 | 130.7079 | 114.3295 | 122.528 | 134.7266 | 85.8014 | 217.1986 | 200.0561 | 232 | 253.4533 | 130.7079 | 128.9065 | 108.472 | 109.0374 | 72.0187 | 75.9439 | 63.7453 | 111.9626 | 183.8388 | 181.5093 | 182.9136 | 173.236 | 122.3785 | 118.8949 | 132.3972 | 133.9626 | 142.5654 | 148.2921 | 201.5584 | 175.1986 | 139.9813 | 135.7266 | 147.764 | 140.8014 | 151.2734 | 164.8388 | 198.6892 | 257.4907 | 144.1986 | 122.0748 | 210.236 | 204.1122 | 161.2173 | 188.6331 | 168.3669 | 387.236 | 140.3669 | 129.764 | 142.3295 | 236.2921 | 113.2547 | 115.3295 | 127.7079 | 151.5093 | 162.8949 | 176.1986 | 201.2547 | 132.3972 | 89.6402 | 96.0187 | 102.3295 | 122.9813 | 209.2921 | 185.3108 | 201.528 | 366.9065 | 218.5654 | 226.6331 | 240.2734 | 318.5654 | 152.6518 | 70.2734 | 78.7827 | 152.6518 | 113 | 99.9626 | 119.8878 | 133 | 93.9323 | 103.6705 | 100.2921 | 126.2921 | 71.9439 | 78.1799 | 69.4533 | 114.2173 | 108.9813 | 97.3037 | 99.9626 | 78.9439 | 219.2173 | 189.764 | 206.9813 | 459.8949 | 144.7827 | 191.6776 | 193.3972 | 398.4533 | 272 | 195.0374 | 199.0748 | 419.5467
1894 | 22.8284 | 153.009 | 176.7375 | 179.6697 | 194.2625 | 121.2125 | 158.7375 | 138.9411 | 165.009 | 134.8009 | 115.0045 | 118.6063 | 150.8688 | 195.0679 | 117.8009 | 161.6697 | 206.475 | 180.7919 | 145.2625 | 129.3348 | 186.8643 | 149.3438 | 318.3803 | 314.9277 | 168.8643 | 186.1268 | 192.7964 | 210.3348 | 298.8732 | 104.9277 | 92.6741 | 113.3303 | 123.4661 | 122.4616 | 127 | 126.7286 | 93.1991 | 212.7375 | 212.7375 | 225.8598 | 253.475 | 124 | 116.9366 | 109.8732 | 115 | 76.3348 | 66.1991 | 68.8732 | 105.733 | 192.9277 | 175.258 | 201.525 | 169.9321 | 112.6107 | 140.3982 | 114.9321 | 124.7964 | 152.733 | 146.7241 | 175.5384 | 163.5295 | 132.7286 | 132.5339 | 151.0679 | 146.8643 | 152.8643 | 174.4616 | 192.2036 | 256 | 136.267 | 138.2714 | 203.9955 | 222.5339 | 162.1402 | 179.0768 | 188.0679 | 386.6018 | 163.5384 | 131.1991 | 152.3303 | 243 | 110.7964 | 131.4705 | 118.0634 | 155.4027 | 181.1946 | 169.3348 | 195.0589 | 120.5384 | 109.3303 | 98.733 | 113.3982 | 115.7286 | 213.0634 | 200.6063 | 204.3982 | 346.3937 | 230.0679 | 209.0679 | 248.2714 | 332.4705 | 137.5339 | 75.6018 | 79.4616 | 137.5339 | 110.3303 | 92.9321 | 99.3303 | 126.5928 | 118.9955 | 88.0634 | 98.1902 | 129.7964 | 68.3393 | 69.4027 | 67.8732 | 105.7964 | 105.733 | 74.8098 | 99.8732 | 68.3982 | 214.5339 | 191.1991 | 200.5295 | 478.9955 | 148.1312 | 212.525 | 179.1357 | 394.7375 | 279.475 | 197.2625 | 196.3169 | 432.3393
1895 | 22.8394 | 169.1356 | 188.3963 | 189.7926 | 177.2979 | 143.6223 | 154.7579 | 149.9016 | 182.2526 | 139.5133 | 124.3245 | 136.5133 | 163.5399 | 197.4415 | 124.7474 | 164.5319 | 230.3511 | 160.1622 | 136.3963 | 129.6489 | 188.8192 | 162.992 | 275.8644 | 301.2793 | 170.8192 | 177.5399 | 182.883 | 216.234 | 314.1622 | 93.7926 | 110.1622 | 107.117 | 125.0718 | 115.8644 | 128.3963 | 116.3963 | 87.0904 | 214.0638 | 210.4335 | 217.0984 | 270.0904 | 125.117 | 126.9548 | 122.3697 | 117.234 | 76.6489 | 64.2793 | 79.6941 | 106.117 | 186.5399 | 162.6569 | 188.8192 | 161.4149 | 133.4149 | 134.883 | 111.7207 | 116.2793 | 151.7207 | 127.6303 | 185.2074 | 159.5851 | 122.6755 | 130.0904 | 151.766 | 141.2793 | 151.4681 | 164.7926 | 196.8112 | 262.7021 | 137.5585 | 146.0904 | 194.3963 | 216.1808 | 176.4867 | 199.2526 | 195.1888 | 383.1808 | 161.1994 | 137.3777 | 145.5585 | 234.6223 | 107.8644 | 136.2526 | 113.0718 | 164.5133 | 170.5133 | 168.5319 | 182.1436 | 128.5319 | 108.1436 | 96.3245 | 108.1622 | 107.9096 | 201.0904 | 212.9282 | 206.9814 | 339.6569 | 234.3963 | 209.2074 | 251.9016 | 339.766 | 143.1888 | 75.5319 | 68.6755 | 143.1888 | 103.2793 | 95.3059 | 94.5133 | 118.774 | 109.117 | 85.3059 | 82.3777 | 122.117 | 75.7393 | 75.7207 | 78.9734 | 99.5133 | 104.4415 | 97.0452 | 110.1356 | 70.9814 | 214.3245 | 186.766 | 195.7474 | 464.0904 | 147.1888 | 190.3245 | 179.4149 | 405.8378 | 296.6489 | 194.2607 | 167.7288 | 441.9734
1896 | 22.8504 | 150.5014 | 184.1282 | 192.6268 | 171.3703 | 128.9971 | 129.9534 | 141.5365 | 163.2944 | 144.2507 | 115.0437 | 129.5888 | 181.2507 | 190.4606 | 141.0817 | 145.7113 | 233.4606 | 158.5831 | 136.5014 | 128.8718 | 215.414 | 137.2478 | 287.793 | 306.9563 | 184.0029 | 201.6648 | 188.4549 | 213.5451 | 308.2099 | 104.207 | 103.3352 | 108.5423 | 133.2915 | 134.207 | 131.4169 | 120.8746 | 80.207 | 197.414 | 187.038 | 227.5043 | 271.6619 | 128.2915 | 124.4198 | 116.7901 | 124.1662 | 68.2915 | 65.5831 | 68.0408 | 106.3761 | 203.0845 | 190.2478 | 209.5831 | 155.7464 | 112.586 | 126.1691 | 111.5831 | 118.1662 | 147.2099 | 142.5394 | 180.9155 | 169.0437 | 122.9183 | 124.6648 | 156.4957 | 141.1254 | 162.1254 | 170.2507 | 184.8338 | 276.793 | 137.8338 | 136.5831 | 199.4577 | 202.1225 | 171.4577 | 181.4606 | 203.6676 | 363.8746 | 131.7464 | 155.2099 | 154.5802 | 219.7056 | 122.4169 | 121.3761 | 122.1662 | 163.7113 | 170.8775 | 162.5394 | 190.4198 | 125.4549 | 120.2099 | 90.5423 | 101.9183 | 119.4577 | 200.3295 | 196.5451 | 225.8746 | 351.5043 | 235.3761 | 200.2507 | 221.0874 | 343.3295 | 151.9592 | 64.8746 | 71.2507 | 151.9592 | 106.3324 | 105.5014 | 103.0408 | 142.8338 | 110.2507 | 106.5802 | 98.7522 | 115.0874 | 60.207 | 73.2507 | 67.4577 | 100.7522 | 104.4577 | 94.3324 | 95 | 82.0029 | 216.414 | 190.6211 | 203.9183 | 451.6676 | 162.0817 | 179.586 | 164.7085 | 398.962 | 284.793 | 197.548 | 196.2507 | 424.6268
1897 | 22.8614 | 138.3539 | 163.7846 | 180.7846 | 192.6194 | 140.8347 | 156.8731 | 164.7079 | 148.2154 | 139.0885 | 109.3806 | 100.6461 | 173.9616 | 177.9115 | 147.6578 | 138.3305 | 221.5192 | 162.4307 | 124.9616 | 146.4808 | 219.758 | 150 | 290.3305 | 321.4307 | 162.1386 | 205.1002 | 210.2154 | 199.0768 | 302.0384 | 108.3539 | 100.2036 | 108.0384 | 128.2271 | 138.0501 | 133.0384 | 124.2154 | 82.5309 | 213 | 207.4808 | 213.5459 | 283.2889 | 124.4424 | 112.9883 | 122.3539 | 121.8347 | 68.3923 | 64.5693 | 76.0384 | 101.5192 | 193.5693 | 189.6311 | 204.9232 | 170.0117 | 103.9883 | 122.0267 | 111.177 | 120.6962 | 141.3422 | 150.8347 | 188.9115 | 164.8998 | 119.9883 | 132.6578 | 175.3156 | 142.6461 | 159.0885 | 166.3038 | 180.7846 | 273.5576 | 137.4307 | 140.4307 | 208.1653 | 212.7463 | 177.4307 | 178.9382 | 191.6194 | 363.2655 | 156.0384 | 144.177 | 167.5309 | 231.1386 | 125.2537 | 117.1269 | 130.4691 | 142.3539 | 159.9499 | 173.8731 | 176.2537 | 141.4424 | 107.3539 | 86.3923 | 93.5192 | 119.9616 | 215.404 | 191.1919 | 222.8347 | 324.7846 | 229.9115 | 197.5192 | 205.8113 | 355.9733 | 146.0885 | 69.4307 | 73.6844 | 146.0885 | 111.0885 | 98.5192 | 106.177 | 138.4808 | 106.6077 | 117.1002 | 89.6461 | 101.9382 | 68.6077 | 66.5693 | 70.0885 | 91.6461 | 106.7846 | 96.6578 | 94.6962 | 69.5576 | 226.2271 | 214.9115 | 206.1535 | 456.9382 | 169.8731 | 168.5576 | 168.8614 | 384.8998 | 277 | 165.1653 | 197.1653 | 414.9115
1898 | 22.8724 | 135.6903 | 164.1613 | 185.271 | 177.3677 | 126.471 | 153.2645 | 153.3742 | 144.3613 | 136.6839 | 125.5806 | 104.5742 | 161.6258 | 184.7419 | 142.8452 | 155.3613 | 216.7806 | 165.7871 | 123.6903 | 142.6839 | 194 | 149.6839 | 304.8323 | 316.8839 | 152.2645 | 198.3807 | 212.3677 | 211.4194 | 309.3161 | 98.7355 | 120.2645 | 117.8452 | 118.1613 | 124.2581 | 138.1032 | 129.8452 | 75.0581 | 214.2645 | 205.2645 | 238.1032 | 249.4258 | 120.0065 | 121.729 | 115.5806 | 111.5806 | 70.4774 | 66.271 | 82.3677 | 103.4194 | 186.4194 | 170.0645 | 195.7419 | 159.3742 | 123.7936 | 131.4064 | 109.8452 | 118.4194 | 150.8968 | 136.471 | 190.0516 | 183.6323 | 134.7355 | 128.4774 | 157.7871 | 148.7871 | 162.3742 | 174.9032 | 182.1097 | 273.2516 | 148.6903 | 148.2129 | 207.0387 | 197.3097 | 185.8452 | 201.471 | 178.8968 | 369.329 | 160.7871 | 144.7419 | 154.6839 | 218.4194 | 131.8387 | 124.529 | 141.4194 | 132.7355 | 167.4194 | 174.3742 | 187.2645 | 124.3613 | 96.1548 | 87.529 | 93.8387 | 115.529 | 197.0516 | 223.471 | 213.529 | 336.8581 | 228.8387 | 201.6323 | 227.8968 | 332.3677 | 156.6452 | 75.6323 | 83.9419 | 156.6452 | 106.471 | 99.1548 | 103.5806 | 131.5226 | 106.4194 | 100.2645 | 99.8968 | 119.4129 | 71.2645 | 70.1677 | 71.4774 | 95.8903 | 108.4258 | 98.1677 | 93.3677 | 78.7355 | 216.4774 | 207.5161 | 229.5742 | 471.2516 | 157.0968 | 179.9484 | 180 | 388.4581 | 275.7355 | 178.2645 | 208.6839 | 408.4645
1899 | 22.8834 | 150.6274 | 178.8821 | 208.2252 | 172.059 | 125.5979 | 157.9705 | 159.7453 | 139.1958 | 139.2842 | 115.2063 | 103.2105 | 154.5979 | 185.4758 | 147.3579 | 142.3137 | 204.9116 | 161.9853 | 130.4169 | 144.2989 | 190.7158 | 147.0295 | 290.9116 | 298.2842 | 156.9705 | 214.7895 | 210.0147 | 200.4463 | 307.3726 | 100.9263 | 115.4463 | 116.1179 | 125.3284 | 122.8673 | 124.8042 | 126.4758 | 91.2547 | 215.6863 | 205.0442 | 220.5348 | 262.941 | 135.9558 | 110.5979 | 118.6716 | 120.5831 | 78 | 79.7011 | 74.1032 | 101.9705 | 195.48 | 191.3726 | 208.2989 | 168.0442 | 130.6863 | 106.2695 | 118.6274 | 119.5979 | 154.6716 | 140.8526 | 184.3874 | 176.1327 | 135.6126 | 126.4758 | 149.7158 | 147.941 | 165.1327 | 183.1327 | 197.5684 | 258.8673 | 146.5495 | 158.2547 | 182.941 | 196.1663 | 189.3726 | 190.0884 | 177.0884 | 385.4905 | 151.7305 | 147.4463 | 151.3726 | 213.3579 | 121.3874 | 135.9116 | 135.0442 | 127.3726 | 170.24 | 179.1032 | 186.059 | 114.5979 | 83.76 | 100.8821 | 99.1516 | 114.76 | 195.3284 | 210.4463 | 222.2842 | 367.6863 | 229.8821 | 210.2252 | 224.4463 | 328.3726 | 177.4316 | 81.9263 | 71.0295 | 177.4316 | 105.9263 | 91.0295 | 107.9853 | 121.7011 | 103.3284 | 99.059 | 110.5684 | 109.5684 | 70.7158 | 86.059 | 75.059 | 96.1958 | 110.1179 | 116.0295 | 91.3432 | 78.9558 | 222.3579 | 187.9853 | 224.5979 | 459.8231 | 144.9263 | 182.6568 | 181.6421 | 372.1073 | 275.6274 | 189.2105 | 195.8484 | 406.7494
1900 | 22.8944 | 150.8893 | 195.6814 | 207.668 | 171.4957 | 137.7257 | 153.8229 | 142.5265 | 169.9249 | 147.3628 | 101.3099 | 116.2522 | 161.615 | 184.3984 | 137.1858 | 153.5178 | 212.9115 | 156.5043 | 116.9557 | 145.5929 | 186.385 | 140.9557 | 299.2522 | 308.4071 | 166.1107 | 197.3186 | 210.7257 | 188.4071 | 299.2743 | 105.8893 | 98.2964 | 103.0664 | 132.8142 | 140.6593 | 110.77 | 115.1771 | 99.9557 | 213.6814 | 201.385 | 199.4735 | 275.0885 | 136.2079 | 115.5708 | 116.2964 | 125.1858 | 87.5399 | 72.2743 | 64.77 | 107.7036 | 215.8893 | 180.2079 | 210.9557 | 166.4292 | 129.3628 | 121.8142 | 127.4071 | 128.3186 | 160.4735 | 153.8451 | 181.77 | 160.385 | 136.5265 | 118.5842 | 147.77 | 157.3628 | 144.9557 | 167.7257 | 199.4158 | 264.7344 | 128.2435 | 153.668 | 197.4735 | 203.668 | 177.8672 | 168.8008 | 176.2435 | 378.9913 | 147.0885 | 137.1107 | 147.0221 | 214.0885 | 123.1992 | 144.2522 | 123.2079 | 126.0885 | 169.0308 | 166.3628 | 174.9336 | 121.9557 | 84.9692 | 112.5708 | 112.7565 | 106.77 | 196 | 202.1549 | 229.3407 | 358.8672 | 239.8672 | 214.7787 | 216.1549 | 331.1771 | 168.4292 | 85.1858 | 63.5929 | 168.4292 | 119.7478 | 98.2435 | 111.3628 | 123.4735 | 104 | 92.7036 | 111.7344 | 126.0664 | 62.2964 | 75.2743 | 67 | 106.8229 | 99.1107 | 107.9115 | 95.4514 | 79.2079 | 221.7257 | 188.6372 | 221.053 | 463.2166 | 155.3407 | 181.615 | 178.5265 | 399.2522 | 287.1328 | 212.1328 | 179.5178 | 422.9423
1901 | 22.9054 | 130.7629 | 195.941 | 182.589 | 186.353 | 142.294 | 132.352 | 137.4131 | 179.5859 | 148.941 | 120.8809 | 117.8251 | 168.2961 | 206.5859 | 129.295 | 163.999 | 210.295 | 153.2961 | 119.707 | 140.412 | 183.5869 | 141.589 | 295.883 | 313.941 | 164.3509 | 199.883 | 207.5279 | 197.824 | 307.0611 | 92.823 | 108.0611 | 112.883 | 142.823 | 131.5279 | 126.001 | 133.471 | 99.53 | 213.941 | 210.942 | 212.353 | 280.176 | 126.412 | 115.8251 | 117.589 | 113.412 | 96.8219 | 63.117 | 75.059 | 103.2339 | 197.1749 | 179.2371 | 212.295 | 166.5279 | 125.999 | 126.1749 | 127.999 | 123.47 | 173.706 | 149.5311 | 181.1159 | 169.236 | 126.471 | 143.824 | 155.941 | 159.294 | 142.412 | 171.235 | 178.412 | 249.8251 | 144.352 | 131.4131 | 207.882 | 186.0021 | 174.1191 | 152.765 | 192.352 | 395.9979 | 151.47 | 140.999 | 145.117 | 225.53 | 133.5869 | 133.823 | 126.8261 | 137.53 | 156.4731 | 173.942 | 174.648 | 130.0021 | 102.4689 | 101.176 | 102.6491 | 123.4131 | 209.7671 | 210.352 | 219.0569 | 347.353 | 238.5901 | 202.7661 | 217.9979 | 331.1149 | 180.177 | 73.059 | 64.7661 | 180.177 | 115.7039 | 109.0569 | 110.47 | 124.7039 | 96.5869 | 98.824 | 97.1781 | 133.059 | 66.4131 | 76.001 | 73.001 | 91.706 | 100.1749 | 99.294 | 101.764 | 76.1191 | 235.1191 | 188.118 | 202.354 | 436.0611 | 162.707 | 186.1781 | 172.001 | 390.588 | 294.764 | 208.8209 | 192.47 | 388.353
1902 | 22.9164 | 120.767 | 193.2694 | 186.6165 | 185.9041 | 138.8082 | 134.42 | 163.42 | 154.2694 | 142.6165 | 110.3835 | 129.6165 | 181.6941 | 183.8265 | 141.5571 | 146.79 | 217.0777 | 169.9818 | 127.9818 | 138.9818 | 176.9406 | 140.1371 | 295.3288 | 314.1918 | 151.767 | 207 | 193.8265 | 206.6529 | 331 | 93.9406 | 124.6941 | 113.0594 | 141.7488 | 117.0959 | 137 | 138 | 98.3288 | 218.5753 | 215.6165 | 217.3835 | 261.8494 | 129.7306 | 130.1735 | 119.0594 | 120.3835 | 78.9041 | 58.1918 | 74.0777 | 87.9041 | 174.0594 | 185.58 | 226.3835 | 148.4429 | 118.6529 | 108.5388 | 117.7306 | 122.1141 | 169.8859 | 167 | 171.4018 | 187.5935 | 135.7488 | 151.5571 | 152.1735 | 150.6941 | 144.6347 | 168 | 187.5753 | 257.5982 | 138.3835 | 156.3241 | 202.1735 | 202.5206 | 183.5206 | 167.3241 | 182.7306 | 387.1506 | 145 | 127.0777 | 146.767 | 247.7076 | 122.1918 | 136.0365 | 152.9041 | 146.5571 | 175.3288 | 172.0412 | 181.3471 | 152.3653 | 91.5753 | 95.6347 | 115.9635 | 138.4612 | 224.4065 | 214.977 | 211.42 | 337.7718 | 240.8312 | 214.4429 | 204.4018 | 315.8812 | 181.9818 | 77.5571 | 76.4429 | 181.9818 | 100.8447 | 99.9588 | 106.5571 | 109.8447 | 91.767 | 103.6347 | 105.9635 | 132.8082 | 74.8859 | 81.5206 | 78.1553 | 96.2876 | 86.9223 | 102.7488 | 99.8447 | 88.0777 | 233.5618 | 194.9223 | 209.6165 | 460 | 160.7536 | 196.79 | 177.1553 | 383.6165 | 276.4065 | 185.0182 | 184.1735 | 390.8265
1903 | 22.9274 | 145.4396 | 190.8672 | 190.3275 | 180.9793 | 128.5811 | 154.1121 | 168.7639 | 149.9793 | 129.7138 | 118.7552 | 131.8172 | 160.3156 | 184.3569 | 139.9586 | 136.3776 | 216.9086 | 166.3983 | 124.3983 | 136.9086 | 186.7345 | 135.0827 | 292.5724 | 317.6224 | 163.2242 | 197.938 | 196.2448 | 207.3362 | 339.6844 | 102.9793 | 108.9793 | 100.2448 | 145.469 | 123.1534 | 138.1328 | 135.7345 | 86.5103 | 224.7138 | 211.7758 | 228.0207 | 242.3983 | 127.6017 | 127.3776 | 114.5517 | 120.0707 | 75.1121 | 66.9086 | 64.0914 | 85.6224 | 165.3983 | 184.012 | 223.4276 | 148.9086 | 117.4483 | 112.4189 | 118.2448 | 126.0914 | 161.3776 | 163.6017 | 190.1534 | 202.9793 | 139.0914 | 147.3156 | 153.9086 | 148.5724 | 154.1948 | 163.8466 | 193.3362 | 249.4189 | 140.3362 | 162.7845 | 194.469 | 195.6431 | 177.0207 | 188.1328 | 188.531 | 406.6017 | 147.2655 | 130.6844 | 155.9586 | 258.3863 | 125.6224 | 137.5397 | 143.4483 | 151 | 164.2655 | 153 | 182.9293 | 141.2949 | 100.7345 | 95 | 97.0914 | 131.1828 | 214.6844 | 228.9173 | 223.938 | 336.2361 | 210.2448 | 211.1328 | 233.7258 | 344.531 | 162.1621 | 71.4276 | 74.6431 | 162.1621 | 110.6224 | 108.8052 | 117.4189 | 115.8466 | 103.6017 | 98.8466 | 93.5103 | 129 | 72.4189 | 70.1121 | 67.2448 | 105.3983 | 85.9586 | 101.5604 | 102.4483 | 79.2242 | 219.7258 | 197.7345 | 211.062 | 447.5397 | 132.6017 | 180.7138 | 175.3069 | 378.2655 | 261.0207 | 195.3983 | 181.7552 | 398.1534
1904 | 22.9384 | 161.7798 | 186.2709 | 206.6607 | 186.1368 | 128.1368 | 155.2886 | 147.4583 | 145.7798 | 122.5595 | 117.6607 | 140.2025 | 147.9494 | 189.7469 | 129.2202 | 142.4583 | 219.9316 | 167.7114 | 137.4076 | 134.4734 | 194.3064 | 138.6252 | 311.1696 | 308.8126 | 149.9822 | 184.1696 | 193.4405 | 200.7291 | 341.1341 | 99.5595 | 104 | 97.4405 | 140.7291 | 124.1519 | 130.2531 | 123.0329 | 86.6607 | 213.6607 | 213.3215 | 223.2531 | 245.6607 | 120.8304 | 124.4911 | 118.4227 | 107.7798 | 72.3899 | 69.9316 | 63.4076 | 94.7469 | 173.3393 | 207.2531 | 209.5089 | 149.5924 | 112.0178 | 126.119 | 114.2709 | 113.7114 | 164.3393 | 150.5924 | 194.6607 | 184.3544 | 136.8481 | 130.2709 | 165.8988 | 160.1519 | 168.6607 | 162.4583 | 187.8988 | 268.1874 | 130.2202 | 142.9494 | 189.3393 | 197.2709 | 182.3899 | 183.762 | 190.9316 | 404.8988 | 148.6607 | 145 | 141.7114 | 241.9165 | 125.3899 | 122.8481 | 133.3393 | 135.0151 | 160.5924 | 157.6785 | 186.643 | 125.119 | 97.3899 | 95.3899 | 95.6279 | 119.2886 | 220.4227 | 210.1696 | 218.357 | 373.5266 | 207.0506 | 233.2735 | 252.9316 | 337.5746 | 143.3393 | 55.5595 | 74.7114 | 143.3393 | 112.7291 | 103.6963 | 125.2709 | 112.5089 | 101.119 | 93.9494 | 102.238 | 123.1519 | 82.6101 | 66.6101 | 64.8304 | 99.6936 | 89.6456 | 89.119 | 95.0684 | 73.7798 | 237.762 | 206.0861 | 211.5746 | 441.4254 | 144.1545 | 183.3064 | 187.8304 | 380.0506 | 271.8481 | 193.2025 | 193.1368 | 398.762
1905 | 22.9494 | 159.3805 | 186.4238 | 194.9567 | 191.9458 | 138.3697 | 151.1415 | 154.3913 | 140.9675 | 122.1848 | 111.9892 | 123.5762 | 159.4455 | 212.239 | 137.652 | 148.1848 | 214.0108 | 158.587 | 143.3697 | 131.6737 | 212 | 123.0542 | 306.1632 | 300.826 | 153.0758 | 196.0542 | 184.5653 | 201.7827 | 322.6087 | 113.2607 | 100.7827 | 108.2715 | 135.348 | 114.5978 | 117.8152 | 119.8585 | 85.413 | 220.8585 | 206.8043 | 212.0217 | 243.2065 | 122.2173 | 127.0433 | 120.6845 | 109.8043 | 75.0108 | 66.4238 | 70.9783 | 107.587 | 181.0217 | 180.7393 | 210.174 | 144.4347 | 137.478 | 131.4022 | 104.3697 | 110.2173 | 157.1415 | 144.2282 | 204.2715 | 154.1523 | 137.9567 | 135.2498 | 167.174 | 156.6303 | 162.9892 | 171.8043 | 186.761 | 267.3263 | 134.2282 | 145.5978 | 197.4238 | 212.304 | 180.9892 | 169.9783 | 184.6087 | 394.5112 | 144.5978 | 144.5978 | 143.0433 | 242.5112 | 132.4347 | 126.7718 | 135.3913 | 122.065 | 154.6303 | 160.9783 | 161.1415 | 130 | 102.0217 | 99.2173 | 103.9892 | 132.4347 | 202.174 | 201.1415 | 219.7285 | 388.7935 | 212.4455 | 254.5437 | 243.3913 | 322.6412 | 161.478 | 64.8368 | 73.2282 | 161.478 | 111.3697 | 119.1848 | 119.7935 | 111.9675 | 99.5653 | 103.8368 | 107.174 | 111.587 | 79.1848 | 68.413 | 63.4022 | 86.826 | 110.5978 | 95.2065 | 94.1523 | 76.2065 | 223.5762 | 212.935 | 186.9892 | 457.1632 | 163.3588 | 188.935 | 190.0217 | 389.8693 | 286.6303 | 193.8693 | 209 | 378.9458
1906 | 22.9604 | 148.1976 | 186.3687 | 197.506 | 187.7759 | 132.4867 | 138.7301 | 151.1711 | 141.1181 | 130.8482 | 113.559 | 113.6843 | 171.5855 | 206.0072 | 143.7108 | 136.9542 | 203.7374 | 158.7301 | 134.1711 | 150.3422 | 191.2771 | 145.0458 | 298.9012 | 300.5398 | 164.4675 | 211.8289 | 185.7759 | 195.3422 | 316.7108 | 125.441 | 96.8289 | 126.7566 | 127.6313 | 120.6313 | 130.2892 | 135.5855 | 87.7566 | 227.4675 | 207.5855 | 218.8289 | 256.6048 | 122.0265 | 129.4675 | 128.2506 | 110.5855 | 77.1711 | 68.0265 | 68.3157 | 102.7566 | 178.2964 | 174.3614 | 198.441 | 165.6048 | 149.053 | 122.4675 | 99.3157 | 122.0458 | 143.0723 | 170.6506 | 196.6458 | 156.3349 | 132.2169 | 145.2699 | 158.7566 | 153.8096 | 168.7036 | 177.559 | 170.612 | 256.5325 | 129.5663 | 135.8819 | 196.1253 | 222.3494 | 173.8554 | 177.6771 | 189.4867 | 382.8482 | 141.5133 | 144 | 157.9012 | 240.7952 | 133.7108 | 126.4602 | 133.4145 | 153.2626 | 158.441 | 160.388 | 159.0916 | 131.6578 | 111.7301 | 103.1711 | 106.8024 | 133.2964 | 215.3084 | 207.3807 | 206.2169 | 374.9807 | 222.9277 | 225.7374 | 227.3229 | 346.0193 | 179.6843 | 74.5855 | 80.1711 | 179.6843 | 108.8024 | 112.5133 | 113.8554 | 102.1711 | 97.0458 | 107.7831 | 100 | 108.4145 | 75.4145 | 68.6843 | 69.8024 | 102.2892 | 109.1711 | 91.612 | 97.5783 | 76.3422 | 214.9277 | 200.388 | 189.388 | 437.053 | 160.2892 | 168.5133 | 190.612 | 404.0265 | 281.7374 | 196.0072 | 212.7301 | 376.4337
1907 | 22.9714 | 133.5988 | 181.6942 | 200.93 | 198.3058 | 133.0128 | 141.4395 | 154.6942 | 160.4012 | 149 | 117.0128 | 112.8407 | 166.3058 | 182.4267 | 135.8407 | 138.5222 | 198.656 | 156.3185 | 135.5605 | 144.586 | 167.9745 | 149.4523 | 303 | 301.0955 | 158.2547 | 197.637 | 186.9173 | 191.293 | 302.0128 | 115.586 | 100.5605 | 111.344 | 140.414 | 130 | 139.0128 | 120.9173 | 85.1465 | 208.4523 | 211.2675 | 218.7198 | 270.8663 | 105.1848 | 114.293 | 117.0955 | 118.1082 | 79.8407 | 65.6942 | 69.1593 | 105.2675 | 161.7325 | 187.293 | 193.2802 | 174.7453 | 132.8663 | 117.9617 | 108.6942 | 114.5033 | 150.2675 | 178.6498 | 161.8407 | 175.8663 | 144.414 | 142.5605 | 151.4523 | 142.6942 | 178.8663 | 165.6498 | 181.7769 | 249.0893 | 132.0572 | 126.414 | 199.0827 | 186.1848 | 170.9872 | 186.758 | 190.4395 | 383.93 | 149.5222 | 144 | 159.4395 | 254.2231 | 122.4267 | 128.8918 | 145.949 | 168.586 | 152.8535 | 161.172 | 167.051 | 127.172 | 107.1848 | 103.2802 | 103.9045 | 111.1848 | 244.4267 | 235.1593 | 215.4267 | 350.7453 | 215.7325 | 220.656 | 219.5222 | 355.3185 | 175.4267 | 81.6815 | 73.0255 | 175.4267 | 113.586 | 109.8535 | 107.5733 | 111.6687 | 109.5605 | 97.7198 | 97.4395 | 115.828 | 73.8663 | 64.8535 | 72.0255 | 106.7453 | 103.3058 | 86.9872 | 105.2103 | 70.586 | 227.7897 | 204.1593 | 196.1465 | 438.7897 | 161.758 | 165.4267 | 183.4267 | 390.5988 | 268.5477 | 188.2165 | 210.3185 | 385.465
1908 | 22.9824 | 135.976 | 184.0486 | 193.3664 | 199.0246 | 126.3658 | 138.8781 | 158.3658 | 154.0726 | 139.7802 | 112.1219 | 125.4637 | 160.439 | 179.9267 | 141 | 145.6582 | 207.3411 | 143.8048 | 137.2438 | 134.2925 | 176 | 133.6582 | 311.3418 | 307.6582 | 153.9507 | 188.415 | 182.4883 | 192.9514 | 317.3178 | 112.3171 | 112.7808 | 98.7075 | 133.585 | 124.7315 | 127.5363 | 117.8054 | 86.6342 | 208.4637 | 206.463 | 218.7562 | 274.1466 | 102.0979 | 113.7562 | 131.561 | 124.1706 | 80.1706 | 70.6829 | 70.1466 | 110.1219 | 167.415 | 192.0246 | 191.9267 | 161.0486 | 130 | 115.9507 | 114.1219 | 108.5616 | 154.2438 | 148.2925 | 170.9514 | 165.9754 | 137.585 | 148.6342 | 151.9027 | 133.1946 | 158.8774 | 144.9514 | 208.7562 | 233.7322 | 147.585 | 131.8781 | 210.5363 | 192.3178 | 189.4883 | 173 | 185.6829 | 367.5856 | 160.1706 | 144.439 | 153.3658 | 236.0246 | 127.8294 | 115.3658 | 154.5363 | 162.2438 | 154 | 146.2925 | 165.9274 | 121.0733 | 102.3418 | 97.0973 | 99.976 | 97.561 | 231.8288 | 210.6822 | 234.0006 | 338.3658 | 213.0733 | 239 | 223.1459 | 337.5363 | 172.9267 | 77.9507 | 68.5123 | 172.9267 | 108.561 | 109.6829 | 105.6829 | 125.1219 | 113.439 | 106.976 | 92.6829 | 121.4877 | 76.2685 | 65.561 | 64 | 100.5123 | 97.8781 | 89.2438 | 94.0979 | 80.9274 | 234.9507 | 192.4144 | 186.2192 | 447.7069 | 144.0486 | 167.7562 | 179.1706 | 366.6336 | 254.9261 | 191.5603 | 201.3171 | 390.2445
1909 | 22.9934 | 142.3274 | 183.5133 | 189.9821 | 209.354 | 122.5487 | 145.416 | 152.292 | 143.124 | 129.354 | 100.6194 | 125.9734 | 163.708 | 183.2213 | 139.646 | 140.8673 | 203.124 | 150.4779 | 134.0973 | 129.2567 | 194.9559 | 124.354 | 308.9114 | 287.5221 | 154.0532 | 201.1947 | 197.3894 | 211.0886 | 329.8141 | 118.9646 | 112.2654 | 112.3186 | 114.4867 | 127.9293 | 112.1327 | 128.5221 | 92.2567 | 223.8673 | 202.4779 | 209.7167 | 280.646 | 113.646 | 121.8673 | 123.7787 | 114.8407 | 68.1327 | 68.0973 | 78 | 105.3894 | 184.708 | 191.9734 | 197.9293 | 167.7346 | 128.1947 | 110.6373 | 108.4867 | 121.9734 | 150.1947 | 146.416 | 178.2567 | 156.0973 | 132.4779 | 146.1327 | 155.23 | 132.4072 | 143.77 | 161.7346 | 194.3008 | 250 | 126.23 | 135.708 | 201.9027 | 205.2654 | 193.1061 | 173 | 193.4779 | 373.292 | 148.584 | 139.1327 | 140.0707 | 238.23 | 137.6106 | 112.9027 | 140.4867 | 157.292 | 164.3806 | 150.2833 | 183.646 | 118.6814 | 103.0707 | 97.8319 | 109.0354 | 103.3186 | 209.1327 | 198.6373 | 241.2034 | 328.6814 | 221.062 | 233.584 | 220.1859 | 339.7346 | 172.1593 | 65.8673 | 72.5487 | 172.1593 | 108.9027 | 103.0354 | 101.292 | 126.708 | 113.5487 | 106.5574 | 98.6726 | 114.7433 | 76.23 | 68.6106 | 68.062 | 103.646 | 100.8053 | 89.708 | 107.4513 | 100.9027 | 224.6726 | 195.7346 | 179.9646 | 447.7966 | 146.6726 | 178.124 | 169.8407 | 369.5665 | 228.4513 | 180.6373 | 204.354 | 415.062
1910 | 23.0044 | 131.6091 | 194.0998 | 183.7902 | 198.0186 | 136.8359 | 147.8274 | 154.5635 | 148.8274 | 145.8359 | 110.2539 | 118.709 | 158.1912 | 182.2639 | 144.0271 | 143.3638 | 215.7817 | 157.3638 | 142.7361 | 132.4636 | 208.7276 | 131.0998 | 293 | 291.2995 | 161.1641 | 195.291 | 206.9087 | 214.9458 | 318.2454 | 125.4636 | 99.8545 | 120.4636 | 132.1811 | 119.137 | 115.1997 | 123.0271 | 95.4636 | 220.8003 | 209.8274 | 202.4907 | 261.3824 | 120.8088 | 121.1184 | 114.0728 | 109.6091 | 72.5905 | 64.6819 | 77.5364 | 97.7546 | 192.6362 | 178.6819 | 203.5093 | 169.9458 | 133.418 | 114.8731 | 105.3181 | 115.1726 | 142.9002 | 147.4365 | 186.5635 | 155 | 136.1184 | 128.3367 | 141.4365 | 149.0457 | 149.6819 | 168.1184 | 167.0457 | 241.1912 | 127.7361 | 135.7546 | 199.7546 | 187.7546 | 183.8088 | 169.291 | 190.6277 | 377.8816 | 150.8088 | 133.3909 | 135.8816 | 243.9087 | 137.3638 | 127.445 | 136.8545 | 147.5093 | 172.3638 | 158.4365 | 182.9272 | 111 | 95.6362 | 106.8274 | 114.1269 | 119.8088 | 204.3181 | 212.1455 | 219.2183 | 333.0542 | 228.3181 | 226.5364 | 235.4636 | 356.7817 | 165.3367 | 70.6819 | 66.4365 | 165.3367 | 112.7817 | 114.154 | 108.1997 | 124.4365 | 106.0457 | 96.418 | 105.2183 | 113.8545 | 66.1455 | 74.3909 | 71.6091 | 98.7546 | 103.9272 | 86.9729 | 102.9002 | 96.9002 | 226.1184 | 200.2639 | 193.8816 | 459.1184 | 150.9002 | 183.8274 | 172.0271 | 409.0542 | 226.6819 | 181.1641 | 196.7276 | 408.4095
1911 | 23.0154 | 145.2292 | 182.3912 | 190.4394 | 184.4277 | 142.5781 | 144.9037 | 158.1445 | 158.2774 | 147.7708 | 126.9577 | 126.8073 | 150.8555 | 170.0482 | 138.1503 | 141.8073 | 220.9037 | 147.2409 | 151.6204 | 139.6628 | 198 | 134.1445 | 296.8073 | 296.1503 | 157.8015 | 191 | 213.4817 | 204.3314 | 327.711 | 107.9153 | 98.6686 | 114.3372 | 142.343 | 113.0847 | 131.2832 | 126.1927 | 89.3372 | 220.8978 | 197.8613 | 222.3737 | 256.7051 | 121.0058 | 114.8555 | 116.3314 | 115.6146 | 76.9577 | 74.8497 | 77.4759 | 98.7591 | 181.343 | 175.5241 | 205.9942 | 168.8497 | 128.1985 | 111.9942 | 119.8978 | 111.9095 | 147.4701 | 157.657 | 191.0963 | 159.7591 | 134.1387 | 112.1927 | 139.2832 | 145.7591 | 158.4219 | 163.2832 | 165.6628 | 243.8497 | 139.9518 | 130.5723 | 189.3372 | 190.6628 | 188.289 | 167.8555 | 178.7591 | 360.8248 | 153.8613 | 141.1869 | 144.5241 | 237.1562 | 122.4817 | 138.2409 | 147.0905 | 136.6686 | 159.3854 | 161.9942 | 179.7168 | 122.8978 | 97.6686 | 95.8131 | 111.7708 | 119.5299 | 221.7533 | 219.5183 | 198.8671 | 340.3372 | 229.0963 | 226.9518 | 247.8978 | 342.3912 | 153.9518 | 68 | 60 | 153.9518 | 113.1445 | 121.1503 | 124.2832 | 114.1927 | 104.6628 | 102.1445 | 103.4277 | 113.6204 | 71.1387 | 66.9577 | 78.5664 | 94.0482 | 95.9577 | 78.5723 | 95.0963 | 88.6204 | 217.9518 | 197.0905 | 214.8978 | 463.8015 | 159.7533 | 178.0482 | 177.0963 | 411.1022 | 244.9402 | 167.8073 | 187.9037 | 402.6146
1912 | 23.0264 | 159.0708 | 169.3467 | 180.7405 | 192.3467 | 141.7169 | 153.3467 | 157.441 | 158.3775 | 134.9056 | 124.3231 | 123.6769 | 164.2523 | 171.441 | 141.5753 | 142.5826 | 220.0708 | 142 | 142.6533 | 144.0708 | 193.6062 | 134.4174 | 298.0708 | 272.7241 | 157.3775 | 205.1579 | 222.04 | 203.6062 | 320.8185 | 97.2759 | 96.4646 | 105.5354 | 128.882 | 138.3702 | 123.7949 | 128.3467 | 92.2523 | 225.7005 | 211.9692 | 232.0944 | 267.1651 | 109.5118 | 119.4646 | 123.9056 | 122.5354 | 71.3938 | 84.118 | 77.5118 | 109.8584 | 173.2523 | 171.0944 | 214.0708 | 168.3539 | 118.8584 | 118.118 | 131.5354 | 114.2287 | 152.6533 | 158.4247 | 184.118 | 160.118 | 137.118 | 122.1579 | 136.1887 | 156.3702 | 165.1416 | 160.6769 | 182.7641 | 261.9056 | 130.2595 | 130.4646 | 188.8349 | 191.6533 | 185.9056 | 176.8584 | 183.5118 | 355.922 | 152.8349 | 147.0236 | 142.0472 | 220.7877 | 113.9528 | 129.5826 | 145.2595 | 134.9528 | 168.0871 | 169.0944 | 179.3938 | 124.7713 | 98.882 | 92.4174 | 95.4882 | 113.3702 | 218.4718 | 215.8421 | 190.2523 | 343.7405 | 215.2831 | 230.9292 | 260.0236 | 326.9056 | 158.9056 | 73.3702 | 63.9056 | 158.9056 | 106.0944 | 103.5826 | 111.4247 | 114.882 | 103.2123 | 92.1651 | 104.0236 | 107.0944 | 71.1887 | 64.3231 | 79.7005 | 95.441 | 95.2759 | 80.9056 | 88.6062 | 83.0708 | 221.441 | 198.1887 | 230.441 | 455.078 | 165.2595 | 184.3231 | 160.8421 | 399.4646 | 272.882 | 175.4174 | 202.6933 | 393.9129
1913 | 23.0374 | 150.4946 | 179.5035 | 193.001 | 204.005 | 143.4896 | 149.9901 | 149.4896 | 143.5015 | 144.5054 | 117.4856 | 115.4995 | 163.9891 | 175.001 | 146.9871 | 144.005 | 212.996 | 143.5015 | 135.4995 | 129.4886 | 177.9891 | 142.004 | 298.5035 | 276.0129 | 157.0158 | 199.4797 | 220.9792 | 195.997 | 337.0297 | 108.001 | 104.5064 | 107.003 | 129.4955 | 135.4916 | 115.004 | 135 | 96.9941 | 216.4995 | 221.9871 | 218.4906 | 257.998 | 113.5045 | 114.4936 | 125.4975 | 116.4955 | 76.5005 | 73.995 | 69.4926 | 111.4955 | 180.4965 | 163.4965 | 209.999 | 149.4965 | 125.5005 | 119.5064 | 130.5005 | 124.4985 | 159.0129 | 148.0079 | 175.4965 | 164.5094 | 140.5084 | 137 | 134.5064 | 153.9921 | 153.995 | 150.998 | 183.4906 | 268.5025 | 124.5054 | 143.0109 | 194.4985 | 185 | 192.5025 | 189.0059 | 170.9881 | 363.9832 | 152.4926 | 142.997 | 146.5064 | 226.997 | 110.995 | 128.5025 | 133.4995 | 139.002 | 171.4856 | 166.001 | 171.9881 | 126.0129 | 98.4946 | 99.5035 | 96.5005 | 110.9921 | 222.0238 | 184.4837 | 202.001 | 365.0099 | 222.0188 | 232.4985 | 233.4747 | 322.9921 | 160.4975 | 66.4876 | 59.9921 | 160.4975 | 109.0069 | 106.5064 | 114.5213 | 115.4955 | 99.005 | 92.0069 | 109.0059 | 103.5005 | 67.004 | 72.5005 | 74.5035 | 82.4847 | 99.995 | 87.5025 | 84.5005 | 80 | 226.002 | 196.0059 | 216.4837 | 448.5183 | 156.002 | 183.9921 | 167.5213 | 403.5025 | 258.9812 | 185.5064 | 203.9881 | 399.0228
1914 | 23.0484 | 147.0511 | 184.0256 | 194.5128 | 186.9502 | 139.1534 | 151.7941 | 153.8708 | 150.1279 | 144.8721 | 105.0511 | 116.0256 | 156.5895 | 180.1023 | 135.0256 | 136.6931 | 212.5895 | 159.358 | 136.5384 | 124.6662 | 189.0498 | 141.3849 | 285.5908 | 291.5639 | 167.8721 | 186.6918 | 198.9744 | 205.3069 | 337.7712 | 110.5384 | 107.4105 | 105.3849 | 135.2557 | 128.0256 | 117.4616 | 138.5895 | 94.0767 | 214.9744 | 208.4872 | 207.9744 | 260.1023 | 112.8721 | 111.0767 | 119.9233 | 128.9219 | 78.0256 | 65.4105 | 67.1279 | 98.7954 | 169.821 | 160.5128 | 198.7443 | 154.7174 | 118.821 | 125.4872 | 124.8466 | 121.9744 | 176.6151 | 156 | 174.5639 | 171.9489 | 135.6676 | 122.642 | 144.5895 | 152.1534 | 142.3338 | 156.6918 | 187.8452 | 253.5653 | 134.1023 | 133.4886 | 201.7174 | 201.9219 | 190.3849 | 186.2826 | 168.7429 | 381.8695 | 132.6931 | 145.1279 | 143.7698 | 225.5384 | 117.2813 | 127.9233 | 125.821 | 149.2046 | 167.2557 | 164.9489 | 158.4616 | 127.2059 | 87.3594 | 103.5128 | 107.2557 | 111.7174 | 225.4886 | 179.7941 | 213.2557 | 358.5908 | 219.463 | 208.4374 | 208 | 330.3836 | 153.8977 | 55.5384 | 65.3324 | 153.8977 | 112.9233 | 108.3849 | 119.5908 | 107.9233 | 107.0767 | 102.0767 | 95.0013 | 107.0767 | 71.5128 | 73 | 82.6151 | 81.8708 | 96.0256 | 87.9489 | 85 | 83.5895 | 225.9489 | 203.5384 | 189.2315 | 449.0525 | 153.8977 | 190.358 | 183.8721 | 414.2046 | 255.8964 | 189.4361 | 180.7187 | 415.3338
1915 | 23.0594 | 141.1238 | 186.0502 | 202.3511 | 160.7492 | 146.5752 | 158.7994 | 149.6223 | 144.4985 | 137.3746 | 111.2006 | 116.4749 | 155.2743 | 192.4012 | 134.9498 | 131.3009 | 233.8526 | 165.1238 | 140.6254 | 136.7758 | 211.5752 | 141.2006 | 271.5752 | 301.8762 | 164.0502 | 186.1238 | 192.2242 | 210.174 | 308.9498 | 109.8997 | 102.9498 | 98.3746 | 131.348 | 130.5752 | 119.1505 | 136.7492 | 92.2743 | 208.7492 | 205.8997 | 222.2272 | 262.4248 | 130.0532 | 106.6489 | 121.2006 | 138.174 | 72.6991 | 63.0502 | 72 | 100.9765 | 170.8762 | 170.9765 | 188.4749 | 153.0235 | 116.7257 | 119.7492 | 121.1003 | 113.6489 | 175.7492 | 152.3245 | 174.8997 | 175.7758 | 125.1003 | 113.7257 | 146.4248 | 156.4248 | 143.8762 | 168.7257 | 192.5988 | 242.7758 | 143.7758 | 122.9263 | 203.174 | 204.8731 | 190.7257 | 180.1003 | 186.4012 | 394.522 | 134.652 | 146.8495 | 141.826 | 235.4012 | 120.1238 | 133.9263 | 120.0502 | 154.3746 | 173.8495 | 165.6254 | 166.4514 | 118.1003 | 89.3511 | 107.1505 | 95.4718 | 110.0235 | 211.2508 | 185.2242 | 219.8495 | 345.6254 | 203.7257 | 194.8762 | 212.2006 | 338.174 | 151.0502 | 66.4514 | 74.8495 | 151.0502 | 95.8229 | 103.4749 | 112.9263 | 113.4012 | 98.4483 | 97.6489 | 81.2508 | 112.6254 | 74.1003 | 65.1238 | 73.348 | 95.4749 | 90.174 | 87.5752 | 88.6755 | 84.8997 | 198.7963 | 198.174 | 187.9263 | 419.3981 | 154.2006 | 186.6725 | 172.6991 | 403.0972 | 269.9498 | 178.5988 | 174.7257 | 424.2272
1916 | 23.0704 | 142.0605 | 173.0284 | 197.1779 | 164.0605 | 140.4768 | 153.3879 | 145.3595 | 143.0605 | 133.3879 | 117.6868 | 122.4484 | 141.8647 | 190.3274 | 140.9858 | 140.2242 | 241.9395 | 151.5516 | 132.2526 | 146.2989 | 199.0284 | 142.8505 | 300.4058 | 306.3132 | 160.7011 | 182.2242 | 198.8221 | 208.2989 | 312.8363 | 117.6726 | 103.6121 | 106.21 | 135.6584 | 126.6263 | 117.1637 | 129.3132 | 90.1495 | 215.8221 | 195.9395 | 217.7295 | 264.7616 | 133.3416 | 107.5232 | 129.2989 | 123.9395 | 76.6726 | 68.2989 | 63.4021 | 115.3737 | 187.1353 | 185.9111 | 189.0747 | 138.6263 | 115.6263 | 120.3737 | 123 | 116.1353 | 169.9253 | 157.0605 | 175.6868 | 182.6121 | 125.9253 | 122.2989 | 141.2384 | 153.3879 | 164.4342 | 166.5516 | 189.2989 | 243.7011 | 138.79 | 118.6405 | 207.21 | 193 | 177.8042 | 180.9253 | 194.5374 | 373.8505 | 146.4626 | 144.5374 | 149.6121 | 244.6121 | 113.5374 | 147.9111 | 124.2242 | 142.3274 | 179.5979 | 168 | 182.5232 | 132.3595 | 92.2384 | 104.0889 | 98.5695 | 117.121 | 221.3737 | 203.6442 | 237.9574 | 370.5695 | 200.4768 | 203.6121 | 215.4626 | 316.4163 | 155.2242 | 71.7758 | 73.6121 | 155.2242 | 98.5837 | 112.1353 | 116.7011 | 127.4484 | 86.9253 | 92.0747 | 93.5232 | 111.2384 | 71.1637 | 70.3595 | 72.2847 | 94.4626 | 101.1958 | 90.0747 | 100.0605 | 77.6263 | 194.2705 | 198.4484 | 199.2242 | 437.0179 | 159.0747 | 182.8221 | 157.3274 | 375.2526 | 267.9253 | 186.5837 | 183.8363 | 400.3842
1917 | 23.0814 | 141.3048 | 152.7551 | 189.1986 | 170.4503 | 129.6027 | 152.5497 | 173.5891 | 143.9537 | 139.1456 | 117.2517 | 117.5565 | 143.8939 | 171.0068 | 150.298 | 145.1986 | 225.6558 | 149.298 | 133.9932 | 145.0531 | 188.0993 | 137.702 | 305.8612 | 299.6027 | 157.5497 | 203.1388 | 206.2517 | 224.6422 | 313.702 | 124.9007 | 113.2449 | 104.0068 | 142.3048 | 124.7483 | 113.5497 | 125.351 | 97.4966 | 212.8082 | 204.3905 | 210.2449 | 270.7483 | 117.9007 | 111.2517 | 132.4503 | 119.1986 | 85.5497 | 71.4503 | 63.6952 | 119.4503 | 199.3973 | 191.5497 | 185.0531 | 149.9401 | 103.3048 | 120.0531 | 121.351 | 113.0068 | 165.702 | 156.8544 | 177.4503 | 180.1524 | 131.5959 | 109.5102 | 133.0531 | 150.9007 | 167.7551 | 163.7483 | 180.3578 | 232.8544 | 140.9932 | 134.9333 | 212.702 | 190.2517 | 180.5891 | 166.2585 | 188.9537 | 398.3836 | 148.7483 | 143.9007 | 146.0531 | 247.649 | 116.7483 | 150.2517 | 122.6027 | 132.9007 | 187.5497 | 164.702 | 179.6558 | 132.5565 | 89 | 100.0993 | 101.5102 | 121.1061 | 210.6095 | 205.8612 | 231.2653 | 378.4572 | 201.1456 | 205.5497 | 222.1456 | 313.4435 | 147.5565 | 71.7483 | 75 | 147.5565 | 109.8014 | 107.3578 | 120.1456 | 111.0136 | 84.9007 | 107.2911 | 95.6027 | 101.4041 | 69.1986 | 80.4503 | 68.2585 | 97.298 | 118.1986 | 95.9469 | 103.702 | 67.5034 | 202.8544 | 222.6884 | 210.7945 | 445.5102 | 141.3116 | 184.7551 | 161.6422 | 369.8476 | 277.9932 | 196.1524 | 187.4503 | 386.1388
1918 | 23.0924 | 151.2966 | 171.2878 | 189.3141 | 183.4868 | 135.5532 | 143.4468 | 167.7698 | 149.1151 | 152.8673 | 121.1815 | 123.0488 | 152 | 174.9249 | 136.7034 | 140.8185 | 228.6771 | 162.6771 | 136.8185 | 132.5707 | 193.4956 | 137.8098 | 317.9736 | 299.9337 | 161.3717 | 206.761 | 196.6946 | 225.4468 | 313.2478 | 117.2566 | 107.6371 | 89.3805 | 135.438 | 123.0663 | 116.2478 | 135.239 | 86.2654 | 204.2478 | 217.562 | 201.2654 | 269.0663 | 124.3054 | 112.9337 | 130.3141 | 126.0576 | 77.5707 | 70.438 | 59.8849 | 101.5795 | 196.2566 | 191.438 | 194.4868 | 174.239 | 106.5532 | 127.239 | 118.8761 | 100.0663 | 153.4468 | 161.1151 | 171.9424 | 165.761 | 124.0751 | 110.6107 | 131.2478 | 148.3141 | 156.5044 | 174.9912 | 193.04 | 243.8585 | 143.8185 | 138.4556 | 188.0839 | 213.2878 | 200.6195 | 184.7834 | 196.9249 | 402.5795 | 149.3141 | 141.8761 | 153.239 | 236.0751 | 131.3629 | 132.8273 | 121.2478 | 150.5444 | 168.8937 | 165.3717 | 166.3805 | 125.1239 | 106.4205 | 92.0088 | 88 | 117.4956 | 204.7434 | 206.1727 | 199.199 | 348.7698 | 214.3054 | 200.3805 | 239.801 | 316.9424 | 162.04 | 72.3141 | 85.1151 | 162.04 | 110.2478 | 98.1239 | 115.3229 | 116.04 | 92.9912 | 109.4468 | 105.4868 | 108.3629 | 71 | 77.1902 | 59.8098 | 93.2566 | 110.4468 | 92.1327 | 106.0576 | 77.0488 | 222.8497 | 223.9512 | 228.1151 | 444.9249 | 142.2966 | 176.3141 | 175.9337 | 381.9912 | 270.1415 | 200.3054 | 208.9161 | 377.96
1919 | 23.1034 | 151.3667 | 178.9409 | 186.2773 | 183.6637 | 145.8712 | 130.2576 | 166.356 | 152.9803 | 151.5349 | 123.1288 | 125.9606 | 162.3363 | 187.8712 | 136.6333 | 137.7227 | 235.2773 | 158.9409 | 134.8712 | 134.6136 | 195.8515 | 145.7424 | 320.6243 | 284.05 | 185.8212 | 188.2379 | 193.297 | 211.6833 | 324.7621 | 107.4061 | 104.3167 | 97.0591 | 135.5742 | 131.4848 | 127.1879 | 132.5152 | 86.6333 | 200.2576 | 204.2182 | 210.2469 | 283.2272 | 128.2773 | 121.1682 | 128.4258 | 123.6833 | 60.6637 | 66.5546 | 65.7818 | 102.356 | 193.297 | 188.703 | 195.8121 | 187.5939 | 124.9106 | 125.6637 | 113.9803 | 101.5939 | 156.9106 | 155.7924 | 166.2773 | 165.0394 | 128.9303 | 119.703 | 140.4651 | 156.7621 | 149.5546 | 171.0894 | 196.644 | 251.8318 | 140.7227 | 125.1485 | 196.8409 | 207.7334 | 194.0894 | 192.4955 | 196.0894 | 384.4061 | 144.5546 | 140.4258 | 140.7531 | 216.8121 | 129.5152 | 134.7818 | 131.0394 | 143.1788 | 152.2773 | 162.8318 | 178.653 | 129.4454 | 110.8121 | 108.5439 | 82.8318 | 116.4061 | 214.5939 | 203.644 | 200.3363 | 351.3757 | 220.5742 | 206.3363 | 223.1591 | 330.2272 | 158.7531 | 60.6637 | 79.7924 | 158.7531 | 117.1682 | 100.7227 | 94.9409 | 130.5546 | 113.7818 | 106.0197 | 98.1985 | 107.0894 | 70.4258 | 66.3864 | 62.5742 | 93.1682 | 94.3864 | 90.0197 | 99.0894 | 78.8121 | 228.0697 | 203.8121 | 221.0697 | 449.8318 | 156.7227 | 172.703 | 176.703 | 372.9212 | 248.9606 | 203.703 | 197.2879 | 381.2469
1920 | 23.1144 | 143.1731 | 177.0384 | 187.9327 | 188.3172 | 130.404 | 134.798 | 166.4424 | 150 | 150.6923 | 118.0673 | 120 | 165.3077 | 180.6155 | 146.5865 | 141.9327 | 224.0289 | 159.9711 | 131.7212 | 129.8559 | 196.1731 | 143.5481 | 296.4424 | 275.2788 | 185.5771 | 201.5286 | 180.9232 | 213.4519 | 332 | 102.2404 | 99.0289 | 100.1347 | 133.0673 | 118.8848 | 119.9232 | 121.0673 | 104.798 | 219.4613 | 211.596 | 217.7885 | 294.2404 | 129.3461 | 134.9711 | 123.3077 | 125.4519 | 65.9037 | 61.6539 | 87.7306 | 108.3077 | 181.5098 | 178.202 | 203.0768 | 171.0579 | 129.4808 | 108.6155 | 116.8653 | 111.4519 | 154.4424 | 142.4808 | 167.3461 | 177.4519 | 132.3751 | 130.3172 | 150.1057 | 156.9616 | 144.0673 | 148.3367 | 203.0094 | 255.6249 | 156.0768 | 145.9421 | 215.4135 | 200.7306 | 197.7306 | 181 | 186.2404 | 407.394 | 146.6923 | 147.6249 | 124.4135 | 226.4229 | 119.8269 | 137.3751 | 138.7596 | 144.0094 | 156.2788 | 151.3751 | 185.1347 | 131.4135 | 108.1057 | 113.0963 | 100.1152 | 122.9711 | 205.0963 | 204.1441 | 212.6923 | 360.3077 | 216.8943 | 192.8848 | 211.0384 | 324.231 | 151.2115 | 65.3172 | 70 | 151.2115 | 103.9906 | 104.3461 | 104.769 | 123.3077 | 106.9906 | 108.4135 | 99.6633 | 103.1057 | 78.798 | 62.3461 | 63 | 88.202 | 104.4229 | 98.8653 | 94.5192 | 73.7596 | 218.7596 | 195.2404 | 201.202 | 441.3077 | 168.5576 | 188.0094 | 177.3461 | 389.1536 | 246.5192 | 203.1731 | 178.9327 | 388.7885
1921 | 23.1254 | 146.9941 | 181.4012 | 178.6224 | 180.8318 | 126.3835 | 157.7671 | 157.8024 | 151.1976 | 153.4012 | 127.9765 | 114.0118 | 166.1918 | 187.1741 | 152.3894 | 152.3835 | 224.1859 | 146.64 | 140.5753 | 130.3835 | 191.6106 | 137.8024 | 290.7965 | 288.5812 | 160.8259 | 204.0235 | 194.9529 | 208.4188 | 328.4071 | 109.9824 | 101.5812 | 111.5694 | 141.7788 | 123.1624 | 107.6047 | 124.9882 | 103.2153 | 215.0412 | 244.36 | 215.5929 | 272.64 | 133.9941 | 121.0412 | 127.7847 | 117.4247 | 66.6165 | 70.7788 | 84.6224 | 106.1976 | 177.3894 | 176.7906 | 201.0235 | 156.4012 | 125.8024 | 112.7788 | 121 | 119.5929 | 145.8024 | 150.18 | 180.9765 | 171.82 | 138.9765 | 132.4129 | 143.4188 | 147.8082 | 148.5871 | 159.5576 | 195.8376 | 265.1918 | 148.6341 | 132.4541 | 213.2035 | 199.4188 | 187.4365 | 162.4365 | 192.7847 | 392.4659 | 140.4188 | 145.2153 | 131.1859 | 219.4365 | 141.7553 | 137.3894 | 141.1976 | 156 | 164.7906 | 171.1506 | 174.4129 | 132.7965 | 107.4071 | 111.1859 | 110.2094 | 124.0118 | 209.7729 | 230.5635 | 207.6165 | 357 | 228.3718 | 199.5576 | 226.7788 | 304.4129 | 159.3953 | 72.2035 | 70 | 159.3953 | 97.9882 | 104.2035 | 113.8082 | 117.6047 | 81.2271 | 100.8141 | 107.6047 | 113.1859 | 77.8141 | 75.9765 | 65.9941 | 92.18 | 112.4071 | 105.9941 | 88.0176 | 68.4129 | 234.9706 | 204.7788 | 207.5753 | 440.3953 | 159.2329 | 168.8612 | 170.0176 | 392.8318 | 267.5635 | 171.0647 | 196.5694 | 387.7906
1922 | 23.1364 | 158.1667 | 177.3333 | 193.0001 | 179.7778 | 148.5001 | 150.6666 | 150.2777 | 147.1111 | 133.4443 | 131.1111 | 116.7223 | 181.2223 | 197.2222 | 138.8888 | 162.2778 | 211.2777 | 161.1113 | 137.9999 | 141.5 | 188.6111 | 151.0556 | 287.1111 | 288.8889 | 158.1111 | 203.9445 | 204.8889 | 206.8889 | 352.2779 | 116 | 107.3889 | 106.7221 | 136.1666 | 124.9999 | 121.2779 | 130.2222 | 102.2778 | 202.2222 | 232.9443 | 227.7778 | 252.2777 | 127.4444 | 107.6111 | 128.7222 | 122.4445 | 65.2778 | 80.4445 | 72.1111 | 105.1667 | 187.1111 | 181.8333 | 199.1111 | 166.3889 | 117.0555 | 127.3334 | 111.2222 | 123.8333 | 152.9445 | 155.7778 | 187.7778 | 168.0556 | 139.0555 | 133.5 | 133.9444 | 141.9444 | 150.5555 | 157.4999 | 192.1667 | 271.6667 | 150.4445 | 123.7778 | 216.2778 | 204.0001 | 175.6111 | 170.7779 | 182.7221 | 368.1667 | 131.5555 | 149.1667 | 145.7778 | 206.3889 | 139.8888 | 127.5555 | 134.6666 | 160.2778 | 163.7222 | 167.2221 | 181.6112 | 129.7222 | 97.6666 | 106.8333 | 100.2777 | 129.1667 | 225.1111 | 221.611 | 202 | 353.3333 | 229.4444 | 204.2222 | 233.3889 | 314.0556 | 159.7778 | 75.2778 | 73.6667 | 159.7778 | 107.5 | 112.1667 | 117.1111 | 114.7778 | 94.0001 | 97.8333 | 96.2222 | 121.0556 | 68.7222 | 69.9444 | 72.8889 | 86.1666 | 115.5 | 101.8889 | 97.2779 | 63.3889 | 240.7222 | 205.8889 | 203.7777 | 427.3332 | 155.3334 | 163.889 | 171.3334 | 380.7778 | 284.8889 | 169.7779 | 208.8333 | 388.5555
1923 | 23.1474 | 145.2978 | 175 | 198.8957 | 180.3893 | 153.3893 | 140.8702 | 161.5851 | 142.7532 | 135.9617 | 125.5064 | 128.4809 | 172.7915 | 199.8702 | 123.0128 | 150.0383 | 210.5979 | 175.2596 | 137.2341 | 151.2341 | 212.6894 | 138.1808 | 272.1553 | 285 | 189.4298 | 218.3511 | 203.4936 | 213.7404 | 322.868 | 104.7787 | 99.5191 | 100.7404 | 132.3638 | 124.2341 | 127.2596 | 123.5191 | 87.5446 | 203 | 212.6362 | 219.6617 | 286.0277 | 113.8957 | 114.2341 | 111.6617 | 123.5191 | 72.3638 | 64.5446 | 72.7404 | 98.3893 | 185.3893 | 192.3511 | 199.2596 | 172.3766 | 119.4809 | 124.5191 | 118.7149 | 127.4936 | 164.8575 | 155.6234 | 204.4554 | 178.7277 | 142.7277 | 128.2723 | 143.8447 | 142.4936 | 152.7404 | 167.5724 | 188.0255 | 256.5446 | 140.2978 | 127.5064 | 197.1808 | 196.6617 | 187.2213 | 184 | 182.3511 | 362.6362 | 123.1425 | 150.0128 | 144.5191 | 221.5851 | 118.4021 | 127.7277 | 139.3511 | 155.5191 | 152.8957 | 163.3511 | 172.7915 | 127.6234 | 97.9872 | 107.2341 | 87.8957 | 123.1553 | 222.7659 | 196.1553 | 189.5319 | 369.0788 | 233.9745 | 214.2085 | 235.4936 | 302.4276 | 152.1425 | 71.1425 | 72.883 | 152.1425 | 123.4681 | 121.7404 | 107.9085 | 119.6107 | 100.5191 | 107.1043 | 100.5979 | 119.2596 | 72.2341 | 80.3256 | 70.3893 | 92.3383 | 111.5191 | 103.6107 | 93.2851 | 79.2085 | 241.7404 | 197.0128 | 194.1298 | 427.3511 | 158.1298 | 164.6489 | 165.4021 | 366.9085 | 290.4936 | 183 | 202.5191 | 380.1425
1924 | 23.1584 | 142.2641 | 186.4425 | 210.5281 | 194.7995 | 144.2787 | 142.6357 | 174.8142 | 166.1565 | 146.2714 | 107.4719 | 119.0147 | 162.3643 | 176.8435 | 137.7995 | 149.8998 | 208.7359 | 163.4645 | 132.1002 | 123.2152 | 215.5575 | 134.5355 | 299.9633 | 297.7139 | 167.8582 | 240.5281 | 205.6357 | 219.8142 | 319.1565 | 115.1638 | 117.8851 | 98.5501 | 151.5281 | 129.9071 | 133.2641 | 122.1785 | 83.9927 | 205.5428 | 229.7066 | 220.5355 | 275.9438 | 113.4499 | 112.9144 | 101.7286 | 119.6357 | 78.1785 | 77.5208 | 68.0073 | 98.1785 | 180.0929 | 197.3643 | 199.5428 | 158.6504 | 118.9144 | 126.357 | 119.3716 | 123.9144 | 174.0856 | 138.8362 | 199.1076 | 169.379 | 149.2714 | 113.4645 | 154.1785 | 139.5501 | 147.3716 | 166.6504 | 195.3496 | 255.5355 | 134.0856 | 138.0782 | 191.6284 | 187.3643 | 193.3643 | 182.0929 | 195.6284 | 375.8924 | 122.1785 | 149.5428 | 148.8998 | 236.0856 | 129.1638 | 123.4645 | 127.2005 | 131.9291 | 157.5355 | 166.4572 | 163 | 141.3496 | 106.7213 | 88.1149 | 93.8068 | 122.9927 | 234.8924 | 198.5355 | 211.242 | 390.1711 | 224.7433 | 204.2934 | 240.8142 | 323.6919 | 157.5355 | 70.1785 | 76.7213 | 157.5355 | 125.9144 | 118.2787 | 113.9853 | 118.5501 | 101.7213 | 112 | 94.9218 | 99.8362 | 86.1711 | 79.9218 | 83.5281 | 100.3643 | 107.6357 | 103.8215 | 104.7066 | 85.1858 | 226.8362 | 193.3643 | 206.9853 | 427.9144 | 163.357 | 177.4352 | 177.4352 | 365.357 | 302.8068 | 189.357 | 190.3716 | 407.7848
1925 | 23.1694 | 141.8161 | 189.7601 | 208.9842 | 184.2644 | 152.6638 | 152.0718 | 175.056 | 161.2084 | 158.6638 | 105.7759 | 109.704 | 168.4799 | 201.8793 | 136.9842 | 157.592 | 213.0718 | 161.2399 | 154.8635 | 116.0158 | 192.1523 | 134.1681 | 292.1365 | 308.2399 | 171.2715 | 209.8248 | 207.944 | 205.8003 | 337.5359 | 117.8721 | 120.6322 | 91.4641 | 139.6164 | 134.2399 | 133.4641 | 115.5762 | 90.592 | 200.5201 | 230.6322 | 229.2399 | 262.1839 | 118.592 | 113.8879 | 125.6235 | 127.1279 | 78.056 | 68.3204 | 66.592 | 106.4799 | 191.9598 | 181.4483 | 193.8721 | 161.3678 | 133.4957 | 120.9282 | 115.648 | 121 | 144.6006 | 134.8319 | 182.2241 | 166.8319 | 137.0402 | 120.9598 | 157.944 | 143.4799 | 161.1437 | 164.1839 | 204.296 | 264.8879 | 127.9282 | 141.7601 | 203.7759 | 196.7198 | 201.4238 | 176.4641 | 186.3923 | 379.1681 | 140.8477 | 158.7759 | 156.592 | 234.4641 | 133.8161 | 121.2399 | 128.6638 | 139.7916 | 174.0158 | 163.056 | 160.408 | 141.2241 | 113.8879 | 95.7356 | 101.296 | 117.9282 | 229.7443 | 208.5359 | 204.0245 | 383.6882 | 212.7601 | 202.0718 | 244.296 | 319.6724 | 165.592 | 70.056 | 74.1681 | 165.592 | 123.648 | 103.3362 | 127.1839 | 117.944 | 90.7443 | 107.4641 | 91.2399 | 113.9755 | 81.6322 | 71.056 | 82.6322 | 102.592 | 113.8319 | 92.2802 | 107.5762 | 84.944 | 211.1681 | 196.8879 | 209.1681 | 433.4238 | 161.1681 | 186.704 | 199.0158 | 388.4396 | 268.8248 | 197.5359 | 173.6882 | 388.4167
1926 | 23.1804 | 144.9425 | 183.378 | 215.5453 | 189.507 | 139.1916 | 147.7561 | 164.0958 | 152.9808 | 143.2108 | 127.8275 | 103.0575 | 176.622 | 215.7369 | 138.2631 | 158.3397 | 221.9617 | 166.9617 | 144.9094 | 131.2439 | 195.5453 | 141.5645 | 293.507 | 290.8519 | 177.7369 | 194.6028 | 180.608 | 217.4686 | 330.7561 | 116.6411 | 116.9808 | 102.2056 | 139.2247 | 130.0575 | 130.3397 | 126.1864 | 99.2631 | 206.9042 | 215.0958 | 219.7753 | 270.2822 | 123.9617 | 108.7369 | 129.0958 | 132.9808 | 77.6603 | 75.1481 | 78.5645 | 111.9808 | 191.7369 | 189.507 | 203.2056 | 160.3972 | 140.3589 | 106.0958 | 138.4495 | 124.3014 | 142.8467 | 139.3206 | 195.1672 | 154.8136 | 127.3589 | 137.2439 | 163.622 | 152.9425 | 158.4547 | 169.6411 | 200.378 | 253.1341 | 135.5453 | 149.2439 | 204.0383 | 213.2247 | 192.1341 | 164.0958 | 175.6986 | 374.0192 | 148.6794 | 151.115 | 147.4355 | 225.3972 | 134.9617 | 122.3397 | 138.9617 | 143.3972 | 181.3206 | 175.2056 | 160.3206 | 132.378 | 108.7369 | 118.2247 | 115.2056 | 126.2056 | 239.1672 | 196.4739 | 214.1098 | 373.0383 | 203.0767 | 207 | 235.0958 | 350.5784 | 161.7178 | 75.6028 | 74.9617 | 161.7178 | 100.8902 | 100.9617 | 122.0767 | 132.2056 | 110.7317 | 94.4355 | 102.9042 | 119.7369 | 80.622 | 66.0383 | 73.6986 | 107.3014 | 103.1341 | 90.9617 | 89.7944 | 84.0192 | 221.8659 | 219.4686 | 212.6028 | 428.7561 | 164.6028 | 189.3014 | 196.4164 | 396.3589 | 274.7317 | 198.6794 | 174.2631 | 365.3589
1927 | 23.1914 | 133.876 | 194.4514 | 208.5486 | 193.9646 | 136.3983 | 146.3628 | 148.9115 | 168.7965 | 141.416 | 122.2035 | 107.3983 | 179 | 190.4778 | 145.3628 | 151.2743 | 221.3097 | 152.8583 | 126.6194 | 129.9469 | 189.8937 | 134.8937 | 274.4246 | 279.6549 | 192.832 | 200.6903 | 176.7611 | 202.4423 | 332.7257 | 114.6372 | 110.6017 | 95.5486 | 128.1858 | 124.9823 | 129.3274 | 136.6903 | 96.2743 | 224.7788 | 224.7965 | 218.708 | 270.9823 | 123.9823 | 124.5045 | 124 | 121.8937 | 63.876 | 66.168 | 75.2566 | 116.3628 | 183.292 | 178.4955 | 224.124 | 161.0354 | 122.1858 | 120.5045 | 125.1149 | 146.1771 | 142.9292 | 155.4691 | 180.4601 | 173.2302 | 118.6017 | 131.2389 | 148.5132 | 167.4337 | 157.3806 | 162.2566 | 193.292 | 268.1948 | 133.9292 | 133.1503 | 197.9646 | 221.6903 | 179.6194 | 181.8674 | 180.0531 | 374.3451 | 133.2035 | 143.6549 | 156.7965 | 227.3806 | 137.6726 | 132.7611 | 124.1858 | 135.9646 | 182.6726 | 172.584 | 166.3806 | 127.9823 | 92.8937 | 113.2389 | 104.5132 | 118.876 | 229.8406 | 220.6108 | 219.9469 | 356.876 | 224.5577 | 211.0354 | 232.6903 | 348.1149 | 158.3274 | 74.292 | 80.3628 | 158.3274 | 105.8142 | 100.3097 | 119.3451 | 113.4423 | 108.1858 | 107.1594 | 110.0177 | 107.2566 | 84.3451 | 74.0885 | 54.5132 | 94.876 | 89.9469 | 91.6549 | 94.4337 | 75.6017 | 247.1594 | 221.9292 | 215.3274 | 438.7965 | 155.2212 | 200.416 | 188.6372 | 390.9646 | 268.8229 | 197.3274 | 186.0708 | 407.717
1928 | 23.2024 | 157.8187 | 186.6179 | 197.206 | 181.7271 | 129.0422 | 137.0422 | 162.4912 | 167.0968 | 145.3151 | 116.3697 | 112.3697 | 182.4243 | 187.5335 | 148.3697 | 134.3028 | 213.8363 | 147.7394 | 137.0124 | 133.1637 | 190.1637 | 142.0124 | 280.861 | 298.861 | 171.8661 | 207.4789 | 186.794 | 213.2853 | 338.7394 | 113.6849 | 107.6849 | 103.3821 | 133.0124 | 132.2183 | 136.5335 | 138.6849 | 82.7271 | 207.7147 | 207.345 | 214.8363 | 281.6426 | 122.3151 | 123.7271 | 117.8363 | 121.4789 | 69.3275 | 59.0546 | 72.3697 | 118.6849 | 189.2183 | 199.8187 | 221.4119 | 169.1637 | 116.0546 | 120.4119 | 124.6426 | 122.4418 | 146.5335 | 159.5757 | 179.2729 | 170.2482 | 134.861 | 136.2729 | 150.2729 | 155.8785 | 169.5881 | 168.2729 | 186.8908 | 280.3697 | 140.2729 | 133.2729 | 197.3697 | 214.7817 | 179.7394 | 177.2482 | 178.8908 | 385.2729 | 129.4243 | 148.4789 | 155.7817 | 228.6303 | 138.6849 | 135.2606 | 118.0546 | 155.2307 | 165.8785 | 173.4789 | 176.5335 | 138.6426 | 117.1339 | 116.2183 | 104.9032 | 120.2183 | 211.4119 | 222.3028 | 208.7817 | 354.794 | 232.206 | 185.6056 | 212.0845 | 324.3574 | 155.2606 | 65.8363 | 88.8486 | 155.2606 | 109.8908 | 111.3275 | 114.5211 | 110.5881 | 100 | 106.4119 | 103.4665 | 113.2729 | 69.2482 | 68.7271 | 65.861 | 94.1637 | 102.0669 | 91.6849 | 115.0669 | 85.0124 | 243.6725 | 191.9753 | 205.4119 | 435.7271 | 157.9032 | 190.6179 | 178.0968 | 412.9701 | 281.6004 | 201.794 | 193.4243 | 422.8363
1929 | 23.2134 | 163.6342 | 187.6687 | 197.0915 | 197.9146 | 145.6117 | 147.3373 | 164.7256 | 151.1484 | 154.063 | 110.0285 | 113 | 170.057 | 187.5142 | 148.3028 | 148.2174 | 211.6972 | 150.3943 | 143 | 147.8516 | 202.063 | 145.2114 | 304.6402 | 316.3658 | 151.0285 | 212.7886 | 202.2459 | 219.1199 | 333.0285 | 110.5142 | 104.5142 | 87.6911 | 136.2114 | 122.7541 | 130.2399 | 136.9085 | 88.4573 | 221.4003 | 207.1544 | 214.7886 | 288.3943 | 130.3658 | 127.3658 | 121.2744 | 124 | 76.3943 | 58.6057 | 75.0915 | 102.2684 | 209.7316 | 199.3598 | 218.3943 | 179.6687 | 126.063 | 111.1199 | 113.2684 | 108.3943 | 185.5548 | 147.5427 | 184.6972 | 168.5772 | 139.8171 | 129.1484 | 173.126 | 154.2744 | 160.7541 | 182.063 | 195.4573 | 286.5772 | 151.2744 | 153.3373 | 203.5772 | 217.9715 | 183.0915 | 165.8171 | 169.3313 | 344.6851 | 144.943 | 136.3598 | 147.1199 | 209.874 | 137.6057 | 122.1484 | 134.3373 | 160.1199 | 188.6747 | 176 | 168.8456 | 134.937 | 123.3313 | 100.4797 | 106.9085 | 107.9655 | 225.8231 | 205.5427 | 199.4228 | 357 | 224.4228 | 200.8861 | 225.7032 | 320.3943 | 151.2114 | 60.9085 | 74.5712 | 151.2114 | 117.063 | 107.2399 | 109.9085 | 115 | 105.5772 | 97.8171 | 101.3943 | 97.7826 | 71.063 | 63.3028 | 59.6627 | 101.1829 | 115.9715 | 96.1829 | 108.7541 | 71.4797 | 227.5427 | 195.3373 | 204.4858 | 438.6687 | 155.7256 | 177.0285 | 184.4573 | 424 | 308.7316 | 194.937 | 216.6117 | 400.4797
1930 | 23.2244 | 170.6416 | 178.23 | 188.0678 | 194.9395 | 152.1622 | 149.7433 | 155.615 | 160.1889 | 155.8717 | 99.9056 | 126.4794 | 166.8378 | 185.2906 | 140.1961 | 163.385 | 215.5472 | 165.1889 | 152.2228 | 145.1961 | 214.5133 | 153.9322 | 303.9056 | 283.8184 | 171.4116 | 202.6489 | 214.385 | 203.5206 | 338.5133 | 102.615 | 98.0339 | 104.2494 | 142.8039 | 131.8983 | 126.7094 | 130.3245 | 85.1961 | 207.1695 | 199.9395 | 223.8039 | 282.615 | 122.6489 | 131 | 114.0678 | 116.1961 | 78.4189 | 57.2906 | 73.1622 | 112.0266 | 211.3245 | 180.9395 | 217.5811 | 173.7772 | 124.3245 | 133.1211 | 124.4455 | 126.7361 | 157.724 | 145.8378 | 172.9395 | 163.9056 | 134.4528 | 137.4794 | 146.9467 | 165.5133 | 155 | 171.1017 | 187.23 | 269.845 | 166.0605 | 138.0073 | 212.385 | 230.2228 | 194.6416 | 172.5133 | 166 | 357.6344 | 145.3245 | 142.77 | 150.6755 | 214.77 | 136.2906 | 134.0266 | 143.8378 | 165.0944 | 180.0073 | 173.1622 | 166.8378 | 147.3172 | 109.3584 | 104.0605 | 95.3584 | 105.2567 | 236.1283 | 205.9661 | 211.1889 | 371.1889 | 220.5811 | 189.5884 | 221.1017 | 333.77 | 157.8039 | 62.1283 | 74.0944 | 157.8039 | 125.2567 | 107.9661 | 106.8717 | 123.5133 | 95.23 | 105.2228 | 95.615 | 104.6077 | 71.4528 | 66.5472 | 67.8983 | 93.7772 | 107.6489 | 102.2567 | 112.9322 | 83.5738 | 232.2228 | 212.6416 | 203.8717 | 439.8717 | 163.6416 | 198.1211 | 188.2906 | 425.4189 | 316.7094 | 206.6077 | 211.1017 | 430.3099
1931 | 23.2354 | 154.0697 | 178.0521 | 181.1131 | 207.3216 | 156.0521 | 153.7739 | 153 | 155.174 | 160.7739 | 108.5477 | 121.174 | 184.5999 | 202.3216 | 165.1476 | 164.5565 | 245.8694 | 144.2958 | 149.5044 | 139.1131 | 197.7914 | 164.4956 | 297.3913 | 311.5823 | 168.7305 | 213.8781 | 210.5044 | 209.5477 | 342.7217 | 109.3825 | 103.2173 | 105.6175 | 127.2349 | 145.2173 | 135.6608 | 133.7739 | 97.8781 | 215.7651 | 203.6608 | 214.7305 | 256.9045 | 133.8781 | 128.8348 | 117.9391 | 116.6087 | 79.7217 | 72.1564 | 66.9479 | 103.1219 | 206.1131 | 191.8781 | 218.4435 | 179.3825 | 132.1043 | 131.4523 | 121.174 | 124.6175 | 155.1564 | 136.174 | 187.4868 | 177.5999 | 140.2173 | 132.174 | 166.9215 | 164.6696 | 170.8781 | 148.4001 | 177.6696 | 264.1652 | 169.5565 | 127.5565 | 219.3304 | 233.2783 | 207.6608 | 167.3392 | 181.1564 | 379.6608 | 145.1652 | 153.0521 | 159.4956 | 237.3216 | 141.7739 | 141 | 129.8436 | 175.2173 | 187.5999 | 152.5132 | 186.7651 | 151.8348 | 97.0609 | 98.8957 | 87.3913 | 109.1652 | 229.7827 | 194.2871 | 215.5565 | 389.2695 | 215.6696 | 219.6954 | 226.5477 | 320.2349 | 157.3913 | 62.2783 | 72.6696 | 157.3913 | 114.0088 | 121.5477 | 103.8348 | 126.2783 | 131.8606 | 109 | 95.8869 | 131.9303 | 72.1652 | 55.7305 | 71.8348 | 82.0609 | 108.7739 | 99.6696 | 105.4523 | 67.461 | 259.0955 | 197.5132 | 197.9479 | 445.4956 | 184.5999 | 190.6784 | 200.2695 | 399.2958 | 317.7217 | 187.7393 | 221.5999 | 396.9221
1932 | 23.2464 | 148.9361 | 199.0847 | 187.3403 | 192.7133 | 152.1278 | 136.1813 | 147.8618 | 141.4577 | 154.1917 | 110.7979 | 111.8618 | 174.8514 | 170.8307 | 168.6597 | 167.6701 | 230.5751 | 162.627 | 131.5855 | 138.734 | 192.9361 | 160.3938 | 289.3938 | 330.9361 | 166.2021 | 207.5216 | 209.4681 | 215.4681 | 333.4577 | 114.4681 | 101.5958 | 97.5958 | 136.8826 | 147.266 | 147.0743 | 130.1278 | 92.9896 | 225.2021 | 209.9361 | 223.2125 | 241.3938 | 140.734 | 125.7979 | 136.4145 | 130.4784 | 71.1917 | 64.0535 | 69.4042 | 99.2021 | 209.4042 | 209.0104 | 216.7979 | 175.6597 | 127.9257 | 136.5423 | 125.8083 | 123.2021 | 169.8083 | 161.361 | 183.2556 | 167.1174 | 126.8514 | 128 | 171.1917 | 161.5319 | 165.2556 | 150.0743 | 193.6166 | 295.8291 | 156.5216 | 131.4042 | 198.2452 | 216.1174 | 215.4042 | 192.627 | 203.1486 | 394.0104 | 151.8722 | 151.3299 | 153.1917 | 254.2764 | 142.5319 | 143.9361 | 127.6701 | 184.6062 | 201.3403 | 183.9033 | 186.6597 | 147.3299 | 95.4681 | 93.5319 | 92.6062 | 121.0104 | 219.6597 | 201.4784 | 218.6701 | 384.4577 | 214.734 | 221.0535 | 217.7875 | 343.095 | 166.2764 | 71.5423 | 69.5319 | 166.2764 | 112.6701 | 133.3403 | 103 | 123.7979 | 128.1813 | 116.3403 | 93.3299 | 131.9257 | 74.4681 | 70.0847 | 84.2125 | 81.2021 | 108.7979 | 103.8722 | 111.2764 | 77.0847 | 224.6925 | 206.1486 | 205.5423 | 451.6701 | 183.6597 | 197.9465 | 193.2556 | 420.5631 | 320.2021 | 173.5958 | 215.5216 | 406.627
1933 | 23.2574 | 156.7168 | 197.0442 | 197.4632 | 190.9705 | 155.9705 | 139.4484 | 175.8526 | 166.36 | 183.0916 | 94.3274 | 108.5074 | 171.2389 | 190.5842 | 152.5663 | 152.581 | 218.7611 | 171.5074 | 120.7758 | 124.82 | 198.4779 | 155.0147 | 286.2537 | 300.6547 | 187.8969 | 226.8821 | 204.7758 | 195.1031 | 299.401 | 108.2832 | 100.7463 | 94.5074 | 134.0442 | 165.6579 | 114.1768 | 139.9411 | 85.2684 | 229.7316 | 214.7316 | 214.5663 | 271.0916 | 139.5074 | 116.7905 | 126.3274 | 127.5368 | 75.4632 | 72.4337 | 59.8053 | 111.1947 | 253.54 | 210.7611 | 212.2684 | 195.3895 | 121.2684 | 115.3716 | 109.5958 | 138.18 | 159.5663 | 160.8053 | 181.2389 | 173.6874 | 140.4042 | 124.2684 | 165.7611 | 163.9853 | 163.9853 | 171.6579 | 193.2832 | 284.6105 | 166.18 | 129.2684 | 198.9558 | 221.1947 | 188.64 | 203 | 191.0884 | 388.2979 | 136.0884 | 158.9558 | 142.5368 | 246.0589 | 152.4484 | 142.0147 | 137.2095 | 167.5958 | 210.7168 | 199.4926 | 204.8969 | 182.5695 | 105.7021 | 94.4926 | 103.9558 | 117.5368 | 255.0621 | 248.54 | 222.2389 | 378.0147 | 227.6874 | 204.8053 | 233.8969 | 330.8642 | 160.2979 | 76.4926 | 77.2095 | 160.2979 | 106.5368 | 111.3716 | 123.1505 | 130.4632 | 107.5663 | 110.0442 | 111.4042 | 117.0589 | 65.2979 | 67.2979 | 69.5958 | 76.7758 | 105.7611 | 98.5368 | 114.2537 | 92.2095 | 235.8674 | 218.7168 | 197.8053 | 447.0295 | 176.5221 | 200.7611 | 186.0147 | 421.5516 | 316.5221 | 189.9116 | 187.1179 | 422.9705
1934 | 23.2684 | 136.2418 | 191.7242 | 186.3451 | 177.8279 | 146.6209 | 158.9308 | 157.9315 | 167.6553 | 179.5865 | 102.6549 | 116.3447 | 163.6553 | 218.6894 | 147.2414 | 148.5172 | 236.9652 | 189.2066 | 140.241 | 144.2754 | 202.2758 | 160.0689 | 258.6901 | 321.1029 | 192.7242 | 219.0692 | 220.448 | 232.7578 | 314.7926 | 127.241 | 110.1033 | 109.9308 | 127.207 | 131.0352 | 130.0685 | 131.1037 | 92.3447 | 230.2414 | 225.8619 | 232.7582 | 277.4484 | 117.759 | 107.1725 | 120.2414 | 128.793 | 72.6897 | 67.1381 | 62.8275 | 122.5861 | 249.7934 | 203.9311 | 234.5168 | 193.1381 | 138.2066 | 119.1377 | 124.241 | 127.0692 | 171.6894 | 145.6209 | 189.5861 | 178.7586 | 143.9656 | 144.9996 | 174.1033 | 178.6549 | 151.3451 | 170.4139 | 199.3447 | 283.0689 | 161.1381 | 137.8619 | 215.6549 | 229.5516 | 201.7582 | 198.4484 | 200.9308 | 397.1377 | 137.5861 | 147.5865 | 144.5516 | 254.1377 | 170.4135 | 144.793 | 145.3103 | 167.8275 | 213 | 205.3103 | 187.7246 | 160.1041 | 115.8275 | 110.9308 | 109.2758 | 121.0689 | 246.759 | 235.6901 | 210.8623 | 373.9656 | 244.1377 | 214.6549 | 244.793 | 341.9652 | 177.4824 | 77 | 74.6897 | 177.4824 | 125.9996 | 135.6201 | 113.3106 | 129.9656 | 116.6549 | 123.6894 | 96.0004 | 115.2758 | 94.6201 | 62.4828 | 63.7586 | 90.1721 | 114.1033 | 99.793 | 93.5176 | 88.1725 | 246.5172 | 224.0344 | 192.4828 | 458.6549 | 181.8275 | 193.9311 | 205.4824 | 404.3451 | 326.3791 | 192.207 | 223.7578 | 441.6894
1935 | 23.2794 | 148.2725 | 180.2074 | 191.2508 | 183.7926 | 142.2291 | 161.6873 | 154.3963 | 181.9598 | 158.0402 | 105.4582 | 122.0836 | 167.9381 | 206.2693 | 144.6873 | 152.0836 | 228.3529 | 155.6842 | 150.8545 | 136.582 | 216.8762 | 155.0619 | 328.6316 | 318.6656 | 170.4148 | 206.291 | 207.4984 | 233.8947 | 316.0619 | 131.6873 | 109.1455 | 106.5201 | 148.3561 | 137.2725 | 128.2074 | 142.418 | 113.5016 | 242.3344 | 235.1672 | 202.9969 | 274.4582 | 137.9814 | 134.2941 | 132.3344 | 108.4148 | 78.709 | 85.5852 | 83.5016 | 108.8111 | 217.0186 | 204.3127 | 226.582 | 218.5232 | 141.6873 | 130.709 | 137.1672 | 118.9164 | 173.9164 | 165.8978 | 202.7926 | 172.0619 | 175.3777 | 140.4365 | 174.6873 | 175.291 | 160.1053 | 171.0836 | 210.4799 | 265.7275 | 164.1672 | 135.6037 | 219.2291 | 233.3127 | 196.6656 | 186.9783 | 191.3529 | 392.5201 | 146.1672 | 152.2508 | 135.2074 | 237.1857 | 147.2477 | 144.4582 | 133.8947 | 166.1455 | 216.0836 | 206.2291 | 169.2074 | 145.1455 | 115.6873 | 102.1238 | 112.3127 | 123 | 236.9164 | 241.6471 | 216.2508 | 360.6656 | 248.7709 | 206.6656 | 213.6223 | 334.8947 | 199.418 | 75.4582 | 74.5418 | 199.418 | 122.9783 | 104.3715 | 120.3344 | 135.9381 | 124.0836 | 110.4984 | 92.8545 | 116.7709 | 81.1022 | 72.0217 | 74.0217 | 91.9164 | 124.709 | 104.0836 | 97.0217 | 89.8545 | 243.9164 | 211.8947 | 223.6068 | 463.7709 | 186.3127 | 204.3344 | 205.8328 | 412.3344 | 323.0619 | 202.5635 | 204.8514 | 437.7492
1936 | 23.2904 | 172.0134 | 177 | 200.2655 | 201.8806 | 150.6151 | 159.4336 | 164.6151 | 172.1194 | 172.5797 | 101.8672 | 117.5177 | 190.3629 | 224.4957 | 135.3849 | 147.5177 | 244.3629 | 151.8319 | 153.5664 | 141.3983 | 206.1194 | 145.1681 | 315.1901 | 329.0974 | 166.3496 | 218.0974 | 212.9647 | 210.4203 | 332.0134 | 128.6504 | 103.3009 | 100.0841 | 127.5884 | 149.2655 | 148.4957 | 143.0841 | 100.9866 | 231.1194 | 215.0707 | 206.0974 | 292.7965 | 140.5177 | 127.3362 | 130.5177 | 112.9647 | 81.7832 | 71.6371 | 73.3362 | 114.9647 | 216.8319 | 193.2522 | 208.6858 | 191.7565 | 148.8319 | 134.5664 | 126.469 | 103.9026 | 147.1548 | 148.7211 | 201.3009 | 180.9647 | 169.3362 | 140.9159 | 188.8806 | 159.6858 | 174.9647 | 176.6991 | 201.2522 | 273.3142 | 158.9513 | 137.9159 | 217.4336 | 235.5664 | 208.6638 | 194.1815 | 193.2655 | 411.1461 | 150.3496 | 141.6858 | 147.6638 | 223.1681 | 157.7965 | 147.1328 | 141.7478 | 181.447 | 211.5177 | 215.3983 | 182.447 | 142.4336 | 130.6638 | 116.7965 | 115.3496 | 131.6151 | 261.8452 | 217.8052 | 212.7345 | 393.8099 | 216.8892 | 229.6284 | 214.1815 | 339.6151 | 193.0353 | 72.6504 | 70.3009 | 193.0353 | 123.1328 | 114.7125 | 124.7832 | 127.0353 | 113.2522 | 102.6504 | 108.0974 | 124.8319 | 71.6504 | 53.8539 | 74.6504 | 104.3142 | 107.4203 | 122.2302 | 102.3496 | 84.7345 | 261.0134 | 229.9293 | 223.6017 | 435.022 | 176.0353 | 235.4116 | 189.1194 | 393.2875 | 326.4823 | 206.7832 | 196.5664 | 469.4603
1937 | 23.3014 | 151.5446 | 183.3638 | 208.3638 | 186.9085 | 145.8407 | 164.5684 | 155.8633 | 187.0915 | 173.2271 | 114.5232 | 108.0452 | 175.3175 | 221.4542 | 129.8181 | 169.8644 | 257.1593 | 157.9774 | 134.9085 | 155.9322 | 208.3638 | 158.9096 | 316.9322 | 324.2497 | 174.1593 | 206.0904 | 220.7729 | 205 | 338.591 | 127.2045 | 113.1367 | 110.1367 | 152.6147 | 136.6814 | 128.7491 | 134.8407 | 115.0915 | 232.5684 | 206.6136 | 214.7729 | 286.0678 | 129.4542 | 114.2497 | 133.7729 | 127.1367 | 86.7729 | 81.9096 | 62.6362 | 111.6362 | 220.591 | 195.5684 | 216.1367 | 180.409 | 132.704 | 133.409 | 115.8407 | 126.2509 | 170.2283 | 157.1141 | 207.9548 | 199.1141 | 166.591 | 130.8633 | 161.9763 | 184.6373 | 162.0904 | 185.1593 | 194.8181 | 299.2735 | 147.4542 | 148.5458 | 207.4542 | 243.9548 | 207.4316 | 189.0452 | 191.8181 | 397.113 | 162.1367 | 149.9322 | 143.2497 | 232.9322 | 149.4768 | 147.2045 | 125.113 | 182.8181 | 215.5684 | 229.1367 | 188.591 | 158.7051 | 123.8633 | 116.4316 | 106.4542 | 122.8633 | 269.7955 | 232.2735 | 213.3864 | 388.8859 | 228.6825 | 224.2723 | 212.2271 | 340.409 | 187.6136 | 75.9774 | 71.3864 | 187.6136 | 127.1819 | 110.6588 | 121.0226 | 125.591 | 123.5232 | 105.1819 | 109.6136 | 137.3412 | 71 | 72.6599 | 74.7955 | 108.7955 | 148.1379 | 126.2045 | 114.9322 | 89.3638 | 231.7943 | 228.0452 | 222.591 | 453.2509 | 183.3412 | 267.6599 | 225.5695 | 404.5006 | 353.4554 | 211.7729 | 211.3186 | 486.1593
1938 | 23.3124 | 134.499 | 179.3456 | 205.9611 | 191.6933 | 123.8057 | 164.3845 | 145.7301 | 177.4601 | 184.1166 | 124.4622 | 113.2699 | 186.1554 | 210.9223 | 136.2699 | 176 | 221.8425 | 190.5031 | 129.1922 | 149.3067 | 205.1534 | 169.4622 | 324.0389 | 338.9632 | 173.5767 | 194.7301 | 212.3067 | 208.2311 | 309.9202 | 122.9611 | 111.1534 | 112.1922 | 146.4601 | 149.9632 | 132.501 | 131.3845 | 86.0736 | 225.1145 | 193.8834 | 223.2699 | 291.8855 | 135.0777 | 138.6565 | 137.4233 | 111.4213 | 83.1534 | 74.6912 | 55.3456 | 116.4622 | 227.4622 | 222.8487 | 200.4213 | 191.3088 | 140.9243 | 117.6523 | 109.1534 | 145.9243 | 160.229 | 181.1943 | 211.6155 | 181.9979 | 161.3456 | 143.3477 | 173.3865 | 185.5378 | 154.7689 | 177.3067 | 202.8855 | 304.1922 | 149.8466 | 157.4622 | 225.1943 | 232.2679 | 203.5767 | 191.8466 | 205.5399 | 406.5399 | 167.4233 | 154.6155 | 165.2331 | 240.8466 | 162.1554 | 168.8098 | 130.501 | 154.5358 | 221.8466 | 221.499 | 197.0777 | 150.0757 | 116.1534 | 115 | 102.3845 | 128.8855 | 236.8814 | 230.7301 | 203.499 | 364.8057 | 254.1943 | 211.3067 | 216.6544 | 338.3845 | 186.1922 | 77.8078 | 85.7321 | 186.1922 | 136.8855 | 118.501 | 115.1534 | 131.6544 | 118.9223 | 100.3456 | 100.1145 | 135.1534 | 67.7689 | 76.5767 | 66.1145 | 121.9243 | 144.6523 | 99.3435 | 119.6155 | 98.2699 | 260.9653 | 212.2679 | 238.3477 | 469.6933 | 180.3456 | 231.9959 | 211.7669 | 415.4622 | 338.1902 | 187.9591 | 208.5378 | 453.2658
1939 | 23.3234 | 140.2006 | 209.1624 | 190.2389 | 187.4395 | 154.2627 | 148.4188 | 154.6608 | 166.6194 | 187 | 124.3599 | 132.2213 | 187.5398 | 209.8201 | 132.2596 | 149.758 | 236.7818 | 174.2182 | 151.1417 | 147.8201 | 209.7404 | 168.5398 | 306.9586 | 299.5367 | 178.7404 | 219.242 | 233.7818 | 218.8408 | 379.2658 | 122 | 120.6608 | 111.1799 | 133.9793 | 152.3599 | 133.3599 | 136.7404 | 91.121 | 243.5016 | 207.4013 | 199.5781 | 283.3392 | 128.7994 | 137.6194 | 126.5191 | 139.8025 | 84.4602 | 76.9204 | 68.7611 | 100.7787 | 229 | 203.5781 | 218.1417 | 216.1417 | 153.0207 | 108.2596 | 126.8614 | 137.5191 | 196.1831 | 162.2182 | 188.2182 | 170.4395 | 152.6194 | 139.6194 | 177.1799 | 181.5398 | 157.2803 | 170.0796 | 223.8614 | 316.3009 | 163.3009 | 161.4602 | 217.6991 | 228.1799 | 212.0207 | 183.1592 | 185.2182 | 391.9586 | 170.4602 | 174.6815 | 143.9379 | 245.2803 | 160.2596 | 151.8583 | 140.3806 | 187.363 | 227.1003 | 235.4013 | 177.6784 | 146.1799 | 119.9204 | 113.3599 | 87.2389 | 122.7994 | 260.1624 | 228.1799 | 211.6608 | 418.2244 | 263.1003 | 215.5605 | 237.6815 | 338 | 215.5223 | 88.6608 | 76.6991 | 215.5223 | 160.3216 | 121.8201 | 107.4395 | 121.5191 | 113.7197 | 110.4809 | 111.9411 | 125.7994 | 80.121 | 66.9793 | 75.4809 | 102.0382 | 133.6194 | 132.363 | 119.1799 | 87.6991 | 248.6784 | 259.8439 | 232.9793 | 435.9172 | 207.7022 | 222.8201 | 210.1003 | 421.9204 | 381.3837 | 213.1624 | 210.2803 | 449.9204
1940 | 23.3344 | 151.1559 | 191.0294 | 212.803 | 205.9765 | 150.3471 | 160.8147 | 156.1676 | 183.3118 | 177.8441 | 134.8206 | 123.5147 | 186.1676 | 232.4735 | 158.4677 | 142.3353 | 248.6588 | 172.3294 | 160.1618 | 149.6647 | 220.9882 | 177.9882 | 342.1206 | 313.3059 | 184.9941 | 221.6706 | 241.4971 | 201.0235 | 338.5676 | 119.5029 | 103.0235 | 120.9882 | 123.6765 | 140.3471 | 142.9882 | 145.4912 | 100.6588 | 254.6588 | 242.6294 | 214.8088 | 289.3235 | 127 | 129.3412 | 133.1559 | 139.5088 | 74.1794 | 78 | 80.3235 | 104.4912 | 229.8324 | 206.3235 | 240.4794 | 204.3529 | 142.5147 | 116.1559 | 123.5088 | 141.6588 | 177.5323 | 162.8265 | 213.7971 | 200.6294 | 173.4735 | 136.3353 | 212.7912 | 166.85 | 179.6412 | 186.4794 | 229.6647 | 319 | 170.1618 | 137.0294 | 205.0118 | 236.3235 | 202.3471 | 206.803 | 189.9882 | 387.1676 | 160.1794 | 176.5029 | 149.6529 | 260.9824 | 153.1735 | 161.15 | 141.1676 | 162.7059 | 224.6706 | 256.4794 | 212.953 | 173.4677 | 112.6765 | 95.5206 | 102.3118 | 127.6588 | 271.1618 | 210.5206 | 216.4971 | 401.0353 | 272.3235 | 218.6647 | 236.1735 | 338.8324 | 196.197 | 70.1912 | 79.8265 | 196.197 | 143.3588 | 110.3471 | 119.3177 | 128.9882 | 130.4794 | 108.8382 | 99.1853 | 126.4971 | 75.5088 | 76.6529 | 68.8441 | 106.9882 | 132.8324 | 127.6823 | 120.6647 | 101.6471 | 273.9647 | 235.2088 | 210.1912 | 431.3294 | 211.5029 | 234.6529 | 217.6588 | 398.0294 | 354.5441 | 216.6706 | 216.8265 | 451.8324
1941 | 23.3454 | 167.359 | 186.8446 | 203.641 | 213.3786 | 157.2911 | 158.0875 | 142.4857 | 157.4374 | 180.2232 | 109.9714 | 126.9125 | 189.3786 | 214.1946 | 148.7964 | 144.5339 | 242.3982 | 196.65 | 154.2428 | 122.9714 | 230.6018 | 169.0196 | 338.175 | 311.2428 | 160.6607 | 219.3107 | 225.1071 | 221.4947 | 372.6108 | 121.5339 | 120.1161 | 111.175 | 143.1161 | 132.9321 | 140.7768 | 144.4661 | 100.3107 | 223.9035 | 238.8643 | 219 | 295.2232 | 148.9608 | 137.2911 | 129.9321 | 117.7285 | 72.8446 | 63.641 | 71.0196 | 102.6214 | 223.2428 | 226.5822 | 255.825 | 192.5536 | 140 | 132.359 | 115.2428 | 141.3107 | 194.8054 | 152.175 | 207.3303 | 189.2625 | 187.2911 | 150.359 | 199.7285 | 202.009 | 178.9321 | 191.6893 | 214.7964 | 290.2821 | 178.6018 | 158.184 | 208.9125 | 234.6214 | 200.8446 | 184.9714 | 182.7089 | 393.7572 | 162.2232 | 183.6018 | 157.9125 | 277.5143 | 152.8446 | 162.3107 | 156.2036 | 186.4072 | 256.0965 | 238.8839 | 202.4178 | 165.4857 | 130.4268 | 123.2519 | 95.0196 | 133.2232 | 284.6697 | 260.2126 | 224.6018 | 376.4178 | 272.3107 | 207.175 | 204.5928 | 355.8929 | 210.4268 | 89.65 | 88.6018 | 210.4268 | 161.8054 | 126.5822 | 125.3786 | 137.7572 | 125.5536 | 124.0482 | 108.6697 | 154.8733 | 71.4661 | 71.3982 | 76.2911 | 120.825 | 112.7285 | 119.9321 | 129.4464 | 91.4857 | 247.9035 | 226.3107 | 231.3393 | 432 | 204.2428 | 233.6214 | 193.6607 | 403.1357 | 386.6983 | 220.2232 | 223.9125 | 451.1554
1942 | 23.3564 | 138.2934 | 206.7095 | 201 | 185.7211 | 160.7139 | 145.0029 | 171.7066 | 185.4205 | 143.2948 | 127.2803 | 122.8584 | 190 | 235.7081 | 140.8584 | 136.4306 | 241 | 201.8569 | 154.7139 | 141.9942 | 228.5723 | 172.9985 | 298.2948 | 335.7081 | 158.5708 | 207.0029 | 244.2803 | 221.7153 | 354.435 | 112.5737 | 124 | 109 | 137.5737 | 170.5621 | 136.5723 | 138.0015 | 96.5723 | 268.5592 | 236.1431 | 207.0029 | 296.8569 | 151.2861 | 127.8598 | 143.5679 | 134.5665 | 74.7139 | 64.4277 | 66.4292 | 105.4277 | 263.1329 | 212.0044 | 243.4321 | 201.2832 | 145.1416 | 118.7182 | 142.2789 | 129.0029 | 191.2876 | 174.8511 | 221.2818 | 180.0015 | 185.5723 | 157.2847 | 191.7153 | 184.1489 | 164.289 | 204.854 | 205.1445 | 274.7168 | 167.146 | 151.8598 | 212.5708 | 236.5708 | 218.1387 | 193.711 | 200.7095 | 400.1416 | 155.2876 | 149.8656 | 169.2832 | 225.1561 | 158.9985 | 155.1445 | 146.146 | 175.7182 | 234.5781 | 267.5636 | 202.4277 | 154.4306 | 112.5766 | 119.5737 | 97.2847 | 129.7153 | 263.8627 | 264.0015 | 210.5752 | 424.4161 | 294.2803 | 216.9971 | 256.4147 | 327.2934 | 198.5752 | 99.9985 | 81.4306 | 198.5752 | 148.8613 | 130 | 120.8584 | 151.854 | 141.9956 | 124.4292 | 123.854 | 150.5737 | 78.7124 | 92.2803 | 72.0015 | 106.7182 | 130.4234 | 104.4321 | 115.5752 | 90.7139 | 270.2789 | 226 | 221.4321 | 434.5708 | 215.854 | 232.1431 | 211.2803 | 415.2832 | 372.5766 | 244.9942 | 211.289 | 444.1445
1943 | 23.3674 | 152.123 | 196.0924 | 201.8692 | 215.7691 | 169.6923 | 172.5537 | 180.4769 | 164.054 | 170.0306 | 131 | 123.7385 | 184.7846 | 205.2309 | 144.3461 | 149.7769 | 266.2076 | 192.4385 | 164.5615 | 146.8692 | 207.1386 | 157.4847 | 314.5999 | 293.0618 | 152.9154 | 228.4691 | 244.5231 | 238.3845 | 340.4385 | 123.1692 | 130.9538 | 115.0846 | 148.1692 | 152.6617 | 134.2615 | 128.3077 | 103.823 | 240.4925 | 236.8692 | 210.2154 | 299.6077 | 128.4001 | 136.4307 | 125.1386 | 155.3845 | 88.9076 | 67.6077 | 67.7385 | 112.0846 | 272.6077 | 227.2538 | 243.6077 | 209.9538 | 147.7385 | 127.2999 | 126.1386 | 124.3923 | 196.9538 | 168.5693 | 182.2771 | 192.9076 | 176.3077 | 163.2154 | 197.9538 | 211.2922 | 165.4769 | 204.3923 | 220.5153 | 294.7307 | 182.3845 | 163.0384 | 215.6077 | 237.8692 | 253.1614 | 210.7769 | 201.3923 | 425.3383 | 167.9076 | 173.5537 | 150.1386 | 252.5075 | 175.6461 | 180.0768 | 149.2154 | 170.3923 | 238.6923 | 225.1925 | 238.6383 | 138.2231 | 111.6077 | 111.0462 | 98.8692 | 140.2999 | 237.4001 | 272.5615 | 212.3461 | 424.3077 | 268.4463 | 238.9922 | 259.0462 | 395.015 | 202.9538 | 95.7846 | 80.8692 | 202.9538 | 142.5231 | 130.8692 | 126.9538 | 134.877 | 102.4079 | 127.4769 | 110.3539 | 150.7385 | 67.8308 | 82.9616 | 73.6077 | 97.0462 | 143.5615 | 122.8614 | 108.6539 | 71.877 | 274.1308 | 216.4385 | 251.1614 | 463.6845 | 210.177 | 217.2231 | 215 | 382.2309 | 364.6539 | 235.9616 | 218.5615 | 454.2999
1944 | 23.3784 | 152.3554 | 183.4218 | 188.7772 | 197.199 | 153.3696 | 180.5261 | 171.3033 | 177.6304 | 166.1848 | 116.0142 | 127.5261 | 177.8294 | 231.7346 | 145 | 163.4597 | 233.8578 | 204.2228 | 168.6446 | 130.2512 | 194.3033 | 178.801 | 340.9195 | 302.7488 | 163.4597 | 246.9858 | 270.4455 | 230.4218 | 351.3412 | 105.6066 | 109.0805 | 126.5782 | 140.3033 | 141.9479 | 131.3554 | 136.6967 | 97.9479 | 220.8957 | 229.9479 | 233.9195 | 292.9479 | 151.4455 | 133.5924 | 129.0521 | 146.5403 | 73.3696 | 81.2228 | 53.8957 | 124.4597 | 233.3317 | 212.3696 | 261.6304 | 193.3696 | 135.6588 | 134.2891 | 118.5924 | 127.5261 | 197.1185 | 160.8294 | 195.3934 | 185.3033 | 182.0521 | 173.6967 | 231.6162 | 190.436 | 180.9858 | 189.8957 | 199.199 | 306.8152 | 188.5261 | 169.4076 | 230.9858 | 229.1848 | 209.5165 | 196.2512 | 194.8294 | 403.436 | 159.4218 | 152.436 | 157.5782 | 265.9337 | 171.8294 | 164.6066 | 153.5261 | 169.1185 | 216.199 | 240.9195 | 194.635 | 181.8389 | 100.5403 | 115.2891 | 102.5261 | 128.7772 | 238.4076 | 230.8056 | 218.2891 | 429.1706 | 302.7868 | 235.8294 | 230.6729 | 360.1611 | 199.5924 | 76.4881 | 91.5782 | 199.5924 | 146.4076 | 126.5924 | 123.5924 | 139.0521 | 130.3792 | 136.8152 | 106.237 | 160.6967 | 70.4076 | 73.0663 | 74 | 93.3554 | 135.3033 | 134.8152 | 122.1043 | 78.6967 | 265.1848 | 221.1706 | 227.7914 | 462.7109 | 214.2891 | 224.6967 | 200.0142 | 396.3934 | 366.6446 | 244.5782 | 221.763 | 465.6967
1945 | 23.3894 | 146.6371 | 200.77 | 204.8762 | 222.6019 | 134.0176 | 177.4248 | 181.6195 | 165.6991 | 181.0886 | 136.3452 | 133.3629 | 202.0266 | 223.4867 | 161.9824 | 158.7433 | 268.3276 | 209.5752 | 165.4248 | 171.7966 | 214.4514 | 163.23 | 348.469 | 316.6195 | 165.8938 | 232.9114 | 229.3096 | 209.3362 | 403.9471 | 122.6638 | 115.8319 | 111.9114 | 138.1062 | 167.8143 | 131 | 140.6814 | 90.7433 | 232.4071 | 241.5133 | 239.6814 | 302.7257 | 168.4071 | 122.2743 | 138.9381 | 130.6991 | 68.3186 | 65.1238 | 67.1948 | 134.9381 | 240.5133 | 227.8762 | 236.2919 | 224.9647 | 149.1948 | 119.8052 | 111.7433 | 132.469 | 185.3805 | 151.9557 | 220.3452 | 176.8495 | 181.2124 | 140.1415 | 193.0972 | 191.469 | 179.4248 | 194.2567 | 234.4338 | 320.5133 | 166.6548 | 186.9824 | 197.2476 | 245.8762 | 216.4071 | 205.6195 | 197.5752 | 394.6371 | 158.8938 | 152.5752 | 152.7433 | 243.761 | 165.6371 | 157.531 | 141.4867 | 167.2124 | 241.6019 | 228.8052 | 236.2657 | 154.0353 | 115.9824 | 107.0619 | 101.2124 | 110.0176 | 242.5752 | 239.3009 | 229.7257 | 387.0972 | 304.4248 | 226.0619 | 266.3276 | 385.2833 | 185.5929 | 74 | 75.1238 | 185.5929 | 147 | 140.3009 | 113.1681 | 135.531 | 129.6371 | 124.5929 | 105.1062 | 126.2476 | 58.4867 | 78.2567 | 81.1505 | 104.6195 | 127.7433 | 113.6548 | 113.2743 | 90.7257 | 265.7876 | 260.4338 | 207.0176 | 467.3629 | 210.531 | 232.2567 | 209.6195 | 373.0795 | 329.46 | 247.7876 | 220.2124 | 472.3629
1946 | 23.4004 | 142.3756 | 204.8122 | 203.3756 | 207.878 | 149.2159 | 180.6244 | 187.5305 | 172.1549 | 164.878 | 129.939 | 134.9061 | 205.9061 | 229.2488 | 156.7512 | 171.5915 | 257.5963 | 186.4414 | 145.9719 | 150.7231 | 192.5353 | 167.3427 | 360.7793 | 338.8403 | 180.4976 | 246.4037 | 228.5305 | 253.2111 | 346.573 | 120.4695 | 118.8122 | 124.4976 | 167.9013 | 162.8451 | 143.6854 | 145.5305 | 81.8451 | 253.9342 | 231.2207 | 207.3804 | 294.939 | 150.9719 | 115.5634 | 115.5353 | 125.3756 | 88.8403 | 88.3708 | 76.2488 | 112.4414 | 259.2159 | 233.6244 | 260.183 | 180.9767 | 144.6573 | 134.3098 | 136.3708 | 124.8451 | 168.5963 | 156.4366 | 216.6573 | 200.4647 | 172.8451 | 175.8684 | 205.2159 | 163.0048 | 180.8122 | 204.061 | 224.5024 | 320.1878 | 169.4366 | 166.3475 | 233.7745 | 248.9061 | 185.3804 | 217.8732 | 217.9342 | 397.6244 | 150.8451 | 146.6573 | 166.4976 | 247.3427 | 165.9061 | 166.9671 | 144.5305 | 184.2159 | 229.5963 | 213.4085 | 246.5305 | 165.4037 | 108.0329 | 114.1549 | 108.2488 | 108 | 264.7464 | 237.3756 | 258.183 | 420.0562 | 327.5586 | 260.3379 | 241.0987 | 411.6525 | 183.0939 | 86.6854 | 92.9342 | 183.0939 | 117.0987 | 114.817 | 145.5257 | 126.8451 | 109.0658 | 105.7841 | 122.2159 | 118.3756 | 76.9342 | 77.1878 | 72.939 | 116.8732 | 127 | 132.7464 | 103.8451 | 82.0329 | 257.8451 | 231.4743 | 234.9013 | 446.2536 | 200.0329 | 211.2536 | 209.1878 | 386.3098 | 391.1453 | 243.4695 | 228.1549 | 492.0281
1947 | 23.4114 | 163.123 | 204.0816 | 202.0816 | 176.6115 | 163.8575 | 166.3057 | 181.5713 | 175.7552 | 158.408 | 128.0816 | 123.0609 | 193.1425 | 207.9586 | 177.123 | 132.5908 | 284.4701 | 183.0816 | 163.2862 | 146.1632 | 219.3885 | 152.3874 | 320.6724 | 320.7954 | 181.0816 | 240.6529 | 294.2058 | 238.7138 | 390.5115 | 130.1023 | 103.3874 | 124.1632 | 149.877 | 168.4287 | 143.1632 | 146 | 95.6943 | 256 | 248.3678 | 260.9402 | 281.1425 | 127.877 | 146.2253 | 126.7759 | 130.5103 | 84.5713 | 79.0609 | 74.2448 | 133.8782 | 205.8965 | 224.8161 | 219.8356 | 195.2862 | 137.5713 | 112.1218 | 150.0207 | 119.408 | 196.3885 | 179.0414 | 187.5299 | 187.3874 | 158.2241 | 168.9793 | 210.6736 | 183.8782 | 169.9793 | 182.0402 | 205.5506 | 284.1828 | 170.9184 | 163.0816 | 220.5506 | 218.6931 | 232.5115 | 212.5713 | 190.6115 | 387.8977 | 171.123 | 143.2448 | 161.5713 | 235.1425 | 154.0609 | 147.7954 | 158.7759 | 196.1023 | 225.2448 | 198.2241 | 199.2437 | 159.6529 | 102.408 | 107.6529 | 116.3471 | 120.8575 | 227.5092 | 267.3069 | 233.4483 | 371.6517 | 307.9586 | 289.7149 | 242.592 | 399.3057 | 183 | 79.7345 | 83.9793 | 183 | 128.6943 | 138.6333 | 99.4069 | 124.1632 | 114.3471 | 115.0207 | 111.1425 | 140.0414 | 81.7552 | 88.9391 | 72.9184 | 107.8977 | 118.7345 | 119.3874 | 118.6126 | 89.2655 | 245.0609 | 228 | 219.6322 | 467.8782 | 191.6529 | 215.4287 | 228.2862 | 395.3471 | 341.0598 | 282.4908 | 229 | 430.631
1948 | 23.4224 | 146.3864 | 192.8318 | 202.9307 | 182.3761 | 179.8909 | 182.683 | 156.8023 | 170.4159 | 154.2773 | 142.8909 | 115.4852 | 170.5943 | 254.3955 | 175.2773 | 160.6432 | 243.2579 | 176.4852 | 147.317 | 149.7227 | 218.2773 | 164.9602 | 371.9102 | 314.3466 | 181 | 245.5841 | 232.9909 | 208.1489 | 341.9511 | 135.6534 | 107.5841 | 123.0693 | 132.1784 | 144.8023 | 142.0693 | 127.3864 | 108.1682 | 275.5443 | 234.1784 | 242.733 | 301.4057 | 131.5841 | 158.3068 | 112.1784 | 139.3761 | 84 | 69.6239 | 73.0693 | 112.733 | 201 | 235.1682 | 237.4057 | 201.6534 | 139.792 | 131.4057 | 126.8023 | 142.267 | 205.5148 | 187.5148 | 234.3261 | 154.3568 | 163.5148 | 181.9602 | 197.0398 | 185.0693 | 184.8216 | 206.0591 | 222.6136 | 301.475 | 159.8318 | 151.8318 | 209.6932 | 220.6534 | 224.9011 | 192.4557 | 188 | 373.9705 | 164.6239 | 136.4852 | 161 | 260.0591 | 171.6136 | 157.1682 | 166.5148 | 170.0102 | 228.7227 | 212.8216 | 211.7523 | 144.1091 | 115.9602 | 84.6636 | 104.9011 | 123.8614 | 247.267 | 216.9511 | 221.6932 | 394.9205 | 326.475 | 295.7227 | 230.9011 | 404.5148 | 203.475 | 83.6534 | 73.6932 | 203.475 | 129.0693 | 120.525 | 116.4057 | 124.9307 | 109.4159 | 121.5841 | 123.0295 | 135.4852 | 81.0693 | 90.9307 | 61.8318 | 114.4455 | 125.4455 | 127.3068 | 120 | 94.6534 | 261.683 | 215.9011 | 204.9705 | 467.208 | 156.5648 | 180.6341 | 221.6239 | 383.9011 | 324.8318 | 245.05 | 248.5443 | 426.8614
1949 | 23.4334 | 155.3727 | 207.0876 | 203 | 232.0346 | 154.5968 | 171.7414 | 188.947 | 175.6578 | 175.6884 | 124.1976 | 131.0305 | 202.947 | 215.5662 | 164.6273 | 155.4562 | 233.3992 | 178.8289 | 168.6314 | 125.4827 | 209.5132 | 157.5132 | 340.167 | 327.2016 | 187.6008 | 269.5743 | 235.5438 | 247.4908 | 356.8595 | 118.0835 | 121.2016 | 106.9695 | 136.6578 | 133.5703 | 140.1141 | 127.8859 | 93.9124 | 243.9959 | 254.6884 | 233.4562 | 322.8024 | 135.7719 | 161.8289 | 138.3462 | 136.2281 | 86.8289 | 60.5132 | 75.8289 | 107.2281 | 207.6008 | 228.4562 | 274.833 | 187.8554 | 128.6843 | 133.943 | 129.7149 | 110.053 | 177.7108 | 186.1141 | 200.2811 | 169.9165 | 195.1181 | 171.6843 | 199.7719 | 173.6843 | 170.9124 | 184.4257 | 220.2281 | 297.3422 | 182.5743 | 159.4868 | 213.7149 | 199.3116 | 208.9124 | 200.4297 | 199.3157 | 385.2586 | 182.8595 | 132.2281 | 144.0265 | 248.7984 | 152.2546 | 154.2281 | 158.5132 | 170.8289 | 217.6843 | 209.2851 | 215.8289 | 147.7149 | 107.5703 | 122.6049 | 113.4297 | 123.057 | 236.7414 | 255.4338 | 243.6314 | 432.833 | 291.2241 | 277.1405 | 267.7189 | 361.6232 | 182.3686 | 79.2851 | 80.5438 | 182.3686 | 118.6273 | 135.9735 | 119.8859 | 125 | 117.4868 | 109.7414 | 117.3992 | 126.5132 | 92.3157 | 87.2281 | 67.6008 | 93.3116 | 150.5173 | 130.8289 | 117.1711 | 96.8859 | 274.3157 | 215 | 203.057 | 446.2546 | 172.8595 | 216.6619 | 205.9124 | 374.5132 | 336.2586 | 291.0346 | 261.3157 | 426.057
1950 | 23.4444 | 154.0895 | 194.6263 | 206.8211 | 220.671 | 169.2395 | 168.1342 | 185.2684 | 161.671 | 164.5816 | 141.15 | 129.1342 | 168.7 | 272.2264 | 171.6421 | 144.4921 | 236.8211 | 178.0447 | 142.2973 | 158.3895 | 230.9711 | 146.4921 | 378.1211 | 306.9842 | 160.2973 | 247.1184 | 213.0737 | 243.3132 | 313.1026 | 134.1948 | 105.7605 | 118.4184 | 122.671 | 141.5974 | 132.3579 | 135.6421 | 89.1789 | 276.3895 | 211.1026 | 255.9263 | 270.5052 | 136.9553 | 152.4474 | 130.4474 | 130.2684 | 77.4474 | 67.6421 | 83.6421 | 109.8658 | 225.1948 | 221.3132 | 256.9395 | 175.5368 | 143.2842 | 161.7027 | 113.7605 | 132.8369 | 177.9105 | 172.6263 | 235.2553 | 162.4026 | 144.4605 | 156.671 | 173.2526 | 179.6868 | 189.1053 | 219.3 | 236.2395 | 304.6421 | 164.8947 | 161.9105 | 209.2237 | 211.3737 | 216.5974 | 210.5526 | 211.4632 | 384.0895 | 162.9842 | 150.15 | 131.5368 | 244.1789 | 131.8947 | 164.5079 | 143.671 | 158.5816 | 231.329 | 221.4184 | 218.8658 | 132.7158 | 110.8211 | 113.5368 | 119.7316 | 125.8658 | 242.6868 | 195.9078 | 213.4763 | 369.0868 | 290.9105 | 283.6421 | 260.4474 | 410.5843 | 167.6263 | 66.5816 | 72.4026 | 167.6263 | 128.5079 | 109.2973 | 142.9263 | 132.6421 | 120.8658 | 113.7763 | 103.6263 | 116.4474 | 74.85 | 82.2237 | 75.6421 | 101.5526 | 116.6552 | 101.3868 | 127.5079 | 80.7605 | 276.9105 | 236.0158 | 222.1053 | 416.3421 | 174 | 213.2684 | 216.4632 | 381.6421 | 328.4026 | 259.6105 | 229.521 | 432.6868
1951 | 23.4554 | 141.4218 | 192.0649 | 208.9351 | 199.6813 | 170 | 199.9293 | 155.9733 | 181.3187 | 158.1947 | 128.4542 | 134.8053 | 172.8053 | 219.8494 | 159.4218 | 179.7995 | 235.0649 | 208.9618 | 148.7404 | 160.9676 | 217.4867 | 155.6756 | 319.044 | 338.8969 | 168.6756 | 246.9676 | 233.2862 | 217.8436 | 345.832 | 107.9085 | 117.5782 | 109.3244 | 141.3511 | 155.5458 | 147.4809 | 112.7787 | 116.0915 | 279.9351 | 241.8969 | 234.7462 | 314.4426 | 143.7729 | 143.292 | 147.416 | 131.9351 | 94.416 | 81.5458 | 59.8111 | 127.416 | 219.2271 | 229.708 | 240.5191 | 190.4809 | 120.7787 | 141.7138 | 129.4484 | 116.584 | 186.708 | 178.7729 | 194.4276 | 172.6431 | 166.1889 | 150.1947 | 185.5458 | 164.5191 | 180.3244 | 183.2653 | 204.1031 | 277.9085 | 146.584 | 144.584 | 209.9676 | 202.3244 | 213.1298 | 194.5516 | 208.1298 | 398.5133 | 157.1622 | 154.8702 | 130.0324 | 262.3835 | 151.3187 | 144.6813 | 136.2271 | 168.6431 | 212.6489 | 192.0058 | 195.7787 | 131.0324 | 106.1622 | 109.1298 | 104.5191 | 112.4542 | 231.3893 | 216.2213 | 257.4751 | 393.0915 | 270.6813 | 274.3244 | 266.7729 | 405.2596 | 178.6107 | 71.8053 | 77.8053 | 178.6107 | 121.2596 | 129.2862 | 124.6489 | 124.292 | 102.6165 | 104.3244 | 126.2213 | 116 | 81.7404 | 61.6813 | 68.2596 | 87.4867 | 111.1298 | 105.8053 | 106.7138 | 94.5133 | 248.9409 | 233.1298 | 223 | 440.1564 | 177.8702 | 202.3569 | 231.5133 | 399.416 | 329.9351 | 245.4218 | 222.1947 | 424.292
1952 | 23.4664 | 132.1814 | 208.6574 | 207.0403 | 221.5365 | 178.8186 | 198.0605 | 167.738 | 180.0403 | 179.5566 | 139.7582 | 118.3426 | 193.5768 | 230.738 | 156.0605 | 171.2015 | 261.4558 | 195.3023 | 141.1612 | 137.4837 | 197.4031 | 155.0202 | 320.9194 | 350.7783 | 188.5969 | 229.3628 | 246.738 | 240.5163 | 351.8992 | 123.6574 | 122.8992 | 117.8186 | 130.2418 | 148.1612 | 126.4434 | 153.1535 | 108.1814 | 261.3829 | 275.3349 | 239.8791 | 285.6248 | 154.7783 | 151.8186 | 134.2822 | 127.1008 | 70.5038 | 71.2217 | 81.5365 | 130.9395 | 197.4434 | 221.1814 | 257.6372 | 190.0202 | 138.6171 | 147.8589 | 123.1411 | 126.7783 | 173.2822 | 196.6372 | 186.1411 | 144.5845 | 141.524 | 179.3954 | 193.8388 | 143.4232 | 197.6372 | 206.4962 | 211.8186 | 257.4031 | 159.7178 | 182.214 | 203.1411 | 221.5969 | 196.3426 | 201.8388 | 201.1411 | 407.8186 | 145.2418 | 163.8186 | 147.6372 | 263.9798 | 162.7783 | 129.3023 | 134.0403 | 144.5038 | 206.1209 | 200.7985 | 222.4357 | 144.7178 | 111.8791 | 89.4031 | 84.4031 | 126.6977 | 226.1008 | 215.0403 | 219.8062 | 359.7054 | 281.7582 | 283.7985 | 293.4558 | 419.6977 | 170.1814 | 79.8388 | 78 | 170.1814 | 120.0202 | 125.1008 | 112.2418 | 133.7985 | 109.8388 | 99.1008 | 111.3225 | 131.6775 | 83.9597 | 81.5768 | 54.2822 | 90.9194 | 120.7985 | 105.0202 | 114.8186 | 99.8992 | 260.738 | 196.7457 | 203.4031 | 445.8992 | 164.2822 | 224.5365 | 219.262 | 402.9395 | 327.0605 | 246.9597 | 266.0931 | 410.2822
1953 | 23.4774 | 164.7405 | 197.0944 | 215.9292 | 215.055 | 183.9607 | 160.2988 | 142.2044 | 190.9135 | 160.1573 | 130.0786 | 142.8034 | 152.3302 | 255.8034 | 160.9607 | 147.1887 | 256.0472 | 182.1022 | 141 | 160.8113 | 229.7405 | 175.8349 | 352.7484 | 305.3617 | 188.0079 | 232.9685 | 216.2437 | 235.0472 | 370.8506 | 116.0629 | 116.055 | 111.055 | 125.0393 | 151.9685 | 135.9214 | 126.2202 | 112.9607 | 250.0865 | 217.4639 | 239.0079 | 297.8978 | 155 | 164.8978 | 128.0472 | 122.0393 | 93.8113 | 65.0472 | 80.0157 | 122.0708 | 226.7641 | 215.0472 | 225.2595 | 177.1022 | 151.8978 | 132.1258 | 119.0315 | 125.0157 | 185.8978 | 184.1022 | 181.0393 | 171.7798 | 159.8506 | 146.2673 | 196.9764 | 172.7641 | 207.9214 | 224.8585 | 204.0629 | 263.945 | 155.0393 | 144.3066 | 227.8034 | 224.9764 | 196 | 185.1337 | 211.9135 | 380.2202 | 162.8585 | 146.1415 | 142.0472 | 222.3302 | 160.0236 | 129.9921 | 136.9764 | 159.8742 | 184.173 | 198.0236 | 217.0472 | 162.8585 | 99.1022 | 124.7169 | 95.9056 | 119.0629 | 236.9135 | 239.8034 | 236.8585 | 416.544 | 276.0472 | 271.1022 | 282.0944 | 371.3853 | 167.0236 | 76.0315 | 81.9685 | 167.0236 | 124.9607 | 135.9135 | 126.8821 | 118.1258 | 126.8663 | 119.8349 | 116.9528 | 125.055 | 80.0315 | 85.9685 | 62.9292 | 89.0157 | 100.1651 | 108.9685 | 114.0079 | 105.9528 | 285.8034 | 198.9764 | 217.8821 | 436.0786 | 170.945 | 202.1808 | 205.1101 | 371.2516 | 336.9214 | 239.0629 | 219.3774 | 418.9292
1954 | 23.4884 | 156.0354 | 203.0089 | 207.938 | 200.9557 | 142.0221 | 170.938 | 164.0753 | 173.9734 | 167.9469 | 142.9026 | 130.0133 | 164.0487 | 246.8052 | 188.9734 | 159.0709 | 259.8627 | 209.9203 | 164.9159 | 138.1107 | 208.8804 | 159.9911 | 329.9026 | 317.1328 | 166 | 240.1904 | 240.9291 | 220.907 | 358.9911 | 142.9159 | 108.0487 | 111.9646 | 130.0487 | 159.9203 | 152.9557 | 154.9911 | 103.9601 | 271.8494 | 268.721 | 243.1063 | 307.9601 | 145.9026 | 150.9779 | 145.9203 | 148.8804 | 70.0487 | 72.0221 | 76.0354 | 129.9469 | 215.8849 | 210.0177 | 232.8849 | 184.9557 | 143.0531 | 120.1328 | 139.9867 | 124.938 | 184.9867 | 189.9779 | 179.0886 | 167.938 | 172.8317 | 171.969 | 163.0841 | 145.0531 | 184.0841 | 199.9159 | 216.9247 | 263.9336 | 174.9203 | 153.0266 | 192.0797 | 196.0753 | 205.0399 | 205.9557 | 192.9557 | 398.1461 | 150.0133 | 161.062 | 170.9513 | 196.1904 | 148.9424 | 147.0089 | 126.0177 | 138.0886 | 194.0709 | 193.0089 | 232 | 150.9336 | 98.0221 | 107.938 | 112.9557 | 114.9646 | 247.9424 | 215.0177 | 220.1727 | 434.7387 | 237.9513 | 317.031 | 263.9911 | 456.8406 | 168.9601 | 78.9159 | 73.9956 | 168.9601 | 107.155 | 107.0841 | 134.9779 | 108.0664 | 114.9646 | 111.093 | 101.9779 | 113.0487 | 78.969 | 76.0443 | 80 | 104.9734 | 93.0576 | 103.062 | 121.0487 | 80.0089 | 282.9601 | 183.2037 | 225.0531 | 415.8937 | 167.907 | 194.938 | 191.1151 | 368.1019 | 297.0133 | 221.093 | 236.1063 | 414.1107
1955 | 23.4994 | 164.2341 | 205.2842 | 194.5852 | 191.0669 | 147.535 | 157.1505 | 180.8328 | 168.0334 | 155.9164 | 121.8694 | 132.9498 | 174.9331 | 203.2675 | 182.6322 | 174.9498 | 230.1369 | 191.883 | 146.2174 | 162.8161 | 182.3009 | 158.1672 | 308.5852 | 346.6991 | 165.9331 | 282.0303 | 225.9864 | 200.4514 | 356.4148 | 123.8161 | 118.6155 | 104.4013 | 141.2341 | 142.1672 | 142.7826 | 152.4483 | 95.0836 | 237.7659 | 206.5852 | 266.2309 | 299.418 | 124.1839 | 146.1338 | 128.0502 | 122.2675 | 80.7659 | 76.9498 | 83.7659 | 118.4347 | 190.1338 | 214.1839 | 207.3344 | 174.9666 | 154.7158 | 149.465 | 137.1839 | 111.4347 | 181.8495 | 184.9164 | 198.7659 | 154.1505 | 135.1672 | 164.5486 | 181.8328 | 156.7325 | 203.3845 | 180.9331 | 200.3009 | 249.2006 | 157.1672 | 158.8328 | 209.7826 | 213.2842 | 213.5486 | 195.7659 | 183.2675 | 430.0136 | 153.0502 | 174.883 | 160.0334 | 239.1003 | 136.3845 | 148.7659 | 130.4681 | 158.0334 | 209.6322 | 194.8997 | 231.4148 | 135.9331 | 102.9833 | 94.4681 | 103.0502 | 107.4849 | 234.4148 | 219.1839 | 258.582 | 376.3344 | 227.1338 | 323.0972 | 262.2675 | 420.7325 | 160.2842 | 60.0167 | 73.1338 | 160.2842 | 141.4483 | 125.7158 | 130.1839 | 122.9498 | 107.0167 | 131.5987 | 97.2341 | 123.8161 | 72.1338 | 85.5319 | 80.0836 | 98.8328 | 106.0167 | 117.0836 | 131.3145 | 81.9331 | 273.1473 | 229.117 | 236.7158 | 392.535 | 147.3177 | 181.7189 | 216.4817 | 389.9968 | 300.0167 | 241.6322 | 258.8464 | 438.0136
1956 | 23.5104 | 177.8549 | 221.971 | 228.5068 | 195.2321 | 178.4488 | 166.1451 | 169.9556 | 169.7679 | 150.1587 | 172.0427 | 129.3037 | 170.7389 | 218.7679 | 161.2611 | 172.2901 | 295.2884 | 185.087 | 158.2747 | 152.9283 | 199.3618 | 167.3328 | 342.5068 | 328.8549 | 162.2321 | 225.1451 | 282.0563 | 226.4778 | 322.4061 | 113.6672 | 97.1024 | 127.971 | 154.3328 | 151.7389 | 130.029 | 120.9283 | 99.7099 | 225.1894 | 242.3055 | 221.029 | 322.9556 | 135.6382 | 153.6229 | 130.3618 | 137.5648 | 67.4932 | 74.2321 | 70.6382 | 142.8106 | 197.884 | 223.9266 | 226.3908 | 173.029 | 137.6519 | 118.1451 | 147.7969 | 137.058 | 173.1741 | 179.4488 | 185.6092 | 162.5648 | 144.2457 | 138.087 | 172.2031 | 140.6229 | 226.5512 | 177.1741 | 217.7389 | 260.6809 | 166.1877 | 149.058 | 197.5512 | 228.8976 | 187.7253 | 181.942 | 199.029 | 372.2611 | 155.6519 | 168.2031 | 161.7969 | 245.3481 | 158.3037 | 135.2321 | 157.3908 | 160.2321 | 187.6809 | 189.2611 | 196.913 | 132.6382 | 102.2611 | 121.6229 | 105.7099 | 135.7389 | 201.2184 | 229.5939 | 233.7969 | 396.116 | 235.2031 | 269.942 | 278.4932 | 404.7969 | 176.5358 | 61.4642 | 80.3328 | 176.5358 | 109.5802 | 109.116 | 140.7969 | 120.1741 | 107.4198 | 107.8259 | 111.2901 | 113.4061 | 80.1741 | 58.4642 | 84.5358 | 89.029 | 106.4778 | 121.2167 | 91.4642 | 77.5068 | 223.7833 | 235.2747 | 219.5358 | 424.1741 | 165.7389 | 223.2747 | 185.913 | 332.1604 | 300.913 | 221.1314 | 192.3925 | 381.1604
1957 | 23.5214 | 173.3717 | 221.9499 | 211.5044 | 203.0826 | 159.7522 | 171.9499 | 135.826 | 161.7935 | 122.3717 | 139.6696 | 107.0738 | 162.1239 | 211.0413 | 169.5044 | 181.2566 | 238.2065 | 187.7109 | 135.1151 | 182.9262 | 179.1977 | 145.5369 | 325.9587 | 323.174 | 169.9174 | 230.1239 | 216.6696 | 209.4543 | 334.6371 | 135.4218 | 133.0088 | 126.9174 | 133.0325 | 143.4956 | 131.0413 | 151.4631 | 90 | 264.4218 | 284.0588 | 223.8586 | 288.4956 | 156.4631 | 140.7109 | 109.8673 | 123.826 | 83.8761 | 81.7522 | 91.2979 | 104.1564 | 193.9174 | 188.9086 | 207.0738 | 174.1239 | 126.8673 | 123.7847 | 140.5044 | 138.2566 | 178.3392 | 160.2979 | 206.1652 | 149.5695 | 119.5369 | 140.7109 | 178.9587 | 128.3717 | 242.4806 | 183.4543 | 209.2065 | 249.174 | 139.5782 | 150.6696 | 216.0413 | 191.5044 | 211.0914 | 180.3717 | 200.7021 | 381.5369 | 145.1564 | 173.2653 | 154.7522 | 257.413 | 135.1239 | 142.2979 | 136.8761 | 167.2566 | 176.9174 | 196.4718 | 194.1652 | 152.6371 | 110.7522 | 108.587 | 96.7847 | 127.7434 | 240.9262 | 216.0826 | 227.7021 | 397.6458 | 241.2979 | 266.9262 | 293.3892 | 397.9174 | 161.826 | 77.1652 | 58.7847 | 161.826 | 128.6696 | 112.6696 | 133.9174 | 126.1239 | 88.9499 | 102.7847 | 120.5044 | 127.0413 | 86.5369 | 73.3392 | 69.5369 | 90.9086 | 89.5782 | 95.826 | 107 | 61.413 | 250.7434 | 210.3805 | 205.0325 | 429.7935 | 157.1239 | 198.7935 | 182.3392 | 371.0826 | 298.1239 | 257.3479 | 238.587 | 420.413
1958 | 23.5324 | 182.3216 | 241.5883 | 199.3569 | 204.8392 | 155.0183 | 192.2314 | 154.8392 | 157.5895 | 131.6431 | 132.3216 | 130.7674 | 164.3033 | 211.8928 | 158.6967 | 164.8039 | 241.7674 | 180.8928 | 159.2314 | 158.9647 | 206.5713 | 158.8575 | 326.2326 | 306.4653 | 168 | 233.3751 | 209.6431 | 220.4823 | 303.6967 | 122.1072 | 110.3216 | 124.9464 | 155.3922 | 155.6967 | 132.2144 | 138.6784 | 89.8928 | 250.7856 | 238.4641 | 266.4105 | 300.1608 | 142.7138 | 135.6078 | 129.8392 | 143.2144 | 80.7856 | 76.3751 | 74.732 | 130.6249 | 191.0889 | 209.4105 | 230.285 | 175.6066 | 145.821 | 141.5713 | 128.7856 | 121.6431 | 164.8221 | 144.8575 | 209.7856 | 184.553 | 133.0719 | 135.0719 | 177.4641 | 137.6431 | 185.5006 | 192.285 | 212.9817 | 230.8575 | 153.5895 | 143.5895 | 217 | 182.3045 | 190.6431 | 189.4287 | 216.8392 | 393.0889 | 171.6249 | 133.6967 | 149.3216 | 266.6249 | 138.1072 | 126.3216 | 133.3569 | 150.4823 | 175.7503 | 162.0183 | 198.1608 | 121.1608 | 105.6431 | 100.2862 | 113.4994 | 143.7674 | 216.3216 | 217.0889 | 242.1779 | 346.8587 | 226.7686 | 243.0183 | 257.9829 | 396.5359 | 179.6602 | 80.2497 | 75.9281 | 179.6602 | 120.0889 | 105.8039 | 131.1425 | 128.1425 | 111.1608 | 120.732 | 108.8928 | 126.4458 | 98.1425 | 58.6967 | 81.8392 | 110.9817 | 102.5177 | 113.6066 | 106.4641 | 71.1608 | 266.3922 | 198.3581 | 228.2497 | 425.6967 | 159.9464 | 192.7674 | 168.2862 | 371.9817 | 300.5177 | 219.9111 | 230.0183 | 428.2314
1959 | 23.5434 | 186.287 | 199.5271 | 190.1083 | 203.1859 | 170.8917 | 160.1435 | 151.2094 | 167.7365 | 143.7247 | 137.6706 | 109.8565 | 153.2518 | 210.0659 | 171.6588 | 179.1318 | 219.8023 | 179.9883 | 128.1318 | 172.3105 | 198.7365 | 172.4188 | 347.1435 | 345.1246 | 166.3529 | 238.3529 | 220.407 | 229.3953 | 316.3953 | 125.1859 | 115.6047 | 123.6706 | 128.7789 | 167.2094 | 135.9341 | 133.3953 | 88.3294 | 245.5506 | 227.6823 | 254.9459 | 301.287 | 120.593 | 161.4305 | 128.12 | 146.407 | 77.4612 | 82.6047 | 68.88 | 123.407 | 176.5153 | 199.7247 | 200.8565 | 152.5153 | 124.4729 | 133.4729 | 124.8023 | 132.407 | 194.6282 | 160.1977 | 203.6941 | 159.3953 | 152 | 152.1553 | 169.3177 | 147.5506 | 213.2518 | 162.593 | 195.3953 | 247.3836 | 165.4494 | 153.1435 | 217.0659 | 222.8682 | 201.4729 | 197.6588 | 212.0894 | 376.0117 | 164.4729 | 146.9883 | 154.7365 | 257.8917 | 139.0117 | 132.6588 | 122.9224 | 158.407 | 186.3647 | 178.2211 | 202.2518 | 124.7906 | 116.7365 | 121.6164 | 87.593 | 121.6047 | 222.1977 | 202.5153 | 210.6588 | 410.5199 | 256.0894 | 259.4847 | 291.7482 | 404.6164 | 155.8023 | 67.8565 | 57.9224 | 155.8023 | 106.833 | 119.7365 | 116 | 113.5271 | 114.4612 | 113.6282 | 108.6471 | 100.6471 | 82.8682 | 70.7365 | 77.88 | 92.0117 | 95.3177 | 89.7247 | 97.4612 | 75.2518 | 239.7789 | 237.3647 | 215.3953 | 436.8141 | 157.7482 | 173.1742 | 191.7247 | 354.4612 | 291.7365 | 235.0117 | 245.6282 | 396.407
1960 | 23.5544 | 163.4853 | 208.1726 | 220.1563 | 220.3909 | 141.469 | 147.8437 | 140 | 162.9837 | 154.6254 | 130.9675 | 122.6254 | 169.1238 | 211.7817 | 181.8599 | 178.4202 | 219.5798 | 193.7655 | 130.6254 | 122.9706 | 193.671 | 150.6254 | 334.2183 | 279.2833 | 146.7524 | 217.1889 | 214.1107 | 236.0163 | 324.7361 | 140.0456 | 109.3746 | 118.2183 | 154.3909 | 154.4365 | 136.0944 | 138.2964 | 91.9056 | 230.3159 | 209.5472 | 239.0619 | 278.329 | 128.1401 | 127.1107 | 143.2183 | 138.3127 | 82.4365 | 77.0782 | 54.798 | 116.2508 | 199.5635 | 211.0163 | 213.7036 | 171.7329 | 119.0325 | 126.3909 | 122.3909 | 124.8599 | 160.9218 | 162.3746 | 171.7036 | 165.5472 | 150.2802 | 126.7817 | 188 | 126.9218 | 229.9837 | 172.9544 | 202.798 | 267.8599 | 183.0294 | 141.8599 | 218.7036 | 221.1563 | 193.531 | 207.5472 | 184.7655 | 362.3127 | 158.2508 | 159.4365 | 150.531 | 228.1563 | 126.2508 | 141.2183 | 135.6873 | 148.8274 | 151.5016 | 154.0163 | 216.7167 | 134.8274 | 113.1563 | 101.2183 | 95.3746 | 116.1563 | 224.2964 | 222.3583 | 219.2964 | 348.798 | 229.3127 | 239.0163 | 275.6417 | 382.1857 | 153.0782 | 80.2345 | 71.6254 | 153.0782 | 144.8111 | 115.2183 | 115.2964 | 120.9218 | 119.4365 | 82.5016 | 129.7329 | 124.7817 | 81 | 66.9218 | 62.6254 | 79.1726 | 113.6873 | 99.6873 | 104.0782 | 92.2183 | 266.0163 | 200.3127 | 222.1726 | 417.8893 | 141.0163 | 204 | 201.8437 | 363.5798 | 290.6579 | 219.5929 | 211.7655 | 388.469
1961 | 23.5654 | 181.9951 | 219.6479 | 223.5379 | 220.4767 | 146.3667 | 148.8949 | 139.9095 | 153.0807 | 160.9144 | 107.4523 | 130.5428 | 147.4523 | 219.3716 | 190.9144 | 151.6186 | 247.8288 | 189.5525 | 138.7237 | 158.3619 | 177.5477 | 159.8998 | 324.1858 | 318.4523 | 187.2004 | 240.9193 | 237.9144 | 245.6479 | 354.6479 | 117.1809 | 102.0905 | 110.8998 | 158.0953 | 136.8142 | 147.9144 | 131.5379 | 79.3619 | 290.0146 | 220.3424 | 228.6333 | 293.8191 | 119.5379 | 151.2763 | 135.1761 | 141.0049 | 64.3619 | 77.5477 | 76.0905 | 127.7432 | 213.2958 | 218.5672 | 220.7335 | 143.8239 | 142.2763 | 130.1858 | 126.1858 | 134.3667 | 162.9854 | 156.7188 | 184.071 | 173.0856 | 130 | 135.4572 | 187.7286 | 129.1663 | 215.3765 | 191.2958 | 223.4572 | 277.6381 | 148.7237 | 151.3667 | 226.2763 | 221.2812 | 189.357 | 212.4621 | 182.8142 | 363.6479 | 172.9095 | 143.8044 | 144.8191 | 229.1858 | 139.5525 | 131.3667 | 133.6284 | 136.6284 | 180.8191 | 165.0049 | 178.6333 | 123.0807 | 113.5525 | 92.9047 | 90.2616 | 118.2714 | 214.5525 | 204.8998 | 213.2616 | 369.9095 | 231.0098 | 250.4572 | 291.7432 | 350.4475 | 152.9144 | 81.8239 | 79.3619 | 152.9144 | 119 | 107.8093 | 107.7237 | 120.1809 | 101.3619 | 111.4572 | 103.7237 | 132.3716 | 80.0953 | 66.0905 | 69.2763 | 92.9095 | 110.8142 | 95.8191 | 105.5428 | 82.0049 | 277.4572 | 204.7237 | 234.643 | 395.4426 | 155.0807 | 201.7384 | 198.8239 | 390.8337 | 321.0953 | 203.0905 | 209.7237 | 392.1907
1962 | 23.5764 | 191.0752 | 196.5138 | 238.589 | 177.2606 | 141.7469 | 176.3559 | 139.1028 | 174.7193 | 150.8221 | 112.2055 | 135.2807 | 151.7945 | 204.7469 | 177.3283 | 185.8421 | 222.609 | 173.4586 | 146.2055 | 159.5338 | 177.4185 | 177.4586 | 314.3559 | 324.0276 | 154.7544 | 219.6917 | 225.8697 | 222.2055 | 327.9173 | 120.1303 | 105.4662 | 127.5338 | 150.5414 | 142.6366 | 138.0276 | 147.2055 | 83.8221 | 260.2055 | 258.4511 | 235.7193 | 292.2055 | 135.411 | 144.7193 | 147.9248 | 131.6165 | 69.8496 | 72.2807 | 77.1028 | 95 | 167.9524 | 175.7469 | 206.5614 | 134.3634 | 133.0476 | 122.8221 | 118.9248 | 126.9724 | 191.7669 | 172.2531 | 223.7669 | 181.1228 | 131.2331 | 129.4862 | 181.8145 | 159.1504 | 189.3083 | 147.6992 | 217.1779 | 275.6441 | 152.609 | 144.589 | 217.8697 | 202.6642 | 204.6441 | 196.6917 | 188.4311 | 341.8496 | 171.0752 | 169.4311 | 144.5414 | 226.1378 | 123.9724 | 127.3634 | 151.2331 | 151.8697 | 178.3835 | 155.7193 | 183.9724 | 141.9449 | 99 | 100.4586 | 111.1504 | 120.8972 | 200.1028 | 225.1303 | 232.609 | 366.7394 | 211.6165 | 244.3835 | 257.2531 | 364.3835 | 141.1779 | 72.0551 | 80.0201 | 141.1779 | 118.3835 | 125.6917 | 115.7193 | 122.411 | 104.6917 | 106 | 110.589 | 116.5138 | 70.1779 | 68.7469 | 63.2331 | 92.6165 | 116.5338 | 96.0551 | 109.4586 | 71.4862 | 271.6917 | 215.2882 | 221 | 420.411 | 174.1504 | 179 | 184.7394 | 359.7193 | 311.1779 | 202.3559 | 217.6165 | 372.3559
1963 | 23.5874 | 180.0088 | 199.9646 | 232.3539 | 197.2301 | 155.0442 | 160.6194 | 142.0709 | 179.2743 | 155.4689 | 114.2301 | 128.0796 | 154.3718 | 222.5311 | 151.4689 | 165.7345 | 193.3451 | 160.6903 | 147.7699 | 139.1505 | 215.1947 | 164.8053 | 303.2213 | 329.6636 | 201.6636 | 218.4956 | 215.1947 | 222.9646 | 300.3097 | 128.7345 | 125.5044 | 107.6107 | 161.584 | 126.2568 | 145.0442 | 147.8495 | 92.7257 | 261.885 | 212.7611 | 237.3981 | 292.2743 | 135.7787 | 148.699 | 153.9291 | 136.7699 | 85.885 | 67.8408 | 78 | 96.3806 | 186.7257 | 189.7345 | 194 | 156.3806 | 116.6903 | 129.6549 | 125.5044 | 118.9204 | 182.3806 | 156.0796 | 212.4248 | 160.3718 | 141.1947 | 125.8053 | 154.9204 | 144.2655 | 189.8141 | 181.1505 | 208.3893 | 285.3981 | 126.1063 | 141.3451 | 207.885 | 191.6903 | 217.9646 | 193.0796 | 165.1947 | 356.1592 | 161.6194 | 147.115 | 158 | 264.6724 | 115.3451 | 147.5398 | 159.584 | 139.1238 | 175.8762 | 165.9116 | 176.2655 | 123.0796 | 102.2213 | 88.3806 | 95.115 | 119.6549 | 205.832 | 232.0088 | 206.2213 | 349.301 | 218.3806 | 237.9646 | 244.416 | 357.0442 | 135.2655 | 87.1238 | 56.301 | 135.2655 | 112.1947 | 120.1238 | 118.8937 | 126.9204 | 103.2655 | 106 | 105.6194 | 121.4602 | 62.7699 | 82.0442 | 74.6903 | 97.885 | 95.4602 | 110.7787 | 96 | 69.301 | 265.0884 | 241.9291 | 221.8053 | 423.5398 | 156.5044 | 181.416 | 168.3364 | 367.2655 | 303.5398 | 191.4514 | 223.115 | 359.4956
1964 | 23.5984 | 160.5655 | 193.5281 | 215.0561 | 195.9439 | 142.1648 | 149.4906 | 157.4906 | 167.7827 | 135.7453 | 115.4906 | 123.8015 | 186.0899 | 240.9813 | 135.5655 | 157.5281 | 198.8015 | 171.6029 | 145.3633 | 147.7453 | 205.0525 | 169.2173 | 326.854 | 304.6367 | 173.5806 | 227.9626 | 208.618 | 214.618 | 302.3857 | 116.1985 | 113.236 | 118.854 | 141.5993 | 155.6892 | 130.2547 | 137.0899 | 104.0899 | 256.9252 | 231.3259 | 204.5468 | 283.4569 | 113.6741 | 131.764 | 137.1086 | 134.1086 | 85.8914 | 79.7079 | 78.5094 | 106.3633 | 201.2734 | 182.6741 | 198.8389 | 167.6367 | 122.8914 | 124.8352 | 113.1086 | 125.1086 | 187.1424 | 152.4381 | 207.1086 | 189.4158 | 135.5094 | 131.7453 | 164.8015 | 152.0899 | 178.0936 | 191.5281 | 195.7453 | 253.0561 | 149.4906 | 142.4719 | 207.1273 | 197.1273 | 208.854 | 186.764 | 161.1648 | 346.0748 | 150.8727 | 147.618 | 153.9252 | 214.7266 | 118.5094 | 141.4532 | 138.0712 | 121.5842 | 196.0899 | 191.6705 | 185.618 | 117.146 | 122.7978 | 98.618 | 97.6554 | 118.4007 | 239.3072 | 210.7827 | 229.9813 | 363.6892 | 228.3633 | 230.2547 | 260.4532 | 341.8727 | 146.5281 | 67.4194 | 73.7453 | 146.5281 | 106.764 | 100.8015 | 97.9288 | 132.5993 | 112.7453 | 105.8727 | 96.9101 | 123.2173 | 61.1273 | 68.2734 | 79.618 | 97.6367 | 99.5094 | 89.0561 | 96.2547 | 88.0187 | 235.6367 | 225.3633 | 226.2173 | 422.4194 | 146.5281 | 197.8352 | 190.4345 | 379.5468 | 298.2173 | 213.2885 | 220.6892 | 374.9475
1965 | 23.6094 | 181.6489 | 202.1848 | 233.347 | 179.1889 | 156.8604 | 145.8604 | 155.2567 | 174.1355 | 134.8378 | 111.0226 | 142.0226 | 167.2752 | 231.2074 | 159.1622 | 166.4641 | 217.1622 | 203.7167 | 137.7885 | 145.7207 | 179.9548 | 157 | 321.5133 | 305.3696 | 157 | 218.4682 | 209.4907 | 214.5133 | 346.7167 | 98.9548 | 108.4189 | 111.4189 | 132.2793 | 130.9056 | 132.1396 | 124.1396 | 93.3737 | 236.6796 | 212.3019 | 221.0226 | 312.0452 | 132.6981 | 138.9548 | 131.5585 | 124.6489 | 90.6037 | 64.6981 | 78.23 | 103.9548 | 208.0452 | 201.1396 | 227.8111 | 171.7207 | 131.3737 | 111.8152 | 108.3963 | 118.0226 | 150.0719 | 173.9056 | 198.2074 | 160.3737 | 139.2793 | 131.9548 | 178.9733 | 140.9548 | 211.0452 | 200.0452 | 193.7207 | 270.23 | 145.0226 | 136.0493 | 207.1622 | 198.6981 | 200.3019 | 188.3696 | 180.6078 | 370.9281 | 150.5585 | 146.6756 | 129.4907 | 210.0493 | 125.3511 | 124.6981 | 120.0944 | 155.3737 | 179.9281 | 164.0944 | 182.0226 | 123.883 | 93.7207 | 96.9774 | 108.1622 | 127.0226 | 217.3511 | 219.23 | 219.9281 | 345.1889 | 224.8378 | 231.0226 | 248.8645 | 340.5811 | 155.883 | 80.7885 | 71.1622 | 155.883 | 113.8152 | 118.6037 | 116.6037 | 123.9774 | 111.6981 | 106.6756 | 111.3963 | 109.0452 | 65.4907 | 74.7659 | 77.8378 | 100.0452 | 102.8604 | 107.6263 | 99.117 | 91.6489 | 242.2341 | 219.6037 | 216.2341 | 433.8337 | 156.7207 | 182.0226 | 171.3059 | 393.5093 | 286.4189 | 184.2115 | 197.0226 | 402.4189
1966 | 23.6204 | 163.7346 | 195.1019 | 213.9616 | 200.4424 | 154.7846 | 144.6962 | 163 | 139.0634 | 140.9115 | 105.7596 | 137.975 | 138.5326 | 215.8231 | 153.6962 | 159.4307 | 214.7346 | 164.8481 | 124.2288 | 145.6711 | 191.3923 | 160.1903 | 334.1135 | 286.4942 | 156.6962 | 244.025 | 232.0634 | 226.6577 | 310.7346 | 112.5192 | 110.5442 | 111.1135 | 136.7346 | 126.2037 | 132.6962 | 127.2788 | 105.4173 | 280.5058 | 211.6461 | 215 | 298.025 | 137.7596 | 150.2404 | 134.2404 | 106.1269 | 80.4808 | 71.1269 | 58.0384 | 114.6327 | 198.2788 | 203.6711 | 203.0634 | 171.0634 | 142.8097 | 125.7346 | 118.2154 | 110.4808 | 169.9115 | 158.5693 | 183.7346 | 171.3539 | 139.6327 | 143.6962 | 148.9115 | 149.3539 | 198.0885 | 184.9616 | 193.975 | 249.1269 | 139.6077 | 159.1769 | 202.9115 | 200.5693 | 196.6077 | 168.8865 | 206.7212 | 359.4441 | 152.1769 | 154.4173 | 151.1519 | 236.2154 | 147.0634 | 131.1269 | 134.6711 | 164.3789 | 162.6711 | 177.7596 | 175.5442 | 116.6962 | 93.6711 | 101.9366 | 101.9366 | 120.2404 | 244.2288 | 201.773 | 203.5827 | 367.5058 | 225.8981 | 222.1135 | 279.5308 | 336.6327 | 149.9115 | 62.6711 | 65.5442 | 149.9115 | 123.025 | 108.6327 | 112.1019 | 126.9616 | 114.4808 | 117.3038 | 117.1135 | 98.975 | 82.1385 | 62.3673 | 80.5693 | 90.8865 | 99.7212 | 92.5442 | 106.1519 | 72.6711 | 254.9366 | 213.2788 | 230.7596 | 405.4942 | 158.0384 | 176.3038 | 200.6577 | 371.0885 | 287.6327 | 206.8865 | 192.975 | 408.1903
1967 | 23.6314 | 179.4926 | 216.6863 | 199.6274 | 184.0317 | 148.4926 | 140.2084 | 163.6568 | 149.2695 | 145.8358 | 112.1348 | 142.9242 | 169.581 | 228.4632 | 148.88 | 166.5957 | 225.5663 | 168.9264 | 157.0295 | 151.5515 | 186.3579 | 175.8652 | 320.1348 | 309.12 | 158.9411 | 232.8358 | 236.0295 | 213.5832 | 319.7599 | 120.1789 | 110.299 | 98.12 | 154.7916 | 169.5368 | 133.6274 | 141.1495 | 90.5074 | 256.7305 | 237.2548 | 213.3579 | 288.3137 | 140.1937 | 144.1937 | 127.0442 | 122.1053 | 71.1789 | 83.6568 | 74.8358 | 106.5074 | 207.2231 | 208.7305 | 209.8211 | 173.7158 | 124.0147 | 140.0147 | 123.8505 | 104.6274 | 171.7158 | 145 | 195.3873 | 152.9705 | 133.4779 | 139.3726 | 159.2548 | 130.8063 | 191.6863 | 170.9558 | 204.1789 | 255.0884 | 144.8063 | 149.1642 | 207.6716 | 189.1348 | 200.3284 | 189.2695 | 190.7158 | 395.9094 | 144.7916 | 142.7916 | 153.1495 | 239.7158 | 148.5663 | 144.1495 | 143.1789 | 135.4632 | 168.88 | 182 | 175.7916 | 116.1495 | 104.3137 | 98.7916 | 96.9853 | 116.3284 | 272.9242 | 227.7599 | 217.5074 | 346.3579 | 212.3601 | 205.8358 | 246.701 | 329 | 156.1495 | 74.1348 | 65.4632 | 156.1495 | 113.9705 | 103.9558 | 131.5515 | 113.7769 | 107.3137 | 118.8358 | 99.0295 | 109.5074 | 57.9558 | 70.8211 | 72.5832 | 108.6421 | 88.6421 | 90.9853 | 108.3137 | 80.3579 | 248.1789 | 223.5368 | 232.0442 | 423.5221 | 170.4021 | 178.1642 | 181.1789 | 365.8358 | 280.1642 | 218.7305 | 202.1937 | 419.7599
1968 | 23.6424 | 181.8235 | 207.176 | 208.5874 | 219.0568 | 150.647 | 133.0601 | 162.4109 | 167.6464 | 144.4705 | 123 | 113.9415 | 145.7066 | 231.9983 | 139.3541 | 142.2956 | 205.2945 | 191.3525 | 148.5885 | 138.9421 | 182.5891 | 167.118 | 329.7645 | 323.235 | 175.6464 | 232.1765 | 226.8825 | 233.5279 | 296.4721 | 112.6464 | 122 | 112.235 | 161.0574 | 155.588 | 144.529 | 146.8235 | 88.1765 | 233.4115 | 260.2934 | 211.8836 | 291.2934 | 135.4126 | 135.8825 | 116.2361 | 139.4699 | 69.6475 | 82.7639 | 71.529 | 105.5885 | 188.353 | 182.2344 | 212.0585 | 157.5295 | 120.9415 | 132.882 | 118.706 | 119.5885 | 155.5295 | 147.824 | 178.412 | 161.7645 | 138.3525 | 131.8246 | 175.5863 | 142.6475 | 185.8825 | 183.529 | 201.059 | 223.7645 | 153.647 | 150.353 | 200.1754 | 198.235 | 202.353 | 210.647 | 175.2355 | 367.176 | 155.1169 | 153.9994 | 158.4705 | 221.7645 | 127.2355 | 144.5284 | 139.5295 | 143.5874 | 157.4126 | 180.941 | 189.4705 | 124.824 | 108.882 | 113 | 103.2355 | 117.8235 | 239.176 | 197.059 | 215.5295 | 335.7055 | 252.9399 | 207.2945 | 232.882 | 327.7645 | 160.7645 | 81.8229 | 77.1175 | 160.7645 | 122.4115 | 118.1175 | 110.6464 | 131.7645 | 113.294 | 113.9404 | 92.0006 | 109.6475 | 72.4705 | 73.7645 | 95.3519 | 111.1754 | 98.412 | 99.5301 | 109.8814 | 74.2945 | 247.5301 | 211.706 | 213.9994 | 417.2355 | 148.059 | 181.471 | 181.5891 | 366.9415 | 282.059 | 198.5885 | 192.647 | 385.882
1969 | 23.6534 | 179.6784 | 190.2241 | 192.7199 | 172.3423 | 147.8672 | 160.9026 | 144.5871 | 153.6991 | 141.4336 | 122.4336 | 128.8526 | 162.2448 | 192.0767 | 166.1681 | 176.0914 | 219.5871 | 180.1328 | 155.2448 | 158.4129 | 199.8465 | 172.2801 | 318.3362 | 319.5311 | 163.0207 | 239.0414 | 230.056 | 197.9647 | 335.6224 | 103.2302 | 120.4896 | 104.1888 | 128.5311 | 146.1681 | 134.1328 | 144.6784 | 88.8112 | 230.3423 | 236.0146 | 238.9026 | 272.6784 | 151.0914 | 139.2448 | 134.1121 | 122.6224 | 78.9793 | 66.2095 | 64.3423 | 114.8879 | 191.8879 | 166.9854 | 205.2655 | 161.6991 | 128.1121 | 122.6224 | 118.6017 | 123.0353 | 158.1888 | 162.8879 | 185.5664 | 157.3216 | 132.0414 | 155.2241 | 133.9854 | 149.7138 | 189.2448 | 172.944 | 205.6224 | 225.9294 | 153.6991 | 151.056 | 177.9086 | 189.8112 | 204.5664 | 208.056 | 185.5311 | 358.9086 | 144.5518 | 142.0767 | 154.8672 | 214.6784 | 137.3423 | 122.7759 | 139.3569 | 129.419 | 181.0207 | 175.4336 | 181.1474 | 135.3569 | 101.2655 | 113.5664 | 109 | 116.4336 | 227.5104 | 209.1742 | 213.0914 | 330.0974 | 227.9294 | 215.5457 | 227.1535 | 334.838 | 156.8879 | 69.4543 | 73.7552 | 156.8879 | 113.8672 | 113.4336 | 97.8319 | 125.8112 | 107.3569 | 94.8112 | 106.7552 | 118.4129 | 72.8319 | 64.7905 | 68.6224 | 83.8112 | 102.9233 | 112.2241 | 89.2095 | 84.2448 | 266.0767 | 211.7905 | 201.1328 | 421.1121 | 159.0414 | 190.9233 | 203.7552 | 375.056 | 285.1121 | 207.5104 | 194.9647 | 384.4957
1970 | 23.6644 | 174.2011 | 184.6575 | 207.7718 | 195.2228 | 148.4293 | 143.0217 | 161.0054 | 163.8152 | 137.5924 | 118.3913 | 148.7718 | 158.3967 | 205.2228 | 156.2119 | 159.2228 | 234.3967 | 185.4022 | 152.2011 | 141.3913 | 183.2065 | 157.6304 | 299.8315 | 327.3371 | 171.9783 | 256.1523 | 228.6141 | 212.1848 | 331.3859 | 125.9783 | 116.6141 | 105.8043 | 146.3913 | 142.0434 | 142.2174 | 141.6141 | 86.3913 | 252.0163 | 208.413 | 212.9783 | 272.8315 | 130.2011 | 139.0163 | 125.3967 | 118.9891 | 69.4184 | 78.1902 | 81.5924 | 121.9946 | 197.587 | 195.9946 | 213.5924 | 168.3967 | 120.8043 | 120.5978 | 107.2174 | 114.2445 | 161.6141 | 164.7664 | 182.5707 | 162.3967 | 150.1577 | 139.6033 | 162.3913 | 131.0217 | 185.7989 | 176.3967 | 202.5924 | 261.4076 | 162.0054 | 143.9837 | 203.0217 | 195.6358 | 204.9891 | 202.1902 | 203.9946 | 382.8152 | 171.9729 | 152.0054 | 152.2119 | 210.6087 | 150.7718 | 137.5924 | 127.5978 | 155.1902 | 190.3805 | 172.3967 | 158.8152 | 124.2011 | 111 | 114.9946 | 110.8098 | 111.7881 | 232.9946 | 236.9837 | 198.0326 | 357.4293 | 257.174 | 206.6087 | 243.1902 | 384.5707 | 161.7826 | 79.7989 | 78.2065 | 161.7826 | 111.413 | 110.1957 | 109.5978 | 126.8098 | 98.6141 | 95.2065 | 108.7935 | 101.9946 | 82.1848 | 54.4184 | 65.7935 | 81.7935 | 95.0054 | 95.3967 | 105.0109 | 81.8043 | 261.7284 | 196.9946 | 213.2391 | 419.6358 | 178.9674 | 184.6141 | 204.9891 | 375.2228 | 273.7826 | 210.3805 | 214.2119 | 414.1631
1971 | 23.6754 | 166.2516 | 217.3981 | 193.828 | 206.6656 | 170.4268 | 159.2134 | 157.3185 | 169.2388 | 133.707 | 118.1879 | 125.6529 | 158.9873 | 225.8408 | 161.7994 | 175.3599 | 235.2006 | 183.586 | 150.4395 | 138.2006 | 187.3599 | 175.7325 | 321.7197 | 285.4809 | 156 | 227.1879 | 238.5732 | 204.414 | 329.9618 | 111.2803 | 126.7866 | 108.7866 | 140.4268 | 165.6783 | 147.7452 | 148.3726 | 80.8535 | 262.293 | 223.3344 | 203.6147 | 291.9459 | 126.7325 | 146.3057 | 123.6401 | 111.6401 | 82.3599 | 72.9204 | 74.293 | 119.2803 | 187.3599 | 196.6943 | 204.7994 | 164.9331 | 124.8535 | 117.9331 | 118.0796 | 143.3726 | 165.172 | 142.1338 | 164.6274 | 162.1338 | 121.3471 | 140.293 | 158.9873 | 144.2261 | 181.7994 | 171.6529 | 193.3726 | 259.9586 | 165.1465 | 132.2134 | 217.7197 | 219.0127 | 202.3344 | 195.2134 | 193.172 | 390.7994 | 152.2134 | 148.5318 | 156.3057 | 219.3471 | 126.586 | 132 | 122.3726 | 148.2134 | 174.5064 | 174.6943 | 173.414 | 124.3599 | 109.0796 | 109.5064 | 115.0127 | 108.3344 | 231.1338 | 223.9331 | 222.8535 | 375.586 | 236.7197 | 212.7866 | 233.6529 | 356.5987 | 148.1465 | 79.8535 | 81.293 | 148.1465 | 119.0796 | 109.7739 | 107.1465 | 131.8662 | 112.8408 | 100.8535 | 100.586 | 103.5478 | 70.7866 | 68.8535 | 63.5605 | 78.9204 | 99.4268 | 95.1338 | 113.2134 | 79.2388 | 219.2421 | 202.3886 | 236.4522 | 436.3981 | 151.586 | 197.5605 | 195.5064 | 384.7452 | 261.4268 | 205.3886 | 218.7325 | 383.1592
1972 | 23.6864 | 140.093 | 190.5754 | 217.1947 | 191.2877 | 162.7476 | 160.9027 | 132.1593 | 149.8053 | 143.8363 | 140.323 | 117.9027 | 164.3583 | 230.0706 | 150.4513 | 172.323 | 247.2257 | 169.1947 | 141.677 | 150.2257 | 183.4203 | 162.031 | 311.584 | 307.6417 | 157.1283 | 248.1947 | 233.615 | 214.292 | 353.7476 | 119.8363 | 125.5487 | 111.1593 | 145.6107 | 136.8053 | 122.354 | 135.4513 | 82.646 | 254.4203 | 233.9733 | 227.5487 | 281.2877 | 114.385 | 129.677 | 126 | 119.8673 | 79.7743 | 78.646 | 73.8673 | 114.522 | 182.5177 | 209.938 | 196.6107 | 160.5487 | 128 | 114.2257 | 112.1283 | 132.031 | 146.77 | 151.354 | 184.77 | 169.5487 | 128.8717 | 138.5133 | 166.615 | 139.1903 | 170.2257 | 166.385 | 179.5487 | 235.8053 | 158.3893 | 131.4203 | 212.5487 | 211.385 | 219.0663 | 196 | 172.5133 | 385.77 | 154.354 | 130.478 | 140.8053 | 227.7743 | 119.6417 | 133.354 | 111.4823 | 151.4823 | 169.677 | 180.2653 | 181.0353 | 120.1947 | 102.9027 | 101.7433 | 104.4513 | 123.7123 | 235.3893 | 226.093 | 222.615 | 356.4556 | 227.0353 | 213.8053 | 228.8363 | 331.1947 | 144.5487 | 82.0973 | 74.7743 | 144.5487 | 107.0353 | 117.969 | 108.062 | 120.8407 | 118.199 | 104.4513 | 90.031 | 118.584 | 70.9027 | 72.9027 | 76.8363 | 86.9027 | 98.7433 | 101.1947 | 112.1947 | 62.5133 | 253.1637 | 226.168 | 211.4867 | 403.4823 | 143.062 | 203.3893 | 182.1016 | 349.4246 | 266.6107 | 225.783 | 207.739 | 374.6417
1973 | 23.6974 | 163.6175 | 218.8102 | 214.8073 | 203.6204 | 141.9461 | 159.4788 | 148.949 | 159.0935 | 157 | 133.4788 | 120.5241 | 147.6175 | 200.5212 | 150.3343 | 167.1445 | 246.5722 | 166.0935 | 144.4759 | 145.7649 | 183.711 | 163.051 | 300.1388 | 324.8584 | 163.6175 | 236.289 | 225.3314 | 207.1416 | 325.0935 | 128.2408 | 124.238 | 122.238 | 154.1927 | 138.7167 | 131.2833 | 138.1898 | 79.6657 | 252.0935 | 213.7081 | 229.5694 | 295.5241 | 127.1898 | 131.5241 | 130.2833 | 136.4306 | 72.8131 | 74.4759 | 87.8131 | 90.5665 | 173.762 | 190.7649 | 212.3314 | 159 | 130.3796 | 116.9037 | 116 | 129.0057 | 170.9518 | 158.3796 | 204.6686 | 168.4759 | 123.5722 | 142.816 | 158.0935 | 153.1927 | 170.5241 | 169.1955 | 186.8045 | 246.5212 | 144.5722 | 127.6657 | 214.0935 | 222.5241 | 202.3372 | 195.2861 | 190.6175 | 394.4845 | 156.1445 | 161.5722 | 146.762 | 227.7139 | 138.2861 | 142.0453 | 117.6204 | 161.9037 | 171.0482 | 143.3314 | 166.1416 | 121.6147 | 104.8102 | 94.7139 | 104.8102 | 112.0935 | 218.2408 | 238.9094 | 220.5183 | 342.5665 | 216.9008 | 220.238 | 236.289 | 321.6686 | 146.8073 | 82.5694 | 74.9518 | 146.8073 | 98.3285 | 114.5694 | 119.3825 | 111.6657 | 96.3285 | 103.8584 | 94.1445 | 107.8526 | 73.5241 | 76.238 | 87.6204 | 94.0453 | 95.5212 | 95.238 | 102.6686 | 76.0963 | 242.3796 | 193.7592 | 196.238 | 410.0963 | 154.1445 | 186.4788 | 164.9943 | 326.9008 | 273.7649 | 175.3343 | 222.3343 | 397.3314
1974 | 23.7084 | 170.7488 | 205.2403 | 224.2473 | 191.996 | 173.7468 | 147.002 | 161.2433 | 166.2473 | 150.4935 | 119 | 121.002 | 160.004 | 221.006 | 138.7438 | 161.7537 | 242.7507 | 171.2453 | 138.4925 | 133.7547 | 192.7408 | 142.999 | 326.003 | 311.994 | 167.7458 | 218 | 231.996 | 217.5035 | 328.994 | 117.5045 | 121.997 | 118.4955 | 144.5025 | 131.006 | 143.7488 | 140.999 | 86.7517 | 258.4965 | 240.4925 | 241.2502 | 282.2383 | 134.2532 | 128.4985 | 135.2413 | 125.7507 | 57.004 | 73.2473 | 71.5035 | 113.7458 | 178.7557 | 177.2532 | 216.7438 | 156.4975 | 138 | 123.7507 | 119.7537 | 109.012 | 175.5015 | 165.7498 | 191.2473 | 174.004 | 125.2562 | 119.5125 | 170.2522 | 138.7478 | 170.5015 | 149.7587 | 215.2502 | 256.4955 | 145.7557 | 138.5055 | 215.2413 | 221.7507 | 188.007 | 196.7537 | 194.4955 | 354.006 | 148.7517 | 154.4975 | 147.001 | 227.7478 | 136.7507 | 150.7458 | 111.001 | 163.7458 | 165.998 | 161.5075 | 177.5045 | 135.7398 | 109.5085 | 97.5005 | 103.7527 | 115.4935 | 213.5105 | 209.997 | 242.2413 | 369.2493 | 228.7398 | 218.4975 | 227.7597 | 315.7547 | 157.999 | 86.7428 | 74.4965 | 157.999 | 126.001 | 123.7478 | 110.2493 | 117.2502 | 120.998 | 102.2552 | 87.7527 | 128.996 | 68.4965 | 74.7478 | 76.4965 | 104.7478 | 108.999 | 99.7537 | 96.004 | 77.5075 | 244.4945 | 214.5055 | 196.7498 | 401.4975 | 149.2542 | 177.7557 | 192.996 | 348.7498 | 266.5095 | 174.004 | 213.997 | 396.2383
1975 | 23.7194 | 174.6137 | 183.0886 | 217.8378 | 186.8261 | 157.699 | 147.2241 | 144.413 | 159.8378 | 136.2508 | 122.6756 | 127.5251 | 166.7492 | 231.9114 | 123.6756 | 168.2743 | 245 | 153.8495 | 120.2006 | 150.1003 | 174.4515 | 147.8762 | 337.8879 | 301.3511 | 159.7257 | 219.5752 | 211.5987 | 229.5752 | 318.0886 | 126.5368 | 110.8997 | 108.1505 | 143.3362 | 150.8378 | 137.8997 | 141.9381 | 90.6873 | 252.2006 | 224.0384 | 235.4365 | 257.2391 | 142.1622 | 126.1003 | 109.5251 | 125.6371 | 71.1003 | 63.9498 | 80.6873 | 102.3127 | 198.3629 | 190.6756 | 198.5251 | 151.3629 | 132.7492 | 121.5368 | 132.0502 | 135.286 | 171.0736 | 163.9498 | 187.7257 | 175.7609 | 142.9498 | 146.2358 | 171.7492 | 130.6873 | 172.8997 | 173.3746 | 212.3244 | 244.0502 | 157.2241 | 158.413 | 198.4515 | 211.3981 | 207.1622 | 203.7994 | 178.8997 | 367.7994 | 151.1121 | 143.8495 | 145.0117 | 219.1622 | 137.9498 | 134.587 | 118.4632 | 147.0619 | 159.2124 | 171.1355 | 188.1121 | 111.8261 | 127.3863 | 96.1121 | 107.2743 | 100.4632 | 241.587 | 202.3127 | 216.7876 | 361.7492 | 201.413 | 215.4632 | 248.8612 | 330 | 157.3629 | 67.6254 | 67.9381 | 157.3629 | 125.3244 | 112.0117 | 108.7876 | 120.8879 | 112.6371 | 114.3244 | 97.3127 | 112.5368 | 61.6756 | 75.8762 | 66.5251 | 101.413 | 99.9616 | 103.6489 | 103.5368 | 94.2241 | 233.5133 | 227.3244 | 205.9765 | 402.1388 | 159.8997 | 196.0502 | 185.2006 | 350.6254 | 288.4365 | 185.4749 | 200.7994 | 364.413
1976 | 23.7304 | 187.7551 | 207.947 | 217.2285 | 204.0762 | 139.4503 | 138.9701 | 150.702 | 163.351 | 142.7551 | 132.7252 | 128.1755 | 154.1987 | 205.9537 | 132.6258 | 152.8013 | 237.3047 | 149.3973 | 133.0993 | 146.3809 | 205.6721 | 163.6789 | 345.7252 | 322 | 168.3279 | 221.5265 | 195.4205 | 226.0299 | 325.6326 | 112.9007 | 108.0231 | 113.702 | 122.5728 | 157.9238 | 133.9238 | 146.9537 | 91.947 | 254.9306 | 227.2584 | 219.1987 | 284.9007 | 139.7483 | 135.0231 | 114.8476 | 119.5497 | 74.5265 | 60.1755 | 78.0993 | 108.4735 | 197.0299 | 205.947 | 208.5198 | 157.7252 | 124.596 | 112.0231 | 132.2517 | 104.4272 | 151.7715 | 160.4503 | 193.3047 | 161.0164 | 139.4503 | 121.4966 | 154.5265 | 133.3211 | 168.0993 | 164.351 | 208.0463 | 243.9769 | 147.3211 | 160.8544 | 217.3047 | 191.1157 | 205.5728 | 189.2517 | 176.8476 | 369.1919 | 152.8939 | 143.5198 | 137.3211 | 218.6721 | 136.649 | 124.7252 | 126.3279 | 140.6721 | 161.947 | 141.8708 | 172.8544 | 130.1755 | 103.8013 | 91.8476 | 98.3973 | 112.1755 | 231.1755 | 210.3973 | 222.8476 | 353.596 | 225.5661 | 237.8939 | 221.3279 | 332.7483 | 159.3279 | 75.5497 | 77.9007 | 159.3279 | 114.7252 | 102.3973 | 104.404 | 134.7715 | 106.2748 | 101.8013 | 101.8245 | 102.7483 | 74.1987 | 89.2054 | 67.7252 | 106.8775 | 88.2218 | 85.7483 | 90.7252 | 82.1225 | 246.5265 | 219.7483 | 222.4571 | 424.1755 | 156.1987 | 194.053 | 187.1061 | 353.8775 | 274.6721 | 182.9007 | 188.4503 | 373.1755
1977 | 23.7414 | 161.6932 | 220.1386 | 200.2773 | 196.9901 | 146.3266 | 155.9803 | 142 | 163.594 | 134.2076 | 134.8712 | 127.4356 | 165.4553 | 197.7523 | 127.8515 | 148.4356 | 216.7129 | 163.0099 | 131.6932 | 127.3167 | 209.0992 | 150.7326 | 333.5151 | 321.4258 | 163.1781 | 212.703 | 219.0197 | 202.1288 | 296.1583 | 113.4455 | 114.2773 | 110.4553 | 138.6038 | 158.1189 | 136.1288 | 129.8515 | 100.1189 | 224.1091 | 201.3364 | 224.4258 | 280.5644 | 141.2575 | 133.2378 | 122.703 | 120.7129 | 72.594 | 60.5841 | 78.4159 | 112.703 | 184.0394 | 211.2477 | 223.8219 | 159.5841 | 138.8417 | 122.8712 | 130.1682 | 104.475 | 158.6734 | 157.5644 | 181.9008 | 188.525 | 133.1189 | 136.8614 | 155.4652 | 143.3962 | 177.6038 | 167.7523 | 217.6833 | 239.7326 | 157.9704 | 137.406 | 197.8614 | 222.0992 | 200.0696 | 188.6038 | 185.8515 | 386.4855 | 164.931 | 168.0099 | 156.8712 | 221.8121 | 143.8712 | 125.7227 | 118.8811 | 152.4258 | 166.6734 | 160.5742 | 160.6038 | 124.8417 | 98 | 110.0394 | 107.703 | 120.6235 | 232.1583 | 229.4652 | 226.3962 | 367.8417 | 248.2083 | 258.8318 | 216.4652 | 327.0795 | 140.9704 | 70.3962 | 69.2575 | 140.9704 | 119.4553 | 113.4258 | 99.6334 | 145.6932 | 100.6833 | 95.7129 | 105.1485 | 111.1485 | 82.0099 | 65.4258 | 66.7129 | 96.5742 | 97.1189 | 86.6833 | 94.0197 | 84.3863 | 237.4159 | 219.2477 | 198.3167 | 430.3266 | 161.7129 | 178.9901 | 169.6137 | 345.297 | 291.594 | 177.1288 | 192.4553 | 374.4455
1978 | 23.7524 | 153.9941 | 208.7182 | 186.7182 | 186.3111 | 158.317 | 148.9941 | 151.2818 | 172.9041 | 160.1135 | 136.3111 | 130.1018 | 175.407 | 214.3053 | 127.6947 | 155.593 | 212.1076 | 175.4853 | 125.1918 | 140.6047 | 189.0959 | 157.6047 | 315.4795 | 313.4129 | 182.7965 | 211.7906 | 227.2035 | 209.6713 | 304 | 118.407 | 112.9941 | 118.3111 | 156.4971 | 145.1018 | 129.5988 | 129.9941 | 86.8023 | 216.8806 | 228.4971 | 226.593 | 279.3953 | 129.3953 | 112.8924 | 123.5871 | 121.7965 | 86.0959 | 70.2935 | 72.2994 | 109.0959 | 209.09 | 198.5871 | 203.0059 | 166.5988 | 134.2935 | 122.8141 | 139.407 | 124.7065 | 147.5812 | 150.7065 | 201.3053 | 160.8082 | 126.9883 | 137.2035 | 168.91 | 133.2877 | 189.2094 | 186.1018 | 207.09 | 252.2935 | 141.8023 | 134.683 | 201.4971 | 201.7965 | 169.5871 | 199.91 | 191.3836 | 359.1683 | 145.8689 | 169.7065 | 161.3053 | 206.4853 | 145.6106 | 135.9883 | 131.8982 | 150.7006 | 142.1076 | 162.5988 | 176.7006 | 120.2935 | 100.3953 | 121.9158 | 104.9941 | 139.5147 | 233.7123 | 247.7006 | 215.9883 | 359.1018 | 211.6947 | 239.4129 | 223.9217 | 306.0783 | 134.9824 | 59.9883 | 61.2877 | 134.9824 | 135.0959 | 111.4012 | 122.0176 | 131.1076 | 90.9883 | 90.8082 | 107.7006 | 116.9941 | 84.3053 | 65.1977 | 70.4912 | 97.1018 | 87.3953 | 77.8865 | 103.7006 | 71.683 | 233.6947 | 207.4853 | 215.4971 | 433.3346 | 157.1076 | 175.7965 | 182.6164 | 352.7965 | 305.6947 | 177.4853 | 196.1194 | 401.2642
1979 | 23.7634 | 170.3511 | 191.3639 | 179.6501 | 185.1565 | 140.1819 | 156 | 167.701 | 166.1298 | 145.2468 | 121.7532 | 122.701 | 162.0127 | 212.6756 | 137.4288 | 162.7532 | 216.4033 | 195.8053 | 138.7405 | 138.7405 | 188.3893 | 151.3766 | 337.5064 | 297.6883 | 185.4415 | 218.3244 | 225.8053 | 232.4669 | 297.4542 | 111.5585 | 119.3766 | 104.6883 | 151.584 | 142.0649 | 137.5458 | 138.2468 | 86.1819 | 250.598 | 233.2468 | 227.5191 | 277.5191 | 129.701 | 126.6756 | 132.1425 | 125.6883 | 83.2074 | 71.7659 | 76.7405 | 113.3766 | 213.8308 | 200.5967 | 197.2468 | 167.3766 | 142 | 114.1043 | 131 | 115.7532 | 168.1298 | 138.6361 | 189.701 | 157.8575 | 140.6883 | 131.7532 | 152.9478 | 138.9606 | 182.9873 | 182.4415 | 218.0649 | 250.3372 | 146.1692 | 150.3244 | 200.3244 | 205.3766 | 183.4415 | 189.2468 | 208.8181 | 382.3499 | 171.1552 | 152.3372 | 151.883 | 224.3117 | 137.9873 | 140.6489 | 133.6883 | 157.1692 | 136.117 | 167.117 | 174.7532 | 128.3117 | 105.0649 | 104.8702 | 114.4936 | 119.1947 | 218.0649 | 243.883 | 226.5712 | 370.0916 | 215.8181 | 229.9224 | 205.4809 | 346.6501 | 150.0776 | 75.8702 | 73.8181 | 150.0776 | 131.8957 | 109.3766 | 102.5585 | 125.117 | 103.7532 | 90.3511 | 98.2723 | 116.2074 | 76.1298 | 74.2341 | 73.8308 | 95 | 97.6756 | 91.5064 | 100.5064 | 87.9478 | 240 | 216.8957 | 217.4415 | 396.6756 | 146.4415 | 182.1819 | 167.0522 | 351.0776 | 305.4542 | 193.7532 | 181.416 | 420.3104
1980 | 23.7744 | 188.732 | 201 | 213.62 | 200.9841 | 145.972 | 155.028 | 156.324 | 158.1121 | 143.4641 | 114.4641 | 111 | 159.0999 | 217.1681 | 155.3679 | 156.436 | 234.3362 | 195.5201 | 148.944 | 145.056 | 184.7479 | 150.972 | 339.1279 | 310.9318 | 179.084 | 208.324 | 220.3362 | 209.592 | 294.3399 | 118.8879 | 112.8161 | 102.704 | 133.972 | 139.028 | 142.9281 | 133.8721 | 89.704 | 264.9281 | 227.5761 | 211 | 270.3959 | 121.8879 | 145.0999 | 122.1839 | 122.408 | 71.8319 | 60.916 | 78.5201 | 116.2119 | 191.5201 | 174.24 | 200.324 | 169.24 | 137.4641 | 132.972 | 127.436 | 120.1279 | 165.944 | 134.5079 | 182.8599 | 170.732 | 138.056 | 139.0439 | 149.1839 | 124.8319 | 179.0962 | 173.4921 | 205.9561 | 230.944 | 153.5761 | 140 | 186.1121 | 204.324 | 186.8039 | 190.38 | 195.9001 | 353.24 | 148.564 | 138.4519 | 146.944 | 231.1559 | 144.1401 | 120.648 | 130.732 | 161.3362 | 148.8319 | 171.408 | 168.1121 | 129.648 | 96.8441 | 102.5201 | 116.436 | 116.8879 | 217.944 | 230.1961 | 219.648 | 385.3922 | 212.944 | 228.7442 | 229.4519 | 368.3081 | 138.296 | 80.972 | 71.268 | 138.296 | 120.1559 | 108.324 | 107.944 | 126.8161 | 101.648 | 110.352 | 85.8039 | 105.4799 | 64.5481 | 82.8161 | 63.8319 | 94.028 | 105.084 | 90.8599 | 92.408 | 82.24 | 241.296 | 211.9598 | 214.972 | 403.436 | 147.5359 | 190.8879 | 172.1121 | 329.2521 | 289.056 | 194.564 | 195.7881 | 387.4397
1981 | 23.7854 | 182.9911 | 198.9823 | 208.9291 | 179.3451 | 147.6637 | 151.6549 | 150.2743 | 148.6549 | 133.6637 | 104.3274 | 111.6726 | 178.6903 | 206.0089 | 170.2743 | 149.3363 | 217.062 | 185.0266 | 147.2831 | 136.9646 | 201.6549 | 143.584 | 350.6371 | 331.9291 | 167.9557 | 211.6903 | 217.1861 | 223.0709 | 309.9291 | 122.6283 | 102.6726 | 109.0797 | 139.3629 | 131.2831 | 126.354 | 131.7789 | 88.3451 | 244.3186 | 216.0532 | 222.0975 | 286.6371 | 125.6283 | 156.9557 | 137.3717 | 122.3806 | 81.9823 | 64.9823 | 68.6992 | 120.9646 | 185.7346 | 171.584 | 196.6283 | 170.2831 | 128 | 129.2831 | 123.3629 | 131.6371 | 163.6106 | 155.416 | 196.6992 | 163.646 | 134.6726 | 148.9025 | 160 | 139.0177 | 162.124 | 159.3097 | 181.9734 | 236.3451 | 151.1329 | 144.708 | 189.77 | 197.6017 | 199.9911 | 185.9911 | 178.6637 | 363.6992 | 156 | 151.3097 | 149.9911 | 237.6106 | 133.6637 | 122.6726 | 122.9734 | 142.0443 | 163.0177 | 179.4515 | 156.6371 | 132.6814 | 90.3894 | 91.0177 | 106.9823 | 118.6106 | 222.3363 | 216 | 214.2743 | 352.6283 | 220.3629 | 215.1418 | 239.2831 | 337.2565 | 138.646 | 77.6194 | 76.6726 | 138.646 | 124.9291 | 112.3629 | 110.3186 | 118.354 | 99.3008 | 105.3008 | 96.3008 | 116.9823 | 54.3806 | 68.9734 | 71.6283 | 94.0177 | 99 | 104.6992 | 90.6992 | 90.6814 | 236.938 | 225.2122 | 214.646 | 407.4338 | 157 | 192.9468 | 168.3717 | 329.8321 | 288.3629 | 202.6726 | 189.6903 | 412.8671
1982 | 23.7964 | 172.6341 | 208.246 | 200.6687 | 187.9258 | 144.9085 | 145.8628 | 140.4858 | 144.6057 | 132.3028 | 113.1088 | 113 | 169.0112 | 213.9258 | 160.1372 | 147.9085 | 232.3943 | 181.5884 | 141.9258 | 130.7429 | 190.6341 | 135.4573 | 338.4228 | 315.3028 | 155.9085 | 217.6972 | 254.8171 | 229.937 | 302.7144 | 116.0915 | 104 | 124.9085 | 141.4685 | 131.1545 | 140.4116 | 149.3313 | 98.6687 | 231.2398 | 229.3943 | 236.6799 | 273.377 | 115.9543 | 138.6341 | 142.5142 | 129.8628 | 78.3486 | 64.1372 | 77.7429 | 113.3486 | 203.7886 | 163.8059 | 183.1199 | 168.4116 | 129.3943 | 120.44 | 125.1199 | 122.2114 | 155.1829 | 168.4228 | 198.7713 | 163.8801 | 133.9085 | 128.7429 | 148.4969 | 145.7886 | 183.0284 | 163.4116 | 182.623 | 234.1199 | 161.1311 | 158.8801 | 209.5142 | 186.1372 | 195.9085 | 186.7886 | 178.6972 | 373.7886 | 154.2571 | 154.3659 | 150.7886 | 219.7713 | 133.6972 | 131.3201 | 122.9258 | 154.0915 | 165.9543 | 206 | 154.1829 | 137.7429 | 108.9258 | 100.2287 | 103 | 116.8059 | 225.44 | 218.7886 | 205.5315 | 348.8801 | 231.8801 | 238.2856 | 219.9827 | 324.8801 | 128.0742 | 79.5488 | 75.9085 | 128.0742 | 116.3201 | 119 | 110.4858 | 124.3943 | 89.2114 | 95.2114 | 92.1372 | 115.7886 | 67.44 | 66.1372 | 64.7429 | 95.2114 | 108.4116 | 109.9085 | 90.6799 | 87.377 | 230.3201 | 203.2287 | 209.6514 | 423.7256 | 161.5315 | 176.8171 | 177 | 350.2683 | 286.9827 | 204 | 193.2571 | 389.2744
1983 | 23.8074 | 161.3304 | 210.6261 | 212.8348 | 191.4218 | 154.713 | 145.413 | 149.5261 | 149.2174 | 133.887 | 127.6696 | 108.3087 | 141.7217 | 220.6696 | 161.6696 | 146.5261 | 227.7826 | 177.713 | 147.587 | 138.3304 | 178.6087 | 152.113 | 333.0522 | 310.3913 | 157.7739 | 217.1957 | 232.5652 | 209.7522 | 319.3609 | 117.113 | 114.1043 | 111.6174 | 135.887 | 143.7 | 146.0913 | 132.113 | 104.3391 | 221.6609 | 243.1869 | 224.8043 | 270.0217 | 116.887 | 125.1652 | 128.7826 | 125.0826 | 77.1957 | 73.9696 | 82.4435 | 114 | 196.0087 | 186.887 | 187.3521 | 180.587 | 129.4739 | 132.2174 | 113.4739 | 123.1348 | 159.0304 | 152.587 | 193.3304 | 173 | 127.4739 | 124.4218 | 136.3826 | 149.5565 | 167.4739 | 177.7522 | 187.4218 | 232.5782 | 130.1348 | 164.3913 | 213.4652 | 195.6087 | 194.5261 | 204.6304 | 183.9696 | 378.2783 | 150.6391 | 167.4739 | 153.4739 | 224.7956 | 133.9174 | 146.8043 | 131.113 | 155.4739 | 171.5782 | 193.3696 | 162.3609 | 134.8652 | 116.0304 | 108.1957 | 106.9696 | 131.587 | 234.3304 | 240.9608 | 208.2261 | 373.1565 | 238.113 | 220.1957 | 202.4739 | 331.4739 | 127.1043 | 92.5348 | 70.1957 | 127.1043 | 115.9261 | 112.8652 | 110.1435 | 121.587 | 89.7739 | 89.6391 | 94.7522 | 116.6696 | 74.5261 | 65.1435 | 65.8348 | 87.8348 | 113.3696 | 105.6391 | 90.713 | 72.113 | 248.3304 | 216.9696 | 210.4435 | 405.5044 | 166.0304 | 172.9696 | 173.7522 | 342.4565 | 281.7435 | 203.2783 | 188.5565 | 392.3391
1984 | 23.8184 | 163.7758 | 191.1506 | 205.6416 | 183.5974 | 171.1636 | 146.7905 | 148.4026 | 164.9853 | 138.6268 | 121.1195 | 109.329 | 158.6727 | 205.5368 | 157.7316 | 143.5368 | 215.3732 | 186.7004 | 138.3732 | 144.8805 | 183.8805 | 139.5368 | 343.1195 | 314.4485 | 150.8346 | 222.9558 | 204.0147 | 199.1489 | 315.5221 | 124.3143 | 125.2831 | 112.1654 | 140.6268 | 129 | 129.1048 | 132.9706 | 86.3879 | 239.9853 | 249.5662 | 239.1948 | 287.8658 | 121.6268 | 129.7463 | 127.568 | 124.761 | 74.3732 | 74.6563 | 80.8952 | 115.4926 | 173.3732 | 184.91 | 218.8364 | 165.0294 | 128.3584 | 147.9853 | 110.8658 | 132.1342 | 150.5074 | 138.8952 | 199.8805 | 167.0294 | 128.9706 | 116.2242 | 148.8952 | 143.6416 | 176.8069 | 176.8511 | 185.1948 | 233.3126 | 147.3437 | 154.6416 | 221.5515 | 200.5074 | 206.8364 | 216.5515 | 196.5974 | 373.2684 | 156.7169 | 164.1195 | 146.761 | 233.4615 | 136.8511 | 142.1636 | 129.7316 | 149.5074 | 174.1784 | 168.0147 | 161.5074 | 129.9706 | 110.1195 | 110.5974 | 108.7758 | 122 | 240.8805 | 262.0442 | 202.1048 | 397.761 | 234.1195 | 220.7316 | 208.4485 | 330.3584 | 153.5827 | 75.1195 | 61.7758 | 153.5827 | 98.8364 | 100.8805 | 100.6121 | 110.5074 | 101.1195 | 95.3437 | 88.2537 | 113.1048 | 78.6268 | 58.2242 | 63.1195 | 86.6121 | 92.4926 | 102.0147 | 105.671 | 62.5221 | 258.9853 | 232.5827 | 208.5221 | 402.5827 | 150.7905 | 184.8511 | 174.3437 | 368.7004 | 297.8805 | 204.239 | 196.8216 | 379.6121
1985 | 23.8294 | 160.0109 | 211.1238 | 206.5527 | 184.1346 | 153.3746 | 141.4799 | 146.3236 | 164.9891 | 145.709 | 119.1455 | 121.1455 | 175.7492 | 196.8545 | 157.8328 | 136.7818 | 214.4582 | 161 | 137.4582 | 141.0727 | 172.3637 | 132.0109 | 335.7492 | 316.582 | 133.5635 | 224.8947 | 211.3344 | 202.5635 | 316.8654 | 134.9783 | 119.7818 | 130.2074 | 144.6254 | 134.0109 | 141.3963 | 141.8436 | 86.6254 | 238.8328 | 235.8654 | 244.1238 | 272.4984 | 121.7709 | 132.1564 | 138.1346 | 126.0109 | 81.5527 | 66.6982 | 77.4691 | 121.4691 | 181.3236 | 182.6363 | 237.0109 | 158.4691 | 138.6254 | 147.6037 | 120.1672 | 124.3746 | 154.1672 | 143.5635 | 189.5201 | 162.3963 | 131.291 | 122.6873 | 142.3855 | 136.0727 | 182.6547 | 174.5928 | 204.3855 | 223.1781 | 150.291 | 153.24 | 198.2182 | 198.3855 | 218.8436 | 194.76 | 194.4365 | 382.418 | 164.9164 | 157.5201 | 137.6037 | 206.3127 | 132.2801 | 127.0727 | 134.0727 | 142.76 | 163.3236 | 163.7709 | 161.3127 | 131.9055 | 115.4691 | 114.2182 | 104.24 | 120.4582 | 240.5418 | 246.2291 | 220.949 | 391.6873 | 235.6145 | 225.4582 | 218.291 | 328.6764 | 162.9891 | 79.3127 | 63.0217 | 162.9891 | 109.6873 | 104.3963 | 112.709 | 106.8436 | 106.4691 | 104.0727 | 92.7818 | 114.9891 | 79.5418 | 66.6145 | 70.0109 | 91.3855 | 98.0836 | 94.6873 | 94.6254 | 63.8654 | 266.6982 | 236.5928 | 202.9273 | 418.1564 | 144.709 | 183.3637 | 190.0109 | 348.7818 | 296.3855 | 194.5092 | 211.5309 | 364.7275
1986 | 23.8404 | 166.8068 | 197.3526 | 212.6248 | 169.2045 | 140.0339 | 157.7842 | 165.9548 | 161.3752 | 158.7955 | 115.7842 | 112.6135 | 158.2158 | 203 | 138.0564 | 142.8181 | 231.8871 | 160.2045 | 137.7842 | 134.0226 | 166.5458 | 145.1706 | 326.5684 | 293.591 | 152 | 205.9887 | 214.6926 | 226.1706 | 324.2384 | 111.0452 | 123.4316 | 116.5797 | 145.2384 | 142 | 144.4316 | 136.4203 | 95.9887 | 228.2045 | 246.0226 | 229.9548 | 256.921 | 131.7503 | 143.1481 | 119.2271 | 130.4203 | 85.2384 | 68.2158 | 75.8406 | 127 | 202.1481 | 192.8294 | 227.4994 | 159.6248 | 134.8632 | 134.6248 | 131.591 | 113.4203 | 170.7616 | 160.0113 | 175.2045 | 160.6587 | 125.3639 | 132.5232 | 146.5797 | 142.1481 | 168.1255 | 173.9097 | 200.2271 | 249.7503 | 149.1368 | 155.2271 | 190.1932 | 198.2045 | 204.2723 | 192.3752 | 179.9774 | 387.1129 | 155.6248 | 142.409 | 134.1706 | 214.7729 | 121.5458 | 130.7616 | 137.7616 | 137.1932 | 176.9887 | 157.0452 | 163.0113 | 131.33 | 118.6135 | 110.5684 | 102.2497 | 119.9887 | 248.1706 | 229.4881 | 222.4994 | 377.261 | 224.6587 | 216.2384 | 192.4768 | 305.8632 | 146.6474 | 75.4429 | 74.6248 | 146.6474 | 111.9661 | 108.2271 | 117.0452 | 110.9661 | 114.7842 | 110.9436 | 101.2045 | 104.6135 | 78.4203 | 71.9661 | 70.8406 | 93.591 | 101.409 | 98.5571 | 96.0339 | 76.8068 | 253.4994 | 219.2045 | 214.7503 | 426.7616 | 151.8294 | 184.3074 | 191.6361 | 364.3639 | 293.0226 | 187.3187 | 204.0113 | 348.3187
1987 | 23.8514 | 167.4602 | 197.699 | 200.6696 | 176.6107 | 145.3512 | 152.5692 | 167.3391 | 178.6609 | 157.9498 | 115.8997 | 106.5398 | 151.1298 | 198.4896 | 127.1003 | 149.0709 | 247.2388 | 159.41 | 137.4896 | 131.2803 | 176.8997 | 145.6194 | 340.7405 | 308.711 | 147.0796 | 212.6903 | 200.6609 | 229.4896 | 298.4187 | 109.2508 | 119.9204 | 117.0796 | 134.1799 | 140.7699 | 126.5692 | 130.1799 | 101.8702 | 225.3599 | 224.0087 | 235.0294 | 276.7699 | 148.2803 | 144.699 | 117.3304 | 129.5104 | 76.6401 | 66.8702 | 64.59 | 116.3391 | 201.6488 | 176.3893 | 201 | 152.59 | 122.1003 | 130.4602 | 126.6194 | 107.59 | 171.5692 | 168.0709 | 185.4809 | 153.6107 | 130.0796 | 146.8495 | 149.5398 | 150.2595 | 181.8495 | 189.7993 | 190.5398 | 254.7993 | 147.1385 | 148.8201 | 197.3304 | 208.891 | 195.6609 | 193.2595 | 185.1799 | 356.59 | 146.1298 | 143.6903 | 144.0502 | 231.8408 | 134.7699 | 139.3599 | 145.5398 | 143.9204 | 162.3686 | 144.1799 | 161.2301 | 142.6194 | 107.6194 | 113.2595 | 90.6401 | 120.5398 | 253.1298 | 215.4308 | 199.6903 | 376.1712 | 217.2007 | 213.7906 | 197.4221 | 311.1418 | 143.8408 | 61.5398 | 64.7197 | 143.8408 | 121 | 99.7699 | 106.2301 | 121.2301 | 112.0294 | 117.3893 | 93.0294 | 104.6903 | 74.6401 | 78.1298 | 66.9706 | 97.6401 | 89.9793 | 98.5191 | 84.1298 | 73.7699 | 231.5104 | 218.41 | 219.3893 | 410.7579 | 148.5104 | 195.0589 | 171.7492 | 369.0796 | 298.891 | 195.0589 | 203.0502 | 376.9706
1988 | 23.8624 | 171.8446 | 180 | 196.8009 | 186.8884 | 162.6214 | 141.5339 | 160.0242 | 174.1508 | 155 | 124.7813 | 103.8446 | 148.267 | 185.6651 | 143.1358 | 157.3982 | 232.3349 | 168.8688 | 136.0679 | 125.0196 | 172.267 | 148.092 | 335.908 | 335.4465 | 147.1795 | 217.5777 | 210.5098 | 220.8884 | 282.7572 | 118.9321 | 118.9758 | 110.4223 | 122.8642 | 149.1358 | 127.359 | 126.0437 | 95.8642 | 228.0679 | 205.2232 | 235.5581 | 269.9321 | 136.9517 | 131.2232 | 123.3107 | 128.3982 | 78.5777 | 64.6651 | 69.0679 | 97.6214 | 186.0679 | 169.9125 | 204.3786 | 155.8009 | 123.3544 | 138.2232 | 117.6893 | 106.5777 | 168.1358 | 171.3303 | 196.0875 | 162.6214 | 128.4902 | 139.267 | 144.733 | 138.6214 | 173.8446 | 161.5294 | 198.8251 | 239.1991 | 133.092 | 149.5777 | 215.9804 | 212.908 | 199.5972 | 184.1554 | 178.5098 | 362.7572 | 139.4661 | 152.8009 | 149.5339 | 229.9517 | 122.8642 | 139.5339 | 134.8205 | 151.8446 | 142.0679 | 153.5581 | 170.1795 | 142.5581 | 104.1795 | 104.1554 | 98.0679 | 116.1554 | 235.0633 | 225.1991 | 209.6456 | 384.4856 | 231.1116 | 209.374 | 224.7089 | 329.0633 | 155.8884 | 63.4902 | 65.9125 | 155.8884 | 113.8205 | 99.267 | 113.0679 | 119.1991 | 104.1116 | 109.6456 | 91.0242 | 112.5339 | 78.267 | 77.3786 | 71.9321 | 94.9321 | 85.4465 | 99.3153 | 88.0242 | 76.2232 | 250.6697 | 231.2474 | 204.0437 | 390.7376 | 141.4856 | 164.8642 | 170.5144 | 362.4223 | 313.8884 | 177.5339 | 207.6893 | 383.6214
1989 | 23.8734 | 166.0462 | 190.4308 | 184.1808 | 184.4346 | 150.1769 | 141.9577 | 168.8269 | 156.9961 | 151.5231 | 125.4423 | 111.5192 | 152.6077 | 188.7346 | 152.6538 | 145.2615 | 227.5269 | 176.2192 | 149.9538 | 127.3846 | 175.7385 | 162.5231 | 336.2538 | 330.4846 | 157.8692 | 214.3923 | 203.8692 | 217.5654 | 301.9961 | 111.5654 | 107.1308 | 108.3923 | 110.7385 | 153.0039 | 140.4808 | 124.1731 | 87.65 | 257.6 | 212.7385 | 243.4808 | 279.0808 | 114.5654 | 136.5654 | 123.6077 | 105.8308 | 82.7808 | 56.4346 | 68.6115 | 96.4769 | 183.4385 | 196.2538 | 210.3038 | 164.0423 | 119.6077 | 131.3962 | 124.7808 | 110.7808 | 169.3962 | 154 | 188.8692 | 169.7346 | 128.6115 | 144.4769 | 145.1731 | 130.5231 | 173.6962 | 157.6885 | 209.7846 | 249.6461 | 146.2192 | 152.9115 | 215.3539 | 188.9154 | 188.3462 | 185.6077 | 169.2615 | 365.0462 | 143.3885 | 155.8269 | 152.5654 | 225.8192 | 121.6038 | 135.1769 | 131.0846 | 147.35 | 161.1692 | 157.1346 | 171.7423 | 146.5692 | 109.6538 | 104.3038 | 106.7385 | 109.35 | 223.3885 | 229.9961 | 221.65 | 364.2615 | 234 | 204.95 | 218.9538 | 319.1269 | 143.4385 | 70.5654 | 69.2192 | 143.4385 | 116.6038 | 107.0846 | 113.0462 | 114 | 98.7423 | 112.0885 | 102.4346 | 111.6538 | 70.4385 | 79.8269 | 75.4308 | 85.8269 | 99.6077 | 126.9115 | 88.5692 | 75.0462 | 265.3923 | 225.8346 | 202.6077 | 413.2615 | 134.7346 | 165.3423 | 202.4731 | 354.3077 | 288.4 | 186.2154 | 200.0039 | 386.3885
1990 | 23.8844 | 154.3186 | 197.7433 | 170.8319 | 187.2345 | 141.3407 | 147.385 | 158.4026 | 155.2521 | 148.3407 | 106.2567 | 127.1505 | 169.4072 | 191.4866 | 145.2124 | 143.4469 | 210 | 167.7655 | 154.084 | 149.5531 | 178 | 155.3186 | 347.4866 | 322.8541 | 158.5531 | 222.1727 | 208.1283 | 216.1062 | 296.6768 | 112.7876 | 108.6814 | 108.1283 | 121.0443 | 154.7479 | 127.9778 | 127 | 89.1902 | 253.9289 | 224.8319 | 227.8495 | 295.9778 | 115.3407 | 130.1902 | 123.4247 | 111.4515 | 79.1681 | 66.8319 | 61.6814 | 100.5531 | 176.3629 | 206.5928 | 198.1459 | 166.2124 | 132.385 | 119.4691 | 119.3804 | 115.2124 | 161.0443 | 155.7876 | 194.4691 | 177.4026 | 125.2567 | 142.7433 | 157.8319 | 136.2788 | 165.2964 | 167.2521 | 200.3186 | 249.8892 | 146.7036 | 145.9335 | 196.3407 | 185.916 | 197.5753 | 193.4691 | 176.8319 | 376.5577 | 153 | 148.5309 | 147.9778 | 243.6371 | 138.8938 | 131.2567 | 133.6371 | 136.8717 | 176.1062 | 151.0443 | 171.2788 | 137.916 | 103.5531 | 109.1283 | 112.1283 | 115.8541 | 237.0222 | 233.615 | 220.5088 | 362 | 238.4691 | 229.6814 | 221.2964 | 343.0665 | 134.1283 | 68.6593 | 73.2788 | 134.1283 | 121.3804 | 121.2567 | 94.1681 | 111.3186 | 100.5134 | 98.5088 | 103.4469 | 108.2345 | 62.0222 | 73.8717 | 80.5088 | 85.6814 | 102.1062 | 117.699 | 89.301 | 73.1505 | 252.615 | 209.9603 | 213.1505 | 399.3804 | 148.2124 | 185.3407 | 188.7166 | 352.385 | 265.2567 | 190.3186 | 191.916 | 396
1991 | 23.8954 | 156.5104 | 199.6431 | 194.0207 | 185.408 | 136.5712 | 158.0816 | 150.3776 | 148.0207 | 143.112 | 119.0512 | 123.1425 | 180.4896 | 189.7759 | 137.4896 | 135.2753 | 224.2351 | 161.7856 | 148.4592 | 143.9488 | 183.5104 | 158.888 | 321.8977 | 329.4385 | 160.296 | 235.5408 | 206.9488 | 220.9696 | 280.4288 | 110.408 | 116.1328 | 112.4592 | 135.592 | 159.6833 | 126.6736 | 127 | 86.592 | 226.3167 | 227.3569 | 226.8064 | 286.8672 | 123.888 | 116.112 | 116.704 | 121.9281 | 66.5408 | 72.5712 | 65 | 100.4592 | 201.8271 | 176.6833 | 189.7247 | 158.9488 | 144 | 126.8368 | 112.8064 | 126.7759 | 167.3264 | 151.5712 | 193.112 | 161.5712 | 131.6224 | 134.0816 | 164.9488 | 140.2753 | 158.8368 | 155.888 | 208.4799 | 225 | 154.0816 | 152.1231 | 201.2144 | 208.0719 | 206.592 | 194.408 | 184.8672 | 381.7953 | 137.3873 | 151.2656 | 136.112 | 240.6736 | 137.704 | 134.408 | 132.0512 | 135.7552 | 163.0609 | 168.8064 | 181.2448 | 134.3569 | 108.9696 | 108.408 | 113.704 | 127.4896 | 238.7856 | 234.3983 | 208.1632 | 371.1839 | 236.6528 | 217.3873 | 209.7856 | 326.3154 | 143.5104 | 75.2656 | 80.0304 | 143.5104 | 116.6431 | 123.4896 | 97.1535 | 117.6431 | 108.1936 | 88.4592 | 98.4896 | 113.7552 | 63.3264 | 71.8368 | 62.6528 | 97.7247 | 105.1328 | 100 | 107.9184 | 81.5408 | 235.9488 | 221.408 | 223.3776 | 392.8064 | 163.4495 | 187 | 166.9903 | 367.2144 | 287.6943 | 187 | 200.296 | 384.0609
1992 | 23.9064 | 159.6996 | 205.3421 | 211.7149 | 187.5438 | 137.0153 | 164.5438 | 151.5285 | 160.5285 | 135.943 | 139.6161 | 113.1864 | 163.6273 | 204.5855 | 132.4145 | 135.8442 | 226.8553 | 148.0987 | 146.1711 | 137.5285 | 188.5855 | 158.9847 | 318.0153 | 316.057 | 166.7719 | 226.9847 | 208.0723 | 210.5549 | 311.5744 | 115.8442 | 113.4562 | 107.3421 | 140.5285 | 135.5132 | 134.7719 | 124.6426 | 89.6426 | 231.1558 | 230.1447 | 230.3421 | 278.6996 | 134.8289 | 112.7149 | 111.6426 | 108.6579 | 64.5855 | 69.2434 | 68.3004 | 101.4715 | 198.5396 | 170.4298 | 203.7719 | 152.057 | 131.7413 | 125.2281 | 120.5855 | 136.3992 | 151.6843 | 150.6008 | 188.7719 | 153.5285 | 134.2434 | 123.0987 | 155.6996 | 132.8289 | 168.0723 | 167.3004 | 204.0834 | 234.9013 | 149.6996 | 170.8707 | 224.8025 | 214.2698 | 214.3574 | 197.4868 | 179.5285 | 370.7302 | 130.3157 | 152.0417 | 134.6008 | 241.6996 | 133.114 | 127.5855 | 133.2851 | 140.4145 | 163.6161 | 171.3992 | 169.0417 | 122.057 | 106.5702 | 102.057 | 111 | 107.3268 | 236.4145 | 247.5855 | 222.0306 | 375.8707 | 220.4562 | 207.4868 | 220.6161 | 300.3157 | 145.7566 | 83.5855 | 79.1293 | 145.7566 | 124.2281 | 128.3157 | 110.2851 | 132.7719 | 94.7566 | 90.886 | 96.2434 | 120.3004 | 57.5855 | 78.7149 | 63.9013 | 100.4562 | 115.1864 | 99.5285 | 95.7983 | 80.057 | 229.057 | 227.7872 | 220.2851 | 401.057 | 171.7413 | 189.3574 | 184.6426 | 338.9388 | 311.6008 | 196.9013 | 203.4715 | 369.057
1993 | 23.9174 | 162.7729 | 209.1594 | 209.7433 | 190.6784 | 131.0028 | 164.292 | 161.1594 | 170.1594 | 144.2567 | 141.4866 | 118.5162 | 160.708 | 199.1298 | 135.9351 | 152.354 | 210.0324 | 148.6107 | 135.7433 | 135.5487 | 191.8378 | 158.2243 | 310.5515 | 319.354 | 176.8054 | 213.1622 | 212.1298 | 207.4485 | 329.7757 | 134.7729 | 112.7405 | 103.9027 | 126.938 | 123.5811 | 136.5811 | 132.6431 | 90.3865 | 225.7729 | 237.7757 | 231.2567 | 250.811 | 135.0973 | 106.8731 | 118.6756 | 115 | 67.8378 | 65.292 | 81.1918 | 108.7729 | 171.9351 | 184.3865 | 205.5811 | 139.8731 | 128.6431 | 120.0324 | 118.0324 | 120.292 | 160.2863 | 152.1946 | 196.3865 | 151.0649 | 143.354 | 113.9351 | 155.8702 | 137.9351 | 168.2595 | 182.1269 | 185.4513 | 255.6756 | 146 | 156.2595 | 231.0353 | 199.0324 | 192.811 | 198.292 | 184.3216 | 386.4513 | 153.6431 | 153.6107 | 143.9351 | 226.8731 | 127.1298 | 126 | 133.3216 | 157.9647 | 180 | 159.1622 | 165.7729 | 131.1594 | 105.1918 | 103.9027 | 111.4838 | 102.1594 | 237.5162 | 229.5515 | 234.6784 | 370.4513 | 227.4809 | 214.0973 | 225.3893 | 303.8082 | 138.0973 | 77.1622 | 77.7757 | 138.0973 | 115.646 | 122.1326 | 117.0945 | 117.6488 | 102.0945 | 87.1946 | 104.8702 | 116.7433 | 59.3865 | 75.7757 | 74.0324 | 93.9351 | 110.8406 | 101.4189 | 83.9027 | 79.9676 | 228.9676 | 234.7109 | 218.3865 | 396.1298 | 153.1622 | 201.6756 | 170.3893 | 301.5487 | 297.7137 | 208 | 209.3216 | 370.9027
1994 | 23.9284 | 165.0393 | 205.1179 | 210.9293 | 181.0314 | 125.4175 | 167.8978 | 159.1022 | 181.9921 | 155.9686 | 122.0314 | 118.9921 | 176.945 | 204.9214 | 143.9528 | 146.0864 | 210.4882 | 162.4882 | 132.4646 | 132.5118 | 181.6218 | 159.0629 | 301.4332 | 311.1022 | 183.9921 | 204.5275 | 207.0079 | 224.0079 | 322.4882 | 133.0707 | 124.9686 | 110.9686 | 116.4725 | 127.4489 | 126.5589 | 148.4646 | 94.9921 | 220.1022 | 215.6061 | 239 | 229.4646 | 125.5511 | 110.8743 | 128.5039 | 120.9528 | 73 | 59.9686 | 84.5511 | 112.0314 | 174.9921 | 188.9921 | 193.5747 | 146.8507 | 121.1493 | 121.4804 | 120.4725 | 105.5432 | 166.6297 | 157.9057 | 195.5354 | 154.4568 | 141.5511 | 124.4332 | 158.5118 | 139.5039 | 156.5275 | 191.5196 | 197.4175 | 249.1336 | 152.4489 | 141.055 | 204.5747 | 200.9764 | 184.3625 | 180.5668 | 198.4332 | 378.0786 | 153.0943 | 157.0707 | 136.0786 | 221.4489 | 126.4725 | 144.3546 | 122.1336 | 164.5825 | 187.9371 | 166.4018 | 164.5668 | 127.614 | 119.4646 | 108.9843 | 104.5589 | 118.4568 | 246.9214 | 211.0079 | 209.6533 | 373.4882 | 243.0079 | 215.4646 | 226.3939 | 299.9528 | 141.4489 | 70.5118 | 77.4332 | 141.4489 | 117.945 | 110.9293 | 121.0707 | 109.4018 | 99.6218 | 95.8821 | 102.055 | 111.9764 | 68.4568 | 73.945 | 73.9921 | 89.5511 | 99.0079 | 101.5196 | 88.9843 | 74.5511 | 221.0707 | 227.5039 | 225.9686 | 390.0157 | 149.4882 | 206.5432 | 168.9136 | 328.7721 | 284.9214 | 211.9686 | 189.7004 | 379.9528
1995 | 23.9394 | 164.0669 | 195.592 | 216.4415 | 178.0167 | 138.0334 | 165.7492 | 139.3913 | 180.9666 | 142.7157 | 119.5251 | 124.0669 | 158.0736 | 205.8495 | 152.0334 | 138.0502 | 229.7927 | 162.9833 | 145.1338 | 126.9331 | 172.1003 | 145.408 | 307.9666 | 300.0334 | 180.4247 | 215.7425 | 195.3244 | 217.408 | 327.0502 | 121.4582 | 119.8495 | 115 | 120.5084 | 132.4749 | 124.592 | 153.5084 | 97.0167 | 221.7425 | 213.6923 | 237.4749 | 245.6923 | 122.5585 | 119.8829 | 119.3579 | 125.9833 | 65.8829 | 69.592 | 76.9833 | 111.5585 | 194.301 | 188.9833 | 195.184 | 150.2408 | 112.6756 | 118.408 | 125.5251 | 105.592 | 137.7993 | 164.408 | 184.3913 | 177.7927 | 143.6421 | 133.5084 | 153.9498 | 136.9666 | 159.1003 | 176.291 | 212.5753 | 230.9833 | 157.9833 | 153.8261 | 204.1505 | 199.4247 | 191.8328 | 191.3177 | 202.4247 | 361.3913 | 150.6589 | 148 | 131.592 | 227.4916 | 135.0836 | 146.7324 | 112.6254 | 146.8829 | 177.699 | 178.4916 | 173.2843 | 126.2174 | 120.4415 | 107.4415 | 96.4916 | 125.0167 | 232.5986 | 228.8094 | 206.7759 | 377.0334 | 222.1739 | 221.5251 | 226.7826 | 307.5251 | 137.3244 | 74.0836 | 75.3244 | 137.3244 | 117.3746 | 115.9331 | 115.0502 | 124.0334 | 97.7258 | 111.5084 | 104.1505 | 112.9666 | 71.4582 | 72.3579 | 71.4415 | 83.5084 | 98 | 104.0836 | 98.1171 | 76.1338 | 216.0669 | 221.408 | 223.8997 | 391.5585 | 147.9498 | 187.7826 | 178.9833 | 330.5485 | 281.7826 | 206.8495 | 188.4014 | 380.9164
1996 | 23.9504 | 167.4793 | 198.3967 | 207.2728 | 180.562 | 140.5207 | 142.1489 | 155.3881 | 177.9587 | 143.7022 | 123.0826 | 118.6283 | 150.1815 | 198.562 | 147.2315 | 140.4793 | 236.5869 | 158.8761 | 147.7935 | 128.2065 | 187.8924 | 154.5783 | 334.1152 | 300.438 | 167.1489 | 219.0663 | 200.6196 | 198.463 | 324.2728 | 128.3717 | 109.9587 | 120.7272 | 129.8511 | 140.3717 | 134.1652 | 137.3391 | 94.3554 | 232.3554 | 233.8511 | 222.9837 | 266.8924 | 131.2065 | 114.0413 | 124.0163 | 138.0163 | 74.6196 | 76.562 | 80.1652 | 113.438 | 196.3804 | 173.9424 | 213.8098 | 156.8674 | 113.6283 | 120.2891 | 121.2728 | 96.4217 | 136.9337 | 165.2478 | 175.9174 | 170.0087 | 147.8348 | 130.3554 | 159.8511 | 150.0989 | 171.2478 | 177.537 | 192.5293 | 239.3717 | 148.6696 | 160.5043 | 202.5869 | 203.3304 | 205.4294 | 194.9011 | 190.7109 | 359.6859 | 158.9587 | 144.8761 | 147.9337 | 224.9174 | 145.2065 | 131 | 127.8098 | 144.1652 | 153.2315 | 171.2315 | 182.7109 | 145.2478 | 117 | 104.5207 | 98.0826 | 121.8348 | 216.8098 | 238.1489 | 213.6283 | 373.2728 | 220.1815 | 214.1489 | 227.0163 | 306.9174 | 132.7272 | 73.7935 | 72.2891 | 132.7272 | 110.5207 | 108.3554 | 103.4217 | 122.8761 | 116.2065 | 107.3141 | 101.025 | 105.7935 | 73.1652 | 67.6446 | 63.3141 | 85.0413 | 102.1652 | 100.6696 | 104.4793 | 76.1902 | 245.512 | 223.2891 | 204.9837 | 402.2891 | 152.2891 | 178.6446 | 187.8924 | 322.2228 | 271.6033 | 199.562 | 203.1489 | 388.4957
1997 | 23.9614 | 157.407 | 191.7365 | 197.2035 | 189.4612 | 143.1318 | 148.3895 | 164.2752 | 181.2635 | 155.2035 | 117.5388 | 119.06 | 169.1977 | 190.407 | 148.4612 | 142.1318 | 212.0776 | 167.1918 | 138.7365 | 142.593 | 201.2635 | 163.2035 | 348.8082 | 291.0058 | 153.6706 | 213.7965 | 216.0659 | 212.1142 | 321.1318 | 127.94 | 98.3411 | 122.8023 | 136.4012 | 139.94 | 114.5505 | 133.7248 | 92.5988 | 226.3353 | 250.5271 | 224.3236 | 276.5329 | 131.2035 | 113.4012 | 140.2635 | 142.5388 | 86.3353 | 73.7365 | 77.0717 | 101.8741 | 193.7248 | 167.9283 | 209.2752 | 162.6105 | 110.8624 | 123.2694 | 118.6647 | 116.0425 | 163.5213 | 173.6647 | 194.2519 | 160.3236 | 147.1977 | 128.0659 | 155.2094 | 156.0058 | 184.4612 | 197.9942 | 187.0542 | 246.4012 | 143.1318 | 154.3294 | 204.1918 | 207.8023 | 221.1376 | 181.5329 | 181.8682 | 357.0717 | 155.8682 | 143.0659 | 156.4012 | 219.8023 | 137.7423 | 127.8023 | 140.3294 | 149.5988 | 161.9224 | 163.9341 | 181.8624 | 137.1435 | 116.4671 | 100.7365 | 94.6706 | 129.7248 | 219.7365 | 251.3177 | 219.3895 | 372.7965 | 225.8741 | 213.9942 | 231.5388 | 315.6589 | 143.3294 | 72.1977 | 72.0717 | 143.3294 | 112.5988 | 109.2635 | 103.8565 | 127.4612 | 116.7365 | 112.593 | 108.653 | 113.2577 | 78.0659 | 77.3953 | 64.3294 | 86 | 95.3411 | 79.1435 | 91.2094 | 75.3953 | 254.0776 | 228.9341 | 213.7848 | 394.0776 | 156.3353 | 182 | 187.407 | 331.5388 | 291.0542 | 203.6647 | 201.3953 | 397.3353
1998 | 23.9724 | 153.9071 | 192.9071 | 191.3643 | 187.2762 | 141.7286 | 150.0953 | 154.5452 | 183.9095 | 148.8191 | 119.7238 | 129.7262 | 172.5452 | 183.6357 | 164.269 | 147.2714 | 205.5428 | 173.1834 | 137.1809 | 154.2714 | 199.0024 | 148.0953 | 319.9168 | 310.7165 | 167.538 | 208.1858 | 225.1785 | 220.2787 | 323.5452 | 117.2738 | 106.9927 | 125.9976 | 135 | 136.9071 | 121.2617 | 132.0048 | 88.0024 | 234.9047 | 237.8263 | 234.9095 | 286.8142 | 138.9952 | 114.7262 | 121.1001 | 128.3667 | 74.731 | 65.6381 | 67.1834 | 100.0881 | 186.0073 | 191.9927 | 206.6333 | 164.1761 | 122.5428 | 115.6381 | 117.7286 | 130.8239 | 166.5501 | 175.4548 | 203.2762 | 168.1834 | 144.5477 | 131.1809 | 147.8166 | 156.0881 | 165.9192 | 194.0953 | 187.2787 | 242.2738 | 148.8166 | 150.8215 | 204.731 | 214.269 | 204.0048 | 178.1834 | 175.0929 | 356.9976 | 149.6381 | 152.7238 | 158.8166 | 218.6357 | 134.0881 | 130.4523 | 146.6357 | 165.1761 | 182.0881 | 161.9095 | 185.9095 | 126.0905 | 106.1858 | 114.4475 | 92.1809 | 124.7334 | 230.7213 | 250.3716 | 215.643 | 359.5525 | 221.3619 | 205.7334 | 221.1834 | 308.643 | 148.5452 | 75.5452 | 78.5404 | 148.5452 | 105.2762 | 112.4548 | 106.731 | 117.0977 | 117.3619 | 110.0953 | 117.9119 | 118.0024 | 77.9095 | 69.3691 | 79.3595 | 86.5452 | 101.8118 | 78.9952 | 85.4523 | 75.5477 | 244.2714 | 224.7286 | 224.3667 | 390.8142 | 159.9976 | 186.3619 | 180.0905 | 333.7238 | 303.8191 | 196.731 | 199.3667 | 375.3716
1999 | 23.9834 | 152.4247 | 190.8672 | 186.6548 | 183.9027 | 135.0973 | 144.9027 | 158.3452 | 174.0796 | 141.4247 | 123.0973 | 128.6548 | 174.6726 | 188.3452 | 151.3715 | 148.8849 | 220.4779 | 163.8672 | 140.6726 | 158.1151 | 194 | 144.5753 | 309.5753 | 311.2565 | 175.7522 | 209.4779 | 221.407 | 214.6903 | 324 | 115.5575 | 118.6548 | 112.6017 | 126.0796 | 137.6548 | 122.814 | 125.3452 | 85.7876 | 223.3715 | 242.7435 | 220.0619 | 267.1237 | 135.6194 | 120.3452 | 122.0709 | 116.4247 | 79.8231 | 67.5753 | 69.0177 | 107.1151 | 174.9027 | 204.7699 | 194.0441 | 159.8318 | 122.4247 | 123.1505 | 122.8054 | 122.9027 | 158.6726 | 169.9823 | 203.8054 | 170.5753 | 143.9027 | 136.9027 | 144.8672 | 151.9645 | 158.9381 | 167.1592 | 186.708 | 242.7876 | 154.2301 | 156.2656 | 192.5398 | 216.9823 | 196.2301 | 178.9027 | 182.1505 | 367.5753 | 154.3629 | 156.6548 | 155.3097 | 221.6726 | 140.5575 | 117.7167 | 144.0973 | 168.0796 | 186.3274 | 168.8054 | 159.9113 | 119.7522 | 110.2656 | 115.0619 | 103.4779 | 117.0177 | 237.4247 | 231.2124 | 210.3629 | 333.8495 | 225 | 203.0355 | 204.6194 | 306.1505 | 149 | 81.5753 | 83.9823 | 149 | 106.9204 | 107.5398 | 107.3629 | 121.9559 | 127.1328 | 96.5398 | 109.7699 | 114.6726 | 70.3097 | 66.3629 | 81.3097 | 89.7876 | 103.292 | 84.5398 | 96.6903 | 69.8849 | 244.7699 | 222 | 224.1328 | 416.8408 | 155.5221 | 195.0177 | 189.9204 | 360.5134 | 302 | 187.8849 | 198.5753 | 369.593
2000 | 23.9944 | 149.7095 | 186 | 195.3963 | 179.0226 | 149.0945 | 151.4189 | 172.3963 | 160.732 | 141.5585 | 114.8716 | 125.4302 | 164.6037 | 196.6982 | 142.4076 | 141.732 | 229.7095 | 172.1058 | 147.6982 | 137.9168 | 185.4527 | 152.4189 | 324.8265 | 317.9323 | 156.3244 | 212.4415 | 209.0113 | 227.9774 | 324 | 107.4527 | 123.9774 | 110.536 | 123.5585 | 143.5473 | 136.2003 | 131.4189 | 87.4302 | 224.0945 | 252.5924 | 211.2793 | 254.1171 | 122.1509 | 124.7095 | 123.7546 | 113.1396 | 75.7546 | 83.9662 | 71.2905 | 109.1396 | 192.8152 | 215.536 | 190.9662 | 143.5811 | 137.9436 | 118.3244 | 116.464 | 117.4527 | 161.1396 | 163.1509 | 220.8265 | 162.464 | 149.2793 | 128.6037 | 141.1396 | 147.4189 | 162.5924 | 182.6305 | 187.4527 | 237.0226 | 138.9055 | 151.7546 | 210.0832 | 201.6037 | 195.1509 | 178.5811 | 189.2905 | 369.1509 | 157.0113 | 141.464 | 156.8378 | 236.6756 | 152.3963 | 129.6418 | 130.1735 | 142.196 | 178.732 | 173.2905 | 165.0719 | 128.2454 | 113.732 | 97.1622 | 108.7207 | 122.1396 | 235.8491 | 231.2793 | 222.6982 | 339.8152 | 210.7546 | 214.9887 | 211.0945 | 301.8942 | 149 | 70.6149 | 84.5585 | 149 | 125.9662 | 113.1171 | 108.8716 | 123.3244 | 113.196 | 100.4076 | 101.732 | 124.5473 | 71.268 | 65.5924 | 77.1396 | 94.2793 | 96.8491 | 83.2793 | 96.3018 | 71.8491 | 254.3963 | 231.1171 | 228.4302 | 425.3018 | 150.8491 | 185.3244 | 178.196 | 356.6262 | 299.1509 | 186.4302 | 197.8716 | 379.1396
2001 | 24.0054 | 159.1495 | 188.9105 | 201.6716 | 193.9559 | 149.7757 | 172.0453 | 153.6409 | 145.5221 | 160.1348 | 126.1348 | 130.8211 | 147.2684 | 194.0147 | 156.1495 | 139.3284 | 230.4179 | 155.223 | 159.5674 | 136.5527 | 187.1495 | 157.9105 | 326.0147 | 332.5074 | 156.4779 | 210.1642 | 200.761 | 210.1336 | 314.6863 | 114.3885 | 122.4326 | 114.1789 | 128.1642 | 138.94 | 146.1936 | 121.7757 | 87 | 237.5821 | 260.3885 | 210.0895 | 265.6569 | 127.5674 | 121.9252 | 117.6569 | 133.2096 | 77.2243 | 83.6863 | 71.7463 | 111.1642 | 197.5968 | 194.7304 | 193.0147 | 156.7169 | 137.8652 | 117.8958 | 112.239 | 117.9853 | 155.0147 | 159.2537 | 215.0294 | 160.5674 | 138.1936 | 129.3137 | 155.9706 | 155.8211 | 164.5674 | 193.94 | 191.4779 | 260.1054 | 135.3137 | 133.4326 | 210.8652 | 188.3431 | 208.1348 | 184.7316 | 181.0147 | 352.4473 | 162.1495 | 138.9853 | 152.1042 | 237.6863 | 159.8358 | 146.6716 | 124.3284 | 154.1949 | 194.3738 | 183.06 | 185.7463 | 139.5821 | 107.2537 | 82.6863 | 100.0147 | 127.0748 | 258.3738 | 227.761 | 229.9105 | 363.2243 | 221.538 | 209.2684 | 220.5074 | 325.1801 | 152.4926 | 71.2243 | 77.5221 | 152.4926 | 136.1642 | 110.6863 | 116.06 | 133.9559 | 92.5968 | 101.9252 | 105.7316 | 122.8505 | 87.06 | 58.2537 | 72.1789 | 92.5074 | 99 | 95.4779 | 99.5674 | 68.761 | 244.3725 | 240.9105 | 221.0147 | 412.8505 | 154.7463 | 186.06 | 183.7917 | 362.9559 | 281.2831 | 188.3284 | 200.4032 | 368.3578
2002 | 24.0164 | 150.7344 | 200.5104 | 200 | 183.4408 | 147.5384 | 179.056 | 144.1048 | 149.2936 | 165.056 | 124.5176 | 122.5176 | 149.916 | 200.888 | 157.8392 | 156.4544 | 228.2168 | 160.9648 | 166.7832 | 142.8392 | 198.944 | 165.944 | 306.14 | 324.056 | 156.8672 | 215.1608 | 205.3216 | 209.6432 | 294.3288 | 106.168 | 127.1048 | 109.4056 | 138.5104 | 134.5664 | 139.3776 | 123.6992 | 88.7832 | 216.0072 | 240.8745 | 232.3704 | 274.3496 | 105.6577 | 126.7272 | 127.5384 | 125.0072 | 76.4896 | 74.6224 | 77.1608 | 117.3496 | 190.028 | 198.5592 | 204.0488 | 156.112 | 128.1888 | 128.5664 | 122.5384 | 106.3568 | 161.2936 | 153.2448 | 194.8952 | 168.972 | 124.8392 | 138.972 | 159.4616 | 148.112 | 175.3496 | 192.5384 | 193.056 | 257.224 | 152.6992 | 132.1608 | 199.4056 | 182.6224 | 200.5736 | 192.1888 | 196.2096 | 351.5104 | 150.168 | 154.6992 | 152.1328 | 233.972 | 144.7344 | 131.9232 | 128.972 | 159.9232 | 189.3848 | 191.5944 | 173.3288 | 126.3288 | 110.7552 | 103.3423 | 93.2168 | 119.8952 | 271.8112 | 253.7199 | 223.6784 | 369.028 | 253.6432 | 216.6712 | 209.084 | 326.4128 | 139.5456 | 73.4616 | 74.7552 | 139.5456 | 123.9232 | 108.1608 | 124 | 131.168 | 93.3496 | 90.084 | 103.084 | 131.86 | 73.6016 | 70.0768 | 71.5664 | 102.4824 | 104.944 | 104.1888 | 87.7624 | 75.6992 | 252.3984 | 235.2728 | 217.7832 | 418.2936 | 163.1328 | 210.6432 | 180.168 | 366.112 | 296.1536 | 184.6504 | 204.4056 | 367.1328
2003 | 24.0274 | 145.2468 | 200.3264 | 209.707 | 182.9873 | 139.8662 | 164.0796 | 152.8201 | 163.0669 | 170.1003 | 126.9204 | 115.2134 | 148.7197 | 210.2134 | 161.6736 | 157.293 | 222.1465 | 162.2261 | 149.7993 | 154.5606 | 201.7866 | 140.2722 | 319.6609 | 296.9458 | 161.1003 | 223.4602 | 217.0669 | 215.5398 | 308.6609 | 108.5606 | 129.1465 | 110.2134 | 145.6067 | 151.5606 | 135.5398 | 139.4937 | 87.5732 | 202.8201 | 236.7739 | 245.3599 | 285.887 | 107.6275 | 130.1799 | 128.3599 | 126.9873 | 70.6067 | 66.9331 | 72.113 | 109.4729 | 188 | 200.8328 | 188.9123 | 159.5271 | 129 | 143.1338 | 123.9666 | 116.2341 | 164.1465 | 145.7532 | 191.6401 | 189.8074 | 127.4602 | 139.1799 | 145.8997 | 146.6736 | 171.7197 | 198.8201 | 193.2468 | 251.3137 | 150.8997 | 138.6401 | 205.6736 | 204.0542 | 190.4268 | 193.6067 | 191.7993 | 381.6609 | 147.707 | 152.293 | 157.6067 | 253.594 | 138.0334 | 138.1672 | 136.4602 | 150.3933 | 172.9666 | 180.4729 | 173.707 | 133.3806 | 114.6067 | 108.6528 | 102.9204 | 117.2468 | 254.0127 | 249.1257 | 218.6067 | 348.7993 | 237.699 | 236.3137 | 201.1799 | 343.7613 | 141.1338 | 70.8201 | 81.0334 | 141.1338 | 110.1465 | 111.6067 | 117.9331 | 119.6067 | 106.707 | 97.9539 | 103.6736 | 119.5525 | 63.8535 | 75.9666 | 80.6736 | 100.6863 | 111.4268 | 95.8997 | 84.4937 | 94.5271 | 254.8662 | 211.7993 | 215.9666 | 410.2261 | 154.6528 | 195.9123 | 173.4602 | 361.6401 | 308.5398 | 175.5398 | 204 | 399.9077
2004 | 24.0384 | 162.2366 | 189.8102 | 218.4759 | 172.9546 | 140.6657 | 155.762 | 158.3329 | 170.7139 | 174.762 | 132.762 | 124.0468 | 158.8555 | 198.0014 | 155.4773 | 147.9051 | 227.0468 | 175.5213 | 146.0468 | 162.1431 | 186.7634 | 129.0468 | 328.1912 | 306.1403 | 166.381 | 217.7153 | 220.381 | 220.6657 | 320.2861 | 116.1431 | 109.9065 | 111 | 137.9532 | 149.8583 | 138.8088 | 127.6685 | 90.3329 | 226.2833 | 265.5694 | 231.2394 | 297.8088 | 136.4745 | 124.0482 | 123.5241 | 138 | 88.3315 | 70.4278 | 70.7139 | 115.6176 | 189.238 | 184.1912 | 193.0454 | 168.8569 | 140.1417 | 144.8102 | 125.7139 | 144.2366 | 162.238 | 149.1898 | 190.9051 | 176.0028 | 143.9986 | 146.6657 | 147.4278 | 150.381 | 157.7153 | 186.3824 | 204.0468 | 275.3315 | 143.762 | 160.1884 | 206.9051 | 209.0963 | 201.9037 | 195.8569 | 189.2847 | 374.0977 | 156.4759 | 162.7125 | 154.9051 | 243.9546 | 153.6176 | 152.9518 | 135.0482 | 149.381 | 167.9051 | 188.4745 | 175.6671 | 141.5241 | 113.762 | 99.381 | 106.9051 | 119.381 | 226.9065 | 242.9037 | 246.2352 | 362.9972 | 204.5255 | 234.3343 | 219.1884 | 337.7648 | 156.4278 | 75.0949 | 87.9518 | 156.4278 | 105.1431 | 109.5241 | 113.381 | 124.9518 | 108.0482 | 98.9532 | 94.3824 | 109.9518 | 79.9986 | 77.0949 | 76.3343 | 97.7139 | 122.9037 | 90.619 | 96.1898 | 79.7167 | 270.5213 | 202.4759 | 215.238 | 404.9518 | 157.1417 | 184.5708 | 194.3315 | 358.4292 | 309.3329 | 186.2366 | 178.0028 | 406.2394
2005 | 24.0494 | 166.5811 | 187.8938 | 197.9055 | 170.7316 | 152.944 | 175.832 | 176.7818 | 167.9498 | 181.5753 | 127.5811 | 125.8436 | 175.5753 | 205.7818 | 154.0502 | 147.2625 | 224.4247 | 183.056 | 140.2684 | 150.8996 | 194.4131 | 146.6254 | 306.9557 | 324.7876 | 153.3746 | 220.2065 | 221.2625 | 205.7993 | 289.5928 | 125.7316 | 111.9942 | 104.6873 | 136.1564 | 144.8938 | 137.9498 | 113.2124 | 94.8938 | 251.3629 | 239.1237 | 213.2684 | 269.8051 | 135.8996 | 115.3186 | 130.8378 | 150.6254 | 78.7993 | 75 | 71.7375 | 126.5251 | 203.2567 | 198.3069 | 198.4247 | 178.8378 | 133.112 | 142.8938 | 124.2124 | 133.4306 | 171.2065 | 147.9498 | 188.3687 | 170.7316 | 159.5753 | 146.9498 | 162.7316 | 148.7375 | 167.1505 | 192.5193 | 197.6371 | 258.224 | 139.2124 | 144.8553 | 200.5811 | 220.0443 | 204.2124 | 191.9498 | 179.8495 | 378.3069 | 164.3127 | 180.5753 | 164.3629 | 251.2007 | 147.4807 | 147.7935 | 128.8436 | 150.8938 | 168.5251 | 206.8378 | 179.944 | 148.2065 | 110.4749 | 109.7316 | 106.2625 | 113.3186 | 234.6756 | 245.8436 | 207.4481 | 358.7993 | 220.2509 | 240.9942 | 232.2625 | 328.5753 | 155.3186 | 77.6313 | 92.8938 | 155.3186 | 110.944 | 123.1505 | 106.6873 | 134.944 | 115.1004 | 100.944 | 86 | 110.4749 | 77.2684 | 73.9498 | 65.9498 | 103.7876 | 103.1178 | 96.0502 | 100.6313 | 88.0945 | 269.8495 | 215.9942 | 205.8996 | 401.056 | 152.6371 | 172.2182 | 195.5309 | 341.7433 | 317.0502 | 191.8436 | 172.1004 | 395.2124
2006 | 24.0604 | 174.297 | 196.7228 | 196.5841 | 173.7822 | 167.9406 | 168.6931 | 168.6238 | 168.2178 | 171.198 | 109.5545 | 119.4951 | 183.8614 | 215.6436 | 145.4159 | 138.9901 | 227.7228 | 174.4951 | 153.9505 | 153.3663 | 211.7228 | 162.3663 | 334.9702 | 311.5545 | 161.4455 | 203.1189 | 218.7822 | 215.2376 | 306.4653 | 131.3564 | 106.1287 | 99.7129 | 141.2178 | 136.3466 | 146.5841 | 126.4455 | 89.5644 | 251.5644 | 231.1584 | 209.5743 | 279.9603 | 141.5841 | 122.297 | 119.2674 | 132.901 | 85.6633 | 68.5644 | 76.1485 | 136.3663 | 191.0495 | 198.703 | 202.3663 | 168.5545 | 128.2178 | 138.8515 | 127.1485 | 118.1386 | 179.8614 | 152.0792 | 192.5049 | 167.3466 | 153.703 | 144.6436 | 167.7129 | 162.802 | 155.406 | 196.4951 | 206.4455 | 258.0891 | 137.6436 | 145.0891 | 205.0792 | 230.6436 | 213.5841 | 199.9406 | 187.0891 | 383.2079 | 175.0792 | 169.5545 | 161.9901 | 250.7723 | 135.1386 | 163.594 | 122.4951 | 163.5841 | 188.0198 | 202.3466 | 176.2772 | 134.3367 | 106.4257 | 101.198 | 105.0693 | 120.9406 | 223.1189 | 239.4951 | 204.6039 | 379.1782 | 229.8515 | 211.3169 | 233 | 324.6337 | 166.802 | 71.5644 | 83.703 | 166.802 | 120.7921 | 135.2178 | 117.802 | 144.1485 | 115.8515 | 108.2178 | 89.2178 | 118.6534 | 81.297 | 67.7822 | 67.5049 | 95.0594 | 91.2178 | 106.0792 | 104.8614 | 98.4257 | 276.4455 | 219.7822 | 206.4356 | 403.4356 | 155.6534 | 183.594 | 181.6337 | 335.9307 | 310.3466 | 204.1584 | 172.2079 | 407.802
2007 | 24.0714 | 172.1297 | 193.7856 | 209.5585 | 181.8378 | 150.3892 | 158 | 153.7532 | 165.409 | 177.4288 | 120.0198 | 115.0324 | 185.3441 | 212.0649 | 138.3441 | 161.5459 | 232.6559 | 177.9027 | 144.0127 | 127.8307 | 216 | 164.3766 | 346.1622 | 307.591 | 163.4415 | 214.6108 | 209.7856 | 224.3766 | 302.701 | 129.0324 | 104.2468 | 100.3117 | 150.8703 | 144.1171 | 143.818 | 129.0973 | 78.7856 | 240.1297 | 239.6559 | 248.3837 | 252.0578 | 134.8829 | 141.7729 | 118.5261 | 113.0973 | 85.5063 | 74.8378 | 77.0324 | 120.6685 | 192.4612 | 192.3441 | 197.818 | 177.0522 | 135.2468 | 131.409 | 127.3766 | 115.3441 | 178.0649 | 175.0198 | 193.0324 | 184.9549 | 154.5585 | 135.818 | 173.5585 | 175.591 | 160.3964 | 188.7532 | 209.0973 | 286.6756 | 157.0198 | 152.6883 | 210.3117 | 216.5712 | 198.3568 | 178.4541 | 192.7207 | 366.8829 | 155.3892 | 149.8505 | 143.8829 | 256.4612 | 140.8703 | 157.9478 | 118.6883 | 168.0324 | 196.6883 | 212.0846 | 175.2793 | 141.7081 | 93.8505 | 96.2793 | 108.591 | 130.2793 | 233.9549 | 225.1946 | 215.7532 | 393.7207 | 221.0973 | 192.3117 | 209.3892 | 317.3766 | 155.9802 | 75.8703 | 79.3117 | 155.9802 | 124 | 125.1946 | 118.1297 | 143.0649 | 114.9676 | 108.6883 | 100.182 | 106.9478 | 81.0973 | 77.8054 | 75.9027 | 92.9351 | 95.6234 | 122.4612 | 101.7532 | 85.1946 | 286.3117 | 238.3315 | 221.8054 | 390.6036 | 172.1495 | 187.7856 | 187.4937 | 369.7927 | 284.0127 | 186.4217 | 188.0198 | 395.0127
2008 | 24.0824 | 163.9911 | 197.3496 | 198.969 | 195.0177 | 160.0488 | 151.3185 | 162.3585 | 164.3363 | 172.9734 | 115.3008 | 134.7125 | 194.354 | 193.2964 | 162.7214 | 151.9468 | 233 | 178.3274 | 138.3452 | 137.3895 | 224.0177 | 149.6371 | 365.0532 | 316.0133 | 153.3185 | 224.9956 | 216.6903 | 213.646 | 321.7347 | 132.0089 | 124.3718 | 109.0177 | 142.3052 | 146.9911 | 141.0044 | 136.6903 | 92.3718 | 252.7036 | 241.3363 | 225.2387 | 266.4161 | 131.3407 | 138.3097 | 119.6593 | 106.6593 | 80 | 65.3008 | 69.9867 | 112.6726 | 203.677 | 202.0222 | 213.7125 | 157.9379 | 133.323 | 129 | 115.3097 | 117.0044 | 173.9956 | 164.2875 | 211.3762 | 181.6327 | 153.323 | 148.3718 | 162.969 | 167.3097 | 158.9778 | 188.0044 | 206.6681 | 302.3452 | 164.3274 | 151.3319 | 211.6681 | 212.3407 | 188.3363 | 177.3629 | 172.9601 | 367.3496 | 150.3496 | 141.3274 | 145.6859 | 248.969 | 139.6548 | 154.3452 | 128.0222 | 174.3496 | 196.6681 | 202.6282 | 183.6815 | 144.9867 | 113.3895 | 107.354 | 100.9778 | 123.9823 | 253.3496 | 245.7258 | 215.6726 | 383.9823 | 251.4073 | 195.0044 | 231.7436 | 334.0399 | 154.6859 | 81.3363 | 71.9823 | 154.6859 | 119.9911 | 116.3274 | 117.3407 | 126.969 | 118.0044 | 111.3407 | 113.6859 | 109.3585 | 73.323 | 72.6415 | 86.3496 | 109.031 | 102.3452 | 111.6282 | 98.3319 | 89.0222 | 266.9556 | 241.6504 | 235.3496 | 395.3629 | 178 | 202.0399 | 196.3407 | 405.0399 | 295.7169 | 201.0621 | 206.6859 | 410.0577
2009 | 24.0934 | 149.0717 | 193.5152 | 212.4131 | 189.4739 | 155.35 | 143.9174 | 159.1543 | 167.7217 | 168.3609 | 119.5674 | 125.9478 | 200.3609 | 196.5674 | 151.8652 | 133.1956 | 242.5261 | 164.7521 | 156.65 | 139.7935 | 204.8652 | 172.2588 | 347.0608 | 327.1652 | 162.2479 | 237.6087 | 221.3196 | 236.2175 | 332.2369 | 131.2783 | 106.463 | 119.8044 | 148.2479 | 139.5565 | 128.3913 | 150.8457 | 101 | 237.8652 | 245.4022 | 242.7849 | 277.5152 | 131.6391 | 137.0826 | 120.0413 | 126.774 | 79.3196 | 79.0935 | 71.0826 | 105.1543 | 207.3609 | 211.0826 | 206.3087 | 144.6804 | 133.7217 | 131.0413 | 102.5152 | 122.7631 | 176.4022 | 162.8457 | 195.8239 | 177.4022 | 150.3196 | 147.4739 | 164.1652 | 153.1543 | 171.0109 | 201.9283 | 194.7521 | 284.5869 | 156.1956 | 165.9696 | 199.7521 | 212.6391 | 182.8761 | 182.6391 | 175.5674 | 363.5152 | 175.0935 | 149.5261 | 145.2369 | 228.3913 | 144.4848 | 150.1956 | 132.3196 | 163.0304 | 211.9696 | 209.65 | 199.9283 | 146.0826 | 108.3087 | 105.8761 | 91.9174 | 111.1543 | 245.4326 | 243.35 | 217 | 371.8348 | 244.1847 | 212.3305 | 243.5565 | 337.5565 | 153.5565 | 79.9587 | 80.2479 | 153.5565 | 124.1239 | 112.2783 | 109.4739 | 145.8566 | 119.6804 | 123.2065 | 108.4739 | 141.537 | 71 | 78.5674 | 67.5456 | 102.3913 | 127.4544 | 107.0826 | 93.9174 | 87.8761 | 259.0413 | 219.6282 | 227.4326 | 409.4848 | 171.1956 | 211.6804 | 202.0826 | 394.9478 | 303.5978 | 182.3391 | 203.5152 | 410.0717
2010 | 24.1044 | 149.2345 | 196.9273 | 220.6145 | 195.3909 | 148 | 151.0055 | 143.9164 | 165.5364 | 189.7818 | 131.9273 | 121.1564 | 198.9218 | 206.1564 | 135.4582 | 147.3182 | 232.4527 | 179.0891 | 149.4527 | 144.0055 | 211.3182 | 154.1345 | 348.24 | 324.7655 | 171.4636 | 229.4527 | 225.8491 | 228.2182 | 300.2127 | 129.3073 | 110.6255 | 120.2291 | 146.38 | 138.3855 | 133.0836 | 151.5364 | 98.2291 | 231.1564 | 244.9218 | 240.2182 | 274 | 147.6255 | 132.7655 | 116.8436 | 131.4582 | 80.3855 | 73.0673 | 63.9945 | 111.3909 | 221.1618 | 202.6091 | 204.2345 | 175.48 | 127.38 | 122.3018 | 110.7764 | 129.1564 | 182.1564 | 178.0836 | 187.4636 | 160.9891 | 162.4691 | 156.1618 | 170.0782 | 158.0055 | 176.9218 | 199.6873 | 205.6255 | 305.48 | 153.6927 | 159.8382 | 220.3236 | 209.2291 | 193.8545 | 177.8436 | 196.24 | 369.6982 | 162.8327 | 144.9945 | 143 | 249.0164 | 160.4691 | 160.8545 | 154.1673 | 165.6982 | 212.0727 | 208.6873 | 190.76 | 165.3182 | 120.0891 | 99.5364 | 91.3855 | 113.9273 | 244.8491 | 256.7818 | 224.62 | 368 | 221.9164 | 228.3127 | 220.9109 | 344.0055 | 149.6145 | 81.7709 | 75.6091 | 149.6145 | 131.8491 | 125.5473 | 122.3182 | 131.6764 | 111.6873 | 128 | 101.9218 | 144.3018 | 75.8491 | 77.0727 | 65.3127 | 83.5309 | 115.14 | 103.4582 | 101.6982 | 88.4636 | 252.38 | 216.5418 | 217.1509 | 390.8327 | 171.4636 | 221.0055 | 198.4582 | 381.1509 | 297.8436 | 205.7927 | 184.0673 | 419.9327
2011 | 24.1154 | 139.3097 | 182.3745 | 214.3599 | 188.7197 | 163.5104 | 174.0355 | 167.2007 | 166.1151 | 186.3097 | 132.1799 | 122.295 | 205.7552 | 225.6255 | 138.6401 | 147.2448 | 224.59 | 190.1151 | 153.5753 | 143.7699 | 228.1652 | 150.5753 | 325.3892 | 303.6695 | 163.1297 | 225.8201 | 214.6047 | 218.295 | 325.776 | 117.0147 | 101.7846 | 101.3745 | 140.8849 | 126.3097 | 143.6401 | 145.0649 | 104.0502 | 255.5606 | 232.0147 | 231 | 273.295 | 155.705 | 137.0502 | 138.9706 | 130.41 | 75.3599 | 77.2803 | 80.4455 | 107.5398 | 220.6548 | 210.6903 | 221.1004 | 186.1799 | 126.8201 | 129.2803 | 127.9853 | 128.1799 | 178.3599 | 171.7197 | 210.1506 | 168.5104 | 168 | 142.9645 | 192.1506 | 173.2803 | 169.6548 | 202.3452 | 221.4602 | 297.8494 | 166.6903 | 151.1799 | 225.6548 | 226.3305 | 207.7552 | 185.8703 | 186.7846 | 353.5545 | 178.3807 | 145.2301 | 151.4602 | 252.5398 | 173.0502 | 157.1297 | 160.4749 | 174.9351 | 230.1506 | 225.4455 | 185.41 | 151.8494 | 126.8849 | 117.0355 | 92.705 | 115.59 | 233.6047 | 254.7197 | 224.4749 | 391.2656 | 223.3452 | 220.0147 | 207.7699 | 328.9645 | 151.8201 | 78.7699 | 77.3452 | 151.8201 | 139.6401 | 122.8348 | 130.705 | 137.6255 | 121.3953 | 125.8849 | 94.6548 | 134.3599 | 70.9498 | 68.3599 | 72.5251 | 90.6903 | 112.0502 | 119.3305 | 101.0649 | 98.4602 | 244.7699 | 221.8201 | 225.5753 | 384.5251 | 168.7699 | 219.3599 | 198.8201 | 383.2301 | 317.1506 | 196.2093 | 187.5753 | 414.3097
2012 | 24.1264 | 156.2366 | 186.477 | 196.2192 | 197.6289 | 157.8057 | 166.9365 | 174.6961 | 168.4346 | 170.2404 | 131.7173 | 119.1308 | 210.4346 | 222.9577 | 135.9788 | 168.3885 | 233.325 | 186.6961 | 168.0423 | 144.1519 | 226.2615 | 147.1096 | 336.6712 | 306.7596 | 180.5193 | 221.2615 | 236.9751 | 226.6077 | 349.1731 | 117.7385 | 115.802 | 104.7596 | 152.9116 | 136.7808 | 135.2404 | 150.1731 | 100.5442 | 226.9826 | 241.3462 | 235.3039 | 280.8904 | 159.5865 | 148.6077 | 121.7422 | 149.65 | 73 | 79.8481 | 77.523 | 122.65 | 222.3039 | 212.4135 | 223.4135 | 193.6077 | 120.1096 | 140.4558 | 133.7173 | 130.5865 | 209.7135 | 163.4135 | 194.6115 | 169.2615 | 180.9116 | 157.2366 | 188.0884 | 171.5442 | 159.1096 | 197.8269 | 208.5019 | 306.2154 | 166.2615 | 147.8481 | 207.9365 | 244.0423 | 218.8904 | 196.4558 | 190.7596 | 385.8866 | 170.35 | 157.0423 | 152.8481 | 283.4309 | 163.8057 | 173.0847 | 148.9577 | 165.523 | 239.7173 | 234.7173 | 191.0212 | 167.3885 | 126.7173 | 110.6538 | 98.7385 | 117.1519 | 222.9788 | 271.5193 | 223.7173 | 399.5654 | 210.2192 | 235.0847 | 211.7385 | 310.2404 | 177.3885 | 72.6961 | 80.7173 | 177.3885 | 130.523 | 135.4981 | 123.1096 | 149.3039 | 119.8269 | 105.6326 | 99.8904 | 125.5442 | 73.7385 | 71.0212 | 79.0212 | 96 | 104.2404 | 137.0423 | 91.8269 | 83.35 | 248.0212 | 249.5404 | 222.8269 | 387.4346 | 180.6289 | 212.6961 | 186.3711 | 364.9153 | 326 | 190.7385 | 194.1519 | 407.5654
2013 | 24.1374 | 161.352 | 195.3776 | 197.296 | 190.5968 | 145.704 | 179.24 | 145.2752 | 186.5104 | 170.3776 | 137.1072 | 122.3776 | 210.2704 | 215.352 | 134 | 154.6528 | 253.7808 | 195.944 | 171.2704 | 155.944 | 230.5664 | 138.1632 | 341.6224 | 277.4384 | 177.3264 | 216.0816 | 250.1888 | 254.0768 | 326.464 | 112.704 | 108.6784 | 107.5408 | 141.2192 | 151.0256 | 131 | 141.3264 | 91.4336 | 226.592 | 247 | 222.408 | 277.4336 | 158.8112 | 128.6528 | 142.3728 | 138.0304 | 83.944 | 79.7296 | 65.704 | 116.1376 | 218.1632 | 202.2448 | 208.1376 | 175.8416 | 135.9696 | 141.5408 | 141.296 | 129.8112 | 194.5456 | 172.2144 | 202.5104 | 175.7552 | 179.4336 | 183.5104 | 185.1888 | 193.8064 | 180.8064 | 213.24 | 223.888 | 313.7296 | 155.2448 | 147.7296 | 208.5664 | 233.408 | 233.6736 | 203.3776 | 187.704 | 369.168 | 165.1888 | 144.9488 | 155.648 | 270.0048 | 176.5104 | 183.1888 | 142.8112 | 178.5104 | 228.3264 | 255.4288 | 171.112 | 168.9744 | 121.8928 | 106.4592 | 106.8368 | 112.8928 | 251.6432 | 267.5968 | 214.5152 | 418.6992 | 221.5104 | 230.5968 | 203.056 | 311.1072 | 166.5712 | 81.2144 | 80.2704 | 166.5712 | 136.2144 | 140.5408 | 127.296 | 149.5408 | 118.4592 | 124.2656 | 111.7552 | 130.296 | 80.3776 | 82.4848 | 74.4336 | 89.4336 | 105.8368 | 116.1936 | 107.9696 | 83.296 | 279.184 | 262.1888 | 232.4032 | 405.5104 | 201.3216 | 235.0768 | 190.4848 | 358.4592 | 310.6784 | 195.9184 | 210.3216 | 400.4336
2014 | 24.1484 | 165.1932 | 185.8716 | 210.3865 | 171.9041 | 146.7095 | 173.3878 | 147.6122 | 179.646 | 183.1284 | 148.6446 | 124 | 213.7095 | 251.0946 | 156.2567 | 158.6122 | 233.2919 | 192.5811 | 183.6122 | 143.6784 | 232.2581 | 158.9986 | 371.1594 | 293.1919 | 161.1297 | 224.6446 | 238.3878 | 263 | 351.8689 | 129.2892 | 122.2892 | 111.4513 | 141.677 | 138.4203 | 144.354 | 150.354 | 93.4513 | 222.3554 | 229.1946 | 213.2905 | 283.1608 | 166.9027 | 132.6122 | 141.3878 | 142.1284 | 75.3878 | 85.1932 | 74.8703 | 138.45 | 223.4189 | 202.7095 | 213.3865 | 187.2892 | 141.5162 | 142.4838 | 122.4851 | 131.2257 | 166.9365 | 174.5162 | 226.8054 | 185.677 | 176.2581 | 166.2595 | 171.9041 | 197.8068 | 176.646 | 213.323 | 209.0014 | 302.8716 | 160.1608 | 144.2905 | 192.4527 | 223.5487 | 215.7433 | 202.0324 | 192.4189 | 404.4811 | 150.4203 | 164.2243 | 148.8392 | 246.646 | 170.3878 | 155.0662 | 144.9676 | 180.5487 | 254.4176 | 246.6784 | 177.8378 | 171.1932 | 107.3878 | 109.9676 | 96.3878 | 113.9676 | 242.9689 | 274.1284 | 204.323 | 385.9379 | 240.6122 | 203.0014 | 194.5487 | 333.773 | 158.2581 | 82.0324 | 76.2905 | 158.2581 | 141.4838 | 128.8716 | 118.1297 | 133.4203 | 107.1297 | 119.1622 | 109.0649 | 137.4513 | 76.0649 | 91.1932 | 87.5797 | 115.1919 | 106.5162 | 125.5473 | 100.9041 | 80.8068 | 288.5162 | 274.8703 | 231.8068 | 423.1284 | 189.1946 | 254.3865 | 188.8068 | 350.0973 | 313.1608 | 182.9041 | 195.9689 | 406.1608
2015 | 24.1594 | 173.7876 | 204.6255 | 223.0502 | 194.1506 | 159.3127 | 177.296 | 153.5084 | 176.5084 | 191.5251 | 155.7709 | 121.7375 | 185.5869 | 243.6371 | 155.704 | 135.0952 | 260.9549 | 211.8713 | 184.2291 | 139.5084 | 219.1789 | 147.3913 | 345.0451 | 311.296 | 158.5084 | 234.7876 | 243.0502 | 220.7658 | 344.3913 | 125.4414 | 132.7709 | 111.4916 | 147.0167 | 141.296 | 142.2124 | 142.9331 | 95.7542 | 232.5753 | 233.5586 | 210.4916 | 273.1789 | 155.6705 | 130.9665 | 133.9833 | 149.7709 | 86.8378 | 74.1789 | 70.704 | 140.4582 | 222.2291 | 204 | 202.6705 | 209.8378 | 141 | 136.9665 | 105.9498 | 137.2793 | 183.6255 | 180.0335 | 205.8494 | 188 | 179.0167 | 164.7876 | 163.9833 | 203.5418 | 165.9665 | 208.7375 | 206.2793 | 302.7709 | 179.592 | 168.6422 | 185.2458 | 255.9382 | 213.2793 | 207.7876 | 201.7876 | 402.8996 | 182.7091 | 183.5586 | 163.3462 | 266.888 | 162.9833 | 193.2677 | 139.9665 | 194.0837 | 225.0283 | 252.3127 | 183.7542 | 168.4749 | 95.4582 | 120.0502 | 87.4749 | 117.2625 | 247.3127 | 264.4247 | 215.5753 | 396.1338 | 242.7375 | 220.6422 | 217.1338 | 329.6873 | 168.5586 | 81.7542 | 79.5251 | 168.5586 | 129.1789 | 122.7375 | 123.0502 | 146.8546 | 120.3462 | 110.9833 | 102.4749 | 146.5418 | 88.3295 | 79.4247 | 87.7207 | 112.9331 | 123.3462 | 127.4749 | 98.2625 | 72.2124 | 288 | 282.0167 | 215.6705 | 419.4582 | 192.8044 | 244.4247 | 201.3295 | 366.6087 | 323.5418 | 200.6255 | 204.8378 | 409.7542
2016 | 24.1704 | 158.3711 | 193.6046 | 210.6705 | 204.5329 | 164.5329 | 176.1676 | 157.8324 | 162.437 | 194.5329 | 154.7006 | 130.9642 | 176.7665 | 253.0959 | 141.5029 | 152.0601 | 246.1734 | 183.7422 | 191.4312 | 137.7006 | 231.0959 | 157.8624 | 374.3896 | 336.2277 | 183.5272 | 260.7607 | 255.1977 | 210.8324 | 363.2936 | 131.9642 | 122.5029 | 116.3653 | 127.3052 | 162.3954 | 144.5988 | 140.5329 | 93.7006 | 265.3595 | 243.8983 | 216.8983 | 281.2636 | 151 | 134.3653 | 128.4012 | 146.4012 | 99.6647 | 76.1318 | 72.5988 | 121.9041 | 229.4312 | 248.4555 | 211.0301 | 182.0416 | 145.5988 | 152.6289 | 109.8983 | 128.2694 | 186.4012 | 182 | 205.4312 | 202.563 | 182.2994 | 177.7306 | 188.2936 | 192.9699 | 182.3954 | 224.8624 | 227.1618 | 303.2335 | 185.7665 | 168.5688 | 213.3595 | 233.2752 | 214.2335 | 188.5387 | 206.2994 | 391.6347 | 162.0416 | 194.6647 | 159.0358 | 242.0416 | 166.5988 | 180.6405 | 167.8925 | 192.8682 | 209.7006 | 274.3954 | 186.2994 | 159.3353 | 123.6589 | 108.437 | 95.1977 | 126.4312 | 228.0058 | 275.3295 | 222.2994 | 391.7364 | 259.6289 | 225.1676 | 231.1318 | 358.1919 | 163.5688 | 68.9699 | 71.8023 | 163.5688 | 142.6289 | 134.2636 | 112.2035 | 149.1676 | 110.6705 | 124.563 | 100.2335 | 123.7064 | 84.5688 | 84.9642 | 72.9699 | 113.5988 | 115.2035 | 112.2035 | 92.1017 | 83.7965 | 295.6647 | 276.8682 | 213.2994 | 388.6405 | 183.7364 | 223.1376 | 186.8382 | 386.0301 | 326.7665 | 219.4971 | 200.8023 | 431.4613
2017 | 24.1814 | 149.1062 | 193.4514 | 194.3185 | 194.8761 | 150.2035 | 167.2124 | 165.2301 | 169.6815 | 173.9734 | 152.4425 | 130.1062 | 187.9027 | 236.9734 | 159.0266 | 157.6637 | 243.4514 | 188.5753 | 191.6637 | 155.6904 | 214.9734 | 165.3363 | 348.8405 | 311.0706 | 176.9822 | 274.2301 | 268.9027 | 254.0533 | 370.2212 | 121.761 | 101.0884 | 113.3274 | 140.4691 | 141.522 | 160.7965 | 144.1151 | 107.7965 | 248.3008 | 238.2124 | 225.2301 | 295.9027 | 147.1062 | 127.4336 | 123.8849 | 152.0089 | 86.4247 | 60.0884 | 68.5486 | 138.0266 | 215.6459 | 238.6371 | 224.3452 | 180.5664 | 137.6548 | 140.0884 | 128.3541 | 115.6548 | 202.9116 | 159.4158 | 193.2035 | 197.6548 | 183.7788 | 169.3185 | 192.8849 | 206.1328 | 179.4336 | 205.0795 | 207.3008 | 275.7432 | 180.5486 | 196.3718 | 222.7788 | 219.1062 | 205.4336 | 209.2568 | 178.9645 | 411.0266 | 162.9027 | 162.7344 | 150.5486 | 263.1506 | 132.9556 | 207.8232 | 141.9556 | 182.4336 | 216.7876 | 265.2035 | 194.0089 | 152.3274 | 118.2035 | 112.5664 | 112.0178 | 113.4247 | 210.0973 | 262.8672 | 199.6371 | 391.1151 | 268.1151 | 224 | 228.3274 | 341.522 | 177.3541 | 70.4514 | 63.5486 | 177.3541 | 135.5398 | 127.0973 | 126.6904 | 144.8849 | 120.7965 | 116.5398 | 95.3274 | 129.239 | 89.7876 | 78.6548 | 75.2301 | 106.4336 | 122.6815 | 141.4869 | 90.7788 | 88 | 282.4247 | 267.9911 | 233.4691 | 402.5842 | 181.5575 | 226.5664 | 187.2301 | 356.5131 | 313.761 | 202.7521 | 194.8849 | 427.8761
2018 | 24.1924 | 165.4032 | 208.4479 | 214.7317 | 199.9106 | 150.7463 | 167.3732 | 170.9553 | 192.7764 | 175.1195 | 148.8358 | 126.6268 | 198.1195 | 227.8358 | 151.5521 | 148.2984 | 263.1943 | 219.8959 | 200.4927 | 149.9252 | 223.239 | 148.5968 | 375.1796 | 328.8959 | 179.3284 | 274.4179 | 253.0147 | 243.0594 | 345.4773 | 127.4927 | 116.5674 | 107.2537 | 153.1195 | 152.1943 | 138.1041 | 164.7764 | 96.9699 | 263.1496 | 291.3739 | 230.1642 | 287.1342 | 136.5073 | 136.8658 | 119.0447 | 141.3431 | 79.6268 | 73.1943 | 68.5821 | 123.4326 | 207.8358 | 214.5968 | 216.7163 | 198.8211 | 130.2537 | 128.6716 | 131.4179 | 121.7016 | 217.4927 | 194.1349 | 210.3585 | 192.6268 | 184 | 187.3585 | 188.0447 | 187.2683 | 180.1642 | 246.2544 | 207.9106 | 291.7317 | 156.0594 | 195.5073 | 226.1642 | 252.0154 | 212.4927 | 196.4326 | 178.9106 | 422.5374 | 181.0301 | 181.478 | 130.0147 | 266.4626 | 144.3585 | 193.4773 | 151.7764 | 175.2537 | 225.3284 | 248.3431 | 205.4927 | 153.3732 | 113.2089 | 110.2537 | 112.0447 | 130.3585 | 233.8959 | 280.1496 | 221.478 | 396.7463 | 258.7163 | 212.1342 | 190.6114 | 322.1342 | 170.9252 | 68.8358 | 65.1642 | 170.9252 | 143.0748 | 139.0301 | 123.2984 | 148.7463 | 125.7911 | 134.3585 | 110.6122 | 139.3284 | 89.6268 | 92.6122 | 84.1195 | 96.0894 | 121.2537 | 138.3431 | 86.2537 | 91.1642 | 246.3578 | 269.1642 | 223.1789 | 411.3732 | 182.7911 | 230.5821 | 187.4179 | 358.0748 | 317.9106 | 231.8065 | 207.4479 | 440.0301
2019 | 24.2034 | 162.7699 | 207.9833 | 212.9665 | 193.1632 | 159.2301 | 171.2134 | 151.9163 | 173.0962 | 185.0335 | 153.6234 | 134.8368 | 192.7699 | 254.3138 | 157.6402 | 151.6234 | 272.0167 | 207.7197 | 199.7866 | 159.8535 | 214.9498 | 161.6736 | 381.7866 | 319.1297 | 164.1297 | 238.6527 | 247.1967 | 229.7699 | 360.6904 | 127.59 | 104.3264 | 103.59 | 134.113 | 142.5398 | 165.5439 | 136.2594 | 95.41 | 253.7364 | 302.7866 | 234.2134 | 278.3766 | 125.9665 | 160.0837 | 143.7071 | 129.9665 | 85.4268 | 69.9665 | 87.477 | 110.7699 | 198.9665 | 238.9205 | 215.6067 | 200.59 | 120.1632 | 139.0502 | 108.5063 | 115.1632 | 196.7029 | 209.8201 | 223.2301 | 178.3431 | 167.9331 | 209.8703 | 186.1967 | 181.8033 | 175.3766 | 249.3598 | 234.9038 | 264.2594 | 170.887 | 193.8033 | 211.7364 | 232.8829 | 220.6234 | 245.6276 | 167.3431 | 394.2762 | 180.1799 | 190.6067 | 133.0335 | 262.59 | 166.0669 | 168.523 | 171.4602 | 172.3933 | 222.1799 | 273.9205 | 220.8535 | 162.8368 | 121.0335 | 133.9038 | 123.0502 | 123.1465 | 226.5398 | 271.5398 | 207.3096 | 388.3598 | 275.2803 | 218.6402 | 218.7573 | 305.3431 | 168 | 76.8368 | 75.6402 | 168 | 126.7197 | 134.1632 | 133.0502 | 139.5565 | 120.3766 | 120.7197 | 114.1967 | 123.3264 | 80.1632 | 83.3431 | 88.41 | 92.3933 | 141.6904 | 132.59 | 94.6402 | 83.1632 | 243.6234 | 265.1799 | 250.3305 | 400.7531 | 179.7866 | 232.6067 | 192.6234 | 367.4268 | 327.2301 | 211.2762 | 202.9665 | 413.4728
2020 | 24.2144 | 160.1844 | 193.9498 | 210.1844 | 213.0222 | 174.0502 | 163.028 | 163.3128 | 180.0502 | 166.6091 | 159.8938 | 148.4189 | 178.7654 | 230.0058 | 153.4749 | 166.0502 | 289.3128 | 207.0782 | 188.3967 | 158.1062 | 231.5753 | 161.9218 | 358.1621 | 305.2123 | 179.5753 | 258.5473 | 272.2848 | 221.4749 | 355.3967 | 127.8156 | 112.2346 | 96.4749 | 166.5193 | 141.4469 | 161.7654 | 111.6872 | 104.972 | 272.8379 | 250.6148 | 226.8436 | 291.6815 | 127.2626 | 129.1119 | 138.5811 | 135.3407 | 97.6033 | 69.6313 | 83.8436 | 126.944 | 210.8658 | 246 | 220.8938 | 170.637 | 136.7597 | 144.4469 | 111.972 | 121.1564 | 182.028 | 207.7374 | 217.6593 | 178.2626 | 152.5811 | 203.3967 | 190.0782 | 161.6091 | 183.7877 | 241.2905 | 226.3185 | 276.3909 | 155.6091 | 182.5811 | 208 | 217.8436 | 212.2123 | 236.9778 | 175.4189 | 389.4469 | 180.6313 | 164.8996 | 153.7597 | 235.084 | 152.056 | 180.3128 | 155.4247 | 158.1342 | 203.056 | 251.637 | 166.9053 | 150.3185 | 110.7654 | 113.084 | 115.3967 | 120 | 220.5531 | 281.0502 | 205.2626 | 394.1564 | 245.7432 | 205.5029 | 233.7095 | 345.2288 | 162.2905 | 79 | 88.6033 | 162.2905 | 150.5473 | 137.7095 | 113.9778 | 139.4469 | 101.056 | 124.972 | 113.1844 | 125.5251 | 85.3407 | 81.6313 | 73.5029 | 105.8658 | 116.8214 | 129.5531 | 86.3967 | 82.6313 | 262.944 | 281.1284 | 258.8156 | 391.4749 | 165.1342 | 212.6091 | 208.6815 | 338.8214 | 328.1844 | 221.1284 | 188.7654 | 414.972
2021 | 24.2254 | 150.0649 | 191.8279 | 227.3865 | 206.409 | 176.1721 | 164.3117 | 173.6234 | 171.409 | 162.8279 | 183.3541 | 141.0649 | 176.8279 | 251.1496 | 145.3766 | 152.4414 | 265.6783 | 186.4738 | 187.6559 | 143.7531 | 203.7108 | 156.8604 | 386.1172 | 274.0225 | 149.227 | 262.5162 | 283.7955 | 228.2793 | 360.4514 | 139.591 | 132.3865 | 97.4838 | 131.1198 | 145.3117 | 171.4189 | 136.9775 | 94.5811 | 276.3441 | 263.8379 | 210.9252 | 277.6135 | 136.2793 | 165.7082 | 125.2369 | 137 | 81.7856 | 85.7306 | 73.7207 | 115.2694 | 210.6883 | 260.0748 | 202.1297 | 198.773 | 136.8604 | 139.2045 | 113.1721 | 126.3117 | 185.7955 | 178.0225 | 190.3342 | 184.7955 | 173.182 | 201 | 181.8928 | 184.3217 | 179.3766 | 276.4289 | 213.8928 | 271.8928 | 147.6883 | 206.4937 | 224.5586 | 198.6135 | 209.1721 | 217.0973 | 199.5262 | 376.7531 | 142.0873 | 192.1172 | 149.7207 | 228.1721 | 138.0649 | 170.7531 | 166.7306 | 172.3865 | 188.2369 | 260.7306 | 196.2244 | 155.2793 | 110.4838 | 116.9351 | 98.0973 | 121.6559 | 212.5486 | 265.7856 | 232.4937 | 397.6559 | 252.9027 | 233.4613 | 215.9576 | 332.6459 | 161 | 72.3766 | 76.9252 | 161 | 140.4414 | 128.2369 | 133.0099 | 144.9676 | 109.4189 | 121.2045 | 105.5486 | 127.8279 | 88.6559 | 91.1072 | 79.1072 | 109.8279 | 119.1072 | 125.6883 | 101.3865 | 81.3441 | 273.6234 | 269.2694 | 237.4738 | 384.2045 | 167.7955 | 218.7631 | 211.1721 | 324.5486 | 325.5162 | 231.6234 | 205.0424 | 429.4189
2022 | 24.2364 | 163.124 | 215.5262 | 231 | 177.113 | 177.6804 | 179.2837 | 161.5565 | 157.2369 | 141.1543 | 163.6336 | 151.6033 | 189.6033 | 278.0055 | 178.449 | 137.2369 | 228.0716 | 221.4903 | 193.0413 | 167.0468 | 204.562 | 174.4848 | 365.2727 | 324.2947 | 174.7686 | 281.3251 | 288.3609 | 257.7273 | 347.7163 | 117.6336 | 130.9587 | 111.4435 | 127.8815 | 153.562 | 150.4738 | 136.2782 | 92 | 244.0716 | 266.4793 | 239.9284 | 294.1653 | 127.9174 | 157.3554 | 132.2424 | 122.7163 | 79.6804 | 74.7163 | 69.4793 | 119.562 | 186.4738 | 257.1185 | 206.4022 | 170.7108 | 126.7576 | 142.2011 | 108.7989 | 111.0358 | 188.6804 | 210.6501 | 200.124 | 186 | 153.6336 | 185.876 | 194.2837 | 190 | 188.0826 | 260.4738 | 201.0771 | 281.7631 | 152.0413 | 179.2314 | 217.9174 | 216.8457 | 203.9587 | 210.6391 | 148.5455 | 376.5207 | 160.887 | 167.9119 | 144.6391 | 247.3251 | 143.562 | 189.8457 | 138.9119 | 165.0771 | 190.2011 | 277.4435 | 199.1185 | 132.6336 | 107.6391 | 107.2369 | 105.9229 | 122 | 215.2011 | 249.3967 | 222.0358 | 408.0826 | 275.3251 | 253.4435 | 229.6446 | 357.4077 | 143.3554 | 74.3609 | 76.5207 | 143.3554 | 121.0358 | 136.9229 | 107.7521 | 111.551 | 114.5207 | 127.562 | 111.562 | 132.2011 | 69.6749 | 73.6749 | 81 | 106.6391 | 118.4793 | 128.3609 | 117.6033 | 87.7218 | 250.6336 | 284.4848 | 208.6336 | 362.8347 | 174.8815 | 237.8044 | 202.5978 | 334.7631 | 313.2369 | 231.3196 | 198.9174 | 406.7934
2023 | 24.2474 | 163.4425 | 235.3452 | 222.4749 | 212.068 | 183.1151 | 199.0502 | 150.4749 | 150.7374 | 177.068 | 158.1475 | 141.2123 | 193.705 | 265.8023 | 173.0648 | 146.9352 | 238.1977 | 211.9498 | 172.6872 | 167.7374 | 213.6726 | 178.8525 | 380.4603 | 306.8671 | 169.9174 | 287.5575 | 284.7374 | 224.637 | 364.6078 | 96.8023 | 128.295 | 123.3776 | 131.5575 | 123.4571 | 153.6726 | 137.5575 | 101.3776 | 274.658 | 282.1977 | 236.6224 | 275.8347 | 137.0826 | 127.5722 | 142.5251 | 125.9676 | 80 | 83.9352 | 67.295 | 125.2626 | 198.1977 | 283.2804 | 215.6726 | 175.9352 | 113.0648 | 136.1799 | 128.4603 | 101.1799 | 188.1475 | 213.7374 | 203.8525 | 179.1799 | 153.2626 | 194.9352 | 188.4749 | 158.4571 | 170.3922 | 266.2301 | 213.4927 | 260.9822 | 156.41 | 180.6726 | 214.295 | 219.295 | 199.59 | 229.6078 | 167.8379 | 364.2123 | 154.9174 | 170.5251 | 144 | 236.5073 | 141.59 | 180.3598 | 144.0826 | 168.9676 | 212.3128 | 248.4571 | 195.4425 | 134.8201 | 111.2626 | 125.4603 | 113.1151 | 121.1475 | 221.1151 | 281.1004 | 236.9027 | 386.9822 | 250.8671 | 247.4749 | 209.9822 | 374.0826 | 143.41 | 86.9352 | 77.8525 | 143.41 | 119.705 | 154.3452 | 123.3128 | 140.8055 | 106.4749 | 122.1799 | 111.295 | 133 | 79.6402 | 68.295 | 81 | 105.1475 | 111.1799 | 118.7699 | 92.7196 | 77.0648 | 248.5575 | 264.9822 | 215.0826 | 392.2479 | 179.41 | 219.6872 | 209.5251 | 328.4749 | 326.3452 | 256.5753 | 186.7699 | 365.342
2024 | 24.2584 | 176.8368 | 191.3007 | 208.028 | 185.2727 | 145.9487 | 183.8392 | 162.8368 | 161.2424 | 168.4336 | 170.972 | 128.6224 | 204.3776 | 241.3799 | 181.3776 | 149.8648 | 255.7016 | 196.8928 | 179.3776 | 153.1632 | 201.1632 | 174.676 | 396.1072 | 295.0816 | 161.9464 | 237.8415 | 289.1888 | 264.6993 | 371.4592 | 102.648 | 152.2145 | 128.4592 | 147.5664 | 127.5128 | 126.4615 | 121.5688 | 88.2984 | 243.8135 | 271.1632 | 216.8392 | 313.5105 | 132.9464 | 143.7553 | 130.1632 | 119.2168 | 73.0816 | 90.324 | 65.2704 | 105.2447 | 213.1072 | 270.704 | 207.4872 | 177.1352 | 131.7553 | 169.5921 | 126.8112 | 119.8904 | 190.5944 | 207.8112 | 223.0256 | 158.9744 | 143.6224 | 203.0536 | 208.6201 | 170.296 | 183.4312 | 250.704 | 222.0536 | 284.6737 | 135.3799 | 179.4056 | 204.4872 | 244.9441 | 191.2168 | 213.9744 | 167.8112 | 361.1352 | 165.972 | 163.352 | 137.0816 | 239.1888 | 140.1352 | 178 | 154.648 | 175.1888 | 203.8928 | 262.0256 | 200.1888 | 135.1352 | 118.9184 | 126.4056 | 88.9207 | 122.7296 | 212.4872 | 274.028 | 220.1096 | 403.7553 | 261.5664 | 245.1352 | 199.0816 | 365.6224 | 150.9184 | 79.4872 | 69.352 | 150.9184 | 113.9464 | 150.9464 | 126.1352 | 144.4056 | 92.8928 | 117.5408 | 102.352 | 129.5408 | 76.8112 | 74.9184 | 82.7296 | 105.8648 | 121.2424 | 118.7296 | 98.3776 | 83.648 | 255.0536 | 243.704 | 230.8368 | 359.0839 | 176.5408 | 228.1072 | 184.1911 | 321.8112 | 292.6783 | 242.8392 | 190.1888 | 410.8881
2025 | 24.2694 | 165.8436 | 225.2234 | 203.3687 | 194.9106 | 129.4749 | 199.419 | 177.2793 | 165.6313 | 157.2291 | 168.6145 | 140.1564 | 168.285 | 245.0167 | 171.5978 | 154.3855 | 234.3185 | 180.9665 | 194.1564 | 153.6313 | 189.352 | 172.2458 | 387.4749 | 336.9776 | 148.7207 | 279.1172 | 301.4022 | 241.3017 | 363.2291 | 119.7877 | 127.933 | 114.0894 | 135.9665 | 138.648 | 152.6983 | 131.2793 | 97.4022 | 229.2291 | 289.1731 | 216.6313 | 318.2458 | 145.1564 | 142.6145 | 138.5251 | 120.6313 | 86.9106 | 72.581 | 80.7877 | 108.1397 | 197.4581 | 275.0167 | 198.9833 | 174.3687 | 122.7207 | 150.4414 | 114.5978 | 122.1229 | 160.3017 | 218.4022 | 193.5476 | 168.2793 | 158.6648 | 178.5643 | 203.2291 | 171.2458 | 195.648 | 248 | 196.6872 | 260.933 | 130.2458 | 180.7542 | 227.5586 | 213.9163 | 218.067 | 203.9833 | 186.296 | 373.2793 | 177.648 | 173.4022 | 143.0167 | 238.2458 | 135.6145 | 186.7709 | 155.1229 | 150.5643 | 187.0894 | 221.1453 | 192.2291 | 131.4916 | 105.9665 | 107.581 | 94.648 | 121.2458 | 222.4022 | 231.6537 | 214.3687 | 385.9498 | 230.6705 | 246.7542 | 212.9106 | 354.352 | 152.8771 | 94.6648 | 65.3687 | 152.8771 | 113 | 149.1229 | 118.1062 | 126.4581 | 104.1564 | 118.7542 | 104.5084 | 110.581 | 81.2626 | 76 | 74.2291 | 97.2291 | 115.1062 | 139.1731 | 114.0335 | 86.7542 | 217.4079 | 255.0335 | 233 | 372.296 | 177.7542 | 205.4414 | 183.5084 | 308.7207 | 291.3855 | 206.6705 | 181.352 | 404.9665
2026 | 24.2804 | 172.8938 | 232.7788 | 203 | 169.4291 | 139.562 | 206.4469 | 147.8716 | 166 | 144.438 | 177.7832 | 164.2346 | 154.9955 | 247.7788 | 153.1017 | 146.9955 | 219.438 | 172.7743 | 183.5486 | 171.7877 | 184.4425 | 147.9866 | 345.0883 | 326.9911 | 175.4603 | 213.9599 | 287.8805 | 235.2212 | 346.8805 | 126.4469 | 130.2257 | 120.0045 | 148.2301 | 141.7788 | 132.9866 | 125.8849 | 100.7788 | 260.9072 | 278.6592 | 251.686 | 303.7699 | 110.5352 | 152.6726 | 121.3229 | 118.3318 | 80.1062 | 70.8894 | 78.5531 | 100.9955 | 207.4514 | 247.5397 | 240.6904 | 177.5575 | 125.4469 | 144.3318 | 121.8938 | 116.6637 | 168.4514 | 220 | 189.8894 | 174.4469 | 142.3229 | 175 | 179.7654 | 146.0972 | 194.3318 | 276.4603 | 192.1106 | 233.876 | 155.7921 | 174.7743 | 217.6592 | 197.438 | 225.5575 | 209.2257 | 196.1151 | 375 | 160.3229 | 159.8805 | 168.0134 | 234.4425 | 132.3318 | 165.7654 | 155 | 168.3452 | 175.2168 | 207.8849 | 210.5665 | 114.9911 | 110.2257 | 122.7877 | 101.3363 | 138.7877 | 209.7699 | 264.2435 | 198.8805 | 383.8894 | 218.8849 | 217.6503 | 209.6637 | 379.6815 | 168.1195 | 81.8805 | 69.4469 | 168.1195 | 119.2257 | 128.5442 | 132.1195 | 122.2212 | 89.1017 | 122.5575 | 105.8894 | 116.0045 | 68.6592 | 60.8805 | 73 | 104.8938 | 98.8805 | 124.2123 | 102.6592 | 83.4425 | 224.4514 | 207.1945 | 231.2212 | 355.4335 | 174.4425 | 195.7743 | 229.3586 | 314.1151 | 296.4469 | 228.6815 | 216.4648 | 378.0972
2027 | 24.2914 | 158.6715 | 230.295 | 215.6235 | 193.0503 | 151.8201 | 175.4414 | 164.7385 | 170.5084 | 152.0168 | 161.8682 | 127.3263 | 147.6883 | 216.4414 | 168.1318 | 156.8201 | 240.5419 | 177.4101 | 163.9665 | 157.7698 | 183.0983 | 163.9352 | 367.0503 | 333.1151 | 150.1464 | 227.5419 | 279.6883 | 214.2615 | 345 | 108.0648 | 125.5899 | 121 | 146.3933 | 143.8034 | 148.0335 | 135.8201 | 80.2615 | 249.6715 | 283.3117 | 234.3598 | 291.1799 | 119.5251 | 144.9832 | 141.5419 | 127.9184 | 87.1151 | 71.9017 | 73.4916 | 126.1486 | 200.8849 | 269.2469 | 237.8849 | 170.7866 | 116.9832 | 155.7218 | 113.9832 | 130.4268 | 166.3933 | 220.9017 | 215.2469 | 177.705 | 149.0168 | 185.8201 | 191.4268 | 154.7218 | 208.4268 | 240.3263 | 201.9184 | 233.705 | 156.295 | 177.6067 | 199.7698 | 218.5419 | 211.5732 | 198.2782 | 191.5899 | 380.4101 | 156.1966 | 167.9184 | 157.4749 | 262.8536 | 132.9017 | 152.1799 | 145.0816 | 158.3765 | 186.6235 | 190.7698 | 181.4414 | 134.6402 | 87.5564 | 125.9017 | 109.2134 | 128.3765 | 242.2637 | 242.8514 | 215.0335 | 403.8369 | 255.8704 | 252.772 | 236.0503 | 377.5899 | 162.7866 | 73.6883 | 75.4101 | 162.7866 | 109.1799 | 146.7385 | 120.4749 | 130.1151 | 110.4436 | 116.6883 | 96.9832 | 113.3933 | 80.5251 | 85.1486 | 79.3117 | 101.4916 | 98.8034 | 115.6883 | 121.7385 | 79.3933 | 259.362 | 219.9352 | 225.5899 | 364.7218 | 147.8514 | 190.4916 | 194.4246 | 365.4938 | 310.5251 | 241.9184 | 192.1464 | 372.295
2028 | 24.3024 | 162.4838 | 245.5374 | 184.0972 | 211.5374 | 156.6559 | 224.0961 | 172.4838 | 156.3765 | 126.495 | 170.0536 | 158.6447 | 145.1721 | 255.0424 | 189.1933 | 160.7419 | 230.2045 | 165.2045 | 162.914 | 169.7095 | 181.1721 | 151.3765 | 391.9352 | 306.581 | 160.7095 | 239.1397 | 299.1073 | 215.6559 | 355.0536 | 108.7419 | 129.5698 | 113.6883 | 129.5486 | 133.0324 | 138.1184 | 145.2257 | 93.5374 | 235.2045 | 246.5274 | 242.0536 | 304.6235 | 123.7419 | 147.6559 | 139.4302 | 126.2581 | 74.2905 | 72 | 68.4302 | 114.3765 | 205.4838 | 221.7319 | 209.581 | 186.4514 | 120.5698 | 166.1397 | 130.3654 | 117.3765 | 163.2581 | 224.6559 | 186.9251 | 174.3441 | 149.086 | 216.2469 | 203.9676 | 139.5486 | 202.6883 | 259.7631 | 212.1397 | 224.8603 | 140.4626 | 160.6346 | 230.9028 | 208.2045 | 199.9464 | 219.8492 | 197.3978 | 335.3017 | 131.3229 | 172.6559 | 139.5486 | 234.0112 | 116.5486 | 144.6022 | 166.8492 | 157 | 158.7531 | 198.1397 | 192.6235 | 146.1397 | 94.1397 | 97.667 | 101.7743 | 113.2905 | 221.3229 | 245.4838 | 199.6346 | 400.5162 | 227.0972 | 256.086 | 205.1832 | 388.8816 | 168.3978 | 65.6883 | 65.0324 | 168.3978 | 126.2793 | 153.5698 | 124.4838 | 121.8603 | 94.7207 | 132.4514 | 114.2793 | 90.1508 | 84.7419 | 75.2045 | 63.5486 | 90.0324 | 106.3117 | 130.5374 | 111.2045 | 80.8279 | 227.3553 | 244.8492 | 211.2905 | 404.3866 | 163.2793 | 192.7419 | 221.0749 | 341.7531 | 278.1832 | 238.4302 | 207.2793 | 393.0212
2029 | 24.3134 | 184.3039 | 279.419 | 211.5665 | 208.3687 | 145.8849 | 191.0234 | 152.6223 | 138.3274 | 129.5575 | 169.1475 | 157.3687 | 144.0737 | 227.5072 | 196.5575 | 190.6402 | 222.5162 | 184.3777 | 145.4011 | 160.8112 | 191.1888 | 130.5486 | 428.2715 | 323.4514 | 161.0737 | 201.0972 | 236.162 | 234.5251 | 318.0234 | 115.4838 | 112.4011 | 111.1475 | 126.1475 | 129.2212 | 123.1061 | 129.3274 | 82.9587 | 264.5665 | 286.5341 | 259.6726 | 305.0737 | 129.5575 | 148.9263 | 148.2626 | 108.4011 | 65.5899 | 81.2626 | 84.6726 | 120.4101 | 164.3184 | 244.7877 | 197.7374 | 184.295 | 133.9676 | 151.2536 | 134.7788 | 110.4425 | 153.7374 | 217.5899 | 219.1978 | 186.0413 | 127.6961 | 210.6637 | 185.5486 | 148.1888 | 202.9263 | 241.6223 | 215.7788 | 257.3453 | 130.6637 | 184.9352 | 234.9263 | 215.3363 | 232.3453 | 213.6637 | 216.5251 | 347.6726 | 132.705 | 162.8112 | 141.705 | 232.8525 | 125.1888 | 138.4425 | 152.3274 | 143.1061 | 150.4425 | 193.4425 | 193.0737 | 144.2212 | 95.9263 | 126.4928 | 105.6313 | 110.1475 | 253.2715 | 245.0737 | 211.8939 | 374.9911 | 237.8939 | 293.9766 | 231.6402 | 352.9497 | 157.8849 | 78.8939 | 76.9676 | 157.8849 | 103.9173 | 145.6637 | 123.1475 | 147.8615 | 95.7788 | 113.6223 | 103.9587 | 105.5989 | 73.8849 | 82.3363 | 61.0737 | 81.5899 | 91.2536 | 108.8436 | 119.2626 | 79.1475 | 247.1564 | 226.6223 | 246.124 | 385.7698 | 162.2212 | 190.2212 | 185.0972 | 312.1385 | 264.8112 | 225.9587 | 191.4011 | 371.8436
2030 | 24.3244 | 162.5363 | 226.6256 | 211.1844 | 180.7821 | 131.8603 | 213.3408 | 155.6927 | 137.9385 | 145.9553 | 164.3073 | 159.8156 | 159.9553 | 213.7374 | 195.1229 | 166.7206 | 241.7095 | 170.0447 | 165.5866 | 151.5531 | 184.4916 | 156.2179 | 365.3016 | 298.7206 | 160.0615 | 202.6927 | 266.6648 | 207.8435 | 326.2626 | 116.9385 | 107.2458 | 127.8939 | 134.4469 | 144.9553 | 128.5698 | 137.3855 | 103.5866 | 229.4581 | 291.8771 | 228.1508 | 310.6313 | 145.0168 | 159.324 | 132.1061 | 113.5698 | 78.1397 | 82 | 95.3855 | 118.1844 | 172.2626 | 272.3408 | 212.9553 | 165.229 | 124.676 | 140.6145 | 125.6145 | 110.9385 | 164.2626 | 219.8156 | 203.229 | 187 | 104.4134 | 186.5363 | 178.3687 | 145.2458 | 216.1397 | 220.2905 | 218.8156 | 241.229 | 134.6927 | 187 | 234.0615 | 193.4748 | 212.4748 | 199.8603 | 172.9497 | 376.2179 | 163.0335 | 149.7989 | 161.7095 | 240.5084 | 126 | 128.6145 | 146.3073 | 126.9832 | 168.771 | 189.2458 | 199.5698 | 131.7989 | 110.0782 | 122.4302 | 96.6145 | 110.9385 | 227.8435 | 248.7542 | 194.229 | 355.1676 | 242.7542 | 279.1676 | 248.0782 | 344.3687 | 144.7989 | 80 | 84.5698 | 144.7989 | 112.324 | 136.5531 | 113.6145 | 134.0447 | 111.9553 | 105.4302 | 98.3073 | 119.2011 | 73.9385 | 83.9385 | 74.1397 | 98.8324 | 107.8324 | 112.6313 | 100.2905 | 63.9832 | 245.2458 | 230.6313 | 246.1844 | 365.229 | 149.7989 | 190 | 202.6481 | 336.2794 | 300.6034 | 201.5363 | 190.9385 | 376.5698
2031 | 24.3354 | 175.2626 | 225.8525 | 203.3933 | 194.2134 | 150.0168 | 203.5899 | 175.9676 | 159.8693 | 137.4916 | 154.4916 | 163.8034 | 148.6391 | 210.1475 | 205.4592 | 181.1643 | 234.4424 | 182.3118 | 153.6882 | 175.7218 | 172.5899 | 136.1307 | 361 | 337.8861 | 169.5084 | 215.3609 | 230.0156 | 240.2302 | 331.7542 | 116.0492 | 113.6559 | 112.8357 | 146.4101 | 136.4916 | 139.4592 | 129.4424 | 85.0324 | 233.6559 | 323.3777 | 227.9017 | 299.5899 | 117.4748 | 153.3441 | 114.8357 | 114 | 79.9508 | 65.8357 | 85.5408 | 118 | 175.8525 | 227.4089 | 208.295 | 162.0983 | 119.2458 | 133.3441 | 128.8034 | 141.4269 | 146.9341 | 214.295 | 201.0492 | 180.3441 | 123.9185 | 179.295 | 182.7542 | 151.6559 | 201.7866 | 215.1966 | 198.0815 | 223.8357 | 150.2134 | 188.9017 | 212.1307 | 211.0168 | 187.229 | 183.7866 | 164.295 | 351.3765 | 138.3765 | 140.4424 | 161.0983 | 203.9173 | 124.0983 | 118.4916 | 130.7866 | 146.9185 | 176.6559 | 169.9832 | 176.229 | 131 | 106.2458 | 103.9341 | 127.3777 | 125.2626 | 225.0492 | 235.6882 | 225.3285 | 330.229 | 237.295 | 275.1475 | 217.6223 | 334.4916 | 153.5084 | 78.0983 | 81.1966 | 153.5084 | 109.1966 | 133.1475 | 115.8525 | 125.3933 | 108.2458 | 98.3441 | 100.8525 | 129.5084 | 69.2458 | 59.2782 | 99.7218 | 97.1475 | 116.6067 | 103.4916 | 101.8525 | 82.9676 | 226.9341 | 237.6559 | 234.5899 | 403.9353 | 142.3441 | 173.8357 | 190.6882 | 326.5899 | 282.0815 | 219.9676 | 191.9508 | 381.7542
2032 | 24.3464 | 176 | 263.5621 | 227.0783 | 203.6682 | 124.0323 | 165.4379 | 171.2212 | 146.553 | 156.2626 | 152.0737 | 150.5162 | 174.0045 | 216.7419 | 188.6637 | 152.143 | 205.1061 | 160.848 | 167.447 | 166.4056 | 180.6682 | 119.5899 | 362.9263 | 331.3318 | 155.553 | 200.5899 | 265.5621 | 258.3732 | 335.8525 | 100.5899 | 101.4793 | 114.8894 | 125.8111 | 103.2536 | 144.8157 | 123.2212 | 82.0737 | 212.8111 | 278.7697 | 232.8157 | 272.0323 | 147.7833 | 141.4424 | 140.0045 | 120.7419 | 70.3687 | 89.0783 | 74.4056 | 107.4056 | 204.8939 | 224.0369 | 208.9631 | 157.1843 | 133.447 | 135.8894 | 140.5576 | 141.0737 | 136.3687 | 210.1475 | 217.3732 | 156.8849 | 123.0737 | 167.4424 | 151.2167 | 145.2581 | 186.553 | 215 | 210.4838 | 236.4838 | 153.8894 | 169.7374 | 231.2257 | 192.7374 | 170.5899 | 197.447 | 182.2995 | 348.0737 | 149.5207 | 151.5576 | 176.4101 | 219.3363 | 133.6313 | 146.8939 | 144.447 | 136.4424 | 183.7419 | 188.2626 | 171.1475 | 156.0414 | 89.6268 | 108.7788 | 97.2167 | 129.8525 | 263.5252 | 241.7419 | 230.8525 | 369.4515 | 245.6682 | 277.8894 | 214.0737 | 334.9631 | 141.4793 | 84.7419 | 77.1475 | 141.4793 | 109.9631 | 130.1106 | 108.295 | 128.8525 | 109.9263 | 124.0045 | 115.447 | 106.8849 | 68.0369 | 61.8525 | 80.7743 | 93.1475 | 111.2212 | 83.7374 | 82.7374 | 74.3687 | 269.3409 | 209.1061 | 214.7374 | 399.2581 | 171.857 | 188.4101 | 205.4101 | 358.7464 | 290.6313 | 220.0369 | 208.3732 | 371.4056
2033 | 24.3574 | 166.2458 | 233.7865 | 233.8525 | 186.4424 | 154.2135 | 223.5007 | 178.8034 | 146.9754 | 146.2704 | 148.0983 | 150 | 135.0077 | 202.3687 | 203.6067 | 171.4839 | 205.9508 | 163.9017 | 145.5653 | 189.4101 | 175.1475 | 114.1229 | 383.4839 | 330.0246 | 142.3195 | 206.828 | 225.0569 | 215.106 | 339.9017 | 110.7296 | 112.7051 | 125.7296 | 142.5576 | 128.3364 | 152.8034 | 122.0246 | 88.828 | 212.9754 | 271.1475 | 229.0983 | 287.5822 | 128.5161 | 154.6559 | 135.1475 | 131.7296 | 71.9508 | 77.3195 | 72.0492 | 111.8771 | 182.5899 | 239.6067 | 224.6067 | 157.9754 | 130.0983 | 150.6313 | 120.5161 | 155.6313 | 140.8771 | 216.828 | 228.7296 | 165.7542 | 133.7296 | 162.1229 | 189.9923 | 151.828 | 225.9923 | 227.6805 | 194.4178 | 217.4916 | 147.172 | 148.5161 | 230.0492 | 202.7296 | 193.4101 | 189.2212 | 179.0983 | 354.828 | 163.6559 | 147.1229 | 161.3933 | 218.0492 | 131.0737 | 132.3933 | 136.2212 | 161.361 | 159.6145 | 173.3933 | 212.9431 | 142.3687 | 92.9017 | 113.8771 | 111.6067 | 118.2949 | 248.4178 | 245.9017 | 223.1966 | 341.7374 | 214.7865 | 289.7051 | 218.8771 | 343.7788 | 135.1475 | 68.4178 | 76.0246 | 135.1475 | 134.3855 | 127.0737 | 117.7542 | 119.2458 | 110 | 109.3933 | 114.0492 | 114.7788 | 64.0983 | 73.7051 | 70.2458 | 80.3195 | 111 | 110.3118 | 98.5822 | 77.9017 | 225.1552 | 199.2212 | 223.7542 | 360.9586 | 137.8848 | 189 | 203.0737 | 360 | 302.7051 | 221.9508 | 192.4178 | 384.6559
2034 | 24.3684 | 163.0369 | 267.5699 | 218.1966 | 183.0369 | 165.8648 | 207.2212 | 155.2949 | 174.656 | 148.9631 | 141.086 | 159.8771 | 163.6314 | 199.0369 | 186.2212 | 190.7665 | 201.0614 | 169.9263 | 164.7543 | 182.0983 | 189.8157 | 144.6191 | 366.2212 | 297.4055 | 166.6928 | 213.914 | 231.9017 | 270.2996 | 318.2703 | 116.9263 | 110.0369 | 106.2457 | 130.1597 | 140.8526 | 129.2949 | 117.0614 | 94.9263 | 207.0737 | 295.6928 | 203.3195 | 262.3195 | 128 | 141.172 | 137.9631 | 134.9631 | 72.9877 | 75.0246 | 50.2703 | 118.914 | 176.0737 | 237.0369 | 183.5161 | 154.0491 | 119.1352 | 124.3318 | 131.8526 | 120.4423 | 170.6314 | 207.1229 | 199.3686 | 162.0491 | 146.8403 | 157.0614 | 172.2335 | 123.3563 | 190.4546 | 228 | 198.9386 | 276.2628 | 154.9017 | 168.742 | 269.5085 | 209.914 | 190.0491 | 176.1597 | 164.1843 | 352.0369 | 154.1229 | 159.8403 | 164.9509 | 223.9263 | 123.0983 | 159.656 | 136 | 139.2826 | 169.8648 | 155.2212 | 181.4055 | 134.0983 | 112.7543 | 110.0491 | 105.086 | 126.8894 | 213.4301 | 235.1352 | 234.8526 | 347.914 | 223.8771 | 276.172 | 232.828 | 324.2457 | 150.8034 | 74.914 | 73.0369 | 150.8034 | 105.3686 | 124.0369 | 103.1843 | 101.2212 | 115.9263 | 107.0246 | 106.0983 | 89.3195 | 64 | 82.8894 | 64.0737 | 87.9017 | 91.2457 | 99.1474 | 98.0123 | 72.0737 | 220.0491 | 218.7543 | 239.8034 | 324.4423 | 162.6805 | 218.6314 | 204.9754 | 307.6512 | 315.8403 | 211.1352 | 194.9631 | 401.7911
2035 | 24.3794 | 179 | 230 | 238.9999 | 188 | 154 | 166 | 159 | 157 | 142.0001 | 144.9999 | 142 | 161 | 198 | 194 | 160.0002 | 233.9999 | 178.9999 | 157 | 162 | 174 | 145 | 336.9999 | 339.9999 | 159.0001 | 222 | 224.0001 | 229.0001 | 310.0001 | 127 | 113 | 113 | 147 | 151.9999 | 145.9999 | 145 | 84.0001 | 217.0002 | 270.9999 | 242 | 286 | 118 | 149.9999 | 142 | 118 | 91 | 79.9999 | 68 | 108 | 173 | 212.0002 | 224 | 179.9999 | 151.9999 | 136 | 119.0001 | 104.0001 | 156.0001 | 208 | 196 | 155 | 138.9999 | 141 | 175.0002 | 144.0001 | 195.9999 | 216.9998 | 223 | 259.9999 | 149.9999 | 172 | 215 | 174.0001 | 191.0001 | 197.9999 | 174.0001 | 380 | 134 | 138.0001 | 178 | 226.9999 | 133 | 121.0001 | 172 | 152 | 167.9999 | 199.9998 | 200 | 137 | 108 | 107 | 98.0001 | 136.9999 | 225.9999 | 230 | 226.0002 | 333.0001 | 193.0002 | 279.9999 | 227.9999 | 335.0001 | 143.0001 | 78 | 82 | 143.0001 | 101.0001 | 103 | 105 | 133 | 102.9999 | 103 | 113 | 102 | 76 | 66 | 93.9999 | 94 | 106.9999 | 75.0001 | 113 | 96.9999 | 248 | 217 | 219 | 426.9998 | 141.0001 | 186.0001 | 156.0002 | 299.0001 | 331.9998 | 184.0002 | 208.0001 | 382
2036 | 24.3904 | 168.7541 | 221.2582 | 215.0738 | 174.9631 | 153.8525 | 161.0984 | 168.1229 | 168.8525 | 156.0861 | 126.8525 | 149.8525 | 169.0492 | 192.0984 | 200.8279 | 205.4713 | 207.2951 | 162.7787 | 164.6803 | 157.8525 | 169.2951 | 136.0861 | 318.2705 | 310.0369 | 174.8402 | 225.0984 | 242.9508 | 262.6926 | 325.0123 | 113.791 | 112.9385 | 112.9508 | 148.9631 | 131.9754 | 124.9877 | 146.9139 | 110.4467 | 270.3361 | 242.6189 | 243.2828 | 285.1721 | 116.0369 | 133.8893 | 136.1229 | 116.1844 | 78.9016 | 63.1352 | 74.9754 | 108.9877 | 177.0984 | 270.0656 | 216.0738 | 144.3074 | 122.0861 | 136.8402 | 143.6557 | 139.7049 | 174.8771 | 198.9385 | 191.4426 | 157.9631 | 122.9385 | 146.2951 | 220.4959 | 166.5943 | 170.1721 | 170.4795 | 210.4426 | 238.9754 | 123.2705 | 171.7172 | 213.8893 | 199.1229 | 218.7664 | 180.1229 | 209.5697 | 367.9508 | 138.0861 | 153.7787 | 165.1352 | 211.3074 | 127.4795 | 139.8525 | 161.9016 | 155.8279 | 151.0738 | 149.1229 | 198.6557 | 125.9262 | 98.2213 | 107.0615 | 127.8033 | 114.8525 | 208.1352 | 229.3565 | 267.2992 | 366.8525 | 239.7664 | 263.2828 | 192.4426 | 355.1352 | 157.8279 | 67.0615 | 79.0123 | 157.8279 | 132.8156 | 108.8648 | 115.9262 | 121.9877 | 85.0369 | 103.0492 | 116.8402 | 112.0738 | 76.9877 | 61.1475 | 64.3197 | 94.9385 | 81.4426 | 98.9631 | 99.8033 | 76.1352 | 258.5574 | 206.3688 | 222.0246 | 379.2582 | 164.8033 | 205.8771 | 211.5451 | 331.0246 | 280.627 | 227.0861 | 221.582 | 370.9139
2037 | 24.4014 | 149.295 | 242.1475 | 221.8851 | 172.1967 | 142.0246 | 168.8033 | 177.6066 | 157.0492 | 163.0246 | 115.4671 | 138.7376 | 172.4099 | 199.6066 | 187.2459 | 163.1475 | 230.705 | 145.5409 | 138.8033 | 145.7541 | 192.4837 | 143.0246 | 339.582 | 312.6066 | 161.1887 | 232.0903 | 238.4345 | 238.1475 | 325.4345 | 97.3442 | 108.418 | 109 | 145.6066 | 130.0492 | 124.4671 | 140.0983 | 66.9588 | 217.3196 | 212.295 | 265.7787 | 298.4099 | 118.705 | 125.4671 | 144.9428 | 130.7787 | 71.0738 | 74.1475 | 73.0246 | 108.418 | 184.7787 | 196.0818 | 221.9262 | 168.9017 | 128.7787 | 123.8525 | 115.8279 | 116.1967 | 164.7296 | 193.8525 | 226.6558 | 155.0246 | 117.7787 | 169.0903 | 180.4917 | 134.2459 | 183.9017 | 208.4837 | 245.1149 | 236.582 | 145.2213 | 149.295 | 205.7622 | 208.6066 | 200 | 190.1229 | 175.1721 | 363.9508 | 145.0246 | 136.3934 | 175.1641 | 235.6066 | 165.2378 | 128.2213 | 153.705 | 141.9262 | 156.9017 | 159.5163 | 171.0983 | 120.5409 | 115.582 | 111.8279 | 112.3196 | 102.8525 | 219.0246 | 257.0903 | 211.3934 | 355.5901 | 222.5981 | 286.295 | 228.4917 | 364.7215 | 143.9754 | 72.1475 | 80.0492 | 143.9754 | 118.0246 | 98.2459 | 109.9017 | 121 | 88.4671 | 107.1967 | 104.4917 | 117.6558 | 75.9017 | 72.9262 | 89.9017 | 90.2704 | 116.9508 | 96.1475 | 84.4425 | 86.9508 | 223.4671 | 235.6066 | 223.5575 | 399.705 | 148.582 | 195.5329 | 174.9754 | 332.5083 | 331.3934 | 233.8525 | 187.9508 | 364.713
2038 | 24.4124 | 162.2538 | 247.8156 | 255.9306 | 180 | 142.8156 | 161.5163 | 162.295 | 158.8894 | 162.6356 | 134.0738 | 167.4837 | 149.9957 | 183.7419 | 196.8156 | 170.4751 | 218.6681 | 166.2993 | 131.7744 | 136.7007 | 172.4794 | 143.705 | 324.2907 | 298.3644 | 130.5488 | 195.2993 | 216.5532 | 244.885 | 304.4382 | 110.41 | 124.1887 | 109.4056 | 130.3688 | 132.7375 | 142.9631 | 143.5944 | 103.9674 | 230.3688 | 225.1801 | 257.6269 | 274.7419 | 107.6269 | 143.2256 | 103.5163 | 122.59 | 74.3688 | 79.3731 | 73.9631 | 125 | 176.3319 | 237.2668 | 219.885 | 164.9631 | 120.3688 | 118.59 | 110.0694 | 124 | 154 | 187.9631 | 213.9957 | 155.6312 | 110.1801 | 133.1106 | 199.2993 | 144.5163 | 179.8894 | 187.6312 | 209.3731 | 220.6638 | 153.3731 | 160.5944 | 236.0738 | 193.295 | 199.5944 | 194.5575 | 182.1844 | 360.488 | 146.885 | 152.2213 | 143.0694 | 219.8894 | 134.7419 | 136.5575 | 141.7787 | 139.1844 | 153 | 180.0369 | 175.4794 | 142.0738 | 99.6638 | 105 | 125.1106 | 97.9957 | 220.5163 | 221.1844 | 227.4056 | 378.2993 | 283.013 | 296.8568 | 246.8568 | 315.2538 | 143.1106 | 77.5575 | 82.2581 | 143.1106 | 118.8525 | 109.4013 | 106.2213 | 120.9262 | 106.2256 | 114.2993 | 123.7787 | 104.3319 | 71.2993 | 70.2213 | 85.5206 | 100.6312 | 114.0412 | 102.0738 | 101.8156 | 85.0369 | 241.7419 | 219.8156 | 207.8438 | 388.0738 | 132.4794 | 176.8894 | 174.7007 | 313.9957 | 345.2299 | 227.2993 | 186.4056 | 391.8568
2039 | 24.4234 | 193.7708 | 240.9841 | 226.9183 | 180.0983 | 137.7542 | 174.8525 | 170 | 156.6884 | 127.7375 | 134.9183 | 153.9017 | 175.5083 | 178.2292 | 190.6233 | 206.6891 | 211.0817 | 149.5734 | 151.41 | 154.41 | 184.8033 | 136.4425 | 357.0166 | 334.6884 | 169.82 | 178.7209 | 231.1967 | 267.0658 | 340.82 | 95.7375 | 103.885 | 120.4917 | 140 | 150.8691 | 141.9017 | 132.705 | 77.4425 | 239.6558 | 257.6717 | 272.525 | 269.2784 | 124.4917 | 122.1641 | 118.5242 | 137.0166 | 83.8033 | 64.0817 | 73.1967 | 124.4592 | 185.4917 | 191.9508 | 241.4758 | 163.6558 | 130.295 | 133.6558 | 137.5575 | 123.9508 | 155.475 | 186.41 | 238.525 | 147.77 | 141.59 | 136.1475 | 179.7708 | 156.4266 | 177 | 179.0325 | 193.3933 | 239.2292 | 138.2292 | 149.7542 | 236.7708 | 201 | 189.8359 | 183.59 | 186.2625 | 322.3275 | 169.5575 | 157.3608 | 171.1475 | 216.5083 | 127.6067 | 125.4425 | 136.7375 | 144.9342 | 153.6884 | 180.0658 | 187.5083 | 144.0983 | 116.6067 | 105.7375 | 126.6233 | 123.9508 | 234.3933 | 226.0983 | 237.2625 | 360.6884 | 207.8026 | 266.6067 | 216.115 | 348.2458 | 146.3933 | 66.59 | 88.2625 | 146.3933 | 113.7708 | 143.4433 | 112.0983 | 119.2458 | 87.1309 | 96.3442 | 117.7542 | 113.2458 | 54.3767 | 75.5083 | 72.9017 | 90.7542 | 89.3933 | 103.3116 | 97.3442 | 85.2625 | 235.5408 | 214.9508 | 254.5741 | 390 | 145.3933 | 174.9342 | 192.5083 | 340.0492 | 299.3933 | 208.0166 | 197.59 | 363.2784
2040 | 24.4344 | 168.8937 | 203.3521 | 207.2906 | 182.5531 | 133.1844 | 170.8323 | 168.9552 | 168.0333 | 140.8937 | 114.676 | 151.4469 | 164.9552 | 202.2458 | 165.2292 | 162.7375 | 230.7094 | 179.0167 | 139.7542 | 143.4302 | 182.2906 | 145.9833 | 339.7823 | 346.9552 | 148.2906 | 210.8937 | 233.5864 | 247.7094 | 317.4469 | 110.676 | 121.3073 | 129.2625 | 138.4635 | 128.1396 | 139.6927 | 125.7708 | 86.3688 | 231.2791 | 286.6198 | 243.8156 | 294.6146 | 133.5698 | 107.5979 | 146.4021 | 118.0615 | 80.4302 | 84.2625 | 76.7542 | 113.5083 | 193.0333 | 191.9833 | 212.5531 | 157.6146 | 136.3688 | 127.3073 | 129.3073 | 124.2906 | 182.2791 | 175.4302 | 209.4469 | 181.5698 | 151.6479 | 139.2458 | 157.8438 | 128.3354 | 176.4469 | 199.0615 | 201.3688 | 261.3406 | 159.9719 | 144.5698 | 212.5698 | 200.0781 | 205.4469 | 193.8323 | 173.6594 | 347.6927 | 160.7542 | 145.1844 | 172.8323 | 207.8604 | 120.8604 | 134.9833 | 150.3854 | 161.8937 | 168.1677 | 162.7375 | 178.3688 | 145.0781 | 108.324 | 119.6927 | 100.676 | 122.8771 | 239.726 | 227.6927 | 222.3854 | 375.2292 | 260.1562 | 256.8489 | 200.8438 | 351.4021 | 153.324 | 76.6479 | 74.6146 | 153.324 | 92.7657 | 94.9833 | 113.7542 | 123.8156 | 109.3688 | 102.324 | 112.3854 | 118.0615 | 80.0781 | 67.475 | 71.0615 | 86.6146 | 97.0615 | 89.6927 | 103.6927 | 71.7375 | 245.4469 | 215.1063 | 227.3073 | 390.1229 | 153.5531 | 191.5864 | 184.3521 | 339.0333 | 306.7542 | 190.4136 | 207.8937 | 387.5698
2041 | 24.4454 | 152.2213 | 223.115 | 227.9588 | 189.5987 | 135.7787 | 154.6225 | 154.1063 | 139.5488 | 124.4425 | 125.59 | 142.8525 | 149.5163 | 205.9262 | 183.705 | 173.8525 | 212.9913 | 164.59 | 135.2625 | 151.6225 | 201.115 | 159.82 | 364.64 | 334.6138 | 169.8438 | 195.18 | 212.9588 | 230.2863 | 308.6312 | 120.0412 | 123.345 | 117.2625 | 116.3275 | 117.1388 | 136.2538 | 109.655 | 91.5575 | 205 | 234.7787 | 241.59 | 302.5987 | 126.8525 | 130.7462 | 123.475 | 119.295 | 86.7787 | 73.7375 | 72.7787 | 106.8113 | 164.4013 | 207.1475 | 218.4925 | 165.8937 | 141.6312 | 130.6725 | 133.41 | 142.525 | 155.6312 | 180.82 | 200.6312 | 173.7462 | 132.1063 | 142.2625 | 184.8937 | 161.7875 | 169.9913 | 198.7462 | 206.0412 | 237.8438 | 129.8113 | 137.41 | 204.8937 | 188.36 | 197.59 | 178.2863 | 200.8438 | 343.59 | 156.2625 | 147.2625 | 155.295 | 223.655 | 132.9675 | 148.8937 | 140.705 | 145.0738 | 184.7462 | 173.41 | 184.5163 | 131.2213 | 99.3275 | 113.9675 | 110.3362 | 120.6312 | 206.8438 | 223.2213 | 213.59 | 393.5575 | 231.41 | 225.1063 | 227.3775 | 327.0738 | 143.3688 | 58.065 | 83.0412 | 143.3688 | 132.7137 | 109.7787 | 110.5163 | 120.8525 | 113.6725 | 92.2213 | 103 | 118.705 | 66.8113 | 88.6725 | 71.705 | 96.9588 | 97.5575 | 86.5488 | 99.2213 | 82.3362 | 236.0412 | 231.4837 | 232.2213 | 391.4837 | 159.8612 | 171.4013 | 203.4512 | 309.8113 | 302.1887 | 211.0412 | 192.0325 | 384.835
2042 | 24.4564 | 155.9465 | 212.086 | 239.914 | 173.8069 | 135.4952 | 172.0746 | 169.7208 | 158.914 | 130.3442 | 132.4838 | 139.6233 | 157.8929 | 205 | 181.3767 | 171.2256 | 237.6558 | 172.3442 | 128.2371 | 171.3117 | 188.7094 | 143.8815 | 336.0325 | 379.5698 | 190.6769 | 211.0325 | 226.8069 | 256.1606 | 304.9465 | 108.3442 | 92.3231 | 108.5162 | 133.6883 | 142.4513 | 150.9675 | 142.7419 | 87.0325 | 207.4952 | 232.3442 | 248.2256 | 285 | 123.7094 | 117.9254 | 140.7094 | 123.6023 | 84 | 81.1931 | 70.8604 | 116.7419 | 182.4302 | 207.3652 | 186.7419 | 151.3977 | 137.4302 | 115.9789 | 126.9465 | 122.1071 | 151.8604 | 166.1721 | 196.7744 | 158.8604 | 144.8815 | 134.4627 | 169.5373 | 134.5162 | 193.4513 | 182.1396 | 194 | 259.6233 | 139.5698 | 133.3556 | 192.0325 | 217.7419 | 203.5373 | 205.8815 | 221.935 | 352.2046 | 145.8815 | 137.8279 | 158.5698 | 255.1071 | 122.5812 | 132.935 | 138.4627 | 145.3977 | 168.0535 | 164.1931 | 189.3652 | 133.4838 | 113.849 | 99.7954 | 102.6023 | 116.086 | 228.7954 | 227.3767 | 221.8604 | 386.0211 | 223.0535 | 237.1931 | 206.3977 | 329.7208 | 146.4513 | 83.5698 | 71.9465 | 146.4513 | 104.6883 | 107.8604 | 117.086 | 117.8815 | 96.7419 | 96.9789 | 103.4302 | 114.1396 | 76.3977 | 72.6883 | 67.4838 | 109.6023 | 94.2371 | 104.8815 | 101.4838 | 75.1185 | 224.6023 | 226.0325 | 234.914 | 384.8279 | 132.3977 | 188.3117 | 186.151 | 325.1625 | 290.4513 | 198.8279 | 205.7744 | 429.9025
2043 | 24.4674 | 165.9017 | 213.295 | 236.4434 | 191.7217 | 159.5417 | 141.7867 | 187.3117 | 156.9183 | 135.6716 | 128.3766 | 126.6716 | 175.1484 | 204.6067 | 193.8367 | 166.4416 | 233.115 | 174.4267 | 149.7384 | 163.0167 | 177.36 | 132.885 | 347.1967 | 372.05 | 164.525 | 222.1967 | 242.9517 | 227.0817 | 316.0817 | 113.6716 | 116.705 | 114.1967 | 139.7217 | 125.115 | 139.8033 | 138.23 | 98 | 232.525 | 235.8033 | 239.0167 | 281.5584 | 112.065 | 147.3284 | 128.5733 | 130.3933 | 84.4917 | 64.1633 | 80.4917 | 112.7217 | 188.18 | 192.36 | 186.7533 | 147.4583 | 140.23 | 137.885 | 133.8534 | 103.6716 | 161.0983 | 168.3933 | 204.9017 | 166.82 | 133.295 | 147.4434 | 156.95 | 139.9017 | 176.0167 | 172.9017 | 194.9833 | 247.8516 | 137.4583 | 165.6567 | 203.7867 | 212.345 | 187.3117 | 194.6883 | 200.0983 | 363.0334 | 135.6716 | 137.18 | 154 | 233.0334 | 147.7384 | 114.835 | 152.1317 | 140.0817 | 160.7533 | 146.8683 | 175.1466 | 127.6067 | 93.4583 | 88.9666 | 110.3766 | 118.0817 | 216.59 | 243.3766 | 231.3933 | 365.1967 | 218.7033 | 217.115 | 202.95 | 344.46 | 132.36 | 77.525 | 81.0167 | 132.36 | 111.9017 | 116.8033 | 117.705 | 106.3933 | 94.885 | 116.4267 | 107.115 | 107.4583 | 69.3117 | 77.935 | 62.885 | 113.935 | 116.525 | 93.4917 | 96.4917 | 87.4917 | 231.7867 | 232.8701 | 233.115 | 381.23 | 128.8516 | 178.2467 | 210.5733 | 377.2467 | 277.9333 | 198.3766 | 213.3117 | 394.5417
2044 | 24.4784 | 162.8982 | 215.0044 | 212.0044 | 181.6593 | 137.4425 | 151.4337 | 179.23 | 148.6593 | 147.4601 | 141.2212 | 142.2212 | 145.9026 | 202.77 | 172.8938 | 194.5707 | 224.115 | 160 | 129.5531 | 153.3363 | 195.3494 | 141.2212 | 344.4645 | 346.5487 | 151.7744 | 222.4513 | 213.1194 | 237.885 | 327.3275 | 127.23 | 114.9956 | 116.9956 | 131.3187 | 118.2169 | 136.8938 | 125.0974 | 97.6681 | 219.4425 | 232.8938 | 231.9912 | 251.77 | 130.4469 | 131.8894 | 146.9824 | 133.5575 | 88.7788 | 82.8938 | 84.8894 | 101.7744 | 194.2432 | 211.4557 | 214.4337 | 167.0132 | 124.1106 | 142.2388 | 105.3319 | 119.885 | 162.6637 | 170.23 | 205.8806 | 154.0088 | 138.3231 | 124.9956 | 182.0088 | 139 | 167.9956 | 173.2169 | 205.3275 | 270.6813 | 157.0132 | 134.4469 | 212.3275 | 187.6681 | 184.4293 | 201 | 201.4425 | 347.7656 | 150.4469 | 148.5531 | 155.8806 | 219.8674 | 126.8894 | 148.4513 | 138.5399 | 149.2256 | 185.5575 | 165.77 | 203.115 | 122.6725 | 115.3363 | 107.1106 | 122.6681 | 126.5619 | 224.7656 | 252.9026 | 235.3319 | 368.1062 | 266.9069 | 212.5399 | 230.3319 | 314.8762 | 152.0088 | 67.5399 | 74.2125 | 152.0088 | 110.6681 | 114.8894 | 115.4425 | 111.4381 | 103.6637 | 100.7831 | 98.5575 | 129.115 | 64.3275 | 59.4425 | 71.885 | 94.7788 | 102.0044 | 98.885 | 99.4513 | 89.6769 | 237.4513 | 198.4293 | 225.2212 | 364.115 | 155.1106 | 156.3187 | 222.345 | 349.5663 | 312.23 | 211.4425 | 207.3319 | 374.5443
2045 | 24.4894 | 145.8771 | 210.4416 | 206.3354 | 192.5417 | 141.4917 | 168.6646 | 165.8604 | 161.7542 | 122.1958 | 141.2792 | 143.7542 | 123.827 | 216.7542 | 162.2792 | 168.2291 | 214.7875 | 159.1396 | 135.9666 | 148.6146 | 168.2125 | 145.5812 | 308 | 356.7875 | 159.8437 | 209.1396 | 199.5979 | 244.5083 | 334.4355 | 114.7208 | 121.0334 | 125.6146 | 156.4021 | 128.7375 | 130.0895 | 146.3126 | 93.6479 | 220.7875 | 229.0395 | 252.0562 | 267.7208 | 125.0167 | 128.6646 | 172.7314 | 129.3854 | 83.3126 | 74.7375 | 84.3687 | 108.1229 | 161.5311 | 188.4188 | 233.0167 | 143.9833 | 124.6313 | 113.2291 | 108.7375 | 124.6646 | 169.2291 | 158.4583 | 218.2959 | 140.3354 | 156.0167 | 135.7041 | 164.2792 | 141.0895 | 172.1897 | 181.2792 | 216.6146 | 246.0895 | 136.4416 | 133.6874 | 222.8771 | 184.8771 | 209.4188 | 201.7375 | 200.0834 | 372.2125 | 146.8604 | 150.6646 | 170.8771 | 257.8938 | 130.302 | 134.0334 | 164.0501 | 144.2291 | 179.4188 | 182.7041 | 191.4522 | 116.4249 | 109.3854 | 108.1229 | 120.7375 | 115.3687 | 245.6479 | 228.3687 | 236.6479 | 377.6146 | 231.5417 | 239.6479 | 233.9833 | 336.7709 | 137.2291 | 93.173 | 90.2792 | 137.2291 | 108.9833 | 113.5083 | 120.7208 | 123.1229 | 109.3687 | 93.9499 | 97.3354 | 119.9105 | 75.2458 | 62.8771 | 79.1229 | 92.6313 | 93.7542 | 104.0334 | 89.5812 | 71.3687 | 223.6479 | 224.525 | 227 | 366.4644 | 156.6146 | 183.4917 | 199.8938 | 326.3126 | 296.2792 | 213.0334 | 208.7709 | 395.7375
2046 | 24.5004 | 145.1352 | 230.2678 | 227.4618 | 178.5153 | 143.5127 | 152.9464 | 168.4847 | 158.1071 | 142.5663 | 129.676 | 139.4311 | 144.9464 | 212.1607 | 152.4337 | 175.1071 | 201.5689 | 152.7296 | 148.3775 | 152.7296 | 185.0816 | 159.2678 | 313.4082 | 339.5127 | 171.3775 | 206.2449 | 209.7832 | 239.3775 | 314.8138 | 122.5382 | 109.1633 | 128.9184 | 146.7577 | 130.8903 | 143.6479 | 121.0816 | 85.3521 | 208.1097 | 258.2397 | 275.051 | 281.3521 | 118.8112 | 111.324 | 129.0816 | 125.1352 | 59.5689 | 78.9184 | 85.5127 | 108.8648 | 199.3775 | 173.921 | 223.9719 | 151.9464 | 125.2449 | 125.3801 | 115.0816 | 111.5153 | 177.4592 | 173.9719 | 199.5408 | 159.8393 | 147.7832 | 147.5637 | 152.8112 | 153.1607 | 147.7041 | 169.2704 | 218.9719 | 257.2143 | 160.1888 | 158.2423 | 223.3521 | 186.5689 | 193.7041 | 207 | 169.7322 | 384.7551 | 153.5408 | 136.2985 | 169.5944 | 247.4311 | 158.7015 | 123.3801 | 143.4056 | 149.6199 | 161.676 | 196.051 | 163.7322 | 143.6734 | 104.5944 | 109.676 | 124.3775 | 117.324 | 233.2959 | 233.7041 | 227 | 375.7806 | 217.1097 | 226.3495 | 239.7832 | 324.8903 | 144.1071 | 72.324 | 77.5944 | 144.1071 | 113.8367 | 110 | 130.8367 | 124.1352 | 109.4311 | 112.0255 | 114.8112 | 104.8648 | 77.9464 | 63.4873 | 80 | 89.5944 | 94.4337 | 91.0816 | 104.2143 | 74.8112 | 216.4337 | 213.7296 | 223.6199 | 427.7806 | 164.5153 | 183.3495 | 192.1352 | 308.2475 | 285.3521 | 197.7832 | 196.6199 | 395.727
2047 | 24.5114 | 147.9175 | 201.18 | 251.4276 | 198.7876 | 134.295 | 160.475 | 150.4249 | 150.2774 | 128.77 | 133.705 | 123 | 151.1475 | 195.475 | 164.7551 | 164.6225 | 214.4601 | 154.8349 | 135.4926 | 149.3775 | 192.295 | 134.7524 | 343.4276 | 333.9824 | 157.9026 | 226.2625 | 208.3423 | 230.9175 | 347.1326 | 98.4249 | 120.9351 | 121.41 | 154.4601 | 115.5724 | 135.59 | 134.6373 | 94.885 | 224.3126 | 219.4249 | 238.3275 | 291.475 | 127.0974 | 109.8024 | 136.59 | 126.295 | 76.1475 | 71.5575 | 74.82 | 105.6401 | 187.3775 | 200.6225 | 211.7375 | 159.0325 | 141.5751 | 146.705 | 126.2774 | 131.0501 | 173.8525 | 161.7375 | 203.885 | 173.7226 | 139.41 | 104.8674 | 157.8525 | 137.5074 | 164.5575 | 170.41 | 205.41 | 238.7524 | 153.3775 | 149.2124 | 230.82 | 199.9026 | 209.0825 | 203.6076 | 203.3924 | 365.0325 | 155.525 | 156.4574 | 163.4601 | 232.18 | 144.1475 | 144.705 | 145.41 | 127.5575 | 166.59 | 158.4425 | 190.1651 | 114.0825 | 102.295 | 114.59 | 116.2124 | 116.5399 | 219.3924 | 248.6401 | 224.9351 | 343.8173 | 237.295 | 205.9499 | 232 | 306.0325 | 134.0649 | 67.8525 | 72.9351 | 134.0649 | 102.6549 | 108.23 | 118.3275 | 124.8525 | 95.8024 | 93.885 | 114.3951 | 102.82 | 81.9351 | 73.1475 | 80.59 | 86.705 | 106.1651 | 98.1475 | 82.5074 | 78.23 | 230.82 | 210.9675 | 204.0649 | 383.8701 | 181.8376 | 163.9824 | 193.295 | 356.2801 | 293.705 | 198.7023 | 182.5223 | 363.18
2048 | 24.5224 | 159.6391 | 205.4434 | 223.7217 | 190.1545 | 132.964 | 167.0826 | 177.5979 | 168.7683 | 141.7164 | 126.727 | 122.2011 | 154.3968 | 202.2423 | 150.9534 | 171.1238 | 200.7111 | 176.6804 | 121.9587 | 143.8349 | 191.763 | 167.0413 | 317 | 358.4381 | 145.4328 | 218.3249 | 241.964 | 240.7217 | 315.6445 | 121.763 | 107.8815 | 121.9947 | 131.2836 | 141.0413 | 135.1651 | 170.9228 | 100.3196 | 209.2317 | 241.6445 | 252.3915 | 296.964 | 143.5619 | 125.0413 | 140.1598 | 124.8043 | 80.6751 | 70.5979 | 69.7577 | 94.7164 | 180.7164 | 209.2011 | 215.2011 | 162.6032 | 113.2836 | 145.3196 | 145.0879 | 118.7936 | 170.1238 | 162.4847 | 209.9587 | 149.9587 | 133.9228 | 135.7164 | 156.3609 | 152.7989 | 158.6445 | 174.51 | 201.5206 | 266.0879 | 141.9228 | 158.4847 | 223.5206 | 185.3555 | 200.8762 | 185.2783 | 218.2064 | 367.1651 | 152.1132 | 177.7736 | 142.6804 | 235.964 | 142.9228 | 143.6391 | 142.9587 | 128.036 | 169.6804 | 163.7164 | 163.2064 | 107.3143 | 101.2836 | 114.964 | 126.763 | 134.4434 | 235.1651 | 233.4021 | 211.0826 | 386.763 | 235.0053 | 222.5206 | 230.8815 | 316.474 | 144.0826 | 66.8402 | 61.4794 | 144.0826 | 119.4381 | 98 | 126.1598 | 121.1238 | 110.5619 | 101.7164 | 87.9121 | 96.4794 | 70 | 74 | 82.7217 | 87.8762 | 84.6391 | 100.2783 | 98.7577 | 67.5206 | 220.4847 | 207.237 | 218.3968 | 364.6338 | 149.2836 | 182.526 | 193.2423 | 325.0053 | 287.526 | 238.7736 | 223.763 | 372.7164
2049 | 24.5334 | 165.5812 | 191.4838 | 219.9253 | 207.8701 | 119.4091 | 154.4188 | 181.8701 | 138.0649 | 157.0325 | 104.5065 | 119.0325 | 164.2467 | 192.4838 | 172.698 | 157.3766 | 233.4188 | 168.1169 | 125.4513 | 159.0097 | 183.0974 | 159.5909 | 323.0227 | 374.0877 | 171.4188 | 200.8929 | 224.6234 | 234 | 298 | 111.0325 | 106.302 | 140.2467 | 119.5812 | 134.1071 | 116.5487 | 155.698 | 102 | 235.8052 | 222.6234 | 284.7955 | 284.9578 | 135.1396 | 122.237 | 140.1396 | 108.6883 | 100.5162 | 74.8701 | 78.1396 | 110.2045 | 193.9675 | 199.3214 | 213.4091 | 148.9675 | 103.1299 | 142.3539 | 116.7533 | 139.526 | 151.3864 | 149.1624 | 215 | 155.1721 | 123.5162 | 147.763 | 152.6558 | 154.4188 | 144.6136 | 208.4935 | 200.7208 | 233.6234 | 134.7857 | 144.1299 | 218.2467 | 204.5487 | 214.7955 | 190.1071 | 196.2695 | 346.6558 | 174.526 | 126.9253 | 146.8507 | 219.6558 | 127.698 | 146.6558 | 147.3117 | 144 | 166.1071 | 177.8279 | 146.4318 | 126.3864 | 89.237 | 100.8929 | 115.6883 | 122.8929 | 218.4416 | 226.5487 | 207.7111 | 374.4838 | 217.4416 | 222.5812 | 228.7857 | 340.3442 | 134.5162 | 64.6234 | 65.8929 | 134.5162 | 125.7955 | 98.1721 | 124.0747 | 110.474 | 103.8604 | 113.2467 | 117.1071 | 103.302 | 70.1721 | 75.3766 | 74.6234 | 101.2792 | 90.4091 | 101.8376 | 106.2792 | 68.0974 | 203.2045 | 217.7955 | 233.4091 | 391.5812 | 138.0974 | 162.6136 | 180.7305 | 304.8604 | 270.1948 | 186.7208 | 211.8279 | 382.526
2050 | 24.5444 | 176.525 | 187.7094 | 241.7435 | 188.5531 | 130.0782 | 144.0281 | 160.7094 | 153.8998 | 160.525 | 130.8156 | 126.2124 | 152.8437 | 192.7655 | 152.1844 | 169.7155 | 221.5531 | 137.2124 | 122.2405 | 136.475 | 194.3126 | 151.1343 | 343.7034 | 403.044 | 160.1062 | 208.7094 | 214.8657 | 233.0782 | 307.0341 | 114.8938 | 120.7314 | 123.8657 | 134.0281 | 129.978 | 121.0501 | 137.0281 | 100.525 | 208.0281 | 224.6654 | 275.3126 | 288.6973 | 127.497 | 130.2345 | 136 | 110.525 | 100.7876 | 59.6092 | 71.2345 | 113.6032 | 190.2906 | 174.0281 | 220.3908 | 146.2124 | 125.0281 | 121.6593 | 128.525 | 118.7375 | 137.3187 | 172.3407 | 211.6813 | 153.2345 | 129.503 | 138.2906 | 154.1343 | 139.6032 | 159.2345 | 181.6313 | 212.3187 | 225.1563 | 151.525 | 160.1282 | 207.2124 | 209.6032 | 211.0281 | 181.3687 | 211.6313 | 345.3687 | 154.2906 | 142.475 | 172.9719 | 221.6874 | 113.0842 | 142.4188 | 140.3126 | 144.3687 | 161.0561 | 177 | 183.6813 | 111.9499 | 99.2625 | 107.6032 | 117.7094 | 131.8156 | 235.1844 | 231.0501 | 244.0561 | 373.8437 | 238.7936 | 237.0281 | 250.503 | 340.1563 | 134.2345 | 60.0281 | 73.3407 | 134.2345 | 118.8938 | 98.6313 | 106.8657 | 129.4188 | 105.6032 | 101.475 | 105.0501 | 121.7876 | 71.1844 | 79.9719 | 70.0281 | 94.6593 | 100.3407 | 81.978 | 94.8657 | 83.3687 | 209.7375 | 211.8156 | 244.2625 | 392.2004 | 146.3627 | 179.4469 | 168.8718 | 309.9218 | 298.7094 | 190.022 | 212.475 | 363.0281
2051 | 24.5554 | 165.4768 | 183.9666 | 265.1837 | 191.1967 | 128.3599 | 152.6067 | 159.9165 | 173.5102 | 152.2301 | 127.4434 | 134.82 | 156.8701 | 200.6735 | 149.6568 | 190.8701 | 222.6234 | 149.3599 | 138.2468 | 143.59 | 185.4731 | 163.82 | 312.1132 | 354.2801 | 164.6067 | 204.5733 | 205.3265 | 235.4898 | 334.6104 | 111.7699 | 95.5566 | 111.5733 | 140.2468 | 150.82 | 116.8497 | 146.1967 | 95.3766 | 219.7532 | 251.987 | 258.6067 | 234.9629 | 114.5566 | 117.6067 | 135.2134 | 108.9165 | 90.2134 | 78.6234 | 63.1299 | 104.3766 | 194.82 | 182.8033 | 204.9165 | 156.1967 | 131.0501 | 130.7699 | 117.6735 | 122 | 153.3766 | 166.7699 | 196.6067 | 147.0965 | 143.6234 | 144.1967 | 168 | 137.2597 | 149.7532 | 182.3599 | 222.6735 | 221.5232 | 147.7532 | 142.41 | 220.3432 | 202.1466 | 218.82 | 183.59 | 206.2634 | 343.6568 | 155.2801 | 146.2468 | 163.6067 | 238 | 140.1967 | 130.6067 | 124.9833 | 144.82 | 175.6568 | 172.6735 | 171.3599 | 123.4267 | 95.8033 | 100.5399 | 110.2301 | 128.8367 | 236.7866 | 219.18 | 252.757 | 388.4898 | 255.2801 | 240.1002 | 255.577 | 334.7532 | 127.5065 | 69.7866 | 65.0167 | 127.5065 | 115.9666 | 99.7532 | 97.3265 | 118.1967 | 98.1466 | 106.6234 | 95.7366 | 113.9666 | 68.8534 | 70.8033 | 75.6568 | 100.4267 | 95.3599 | 104.7866 | 87.0965 | 78.3135 | 216.7366 | 212.3766 | 238.8367 | 346.2931 | 124.6698 | 179.1633 | 187.7699 | 316.3599 | 288.8701 | 172.1299 | 221.9499 | 377.7032
2052 | 24.5664 | 150.1343 | 194.5074 | 224.5522 | 184.0598 | 137.791 | 152.4627 | 179.3731 | 154.1044 | 151.0596 | 119.5074 | 128.3284 | 139.791 | 188.0148 | 139.9701 | 172.3284 | 222.4327 | 160.6716 | 132.8506 | 147.2537 | 214.3284 | 160.4627 | 334.4627 | 332.7015 | 163.6269 | 210.791 | 222.791 | 255.5225 | 293.6118 | 119.8358 | 104.5373 | 119.4627 | 130.4627 | 139.3135 | 149.0299 | 145.3284 | 98.2836 | 226.821 | 212.4327 | 254.9105 | 290.1194 | 121.2389 | 115.5821 | 130.1194 | 131.7164 | 86.791 | 76.5522 | 78.7463 | 113.5522 | 191.4627 | 184.0895 | 224.791 | 155.5373 | 118.1642 | 134.8657 | 109.4028 | 118.8657 | 157.1194 | 171.7015 | 193.3284 | 170.1194 | 134.2389 | 138.3135 | 163.4031 | 162.433 | 159.5373 | 192.4179 | 206.6716 | 240 | 157.1194 | 139.5821 | 231.4926 | 209.9852 | 212.3284 | 184.9552 | 191 | 332.5074 | 142.6865 | 136.6716 | 162 | 236.5373 | 133.4777 | 128.791 | 132.9701 | 142.7164 | 161.5821 | 156.6716 | 178.4926 | 119.2985 | 94.3731 | 116.4627 | 100.7015 | 120.6269 | 241.6716 | 223.5821 | 224.6118 | 409.3583 | 244.776 | 224.9849 | 223.4476 | 354.567 | 146.6568 | 72.3731 | 64.1343 | 146.6568 | 121.4926 | 111.8358 | 117.5074 | 118.3731 | 109.5373 | 102.0448 | 107.0299 | 121.3731 | 60.8059 | 72.0895 | 65.3432 | 95.2537 | 106.6716 | 107.5821 | 109.0747 | 56.6417 | 229.0747 | 222.388 | 229.3731 | 368.4478 | 148.6867 | 190.2985 | 196.0448 | 336.2386 | 277.0148 | 195.6865 | 231.0747 | 395.5673
2053 | 24.5774 | 161.115 | 209.9736 | 239.1062 | 159.5311 | 134.1238 | 152.2477 | 176.115 | 180.7788 | 159.8405 | 126.7876 | 124.2124 | 141.6549 | 207.885 | 152.3451 | 168.8762 | 238.3539 | 166.115 | 148.8052 | 155.9824 | 200.699 | 159.3627 | 336.2389 | 331.7435 | 163.7876 | 211.3275 | 218.4601 | 230.7788 | 324.4689 | 120.1238 | 100.7699 | 126.5487 | 137.5487 | 119.5311 | 131.5663 | 136.5663 | 93.3098 | 212.7699 | 230.7876 | 247 | 275.4777 | 108.7699 | 116.2124 | 130.9647 | 141.5575 | 87.5487 | 85.5751 | 76.8762 | 123.2389 | 193.4601 | 192 | 222.0088 | 149.3363 | 118.0974 | 125.6549 | 145.885 | 105.0088 | 154.6461 | 165.2124 | 188.1062 | 163.885 | 121.3275 | 119.416 | 152.8585 | 133.0974 | 156.2124 | 192.0088 | 213.6637 | 246.416 | 145.5751 | 137.7788 | 223.9912 | 191.2212 | 207.3275 | 184.5399 | 197.8585 | 344.6549 | 164.2389 | 145.8762 | 164.2124 | 230.5575 | 110.0886 | 125.5663 | 143.1327 | 152.7788 | 158.4513 | 166.3186 | 171.8762 | 130.8762 | 95.5399 | 116.4689 | 95.9824 | 119.0176 | 246.0088 | 223.3275 | 253.4777 | 383.5751 | 270.9202 | 255.2389 | 262.5575 | 377.0441 | 132.0088 | 73.9824 | 75.5575 | 132.0088 | 109.7876 | 116.5487 | 127.4425 | 119.3186 | 101.7876 | 107.1062 | 98.1948 | 119.2212 | 77.4513 | 78.4513 | 75.1238 | 101.5487 | 108.3539 | 102.2389 | 96.8938 | 78.2212 | 222.4249 | 230.814 | 226.5575 | 363.8497 | 130.7347 | 192.3627 | 196.2389 | 360.8405 | 298.2124 | 217.2389 | 215.5751 | 378.5399
2054 | 24.5884 | 153.7964 | 224.6946 | 235.994 | 173.8623 | 127.2695 | 139.4731 | 168.0958 | 173.6946 | 117.006 | 121.1018 | 129.8982 | 151.7006 | 201.8922 | 142.0659 | 178.2994 | 221.2994 | 157.8623 | 129.1018 | 166.9641 | 180.3713 | 133.4311 | 323.7006 | 356.8982 | 156.2335 | 214.3653 | 206.7006 | 232.8024 | 310.8383 | 112.3353 | 110.0359 | 129.4311 | 140.8982 | 142.503 | 126.0359 | 133.6048 | 104.5269 | 222.503 | 222.7665 | 250.9701 | 251.5689 | 113.8323 | 123.7665 | 149.4251 | 138.5988 | 88.0958 | 69.2994 | 83.497 | 113.9701 | 182.4012 | 191.5329 | 213.1317 | 145.1317 | 125.2635 | 129.3952 | 139.8922 | 103.1377 | 160.491 | 167.8623 | 198.5389 | 165.5988 | 125.2994 | 139.4311 | 170.9281 | 143.7665 | 163.7665 | 181.9641 | 209.2275 | 259.6587 | 134.2036 | 138.4012 | 225.8623 | 194.1018 | 208.7305 | 197.4671 | 217.0958 | 353.5329 | 149.5988 | 156.7006 | 174.3353 | 231.3353 | 130.1018 | 120.2695 | 127.6347 | 152.4671 | 152.0659 | 173.7964 | 179.6647 | 133.0599 | 105.6647 | 97 | 109.2994 | 111.0719 | 236.1976 | 219.8263 | 230.7365 | 370.3353 | 249.2695 | 234.7605 | 262.6347 | 336.3952 | 130.6048 | 82.3952 | 73.7665 | 130.6048 | 105.9701 | 118.7305 | 126.4311 | 122.3593 | 99.1377 | 103.994 | 115.3293 | 116.9641 | 74.8024 | 71.8323 | 68.2695 | 100.9281 | 89.1976 | 90.1677 | 96.7365 | 77.1317 | 234.497 | 206.9042 | 228.9701 | 389.6287 | 159.8623 | 169.9341 | 179.9641 | 312.8683 | 298.0599 | 202.5988 | 206.3054 | 388.6647
2055 | 24.5994 | 142.7041 | 206.2124 | 214.9666 | 163.8831 | 141.1957 | 158.346 | 160.8497 | 153.9833 | 137.7876 | 129.9666 | 121.5251 | 155.4749 | 180.8497 | 142.1957 | 175.0167 | 220.4749 | 149.1122 | 138.2124 | 157.4916 | 195.5418 | 133.6039 | 321.3293 | 365.2411 | 156.5084 | 205.5585 | 213.9165 | 238.5585 | 327.7876 | 117.5418 | 114.346 | 125.179 | 129.821 | 148.5919 | 140.4248 | 146.5919 | 86.6373 | 232.5251 | 227.1623 | 253.4295 | 261.7208 | 113.1957 | 121.0334 | 123.4081 | 127.3627 | 73.4749 | 71.4248 | 76.179 | 128.7208 | 185.5251 | 195.4081 | 211.179 | 139 | 117.6707 | 118.1289 | 114.179 | 124.4248 | 137.5538 | 155.179 | 210.9045 | 162.2291 | 129.3914 | 132.9666 | 154.1957 | 141.0334 | 164.9666 | 174.2124 | 193.6373 | 240.0621 | 147.7542 | 144.4749 | 220.7996 | 201.4916 | 200.4582 | 189.9045 | 197.3126 | 358.1456 | 145.7375 | 159.4916 | 178.0668 | 230.8878 | 137.7375 | 127.8043 | 133.2458 | 151.0501 | 147.2792 | 160.9833 | 163.8878 | 108.7208 | 98.2458 | 97.9833 | 107.7375 | 131 | 226.2625 | 205.5371 | 235.1336 | 377.0167 | 254.5919 | 211.7375 | 265.0501 | 326.358 | 147.0334 | 66.9499 | 73.9833 | 147.0334 | 116.2959 | 103.5752 | 121.6874 | 102.358 | 113.7876 | 80.5084 | 100.7375 | 112.4081 | 81.2959 | 67.0167 | 77.0334 | 84.4415 | 85.654 | 94.2458 | 104.0835 | 71.9833 | 221.2792 | 222.2625 | 238.3126 | 368.0668 | 146.9332 | 176.4415 | 168.2792 | 322.6874 | 286.9953 | 196.7709 | 224.5418 | 378.2959
2056 | 24.6104 | 150.7419 | 207.5796 | 221 | 192.4854 | 150.4838 | 144.8717 | 188.2581 | 157 | 135.1299 | 136 | 113.1283 | 155.0958 | 201.0309 | 151.4838 | 174.323 | 221.3864 | 172.7094 | 148.6136 | 149.1916 | 185.677 | 157.4188 | 310.8717 | 390.8345 | 161.1948 | 197.8393 | 224.612 | 230.0649 | 323.0649 | 111.0325 | 98.0325 | 131.5471 | 124.2272 | 138.4529 | 145.9984 | 132.0649 | 104.677 | 228.7743 | 240.1607 | 231.8409 | 264.9351 | 121.7094 | 117.0649 | 139.7419 | 116.0341 | 79.0325 | 84.2257 | 84.612 | 129.612 | 180.4838 | 202.7078 | 219.612 | 143.6461 | 137.4204 | 134.3215 | 122.0958 | 133.3539 | 170.1607 | 167.4838 | 189.9691 | 162.6445 | 142.289 | 143.9042 | 161.677 | 140.1623 | 175.1299 | 181.5162 | 212.4513 | 257.9968 | 148.8068 | 156.2272 | 239.19 | 206.8717 | 204.3864 | 171.0341 | 190.6461 | 366.9335 | 149.0325 | 152.4822 | 163.1607 | 214.0016 | 136.1283 | 116.1932 | 127.0309 | 141.4838 | 149.7435 | 161.1607 | 144.9367 | 115.0325 | 102.6136 | 107.1948 | 103.289 | 129.7094 | 222.9675 | 240.3864 | 220.5503 | 372.7094 | 242.6461 | 220.711 | 250.0633 | 346.5796 | 136.612 | 73.4188 | 79.323 | 136.612 | 107.2257 | 97.4854 | 133 | 119.2241 | 104.9351 | 85.4529 | 105.323 | 120.4822 | 72.7419 | 63.4838 | 68.677 | 93.9675 | 98.8701 | 93.4513 | 90.0325 | 77.323 | 219.0974 | 220 | 227 | 357.2906 | 165.1948 | 178.2241 | 172.8084 | 328.8377 | 327.0633 | 194.2906 | 214.9351 | 375.9367
2057 | 24.6214 | 147.2959 | 198.2194 | 225.8673 | 202.1021 | 151.9745 | 158.9745 | 186.0255 | 156.4592 | 140.5102 | 134.9184 | 106.0561 | 145.8673 | 187.2194 | 146.2143 | 172.6174 | 208.5663 | 170.8928 | 151.699 | 123.1633 | 175.4847 | 146.7551 | 319.8367 | 334.3265 | 174.4031 | 209.5153 | 213.352 | 230.5919 | 321.1582 | 118.3265 | 106.949 | 112.1888 | 142.8367 | 151.051 | 127.5969 | 135.0255 | 94.7551 | 236.1377 | 240.1122 | 272.1938 | 267.653 | 128.0051 | 128.6785 | 133.5919 | 140.1072 | 86.3265 | 75.7806 | 71.1888 | 117.8112 | 177.9184 | 173.2959 | 201.3215 | 161.5969 | 145.699 | 118 | 109.0816 | 122.7041 | 163.8928 | 166 | 206.7551 | 158.4081 | 128.949 | 155.5663 | 160.1377 | 143.1021 | 176.4541 | 180.0255 | 204.4847 | 224.4387 | 151.0255 | 169.1582 | 203.6836 | 213.6735 | 195.6224 | 198.8928 | 204 | 340.9745 | 153.8928 | 136.9234 | 159.8673 | 235.1072 | 124.1888 | 111.8112 | 119.4387 | 139.1888 | 164.9694 | 161.9234 | 166.0561 | 126.1122 | 108.6735 | 118.5102 | 95.3571 | 119.7806 | 232.9795 | 223.7806 | 244.4847 | 368.7296 | 248.4286 | 233.5102 | 242.4642 | 329.9184 | 123.4592 | 65.1888 | 84.648 | 123.4592 | 101.7551 | 112.7806 | 128.9439 | 100.4898 | 101.4337 | 106.1633 | 108.7551 | 98.7041 | 70.1072 | 65.2449 | 66.3265 | 89.9184 | 88.0816 | 92.2449 | 94.6224 | 79.4031 | 225.2959 | 216.2143 | 225.3776 | 362.8928 | 176.7806 | 164.6275 | 196.9439 | 308.5919 | 298.3725 | 192.3215 | 209.8112 | 391.648
2058 | 24.6324 | 147.9155 | 207.9326 | 238.4346 | 170.8481 | 152.6499 | 158.8018 | 173.4768 | 149.6287 | 121.8692 | 128.6076 | 120.3924 | 156.173 | 199.1941 | 141.3713 | 157.6328 | 208.2194 | 176.5654 | 134.5443 | 135.2194 | 173.9155 | 139.9789 | 318.8059 | 345.7173 | 172.1771 | 214.1941 | 221.5232 | 221.9367 | 302.5443 | 126.0422 | 112.5402 | 113.6751 | 130.2152 | 135.5654 | 135.4768 | 134.9155 | 86.8481 | 244.6287 | 260.8692 | 291.6287 | 281.671 | 145.9537 | 138.346 | 118.7173 | 133.8692 | 94.6076 | 61.827 | 73.5232 | 119.1519 | 173.8692 | 171.654 | 189.9578 | 166.3672 | 125.7173 | 124.5021 | 114.827 | 122.0845 | 170.9789 | 164.8692 | 197.1519 | 164.8018 | 145 | 147.3038 | 170.8903 | 121.7425 | 149.9789 | 175.3924 | 200.654 | 248.671 | 144.6962 | 144.8903 | 234.4094 | 201.4346 | 196.6076 | 210.7848 | 194.9537 | 347.8692 | 160.1308 | 158.7384 | 171.5865 | 227.4557 | 124.827 | 123.8944 | 145.9326 | 139.5443 | 143.7384 | 176.9537 | 181.2405 | 145.4557 | 98.4135 | 104.1097 | 120.9114 | 103.8481 | 253.017 | 209.2616 | 243.4809 | 360.3672 | 227.7173 | 213.173 | 256.1267 | 332.9367 | 133.3079 | 69.7848 | 80.4346 | 133.3079 | 98.9367 | 101.3713 | 114.3249 | 117.1519 | 105.7384 | 109.173 | 97.7384 | 109.6751 | 68.9578 | 71.7384 | 74.0422 | 88.1308 | 90.4346 | 102.6962 | 99.8481 | 68.6962 | 216.8692 | 214.7636 | 217.6076 | 367.1519 | 162.2616 | 196.5654 | 180.6287 | 294.2827 | 271.8059 | 180.1097 | 209.7384 | 385.7384
2059 | 24.6434 | 152.9556 | 187.3451 | 225.2004 | 179.7847 | 138.4248 | 140.9499 | 165.8349 | 151.0444 | 130.2096 | 116.7551 | 123.6902 | 151.3599 | 190.2449 | 151.1651 | 174.9351 | 220.7551 | 168.2654 | 138.3451 | 150.705 | 186.6253 | 148.2449 | 337.4898 | 349.1298 | 189.1503 | 208.4898 | 222.6754 | 236.41 | 323.4545 | 119.2449 | 93.0797 | 115.0353 | 112.705 | 137 | 122.82 | 149.6902 | 94.3098 | 235.7198 | 248.8553 | 263.5455 | 263.0148 | 122.41 | 121.3451 | 117.115 | 129.82 | 79.5102 | 67.82 | 78.5102 | 120.2449 | 175.1298 | 192.4898 | 209.7346 | 158.2096 | 120.2802 | 138.3451 | 120.5251 | 107.3098 | 179.8349 | 159.0501 | 201.1948 | 148.7198 | 138.5102 | 136.5102 | 167.4248 | 149.295 | 159.7198 | 173.9704 | 219.7198 | 227.0649 | 152.2449 | 140.8349 | 210.3098 | 198.2301 | 187.705 | 217.9704 | 175.8349 | 341.4601 | 161.82 | 152.705 | 169.7699 | 223.3599 | 133.18 | 144.5752 | 125.0501 | 149.5399 | 142.7198 | 160.4898 | 191.705 | 137.82 | 105.5399 | 108.1651 | 106.5251 | 112.4898 | 222.3747 | 220.0148 | 257.6253 | 356.6344 | 221.6902 | 203.6401 | 225.41 | 348.885 | 150.7904 | 81.3952 | 78.41 | 150.7904 | 116.0649 | 114.41 | 113.8497 | 118.8349 | 100.295 | 102.59 | 95.5399 | 113.3952 | 78.705 | 69.5399 | 71.6697 | 92.18 | 89.295 | 103.4601 | 104.3599 | 76.2449 | 219.6048 | 230.2153 | 211.0649 | 368.8349 | 168.295 | 190.3303 | 179.6845 | 302.0797 | 292.8497 | 194.7847 | 207.8349 | 373.8052
2060 | 24.6544 | 124.3747 | 187.7602 | 201.3019 | 190.1617 | 146.2345 | 148.3073 | 171.9272 | 183.3126 | 155.7709 | 114.8383 | 111.3073 | 154.5418 | 190.9326 | 142.9218 | 168.1563 | 205.9326 | 159.1349 | 139.3747 | 146.3126 | 174.2238 | 153.5418 | 347.469 | 358.6927 | 174.4529 | 220.6199 | 207.2184 | 235.6146 | 348.167 | 125.1563 | 108.1563 | 112.9004 | 121.2184 | 135.1563 | 123.3801 | 141.6092 | 103.3126 | 239.9326 | 233.4529 | 238.833 | 261.787 | 130.833 | 118.0728 | 118.3801 | 130.3801 | 68.3019 | 70.531 | 63.6146 | 124.0054 | 189.6092 | 202.7762 | 235.151 | 166.2559 | 109.9218 | 132.9218 | 124.3747 | 122.7655 | 180.3962 | 158.4529 | 208.318 | 159.6927 | 113.4743 | 125.3019 | 167.5525 | 145.3908 | 171 | 160.7548 | 217.787 | 232.6146 | 158.7709 | 146.0054 | 224.8437 | 185.0889 | 183.6199 | 184.167 | 190.8383 | 334.5364 | 155.3126 | 148.9272 | 170.3126 | 230.2291 | 132.9272 | 131.2345 | 123.2238 | 151.8544 | 161.9893 | 168.318 | 179.6306 | 127.0107 | 115.5364 | 102.6874 | 102.6927 | 121.2398 | 241.2291 | 228.6199 | 241.5364 | 379.8651 | 208.3854 | 211.213 | 229.531 | 356.2238 | 132.9111 | 68.6146 | 73.3126 | 132.9111 | 117.3126 | 113.6146 | 123.3234 | 125.2345 | 107.1456 | 104 | 103.0782 | 110.1402 | 73.6981 | 77.0782 | 84.1617 | 95 | 88.4636 | 95 | 106.0054 | 84.6146 | 231.4636 | 220.4529 | 218.4582 | 375.5418 | 158.8598 | 175.6874 | 206.3126 | 319.3073 | 307.8544 | 207.3126 | 209.318 | 366.9893
2061 | 24.6654 | 135.2477 | 204.5979 | 223.7828 | 177 | 138.0826 | 154.1131 | 159.8869 | 172.4434 | 148.2064 | 125.1239 | 112.6391 | 151.5566 | 181.0719 | 148.8349 | 158.5673 | 207.896 | 189.0826 | 153.4434 | 143.4327 | 186 | 146.4021 | 336.5979 | 344.8976 | 185.4847 | 207.2477 | 222.2477 | 234.4434 | 331.7523 | 116.526 | 108.1544 | 139.4954 | 137.8456 | 135.1544 | 131 | 149.4434 | 94.6804 | 229.1131 | 250.5566 | 266.4327 | 256.9265 | 143.4129 | 126.9174 | 123.1239 | 134.4847 | 78.9587 | 79.4847 | 66.2783 | 115.3196 | 194.2477 | 187.8043 | 237.763 | 139.659 | 121.5871 | 126.3716 | 140.6804 | 127.0826 | 160.3609 | 168.8456 | 194.1957 | 153.8869 | 110.5351 | 131.8043 | 150.0719 | 134.0413 | 165.5673 | 181.4847 | 192.7936 | 237.8349 | 148.0107 | 138.9174 | 229.6391 | 181.2981 | 180.1544 | 177.659 | 196.6499 | 353.0198 | 155.3088 | 146.474 | 160.7217 | 225.0107 | 127.2783 | 125.3196 | 134.6804 | 138.3196 | 175.2064 | 153.8761 | 156.8761 | 121.4633 | 112.289 | 109.4021 | 105.5153 | 107.5979 | 230.8976 | 224.1957 | 240.526 | 357.5871 | 215.6284 | 237.4847 | 231.4542 | 356.4954 | 154.5979 | 70.9587 | 66.2783 | 154.5979 | 114.1131 | 111.4847 | 109.1437 | 118.0413 | 111.4129 | 100.8043 | 110.9067 | 125.0107 | 64.9587 | 73.7217 | 74.5153 | 99.474 | 85 | 102.6697 | 94.4434 | 84.4021 | 228 | 234.6804 | 222.7217 | 362.3303 | 147.5046 | 186.237 | 189.052 | 338.5351 | 290.8043 | 203.474 | 198.711 | 384.3609
2062 | 24.6764 | 127.6461 | 204.6814 | 240.3804 | 174.677 | 137.6505 | 163.3407 | 161.9426 | 167.9956 | 144.9691 | 119.9956 | 118.3142 | 158.3274 | 195.6593 | 149.699 | 151.9779 | 241.3539 | 180.3539 | 144.3495 | 145.376 | 192.6372 | 151.2966 | 345.6417 | 325.9603 | 171.3628 | 193.6726 | 200.7079 | 232.3186 | 342.3407 | 107 | 118.6593 | 120.9735 | 135.9691 | 135.7034 | 124.0309 | 147.9823 | 93.3363 | 228.3848 | 256 | 268.7079 | 293.4069 | 120.0132 | 130.3363 | 123.6372 | 125.0088 | 75.0265 | 72.3319 | 73.6461 | 116 | 184.6549 | 185.6461 | 229.6814 | 177.3142 | 135.3628 | 111.6505 | 130.7079 | 122.3363 | 162.6505 | 161.6593 | 198.3451 | 148.3098 | 144.6902 | 133.6284 | 159.6814 | 145.9382 | 160.6372 | 175.9912 | 205.6549 | 234.7167 | 134.9823 | 145.9868 | 223.699 | 214.6726 | 187.6726 | 203.3672 | 184.9823 | 380.0221 | 165.699 | 150.0265 | 167.9559 | 214.6372 | 125.0221 | 132.3054 | 136.323 | 144.6417 | 163.3407 | 159.6681 | 163 | 138.0441 | 103.9735 | 107.6593 | 114.6593 | 105.6902 | 213.9515 | 232.3274 | 231.3319 | 384.9691 | 229.6858 | 233.3186 | 225.925 | 327.6858 | 160.323 | 76.3186 | 67.3407 | 160.323 | 124.0044 | 101.677 | 125.7079 | 110.3583 | 90.6681 | 94 | 126.6902 | 112.9779 | 69.9868 | 73.6549 | 82.9956 | 107.6726 | 93.6284 | 105.0618 | 93.3142 | 77.3495 | 224.3495 | 245.947 | 230.2877 | 346.9956 | 168.3495 | 193.6549 | 182.6108 | 355.7652 | 278.6902 | 207.6902 | 219.6372 | 384.3274
2063 | 24.6874 | 139.0649 | 195.5909 | 227.2597 | 175.1623 | 148.7857 | 163.0974 | 177.6754 | 172.4091 | 149.7079 | 125.4415 | 122.8702 | 157.9026 | 201.0649 | 136.0325 | 156.4935 | 222.7079 | 173.0649 | 144.9155 | 128.0649 | 198.4286 | 158.7079 | 356.6558 | 338.1494 | 160.7532 | 186.8702 | 189.2272 | 250.3571 | 327.2988 | 104.2468 | 116.4935 | 129.5585 | 147.2468 | 127.948 | 117.9155 | 155.6558 | 93.3766 | 212.5714 | 243.2663 | 240.3637 | 273.422 | 116.7532 | 132.7857 | 133.2143 | 126.8506 | 68.8506 | 74.3766 | 78.1819 | 113.2468 | 185.526 | 190.1819 | 220.2468 | 173.9546 | 123.7208 | 120.3766 | 117.1623 | 121.3442 | 159.6754 | 166.0325 | 197.8181 | 157.9675 | 128.4935 | 144.4611 | 155.7532 | 165.3962 | 173.6558 | 180.3442 | 210.3117 | 217.5065 | 139.2143 | 152.6883 | 209.6883 | 223.3571 | 191.8831 | 180.7728 | 194.0325 | 369.6558 | 156.8506 | 133.8702 | 184.5585 | 231.0974 | 124.2921 | 140.8702 | 135.4935 | 152.9026 | 158.6234 | 159.6883 | 166.7857 | 115.9351 | 114.2792 | 107.2143 | 113.1819 | 104.9804 | 235.3117 | 230.526 | 232 | 383.169 | 216.1819 | 236.9026 | 255.8702 | 321.7532 | 157.4286 | 81.2792 | 63.6558 | 157.4286 | 119.9351 | 103.1948 | 111.474 | 98.3117 | 92.4091 | 99.5065 | 114.2792 | 121.9675 | 74.2792 | 74.1819 | 79.4935 | 97.7728 | 108.5909 | 85.2597 | 96.4935 | 73.4415 | 226.6363 | 255.5456 | 240.3313 | 366.2077 | 160.6558 | 203.1298 | 201.8052 | 316.1623 | 280.0454 | 198.3766 | 227.1494 | 389.4091
2064 | 24.6984 | 146.9209 | 183.5147 | 231.2373 | 194.2676 | 152.0791 | 164.7227 | 156.9307 | 169.8711 | 139.8418 | 130.5742 | 119.9902 | 149.1484 | 207.8516 | 138.7129 | 152.416 | 206.0693 | 171.2969 | 154.6533 | 138.7724 | 180.4356 | 138.5049 | 350.6533 | 332.3467 | 169.9209 | 193.9707 | 212.9804 | 267.0098 | 306.4258 | 108.624 | 106 | 123.3564 | 146.9902 | 140.0098 | 128.0098 | 141.0987 | 94.9307 | 238.4062 | 222.5644 | 239.4658 | 296.8613 | 122.3564 | 137.8613 | 117.0889 | 128.376 | 82.4951 | 72.3662 | 69 | 109.4258 | 180.208 | 182.4258 | 215.7129 | 152.0693 | 118.8018 | 117.6533 | 119.8711 | 122.3564 | 143.9209 | 162.6533 | 204.5049 | 164.1982 | 127.624 | 130.5742 | 154.2276 | 155.0596 | 164.4453 | 183.4951 | 210.0693 | 229.7822 | 134.1387 | 161.4853 | 203.5147 | 235.7324 | 202.1484 | 179.1191 | 190.6533 | 369.7129 | 161.9404 | 133.1289 | 174.4356 | 248.7627 | 145.9209 | 139.0596 | 131.7724 | 147 | 159.9209 | 170.9804 | 170.4356 | 119.1289 | 110.2871 | 98.9307 | 106.1387 | 118.6533 | 234.3564 | 226.9902 | 232 | 356.9209 | 226.248 | 220.6631 | 244.792 | 328.7822 | 135.5147 | 78.3564 | 64.7822 | 135.5147 | 113.1484 | 111.792 | 120 | 108.7627 | 96.6436 | 108.2178 | 107.0791 | 118.5742 | 73.8516 | 63.9307 | 69.3564 | 85.4258 | 98.6533 | 97.0791 | 92.416 | 80 | 234.6631 | 229.4258 | 211.4453 | 388.3067 | 154.2969 | 199.9502 | 189.2871 | 317.8018 | 309.416 | 210.9804 | 207.4453 | 390.792
2065 | 24.7094 | 152.5251 | 172.5811 | 210 | 206.2064 | 143.8938 | 153.3126 | 163.1122 | 161.4248 | 151.0501 | 121.7315 | 125.6814 | 148.4248 | 206.3628 | 136.3126 | 166.2124 | 222.3807 | 166.8998 | 143.8938 | 152.8437 | 183.2184 | 135.8437 | 332.5191 | 347.531 | 169.994 | 212.8878 | 217.3568 | 248.4129 | 304.2064 | 127.4749 | 115.2184 | 116.994 | 149.7315 | 135.7936 | 113.469 | 127.0621 | 86.3628 | 251.2566 | 232.6874 | 272.8437 | 280.8317 | 124.4749 | 132.5251 | 110.5871 | 112.4749 | 81.469 | 68.0561 | 70.4749 | 116.0561 | 179.7315 | 188.3186 | 221.0561 | 154.7375 | 112.1122 | 115.006 | 118.4308 | 125.2124 | 151.7375 | 152.2625 | 195.9439 | 175.6814 | 139.8377 | 134.6372 | 146.4248 | 149.7756 | 153.7936 | 186.1563 | 212.3687 | 235.9499 | 148.6933 | 167.2566 | 194.7936 | 214.5752 | 209.9057 | 188.8317 | 192.7995 | 363.2566 | 156.531 | 143.0501 | 169.1062 | 257.463 | 140.463 | 137.463 | 144.3687 | 148.4749 | 171.4248 | 181.6253 | 164.3687 | 127.2064 | 107.8938 | 97.7375 | 112.5811 | 116.006 | 233.5251 | 228.6253 | 235.6874 | 382.068 | 248.463 | 230.7614 | 250.4308 | 327.2064 | 137.4869 | 78.2625 | 69.8437 | 137.4869 | 111.3186 | 107.4749 | 116.6814 | 124.1002 | 100.7936 | 111.6372 | 101.1062 | 114.1563 | 79 | 69.7434 | 77.3747 | 85.4189 | 95.6372 | 98.4809 | 102.8938 | 76.3126 | 231.4869 | 229.0501 | 206.3246 | 363.8377 | 143.2625 | 177.7876 | 198.6933 | 296.3628 | 312.8878 | 219.7816 | 198.2685 | 371.7255
2066 | 24.7204 | 154.1913 | 184.6205 | 216.8585 | 193.0467 | 131.7123 | 157.6687 | 172.8087 | 158.7123 | 153.6205 | 114.0482 | 115.4277 | 152.1897 | 197.2862 | 131.5241 | 158.9502 | 236.3313 | 180.0482 | 142.7621 | 149.5226 | 185.6641 | 138.619 | 323.3344 | 364.6205 | 166.3826 | 199 | 199.3344 | 241.5257 | 307.0498 | 112.8538 | 126.4277 | 109.9533 | 139.381 | 132.9502 | 108.9533 | 147.9051 | 81.4774 | 226.9502 | 243.1913 | 269.1415 | 266.0964 | 126.619 | 121.6174 | 122.4277 | 121.0964 | 79.2395 | 66.9984 | 77.1913 | 121.4759 | 176.2395 | 185.9984 | 225.7138 | 153.3328 | 119.5226 | 127.381 | 143.3344 | 125.9518 | 158.7154 | 153.2862 | 195.0964 | 166.9518 | 136.3344 | 140.2846 | 156.2862 | 132.6687 | 161.2379 | 174.2364 | 215.2862 | 252.8118 | 158.6174 | 141.0451 | 199.9518 | 212.7154 | 225.9036 | 179.8118 | 208.8087 | 341.6174 | 156.8554 | 141.0482 | 163.7605 | 233.7138 | 123.5723 | 120.5723 | 140.4277 | 150.2379 | 174.0467 | 161.3795 | 164.2379 | 111.7605 | 106.7621 | 105.0964 | 117.381 | 130.6672 | 231 | 212.9518 | 234.3795 | 390.9005 | 228.1431 | 254.4712 | 258.1882 | 312.9036 | 157.9036 | 71.6656 | 69.1897 | 157.9036 | 110.9036 | 103.9036 | 111 | 117.4292 | 104.8087 | 113.8554 | 104.9051 | 112.1431 | 75.9518 | 79.5708 | 86.9518 | 89.7636 | 103.1897 | 106.1415 | 98.3328 | 75.7154 | 263.3344 | 219.0467 | 229.7154 | 367.9549 | 153.4308 | 173.2379 | 204.4261 | 300.6221 | 294.4277 | 195.6174 | 213.5723 | 380.6252
2067 | 24.7314 | 159.82 | 196.2134 | 216.9871 | 202.6993 | 130.3394 | 167.9203 | 165.6067 | 152.6196 | 148.3007 | 117.0334 | 117.4396 | 152.2931 | 197.4601 | 133.2134 | 146.5064 | 216.2468 | 166.9074 | 142.4267 | 139.7866 | 173.4396 | 139.18 | 333.5733 | 366.3804 | 178.4267 | 197.82 | 213.1131 | 234.9742 | 328.5399 | 92.473 | 116.2468 | 112.8869 | 138.0334 | 124.8329 | 122.5064 | 146.2802 | 92 | 214.8998 | 247.6401 | 261.1465 | 281.5064 | 118.9203 | 115.473 | 125.6196 | 124.3136 | 91 | 59.5064 | 82.0334 | 113.1002 | 193.1131 | 172.6067 | 212.1671 | 148.6067 | 110.9666 | 128 | 144.1337 | 128.0797 | 172.3265 | 155.82 | 208.9332 | 161.2134 | 146.9666 | 129.82 | 158.0334 | 151.18 | 161.18 | 160.3933 | 226.473 | 256.2673 | 156.0129 | 133.3728 | 198.5399 | 220.82 | 204.5938 | 198.6067 | 196.1002 | 353.5656 | 150.2597 | 140.4936 | 154.7532 | 220.1671 | 136.8998 | 131.9332 | 131.0334 | 142.7069 | 171.5064 | 150.9666 | 161.82 | 107.473 | 111.9332 | 106.7403 | 111.7069 | 140.5064 | 228.6401 | 218.6196 | 240.8792 | 362.0797 | 232.3599 | 227.5527 | 241.7532 | 324.6325 | 145.2468 | 64.18 | 71.2597 | 145.2468 | 110.4396 | 97.9332 | 106.6735 | 113.9537 | 99.5733 | 97.4267 | 104.4601 | 106.1337 | 68.6401 | 74.3933 | 75.3136 | 98.4936 | 94.2468 | 104.4396 | 96.7532 | 81.4601 | 257.4473 | 217.6864 | 238.6067 | 382.7069 | 169.6401 | 180.653 | 189.5064 | 318.9203 | 295.2597 | 187.5064 | 209.2005 | 399.748
2068 | 24.7424 | 152.7271 | 202.301 | 209.329 | 209.7196 | 145.8878 | 149.2934 | 172.7066 | 168.5944 | 131.1402 | 115.2168 | 122.2654 | 162.8112 | 197.2729 | 141.7346 | 158.6505 | 217.2729 | 157.2524 | 141.6224 | 138.1607 | 187.5944 | 150.7346 | 328.699 | 338.2374 | 169.1047 | 202.4897 | 218.972 | 206.6785 | 327.5103 | 105.6224 | 116.4617 | 113.1122 | 133.7832 | 140.4056 | 124.1047 | 133.7271 | 89.1607 | 231.7066 | 239.9439 | 255.4542 | 279.4542 | 105.5663 | 131.8673 | 134.2934 | 119.6785 | 86.1327 | 72.4617 | 79.4056 | 108.8673 | 191.2654 | 175.2449 | 205.9234 | 152.8673 | 118.4617 | 128.4056 | 128.3495 | 125.2093 | 172.9159 | 152.7832 | 195.9364 | 153.5103 | 145.9439 | 128.4056 | 152.1607 | 148.5383 | 162.5944 | 165.8673 | 229.0766 | 239.301 | 149.8803 | 154 | 211.301 | 210.8878 | 190.2449 | 191.9159 | 186.5944 | 378.3981 | 161.3776 | 146.1963 | 155.7551 | 217.9795 | 129.7757 | 139.2168 | 133.2729 | 152.4692 | 182.8393 | 159.6785 | 167.7066 | 125.0841 | 119.2168 | 97.6505 | 109.301 | 133.9925 | 230.6785 | 243.9234 | 260.671 | 366.9159 | 236.4056 | 239.4542 | 246.8112 | 343.8037 | 146.2729 | 72.4897 | 79.9439 | 146.2729 | 116.0766 | 100.3495 | 104.8673 | 112.1683 | 103.2168 | 101.4897 | 98.1888 | 103.329 | 70.6785 | 72.972 | 65 | 92.4337 | 91.6224 | 108.8598 | 101.8112 | 74.3776 | 232.5663 | 215.8317 | 221.7757 | 375.028 | 157.0766 | 181.8317 | 193.9439 | 312.4617 | 312.4617 | 196.4056 | 206.4131 | 369.8748
2069 | 24.7534 | 149.5969 | 198.3735 | 207.1198 | 188.7759 | 137.7611 | 142.5969 | 172.7167 | 168.4179 | 142.6568 | 120.3432 | 108.4179 | 163.1642 | 199.1494 | 146.8062 | 171.1043 | 222.4926 | 179.821 | 141.9105 | 132.3284 | 186.1642 | 156.2241 | 328.8654 | 330.1938 | 154.4031 | 209.0747 | 213.9105 | 213.7463 | 327.6117 | 107.5821 | 124.2537 | 123.3284 | 133.6716 | 146.0148 | 136.1487 | 130.5969 | 93.358 | 237.1494 | 233.1642 | 252.1191 | 267.7611 | 108.2685 | 131.0599 | 127.179 | 120.4778 | 83.5969 | 70.2241 | 75.4031 | 109.7315 | 187.4475 | 185.8506 | 223.179 | 162.0895 | 124.1642 | 124.4031 | 128.0148 | 122.1339 | 173.1938 | 163.9549 | 175.5222 | 149.0148 | 136.6568 | 131.9253 | 153.0148 | 143.2537 | 149.0451 | 174.6716 | 210.9852 | 253.7611 | 138.5666 | 151.4926 | 225.3432 | 201.5074 | 198.3432 | 189.6864 | 193.1043 | 353.2389 | 164.4327 | 163.2389 | 145.9852 | 237 | 122.716 | 135.5673 | 142.6716 | 161.3587 | 180.3136 | 169.2537 | 176.9105 | 136.0895 | 120.5821 | 103.4179 | 114.5673 | 122.8654 | 252.3728 | 241.1198 | 240.7315 | 356.3284 | 249.537 | 228.179 | 240.8957 | 329.6864 | 148.9852 | 74.0599 | 67.7315 | 148.9852 | 108.4327 | 100.0148 | 109.0747 | 123.7315 | 106.2537 | 103.4778 | 101.179 | 119.0895 | 72.3136 | 72.5074 | 71.2685 | 89.3136 | 94.4179 | 97.3432 | 100.0747 | 77.4327 | 222.7315 | 201.3432 | 204.2685 | 378 | 147.7611 | 163.1642 | 186.7463 | 331.9549 | 313.5673 | 199.0895 | 228.5821 | 392.3136
2070 | 24.7644 | 156.4302 | 187.4642 | 211.6718 | 196.2679 | 152.034 | 162.7661 | 157.8604 | 173.3019 | 147.4642 | 124.5698 | 117.1736 | 164.5811 | 188.4189 | 146.9057 | 162.932 | 222.0113 | 166.6868 | 130.3962 | 145.4869 | 185.8604 | 150.4529 | 339.6981 | 354.185 | 156.1736 | 206.7208 | 220.8944 | 216.7321 | 334.5471 | 110.8717 | 126 | 111.5358 | 132.9887 | 143.1056 | 146.3509 | 134.4189 | 97.2566 | 241.9057 | 228.9887 | 264.2679 | 276.4415 | 121.3019 | 123.0113 | 118.8037 | 110.4302 | 82.6868 | 63.1623 | 73.7321 | 107.0227 | 202.4302 | 185.8264 | 214.3735 | 171.4529 | 122.5698 | 122.7321 | 127.2566 | 125.3509 | 170.5131 | 170.815 | 186.8604 | 162.0227 | 135.0113 | 132.1283 | 155.6981 | 140.2679 | 136.5925 | 176.5698 | 215.185 | 245.6641 | 160.5358 | 149.2906 | 216.6641 | 195.966 | 203.4302 | 194.6981 | 200.849 | 349.7208 | 168.5585 | 156.2792 | 147.1736 | 232.6981 | 135.6754 | 137.034 | 141.5585 | 136.5811 | 169.7094 | 163.6868 | 170.1283 | 136.849 | 120 | 100.5585 | 108.7208 | 127.2452 | 257.6641 | 222.0113 | 232 | 361.7434 | 254.5245 | 238.7321 | 242.185 | 339 | 140.7094 | 66.8717 | 68.034 | 140.7094 | 120.7321 | 116.034 | 107.5811 | 112.4189 | 107.5698 | 103.3246 | 104.849 | 122 | 70.3132 | 79.4415 | 76.1283 | 82.151 | 102.3132 | 100.7094 | 99.8717 | 82.4189 | 225.185 | 198.2566 | 222.034 | 381.0113 | 162.4642 | 174.4755 | 191.0227 | 326.7697 | 293.0943 | 206.7321 | 221.9773 | 396.7661
2071 | 24.7754 | 147.2654 | 197.7876 | 232.8406 | 187.7876 | 157.3628 | 165.9556 | 165.6372 | 179 | 158.6726 | 127.5399 | 124.9026 | 159.5928 | 189.4248 | 157.8053 | 151.832 | 224.6726 | 155.2301 | 125.6372 | 150.0707 | 196.7346 | 156.7876 | 339.3098 | 360.1504 | 175.8496 | 208.115 | 231.3451 | 227.8673 | 318.0353 | 112.9026 | 120.6902 | 105.1947 | 134.3098 | 129.115 | 130.062 | 125.2477 | 87.4425 | 243.0707 | 229.8673 | 252.6902 | 292.947 | 133.1947 | 131.4248 | 105.7523 | 123.832 | 76.9823 | 74.4248 | 76.4601 | 112.7876 | 198.5752 | 177.6549 | 202.5222 | 176.0177 | 130.8496 | 135.6372 | 113.4601 | 118.3542 | 159.7346 | 158.6549 | 195.0797 | 171.7699 | 142.5399 | 129.2124 | 145.5752 | 131.7876 | 150.1947 | 167.1504 | 215.8406 | 235.9647 | 160.3895 | 159.8496 | 214.9293 | 198.1594 | 198.6902 | 189.4425 | 204.1947 | 362.6196 | 156.4778 | 163.2921 | 161.5399 | 236.7346 | 125.9823 | 146.7876 | 129.4778 | 149.7346 | 166.2301 | 156.2124 | 163.6726 | 137.0974 | 115.5752 | 95.5575 | 123.3895 | 124.062 | 226.7256 | 240.1594 | 234.2124 | 379.0797 | 242.8673 | 233.938 | 261.4248 | 342.5399 | 142.5399 | 68.9026 | 74.2477 | 142.5399 | 123.0177 | 116.4955 | 114.0974 | 114.3098 | 105.2301 | 113.0177 | 109.0797 | 117.1327 | 80.885 | 77.3628 | 72.3274 | 93.8496 | 100.053 | 100.3274 | 104.5575 | 79.885 | 242.6549 | 195.5222 | 229.1327 | 376.5928 | 170.5752 | 180.3804 | 207.4072 | 307.062 | 293.2568 | 212.115 | 209.1327 | 407.4778
2072 | 24.7864 | 149.098 | 199.093 | 217.7286 | 194.5502 | 145.3618 | 149.2739 | 156.7186 | 174.907 | 149.2689 | 128.8166 | 113.2689 | 153.093 | 200.912 | 145.907 | 163.907 | 225.7286 | 157.093 | 143.0955 | 128.1809 | 205.8166 | 148.9975 | 341.4523 | 349.4548 | 176.5402 | 207 | 220.2689 | 219.902 | 305.2714 | 116.7311 | 115.3643 | 102.0855 | 141 | 135.7311 | 138.088 | 128.005 | 97.5502 | 233.4598 | 231.4523 | 239.1784 | 305.7261 | 143.7286 | 137.4548 | 125.2789 | 133.1784 | 76.5427 | 80 | 78.3668 | 113.6382 | 187.5427 | 178.2714 | 213.3643 | 170.5452 | 153.8241 | 136.2689 | 119.9145 | 132.6307 | 171.0905 | 161.0905 | 211.2764 | 159.902 | 144.2714 | 122.4498 | 137.7261 | 130.8191 | 152.9975 | 164.1859 | 204.3668 | 240.5427 | 161.902 | 151.8995 | 219.5377 | 219.1834 | 185.6332 | 195.4573 | 207.907 | 386.005 | 157.9145 | 161.3568 | 166.9095 | 248.0905 | 137.8216 | 139.907 | 131.8241 | 156.9975 | 161.7261 | 160.8191 | 167.4573 | 135.088 | 112.2739 | 97.7286 | 128.995 | 125.7211 | 213.1834 | 237.5352 | 237.4548 | 395.7311 | 258.0955 | 216.8166 | 261.9975 | 351.5477 | 156.0955 | 70.912 | 70.6382 | 156.0955 | 109.8141 | 96.088 | 110.2689 | 122.3643 | 102.5452 | 102.9975 | 107.5402 | 115.5477 | 78.3618 | 63.9975 | 79.912 | 102.2714 | 89.093 | 103.8191 | 102.6357 | 78.2714 | 243.7261 | 219.5527 | 225.2739 | 371.4573 | 169.093 | 179.2789 | 204.3568 | 336.4623 | 305.9045 | 204.6332 | 196.6332 | 398.9095
2073 | 24.7974 | 158.527 | 196.0599 | 208.3893 | 207.8683 | 145.6706 | 161.8083 | 149.5448 | 180.2754 | 135.7964 | 118.9282 | 109.6048 | 168.7425 | 200.994 | 144.2694 | 157.1317 | 233.6706 | 167.6048 | 141.8563 | 125.5329 | 214.1437 | 140.6647 | 338.8024 | 361.2095 | 171.006 | 201.8623 | 207.7306 | 207.2036 | 311.8083 | 120.3952 | 111.8623 | 102.0778 | 138.6647 | 149.2036 | 134.2694 | 132.0599 | 101.994 | 246 | 209.1857 | 255.2873 | 293.6587 | 139.527 | 138.4671 | 146.0718 | 125.4671 | 83.0778 | 71.5929 | 83.6647 | 123.6048 | 192.6766 | 189.946 | 212.6647 | 149.9163 | 153.054 | 133.5389 | 130.1976 | 123.2036 | 166.7306 | 159.5329 | 203.1857 | 157.0119 | 139.5988 | 115.2036 | 145.7425 | 135.8024 | 148.4012 | 180.0718 | 212 | 233.0659 | 151.0718 | 142.0778 | 220.6167 | 220.6647 | 193.8802 | 201.0659 | 188.9881 | 372.7786 | 170.5329 | 146.4012 | 167.0659 | 231.5869 | 149.7365 | 129.3952 | 133.3893 | 153.8024 | 166.006 | 170.006 | 167.4611 | 121.4611 | 109.8623 | 108.006 | 111.3952 | 118.006 | 220.7365 | 214.4012 | 242.2036 | 387.7186 | 242.8444 | 228.4131 | 247.5929 | 344.3893 | 163.2635 | 71.9282 | 70.7964 | 163.2635 | 101.4012 | 99.2754 | 105.6706 | 123.0659 | 100.5988 | 94.1976 | 100.1377 | 109.7905 | 71.1976 | 63.6048 | 74.8563 | 95.2635 | 105.2095 | 104.1317 | 98.6647 | 78.2694 | 245.6706 | 225.3234 | 218.1917 | 374.7306 | 160.4552 | 187.9282 | 186.1317 | 342.3893 | 308.5448 | 202.1377 | 193.2036 | 389.2575
2074 | 24.8084 | 155.2729 | 191.3142 | 200.711 | 207.1239 | 143.8097 | 154.7845 | 168.3555 | 184.8097 | 150.6535 | 118.2316 | 122.711 | 177.1652 | 200.1903 | 144.6445 | 156.9174 | 241.8761 | 181.1903 | 147.6535 | 132.1903 | 216.4955 | 158.6123 | 330.9748 | 371.6032 | 172.289 | 210.3806 | 217.9013 | 221.1077 | 313.8923 | 115.9174 | 117.0252 | 114.1239 | 138.3968 | 145.851 | 136.0826 | 138.8187 | 92.5619 | 233.0574 | 207.7361 | 278.6858 | 284.4381 | 125.2477 | 138.0413 | 155.8761 | 123.1239 | 90.8097 | 67.7523 | 74.289 | 133.8348 | 205.0413 | 187.8258 | 213.8348 | 141.3806 | 133.6032 | 136.6858 | 129.3968 | 122.2477 | 166.8348 | 159.4794 | 181.6032 | 183.0665 | 136.0826 | 126.711 | 156.5619 | 147.149 | 156.2316 | 180.289 | 213.4381 | 243.9839 | 158.4794 | 148.851 | 229.0987 | 220.3968 | 213.9174 | 201.4381 | 186.4219 | 364.6123 | 158.9748 | 151.3555 | 172.711 | 218.7935 | 143.9335 | 134.0252 | 132.9013 | 155.0826 | 170.6445 | 167.4542 | 161.9174 | 122.149 | 106.8761 | 116 | 103.5619 | 118.3303 | 235.0665 | 227.0252 | 231.6606 | 383.8187 | 230.2568 | 240.7271 | 234.6445 | 329.5206 | 151.8097 | 68.8348 | 73.1903 | 151.8097 | 104.9174 | 100.9335 | 101.4129 | 128.2316 | 100.9174 | 95.7935 | 106 | 111.86 | 71.3555 | 70.4794 | 64.8348 | 92.4381 | 104.0987 | 96.7271 | 93.6032 | 82 | 253.8761 | 216.7935 | 217.5458 | 359.9748 | 148.3555 | 177.6445 | 182.0826 | 339.5045 | 314.8923 | 205.1239 | 206.6284 | 373.6445
2075 | 24.8194 | 156.5752 | 205.8329 | 203.5752 | 197.6086 | 139.9332 | 148.7661 | 164.1337 | 177.4916 | 162.1671 | 120.0835 | 126.5585 | 167.5919 | 215.3747 | 135.5919 | 151.6253 | 224.3509 | 183.5919 | 153.759 | 130.6587 | 204.9833 | 166.7423 | 318.5084 | 356.2172 | 158.117 | 226 | 228.5752 | 255.5823 | 304.9833 | 108.1671 | 119.6587 | 111.9833 | 136.0835 | 141.4248 | 137.4415 | 138.7924 | 86.5752 | 218.5084 | 230.4415 | 251.8759 | 287.9666 | 127.8496 | 122.7423 | 138.8425 | 132.7828 | 86.9332 | 83.8329 | 75.8496 | 133.0835 | 198.5919 | 191.5823 | 212.1002 | 156.5084 | 125.1002 | 123.1504 | 126.5919 | 114.0334 | 155.2673 | 162.4582 | 192.7661 | 202.358 | 148.2577 | 130.5585 | 151.0668 | 149.117 | 145.7924 | 166.117 | 220.4081 | 244.2172 | 158.5084 | 148.358 | 215.9666 | 217.5919 | 207.1504 | 197.5919 | 198.5418 | 360.9428 | 153.3914 | 139.2673 | 178.0334 | 215.1504 | 126.117 | 144.5251 | 135.2172 | 148.0835 | 164.0501 | 155.5084 | 167.9332 | 122.642 | 118.358 | 109.117 | 108.3914 | 109.0167 | 244.5251 | 238.5084 | 221.883 | 408.8663 | 233.2673 | 230.5752 | 250.1742 | 346.6993 | 144 | 76.4415 | 76.5752 | 144 | 101.1002 | 110.1575 | 107.716 | 112.3843 | 103.4916 | 101 | 99.6086 | 125.5418 | 74.5084 | 74.9332 | 75.3914 | 101.3747 | 95.883 | 103.2911 | 95.9165 | 73.642 | 252.0835 | 198.8926 | 228.5084 | 353.4081 | 151.0167 | 185.3079 | 178.0334 | 336.2673 | 302.0501 | 189.7089 | 207.1504 | 396.0572
2076 | 24.8304 | 169.6378 | 202.3937 | 207.063 | 188.9843 | 143.4961 | 164.0079 | 164.5669 | 156.3307 | 150.4882 | 118.0236 | 119.4724 | 168.5512 | 215.9449 | 136.5512 | 148.5354 | 221.1417 | 165.9055 | 132.9606 | 129.063 | 193.4016 | 150.9921 | 329.0866 | 354.0866 | 172.6693 | 212.3937 | 213.9213 | 262.8583 | 308.5197 | 96.4882 | 104.9606 | 110.4803 | 137.0472 | 134.9134 | 144.0236 | 136.1181 | 91.0709 | 227.5748 | 248.1102 | 245.126 | 278.4095 | 137.5039 | 115.0551 | 125.0551 | 131.3858 | 78.9055 | 95.0079 | 82.9843 | 120.9449 | 185.9449 | 200.874 | 215.5748 | 163.5591 | 120.0079 | 123.5748 | 124.0236 | 115.0236 | 153.6142 | 156.937 | 207 | 194.3701 | 154.937 | 133.0472 | 155.5669 | 145.0236 | 139.5905 | 170.0866 | 231.0394 | 235.5354 | 154.9764 | 158.5118 | 212.4567 | 197.3858 | 205.5591 | 195.5276 | 196.5039 | 353.1496 | 161.0079 | 144.1653 | 174.9921 | 222.126 | 122.0236 | 148.0394 | 123.0079 | 153.0787 | 165.5354 | 163.5669 | 175.5276 | 119.5433 | 112.8898 | 103.0079 | 111.4724 | 120.5984 | 228.3858 | 243.0394 | 234.0394 | 403.8976 | 230.1024 | 225.4961 | 245.811 | 347.3622 | 152.5669 | 73.4488 | 70.4882 | 152.5669 | 99.0315 | 115.8819 | 116.937 | 102.1024 | 101.9843 | 105.5354 | 93.5039 | 125.5197 | 71.4803 | 77.9921 | 85.5276 | 98.4173 | 106.0236 | 118.5197 | 94.9528 | 73.5669 | 237.4252 | 194.1496 | 222.4567 | 348.9213 | 156.5512 | 180.3701 | 174.4882 | 333.1024 | 300.0079 | 197.6614 | 214.126 | 402.3386
2077 | 24.8414 | 161.7048 | 184.3539 | 212.4189 | 189.0649 | 149.1947 | 163.4513 | 177.6461 | 167.0061 | 146.4189 | 114.289 | 129.4219 | 176.0325 | 210.0325 | 140.9351 | 149.3864 | 226.0943 | 149.3539 | 124.3864 | 155.5842 | 197.0031 | 148.9675 | 327.6105 | 338.6724 | 176.4482 | 205.711 | 216.9057 | 242.9351 | 312.0325 | 104.8083 | 101.0325 | 112.1298 | 138.3539 | 140.5193 | 130.4807 | 155.6461 | 93.8053 | 238.5487 | 251.6754 | 269.2596 | 285.068 | 132.8377 | 122 | 138.1947 | 124.7434 | 78.357 | 86.1917 | 83.0649 | 115.5487 | 179.5162 | 184.4838 | 219.3215 | 171 | 132.357 | 141.2596 | 134.2272 | 118.5162 | 163.8702 | 152.0974 | 210.6136 | 178 | 158.357 | 146.2272 | 162.9675 | 141.8053 | 156.6785 | 179.9675 | 217.932 | 227.6105 | 157.6785 | 148.643 | 198.2241 | 190.7434 | 200.0943 | 191.2566 | 201.1298 | 383.357 | 155.8053 | 162.9351 | 163.1592 | 233.3539 | 138.9382 | 146.8053 | 133.2921 | 159.9026 | 169.4838 | 173.5487 | 180.5487 | 122.4189 | 109.3246 | 109.1623 | 118.3246 | 127.3215 | 218.1298 | 242.8377 | 249.8408 | 387.9026 | 231.643 | 215.1917 | 243.1653 | 335.1623 | 149.643 | 74.7434 | 72.6136 | 149.643 | 108.6785 | 107.711 | 112.6136 | 125.8408 | 103.6136 | 107.9351 | 97.6136 | 120.2566 | 64.3539 | 74.4189 | 80.2241 | 89.5487 | 105.3864 | 111.1917 | 94.6785 | 79.9351 | 235.7434 | 236.7465 | 228.8732 | 349.3246 | 157.8377 | 168.6461 | 181.2596 | 326.8996 | 309.7759 | 208.1917 | 218.643 | 376.8377
2078 | 24.8524 | 148.0419 | 182.1141 | 222.1559 | 194.6996 | 149.7148 | 144.0304 | 179.8859 | 195.4144 | 149.8137 | 118.5704 | 128.787 | 169.0722 | 193.5589 | 149.5704 | 140.7148 | 225.6274 | 155.0419 | 136.3271 | 150.8022 | 200.3726 | 159.6274 | 323.9278 | 317.6996 | 167.3993 | 212 | 225.3004 | 230.4296 | 313.5285 | 117.6996 | 99.8859 | 124.9848 | 129.2433 | 160.7567 | 131.3841 | 145.73 | 96.4563 | 229.9581 | 249.9278 | 261.1444 | 294.6007 | 126.943 | 126.2282 | 141.8859 | 136.7567 | 83.7148 | 75.943 | 81.8289 | 121.7567 | 190.0419 | 181.8859 | 218.2282 | 174.6996 | 136.6578 | 133.6729 | 139.4144 | 115.8289 | 168.9848 | 153.4144 | 206.0722 | 171.1293 | 155.787 | 141.3726 | 157.2433 | 129.6578 | 159.3574 | 193.27 | 204.6996 | 230.27 | 154.5437 | 148.0419 | 196.8707 | 206.4563 | 188 | 182.943 | 199.7148 | 372.3307 | 146.3004 | 148.8441 | 147.7148 | 233.2282 | 135.0874 | 133.6007 | 143.057 | 166.5133 | 168.4715 | 173.943 | 171.9581 | 133.213 | 106.8441 | 114.5285 | 111.6159 | 125.1711 | 242.6122 | 234.3004 | 242.5589 | 395.5704 | 241.0837 | 219.213 | 240.2737 | 348.9848 | 132.6578 | 74.6007 | 83.3993 | 132.6578 | 117.6996 | 101.8441 | 107.5437 | 128.0722 | 112.2852 | 114.9848 | 107.8707 | 116.1711 | 74.27 | 70.943 | 72.5285 | 88.3574 | 97.2433 | 94.6007 | 94.1863 | 73.7718 | 253.5704 | 233.1026 | 224.6729 | 342.0874 | 159.8707 | 178.8137 | 191.6426 | 346.5248 | 312.1863 | 205.3422 | 208.5285 | 399.483
2079 | 24.8634 | 155.3776 | 187.2449 | 229.7551 | 199.0816 | 152.5713 | 143.6835 | 166.1022 | 183.1838 | 149.5919 | 128.5408 | 122.4081 | 170.1122 | 198.01 | 155.7551 | 141.9489 | 227.8878 | 166.7041 | 147.2959 | 134.2449 | 196.4897 | 171.0816 | 331.5408 | 316.1327 | 171.2959 | 218.4897 | 201.99 | 226.9489 | 323.7346 | 116.6735 | 111.5203 | 122.7246 | 135.2754 | 162.2959 | 137.347 | 135.1632 | 104.5408 | 229.653 | 252.6735 | 233.4797 | 293.9489 | 135.1938 | 125.1327 | 141.0816 | 137.7551 | 85.4897 | 72.8368 | 79.5408 | 124.3776 | 199 | 165.3981 | 213.347 | 165.5614 | 129.3776 | 127.5713 | 134.7551 | 126.5203 | 166.7246 | 167.1427 | 195.7551 | 167.1632 | 141.8368 | 143.4386 | 151.9184 | 126.7041 | 162.9489 | 194.643 | 216.653 | 263.01 | 156.7346 | 152.1327 | 215.4386 | 215.6224 | 192.3265 | 178.7551 | 186.8878 | 356.7857 | 148.9489 | 155.3059 | 155.9794 | 262.4181 | 125.9489 | 132.9489 | 144.5408 | 163.643 | 180.9794 | 164.347 | 166.2449 | 121.2044 | 97.0816 | 113.3776 | 104.5713 | 131.2449 | 249.643 | 239.1122 | 222.1327 | 405 | 254.1327 | 221.8062 | 230.8162 | 354.2959 | 140.5203 | 69.6224 | 85.6735 | 140.5203 | 110.1838 | 100.7346 | 95.1327 | 118.2959 | 106.1838 | 109.4797 | 113.4592 | 110.8878 | 79.9695 | 67.2959 | 76.2449 | 88.7041 | 89.7551 | 95.0305 | 95.4897 | 70.5408 | 242.99 | 221.8978 | 221.2754 | 334.5713 | 151.9389 | 184.5408 | 192.3776 | 359.99 | 321.0611 | 206.6735 | 205.2143 | 405.0716
2080 | 24.8744 | 151.2345 | 185.0221 | 231.4248 | 201.1062 | 170.6151 | 158.1062 | 157.6593 | 171.7876 | 144.4469 | 128.4469 | 112.6194 | 165.938 | 206.1504 | 146.9159 | 164.6992 | 235.9337 | 180.062 | 138.9159 | 155.8054 | 202.5531 | 164.8097 | 314.3008 | 331.9116 | 173.4248 | 211.2787 | 199.9337 | 217.0663 | 304.898 | 114.6593 | 116.9159 | 109.5752 | 127.407 | 155.5752 | 129.4469 | 138.1062 | 90.6194 | 241.4248 | 241.2566 | 247.4427 | 299.5531 | 133.5973 | 117.6814 | 137.5752 | 129.469 | 83.8097 | 78.7434 | 72.2566 | 126.8717 | 196.2345 | 191.164 | 203.7876 | 168.2744 | 140.1682 | 121.8318 | 123.1504 | 130.8097 | 173.4868 | 163.407 | 205.7213 | 158.4911 | 137.2345 | 147.9159 | 157.0841 | 146.1461 | 163.5752 | 176.1504 | 202.4291 | 263.9601 | 160.0221 | 140.8097 | 214.3849 | 221.4248 | 189.3628 | 190.9337 | 182.2124 | 371.062 | 156.2124 | 171.1062 | 159.2566 | 260.2167 | 139.2965 | 136.3407 | 149.9779 | 162.2965 | 195.3186 | 168.6151 | 182.2744 | 119.2523 | 117.9116 | 111.4469 | 107.1283 | 122.938 | 233.3628 | 239.3628 | 224.6372 | 418.2744 | 248.8938 | 232.8054 | 242.2124 | 342.5973 | 150.8938 | 74.3186 | 82 | 150.8938 | 110.4027 | 107.3407 | 111.4647 | 129.8275 | 98.1062 | 101.3186 | 107.469 | 112.8496 | 76.7655 | 77.1682 | 77.8938 | 94.8717 | 97.5089 | 99.3407 | 110.9558 | 77.6372 | 238.1682 | 224.8097 | 233.3186 | 342.6593 | 151.6151 | 179.469 | 202.6151 | 339.1283 | 323.0442 | 207.4248 | 200.8938 | 420.5268
2081 | 24.8854 | 163.7002 | 194.5694 | 205.5996 | 202.5654 | 168.6922 | 156.7384 | 147.6922 | 156.4306 | 149.6539 | 119.5191 | 114.5694 | 147.9537 | 203.5231 | 143.1308 | 150.165 | 241.7807 | 170.3421 | 136.2616 | 161.3883 | 201.3038 | 164.6539 | 307.9155 | 334.4306 | 179.827 | 207.2193 | 197.8652 | 219.1348 | 297.1348 | 114.3038 | 102.9537 | 106 | 123.0463 | 149.173 | 124.4769 | 129.3883 | 79 | 239.9115 | 236.1308 | 256.1268 | 277.9497 | 122.6076 | 131.3964 | 132.3038 | 123.8692 | 87.6117 | 74.8229 | 70.5231 | 127.173 | 183.2575 | 212.9537 | 203.6962 | 168.8229 | 139.2575 | 119.3501 | 123.9155 | 131.2193 | 165.0805 | 162.4386 | 208.6499 | 157.3501 | 133.8692 | 134.5191 | 156.3461 | 161.1308 | 157.173 | 174.6539 | 195.1771 | 256.0885 | 143.5614 | 144.0463 | 209.5231 | 198.992 | 203.7887 | 195.6499 | 190.7847 | 372.0845 | 156.8692 | 156.169 | 152.4346 | 233.0845 | 143.7384 | 132.7384 | 162.4809 | 154.8652 | 191.2153 | 183.6539 | 183.3883 | 136.2616 | 126.6499 | 105.3461 | 105.1308 | 121.9155 | 231.3924 | 243.0463 | 235.0885 | 408.0805 | 241.7807 | 228.2113 | 253.0463 | 343.4809 | 163.5694 | 62.8652 | 82.5654 | 163.5694 | 121.9577 | 108.2616 | 124.6076 | 136.4769 | 104.0885 | 111.3501 | 106.3924 | 106.996 | 64.2998 | 87.5654 | 73.6076 | 99.6962 | 96.9537 | 90.6499 | 103.169 | 73.3883 | 253.6539 | 232.004 | 231.4769 | 394.8853 | 165.5231 | 185.1771 | 207.4769 | 333.173 | 317.2616 | 194.6036 | 206.2193 | 430.169
2082 | 24.8964 | 165.7571 | 190.5579 | 206.1073 | 191.4463 | 151.3348 | 150.3785 | 159.2189 | 152.3545 | 148.2232 | 113.5777 | 122.6893 | 154.8644 | 219.1313 | 143.4223 | 144.1751 | 232.9562 | 177.3743 | 129.3348 | 149.3785 | 205.7768 | 176.5099 | 326.1313 | 321.6893 | 185.4661 | 209.6893 | 210.685 | 219.6017 | 320.709 | 108.3785 | 116.7966 | 99.6455 | 152.5296 | 141.8009 | 126.1991 | 135.8644 | 94.0198 | 220.4026 | 250.8644 | 254.5099 | 277.7528 | 109.6017 | 140.5339 | 131 | 125.8884 | 85.7571 | 70.4661 | 71.6893 | 131.3545 | 182.5099 | 200.2232 | 204.4223 | 169.0876 | 146.0198 | 124.4223 | 124.8009 | 134.267 | 147.8446 | 180.0876 | 195.4901 | 164.1554 | 131.267 | 120.4901 | 153.1554 | 157.3785 | 165.3983 | 185.9322 | 212.8206 | 262.3107 | 143.8206 | 134.8249 | 212.4223 | 193.4421 | 211.4901 | 183.6455 | 196.5777 | 368.8884 | 153.6893 | 141.1116 | 145.6455 | 240.2867 | 154.709 | 135.0438 | 158.024 | 149.7768 | 177.9124 | 164.8927 | 176.5777 | 147.8206 | 123.3107 | 101.5777 | 108.3107 | 121.0678 | 257.1073 | 245.9562 | 258.6412 | 407.5974 | 230.0678 | 226.2867 | 243.8249 | 352.3107 | 152.6257 | 66.4421 | 64.5141 | 152.6257 | 125 | 111.1554 | 115.6455 | 128.3785 | 110.8884 | 98.5141 | 98.024 | 110.8644 | 59.9322 | 75.8687 | 75.6215 | 105.6215 | 93.4661 | 93.0876 | 93.3107 | 69.4661 | 251.6455 | 226.4026 | 239.5537 | 402.8051 | 159.7571 | 186.0678 | 193.024 | 335.6215 | 319.5777 | 195.2867 | 205.2232 | 418.5777
2083 | 24.9074 | 181.419 | 192.3895 | 218.8702 | 181.8201 | 147.9499 | 159.3895 | 153.3509 | 162.8997 | 155.2095 | 122.2596 | 114.5103 | 156.1504 | 206.171 | 145.3599 | 144.2005 | 227.0501 | 175.3008 | 135.9794 | 138.3303 | 204.1003 | 169.6105 | 324.9704 | 330.6195 | 200.3895 | 224.5193 | 221.5103 | 227.9293 | 332.3599 | 123.8792 | 121.9704 | 116.829 | 160.1504 | 143.8997 | 131.7699 | 142.4602 | 98.5103 | 237.3188 | 241.5308 | 243.4807 | 292.9499 | 117.3393 | 126.7905 | 130.41 | 133.3098 | 86.0797 | 83.0296 | 70.59 | 121.8406 | 194.4897 | 190.6902 | 190.4306 | 172.6401 | 150.5103 | 118.1003 | 131.0296 | 122.3805 | 150.5398 | 167.1208 | 198.8496 | 171.4897 | 142.3895 | 138.009 | 159.8997 | 145.7404 | 152.9409 | 175.0707 | 220 | 261.0501 | 153.3599 | 130.6697 | 216.7198 | 226.419 | 210.7198 | 214.3985 | 196.41 | 349.7609 | 163.7995 | 139 | 150.4396 | 257.0797 | 162.2301 | 130.9203 | 140.5604 | 148.1003 | 164.9203 | 185.1684 | 185.2596 | 150.8702 | 117.9203 | 103.1799 | 107.6401 | 123.6697 | 257.4807 | 236.5103 | 248.8111 | 408.3805 | 225.59 | 224.7905 | 237.3098 | 349.8702 | 152.3393 | 68.7404 | 64.5694 | 152.3393 | 120.8702 | 116.1298 | 101.5604 | 114.9704 | 123.6195 | 90.3098 | 94.1298 | 113.9203 | 65 | 63.4602 | 83.1298 | 100.1504 | 111.9293 | 100.7699 | 99.7198 | 79.6697 | 244.0501 | 211.4602 | 243.8702 | 377.0501 | 175.419 | 184.5398 | 181.4602 | 338.41 | 327.0797 | 199.6902 | 193.9203 | 401.3008
2084 | 24.9184 | 172.9094 | 193.7728 | 221.4204 | 197.2612 | 149.3977 | 160.1706 | 153.4544 | 180.2498 | 149.1479 | 130.4091 | 106.9887 | 154.2159 | 199.4317 | 138.5683 | 152.8634 | 232.2272 | 161.7955 | 130.5456 | 137.8181 | 207.2272 | 162.2045 | 310.7728 | 325.9547 | 188.523 | 224.557 | 189.2958 | 239 | 328.5796 | 126.1592 | 124.6362 | 128.9887 | 137.7388 | 153.4204 | 134.2045 | 144.2159 | 97.6136 | 246.7615 | 247.2725 | 234.4091 | 287.7728 | 113.5456 | 131.6476 | 125.1819 | 138.8068 | 78.9547 | 85.7841 | 69.1932 | 107.1819 | 199.6023 | 193.0227 | 197.2612 | 179.4317 | 124.3184 | 110.9887 | 131.9773 | 137.2838 | 168.0566 | 185.3178 | 210.4204 | 171.1819 | 144.9773 | 146.7728 | 159.1819 | 147.2272 | 155.2612 | 186.8861 | 210.966 | 257.7955 | 151.3864 | 143.8294 | 218.1932 | 219.7162 | 197.1366 | 201.0573 | 183.3524 | 356.6816 | 164.1706 | 143.2159 | 150.9773 | 250.557 | 144.1366 | 130.8181 | 130.3864 | 148.2159 | 179.2725 | 183.2958 | 196.4204 | 145.5909 | 106.9773 | 109.4204 | 104.7955 | 141.0453 | 242.9887 | 248.2612 | 234.4091 | 417.8634 | 248.2838 | 218.8068 | 245.2159 | 348.2045 | 138.9094 | 70.8294 | 69.1819 | 138.9094 | 110.7728 | 110.5683 | 99.8181 | 109.8068 | 115.3411 | 105.443 | 95.7955 | 117.4317 | 68.0113 | 68.8294 | 78.1706 | 92.7955 | 110.9547 | 111.034 | 101.7955 | 73.557 | 263.6816 | 245.7382 | 245.2159 | 385.8408 | 183.1706 | 177.3638 | 191.6476 | 341.6136 | 338.6249 | 195.3977 | 210.0793 | 407.6702
2085 | 24.9294 | 150.8544 | 203.1348 | 223.7709 | 193.2506 | 153.3019 | 159.3019 | 158.6981 | 184.6981 | 145.531 | 128.9272 | 118.5202 | 165.6038 | 215.4475 | 141.6038 | 145.5633 | 228.3962 | 172.6766 | 141.8949 | 144.0728 | 206.469 | 154.3962 | 303.5418 | 319.8437 | 174.3854 | 217.6146 | 193.7386 | 228.5525 | 340.9785 | 125.7601 | 129.7709 | 130.0728 | 133.7601 | 146.5525 | 139.3019 | 141.469 | 111.062 | 227.0943 | 226.1994 | 234.7709 | 291.6038 | 121.8221 | 135.0835 | 127.531 | 132.0108 | 76.531 | 72.5525 | 68.6146 | 113.2183 | 196.9272 | 193.3127 | 199.3962 | 172.0943 | 129.6658 | 125.593 | 120.0108 | 131.7196 | 165.7089 | 185.1051 | 196.7924 | 171.0728 | 148.3747 | 124.1779 | 152.3127 | 146.469 | 159.2399 | 181.1051 | 201.9272 | 252.0835 | 142.8544 | 153.3019 | 226.2183 | 212.9892 | 200.7494 | 196.2804 | 195.8949 | 393.124 | 154.6981 | 141.6981 | 149.4582 | 254.6766 | 138.531 | 133.3854 | 152.5823 | 147.9272 | 188.9272 | 176.6766 | 172.9595 | 137.2399 | 109.7601 | 113.6146 | 105.2291 | 137.3234 | 260.6658 | 266.3747 | 237.8437 | 420.3962 | 266.6873 | 230.4475 | 237.5525 | 345.3127 | 123.6146 | 70.469 | 77.6766 | 123.6146 | 101.0835 | 107.4582 | 101.1563 | 109.7709 | 105.6146 | 119.7601 | 109.1348 | 123.6146 | 74.9165 | 80.7601 | 69.9272 | 87.6981 | 95.6253 | 110.8544 | 98.5418 | 81.9785 | 269.3962 | 220.8355 | 234.4798 | 388.0835 | 174.9272 | 161.1671 | 189.5525 | 334.7817 | 347.9165 | 213.4367 | 208.0215 | 397.2614
2086 | 24.9404 | 155.7759 | 205.7315 | 229.8953 | 178.3584 | 154.7463 | 168.8953 | 163.5222 | 187.0148 | 147.1195 | 132.0148 | 125.1195 | 173.5074 | 227.0148 | 148.2537 | 132.7463 | 238.6712 | 189.4027 | 166.2833 | 150.3879 | 202.3732 | 172.1934 | 318.9249 | 349.2081 | 181.5222 | 221.7611 | 202.3584 | 201.9409 | 355.3879 | 130 | 115.2094 | 128.2389 | 134.2389 | 136.2389 | 143.2537 | 125.4631 | 108.5973 | 227.149 | 243.8645 | 249.6712 | 289.4778 | 131.7463 | 118.2094 | 124.7315 | 100.0456 | 84.3879 | 71.6416 | 76.5222 | 125.2685 | 196.2537 | 180.9704 | 205.2833 | 175.2833 | 140.1195 | 135.3732 | 126.9101 | 115.2389 | 163.6416 | 171.3732 | 196.6564 | 170.4926 | 147.3584 | 120.5222 | 153.7611 | 161.1786 | 149.3584 | 165.4926 | 224.447 | 260.9101 | 142.1342 | 135.314 | 220.7167 | 234.298 | 203.1047 | 206.4926 | 192.0751 | 381.2993 | 152.6268 | 152.7906 | 155.3879 | 249.4631 | 150.7759 | 140.5222 | 155.4631 | 164.1786 | 175.7167 | 173.3436 | 172.298 | 138.6416 | 98.9557 | 123.4027 | 105.3732 | 134.3879 | 262.3436 | 245.2993 | 242.7611 | 433.8054 | 253.3288 | 235.1195 | 251.0591 | 354.2833 | 136.5369 | 70.1342 | 68.702 | 136.5369 | 101.1342 | 102.1047 | 113.7906 | 126.5517 | 119.1638 | 103.314 | 101.702 | 124 | 72.3584 | 73.0899 | 76.7759 | 98.7906 | 89.1195 | 100.7315 | 100.1342 | 91.3732 | 267.7611 | 210.0591 | 239.1638 | 378.1047 | 168.6121 | 173.8054 | 181.7463 | 335.2685 | 326.553 | 219.9852 | 206.7759 | 397.4175
2087 | 24.9514 | 162.9174 | 192.4129 | 230.1652 | 188.4219 | 139.939 | 165.9695 | 164.1652 | 177.2172 | 151.1131 | 137.5566 | 131.0305 | 181.3914 | 203.1562 | 143.8869 | 155.009 | 240.3303 | 178.3824 | 169.8043 | 147.8869 | 198.8043 | 164.9084 | 305.991 | 327.0126 | 164.9084 | 229.7522 | 201.5566 | 209.1742 | 349.0521 | 127.4434 | 107 | 108.1041 | 147.3393 | 139.7522 | 124.1867 | 130.1436 | 102.3609 | 252.4524 | 239.517 | 239.8258 | 295.8654 | 134.1957 | 127.8959 | 126.7522 | 108.9264 | 87 | 85.2262 | 68.8564 | 130.9174 | 207.2262 | 183.3088 | 225.4219 | 186.474 | 126.8564 | 133.0826 | 122.4955 | 124.5045 | 168.2783 | 156.2998 | 217.7002 | 178.5871 | 129.939 | 124.3609 | 152.8043 | 174.5566 | 161.3393 | 178.061 | 222.3824 | 275.0305 | 149.1131 | 146.6481 | 229.3393 | 256.1436 | 189.4129 | 213.3088 | 194.7828 | 371.1131 | 154.9174 | 166.1131 | 161.8348 | 270.7612 | 146.4129 | 149.474 | 137.1346 | 164.7738 | 161.3303 | 176.5871 | 167.2998 | 145.8348 | 114.9264 | 117.4955 | 100.526 | 138.9174 | 263.6697 | 235.7522 | 220.7127 | 437.3303 | 242.7217 | 219.2998 | 236.1562 | 348.8564 | 147.8348 | 84.7828 | 69.2262 | 147.8348 | 110.0305 | 106.9479 | 111.1346 | 126.1346 | 112.2998 | 97.3914 | 99.6697 | 118.8869 | 79.8654 | 71.1131 | 81.3609 | 114.0305 | 86.0826 | 96.3609 | 112.2262 | 80.1346 | 243.1562 | 232.2262 | 255.9479 | 380.3914 | 161.526 | 193.9479 | 182.9174 | 337.0826 | 323.5045 | 225.3088 | 224.7828 | 411.1131
2088 | 24.9624 | 154.88 | 186.3486 | 240.3772 | 205.7714 | 145.0286 | 153.5314 | 152.88 | 166.7429 | 165.0743 | 128.5771 | 122.6228 | 176.5314 | 206.8915 | 143.6057 | 145.8514 | 236.6514 | 174.8629 | 154.32 | 152.8172 | 204.8172 | 164.7257 | 306.0286 | 351.9029 | 175.8 | 219.9714 | 203.6514 | 212.1828 | 349.8629 | 116.2286 | 121.3314 | 114.9372 | 151.44 | 141.6971 | 122.12 | 135.6971 | 101.3486 | 271.1657 | 259.1772 | 237.4686 | 307.2114 | 134.0457 | 113.2457 | 139.7714 | 127.56 | 81.1371 | 76.6686 | 72.4229 | 130.0457 | 193.4571 | 195.1657 | 230.3943 | 181.8343 | 131.0743 | 135.9086 | 123.1657 | 134.2571 | 154.6686 | 168.8457 | 240.9829 | 174.88 | 152.6172 | 134.4229 | 165.3314 | 151.2457 | 178.4686 | 169.0628 | 187.5942 | 276.3943 | 157.2114 | 148.9257 | 238 | 245.4114 | 217.8743 | 200.4114 | 198.7429 | 369.44 | 176.1943 | 139.0342 | 147.0628 | 244.0057 | 141.6514 | 145.4857 | 152.4972 | 162.9086 | 171.3314 | 174.8343 | 159.6514 | 134.32 | 116.6228 | 115.56 | 119.4972 | 129.5771 | 255.6228 | 229.88 | 250.6458 | 404.3371 | 255.68 | 220.7714 | 243.8 | 342 | 148.0457 | 77.6686 | 71.7429 | 148.0457 | 114.6514 | 125.0286 | 114.12 | 121.3029 | 109.8629 | 104.2571 | 118.3314 | 114.0457 | 81.44 | 83.12 | 78.3943 | 101.0628 | 100.6343 | 91.44 | 107.5771 | 78.56 | 226.6971 | 236.0457 | 244.0628 | 393.12 | 156.3943 | 191.8343 | 189.8629 | 332.0914 | 310.4571 | 218.9257 | 221.5771 | 412.0457
2089 | 24.9734 | 152.6549 | 178.6991 | 233.7256 | 207.0177 | 151.3363 | 147.6726 | 153.9735 | 169.646 | 182.2833 | 121.0088 | 112.0177 | 173.3274 | 223.6372 | 162.9205 | 153.2479 | 238.9912 | 173.354 | 140.3628 | 158.9912 | 197.7167 | 165.0265 | 322.2921 | 331.2033 | 168.7521 | 205.6991 | 215.947 | 225.2567 | 343.0442 | 132.9028 | 109.7521 | 117.3628 | 139.7079 | 144.3186 | 127.6637 | 144.2921 | 104.3186 | 249.7786 | 231.5396 | 253.947 | 304.6902 | 131.6726 | 117.9205 | 136.3716 | 118.053 | 66.053 | 81.6107 | 83.9735 | 119.7079 | 200.9205 | 189.0442 | 224.354 | 171.3628 | 120.407 | 138.6637 | 119.0354 | 149.2744 | 154.9646 | 165.3981 | 233.7344 | 179.2921 | 147.7609 | 128.7167 | 149.1061 | 146.6284 | 179.0177 | 169.2921 | 204.5311 | 247.7875 | 157.3451 | 142.3363 | 230.0354 | 211.4423 | 229.3628 | 204.2744 | 184.3893 | 368.3274 | 163.7698 | 144.2391 | 165.2125 | 244.8939 | 135.3628 | 153.947 | 140.7698 | 171.6372 | 177.0088 | 168.3451 | 167.3009 | 135.6284 | 108.0088 | 122.646 | 113.7433 | 120.0177 | 243.6902 | 254.8674 | 250.1061 | 402.2744 | 273.6195 | 231.3098 | 242.7256 | 369.2125 | 172.8851 | 68.6726 | 70 | 172.8851 | 129.6019 | 121.3805 | 120.3274 | 131.9558 | 103.7079 | 106 | 117.3716 | 116.3186 | 75.0177 | 70.7433 | 61.0707 | 99.3009 | 94.3981 | 103.6019 | 111.2921 | 81 | 245.9116 | 235.6637 | 238.3186 | 385.3893 | 150.354 | 195.9646 | 185.6991 | 322.6991 | 323.8939 | 222.2921 | 216.6637 | 414.3186
2090 | 24.9844 | 158.7321 | 198.6604 | 237.1402 | 201.6199 | 151 | 131.4517 | 159.704 | 174.028 | 173.1277 | 126.0841 | 112.704 | 163.1838 | 221.5919 | 152.3957 | 146.0997 | 224.4517 | 157.4798 | 143.7601 | 158.648 | 203.4922 | 165.3801 | 346.6043 | 316.7882 | 156.972 | 211.4642 | 223.352 | 229.9439 | 328.5358 | 123.0436 | 106.0841 | 103.1838 | 142.4361 | 136.5358 | 142.1962 | 130.0717 | 98.5639 | 263.3645 | 232.3364 | 250.5358 | 299.972 | 126.2679 | 138.4922 | 126.5919 | 128.9003 | 72.1682 | 82.5919 | 85.648 | 130.5483 | 207.296 | 193.4642 | 222 | 178.8162 | 130.2523 | 137.648 | 119.0561 | 125.5795 | 165.0841 | 153.2679 | 211.1277 | 186.704 | 149.8442 | 131.7882 | 175.5327 | 145.5919 | 160.7757 | 173.324 | 213.2399 | 256.3084 | 164.1121 | 159.5763 | 224.648 | 218.6043 | 204.3676 | 193.4237 | 199.6324 | 360.8879 | 164.1682 | 151.6199 | 164.8038 | 265.7882 | 137.4081 | 151.2118 | 133.7321 | 178.3801 | 176.676 | 182.5483 | 175.7321 | 141.352 | 115.1121 | 106.7477 | 108.3801 | 122.7321 | 231.5358 | 253.0997 | 250.8442 | 424.5483 | 267.5078 | 223.8598 | 210.3116 | 374.2118 | 151.5234 | 73.4081 | 85.5483 | 151.5234 | 132.2679 | 114.648 | 114.9159 | 126.1838 | 98.324 | 101.2679 | 101.5078 | 117.324 | 77.7321 | 64.704 | 77.3364 | 96.2399 | 97.8162 | 89.3676 | 108.5639 | 83.704 | 246.5358 | 229.9159 | 220.3957 | 365.4798 | 152.7321 | 179.7197 | 199.5763 | 341.9844 | 350.8723 | 231.7321 | 204.1558 | 382.8754
2091 | 24.9954 | 160.3117 | 203.8053 | 213.7796 | 199.3117 | 139.9872 | 141.8959 | 160.3117 | 181.883 | 163.2468 | 123.4936 | 120.883 | 166.9479 | 210.0521 | 140.2468 | 145.5713 | 222.5064 | 183.3502 | 144.9351 | 140.7924 | 204.8702 | 165.6234 | 357.3766 | 318.2468 | 167.7011 | 220.1298 | 225.3766 | 220.4287 | 343.9608 | 125.3894 | 128.2725 | 104.1947 | 146 | 147.1427 | 132.4807 | 133.3894 | 81.9222 | 250.5328 | 239.8702 | 295.5578 | 301.0649 | 139.831 | 131.6106 | 130.1947 | 154.8959 | 89.7011 | 84.1298 | 91.883 | 120.7924 | 219.0778 | 198.6883 | 226.1298 | 168.169 | 130.7404 | 146.6362 | 119.6234 | 122.7011 | 179.7011 | 153.0649 | 217.0778 | 188 | 156.6883 | 159.4023 | 168.5328 | 152.6362 | 168.1427 | 179.883 | 210.6883 | 261.4936 | 166.6234 | 156.2989 | 211.6106 | 221.117 | 208.4544 | 185.6883 | 216.1947 | 390.7268 | 152.7924 | 141.0521 | 173.831 | 247.9743 | 141.3766 | 161.4544 | 136.6883 | 180 | 185.9479 | 201.7011 | 184.883 | 122.0392 | 117.6234 | 111.7661 | 108.6234 | 126.3766 | 229.7532 | 236.7404 | 241.169 | 447.1427 | 270.2596 | 227.9479 | 218.6491 | 370.6883 | 143.2596 | 71.1819 | 72.3509 | 143.2596 | 116.2339 | 108.4936 | 105.117 | 136.831 | 132.4151 | 90.0521 | 103.5713 | 110.117 | 75.1819 | 79.0778 | 89 | 97.8181 | 112.6362 | 98.2725 | 108.4415 | 78.117 | 270.2204 | 218.7404 | 232.3374 | 359.6883 | 143.9872 | 192.0257 | 192.8573 | 363.3245 | 330.4679 | 216.1041 | 198.6883 | 422.0642
2092 | 25.0064 | 151.5929 | 209.4071 | 213.6107 | 203.2036 | 136.4012 | 150 | 175.4131 | 178.6947 | 164.1018 | 146.2214 | 124 | 175.2036 | 194.0899 | 154.4131 | 166.5148 | 225.7006 | 184.6887 | 134.8923 | 139.3053 | 205.1018 | 169.2036 | 337.6828 | 344.3232 | 173 | 206.5869 | 235.1077 | 195.9822 | 351.5988 | 142.9101 | 110.3774 | 121.0119 | 140.3952 | 158.2036 | 135.5089 | 137.5988 | 90.7125 | 236.1976 | 229.5929 | 282.2697 | 309.7065 | 126.6828 | 135.8083 | 140.0059 | 133.5751 | 80.9881 | 76.8923 | 71.8804 | 124.2095 | 210.2875 | 191.2935 | 217.4911 | 168.7065 | 142.4131 | 132.084 | 120.4012 | 137.1077 | 182.8982 | 166.6107 | 203.3834 | 198.5089 | 160.503 | 141.9762 | 163.3053 | 167.5089 | 175.7006 | 186.503 | 194.1857 | 247.0899 | 173.7065 | 154.503 | 213.0059 | 206.0899 | 199.5869 | 201.4131 | 193.7786 | 377.2757 | 156.2095 | 153.8143 | 162.084 | 256.916 | 125.8863 | 142.078 | 134.8982 | 170.8923 | 173.8863 | 188.084 | 198.5089 | 115.8024 | 109.9941 | 111.6947 | 104.497 | 117.8923 | 235.9042 | 263.8261 | 230.497 | 456.1018 | 264.1917 | 215.8863 | 212.5869 | 365.3952 | 147 | 73.9042 | 54.5929 | 147 | 119.1077 | 104.5988 | 116.0119 | 134.1917 | 124.1798 | 104.916 | 109.1018 | 114.0059 | 84.2095 | 79.3952 | 71.4852 | 108.4071 | 107.8923 | 105.5988 | 111.4012 | 84.1077 | 278.0958 | 207.2935 | 216.078 | 357.1976 | 162.8202 | 202 | 186.7006 | 350.4852 | 318 | 219.2095 | 203.9042 | 391.6531
2093 | 25.0174 | 143.7227 | 195.8908 | 227.5546 | 212.1288 | 144.8417 | 164.2577 | 181.2871 | 166.7325 | 155.7325 | 145.5938 | 116.1583 | 192.6835 | 213.951 | 143.1779 | 146.9104 | 228.1387 | 183.2773 | 125.2969 | 146.9902 | 217.4062 | 175.2773 | 324.7227 | 337.465 | 164.4454 | 197.8613 | 221.8908 | 187 | 348.1485 | 121.6233 | 110.4062 | 118.4454 | 133.7227 | 151.4454 | 124.3165 | 130.584 | 92.7227 | 263.5154 | 234.5546 | 262.7227 | 303.0196 | 124.416 | 147.1288 | 130.1681 | 122.4258 | 79.9902 | 76.5644 | 77.6835 | 126.8613 | 212.5546 | 180.1583 | 221.5546 | 169.1485 | 136.1681 | 141.1092 | 131.6933 | 128.1681 | 182.7129 | 152.7521 | 212.8221 | 187.3165 | 144.8908 | 136.416 | 180.9706 | 175.5644 | 151.7619 | 192.9902 | 204.1092 | 258.3964 | 172.7227 | 148.8712 | 228.8319 | 225.951 | 198.7031 | 197.3067 | 216.5056 | 378.5448 | 156.0098 | 151.7325 | 156.8515 | 271.416 | 149.654 | 148.8221 | 143.9804 | 167.7129 | 180.5448 | 192.119 | 189.4552 | 114.8613 | 104.8613 | 114.7031 | 112.2675 | 132.535 | 242.9902 | 267.7325 | 219.7325 | 427.7717 | 226.4944 | 236.0896 | 243.7829 | 335.9104 | 149.1387 | 73.1485 | 81.654 | 149.1387 | 117.2969 | 110.416 | 113.4454 | 115.0294 | 111.8613 | 115.1387 | 110 | 105.5938 | 78.3067 | 67.0196 | 70.416 | 113.4258 | 102.5742 | 98.584 | 109.8613 | 83.7227 | 243.2073 | 227.5252 | 226.3865 | 338.8908 | 161.5938 | 199.1485 | 177.7325 | 355.8319 | 316.5742 | 225.4258 | 223.8221 | 406.4958
2094 | 25.0284 | 164.4805 | 189 | 228.8245 | 185.2678 | 143.6489 | 175.0762 | 168.672 | 163 | 168.679 | 130.1224 | 113 | 188.8475 | 222.5496 | 133.8245 | 138.9007 | 243.5035 | 184.2748 | 141.1294 | 149.7252 | 206.7713 | 172.6489 | 339.679 | 339.4273 | 176.2287 | 217.3049 | 210.6489 | 215.2819 | 325.9699 | 111.7252 | 118.6259 | 129.5035 | 151.5797 | 142.1986 | 137.5797 | 122.1986 | 95.3511 | 248.1685 | 235.8245 | 271.8776 | 288.1224 | 125.5496 | 122.4433 | 119.1986 | 136.7783 | 78.3741 | 70.7482 | 70.9468 | 130.3511 | 190.6188 | 205.2819 | 218.4734 | 199.9077 | 117.2217 | 140.0231 | 130.1986 | 133.1525 | 180.8245 | 156.6028 | 217.0993 | 198.4042 | 143.8014 | 154.2287 | 165.9699 | 160.321 | 167.3812 | 184.1224 | 215.3511 | 290.3812 | 171 | 152.5266 | 215.8706 | 243.2518 | 211.1525 | 180.672 | 199.5426 | 378.9238 | 168.5035 | 165.4042 | 157.2748 | 265.2979 | 140.9699 | 150.1986 | 146.5496 | 171.6259 | 183.8245 | 194.0993 | 194.1525 | 117.6259 | 102.5496 | 106.8475 | 114.5496 | 122.5958 | 264.5797 | 241.5195 | 239.9308 | 421.0762 | 247.0833 | 235.3972 | 257.5496 | 369.9609 | 160.1525 | 67.6489 | 76.5958 | 160.1525 | 112.8245 | 119.5266 | 115.0762 | 132.9308 | 134.2056 | 106.5727 | 104.9238 | 109.7021 | 91.4042 | 68.8014 | 68.6489 | 106.0231 | 99.8245 | 102.5266 | 102.4734 | 72.5727 | 249.5797 | 228.2979 | 226.2979 | 351.5797 | 138.1454 | 184.2217 | 205.9077 | 347.9468 | 334.1294 | 202.0692 | 208.5195 | 396.6188
2095 | 25.0394 | 164.1504 | 209.6491 | 209.5633 | 211.6108 | 164.8615 | 165.938 | 162.5251 | 165.2124 | 177.2124 | 140.0119 | 122.5871 | 176.8879 | 230.8496 | 138.1623 | 151.7995 | 242.3628 | 181.0501 | 160.5369 | 132.3008 | 207.6372 | 190.1741 | 346 | 310.5515 | 179.0501 | 219.8377 | 246.6108 | 230.4248 | 315.7876 | 125.2744 | 111.8879 | 125.4129 | 153.8377 | 142.2124 | 147.2124 | 140.6491 | 105.1121 | 252.7493 | 227.6253 | 265.6755 | 295.7995 | 145.6491 | 130.4367 | 145.7612 | 144.9499 | 75.5251 | 82.0119 | 68.9499 | 122.4129 | 188.3747 | 204.938 | 221.8997 | 189.1385 | 137.0739 | 124.4631 | 123.5752 | 128.8879 | 171.8879 | 177.2243 | 210.1003 | 213.1121 | 137.8879 | 153.3628 | 165.3747 | 171.6992 | 179.2124 | 191.062 | 223.6372 | 302.2124 | 156.2507 | 156.4749 | 234.0739 | 215.7639 | 207.6253 | 198.124 | 204.9617 | 368.8879 | 158.5132 | 160.2005 | 167.3245 | 274.5369 | 135.9499 | 151.6873 | 144.5251 | 175.2124 | 177.8377 | 185.6253 | 201.6873 | 136.6992 | 116.0119 | 98.5752 | 109.5752 | 116.7375 | 281.5871 | 244.062 | 238.6755 | 449.5488 | 269.5871 | 259.0237 | 242.2507 | 395.8496 | 157.3628 | 74.1121 | 75.8997 | 157.3628 | 112.7375 | 112.4129 | 101.5132 | 150.8496 | 117.9261 | 105.9499 | 113.3245 | 104.1504 | 86.938 | 69.5251 | 69.2124 | 100.0501 | 110.7995 | 99.8377 | 104.4248 | 80.062 | 259.2124 | 253.7612 | 254.7111 | 379.6491 | 164.9235 | 195.9617 | 192.1887 | 342.2625 | 342.4749 | 205.5369 | 210.3245 | 412.0739
2096 | 25.0504 | 164.7488 | 202.0048 | 230.741 | 198.7585 | 164.0029 | 156.0019 | 153.7527 | 161.5015 | 157.0068 | 124.0068 | 123.001 | 179.2483 | 221.2541 | 134.7517 | 155.2502 | 218.257 | 176.2512 | 148.5053 | 128.999 | 209.2502 | 176.0068 | 322.0077 | 337.9874 | 180.2493 | 224.7478 | 236.7575 | 236.4985 | 354.7372 | 109.0068 | 107.5005 | 107.0048 | 165.4956 | 130.2541 | 139.7527 | 133.7546 | 96.7536 | 269.2464 | 251.9913 | 234.259 | 310.4966 | 139.5044 | 113.7575 | 133.5073 | 120.5082 | 83.2473 | 86.2502 | 79.7469 | 120.4995 | 196.9981 | 203.999 | 199.258 | 167.5044 | 143.7507 | 121.4995 | 110.0039 | 132.7478 | 163.0019 | 179.2512 | 213.2483 | 192.7575 | 162.741 | 126.258 | 182.2454 | 150.259 | 181.4995 | 184.5034 | 217.7527 | 283.5063 | 143.5024 | 175.7439 | 236.2522 | 219.2454 | 186.2561 | 188.0058 | 206.5015 | 394.4908 | 158.9981 | 167.2464 | 162.0029 | 240.7623 | 142.9981 | 139.5044 | 161.2444 | 190.2454 | 206.7401 | 166.5053 | 196.2522 | 139.2512 | 114.2522 | 96.2502 | 102.0019 | 128.2464 | 295.4966 | 254.7478 | 263.4908 | 435.7575 | 255.7556 | 248.7565 | 252.7449 | 348.7662 | 129.5082 | 67.2531 | 85.4976 | 129.5082 | 124.2464 | 106.7507 | 134.2377 | 136.7556 | 108.2502 | 120.4956 | 114.001 | 104.7488 | 86.7488 | 63.0019 | 70.7498 | 104.9981 | 97.7556 | 111.4956 | 95.5034 | 74.2531 | 258.5005 | 234.7594 | 253.2541 | 357.0097 | 156.7565 | 197.5015 | 186.7488 | 346.4985 | 325.0058 | 218.9971 | 194.5063 | 415.0019
2097 | 25.0614 | 155.3313 | 213.003 | 219.4247 | 188.4759 | 144.997 | 160.0964 | 157.0964 | 160.762 | 148.4759 | 122.3343 | 118.9518 | 168.8072 | 202.5211 | 141.3825 | 155 | 239.9579 | 176.5241 | 159.7651 | 126.1898 | 209.762 | 173.5723 | 347.5302 | 309.8494 | 193.1928 | 249.0994 | 219.8554 | 241.0482 | 350.4729 | 124.0994 | 125.2892 | 96.6657 | 163.1416 | 129.8102 | 137 | 117.5693 | 89.9518 | 255.4729 | 244.997 | 237.1928 | 314.762 | 151.7651 | 113.6205 | 147.3374 | 145.5302 | 77.6175 | 81.4277 | 83 | 141.5753 | 188.3313 | 185.9488 | 210.8133 | 178.241 | 140.7139 | 123.5241 | 123.5271 | 131.1898 | 179.2892 | 182.5723 | 199.759 | 180.4277 | 155.2349 | 139.3374 | 176.3313 | 154.7169 | 178.1898 | 181.762 | 222.6205 | 279.2861 | 153.1928 | 143.8975 | 232.4759 | 243.0512 | 188.3825 | 188.8584 | 201.9518 | 399.4277 | 160.238 | 156.5211 | 146.8072 | 272.1988 | 142.7139 | 154.0512 | 139.5662 | 196.5241 | 177.3735 | 173.1928 | 190.9518 | 148.6687 | 123.4307 | 116.5753 | 100.762 | 137.3343 | 275.3765 | 264.6205 | 275.2861 | 460.7681 | 266.003 | 258.7651 | 261.0482 | 368.5784 | 137.003 | 65.5241 | 71.997 | 137.003 | 112.759 | 99.9036 | 123.3765 | 142.4307 | 115.6205 | 115.8554 | 103.8554 | 121.241 | 78.8554 | 64.0482 | 78.6205 | 94.0452 | 101.9066 | 102.2831 | 94.2861 | 80.1446 | 268.6687 | 208.997 | 241.3795 | 394.247 | 189.6265 | 189.1416 | 184.1898 | 327.4247 | 313.6657 | 193.8042 | 210.0994 | 441.1958
2098 | 25.0724 | 164.3894 | 202.5132 | 203.7079 | 184.1283 | 146.9691 | 147.2875 | 152.0309 | 164.0974 | 168.1328 | 139.4868 | 114.1283 | 171.9691 | 187.9336 | 146.323 | 138.7389 | 241.2566 | 168.4823 | 157.2566 | 146.6815 | 222.3894 | 172.677 | 351.8053 | 326.4249 | 187.7079 | 256.7743 | 236.3585 | 228.8362 | 338.8053 | 121.7079 | 119.3849 | 119.7789 | 143.1902 | 146.4868 | 147.0664 | 129.4868 | 91.323 | 243.8053 | 224.5132 | 260.3585 | 289.4468 | 140.7389 | 126.8408 | 151.677 | 133.5441 | 83.5177 | 74.5796 | 80.677 | 124.7698 | 189.646 | 197.8098 | 209.2566 | 179.1283 | 128.3849 | 137.1638 | 129.7743 | 116.0619 | 184.2257 | 163.8672 | 216.6815 | 193.7125 | 140.7079 | 154.5177 | 161.3849 | 165.9691 | 164.6106 | 195.9381 | 237.3894 | 280.7743 | 157 | 155.2302 | 252.9072 | 249.7743 | 204.1638 | 220.4249 | 197.9026 | 401.8717 | 162.323 | 167.4868 | 169.3276 | 265.0928 | 138.9026 | 146.0619 | 146.4868 | 189.2566 | 182.0355 | 194.8098 | 185.354 | 149.677 | 128.5487 | 106.7389 | 98.677 | 138.2257 | 277.2921 | 269.323 | 241.1547 | 445.4468 | 250.0928 | 237.6724 | 269.7434 | 380 | 161.3585 | 69.8717 | 77.8408 | 161.3585 | 114.1947 | 112.7125 | 128.3894 | 125.8672 | 116.4513 | 122.2921 | 121.9072 | 126.7743 | 86.8408 | 68.8717 | 74.8053 | 102.3894 | 100.354 | 97.2257 | 106.6151 | 85.323 | 247.2211 | 214.8408 | 225.8362 | 367.1856 | 182.6415 | 196.2921 | 208.5532 | 329.5177 | 352.2657 | 202.0355 | 213.1283 | 424.4468
2099 | 25.0834 | 160.9203 | 210.5861 | 207.293 | 208.1722 | 157.6529 | 173.6786 | 156.293 | 169.7198 | 163.7738 | 140.0668 | 123.2262 | 166.1337 | 210.1722 | 139.9203 | 153.6658 | 237.4267 | 169.1465 | 142.4139 | 154.5733 | 208.6941 | 153.1209 | 333.4807 | 338.9332 | 181.0668 | 229.4679 | 246.7198 | 228.1465 | 339.3599 | 121.3599 | 123.0797 | 106.5476 | 140.3599 | 150.2134 | 142.1337 | 128.4936 | 91.2134 | 234.9203 | 229.4396 | 258.9203 | 281.2134 | 121.8406 | 130.7866 | 147.8535 | 122.2802 | 92.293 | 77.7198 | 76.0668 | 124.293 | 220.892 | 210.8663 | 204.6401 | 187.4396 | 132.0797 | 138.6401 | 138.6529 | 118.293 | 175.3471 | 165.0797 | 208.8406 | 197.2134 | 160.8123 | 153.8535 | 164.293 | 173.5064 | 154.707 | 166.9615 | 215.8278 | 269.9872 | 157.7866 | 154.9203 | 229.108 | 262.5861 | 198.5604 | 186.5219 | 182.8406 | 371.5348 | 184.239 | 179.0797 | 154.1877 | 247.9872 | 146.6529 | 150.6529 | 154.1465 | 164.9743 | 187.7866 | 210.2262 | 195.0128 | 149 | 116.4139 | 116.946 | 122.3856 | 149.0128 | 283.1465 | 242.4679 | 227.8535 | 474.185 | 262.8791 | 260.6786 | 280.6529 | 364.2674 | 138.6813 | 89.8791 | 72.3471 | 138.6813 | 118.3599 | 105.9872 | 129.6401 | 129.4396 | 117.5733 | 106.1209 | 109.9075 | 139.5861 | 88.4267 | 81.7995 | 65.9203 | 102.8535 | 91.1337 | 100.1465 | 105.2802 | 80.4936 | 262.0257 | 226.6529 | 229.8663 | 369.946 | 178.5733 | 208.4396 | 201.8406 | 318.6273 | 322.3085 | 219.5861 | 225.3727 | 423.293
2100 | 25.0944 | 172.5817 | 210.0054 | 197.8151 | 216.5978 | 172.7828 | 171.614 | 182.7666 | 186.9784 | 161.7989 | 142.9946 | 122.8043 | 167.9946 | 205.8151 | 139.5978 | 137.4291 | 241.7935 | 188.3752 | 158.9731 | 148.6086 | 209.5925 | 155.1903 | 325.0054 | 347.1903 | 177.6032 | 231.5871 | 243.2065 | 238.5871 | 364.7666 | 118.0054 | 96.2388 | 115.3806 | 144.9946 | 147.6032 | 125.6194 | 131.7935 | 82.2119 | 270.5493 | 228.0054 | 257.7989 | 297.7774 | 119.5978 | 138.9892 | 132.6194 | 122.5978 | 77.2226 | 72.6086 | 73.4022 | 137.9838 | 206.6302 | 199.4183 | 237.5547 | 165.2334 | 159.5655 | 146.7881 | 137.0054 | 135.1795 | 168.2065 | 159.8097 | 208.9946 | 193.8043 | 159.8097 | 145.0108 | 185.9731 | 147.8366 | 182.5601 | 173.1795 | 233.7666 | 287.772 | 159.5978 | 160.9892 | 228.1903 | 271.5925 | 219.9677 | 200.5655 | 193.3806 | 390.1634 | 187.6032 | 166.6194 | 148 | 286.544 | 137.0162 | 131.4291 | 130.2334 | 159 | 191.9946 | 200.2172 | 179.6248 | 155.3914 | 128.9784 | 113.0108 | 101.0377 | 142.4129 | 252.0431 | 262.9623 | 240.5817 | 452.0431 | 266.4022 | 265.0054 | 263.0269 | 336.8313 | 142.9838 | 75.0269 | 79.5871 | 142.9838 | 114.2065 | 122.9731 | 136.9892 | 148.7774 | 88.4399 | 118.5763 | 120.1795 | 144.5978 | 78.4129 | 86.5978 | 83.1741 | 113.1849 | 110.5709 | 105.7935 | 98.4075 | 89.386 | 265.6032 | 229.7989 | 227.2065 | 360.4183 | 173.4075 | 203.0108 | 207.5871 | 311.0054 | 302.2119 | 220.6032 | 236.9892 | 452.1634
2101 | 25.1054 | 158.1533 | 213.0561 | 198.2449 | 216.1888 | 154.9085 | 164.1327 | 160.6075 | 181.2654 | 162.8112 | 122.0972 | 112.2654 | 183.6018 | 199.7551 | 158.6579 | 153.9028 | 234.0767 | 197.0561 | 155.8879 | 162.4131 | 205.3215 | 159.4336 | 333.7346 | 313.3066 | 189.9794 | 229.9439 | 257.4131 | 260.4691 | 391.2803 | 115.3776 | 112.5252 | 123.0561 | 141.1327 | 147.8112 | 130.1121 | 155.714 | 99.4691 | 255.6636 | 226.1888 | 261.2449 | 276.8524 | 122.4336 | 138.5664 | 134.6785 | 128.6785 | 79.4897 | 79.9233 | 81.9233 | 116.6636 | 209.9233 | 210.6018 | 196.516 | 176.0355 | 146.5309 | 138.4542 | 129.5103 | 130.0767 | 164.5664 | 176.6579 | 220.5458 | 163.7963 | 178.2803 | 151.9233 | 162.6075 | 147.4897 | 161.6075 | 189.1682 | 229.4542 | 284.8879 | 145.3982 | 138.6636 | 243.7906 | 240.5515 | 202.4748 | 227.2803 | 186.4542 | 382.3982 | 166.7197 | 152.4542 | 173.1476 | 297 | 161.5813 | 144.6579 | 148.3364 | 171.9794 | 204.357 | 214.8467 | 200.9588 | 150.5103 | 104.6075 | 101.2654 | 106.9794 | 124.5869 | 261.8467 | 258.643 | 231.8318 | 431.0206 | 267.6224 | 248.5869 | 261.2449 | 374.8055 | 152.4897 | 76.4897 | 79.5664 | 152.4897 | 125.1682 | 110.1533 | 130.0767 | 131.9085 | 103.714 | 114.8879 | 102.9085 | 140.1327 | 78.4336 | 69.1533 | 69.3421 | 97.3421 | 107.0767 | 105.3776 | 101.0561 | 83.0767 | 274.7346 | 234.0561 | 206.5309 | 379.714 | 154.1533 | 202.6224 | 215.6785 | 312.4336 | 317.0355 | 214.3215 | 242.2449 | 437.9085
2102 | 25.1164 | 158.1175 | 213.1765 | 197.353 | 191.2949 | 140.9415 | 149.0004 | 156.4705 | 211.9402 | 170.4115 | 135.1171 | 151.1752 | 192.7645 | 203.1175 | 156.412 | 158.1765 | 223.765 | 187.2944 | 154.8235 | 158.5885 | 204 | 155.8825 | 339.294 | 343.7047 | 185.5885 | 242.9996 | 214.0603 | 240.2949 | 378.7064 | 130.6466 | 121.294 | 130.588 | 153.9996 | 152.941 | 143.5291 | 152.765 | 94.9415 | 264.8231 | 263.8812 | 243.0594 | 294.0581 | 132.0585 | 154.4701 | 143.4115 | 122.5885 | 79.353 | 70.4709 | 73.2944 | 120.8821 | 224.3526 | 213.1765 | 198.9996 | 199.7641 | 151.8821 | 150.8231 | 141.1761 | 127.1765 | 161.5295 | 204.0581 | 223.8235 | 177.5876 | 147.5893 | 149.059 | 158.4705 | 145.706 | 174.7641 | 178.8239 | 236.8821 | 295.3526 | 174.1167 | 144.5291 | 235.4709 | 238.7645 | 220.8816 | 204.8244 | 176.5885 | 391.3526 | 183.4111 | 131.0594 | 150.1773 | 296.1765 | 144.9419 | 144.8825 | 145.765 | 184.0585 | 225.9406 | 204.1769 | 188.8829 | 158.8821 | 105.4115 | 110.5291 | 119.0585 | 141.5876 | 262.706 | 274.9406 | 242.1761 | 439.5291 | 251.5299 | 243.353 | 239.7654 | 385 | 167.1761 | 82.941 | 68.2944 | 167.1761 | 123.059 | 135.6462 | 125.5295 | 130.294 | 97.4709 | 117.941 | 111.9996 | 121.7064 | 88.8821 | 74.8821 | 78.1761 | 115.2346 | 107.4705 | 103.353 | 97.8825 | 91.7056 | 270.412 | 259.7051 | 249.7632 | 357.8244 | 173.8816 | 216.1761 | 210.412 | 313.8235 | 345.7051 | 212.1765 | 214.1773 | 463.4697
2103 | 25.1274 | 144.7915 | 219.6864 | 226.2528 | 189.3432 | 144.0148 | 176.0886 | 159.5074 | 203.1199 | 197.0739 | 126.4631 | 140.7767 | 196.5074 | 225.7307 | 170.0443 | 157.1642 | 236.2085 | 183.3284 | 154.1642 | 143.6273 | 221.5517 | 181.7455 | 345.0148 | 332.7767 | 188.179 | 239.3136 | 220.7159 | 242.5222 | 370.821 | 116.4483 | 132.0295 | 126.1494 | 154.4926 | 154 | 141.821 | 161.8653 | 100.5222 | 241.2545 | 241.0756 | 241.5074 | 307.358 | 145.7011 | 139.6125 | 142.4926 | 155.2676 | 81.5074 | 75.5222 | 83.5369 | 125.5074 | 211.9557 | 207.1494 | 207.8506 | 191.4631 | 128.9261 | 123.9114 | 114.7472 | 157.0886 | 175.2085 | 176.404 | 219.821 | 182.8358 | 143.3432 | 156.358 | 151.4778 | 155.0295 | 184.8506 | 193.5517 | 209.7472 | 296.3284 | 158.4335 | 154.5222 | 227.9852 | 234.9852 | 210.1199 | 210.7011 | 184.1937 | 381.4631 | 189.6716 | 166.2824 | 149.0148 | 281.7915 | 137.4926 | 145.6716 | 144.8358 | 155.0756 | 207.4335 | 206.0148 | 200.8801 | 161.8358 | 100.3136 | 122.1937 | 92.583 | 136.8063 | 288.7455 | 239.7176 | 270.0739 | 446.179 | 269.7307 | 219.5977 | 280.9687 | 369.9557 | 152.4483 | 73.1347 | 91.9097 | 152.4483 | 124.5074 | 128.6273 | 125.8358 | 133.5074 | 100.0148 | 119.8358 | 109.9852 | 112.1494 | 80.9705 | 70.3136 | 69.2989 | 94.9261 | 115.5222 | 93.8063 | 97.8358 | 93.1642 | 280.7011 | 236.583 | 223.225 | 387.1034 | 171.4778 | 224.0148 | 216.5222 | 310.6568 | 338.4631 | 228.7159 | 235.5812 | 434.0608
2104 | 25.1384 | 164.8983 | 234.5694 | 204.8612 | 188.3038 | 156.0251 | 154.0131 | 162.5443 | 200.8481 | 187.5825 | 145.2022 | 147.177 | 199.5443 | 256.2907 | 165.3672 | 176.506 | 209.317 | 201.6579 | 139.5825 | 146.0885 | 214.823 | 163.2536 | 327.3421 | 329.8481 | 190.6962 | 229.5191 | 221.4557 | 262.6579 | 361.5191 | 121.4809 | 118.7344 | 112.2787 | 164.177 | 148.0634 | 139.3038 | 146.1902 | 99.4557 | 258.0502 | 246.8732 | 264.8983 | 324.2656 | 114.0765 | 141.0885 | 135.2153 | 139.1017 | 76.9115 | 82.9366 | 69.0383 | 125.1519 | 231.0502 | 216.177 | 231.0502 | 183.9115 | 119.7596 | 133.2656 | 114.0885 | 138.4055 | 199.2022 | 142.0131 | 211.3672 | 185.5443 | 155.8732 | 144.4306 | 175.4426 | 162.0885 | 187.6962 | 211.4175 | 228.5945 | 312.1136 | 170.1136 | 166.177 | 225.3038 | 203.4689 | 210.3923 | 190.9498 | 207.2022 | 409.5311 | 193.3923 | 152.7978 | 168.6579 | 274.7596 | 140.3923 | 168.0502 | 142.4557 | 151.5443 | 202.1519 | 217.177 | 205.6962 | 161.1519 | 109.177 | 114.671 | 105.6579 | 131.6077 | 267.7093 | 238.7847 | 236.8361 | 487.7081 | 245.1651 | 217.5443 | 278.1268 | 355.9749 | 149 | 78.6328 | 84.2787 | 149 | 122.4557 | 109.0383 | 120.9115 | 151.8098 | 116.2656 | 125.0885 | 106.4557 | 118.6328 | 83.2404 | 77.4809 | 77.177 | 103.5694 | 103.4306 | 92 | 98 | 93.8481 | 307.5945 | 263.2273 | 238.8983 | 383.823 | 186.1136 | 226.6962 | 205.2787 | 293.8864 | 346.177 | 220.1268 | 244.3923 | 455.8349
2105 | 25.1494 | 150.9321 | 226.6755 | 213.766 | 197.4642 | 147.6755 | 149.8604 | 150.9547 | 168.3057 | 161.7698 | 153.3019 | 120.6075 | 160.4226 | 240.3509 | 143.3509 | 164.5132 | 265.0868 | 179.1887 | 179.1585 | 148.7208 | 230.2075 | 162.4415 | 314.5358 | 349.7887 | 168.6302 | 240.0453 | 235.6264 | 266.8604 | 348.8151 | 113.5358 | 126.3245 | 107.4189 | 150.5132 | 144.4189 | 138.1396 | 143 | 97.2792 | 237.0491 | 267.0679 | 267.2792 | 322.6981 | 145.8566 | 142.8604 | 131.4189 | 147.0453 | 84.6038 | 76.2566 | 70.3019 | 125.8604 | 218.6528 | 208.5358 | 233.2792 | 189.883 | 146.5321 | 142.883 | 132.2075 | 144.3245 | 180.6302 | 163.6717 | 212.5811 | 203.2075 | 171.766 | 164.3698 | 158.4906 | 175.0453 | 190.5811 | 199.3736 | 227.8377 | 300.3736 | 173 | 169.7208 | 221.5585 | 259.0868 | 204.9774 | 211.0906 | 193.7925 | 377.1434 | 183.6755 | 163.6264 | 183.1849 | 285.1849 | 123.7925 | 157.3736 | 160.0679 | 198.4604 | 201.1396 | 216.4189 | 209.4415 | 160.1396 | 111.8604 | 107.8377 | 97.8151 | 149.9283 | 263.8604 | 253.766 | 274.7396 | 481.234 | 279.5774 | 224.883 | 274.2792 | 385.834 | 147.2792 | 75.6981 | 73.3962 | 147.2792 | 109.9547 | 117.1849 | 115.6981 | 133.4906 | 125.0226 | 113.9547 | 118.0453 | 119.1396 | 69.3736 | 76.4189 | 73.8377 | 104.2792 | 107.0226 | 110.0679 | 104.0226 | 86.2566 | 283.6075 | 233.7245 | 241.2792 | 382 | 203.6264 | 209.7925 | 212.4642 | 334.8793 | 346.2792 | 177.5623 | 209.7245 | 425.7245
2106 | 25.1604 | 179.4161 | 187.4753 | 191.5653 | 208.5994 | 149.4907 | 171.8167 | 178.6707 | 184.8167 | 160.618 | 137.4193 | 139.562 | 191.5247 | 231.764 | 170.5434 | 168.1087 | 267.146 | 191.5807 | 176.4006 | 138.528 | 205.9473 | 159.5093 | 333.0714 | 323.3293 | 178.09 | 203.6026 | 237.1273 | 278.3447 | 344.382 | 115.4907 | 135.854 | 127.944 | 165.4534 | 144 | 135.382 | 149.9813 | 105.7267 | 276.6334 | 250.8013 | 259.146 | 319.382 | 154.618 | 129.0373 | 136.236 | 150.7453 | 86.8727 | 80.236 | 66.6367 | 137.3447 | 219.4907 | 227.0714 | 219.0373 | 174.4193 | 139.6553 | 129.1646 | 135 | 144.2547 | 186.5994 | 171.4907 | 221.7267 | 204.2547 | 150.438 | 154.91 | 165.472 | 160.4193 | 173.5466 | 187.4006 | 228.7453 | 329.4161 | 168.6367 | 184.8354 | 251.5434 | 238.4566 | 218.8354 | 213.2547 | 188.382 | 354.4193 | 175.8913 | 153.7826 | 165.8013 | 281.764 | 155.9067 | 165.472 | 135.9473 | 177.2019 | 203.618 | 195.9286 | 189.9286 | 146.91 | 113.7453 | 132.3074 | 111.708 | 123.3293 | 250.91 | 274.326 | 270.6553 | 452.8199 | 272.91 | 227.7453 | 253.9286 | 349.9846 | 149.618 | 70.6367 | 64.146 | 149.618 | 112.3633 | 122.4907 | 101.91 | 151.8167 | 116.4006 | 118.9813 | 117.382 | 132.09 | 90.562 | 64.6553 | 66.8913 | 104 | 112.3633 | 115.618 | 106.7453 | 100.708 | 269.4006 | 234.9813 | 263.6894 | 390.7267 | 184.1833 | 187.8013 | 207.0187 | 318.438 | 360.8354 | 217.2514 | 210.1087 | 398.1833
2107 | 25.1714 | 150.3715 | 184.6549 | 218.0886 | 204.6902 | 180.0886 | 175 | 177.6902 | 185.3451 | 176.9293 | 135 | 119.9911 | 181.0707 | 243.3894 | 181.1947 | 166.3451 | 257.1504 | 202.8496 | 157.3008 | 137 | 225.0089 | 182.0089 | 304.1415 | 322.5398 | 188.8496 | 240.4781 | 255.5842 | 259.6459 | 344 | 128.3894 | 137 | 102.6813 | 145.876 | 166.124 | 136.7699 | 148.3451 | 81.3362 | 269.7256 | 246.2301 | 273.9293 | 332.2744 | 143.4955 | 146.4691 | 144.9646 | 150.115 | 93.1947 | 80.115 | 74.8496 | 117.761 | 218.2301 | 230 | 259.4781 | 178.1947 | 137.115 | 155.3187 | 150.9293 | 141.3451 | 179.1504 | 178.1947 | 234.5045 | 180.1061 | 158.5045 | 118.4866 | 188.239 | 193.3984 | 162.1504 | 188.6549 | 235.1947 | 272.9378 | 153.8406 | 154.2565 | 239.1858 | 234.885 | 191.7963 | 205.9203 | 190.6549 | 394.4781 | 162.6106 | 148.4602 | 172.7346 | 278.3451 | 149.4955 | 139.5663 | 114.3008 | 158.8406 | 204.885 | 217.7789 | 196.7346 | 143.2301 | 137.0089 | 114.761 | 123.7346 | 125.1947 | 278.2037 | 260.1858 | 258.3805 | 456.0797 | 261.2654 | 273.1329 | 241.2654 | 336.0354 | 158.8496 | 59.3805 | 71.8496 | 158.8496 | 113.885 | 120.3451 | 101.7699 | 134.6459 | 114.115 | 129.7346 | 102.8406 | 114.5309 | 75.4158 | 85.124 | 73.0797 | 107.5398 | 99.7256 | 111.5752 | 107.885 | 115.3894 | 266.2301 | 237.7699 | 243.9911 | 353.9467 | 167.7256 | 200.0443 | 211.3098 | 324.7346 | 355.9203 | 208.9557 | 219.8496 | 418.8939
2108 | 25.1824 | 149.589 | 205.6368 | 208.5412 | 235.4038 | 146.3155 | 165.1302 | 167.1302 | 157.1852 | 179 | 131.411 | 134.0478 | 189.767 | 234.233 | 160.466 | 153.4385 | 244.3357 | 205.7945 | 151.411 | 149.5615 | 199.288 | 159.877 | 346.657 | 340.945 | 185.5137 | 240.6165 | 250.822 | 235.466 | 376.301 | 130.8973 | 141.4863 | 135.7873 | 138.5137 | 147.466 | 141.4863 | 156.9725 | 103.123 | 287.7395 | 279.1983 | 259.8495 | 290.0347 | 145.589 | 141.822 | 128.9522 | 157.178 | 87.7192 | 79.1027 | 88.5615 | 135.6368 | 208.1302 | 224.6165 | 231.8017 | 193.356 | 134.3082 | 142.8495 | 145.822 | 121.2605 | 169.9247 | 161.055 | 209.0825 | 196.7395 | 146.5412 | 160.657 | 187.411 | 174.6715 | 157.411 | 180.0275 | 228.822 | 311.657 | 175.3285 | 161.6643 | 218.1577 | 243.9725 | 190.6918 | 204.1027 | 200.8698 | 392.822 | 186.123 | 175.8148 | 161.4385 | 270.822 | 152.4863 | 151.2533 | 146.0955 | 178.534 | 214.8698 | 187.8017 | 172.877 | 151.0753 | 111.288 | 121.8698 | 121.411 | 134.9725 | 267.644 | 293.89 | 282.123 | 457.8973 | 279.7395 | 232.343 | 257.945 | 384.3488 | 148.3357 | 78.6368 | 88.2533 | 148.3357 | 130.1505 | 117.3082 | 122.6368 | 140.9725 | 104.1302 | 113.9522 | 103.6918 | 117.3835 | 88.2533 | 78.1302 | 80.2808 | 96.3357 | 113.2533 | 119.0753 | 104.411 | 93.6715 | 241.7742 | 238.8973 | 258.0478 | 336.4385 | 175.8698 | 209.178 | 196.7467 | 336.767 | 340.644 | 223.1505 | 226.3835 | 437.2533
2109 | 25.1934 | 170.9195 | 225.2813 | 202.4523 | 218.0805 | 162.9195 | 161.2714 | 152.3569 | 176.7385 | 136.2515 | 114.6282 | 123.2664 | 195.5477 | 234.8191 | 148.9046 | 158.3668 | 220.2615 | 179.6233 | 149.1809 | 168.2813 | 227.834 | 167.005 | 324.7137 | 349.3668 | 178.6332 | 234.5427 | 241.8141 | 202.9851 | 393.6431 | 133.7286 | 117.4424 | 108.166 | 147.0954 | 149.5477 | 136.5427 | 156.1809 | 95.995 | 254.5279 | 260.2615 | 238.8996 | 316.834 | 146 | 139.1809 | 127 | 140.7187 | 97.005 | 81.7286 | 57.2565 | 145.2763 | 209.7286 | 210.3569 | 229.8191 | 168.6233 | 120.3569 | 129.176 | 143.1809 | 120.8191 | 152.6282 | 171.7336 | 236.0149 | 201.7286 | 150.4573 | 158.7237 | 165.171 | 165.6332 | 166.0954 | 185.3668 | 225.2714 | 303.3569 | 178.9095 | 180.2813 | 266.9343 | 260.4622 | 221.9244 | 214.9145 | 200.1809 | 391.0905 | 159.8947 | 170.8141 | 146.3569 | 307.2912 | 163.005 | 153 | 146.3618 | 164.6282 | 214.1809 | 191.2763 | 162.7237 | 162.9145 | 102.5427 | 130.2763 | 124.6382 | 129.6332 | 306.9294 | 234.3321 | 275.9046 | 450.7237 | 246.5279 | 248.829 | 270.005 | 381.8141 | 151.5477 | 72.8141 | 81.8141 | 151.5477 | 129.2714 | 133.3718 | 118.6332 | 138.3618 | 122.1004 | 109.2714 | 106.7286 | 138.9195 | 75.4473 | 66.0855 | 83.7286 | 102.2763 | 114.0905 | 102.7187 | 114.005 | 100.0954 | 282.658 | 228.995 | 235.4424 | 352.2813 | 160.6282 | 211.8191 | 184.0855 | 335.2714 | 352.6431 | 239.5527 | 217.9046 | 430.8141
2110 | 25.2044 | 188.6712 | 209.4852 | 220.4367 | 192.9542 | 140.1105 | 163.7655 | 182.3423 | 198.3585 | 170.717 | 130.5148 | 116.469 | 191.3908 | 235.9218 | 148 | 167.2965 | 297.2776 | 209.2641 | 165.593 | 149.7197 | 226.3908 | 148.6415 | 316.469 | 352.7655 | 199.2022 | 240.4528 | 232.7035 | 232.2641 | 355.3612 | 148.7493 | 135.2803 | 116.062 | 169.2022 | 156.4528 | 161.8113 | 144.938 | 96.8437 | 264.8275 | 231.2669 | 241.6092 | 303.407 | 146 | 144.531 | 148.2022 | 127.0162 | 87.8598 | 74.6253 | 63.2183 | 131.2507 | 223.8275 | 227.4367 | 236.4528 | 180.7493 | 130.9838 | 161.186 | 116.2669 | 125.6092 | 173.124 | 161.938 | 186.4555 | 185.407 | 152.8437 | 170.9057 | 174.9838 | 183.4367 | 186.3585 | 198.9057 | 210.2507 | 310.2965 | 151.345 | 154.345 | 225.9084 | 242.6415 | 229.6092 | 206.7817 | 201.8437 | 402.9838 | 177.2803 | 165.3908 | 160.6712 | 302.7035 | 151.0943 | 171.4367 | 154.2965 | 190.655 | 190.9542 | 188.3127 | 202.5606 | 173.2183 | 103.8437 | 114.407 | 96.4232 | 138.2183 | 287.0323 | 264.8733 | 251.0323 | 469.3585 | 259.593 | 245.469 | 235.9704 | 388.3747 | 144.6253 | 68.3127 | 79.1563 | 144.6253 | 113.3288 | 112.876 | 131.8275 | 138 | 110.1725 | 112.6873 | 114.3747 | 143.7655 | 79.531 | 65.9218 | 74.7817 | 99.3127 | 93.7197 | 125.8895 | 84.5795 | 89.938 | 286.0782 | 237.2183 | 221.938 | 392.717 | 186.655 | 201.8598 | 214.3423 | 335.9218 | 329.1105 | 220.7197 | 238.2022 | 463.186
2111 | 25.2154 | 175.054 | 193.054 | 236.0119 | 234.9038 | 153.8801 | 137.8445 | 161.6469 | 171.0422 | 174 | 141.3412 | 126.2754 | 171.3834 | 215.4493 | 157.3412 | 176.4071 | 275.0422 | 170.8986 | 141.7787 | 145.1976 | 224.1318 | 175.9578 | 366.4427 | 318.4374 | 206.6047 | 244.7365 | 241.3412 | 237.8024 | 381.892 | 131.3175 | 117.3834 | 141.2872 | 166.3294 | 169.1436 | 136.9104 | 126.2516 | 98.8682 | 232.3715 | 267.2991 | 268.1555 | 306.6706 | 137.5929 | 123.5152 | 135.9881 | 140.946 | 87 | 79.6047 | 72.4071 | 143.0777 | 209.1199 | 215.9223 | 245.4071 | 175.4611 | 137.6047 | 145.3175 | 132.6825 | 132.5389 | 185.2754 | 163.8024 | 220.2991 | 179.3294 | 142.7246 | 161.7246 | 158.2516 | 149.5033 | 199.2095 | 181.3175 | 212.7365 | 316.6047 | 156.473 | 159.473 | 236.0119 | 257.8142 | 194.5033 | 205.0659 | 187.054 | 374.108 | 152.8445 | 150.054 | 165.7365 | 275.8445 | 167.7484 | 156.1858 | 165.2754 | 170.581 | 211.419 | 187.0659 | 208.8024 | 156.2516 | 93.7246 | 113 | 97.7365 | 139.9341 | 288.7365 | 290.419 | 284.4967 | 488.7484 | 259.1318 | 255.2754 | 258.2213 | 384.3294 | 139.3294 | 83.8801 | 89.2754 | 139.3294 | 127.8801 | 128.7484 | 126.4611 | 117.4493 | 118.3412 | 106.4611 | 107.527 | 126.2516 | 77.1976 | 70.6706 | 70.2635 | 100.8682 | 115.3531 | 108.3834 | 93.2095 | 81.527 | 269.1858 | 240.8024 | 220.0659 | 363.3057 | 169.3834 | 220.6166 | 202.054 | 358.419 | 321.3953 | 221.8024 | 203.5692 | 451.054
2112 | 25.2264 | 160.7502 | 210.9283 | 219.9645 | 226.643 | 168.2498 | 175.7494 | 172.3034 | 151.9645 | 164.5359 | 140.1072 | 121.3215 | 191.7675 | 219.6785 | 175.0355 | 161.8574 | 271.1072 | 179.357 | 152.3034 | 169.6068 | 227.7857 | 160.0181 | 363.3751 | 333.0355 | 189.9645 | 244.0536 | 234.4287 | 244.6249 | 375.4823 | 111.0717 | 123.5713 | 112.7147 | 160.3215 | 173.7857 | 152.0355 | 144.8747 | 103.7321 | 259.3389 | 230.2506 | 269.0536 | 314.5713 | 148.357 | 125.7857 | 133.1072 | 140.1072 | 87.9464 | 91.357 | 87.1962 | 126.0181 | 254.3743 | 204.5894 | 227.0717 | 174.0536 | 122.8574 | 134.5359 | 136.8392 | 145.3034 | 182.2143 | 167.7857 | 214.4823 | 190.357 | 145.7857 | 176.1426 | 157 | 168.7675 | 200.9464 | 189.4641 | 265.9992 | 290.5004 | 184.446 | 171.357 | 259.714 | 259.9464 | 213.7675 | 203.1072 | 203.0355 | 405.1245 | 168.0355 | 174.5532 | 164.1072 | 304.2853 | 155.7502 | 151.2143 | 152.7502 | 178.4641 | 203.5359 | 196.4641 | 212.7857 | 151.2143 | 139.3743 | 117.7321 | 113.1426 | 135.2679 | 265.3396 | 262.6611 | 263.3396 | 463.5004 | 247.643 | 267.357 | 274.1962 | 428.4815 | 143.7321 | 68.9109 | 68.2325 | 143.7321 | 140.357 | 112.9645 | 116.5359 | 125.4641 | 101.0181 | 113.5713 | 109.8392 | 111.7502 | 80.7857 | 78.5713 | 73.7857 | 101.9464 | 114.1608 | 101.3215 | 113.8747 | 87.6249 | 242.4468 | 223.9645 | 214.3215 | 377.0891 | 179.357 | 205.9109 | 197.2143 | 311.7328 | 361.6958 | 228.6249 | 201.9464 | 423.5004
2113 | 25.2374 | 153.2891 | 188.0324 | 209.4129 | 214.4955 | 156.5368 | 186.6284 | 169.1652 | 189.3482 | 164 | 152.4632 | 126.7522 | 165.1975 | 220.9587 | 154.9085 | 177.298 | 269.0826 | 174.2478 | 166.4219 | 112.5189 | 258.6786 | 150.3716 | 333.2801 | 316.7433 | 201.4632 | 238.2478 | 230.1652 | 216.2388 | 364.4542 | 129.1741 | 123.0413 | 106.2065 | 163.8348 | 127.9821 | 112.7344 | 125.8672 | 105.9174 | 298.3895 | 234.7109 | 267.0826 | 313.0826 | 125.9911 | 157.6373 | 138.7522 | 129.4542 | 87.0413 | 68.0324 | 73.6194 | 120.2065 | 230.1562 | 198.2478 | 243.2567 | 178.7935 | 126.7935 | 139.7522 | 132.2065 | 116.2801 | 201.1741 | 191.9676 | 201.5368 | 200.5871 | 153.6696 | 167.4129 | 164.6696 | 155.6194 | 189.4955 | 186.1652 | 230.6518 | 302.4219 | 175.4542 | 148.0324 | 231.2801 | 250.4129 | 207.3304 | 203 | 223.1741 | 422.3393 | 157.4955 | 190.3806 | 165.9174 | 246.5602 | 145.4129 | 151.9587 | 146.2478 | 204.8851 | 192.4542 | 183.5781 | 191.9085 | 156.7522 | 102.6931 | 105.5368 | 114.9587 | 129.2478 | 250.5781 | 245.6607 | 242.8259 | 502.2656 | 261.3806 | 282.3806 | 256.7846 | 433.8761 | 140.1652 | 76.6284 | 74.6696 | 140.1652 | 123.7433 | 108.1652 | 109.2891 | 146.1328 | 93.2891 | 114 | 121.5045 | 129.2154 | 84.8348 | 76.1239 | 72.0826 | 105.8348 | 106.3304 | 122.0915 | 107.3304 | 78.4129 | 258.2567 | 254.6373 | 214.9587 | 357.8672 | 202.0502 | 213.6284 | 208.5045 | 310.9174 | 336.1975 | 214.6194 | 215.4219 | 453.6373
2114 | 25.2484 | 167.565 | 201.565 | 191.5221 | 216.913 | 156 | 160.7831 | 174.826 | 176.435 | 161.087 | 131.6381 | 124.087 | 180.5069 | 232.652 | 136.5221 | 147.8991 | 277.739 | 201.1879 | 146.6091 | 167.2888 | 238.6381 | 199.5208 | 355.3039 | 350.9559 | 196.174 | 265.1879 | 261.0719 | 243.1589 | 391.1879 | 119.319 | 140.4779 | 131.2459 | 178.565 | 123.087 | 141.1009 | 154.1299 | 84.6381 | 274.7541 | 269.9559 | 264.087 | 321.739 | 166.7529 | 159.971 | 134.145 | 147.4489 | 68.5511 | 100.9849 | 74.942 | 147.1879 | 226.087 | 205.768 | 259.536 | 173.174 | 158.0719 | 131.261 | 120.348 | 145.1009 | 175.7831 | 173.5801 | 212.652 | 201 | 150.116 | 161.174 | 180.536 | 157.913 | 188.029 | 226.7819 | 189.1891 | 328.2459 | 165.29 | 172.2459 | 237.768 | 264.565 | 194.377 | 226.3039 | 194.8701 | 398.7251 | 189.0429 | 156.0441 | 199.9849 | 308.0858 | 151.797 | 118.9861 | 124.6381 | 188.5221 | 197.826 | 203.3909 | 181.29 | 145.348 | 109.739 | 108.884 | 105.29 | 153.2749 | 279.1299 | 266.3619 | 276.9559 | 478.7541 | 272.681 | 247.0731 | 247.261 | 430.116 | 158.4489 | 78.942 | 84.71 | 158.4489 | 122.029 | 133.2459 | 140.0719 | 125.6381 | 115.3329 | 114 | 103.5511 | 129.029 | 70.435 | 98.3329 | 76.855 | 94.348 | 124.4489 | 106.4931 | 105.058 | 73.145 | 238.6091 | 227.8411 | 232.4779 | 344.377 | 157.3631 | 201.377 | 208.029 | 290.6091 | 367.0429 | 203.319 | 190.7541 | 443.348
2115 | 25.2594 | 159.1504 | 215.766 | 218.5321 | 205.2005 | 170.7493 | 169.8329 | 154.351 | 161.2507 | 182.6323 | 153.6156 | 133.8329 | 182.9666 | 244.7995 | 163.5321 | 141.1003 | 306.5154 | 220.6825 | 156.8162 | 157.2005 | 238 | 172.4846 | 383.5321 | 308.7353 | 219.5989 | 218.8022 | 246.2674 | 205.6518 | 396.9164 | 127.8496 | 126.2507 | 109.3844 | 138.6852 | 168.2312 | 143.9666 | 124.5181 | 99.7326 | 262.2005 | 257.234 | 262.0334 | 337.7326 | 166.0334 | 134.4345 | 133.0167 | 159.7995 | 97.4986 | 71.5181 | 72.0501 | 134.234 | 189.6183 | 238.4485 | 221.6518 | 209.3817 | 132.4512 | 140.8329 | 140.649 | 149.9332 | 184.8329 | 181.8496 | 208.0836 | 184.2841 | 166.7159 | 156.0836 | 182.9666 | 171.766 | 185.0501 | 186.7019 | 247.9806 | 313.2674 | 172.8663 | 166.117 | 204.5682 | 240.4178 | 199.8997 | 214.2173 | 194 | 374.4011 | 152.6351 | 156.9666 | 172.4846 | 293.2841 | 135.2841 | 149.4652 | 166.2814 | 186.0334 | 206.8496 | 190.234 | 195.7493 | 146.9666 | 110.9833 | 125.7159 | 103.0334 | 142.2005 | 301.6323 | 269.9499 | 267.1838 | 526.1811 | 273 | 274.5154 | 243.0668 | 380.8356 | 156.0501 | 77.0334 | 87.9499 | 156.0501 | 112.1671 | 122.2005 | 116.4178 | 146.6323 | 102.234 | 123.8329 | 111.8496 | 125.0668 | 71.9666 | 65.5682 | 90.766 | 90.0668 | 93.5348 | 95.1838 | 99.1003 | 80.8663 | 264.5488 | 240.766 | 240.8663 | 388.248 | 173.6992 | 204.9332 | 191.2841 | 342.1143 | 309.0027 | 213.8162 | 217.5321 | 462.6658
2116 | 25.2704 | 159.9956 | 197.084 | 218.0044 | 202.0133 | 171.9956 | 161.0398 | 157.9823 | 196.8408 | 171.0531 | 160.9691 | 123.0486 | 174.0398 | 198.2078 | 146.0796 | 151.9514 | 288.084 | 202.084 | 153.0177 | 145.0531 | 195.1901 | 193.9027 | 345.1724 | 318.9514 | 189.137 | 212.0265 | 238.0354 | 233.8718 | 386.0486 | 121.0309 | 126.9956 | 108.0044 | 164.8806 | 127.1857 | 134.0442 | 150.8806 | 108.9602 | 228.1503 | 280.8939 | 262 | 331.0309 | 145.0928 | 139.9735 | 137.9779 | 135.1105 | 94.0177 | 84.9381 | 75.9823 | 121.0575 | 196.9646 | 219.0884 | 245.8895 | 170.1768 | 153.9027 | 150.9558 | 131.0442 | 119.137 | 185 | 183.9912 | 206.0088 | 177.0309 | 121.2034 | 173.9204 | 141.1857 | 172.9956 | 163.0973 | 184.0088 | 191.2564 | 298.0663 | 176.9823 | 187.9027 | 223.9116 | 276.8364 | 227.8762 | 225.9469 | 189.0221 | 421.7878 | 152.9956 | 157.9956 | 163.0398 | 298.9735 | 145.9514 | 149.0044 | 147.0884 | 176.0442 | 188.084 | 179.0486 | 194.0088 | 147.9956 | 109.0088 | 104.0973 | 111.9602 | 167.8851 | 257.1989 | 271.9912 | 256.0486 | 478.2166 | 281.9602 | 267.0354 | 277.8453 | 410.863 | 166.9514 | 94.9204 | 98.9514 | 166.9514 | 128.9248 | 106.0707 | 119.9823 | 126.0928 | 125.8939 | 114.0442 | 117.9735 | 124.0044 | 67.0221 | 73.9602 | 75.0707 | 98.9602 | 106.9381 | 102.9646 | 87.0531 | 96.9293 | 260.0221 | 248.9646 | 224.0752 | 382.0309 | 176.9867 | 208.9823 | 198.9646 | 308.1547 | 371.7171 | 209.0221 | 218 | 480.9204
2117 | 25.2814 | 162.0708 | 228.6222 | 206.0315 | 206.1338 | 165.8741 | 162.7639 | 145.1181 | 182.0236 | 162.9843 | 138.8505 | 121.2282 | 212.8662 | 228.8977 | 170.8583 | 144.1338 | 290.9685 | 179.2676 | 122.3778 | 136.1968 | 217.0708 | 203.6144 | 349.8347 | 314.1653 | 202.0079 | 224.2597 | 256.7875 | 272.5514 | 375.1889 | 135.756 | 114.1338 | 132.8269 | 159.0866 | 147.0236 | 138.1417 | 140.9056 | 114.063 | 268.0472 | 271.8977 | 224.362 | 369.0394 | 139.937 | 130.1102 | 127.063 | 161.7639 | 86.8505 | 94.8583 | 89.0472 | 134.937 | 231.9213 | 204.362 | 238.9292 | 169.2282 | 117.1338 | 147.7088 | 157.7245 | 127.819 | 191.9921 | 182.9528 | 203.8898 | 168.2518 | 168.1181 | 149.9764 | 167.1338 | 203.8505 | 177.0315 | 204.7954 | 221.0394 | 298.2125 | 167.937 | 177.9921 | 217.9606 | 264.6065 | 226.0315 | 208.1653 | 200.9056 | 389.1338 | 175.9606 | 172.9056 | 160.0394 | 293.8111 | 152.1338 | 147.9685 | 155.9528 | 172.9843 | 193.1102 | 206.8819 | 186.0236 | 157.9843 | 104.0157 | 111.9134 | 128.8505 | 106.181 | 278.0157 | 248.1023 | 229.1417 | 510.7718 | 267.8269 | 241.0787 | 270.8347 | 360.543 | 148.8977 | 74.8819 | 87.0866 | 148.8977 | 116.1023 | 120.9921 | 110.2282 | 156.756 | 108.1968 | 127.8505 | 99.0315 | 115.1417 | 70.0551 | 68.0394 | 85.7403 | 124.8505 | 113.8977 | 89.1259 | 96.063 | 99.7796 | 281.0315 | 224.7324 | 245.8505 | 373.6931 | 184.756 | 224.6537 | 203.0708 | 375.7639 | 357.6852 | 220.0866 | 246.8347 | 465.9213
2118 | 25.2924 | 170.8185 | 181.7863 | 210.4839 | 222.5565 | 150.4032 | 133.9274 | 160.2621 | 185.5645 | 161.3024 | 120.1411 | 150 | 195.8185 | 216.1613 | 153.2016 | 161.2218 | 286.6976 | 212.5968 | 169.7581 | 160.8185 | 226.0806 | 155.3831 | 329.2016 | 334.7581 | 203.3226 | 256.4758 | 229.7782 | 215.9798 | 398.7177 | 105 | 130.9194 | 110.8185 | 169.3548 | 150.2218 | 155.9194 | 129.3226 | 121.6573 | 273.2339 | 258.2137 | 269.7177 | 373.5766 | 131.6371 | 143.8185 | 134.9597 | 132.2823 | 68.2419 | 77.1815 | 94.6976 | 127.2621 | 221.7581 | 249.5565 | 230.121 | 197.7984 | 134.4839 | 111.4839 | 123.0806 | 105.6855 | 190.7379 | 177.1613 | 190.8266 | 199.8185 | 182.6371 | 146.9798 | 184.3831 | 184.5565 | 181.0806 | 179.2823 | 225.3952 | 324.6976 | 159.8992 | 177.1411 | 213.1815 | 215.9274 | 229.2339 | 228.5968 | 189.3024 | 405.3347 | 170.8992 | 161.1613 | 165.0605 | 270.9677 | 168.5161 | 144.3024 | 149.9798 | 170.7581 | 207.0806 | 192.1613 | 189.121 | 155.5363 | 106.0403 | 101.4032 | 109.879 | 129.5242 | 279.5363 | 261.121 | 247.2218 | 481.8992 | 246.6048 | 251.2016 | 249.7379 | 428.3952 | 136.4839 | 60.3629 | 97.7581 | 136.4839 | 128.9798 | 119.9798 | 138.6976 | 126.2218 | 132.3347 | 108.7984 | 102.9395 | 132.8387 | 77.1008 | 73 | 53.6048 | 105.6573 | 100.879 | 105 | 104.6653 | 72.7661 | 284.0524 | 191.6653 | 226.879 | 335.9677 | 154.8468 | 181.7056 | 211.7177 | 345.6371 | 318.625 | 230.6976 | 226.625 | 455.9798
2119 | 25.3034 | 161.6106 | 219.1562 | 233.1562 | 202.1034 | 170.0325 | 178.2211 | 172.0913 | 211.7668 | 176.1298 | 126.7728 | 149.1562 | 186.2861 | 224.2272 | 163.4543 | 171.5132 | 271.643 | 193.0649 | 157.4808 | 151.9675 | 230.0974 | 174.2272 | 338.5457 | 322.7728 | 219.1623 | 231.5517 | 219.8438 | 215.649 | 384.7728 | 105.4543 | 126.9026 | 102.3245 | 139.1358 | 160.6106 | 151.0913 | 145.2596 | 104.9675 | 237.2332 | 219.6755 | 256.0325 | 352.8053 | 115.2332 | 135.4219 | 132.643 | 145.351 | 80.0649 | 85.8702 | 79.9026 | 139.5457 | 211.4928 | 227.351 | 234.9615 | 188.8438 | 157.6106 | 135.6166 | 126.5132 | 139.4219 | 179.1034 | 184.351 | 229.994 | 191.7464 | 165.6166 | 145.7079 | 201.8966 | 163.4868 | 185.4868 | 193.0974 | 197.5902 | 310.5517 | 155.5192 | 182.3125 | 221.5457 | 262.006 | 193.2332 | 209.357 | 203.4483 | 373.3245 | 166.357 | 169.0974 | 167.4808 | 316.7993 | 145.7789 | 159.1623 | 149.4219 | 159.9087 | 211 | 199.0913 | 195.7789 | 133.7464 | 108.0325 | 120.643 | 104.8762 | 154.8053 | 258.0709 | 267.3894 | 257.9675 | 478.3305 | 276.2272 | 260.351 | 238.4603 | 399.0325 | 159.4483 | 77.8377 | 85.9026 | 159.4483 | 127.9675 | 119.9736 | 124.0325 | 136.7728 | 100.0974 | 99.4219 | 100.2921 | 124.6755 | 81.9026 | 72.7079 | 82.5781 | 89.9087 | 95.9087 | 104.4483 | 135.28 | 109.1238 | 238.8438 | 224.5517 | 221.3245 | 382.2211 | 195.2211 | 216.3894 | 197.8702 | 328.8113 | 349.1623 | 215.9675 | 256.6106 | 455.6815
2120 | 25.3144 | 152.1029 | 193.6421 | 208.4027 | 234.3289 | 170.9553 | 155.4922 | 145.6264 | 175.9844 | 178.3892 | 121.0291 | 123.5526 | 166.9687 | 230.6421 | 176.9105 | 157.0895 | 260.8658 | 195.2237 | 142.5817 | 150.3736 | 233 | 180.1946 | 327.3714 | 316.1342 | 223.9105 | 248.6264 | 243.8814 | 234.8658 | 377.9553 | 118.8658 | 124.4922 | 112.9844 | 172.1051 | 150.0738 | 125.0738 | 152.2394 | 105.1633 | 272.3892 | 211.4765 | 254.9418 | 347.0895 | 152 | 147.7315 | 121.6421 | 127.0738 | 82 | 82.2685 | 76.6868 | 126.1342 | 256.179 | 208.0895 | 206.3266 | 213.2394 | 145.1946 | 153.7763 | 112.6264 | 151.4631 | 210.7024 | 165.7159 | 200.9844 | 212.7472 | 184.4922 | 137.0895 | 169.7606 | 177.2841 | 200.4027 | 195.9553 | 243.8367 | 325.8814 | 170.5078 | 133.2081 | 210.3714 | 289.9418 | 228.6577 | 218.4787 | 187.6711 | 383.2685 | 176.2394 | 171.8658 | 151.1051 | 280.7763 | 168.8658 | 163.4631 | 161.4631 | 186.5526 | 211.5817 | 173.4318 | 218.0156 | 155.5973 | 109.2685 | 110.4027 | 130.1499 | 149.3579 | 289.5973 | 278.5078 | 255.8814 | 517.0156 | 280.8523 | 241.3132 | 281.5973 | 398.5235 | 143.9844 | 73.3579 | 82.821 | 143.9844 | 127.9844 | 147.076 | 124.5526 | 130.179 | 104.566 | 112.3579 | 110.2528 | 115.4027 | 78.5078 | 64.3579 | 70.6711 | 116.821 | 122.3736 | 89.2528 | 85.745 | 83.3579 | 263.7763 | 241.3132 | 230.9105 | 360.8792 | 173.387 | 227.1499 | 194.1342 | 352.2841 | 353.8658 | 215.0895 | 244.3289 | 475.6868
2121 | 25.3254 | 195.1179 | 187.8821 | 215.8023 | 221.3118 | 170.2281 | 166.1711 | 160.0266 | 197.7414 | 145.0266 | 142.5437 | 114.2852 | 208.7719 | 223.1141 | 173.9734 | 159 | 256.6882 | 200.6844 | 154.3156 | 137.5703 | 234.4829 | 162.8289 | 377.0304 | 319 | 220.5741 | 261.2015 | 221.3688 | 232.6274 | 375.6312 | 117.2548 | 134.3726 | 133.7148 | 153.8555 | 170.3764 | 146.9734 | 136.8555 | 127.4335 | 238.4563 | 243.9125 | 211.6844 | 348.0304 | 151.5437 | 142.1711 | 115.9392 | 149.0875 | 82.1711 | 87.3156 | 70.6274 | 129.3422 | 258.7452 | 210.6844 | 253.2624 | 195.5171 | 128.8555 | 147.7452 | 126.2281 | 140.1711 | 183.2281 | 181.2281 | 220.4601 | 187.3688 | 192.0912 | 140.3118 | 184.2319 | 163.1977 | 210.0837 | 195 | 218.2586 | 304.3954 | 161.2548 | 158.4867 | 259.1179 | 246.6274 | 201.8821 | 185.7148 | 203.3118 | 388.3726 | 160.2852 | 169.6274 | 134.2814 | 276.2852 | 164.6882 | 152.3992 | 149.3726 | 177.5703 | 223.943 | 204.4563 | 197.1141 | 148.0266 | 114.6008 | 118.8859 | 111.3726 | 155.8023 | 282.654 | 266.1749 | 232.1711 | 495.6578 | 238.1939 | 248.6274 | 271.5741 | 367.6844 | 164.2015 | 80.3156 | 78.4297 | 164.2015 | 148.943 | 106.3422 | 115.6844 | 134.057 | 135.7757 | 119.7148 | 135.9163 | 124.9696 | 68.4563 | 72.3422 | 84.6578 | 113.057 | 109.4563 | 115.4297 | 122.6008 | 90.8289 | 260.8251 | 247.0304 | 229.3422 | 402.597 | 201.1179 | 208.0875 | 196.8859 | 336.3156 | 351.1711 | 217.057 | 228.289 | 467.7452
2122 | 25.3364 | 165.1092 | 218.8613 | 197.2479 | 240.6429 | 173.3067 | 170.3172 | 176.2374 | 209.9307 | 162.6239 | 135.208 | 120.3172 | 204.3761 | 225.2773 | 157.8319 | 158.3761 | 238.2583 | 210.1975 | 144.0399 | 148.6638 | 257.4055 | 161.8718 | 358.5043 | 321.4264 | 197.416 | 247.5147 | 244.3466 | 243.2773 | 354.7331 | 137.9307 | 124.8319 | 128.5147 | 167.7227 | 129.0798 | 130.2773 | 150.8613 | 86.8718 | 246.1387 | 257.6828 | 223.6239 | 332.6933 | 143.8613 | 143.7521 | 146.6135 | 136.0798 | 86.8025 | 75.1681 | 80.6534 | 135.3466 | 236.4748 | 220.8908 | 210.4957 | 173.3571 | 143.4853 | 127.0693 | 129.4454 | 141.3362 | 186.5147 | 183.7521 | 197.0104 | 210.4853 | 148.0902 | 160.1282 | 157.1491 | 181.7521 | 226.6828 | 195.8319 | 206.6239 | 344.6933 | 182 | 149.9307 | 226.7521 | 255.3362 | 232.5147 | 182.4558 | 220.8405 | 379.6638 | 165.5546 | 179.1681 | 171.9307 | 281.7626 | 144.8025 | 159.208 | 139.8319 | 186.5147 | 221.4748 | 212.416 | 199.208 | 164.3067 | 110.0104 | 117.3466 | 102.7331 | 136.5546 | 309.3067 | 237.8025 | 234.0294 | 491.1785 | 288.6429 | 257.4055 | 249.9411 | 378.7227 | 149.0589 | 68.6534 | 71.2184 | 149.0589 | 145.7123 | 112.4853 | 126.0988 | 135.1387 | 99.8319 | 116.4558 | 117.792 | 140.8613 | 77.2479 | 77.792 | 79 | 113.4454 | 115.6828 | 106.208 | 115.0294 | 87.3761 | 291.8319 | 231.6239 | 235.4853 | 425.3258 | 171.1092 | 193.6933 | 194.3067 | 325.3466 | 352.1975 | 218.5546 | 200.6933 | 448.1785
2123 | 25.3474 | 180.5713 | 216.0206 | 210.7353 | 208.2655 | 164.8978 | 186.4493 | 167.3059 | 208.3471 | 171.2448 | 139.4691 | 136.6125 | 194.5309 | 229.2448 | 168.0206 | 152.2853 | 283.1426 | 187.1426 | 157.5103 | 166.7559 | 237.7956 | 186.5103 | 327.0404 | 351.225 | 203.1632 | 240.5103 | 246.3883 | 245.2861 | 379.9588 | 134.9596 | 134.2044 | 124.122 | 165.4691 | 157.1632 | 135.8772 | 147.4493 | 111.9184 | 249.7956 | 240.8162 | 233.061 | 341.5919 | 142.5713 | 128.1426 | 127.4287 | 164.1632 | 109.1228 | 65.1426 | 75.1838 | 138.1228 | 214.1228 | 206.7345 | 244.2448 | 203.4493 | 149.5919 | 130.2853 | 122.8162 | 120.9588 | 180.8978 | 167 | 223.3471 | 214.6331 | 197.6133 | 136.061 | 184.6125 | 166.061 | 208.7552 | 206.3471 | 217.0404 | 352.2044 | 179.7964 | 148.3471 | 210.8162 | 234.1426 | 225.5103 | 203.1426 | 168.1014 | 401.5103 | 171.3677 | 169.2242 | 171.0816 | 292.6529 | 169 | 160.5309 | 148.9596 | 181.4691 | 201.5713 | 217.5919 | 201.8368 | 155.9794 | 134.7147 | 122.4081 | 122.7353 | 143.3471 | 299.1838 | 263.7956 | 222.5507 | 504.3883 | 257.9794 | 236.8978 | 272.4905 | 375.8162 | 124.9588 | 65.061 | 99.0412 | 124.9588 | 115.8978 | 121.0404 | 114.4897 | 136.2655 | 110.5919 | 133.3883 | 114.6735 | 138.4287 | 93.4287 | 75.4897 | 80.3875 | 106.3265 | 99.2242 | 110.061 | 102.1632 | 81.4485 | 300.5515 | 242.6117 | 243.6323 | 378.5301 | 184.3677 | 204.6323 | 187.2036 | 328.939 | 328.1632 | 226.0816 | 210.0816 | 467.7544
2124 | 25.3584 | 185.7792 | 207.7233 | 176.066 | 203.3807 | 174.7183 | 172.8502 | 179.7462 | 200.8122 | 174.3756 | 156.8452 | 120.7183 | 183.1649 | 232.6574 | 156.1548 | 175.6523 | 297.5964 | 198.5914 | 151.9061 | 122.7563 | 258.1878 | 178.8401 | 349.4366 | 320 | 202.0889 | 237.6294 | 219.2538 | 226.3756 | 404.5355 | 112.9391 | 116.066 | 142.0838 | 178.4315 | 156.8401 | 156.3426 | 130.4695 | 109.7792 | 266.995 | 250.8452 | 245.2817 | 337.1878 | 148.4366 | 139.3097 | 124.005 | 164.1219 | 87.4366 | 76.7792 | 65.8122 | 119.3477 | 195.4416 | 213.6853 | 245.7792 | 187.7842 | 143.1219 | 153.7462 | 131.3426 | 144.3147 | 191.2817 | 169.7233 | 216.3756 | 190.5355 | 141.1319 | 148 | 169.8452 | 177.2487 | 204.8731 | 198.8122 | 236.7132 | 327.6803 | 156.1269 | 142.9721 | 220.8452 | 249.2538 | 221.1269 | 213.8401 | 201.4036 | 397.5964 | 156.2868 | 182.3426 | 169.934 | 301.8781 | 166.6523 | 143.6244 | 126.7513 | 196.2158 | 210.7233 | 211.4975 | 199.2487 | 166.5305 | 110.3147 | 126.6244 | 88.8172 | 136 | 289.9061 | 282.2158 | 241.7842 | 476.6903 | 265.1498 | 244.8401 | 235.1599 | 384.9061 | 147.8452 | 76.7183 | 77.2817 | 147.8452 | 128.1599 | 141.3706 | 120.2817 | 129.7842 | 108.3477 | 105.8781 | 110.4366 | 132.0939 | 85.1219 | 80.1548 | 96 | 109.9061 | 113.5634 | 120.7792 | 104 | 106.6523 | 270.2767 | 266.8351 | 256.8351 | 403.6193 | 166.4695 | 220.5584 | 210.4975 | 317.7513 | 333.0989 | 227.0939 | 210.0609 | 509.5305
2125 | 25.3694 | 173.7876 | 233.0442 | 195.6814 | 235.5752 | 170.938 | 205.7521 | 159.8054 | 197.7256 | 177.0442 | 162.1326 | 116.8318 | 228.9645 | 239.6372 | 150.8674 | 153.3186 | 305.9909 | 186.062 | 152.1682 | 157.0884 | 263.186 | 160.655 | 347.7169 | 318.5132 | 180.7967 | 258.0087 | 238.0264 | 231.3806 | 424.1149 | 124.124 | 135.5752 | 89.5665 | 146.5488 | 138.3364 | 151.6992 | 135.531 | 97.7876 | 246.3897 | 256.876 | 243.4335 | 337.5132 | 144.593 | 123.8496 | 152.469 | 146.2124 | 82 | 67.4426 | 63.7876 | 143.3806 | 221.3364 | 202.2744 | 233.469 | 200.3273 | 127.0178 | 132.2124 | 125.531 | 155.3008 | 195.4868 | 193.6636 | 218.1946 | 213.0884 | 177.4954 | 147.469 | 177.7876 | 174.0355 | 195.6992 | 195.407 | 207.655 | 293.8587 | 164.345 | 161.3628 | 228.2566 | 268.0264 | 229.0264 | 197.2479 | 185.469 | 413.4248 | 188 | 172.4954 | 151.1682 | 319.0442 | 144 | 141.1682 | 137.5046 | 177.7256 | 232.9645 | 197.4248 | 193.6992 | 159.9822 | 124.062 | 121.5132 | 114.6194 | 139.6108 | 285.814 | 264.7876 | 257.3008 | 497.1417 | 233.9471 | 228.8851 | 255.0442 | 381.6636 | 154.0884 | 76.2302 | 79.469 | 154.0884 | 146.876 | 115.6814 | 122.7876 | 143.9202 | 129.938 | 123.3628 | 104.7876 | 134.2744 | 69.7612 | 74.124 | 95.0442 | 106.345 | 118.469 | 109.4248 | 103.3628 | 84.3186 | 251.3277 | 221.5665 | 211.0355 | 374.5046 | 173.2302 | 201.8851 | 196.531 | 329.3541 | 364.4868 | 227.6814 | 201.2124 | 500.8583
2126 | 25.3804 | 172.7109 | 224.2891 | 193.9479 | 230.2227 | 163.3034 | 169.8815 | 175.3555 | 189.9622 | 169.8294 | 135.3932 | 112.332 | 194.4076 | 242.7487 | 176.3034 | 158.7253 | 249.3698 | 195.7109 | 162.237 | 141.8294 | 287.7487 | 179.4453 | 376.6302 | 307.1849 | 233.7344 | 222.3984 | 214.0664 | 248.569 | 383.1185 | 139.9857 | 131.1706 | 125.7487 | 169.5404 | 156.3411 | 160.7865 | 141.4219 | 96.1185 | 310.957 | 239.5924 | 211.7539 | 328.4362 | 156.3411 | 131.8294 | 146.2227 | 148.3555 | 81.8815 | 85.8672 | 64.8438 | 157.3698 | 238.0378 | 212.0521 | 233.1471 | 165.0286 | 141.5117 | 134.5924 | 131.5404 | 139.3411 | 206.6966 | 179.0951 | 205.3698 | 197.5924 | 153.2135 | 144.4219 | 174.4596 | 205.5117 | 206.8008 | 181.4076 | 231.0664 | 355.6159 | 148.3984 | 156.9479 | 221.2891 | 247.8581 | 204.5924 | 229.8008 | 181.474 | 413.9193 | 187.1706 | 147.5026 | 160.0521 | 312.4219 | 146.0143 | 151.5924 | 164.5117 | 169.9622 | 201.2513 | 202.8958 | 207.4076 | 145.6068 | 129.6823 | 109 | 103.3555 | 169.763 | 262.0807 | 263.9479 | 244.7773 | 533.0664 | 284.1185 | 267.2747 | 246.4076 | 369.5833 | 138.3555 | 93.2227 | 77.1328 | 138.3555 | 129.1185 | 115.2513 | 122.1849 | 122.2656 | 121.9479 | 123.6875 | 104.8958 | 143.1042 | 91.4596 | 92.474 | 85.9336 | 86.2513 | 113.4076 | 114.7773 | 99.4219 | 88.4219 | 309.5638 | 262.9622 | 243.6966 | 397.7201 | 189.1562 | 240.2747 | 203.6068 | 364.1042 | 372.8529 | 224.526 | 202.6445 | 473.9388
2127 | 25.3914 | 176.4306 | 217.1691 | 201.1308 | 219.6156 | 172.6077 | 170.8309 | 176.1691 | 211.2153 | 179.5311 | 121.1004 | 150.2999 | 189.8692 | 227.3078 | 185.2153 | 179.5231 | 267.2616 | 204.4003 | 165.177 | 148.2616 | 266.6842 | 158.6236 | 355.7305 | 316 | 205.4003 | 268.2999 | 221.6077 | 281.8916 | 390.8006 | 129.8389 | 126.3078 | 110.4386 | 178.9075 | 146.6156 | 126.5773 | 150.2999 | 97 | 265.7384 | 243.4769 | 262.3923 | 355.5694 | 143.6077 | 139.177 | 135.2233 | 151.1308 | 84.008 | 70.2616 | 83.6459 | 168.3302 | 216.6539 | 208.5311 | 260.7305 | 192.2536 | 128.1004 | 138.3461 | 139.7305 | 127.7847 | 197 | 212.2919 | 218.6842 | 204.8772 | 188.8613 | 157.4848 | 163.7847 | 186.4769 | 181.2999 | 179.4848 | 237.4306 | 329.4083 | 190.3764 | 165.4386 | 215.7384 | 282.6922 | 209.5231 | 204.5614 | 187.4848 | 394.7926 | 183.6156 | 177.8613 | 152.3461 | 319.4689 | 162.3078 | 158.2233 | 147.4386 | 191.2153 | 218.7847 | 218.9617 | 204.0462 | 165.2616 | 105.3923 | 109.9155 | 108.2233 | 164.8613 | 287.2233 | 269.0383 | 256.9538 | 545.8389 | 283.0383 | 250.1387 | 241.2153 | 424.0383 | 140.8692 | 80.1308 | 92.3461 | 140.8692 | 131.9617 | 129.7767 | 130.8692 | 153.4769 | 127.1691 | 159.3302 | 118.4769 | 131.7464 | 80.5231 | 94.1691 | 76.4306 | 104.4386 | 109.6539 | 127.4769 | 109.3461 | 99.7847 | 285.9617 | 282.9075 | 235.177 | 359.7623 | 169.177 | 221.9617 | 221.1691 | 353.1387 | 344.0925 | 221.1308 | 202.0925 | 529.3381
2128 | 25.4024 | 165.2846 | 207.7184 | 202.8584 | 238.0015 | 165.1356 | 179.9955 | 166.003 | 208.7184 | 203.2861 | 154.4277 | 136.997 | 190.8599 | 237.2876 | 175.1356 | 177.7063 | 273.1491 | 224.8539 | 170.2816 | 149.1416 | 243.4367 | 191.7078 | 339.0075 | 319.2906 | 224.8524 | 255.4262 | 224.7229 | 232.8644 | 432.2801 | 155.5618 | 135.7154 | 120.4292 | 167.1461 | 161.4247 | 159.4322 | 139.2861 | 97.8584 | 266.8614 | 248.7274 | 259.4202 | 365.2846 | 143.003 | 145.7124 | 150.4292 | 155.1476 | 100.9955 | 73.7169 | 70.003 | 126.5753 | 224.1476 | 230.2831 | 239.4292 | 176.5753 | 162.7154 | 133.7139 | 120.2891 | 135.7184 | 196.1416 | 183.8614 | 197.8689 | 224.5723 | 167.7139 | 172.5693 | 169.5723 | 181.4262 | 172.1461 | 199.7199 | 224.4247 | 366.8494 | 155.8614 | 171.7094 | 217.7199 | 273.7139 | 212.1416 | 198.7199 | 207.4337 | 414.7063 | 194.4187 | 160.5768 | 153.5798 | 294.9985 | 171.8584 | 170.8539 | 158.0015 | 185.8569 | 225.5738 | 230.8554 | 212.0015 | 167.4292 | 115.5828 | 119.0045 | 119.1371 | 141.8539 | 300.4262 | 260.4352 | 252.1461 | 569.4157 | 272.003 | 269.2831 | 237.7169 | 413.7184 | 143.4337 | 79.8554 | 87.7139 | 143.4337 | 144.8569 | 117.5753 | 131.2876 | 152.8614 | 115.4292 | 116.7169 | 114.7139 | 147.5678 | 86.003 | 80.7124 | 68.2891 | 111.4247 | 109.7078 | 120.7094 | 106.8599 | 101.9955 | 299 | 263.9925 | 246.1476 | 418.2861 | 176.4277 | 238.7199 | 210.8599 | 368.1341 | 359.8599 | 220.9985 | 218.1461 | 503.7124
2129 | 25.4134 | 162.709 | 217.6301 | 210.1751 | 243.3785 | 139.5733 | 162.4661 | 180.1751 | 222.2035 | 204.6893 | 151 | 126.8644 | 204.4858 | 240.4945 | 150.9716 | 147.2429 | 289.455 | 208.2713 | 158.5055 | 148.9716 | 273.2035 | 164.9125 | 367.4464 | 339.4661 | 209.418 | 248.3304 | 260.0875 | 257.1838 | 412.3304 | 121.1948 | 142.0197 | 124.8644 | 177.6696 | 147.0875 | 166.3872 | 141.4661 | 104.3983 | 280.4267 | 298.6696 | 232.9913 | 361.9322 | 154.0678 | 141.1073 | 153.4661 | 172.9125 | 84.0875 | 83.3785 | 82.6215 | 143.8644 | 244.0875 | 219.8446 | 240.6017 | 192.9322 | 164.0481 | 133.5536 | 130.582 | 152.9322 | 192.2429 | 198.0481 | 242.6696 | 229.2429 | 168.4858 | 162.4464 | 176.1073 | 173.709 | 185.4661 | 213.9211 | 210.0875 | 338.0197 | 170.8249 | 157.4377 | 237.0481 | 274.4858 | 211.0394 | 220.6893 | 220.5426 | 386.418 | 159.0394 | 181.8446 | 183.1159 | 289.3107 | 174.5142 | 157.5339 | 164.4661 | 184.8446 | 229.2516 | 217.7855 | 216.6017 | 168.4464 | 156.1838 | 132.8052 | 96.7768 | 133.1948 | 291.7768 | 284.8249 | 263.6017 | 525.302 | 284.3107 | 257.291 | 245.8249 | 428.1356 | 166.6412 | 75.0197 | 85.3785 | 166.6412 | 141.2035 | 134.3983 | 133.2516 | 172.0197 | 117.2232 | 128.3983 | 112.3785 | 134.709 | 94.7374 | 72.8446 | 81.0678 | 100.6609 | 91.6128 | 102.2429 | 114.1159 | 85.3983 | 302.8841 | 237.8841 | 259.2516 | 419.3785 | 175.9519 | 250.5907 | 220.9125 | 332.8446 | 366.3391 | 215.3107 | 230.0678 | 496.9322
2130 | 25.4244 | 172 | 185.0118 | 216.8027 | 236.3116 | 159.1677 | 171.0355 | 184.2879 | 212.0296 | 201.6588 | 151.1677 | 135.1558 | 215.1499 | 224.718 | 173.9763 | 156.6825 | 251.718 | 197.5563 | 181.3235 | 173.1499 | 251.6292 | 159 | 363.6943 | 339.6528 | 228 | 264.856 | 259.7968 | 229.1558 | 420.8087 | 143.3294 | 149.4793 | 130.8087 | 163.4911 | 154.3412 | 130.3649 | 140.9822 | 110.8264 | 262.5089 | 282.9822 | 268.8205 | 368.5089 | 150.8442 | 154.4793 | 149.9645 | 168.1736 | 91.5089 | 80 | 83.6528 | 151.8205 | 248.3235 | 223.359 | 233.3353 | 196.8264 | 142.9704 | 141.3294 | 116.5267 | 157.3294 | 201.1795 | 180.1558 | 229.8323 | 232.9822 | 181.1617 | 158.5267 | 188.3057 | 183 | 185.9882 | 176.7062 | 218.8501 | 351.1795 | 162.5207 | 183.9822 | 227.3708 | 277.9408 | 234.5089 | 221.0118 | 179.1617 | 415.2269 | 183.0118 | 181.3353 | 163.8442 | 295.0237 | 161.503 | 152.9882 | 161.9704 | 180.9822 | 202.5267 | 188.359 | 206.4852 | 160.6706 | 127.4852 | 110.8383 | 101 | 157.0059 | 297.5089 | 270.9882 | 253.9882 | 563.9941 | 282.3116 | 253.7002 | 236.5148 | 414.4793 | 178.1499 | 87.0059 | 82.8383 | 178.1499 | 126.6706 | 142.8383 | 107.8679 | 179.144 | 114.6765 | 131.138 | 108.3294 | 142.6588 | 77.1677 | 74.6825 | 74.9941 | 122.6469 | 127.4674 | 111.8501 | 94.5089 | 94.3412 | 323.1617 | 262.0178 | 234.3708 | 413.3175 | 191.3294 | 200.3708 | 217.5148 | 329.4852 | 339.3235 | 223.8738 | 224.497 | 504.8501
2131 | 25.4354 | 177.0385 | 200.5783 | 196.2184 | 220.1593 | 161.7994 | 196.1415 | 147.9794 | 229.4423 | 192.4808 | 148.761 | 121.6813 | 200.6401 | 258.283 | 161.3777 | 168.0206 | 301.8379 | 243.9203 | 170.8407 | 156.3008 | 235.6552 | 166.3777 | 384.8407 | 331.0591 | 227.4602 | 285.4396 | 231.5192 | 218.9203 | 407.1978 | 139.2802 | 130.2006 | 110.7198 | 160.6786 | 160.4602 | 156.4602 | 127.4396 | 104.8201 | 272.6992 | 264.4011 | 254.581 | 378.3393 | 160.5398 | 136.8201 | 122.8791 | 173.1003 | 100.4396 | 81.0797 | 72.3599 | 142.6195 | 238.9203 | 242.6607 | 234.4602 | 196.0385 | 123.2184 | 133.6813 | 136.8407 | 156.1593 | 212.1799 | 174.419 | 228.8201 | 226.0975 | 176.1003 | 175.7816 | 168.5989 | 187.3187 | 182.1181 | 208.2006 | 232.8201 | 357.6813 | 179.6401 | 170.0797 | 259.7198 | 238.0769 | 239.3008 | 223.2624 | 179.6786 | 462.9409 | 189.4011 | 179.9409 | 175.3393 | 318.5989 | 166.8791 | 148.3983 | 137.8997 | 169.9588 | 226.9794 | 214.8585 | 199.2184 | 164 | 118.9588 | 116.7994 | 102.6195 | 164.3393 | 303.3805 | 262.2596 | 242.9203 | 560.2596 | 267.0591 | 280.0206 | 256.3777 | 400.7788 | 163.8201 | 89.1209 | 83.581 | 163.8201 | 132.6992 | 141.9615 | 131.6017 | 162.1593 | 124.4396 | 109.1593 | 109.1387 | 141.3983 | 80.6992 | 86.0206 | 72.3393 | 106.8997 | 105.7582 | 125.1003 | 103.4396 | 99.3805 | 316.4808 | 275.3805 | 258.9822 | 392.2624 | 187.4602 | 228.2212 | 224.9615 | 320.239 | 333.3393 | 257.6401 | 219.6607 | 512.522
2132 | 25.4464 | 191.3494 | 224.8466 | 215.963 | 236.536 | 166.9242 | 174.1922 | 157.5767 | 198.8447 | 179.0776 | 136.1146 | 105.268 | 200.9224 | 223.4196 | 185.5804 | 157.8466 | 325.1183 | 230.9261 | 163.8835 | 146.4991 | 298.3494 | 193.8485 | 373.8466 | 335.3882 | 224.6155 | 287.1941 | 242.6155 | 222.6506 | 441.9224 | 136 | 122.3845 | 115.7301 | 182.1922 | 160.6913 | 153.4233 | 134.3845 | 105.9224 | 285.5767 | 258.7652 | 233.8096 | 385.1553 | 159.732 | 138.8835 | 139.8447 | 165.1553 | 102.7708 | 81.1941 | 74.3845 | 148.0776 | 233.4233 | 229.6932 | 230.4621 | 218.8078 | 150.8447 | 118.4214 | 133.536 | 163.8854 | 212.0388 | 190.9612 | 226.0776 | 256.8485 | 172.1922 | 157.3826 | 179.0407 | 206.0388 | 213.4602 | 200.7689 | 235.268 | 340.3068 | 178 | 169.2329 | 261.6543 | 285.5379 | 230.0758 | 194.9205 | 199.2699 | 455.3438 | 172.8078 | 166.3845 | 181.1941 | 335.3845 | 180.3845 | 173 | 144.6913 | 193.3826 | 232.1553 | 237.7689 | 215.8873 | 164 | 133.732 | 129.037 | 112.1146 | 170.0019 | 296 | 273.3826 | 242.2292 | 572.9205 | 285.037 | 270.2311 | 279.8115 | 416.6543 | 164.9224 | 76.9612 | 70.1146 | 164.9224 | 147.4991 | 121.1146 | 110.6525 | 165.8485 | 130.6155 | 115.3475 | 128.5767 | 163.5009 | 90.3087 | 75.6543 | 82.0388 | 114.6525 | 134.6913 | 119.0776 | 114.037 | 92 | 302.5009 | 271.8447 | 225.4621 | 373.3401 | 189.037 | 211.7689 | 204.6913 | 345.3807 | 352.2661 | 252.5398 | 209 | 497.0332
2133 | 25.4574 | 155.4091 | 222.6589 | 204.7268 | 259.0914 | 155.4317 | 173.3638 | 161.2272 | 216.4317 | 176.1132 | 142.9547 | 123.4996 | 210.8407 | 268.5222 | 157.0445 | 153.2045 | 298.3403 | 192.9095 | 175.5909 | 157.8181 | 259.3411 | 162.0687 | 373.0905 | 300.7268 | 219.7276 | 268.2272 | 247.3403 | 254.7049 | 450.2045 | 135.7955 | 120.7276 | 127.0905 | 185.0453 | 173.4317 | 151.8181 | 137.0226 | 120.1358 | 291.0679 | 297.8642 | 216.0905 | 371.8634 | 147.4317 | 151.2045 | 155.3864 | 150.2272 | 85 | 64.6815 | 82.7494 | 145.9313 | 239.1811 | 215.2959 | 228.7041 | 212.0687 | 164.7502 | 135.9774 | 156.2959 | 149.8868 | 208.2045 | 193.9774 | 221.886 | 232.6362 | 172.3864 | 178.5683 | 166.2716 | 201.7955 | 217.9782 | 210.9539 | 242.2506 | 347.7502 | 181.477 | 148.477 | 251.7049 | 296.2951 | 239.523 | 227.477 | 196.4996 | 482.2045 | 176.0905 | 176.3856 | 156.7049 | 338.4317 | 179.1366 | 170.7502 | 156 | 213.9547 | 216.6136 | 238.1366 | 192.7494 | 162.5683 | 125.9313 | 141.9095 | 121.4091 | 151.0453 | 299.886 | 293.1366 | 264.7041 | 602.2045 | 301.1819 | 262.5004 | 238.2272 | 418.9766 | 167.8868 | 80.1819 | 82.2724 | 167.8868 | 155.5457 | 131.8407 | 126.6136 | 139.5909 | 129.4091 | 100 | 130.5909 | 157.2951 | 77.7728 | 70.6136 | 79.8407 | 123.6597 | 139.6597 | 111.8181 | 129.9774 | 91.7955 | 300.9992 | 294.3403 | 219.8181 | 436.6823 | 204.7728 | 217.8634 | 214.9774 | 380.4099 | 393.1358 | 235.1366 | 216.772 | 556.2045
2134 | 25.4684 | 160.0355 | 231.9159 | 222.6636 | 241.5177 | 158.615 | 175.2391 | 169.2523 | 218.9645 | 190.5795 | 138.0355 | 139.0486 | 218.2168 | 282.845 | 196.5309 | 157.9027 | 324.7345 | 184.8054 | 175.7345 | 159.0486 | 257.8895 | 153.31 | 388.7832 | 324.9514 | 204.615 | 275.6018 | 273.7345 | 237.5664 | 445.1459 | 137.385 | 114.0709 | 133.6768 | 192.1327 | 177.6991 | 151.0973 | 139.6991 | 136.845 | 301.6991 | 276.8186 | 239.155 | 381.6991 | 159.5045 | 154.1682 | 149.3141 | 155 | 85.6505 | 78.3495 | 103.7964 | 162.7123 | 263.6282 | 205.6727 | 252.6373 | 198.9736 | 158.1682 | 136.9868 | 139.3009 | 140.1904 | 206.8318 | 189.3495 | 237 | 238.1327 | 178.2391 | 172.5664 | 210.7477 | 205.77 | 182.9027 | 232.3627 | 221.2036 | 344.6373 | 189.3627 | 163.3009 | 234.5664 | 314.0841 | 226.8673 | 235.1459 | 211.615 | 475.4114 | 192.6505 | 203.7964 | 151.0577 | 348.9868 | 167.1327 | 162.6505 | 156.8673 | 200.7964 | 229.4073 | 222.0132 | 212.7123 | 160.0355 | 147.0486 | 126.6505 | 119.0973 | 159.6505 | 326.925 | 285.6859 | 271.075 | 609.7214 | 298.4336 | 254.3232 | 254.7082 | 448.2655 | 147.3495 | 79.1682 | 90.23 | 147.3495 | 143.8318 | 133.3627 | 134.8541 | 141.6859 | 127.7477 | 102.6018 | 127.9159 | 168.7964 | 73.2168 | 74.3009 | 82.8805 | 106.8054 | 121.2877 | 118.0355 | 120.1459 | 89.2655 | 327.7609 | 314.6636 | 224.7345 | 412.7477 | 207.8054 | 225.9645 | 210.1327 | 350.0355 | 417 | 224.4336 | 242.23 | 557.2168
2135 | 25.4794 | 167.7925 | 230.9784 | 206.3127 | 245.8545 | 154.124 | 194.6901 | 184.8949 | 207.0836 | 179.7061 | 144.4178 | 134.9784 | 216.2507 | 264.7897 | 179.1887 | 168.7925 | 318.8572 | 217.3531 | 178.7925 | 153.8329 | 286.5634 | 204.3127 | 388.4582 | 328.1051 | 211.8085 | 283.8545 | 275.6469 | 234.8545 | 432.2479 | 146.6873 | 138.7493 | 99.0836 | 186.4798 | 170.2507 | 141.5822 | 138.2075 | 111.4582 | 298.6038 | 284.0216 | 258.3558 | 374.938 | 173.6065 | 157.4178 | 136.9164 | 156.8329 | 91.2075 | 80.8113 | 80.4178 | 141.8329 | 237 | 234.2507 | 262.461 | 226.9812 | 165.0836 | 154.0836 | 142.6253 | 170.8141 | 205.6442 | 185.6253 | 237.9164 | 235 | 206.6253 | 167.5634 | 197.3962 | 214.523 | 199.062 | 210.6253 | 236.3343 | 362.4798 | 172.4366 | 169.1455 | 235.9784 | 313.1887 | 232.062 | 214.6873 | 204.1455 | 450.062 | 192.7089 | 180.1887 | 183.6685 | 357.1483 | 170.8733 | 163.6253 | 162.7493 | 182.6873 | 260.3558 | 212.7061 | 194.3531 | 173.7493 | 138.8545 | 128.5418 | 105 | 162.4582 | 364.5014 | 295.3343 | 232.124 | 619.337 | 297.938 | 288.2291 | 287.8356 | 441.0836 | 147.5202 | 92.7277 | 70.4798 | 147.5202 | 135.3127 | 118.0404 | 146.8356 | 151.1051 | 122.4151 | 108.3343 | 123.0836 | 147.0216 | 77.4366 | 78.7709 | 69.6038 | 134.5418 | 138.0432 | 127.6253 | 101.062 | 86.6657 | 301.4366 | 296.0216 | 237.1859 | 407.6442 | 228.6685 | 222.7897 | 214.3315 | 363.7493 | 418.3747 | 226.2291 | 224.7709 | 563.2695
2136 | 25.4904 | 154.586 | 241.2414 | 205.207 | 253.8286 | 165.1032 | 171.3458 | 191.2408 | 205.9312 | 205.8618 | 129.2414 | 138.7236 | 215.2076 | 283.758 | 168.7936 | 160.8968 | 285.7248 | 222.1026 | 170.6554 | 162.1726 | 281.2764 | 197.8962 | 390.2414 | 320.5522 | 265.2414 | 277.8274 | 264.8624 | 238.9662 | 460.4828 | 140.551 | 140.758 | 107.1044 | 176.793 | 155.4478 | 156.7586 | 149.2414 | 115.6554 | 298.6892 | 270.8618 | 241.7936 | 365.3446 | 140.0688 | 142 | 139.2758 | 170.0006 | 95.4822 | 91.4478 | 70.3108 | 149.207 | 231.2064 | 225 | 226.7586 | 199.0344 | 155.4822 | 153.414 | 135.8274 | 154.1038 | 219.7924 | 181.4828 | 237.6204 | 233.793 | 198.1376 | 160.5866 | 194.1726 | 190.1038 | 201.172 | 214.6904 | 227.8624 | 363.1732 | 186.8968 | 174.4484 | 251.3108 | 299.414 | 246.242 | 223.7592 | 204.6204 | 467.1388 | 186.4484 | 169.3108 | 160.207 | 326.414 | 188.2064 | 166.4834 | 164.758 | 193.9318 | 239.6898 | 245.8624 | 205.6204 | 177.6892 | 135.9656 | 122.6548 | 106.6898 | 159.8962 | 334.4128 | 280.1032 | 248.6554 | 580.6216 | 291.4828 | 286.1032 | 257.2414 | 441.8624 | 158.6554 | 105.4822 | 66.3452 | 158.6554 | 135.1726 | 137.1726 | 123.242 | 143.3108 | 141.1032 | 99.3796 | 123.1382 | 135.5516 | 87.3102 | 76.3102 | 75 | 130.3446 | 123.0694 | 123.7242 | 107.7586 | 97.793 | 323.1388 | 284.5516 | 250.9994 | 434.828 | 211.242 | 254.3108 | 232.7236 | 368.172 | 429.035 | 229.8968 | 222.5516 | 578.8274
2137 | 25.5014 | 152.3275 | 233.8819 | 206.9705 | 279.3302 | 161.5812 | 202.1799 | 165.9409 | 219.8644 | 207.4161 | 133.8054 | 130.6698 | 238.7168 | 256.4336 | 177.6134 | 163.6577 | 317.6107 | 194.5517 | 171.8966 | 168.7463 | 294.6577 | 183.2242 | 388.7168 | 335.7168 | 260.4188 | 278.6107 | 270.404 | 252.0765 | 459.2094 | 117.298 | 125.357 | 133.6282 | 179.0886 | 149.1624 | 154.9852 | 147.4631 | 125.2685 | 287.5812 | 270.1329 | 260 | 361.8201 | 137.0443 | 149.357 | 134.4631 | 183.1208 | 80.0738 | 77.298 | 82.9409 | 154.2685 | 224.9235 | 222.4631 | 223.7168 | 196 | 140.5812 | 154.9114 | 127.7315 | 167.5074 | 202.6255 | 187.0591 | 224.9705 | 229.7463 | 191.0738 | 175.5074 | 203.2832 | 200.7168 | 188.5369 | 234.3893 | 234.4188 | 397.5812 | 195.7463 | 176.2094 | 252.5248 | 310.298 | 263.1651 | 234.0913 | 192.7315 | 488.5987 | 198.6107 | 185.7463 | 162.2242 | 331.9705 | 174.3128 | 180.8671 | 141.7463 | 224.2832 | 250.0738 | 259.7758 | 203.3718 | 167.8497 | 144.8349 | 114.8349 | 110.2242 | 153.596 | 317.2805 | 286.9825 | 258.0148 | 618.0591 | 305.9383 | 272.1799 | 255.9705 | 449.9409 | 168.7758 | 83.9852 | 80 | 168.7758 | 138.4483 | 145.2685 | 139.9114 | 155.6872 | 133.2685 | 116.596 | 130.7168 | 132.2685 | 85.8054 | 78.8644 | 76.2685 | 119.9705 | 144.3275 | 125.2389 | 106.7463 | 97.8054 | 350.4336 | 285.3275 | 232.0443 | 436.047 | 230.9557 | 264.404 | 219.596 | 353.5074 | 441.6577 | 245.8497 | 229.1624 | 573.5221
2138 | 25.5124 | 168.532 | 212.3938 | 204.1918 | 249.7877 | 183.3196 | 188.266 | 158.7877 | 236.3504 | 234.0845 | 142.3402 | 158.3504 | 230.9361 | 254.4372 | 151.8619 | 148.9897 | 326.7009 | 203.9361 | 165.0971 | 163.4783 | 279.9897 | 183.0537 | 383.8619 | 332.4578 | 235.2226 | 289.2968 | 257.1176 | 222.5856 | 464.5651 | 127.532 | 139.5525 | 114.8516 | 190.3504 | 178.8082 | 155.7238 | 137.6062 | 128.2021 | 300.8082 | 292.2123 | 254.6804 | 378.9258 | 144.6701 | 168.0742 | 129.6598 | 168.1713 | 93.1381 | 87.7979 | 75.2557 | 160.6598 | 261.8619 | 226.7032 | 222.3196 | 208.2352 | 162.0537 | 138.8619 | 127.7238 | 172.1918 | 228.532 | 194.2123 | 217.6701 | 238.3094 | 201.4783 | 166.6267 | 216.6496 | 204.1073 | 194.9361 | 224.2352 | 217.2021 | 407.0845 | 197.1279 | 174.3836 | 221.629 | 322.9258 | 243.8721 | 209.7135 | 193.7877 | 459.0639 | 211.6906 | 184.734 | 158.0639 | 332.6496 | 182.2123 | 156.3938 | 142.3299 | 229.9361 | 263.6701 | 251.1689 | 228 | 187.468 | 154.2865 | 127.2123 | 110.8516 | 168.4783 | 368.4144 | 332.9361 | 257.0103 | 621.2226 | 329.371 | 268.3733 | 245.2123 | 439.8619 | 179.8516 | 78.3402 | 82.1279 | 179.8516 | 137.5754 | 147.6701 | 125.9897 | 148.0537 | 141.7877 | 136.7979 | 120.5422 | 143.7135 | 95.4041 | 94.5525 | 78.4041 | 118.7877 | 156.2762 | 120.8721 | 102.8082 | 105.5422 | 326.8927 | 299.6701 | 236.4783 | 398.9897 | 217.2123 | 243.4041 | 247.5114 | 350.4783 | 424.8619 | 260.6804 | 254.5525 | 578.266
2139 | 25.5234 | 178.9047 | 222.3393 | 216.3393 | 262.1652 | 195.4955 | 187.0521 | 168.1042 | 223.5655 | 220.5655 | 140.0089 | 138.6086 | 245.2529 | 285.7217 | 162.8437 | 157.2173 | 301.1309 | 211.2961 | 198.1563 | 152.6696 | 284.3214 | 196.1652 | 387.0521 | 346.7306 | 217.6786 | 260.8958 | 284.8527 | 248.5744 | 493.9301 | 132.8958 | 120.9479 | 132.3914 | 172 | 169.4435 | 163.4955 | 135.1741 | 123.7738 | 279.4777 | 281.5045 | 239.1652 | 393.9568 | 144.8958 | 165.0089 | 137.2783 | 193.1563 | 85.4955 | 87.7738 | 88.1131 | 161.3214 | 271.7306 | 263.7306 | 252.0268 | 231.4256 | 174.8869 | 147.8958 | 140.5045 | 173.7649 | 227.4955 | 176.8259 | 222.6265 | 259.8259 | 190.6696 | 151.9137 | 232.7738 | 228.6607 | 187.5476 | 252.7128 | 217.5045 | 385.2173 | 207.1741 | 183.6518 | 257.2694 | 330.7217 | 240.7827 | 232.3914 | 206.7217 | 456.1563 | 189.6176 | 183.1652 | 166.8437 | 355.1741 | 176.5134 | 163.8348 | 151.0089 | 227.5565 | 264.7306 | 226.3571 | 221.3214 | 188.5045 | 132.7217 | 119.8437 | 126.1652 | 160.4524 | 340.4777 | 330.2783 | 248.7827 | 599.7827 | 290.1131 | 293.3824 | 236.4524 | 448.061 | 194.3304 | 71.2783 | 85.4955 | 194.3304 | 167.8958 | 138.7128 | 141.1741 | 161.1652 | 150.8259 | 138.7217 | 109.3304 | 163.3304 | 95.4524 | 77.6176 | 77.6176 | 132 | 139.0432 | 118.3393 | 99.8437 | 98.7827 | 311.5223 | 297.1131 | 231.2351 | 406.1042 | 211.5134 | 241.2262 | 257.6429 | 351.3571 | 428.0521 | 240.9911 | 227.5997 | 579.2783
2140 | 25.5344 | 202 | 243.2037 | 222.3849 | 261.1623 | 191.6151 | 179.6755 | 162.3585 | 235.6755 | 228.8981 | 149.8038 | 132.6755 | 274.3095 | 281.8038 | 178 | 160.7509 | 331.1623 | 237.4189 | 189.1019 | 159.6151 | 274.5547 | 194 | 386.3585 | 355.5132 | 235.2906 | 285.2377 | 297.8907 | 273.6755 | 478.8113 | 148.2302 | 139.6565 | 130.7698 | 174.034 | 174.1019 | 163.1019 | 144.8377 | 120.3245 | 263.8527 | 286.5472 | 246.5886 | 402.9925 | 154.4189 | 173.351 | 142.0679 | 177.1283 | 82.1962 | 82.8717 | 86.351 | 149.5207 | 277.8981 | 260.6001 | 277.9246 | 211.2641 | 165.0265 | 159.4528 | 144.966 | 159.3585 | 226.5207 | 175.0075 | 239.1887 | 246.3849 | 203.4264 | 181.317 | 221.1887 | 219.7094 | 185.2981 | 231.4453 | 230.1019 | 396.015 | 211.317 | 167.5207 | 241.8038 | 318.3774 | 242.4793 | 241.8113 | 218.2037 | 449.1358 | 201.034 | 183.034 | 178.5132 | 359.8981 | 197.4528 | 164.5472 | 163.1283 | 221.4453 | 284.5547 | 263.9056 | 213.2981 | 191.5132 | 131.4189 | 132.6755 | 124 | 176.3585 | 320.4943 | 327.8038 | 258.6151 | 615.1358 | 306.9471 | 277.8981 | 248.8717 | 455.1283 | 187.966 | 76.0679 | 78.1283 | 187.966 | 173.6415 | 126.1623 | 147.351 | 164.2302 | 138.2566 | 137.1283 | 111.1019 | 158.7094 | 95.0869 | 87.5811 | 91.0679 | 138.1019 | 124.7773 | 133.3924 | 111.5132 | 96.9925 | 342.2566 | 301.7434 | 259.8792 | 402.683 | 228.3849 | 242.351 | 229.683 | 373.7963 | 434.332 | 241.6565 | 217.1019 | 581.1623
2141 | 25.5454 | 195.0343 | 253.9201 | 215.3943 | 266.4228 | 187.4001 | 168.2457 | 174.5143 | 230 | 221.2685 | 152.2971 | 130.0286 | 236.3029 | 286.1142 | 188.2971 | 183.0629 | 351.5656 | 233.5772 | 188.8572 | 162.0629 | 304.2056 | 196.4228 | 407.2971 | 348.8171 | 235.0914 | 311.8229 | 276.1428 | 265.88 | 486.64 | 156.4572 | 163.6115 | 121.9371 | 174.76 | 180.8229 | 171.3372 | 134.1258 | 126.6057 | 306.2114 | 292.0857 | 273.3314 | 388.2685 | 155.12 | 164.12 | 149.2743 | 177.7257 | 90 | 84.6971 | 76.8171 | 156.0343 | 274.8114 | 225.12 | 256.9028 | 230.8171 | 151.9371 | 153.9143 | 143.3314 | 176.9657 | 237.88 | 202.7543 | 219.8515 | 251.2056 | 219.4572 | 171.36 | 210.9371 | 220.5143 | 204.9714 | 215.4228 | 244.6971 | 421.223 | 200.1485 | 177.6686 | 243.0857 | 306.657 | 244.7485 | 242.0687 | 235.5829 | 475.36 | 196.9143 | 189.2114 | 177.5714 | 356.5085 | 192.2172 | 169.7828 | 157.9714 | 210.2685 | 302.3943 | 270.0914 | 243.5714 | 189.6628 | 130.3029 | 128.8171 | 125.2114 | 179.7315 | 345.8515 | 321.8171 | 268.0286 | 635.6057 | 349.1771 | 267.24 | 259.7828 | 456.6343 | 179.9714 | 89.3314 | 80.5428 | 179.9714 | 159.6686 | 140.5085 | 137.2114 | 169.7315 | 138.24 | 137.4228 | 126.3029 | 158.9086 | 73.8742 | 92.3314 | 99.1829 | 145.7315 | 130.3657 | 142.3372 | 110.5714 | 78.3314 | 341.9371 | 302.3657 | 270.9772 | 410.2172 | 218.0629 | 243.1142 | 236.3086 | 360.3542 | 457.4572 | 272.2743 | 234.7257 | 589.4514
2142 | 25.5564 | 185.6181 | 207.1637 | 216.5212 | 277.3575 | 172.1392 | 161.3029 | 169.491 | 232.5212 | 237.6849 | 166.8608 | 139.8363 | 248.6604 | 303.9153 | 195.6124 | 168.9577 | 376.873 | 236.8241 | 162.509 | 157.7272 | 327.4788 | 209.5635 | 446.4422 | 348.9031 | 248.4422 | 296.1515 | 302.1759 | 270.1392 | 473.5757 | 144.4243 | 142.3151 | 129.4243 | 182.0179 | 172.3998 | 165.7695 | 129.5513 | 132.7272 | 307.4243 | 311.4243 | 269.9699 | 407.8363 | 166.6181 | 170.5757 | 146.4666 | 179.9031 | 86.5334 | 80.8485 | 79.1091 | 138.4243 | 296.9577 | 228.4243 | 282.9454 | 239.7272 | 163.5212 | 144.0301 | 146.2728 | 181.3396 | 237.7272 | 208.4365 | 220.2907 | 268.1759 | 208.3697 | 190.0969 | 224.7272 | 223.0546 | 197.6849 | 221.9454 | 244.9454 | 389.3453 | 210.1091 | 171.5757 | 263.0546 | 333.0244 | 280.0546 | 220.0847 | 197.2305 | 486.2182 | 191.7573 | 188.2182 | 193.4243 | 392.4544 | 177.3575 | 182.1213 | 161.3998 | 233.9332 | 302.5757 | 269.8909 | 254.0668 | 202.1637 | 131.6303 | 130.4788 | 126.4243 | 167.0969 | 332.7394 | 335.4544 | 271.2182 | 633.8485 | 352.5578 | 279.206 | 274.6425 | 474.1392 | 178.3575 | 93.5334 | 91.6303 | 178.3575 | 153.2606 | 166.0546 | 141.2606 | 158.9878 | 143.2728 | 145.8363 | 127.9454 | 152.8062 | 90.4243 | 96.2182 | 86.8062 | 127.7394 | 132.3998 | 128.8909 | 122.3274 | 86.6303 | 346.2728 | 295.8909 | 237.8485 | 398.1938 | 224.4422 | 263.1213 | 218.0969 | 390.4365 | 452.6726 | 266.3151 | 246.0423 | 592.5456
2143 | 25.5674 | 200 | 210.1683 | 219.0531 | 279.5043 | 188.0707 | 175 | 161.8054 | 240.2921 | 233.3982 | 156.4957 | 146.9823 | 269.4159 | 288.8762 | 170.8054 | 155.7079 | 346.9469 | 252.0707 | 176.4159 | 163.4159 | 324.6195 | 218.1415 | 437.7606 | 332.4689 | 268.5751 | 305.2921 | 290.2036 | 279.1946 | 482.5664 | 147.2213 | 137.1061 | 123.0531 | 209.5043 | 194.3451 | 189.2921 | 153.4336 | 130.8848 | 308.5841 | 309.9646 | 250.5841 | 405.8141 | 179.6902 | 167.7787 | 149.4159 | 179.1859 | 82.6018 | 79.8939 | 74.6726 | 146.4603 | 285.6726 | 250.8676 | 280.7433 | 253.2744 | 171.2921 | 157.522 | 132.3982 | 166.1506 | 245.0531 | 196.2567 | 250.2744 | 253.2567 | 211.2567 | 199.3274 | 225.5043 | 233.1152 | 202.8939 | 221.7079 | 252.5664 | 428.4073 | 205.3451 | 178.2744 | 256.0884 | 315.5134 | 276.3628 | 239.7079 | 176.3274 | 486.1859 | 232.2213 | 183.9293 | 180.5043 | 406.1238 | 198.8231 | 174.1683 | 184.9823 | 253.9116 | 307.6372 | 270.7256 | 237.3628 | 195.3451 | 128.4159 | 131.5488 | 121.6902 | 175.0177 | 342.7701 | 345.8494 | 262.3451 | 640.7524 | 335.0444 | 286.6195 | 267.5841 | 483.522 | 185.4159 | 86 | 89.3982 | 185.4159 | 161.5664 | 152.8671 | 142.3628 | 171.3451 | 133.6549 | 154.2921 | 128.6902 | 142.5311 | 88.6372 | 88 | 68.0177 | 123.3628 | 146.8141 | 134.9646 | 112.9469 | 96.8408 | 351.0621 | 299.3451 | 242.7878 | 417.3098 | 255.3805 | 255.1683 | 232.239 | 372.6902 | 460.4159 | 273.2213 | 258.6372 | 583.1329
2144 | 25.5784 | 194.5643 | 225.6027 | 220.474 | 268.8329 | 188.1535 | 175 | 181.2822 | 238.2054 | 245.0248 | 171.9233 | 141.8657 | 276.307 | 289.7506 | 182.8081 | 158.474 | 360.1343 | 254.1919 | 173.7946 | 153.3205 | 331.0192 | 207.7754 | 370.0248 | 316.6603 | 264.4413 | 309.6603 | 299.9041 | 260.3973 | 499.7754 | 155.5835 | 127.7178 | 110.8849 | 193.0576 | 198.0959 | 193.6603 | 147.0959 | 111.4357 | 332.9481 | 311.2822 | 261.3589 | 393.7562 | 166.8081 | 161.1151 | 139.6603 | 192.6027 | 94.0767 | 97.4357 | 82.4932 | 162.1862 | 294.8521 | 278.1806 | 273.3589 | 255.9233 | 166.1479 | 134.3014 | 140.9672 | 180.5452 | 252.5891 | 208.0576 | 254.2822 | 266.0767 | 214.5643 | 200.6795 | 220.2686 | 244.7506 | 209.5643 | 230.5891 | 273.8521 | 451.5452 | 207.0576 | 180.9233 | 249.3973 | 341.1535 | 275.7178 | 226.1671 | 195.6851 | 499.263 | 240.2438 | 199.474 | 165.0767 | 387.2438 | 214.2438 | 200.8138 | 188.0192 | 255.8465 | 314.0576 | 262.3014 | 235.6986 | 194.3397 | 135.6467 | 134.1287 | 120.3589 | 176.0576 | 378.5124 | 336.211 | 271.1919 | 664.2438 | 357.1287 | 283.1671 | 279.3781 | 464.7178 | 177.3589 | 82.9424 | 83.3589 | 177.3589 | 164.5068 | 142.9097 | 141.7178 | 182.57 | 141.1151 | 150.8465 | 127.0192 | 162.7562 | 90.3014 | 89.6986 | 69.3973 | 124.0767 | 133.3973 | 147.4932 | 120.0192 | 107.8657 | 377.2438 | 313.2878 | 267.6795 | 422.3781 | 248.9808 | 274.3397 | 240.9616 | 372.3781 | 464.2494 | 289.737 | 250.4492 | 607.6986
2145 | 25.5894 | 180.1278 | 224.2803 | 233.2399 | 289.1682 | 204.352 | 176.4081 | 169.6636 | 225.352 | 279.8879 | 172.0965 | 132.9283 | 298.4955 | 272.8005 | 167.9283 | 175.8162 | 374.704 | 264.8475 | 185.5919 | 157.6323 | 334.0561 | 216.5359 | 411.2242 | 320.5763 | 293.1031 | 314.2803 | 297.7444 | 260.6636 | 505.0157 | 151.9283 | 136.1682 | 110.6166 | 206.3924 | 197.7914 | 183.8475 | 151.7197 | 123.0561 | 355.6636 | 300.0157 | 265.7601 | 416.7288 | 160.7288 | 179.6884 | 143.9283 | 196.2086 | 91.7914 | 100.2556 | 89.1435 | 154.7444 | 310.8318 | 237.8565 | 271.4237 | 245.8879 | 160.0247 | 147.1929 | 153.343 | 169.7601 | 275.8318 | 211.5359 | 244.7758 | 272.2399 | 212.0965 | 195.3116 | 251.8722 | 223.2242 | 207.4485 | 252.0717 | 288.4237 | 428.7914 | 218.6323 | 177.2242 | 264.0965 | 362.6323 | 275.0157 | 218.1525 | 228.1278 | 495.8722 | 237.6884 | 202.7354 | 171.2399 | 388.9126 | 206.4081 | 206.8475 | 181.5515 | 247.3924 | 309.0874 | 269.5606 | 246.0404 | 195.704 | 166.9439 | 118.6323 | 125.8162 | 172.8475 | 380.4551 | 344.2712 | 290.648 | 676.4732 | 345.8565 | 283.953 | 274.4955 | 470.704 | 188.0965 | 75.9439 | 84.2399 | 188.0965 | 164.5449 | 169.0561 | 148.0561 | 193.9753 | 153.352 | 134.296 | 125.8318 | 174.4641 | 96.5045 | 90.8879 | 78.8162 | 140.8005 | 141.4081 | 156.9596 | 122.352 | 98.2242 | 370.1121 | 325.751 | 261.6076 | 431.2242 | 254.0404 | 271.4798 | 238.4081 | 392.8409 | 482.9596 | 288.9035 | 241.3767 | 609.5919
2146 | 25.6004 | 176.6431 | 229.264 | 231.4572 | 297.1859 | 207.1859 | 189.565 | 176.2155 | 234.3791 | 277.4572 | 182.6283 | 131.7993 | 290.8364 | 291.5502 | 175.5428 | 185.3717 | 377.8216 | 257.1078 | 199.7583 | 175.2862 | 338.1859 | 216.3717 | 431.7509 | 343.5724 | 273.6579 | 309.7919 | 302.8922 | 254.4647 | 483.435 | 163.5502 | 144.1859 | 134.1859 | 203.2417 | 180.4572 | 168.0929 | 150.7435 | 120.9926 | 338.171 | 294.4647 | 270.4572 | 428.5205 | 180.5353 | 186.8997 | 148.9926 | 209.8997 | 75.5502 | 85.6357 | 76.829 | 169.3643 | 306.4572 | 245.7993 | 260.8141 | 252.9293 | 178.171 | 166.7993 | 132.4795 | 181.7435 | 259.435 | 208.8216 | 247.2007 | 267.9071 | 218.2565 | 184.0929 | 260.0929 | 236.829 | 223.3643 | 243.3643 | 280.3643 | 428.9442 | 220.9852 | 195.9293 | 281.9145 | 375.9145 | 271.2862 | 239.3717 | 212.6209 | 507.3717 | 235.4276 | 182.7286 | 169.4572 | 403.8997 | 210.4572 | 199.1078 | 166.3643 | 260.6357 | 306.4869 | 294.4647 | 254.264 | 210.8438 | 158.4424 | 128.2714 | 125.8997 | 164.7435 | 357.6505 | 322.2788 | 278.7064 | 700.699 | 354.1636 | 313.3569 | 266.1078 | 467.6283 | 203.7286 | 86.7509 | 76.264 | 203.7286 | 193.5428 | 172.8216 | 144.5353 | 180.2007 | 160.5576 | 138.1003 | 122.1859 | 180.4572 | 110.7212 | 92.1003 | 87.2788 | 151.5353 | 158.5724 | 151.2862 | 129.5576 | 106.7286 | 369.264 | 322.0519 | 256.3791 | 433.171 | 270.6431 | 269.736 | 240.2714 | 407.4276 | 475.4647 | 276.9145 | 257.6209 | 605.1784
2147 | 25.6114 | 177.351 | 232.2981 | 232.117 | 310.0384 | 200.0786 | 206.8702 | 195.1955 | 256.1554 | 281.1298 | 174.2468 | 128.3109 | 312.1426 | 292.234 | 179.4022 | 199.4022 | 389.6618 | 268.8574 | 211.6763 | 184.6106 | 351.7916 | 230.0256 | 441.5721 | 363.7276 | 287.9744 | 298.5577 | 299.4279 | 253.4552 | 476.6747 | 162.351 | 144.6106 | 127.4552 | 187.1811 | 191.2852 | 176.0256 | 167.883 | 119.649 | 340.0513 | 320.1939 | 281.6618 | 432.1153 | 175.5192 | 173.4808 | 143.5064 | 202.5064 | 90.5448 | 85.3766 | 80.9744 | 175.2724 | 312.0128 | 239.6747 | 271.6875 | 258.4423 | 179.6747 | 154.649 | 157.1298 | 189.468 | 256.8173 | 215.766 | 252.5721 | 268.6362 | 224.8429 | 195.0384 | 256.351 | 253.7788 | 221.3638 | 251.1554 | 281.7532 | 449.1955 | 230.7259 | 204.0786 | 281.7276 | 372.3382 | 273.4552 | 237.2084 | 201.3894 | 517.6362 | 219.0384 | 198.6875 | 184.8045 | 396.1298 | 232.5833 | 211.6106 | 176.415 | 250.2084 | 318.5465 | 296.468 | 255.7916 | 230.1042 | 143.9872 | 128.5064 | 113.6106 | 183.766 | 367.6106 | 332.1426 | 265.7788 | 684.1682 | 341.8958 | 300.2212 | 281.6234 | 480.7131 | 186.9231 | 98.8318 | 70.6362 | 186.9231 | 184.9744 | 167.3382 | 143.6618 | 195.3638 | 172.3766 | 140.2212 | 123.7404 | 190.1554 | 101.6106 | 105.1426 | 93.7532 | 141.6234 | 180.234 | 159.8574 | 136.8574 | 111.3894 | 367.0256 | 351.0401 | 271.3766 | 414.7148 | 275.4936 | 278.3766 | 234.4808 | 386.5192 | 479.351 | 285.3894 | 265.2195 | 609.532
2148 | 25.6224 | 160.9999 | 237.7776 | 228.2222 | 321.5 | 195.2224 | 208.5 | 179.4445 | 270.7222 | 284.5556 | 175.7222 | 151.6111 | 323.6667 | 298.4446 | 177.1665 | 197.1665 | 396.2221 | 253.0555 | 193.889 | 179 | 364.8889 | 225.2221 | 429.9445 | 355 | 283.8334 | 325.2778 | 307.1665 | 246.889 | 508.0557 | 158.4445 | 143.6667 | 113.1111 | 216.0001 | 199.1666 | 180.9444 | 160.111 | 124.3333 | 359.2222 | 336.111 | 287.0555 | 463.7776 | 185.3334 | 172.2223 | 146.7778 | 199.1666 | 94.7776 | 91.0556 | 70.2222 | 177.7777 | 314.2778 | 262.5 | 283.1665 | 237.1667 | 184.611 | 159.7222 | 155.1111 | 173.5 | 262.7776 | 236.7779 | 236.2778 | 274.9445 | 255.5555 | 210.7778 | 245.8333 | 250.2221 | 227.1112 | 250.1665 | 279.5 | 446.6669 | 265.7222 | 196.5001 | 274.9445 | 377.4446 | 268.8335 | 238.8891 | 214.0001 | 520.4444 | 239.0556 | 215.9999 | 190.0554 | 401.5 | 250.9999 | 200.1666 | 184.611 | 244.5001 | 313.6669 | 296.4446 | 262.6666 | 218.1111 | 137.4444 | 137.6112 | 102.9444 | 191.1667 | 364.3334 | 333.9444 | 271.1666 | 703.3334 | 360.889 | 296.389 | 284.8889 | 495.9443 | 175.6669 | 87.8334 | 76.1667 | 175.6669 | 169.1666 | 156.8889 | 158 | 188.2778 | 165.9999 | 143.7779 | 131.2779 | 194.2221 | 93.2778 | 108.8889 | 92.6666 | 136.3333 | 174.7778 | 149.1111 | 130 | 115.4444 | 379.3334 | 329.6668 | 280.1667 | 415.0001 | 269.5 | 287.9445 | 235.1667 | 389.3333 | 475.8334 | 290.6111 | 274.3886 | 618.1111
2149 | 25.6334 | 160.832 | 225.4249 | 227.4012 | 316 | 216 | 205.9763 | 186.7964 | 270.587 | 298.6343 | 173.1799 | 145.7846 | 328.5929 | 311.1799 | 165.0237 | 179.8083 | 392.4249 | 239.8024 | 206.9882 | 183.0118 | 357.5929 | 206.413 | 448.413 | 363.0237 | 307.0414 | 331.5988 | 309.4663 | 280.0474 | 512.5574 | 176.6284 | 144.9822 | 113.587 | 228.2036 | 192.0059 | 170.5929 | 149.8201 | 132.245 | 346.7669 | 319 | 276.6107 | 454.2331 | 186.9941 | 178.7905 | 153.2154 | 189.5988 | 83.6284 | 94.587 | 71.8083 | 171.6107 | 313.2095 | 270.8083 | 260.1917 | 250.6343 | 176.4426 | 159.0178 | 147.7964 | 173.6166 | 245.4426 | 243.5515 | 235.3775 | 292.0414 | 252.2154 | 215.2095 | 245.6047 | 255.8852 | 242.5988 | 245.245 | 257.5515 | 491.245 | 246.1621 | 212.5988 | 272.7846 | 389.9467 | 287.3953 | 258.5751 | 214.9704 | 541.2687 | 247 | 221.0474 | 164.5988 | 409.8083 | 244.245 | 203.8379 | 169.3893 | 267.8201 | 353.5988 | 314.3953 | 246.7964 | 216.413 | 142.6225 | 157.6284 | 113.8556 | 195.4071 | 374.4189 | 329.2095 | 256.7846 | 718.2331 | 371.1799 | 322.0414 | 288.9941 | 489.2687 | 198.9822 | 100.6107 | 80.0059 | 198.9822 | 166.4189 | 150.5633 | 149.9763 | 179.1858 | 160.6166 | 160.9822 | 130.5455 | 174.7787 | 96.6225 | 99.1858 | 83.3834 | 137.4249 | 167.5751 | 140.5929 | 124.7846 | 113.4012 | 387.413 | 352.413 | 290.9941 | 444.6225 | 261.1917 | 295.5811 | 245.9941 | 390.826 | 483.7964 | 301.6343 | 253.8675 | 620.0237
2150 | 25.6444 | 172.8653 | 242.1347 | 226.3461 | 311.4518 | 211.8653 | 199.7886 | 181.6922 | 269.3752 | 327.5773 | 177.4716 | 140.8943 | 356.4241 | 315.8851 | 190.6447 | 173.6631 | 402.9326 | 257.9525 | 212.5773 | 177.4227 | 365.5773 | 227.4716 | 454.5865 | 352.2589 | 311.461 | 321.9036 | 341.7596 | 297.7979 | 501.3369 | 179.0092 | 139.7213 | 109.4808 | 226.6922 | 195.4135 | 180.2312 | 148.8362 | 140.9419 | 350.2603 | 312.3844 | 277.4518 | 446.1915 | 185.6539 | 178.1347 | 164.3078 | 198.0964 | 95.1057 | 91.7213 | 83.0291 | 172.0383 | 310.9709 | 270.4518 | 251.4518 | 260.5575 | 188.4901 | 166.7596 | 135.6631 | 194.8177 | 266.5865 | 215.2787 | 245.9142 | 293.1532 | 281.9142 | 211.317 | 245.1057 | 290.7305 | 249.4425 | 258.9036 | 247.8851 | 504.4901 | 248.9142 | 216.9617 | 278.2312 | 371.683 | 293.683 | 247.2404 | 204.5482 | 557.3936 | 251.1347 | 237.144 | 171.0291 | 408.6248 | 271.9617 | 216.3844 | 175.1638 | 286.2695 | 357.5482 | 310.7596 | 253.683 | 240.7794 | 146.5575 | 151.7397 | 132.9036 | 188.6631 | 387.3078 | 342.2695 | 253.5482 | 740.7886 | 371.7503 | 331.0092 | 294.2695 | 534.3369 | 191.6539 | 94.8362 | 85.0674 | 191.6539 | 175.5865 | 151.4333 | 144.2021 | 188.1255 | 160.7305 | 149.9326 | 121.1929 | 172.2695 | 101.7979 | 93.6539 | 81.2021 | 154.1347 | 157.0674 | 147.3369 | 132.712 | 114 | 398.5482 | 377.1929 | 286.7596 | 461.3752 | 259.4808 | 292.2021 | 245.0674 | 403.1731 | 477.4518 | 312.3844 | 270.2206 | 624.9709
2151 | 25.6554 | 178.9214 | 254.3865 | 224.8515 | 309.6834 | 214.9411 | 213.535 | 177.4258 | 269.762 | 323.1375 | 198.8515 | 147.9804 | 389.5043 | 353.6049 | 215.1092 | 176.4564 | 400.8515 | 282.8515 | 225.5939 | 181.4368 | 386.2577 | 235.227 | 460.8123 | 340.4368 | 289.7031 | 336.249 | 322.5436 | 287.7227 | 541.546 | 165.4061 | 145.4258 | 117.6834 | 221.1485 | 205.3669 | 189.6331 | 142.7926 | 124.8319 | 379 | 321.7337 | 277.3865 | 438.6748 | 191.8319 | 179.3166 | 152.3951 | 211 | 82.9104 | 93.1681 | 98.1092 | 177.3472 | 327.3975 | 271.238 | 266.2883 | 255.9607 | 189.476 | 168.8319 | 126.1092 | 201.9301 | 288.5656 | 213.8319 | 261.2467 | 272.6638 | 304.9104 | 210.9214 | 254.1987 | 290.9214 | 248.5043 | 260.0503 | 265.1681 | 499.9411 | 274.8908 | 214.6331 | 291.8908 | 395.4957 | 307.7031 | 260.0699 | 218.238 | 534.4061 | 249.3362 | 221.3166 | 184.8319 | 409.9411 | 276.0196 | 211.2577 | 199.0896 | 307.7926 | 374.2184 | 318.3669 | 254.0786 | 257.9104 | 143.6638 | 126.8319 | 120.8515 | 181.238 | 388.5939 | 359.9607 | 259.1485 | 747.7227 | 382.7423 | 320.3865 | 298.762 | 540.4847 | 193.1485 | 86.238 | 89.2773 | 193.1485 | 181.8123 | 175.5939 | 149.1681 | 213.9804 | 157.0896 | 154.238 | 133.3755 | 190.8123 | 98.1092 | 93.0196 | 87.4454 | 159.1485 | 160.4258 | 155.6135 | 143.0786 | 115.2773 | 439.487 | 383.8405 | 283.7227 | 465.1681 | 268.9607 | 313.3472 | 257.3669 | 432.6551 | 473.5546 | 303.4258 | 250.9607 | 626.0699
2152 | 25.6664 | 181.4202 | 269.1328 | 234.3231 | 311.1814 | 227.8761 | 215.3053 | 176.2478 | 285.8454 | 302.885 | 205.2567 | 151.1239 | 378.8186 | 370.3053 | 208.4202 | 189.0486 | 409.885 | 301.0842 | 217.0575 | 192.3538 | 406.8939 | 239.9028 | 466.9336 | 364.4955 | 295.5045 | 370.8097 | 314.7701 | 275.4866 | 557.1635 | 165.6194 | 141.6194 | 122.2478 | 219.5619 | 212.7433 | 180.5045 | 154.2478 | 122.553 | 374.6194 | 341.7433 | 280.7433 | 485.1992 | 207.0753 | 182.1992 | 138.7611 | 215.8186 | 88.2081 | 85.3717 | 98.4291 | 165.5263 | 354.6769 | 288.447 | 289.3053 | 278.4559 | 206.8097 | 156.4777 | 140.447 | 195.3231 | 303.668 | 227.7611 | 252.2656 | 302.2745 | 290.9336 | 230.0664 | 278.5709 | 310.9425 | 236.9425 | 270.6769 | 260.4381 | 506.7433 | 268 | 203.3142 | 279.3053 | 412.3142 | 306.9336 | 271.0575 | 216.6105 | 545.1328 | 242.0664 | 213.6858 | 178.1725 | 438.6462 | 275.5045 | 220.0664 | 217.885 | 324.5045 | 388 | 337.1328 | 250.2656 | 245.2478 | 147.4291 | 128.9336 | 123.7522 | 193.1903 | 389.6947 | 367.5619 | 260.1903 | 740.3053 | 373.0575 | 345.6462 | 291.6283 | 554.8454 | 185.8672 | 102.1328 | 82.2389 | 185.8672 | 197.1328 | 174.9425 | 139.6194 | 222.3806 | 168.8761 | 157.4291 | 130.7701 | 193.8672 | 107.1903 | 92.0664 | 83.5619 | 161.9425 | 166.6947 | 150.9425 | 131.3806 | 113.0575 | 471.9822 | 364.1328 | 276.7433 | 449.9247 | 281.8186 | 325.1814 | 270.8761 | 453.4113 | 495.8365 | 310.1328 | 257.6858 | 620.4113
2153 | 25.6774 | 177.9608 | 261.2378 | 231.4896 | 322.5636 | 222.2448 | 208 | 181.2056 | 297.8336 | 326.0601 | 206.5958 | 156.3049 | 401.6629 | 372.4573 | 201.358 | 180.4042 | 413.2448 | 321.649 | 234.5174 | 178.6559 | 424.649 | 247.284 | 457.5958 | 362.2517 | 317.1132 | 368.9007 | 343.1063 | 266.9538 | 549.6629 | 164.0531 | 148.1594 | 126.7552 | 230.2587 | 216.8084 | 181.8476 | 152.9007 | 121.1063 | 373.9538 | 325.0923 | 295.1664 | 490.6951 | 227.2056 | 184.5427 | 157.2056 | 226.5035 | 111.2517 | 79.7483 | 85.9469 | 140.709 | 351.6559 | 284.0392 | 286.5035 | 307.358 | 209.4042 | 161.4712 | 146.3972 | 203.7483 | 290.7552 | 237.1986 | 267.0993 | 325.649 | 307.2657 | 245.2587 | 285.649 | 334.0139 | 244.2517 | 266.3049 | 274.9608 | 522.0671 | 273.4042 | 206.351 | 285.0601 | 432.4643 | 315.6097 | 266 | 210.5566 | 570.5636 | 262.2125 | 225.9608 | 158.4965 | 439.284 | 296.6629 | 238.4111 | 223.0462 | 339.358 | 400.6097 | 348.6028 | 264.1986 | 254.2587 | 158.8153 | 130.4042 | 132.7552 | 204.1063 | 417.2657 | 360.6951 | 272.007 | 757.3188 | 371.6028 | 353.4896 | 297.7021 | 570.4364 | 169.8937 | 100.5427 | 71 | 169.8937 | 205.9007 | 176.3972 | 149.3118 | 215.8406 | 174.9538 | 156.2056 | 138.3903 | 193.2056 | 104.1455 | 100.5035 | 82.0993 | 160.6951 | 168.1455 | 160.5035 | 145.5566 | 108 | 468.4111 | 387.3118 | 273.1524 | 462.6698 | 291.6028 | 322.1524 | 257.5889 | 429.2517 | 533.9608 | 316.649 | 260.0531 | 623.723
2154 | 25.6884 | 190.3736 | 255.4165 | 218.7011 | 317.6296 | 218.6264 | 206.1494 | 193.7011 | 281.7154 | 370.221 | 201.8506 | 157.9857 | 415.9857 | 381.2242 | 204.2846 | 181.4483 | 412.8649 | 318.6121 | 252.7472 | 180.5978 | 435.0286 | 252.0715 | 454.7011 | 359.911 | 320.4197 | 374.6264 | 353.9253 | 274.3879 | 593.132 | 178.2528 | 168.089 | 133.1494 | 246.7758 | 223.5231 | 185.4022 | 160.4769 | 141.6407 | 378.1494 | 327.3418 | 302.0318 | 485.7758 | 231.8362 | 176.7011 | 152.1208 | 238.4769 | 105.6725 | 78.3879 | 79.1494 | 168.8792 | 382.2814 | 271.1638 | 301.7154 | 322.3132 | 234.9682 | 192.2242 | 146.6264 | 194.5231 | 313.8045 | 231.2989 | 285.4308 | 335.1033 | 322.7472 | 258.5374 | 290.9396 | 361.5517 | 238.2099 | 281.8649 | 306.8045 | 574.4452 | 280.4626 | 217.7154 | 299.6121 | 448.9857 | 321.2846 | 275.7154 | 239.9682 | 600.3275 | 285.7011 | 240.2242 | 159.9396 | 432.0429 | 318.8506 | 249.9857 | 212.3736 | 342.7472 | 440.9825 | 358.5517 | 264.3132 | 271.2385 | 173.7472 | 147.1781 | 145.1638 | 206.5978 | 463.2814 | 363.1781 | 287.6264 | 793.9396 | 390.3418 | 359.6693 | 315.3418 | 570.2814 | 182.2067 | 93.1638 | 75.6264 | 182.2067 | 216.2528 | 167.3736 | 162.911 | 204.2385 | 183.7758 | 164.5374 | 127.6264 | 197.8362 | 109.0286 | 109.2385 | 87.4769 | 174.2814 | 187.3704 | 166.9253 | 148.5978 | 116.3275 | 477.2099 | 420.3418 | 277 | 484.7472 | 307.566 | 317.2099 | 261.744 | 431.0747 | 561.1781 | 322.8649 | 283.5056 | 659.5374
2155 | 25.6994 | 190.6992 | 244.8788 | 235.3482 | 320.2451 | 242.5223 | 213.9735 | 189.4513 | 295.3747 | 418.1978 | 204.9499 | 150.0501 | 447.9443 | 387.0237 | 194.4248 | 184.4986 | 453.8203 | 318.4248 | 245.3245 | 181.0237 | 450.4248 | 278.9262 | 470.3983 | 358.2744 | 321.3719 | 379.0501 | 358.7994 | 262.702 | 605.2772 | 190.8997 | 176.8496 | 127.2006 | 265.1978 | 223.5989 | 188.3008 | 160.876 | 141.6518 | 375.5251 | 339.2507 | 271.9262 | 496.5989 | 235.0738 | 183.8496 | 149.298 | 241.2507 | 107.3983 | 88.0738 | 79.3747 | 184.5251 | 401.3008 | 278.8496 | 297.8022 | 317.4011 | 230.8788 | 181.876 | 144.876 | 212.6462 | 336.2242 | 231.7493 | 308 | 359.8733 | 340.9707 | 226.6546 | 297.1003 | 365.6253 | 246.0473 | 293.8496 | 306.9025 | 580.2535 | 297.6226 | 219.5014 | 292.3008 | 445.7994 | 315.2242 | 292.6992 | 250.1267 | 615.2242 | 284.3008 | 241.7493 | 171.3245 | 466.773 | 326.2242 | 264.8468 | 202.7256 | 344.8231 | 471.8997 | 381.1476 | 264.6253 | 291.1476 | 164.8997 | 158.0501 | 139.5515 | 194.6755 | 504.1476 | 385.9234 | 270.2034 | 805.7994 | 401.3008 | 370.3273 | 313.4777 | 585.5014 | 204.6755 | 92.3008 | 77.6755 | 204.6755 | 219.4011 | 172.4485 | 164.5989 | 213.2242 | 183.6755 | 164 | 132.0501 | 215.7966 | 126.7994 | 110.6253 | 91.6755 | 187.1267 | 203.7521 | 161.351 | 148.0738 | 122.2006 | 488.3719 | 433.6755 | 283.1741 | 493.4721 | 315.3008 | 327.8969 | 272.7521 | 448.9972 | 560.6518 | 332.4749 | 265.5808 | 696.9944
2156 | 25.7104 | 192.1278 | 230.6166 | 241.2045 | 363.8211 | 244.4856 | 220.1278 | 182.9233 | 306.77 | 439.385 | 204.5383 | 153.8722 | 492.9489 | 402.8467 | 197.4617 | 201.8211 | 480.23 | 338.5144 | 258.7444 | 183.2045 | 483.1822 | 268.0511 | 493.7189 | 352.2811 | 331.3578 | 398.9505 | 375.1805 | 277.2333 | 604.5911 | 181.7939 | 172.2045 | 125.9233 | 278.5895 | 244.1805 | 185.4105 | 164.6933 | 137.2572 | 406.6694 | 337.8978 | 274.2061 | 485.5112 | 252.77 | 195.77 | 158.6406 | 250.1294 | 111.23 | 97.9744 | 90.77 | 192.77 | 399.385 | 289.7955 | 286.3594 | 314.8467 | 212.7189 | 168.6406 | 133.1022 | 228.3322 | 361.4888 | 247.2316 | 302.1533 | 373.4872 | 377.8722 | 212.4633 | 309.1294 | 380.54 | 277.6933 | 289.6917 | 296.77 | 584.2077 | 329.155 | 220.6933 | 293.3083 | 469.9761 | 323.9233 | 293.1533 | 238.4617 | 624.8978 | 297.4888 | 242.1278 | 174.0256 | 471.2811 | 346.6166 | 279.2556 | 212.1294 | 375.0783 | 482.77 | 406.7955 | 270.77 | 329.4633 | 173.8211 | 145.3067 | 125.0767 | 197.8211 | 515.1789 | 407.2572 | 284.4649 | 830.9505 | 383.3067 | 347.6406 | 281.3067 | 577.9233 | 217.5655 | 102.0783 | 73.5128 | 217.5655 | 227.0783 | 184.9744 | 174.9489 | 239.9505 | 196.5655 | 165.9489 | 129.0511 | 233.2045 | 116.3834 | 106.5383 | 93.8467 | 180.8211 | 182.6406 | 159.8467 | 144.3322 | 127.2572 | 512 | 449.9761 | 313.3866 | 524.7444 | 330.925 | 370.4888 | 276.0016 | 499.2333 | 568.925 | 330.5639 | 229.7955 | 721.46
2157 | 25.7214 | 196.4907 | 263.4829 | 231.002 | 367.5034 | 232.4927 | 224.4907 | 182.5034 | 333.4829 | 479.9649 | 224.479 | 161.9971 | 524.4712 | 423.4859 | 197.5015 | 207.002 | 507.9639 | 349.0088 | 257.0117 | 180.4946 | 518.9971 | 288.9776 | 515.9893 | 336.0049 | 339.4781 | 424.9961 | 377.5102 | 301.5063 | 632.9883 | 186.9834 | 168.4956 | 137.9912 | 261.0137 | 262.9941 | 196.4927 | 180.9961 | 151.9951 | 412.0273 | 364.4781 | 287.4985 | 484.9805 | 276.9854 | 201.0039 | 149.0039 | 262.5024 | 104.4976 | 87.0117 | 94.5054 | 202 | 429.4751 | 309.4888 | 295.4966 | 328.4927 | 236.9873 | 174.4888 | 131.4976 | 221.9932 | 384.4976 | 249.0088 | 285.5102 | 384.4898 | 393.9893 | 235.998 | 336.4878 | 408.4907 | 273.0166 | 296.4859 | 306.999 | 630.9854 | 343.999 | 236.4966 | 304.9961 | 501.9893 | 367.96 | 295.4917 | 233.5063 | 648.481 | 330.4878 | 259.4781 | 170.5024 | 494.4663 | 384.4781 | 282.4868 | 233.9932 | 411.9805 | 521.9707 | 419.9951 | 289.4907 | 362.9883 | 189.4917 | 144.9883 | 132.4898 | 223.9815 | 508 | 423.4937 | 313.5122 | 863.9893 | 401.4702 | 344.4976 | 255.5132 | 569.5112 | 251.4839 | 106.5122 | 81.4917 | 251.4839 | 241.5024 | 192.9932 | 173.0039 | 253.5083 | 218.9951 | 179.4888 | 138.4888 | 229.4956 | 107.9932 | 110.9941 | 103.9961 | 184.5034 | 208.4693 | 177.9883 | 138.4927 | 137.4995 | 541.4712 | 465.0059 | 327.0107 | 534.001 | 348.5054 | 400.4907 | 292.0039 | 521.5083 | 586.0059 | 319.0088 | 242.4956 | 748.4556
2158 | 25.7324 | 206.5118 | 274.8584 | 229 | 371.1652 | 256.3777 | 238.0944 | 174.9056 | 349.4646 | 541.5902 | 240.3702 | 165 | 561.1652 | 439.0236 | 198.0472 | 198.8584 | 555.7479 | 357.4013 | 258.3069 | 198.7951 | 539.9131 | 327.3541 | 524.441 | 334.0708 | 359.9528 | 443.8423 | 389.5193 | 307.2833 | 645 | 204 | 181.7007 | 135.2285 | 260.3069 | 272.5826 | 201.441 | 170.6695 | 145.2285 | 416.2436 | 363.457 | 274.6695 | 516.2597 | 301.2125 | 186.7639 | 136.2993 | 259.4882 | 113.6534 | 81.6534 | 91.0472 | 211.2125 | 441.1813 | 327.6534 | 295.4174 | 347.2597 | 241.8111 | 180.882 | 139.6298 | 226.9528 | 388.0236 | 236.4174 | 286.2597 | 409.8423 | 418.8187 | 248.2361 | 330.0633 | 444.6138 | 269.3069 | 314.5826 | 312.6062 | 635.2521 | 371.102 | 241.5354 | 326.9131 | 538.5902 | 419.2361 | 318.8423 | 247.4721 | 692.5666 | 319.457 | 285.5826 | 179.2597 | 519.7875 | 414.1652 | 309.3069 | 214.3862 | 436.6062 | 622.1171 | 419.3702 | 313.8423 | 363.7403 | 202.0944 | 162.118 | 120.9925 | 253.7479 | 531.543 | 446.3777 | 302.0236 | 881.6534 | 416.1341 | 342.9056 | 272.1964 | 575.4013 | 268 | 102.1889 | 90.5118 | 268 | 256.4013 | 192.8348 | 178.2125 | 273.1492 | 242.9367 | 188.441 | 139.2521 | 234.5118 | 125.7479 | 117.5118 | 100.8348 | 187.1416 | 225.1577 | 199.2125 | 154.1889 | 144.1416 | 586.3541 | 470.2597 | 313.441 | 521.2285 | 365.5193 | 423.8187 | 282.3702 | 525.2833 | 598.4249 | 313.0708 | 251.6062 | 803.2125
2159 | 25.7434 | 185.5121 | 253.2772 | 231.0964 | 395.8192 | 282.4819 | 241.4759 | 173.0964 | 366.8674 | 584.8674 | 236.5723 | 152.9458 | 597.8734 | 453.6265 | 205.2409 | 189.8554 | 569.1446 | 365.6145 | 284.1024 | 193.7049 | 556.4759 | 343.0482 | 536.1505 | 338.0482 | 362.7168 | 432.3194 | 406.8072 | 317.9518 | 649.1928 | 190.8976 | 194.1928 | 127.1446 | 289.771 | 280.1928 | 215.2469 | 161.7168 | 130.3313 | 466.3915 | 369.3011 | 275.6746 | 520.1868 | 305.8072 | 200.5842 | 146.8674 | 277.8674 | 114.7591 | 90.6205 | 85.1386 | 213.1868 | 465.2107 | 334.5241 | 301.4337 | 351.1868 | 261.777 | 176 | 135.0422 | 241.2469 | 405.771 | 227.7591 | 319.012 | 434.4819 | 451.9156 | 260.0964 | 346.0661 | 484.1505 | 288.2891 | 320.6205 | 311.235 | 692.0841 | 410.6746 | 226.229 | 331.4217 | 567.1928 | 418.5181 | 333.5241 | 250.229 | 735.3915 | 331.5902 | 297.3855 | 206.771 | 557.6444 | 438.8192 | 334.0542 | 224.1143 | 428.9458 | 660.1748 | 441.777 | 336.9096 | 377.6324 | 210.7168 | 179.5783 | 102.6205 | 252.994 | 579.6806 | 472.4819 | 298.2832 | 919.4457 | 402.0482 | 351.0542 | 310.9578 | 575.7531 | 248.6085 | 107.3795 | 82.6145 | 248.6085 | 287.1505 | 192.2891 | 191.1928 | 302.0964 | 246.8495 | 179.1868 | 130.0482 | 251.2469 | 132.8554 | 123.765 | 100.2891 | 219.7288 | 240.8734 | 201.1868 | 153.0904 | 147.9036 | 661.2709 | 512.9698 | 318.3373 | 532.5361 | 385.9518 | 441.7168 | 278.0482 | 563.7288 | 646.9216 | 319.1446 | 236.0844 | 810.9518
2160 | 25.7544 | 191.2101 | 248.6549 | 238.9002 | 406.6361 | 294.6819 | 253.8733 | 181.4366 | 375.4178 | 611.1186 | 254.6279 | 146.8275 | 677.5658 | 472.973 | 236.2829 | 196.1186 | 600.4284 | 403.0564 | 311.0188 | 181.2183 | 617.6843 | 346.6819 | 550.6092 | 344.9002 | 386.3099 | 454.5928 | 420.7007 | 316.4636 | 673.919 | 197.7735 | 200.6819 | 126.7817 | 298.027 | 296.3369 | 209.081 | 159.027 | 128.6092 | 490.973 | 397.6819 | 296.5094 | 541.892 | 315.946 | 217.7089 | 148.9812 | 308.946 | 101.4178 | 101.0458 | 78.5364 | 201.0998 | 510.2641 | 324.8087 | 299.8087 | 369.1373 | 267.1537 | 179.2183 | 130.2911 | 273.1643 | 445.6737 | 248.7465 | 344.9002 | 462.2371 | 516.1291 | 269.5094 | 381.8275 | 517.919 | 282.1995 | 339.6279 | 316.7277 | 758.3639 | 427.7547 | 227.0916 | 322.1455 | 566.1725 | 422.946 | 331.8545 | 232.8545 | 740.1267 | 357.0998 | 305.5364 | 210.1995 | 627.892 | 464.6549 | 355.1913 | 260.2911 | 461.4472 | 692.466 | 488.4554 | 349.2911 | 420.919 | 212.8545 | 182.4178 | 122.1643 | 260.1643 | 633.574 | 476.6361 | 299.3639 | 986.6198 | 420.7007 | 361.4636 | 308.7359 | 562.6092 | 259.4284 | 113.0458 | 76.0728 | 259.4284 | 321.4554 | 209.8005 | 183.1995 | 309.9002 | 236.1455 | 179.973 | 137.4366 | 275.6549 | 132.6819 | 122.5634 | 99.027 | 237.5634 | 261.5634 | 206.8005 | 148.2183 | 150.8275 | 696.1537 | 554.8733 | 335.7277 | 561.0458 | 396.2641 | 449.7547 | 296.7007 | 593.9002 | 704.5012 | 318.7817 | 233.5552 | 834.1373
2161 | 25.7654 | 216.8407 | 278.3984 | 230.2831 | 426.6904 | 293.1593 | 271.0354 | 190.2921 | 405.31 | 629.9735 | 275.0354 | 158.1328 | 724.0974 | 485.0354 | 240.8937 | 191.3804 | 642.009 | 446.3894 | 341.5576 | 204.8497 | 692.9471 | 386.31 | 569.5576 | 362.6196 | 412.8762 | 510.4603 | 431.6106 | 312.1593 | 731.5047 | 241.4338 | 211.2301 | 125.646 | 322.1772 | 309.7434 | 208.009 | 166.1682 | 140.4249 | 518.3984 | 388.4778 | 312.8762 | 561.0619 | 331.2921 | 225.5222 | 159.5931 | 319.354 | 107.1682 | 99.7699 | 88.3275 | 201.4868 | 569.9296 | 349.4693 | 278.3804 | 407.5576 | 295.5401 | 190.2301 | 134.5487 | 287.4513 | 473.5841 | 259.4778 | 343.9646 | 509.2746 | 583.4338 | 293.009 | 410.6904 | 549.1682 | 267.6106 | 360.0354 | 330.9381 | 805.7968 | 444.7169 | 266.7878 | 336.6196 | 606.4428 | 457.4957 | 333.2921 | 217.2831 | 759.7259 | 369.5576 | 323.009 | 193.9646 | 670.1063 | 496.5931 | 395.8232 | 256.3185 | 546.7349 | 786.3454 | 542.664 | 353 | 485.0888 | 228.009 | 177.7434 | 124.3804 | 273.354 | 677.009 | 474.1947 | 290.2831 | 1011.2212 | 467.8232 | 343.991 | 293.5132 | 587.0444 | 287.2921 | 104.6371 | 79.7434 | 287.2921 | 344.3894 | 225.4868 | 177.9381 | 340.7878 | 272.5666 | 209.0444 | 165.4957 | 306.4957 | 138.8407 | 132.9116 | 91.354 | 239.1328 | 280.6904 | 215.354 | 151 | 155 | 727.2835 | 570.9381 | 337.3185 | 578.4249 | 430.1418 | 490.31 | 307.6106 | 643.9915 | 736.0354 | 318.7434 | 244.8053 | 868.1682
2162 | 25.7764 | 213.2684 | 298.8048 | 253.78 | 453.7316 | 309.0729 | 267.5855 | 186.8294 | 446.6584 | 693.9748 | 272.1465 | 157.8294 | 795.6819 | 512.9997 | 229.3658 | 175.9513 | 669.4632 | 471.8777 | 349.2684 | 206.2929 | 731.1948 | 407.4636 | 625.5116 | 353.3661 | 421 | 544.1948 | 450.9997 | 328.0729 | 778.0242 | 252.7806 | 198.3661 | 135.9755 | 336.9026 | 324.3413 | 223.1219 | 166.9026 | 145.0732 | 551.7071 | 394.7071 | 303.049 | 561.4877 | 337.3658 | 229.9513 | 175.0732 | 333.1461 | 111.2684 | 94.1219 | 92.0732 | 223.39 | 678.8771 | 374.7561 | 321.2919 | 444.9997 | 334.1219 | 189.7074 | 140.0487 | 284.7561 | 503.6826 | 268.5119 | 372.0971 | 527.7074 | 648.0726 | 317.0974 | 432.1219 | 588.6093 | 270.1706 | 384.0729 | 346.7803 | 859.5119 | 487.9751 | 285.7561 | 343.6342 | 676.0971 | 465.3416 | 389.8529 | 222.829 | 808.5606 | 419.341 | 313.4394 | 196.8535 | 671.5368 | 567.048 | 406.8539 | 283.0971 | 597.2194 | 867.9993 | 609.5848 | 373.7558 | 552.1703 | 238.6342 | 187.8535 | 137.5606 | 276.4877 | 733.0726 | 506.8532 | 272.2684 | 1065.7064 | 490.0732 | 345.7071 | 297.9755 | 617.7803 | 275.9758 | 108.829 | 84.2439 | 275.9758 | 378.2922 | 228.8781 | 196.0242 | 359.1952 | 319.5119 | 225.1952 | 176.1465 | 351.3655 | 153.7803 | 138.1465 | 101.7803 | 270.2435 | 297.6342 | 216.2439 | 145.9513 | 169.5852 | 811.8774 | 602.4874 | 333.8048 | 625.7068 | 481.4142 | 524.9268 | 304.561 | 693.9023 | 756.1461 | 325.4877 | 245.8048 | 939.5845
2163 | 25.7874 | 224.6245 | 283.229 | 251.7638 | 444.0953 | 341.771 | 274.7582 | 192.8919 | 461.8791 | 787.4468 | 284.3443 | 148.9872 | 879.6629 | 564.249 | 241.1721 | 202.956 | 708.0897 | 492.4396 | 383.5548 | 212.3443 | 824.3899 | 416.6245 | 641.9488 | 348.3187 | 458.2618 | 562.293 | 456.9616 | 327.522 | 799.8919 | 257.7582 | 198.478 | 144.1465 | 354.7966 | 339.7326 | 225.1465 | 154.9872 | 150.5989 | 549.0953 | 425.9175 | 308.4908 | 578.1849 | 360.0641 | 225.4267 | 162.2546 | 347.293 | 116.3187 | 95 | 87.2802 | 252.478 | 725.1209 | 381.0256 | 337.7766 | 476.7582 | 353.3443 | 178.8407 | 156.0384 | 293.8919 | 562.5548 | 292.1721 | 402.0128 | 561.5548 | 689.0256 | 327.293 | 457.0769 | 655.1409 | 278.4396 | 413.9047 | 368.6117 | 898.6501 | 544.5292 | 262.7894 | 350.1721 | 720.3443 | 491.9688 | 387.2858 | 233.707 | 860.249 | 456.7454 | 349.4596 | 216.1849 | 675.9047 | 623.0512 | 406.4652 | 296.9616 | 613.4524 | 926.9303 | 642.4396 | 378.5348 | 587.0384 | 258.9303 | 198.5861 | 143.3883 | 302.3571 | 809.5677 | 532.7198 | 270.4396 | 1108.2674 | 491.0128 | 347.1081 | 300.9872 | 625.8535 | 305.7014 | 117.9872 | 78.5476 | 305.7014 | 428.3699 | 245.771 | 202.4139 | 391.9688 | 343.1721 | 210.3883 | 170.5733 | 370.8535 | 172.1721 | 130.8535 | 115.5861 | 302.0384 | 303.0256 | 230.0384 | 161.4908 | 189.0256 | 882.542 | 663.5548 | 352.0512 | 685.2362 | 539.249 | 558.956 | 322.1978 | 735.5292 | 835.0713 | 333.5861 | 243.9872 | 1017.956
2164 | 25.7984 | 224.7035 | 284.6387 | 241.1254 | 467.5474 | 343.8746 | 278.559 | 229.1064 | 471.0988 | 881.7983 | 318.4942 | 163.5665 | 962.4827 | 629.9885 | 245.6578 | 206.0913 | 751.2509 | 521.3497 | 392.5058 | 217.8023 | 887.8555 | 442.4676 | 643.4676 | 359.5133 | 493.0266 | 596.7908 | 498.9428 | 325.9543 | 860.6995 | 290.8364 | 219.9543 | 126.2624 | 374.5855 | 384.4029 | 247.0798 | 150.8289 | 155.8289 | 576.0607 | 437.6578 | 318.3156 | 584.4867 | 400.7642 | 248.422 | 183.6919 | 356.0266 | 125.1711 | 115.4942 | 87.7566 | 252.8746 | 756.3041 | 399.8821 | 321.9277 | 509.8364 | 385.152 | 187.2965 | 156.2168 | 309.6121 | 615.422 | 294.9012 | 400.3156 | 622.6197 | 746.5284 | 330.7566 | 473.8289 | 776.418 | 296.8098 | 396.3081 | 388.5399 | 986.6804 | 571.5133 | 281.5474 | 339.4335 | 769.7185 | 536.152 | 394.4029 | 237.8555 | 903.7376 | 484.4942 | 368.9202 | 205.5058 | 756.0949 | 714.0492 | 452.4752 | 312.0075 | 665.844 | 1033.5891 | 695.9428 | 385.2243 | 646.3573 | 264.289 | 201.4145 | 135.7566 | 333.6387 | 878.8896 | 567.3763 | 285.2965 | 1194.1746 | 499.8555 | 340.7566 | 280.4335 | 600.4069 | 328.3879 | 116.7566 | 81.1977 | 328.3879 | 452.0988 | 276.5665 | 216.5665 | 420.9277 | 350.5855 | 229.6919 | 167.4867 | 377.6121 | 187.7832 | 134.0988 | 104.9468 | 321.5399 | 337.6919 | 259.4942 | 164.289 | 193.8289 | 932.422 | 711.7376 | 362.8289 | 702.3156 | 590.9353 | 579.0723 | 351.3954 | 780.0798 | 892.6121 | 328.8023 | 242.7566 | 1036.9012
2165 | 25.8094 | 230.946 | 285.4345 | 249.5978 | 506.946 | 350.3482 | 274 | 227.4777 | 527.8057 | 960.2618 | 326.4238 | 186.1633 | 1065.1862 | 688.2942 | 235.0324 | 205.1525 | 817.2186 | 563.3158 | 415.664 | 217.9784 | 988.2524 | 482.4899 | 709.7949 | 369.1741 | 508.7611 | 666.6316 | 529.0216 | 332.4022 | 870.4993 | 344.0661 | 219.8259 | 125.5547 | 383.9676 | 436.9136 | 240.4777 | 148.2065 | 162.1741 | 613.7611 | 452.1417 | 310.8151 | 569.4453 | 433.946 | 260.2173 | 177.0972 | 378.9352 | 116.4345 | 109.6734 | 86.6086 | 246.7935 | 838.7949 | 398.8367 | 330.5762 | 563.0661 | 380.0648 | 222.9028 | 150.5978 | 332.3374 | 669.0661 | 299.1741 | 416.1309 | 652.7719 | 798.5439 | 364.2834 | 504.6856 | 846.7719 | 305.6086 | 420.2402 | 401.9784 | 1108.3387 | 631.9905 | 315.5655 | 362.2726 | 855.3819 | 589.0553 | 418.8151 | 238.413 | 963.6316 | 526.5007 | 354 | 189.2065 | 816.9568 | 792.7179 | 498.7287 | 337.3698 | 742.0445 | 1163.3603 | 766.6748 | 411.5439 | 699.3374 | 277.5223 | 235.0769 | 128.6302 | 345.7935 | 941.4683 | 619.6964 | 310.7395 | 1327.3495 | 517.1525 | 357.5439 | 281.7503 | 583.5978 | 356.4791 | 120.9892 | 91.7827 | 356.4791 | 458.587 | 290.1957 | 252.9136 | 473.2186 | 365.946 | 262.5547 | 177.5547 | 421.2618 | 196.9892 | 139.3914 | 95 | 337.9676 | 371.1525 | 309.8704 | 176.9244 | 216.3158 | 1023.1322 | 744.1309 | 366.7827 | 725.9028 | 623.7503 | 635.5884 | 368.3806 | 803.9676 | 938.6532 | 336.7503 | 254.1633 | 1088.4035
2166 | 25.8204 | 240.3899 | 296.6938 | 247.5595 | 534.0837 | 400.8788 | 302.0661 | 209.3392 | 573.152 | 1102.5352 | 331.152 | 185.6784 | 1111.7114 | 746.4383 | 250.3546 | 198.7467 | 854.6608 | 643.0638 | 468.4559 | 223.2203 | 1065.2026 | 510.9317 | 757.7621 | 327.0199 | 522.1013 | 744.37 | 583.5726 | 342.3722 | 936.8083 | 337.8833 | 215.6101 | 134.6101 | 404.033 | 466.4229 | 238.8128 | 159.2026 | 185.304 | 656.9978 | 448.967 | 319.5925 | 608.5903 | 439.1189 | 282.0154 | 187.0661 | 409.9141 | 117.3216 | 108.8128 | 92.3216 | 255.3216 | 910.5572 | 420.0661 | 350.2533 | 607.5242 | 403.9978 | 254.4229 | 149.7797 | 355.5925 | 715.3546 | 339.3876 | 437.5418 | 726.9449 | 876.1145 | 412.8458 | 542.4736 | 897.1497 | 305.2203 | 449.3899 | 399.2886 | 1220.4537 | 699.0661 | 339.6938 | 378.6784 | 894.7797 | 642.8458 | 442.8458 | 259.7951 | 1051.7423 | 591.9801 | 403.4207 | 189.2203 | 856.7269 | 829.3546 | 543.7269 | 363.1343 | 790.5242 | 1295.2488 | 841.2511 | 449.4559 | 728.6938 | 296.2709 | 227.4934 | 121.7797 | 375.6762 | 974.5418 | 650.5925 | 315.8987 | 1392.9141 | 561.9978 | 369.2203 | 292.2203 | 607.185 | 408.2863 | 123 | 99.8811 | 408.2863 | 524.4537 | 307.4736 | 271.4405 | 516.7621 | 407.1167 | 281.9824 | 176.2379 | 467.1343 | 208.152 | 167.8458 | 111.4736 | 351.3216 | 392.8128 | 319.967 | 172.0859 | 236.2709 | 1102.8965 | 822.8942 | 384.033 | 764.7445 | 671.4383 | 670.7291 | 373.2203 | 828.304 | 1006.8788 | 360.033 | 243.9163 | 1187.2334
2167 | 25.8314 | 218.8376 | 312.9794 | 241.643 | 568.6521 | 405.2152 | 323.1121 | 189.5722 | 631.9059 | 1140.2771 | 374.3454 | 213.9884 | 1209.3634 | 796.5193 | 273.3363 | 200.826 | 947.2513 | 664.7049 | 497.826 | 236.4485 | 1166.3634 | 568.2835 | 788.8969 | 331.8557 | 555.2539 | 797.6521 | 611.0412 | 367.6727 | 1055.732 | 336.6727 | 225.3363 | 137.4897 | 417.9794 | 481.8467 | 276.2539 | 178.826 | 182.6843 | 693.0709 | 491.0387 | 312.0619 | 643.357 | 464.6314 | 299.2448 | 201.8879 | 465.5696 | 120.1327 | 104.5515 | 88.9085 | 283.6521 | 944.4485 | 462.8969 | 352.5309 | 651.7436 | 435.0915 | 243.0825 | 176.3866 | 382.2951 | 787.5077 | 344.5722 | 446.1121 | 787.8054 | 941.4072 | 444.3157 | 579.7642 | 983.8853 | 322.8054 | 482.6108 | 428.6108 | 1356.2307 | 735.0503 | 357.8467 | 387.6933 | 949.3956 | 714.7732 | 450.0412 | 263.7964 | 1107.2951 | 646.2745 | 448.0709 | 205.5606 | 925.549 | 839.8879 | 565.2448 | 367.2861 | 877.0683 | 1349.7552 | 866.5309 | 463.8879 | 784.8144 | 305.8467 | 224.183 | 126.6018 | 382.7964 | 1028.549 | 645.5515 | 324.4485 | 1471.5992 | 616.7436 | 353.1946 | 284.2861 | 642.54 | 432.6018 | 120.5103 | 116.6933 | 432.6018 | 596.1211 | 337.9175 | 250.5928 | 555.366 | 467.4575 | 307.0503 | 178.0915 | 488.0915 | 235.1624 | 173.174 | 123.8673 | 382.1418 | 423.4072 | 316.6018 | 166.8467 | 233.3982 | 1143.7848 | 880.5606 | 386.1533 | 804.1418 | 708.4485 | 710.3248 | 360.9291 | 825.6843 | 1116.4047 | 376.4691 | 235.4897 | 1266.2126
2168 | 25.8424 | 214.886 | 332.5026 | 235.1917 | 594.4248 | 458.0102 | 347.8497 | 214.544 | 707.1605 | 1168.1424 | 426.1968 | 216.8627 | 1259.0181 | 871.0284 | 290.4248 | 213.0777 | 1052.1424 | 732.9739 | 515.1554 | 260.5026 | 1184.2825 | 550.4923 | 805.9041 | 347.2876 | 606.0051 | 869.1242 | 657.0465 | 388.5207 | 1152.1242 | 365.1373 | 230.3653 | 135.1917 | 456.544 | 502.5026 | 290.8264 | 179.0181 | 193.6761 | 723.215 | 549.9274 | 316.329 | 651.7124 | 524.5258 | 299.3653 | 196.8264 | 503.4248 | 129.1554 | 105.886 | 77.0181 | 305.6166 | 1020.5491 | 509.101 | 362.5984 | 683.6943 | 440.6347 | 251.3108 | 180.3057 | 404.1554 | 873.1424 | 354.3108 | 497.5077 | 826.5621 | 1037.5128 | 477.1191 | 608.6943 | 1093.2615 | 324.7487 | 494.114 | 438.2099 | 1516.1889 | 780.2745 | 372.7901 | 397.2694 | 1062.3392 | 711.8213 | 466.8497 | 249.3834 | 1197.0698 | 652.7487 | 426.8032 | 215 | 999.43 | 912.8962 | 603.4481 | 377.8678 | 932.7901 | 1413.1061 | 932.4532 | 472.1554 | 843.7357 | 335.3886 | 232.7306 | 128.7306 | 379.8083 | 1094.1787 | 683.6994 | 319.9404 | 1562.1424 | 698.8367 | 364.5803 | 273.9223 | 684.1968 | 481.0647 | 129.1554 | 109.7669 | 481.0647 | 634.0414 | 372.215 | 272.544 | 620.0647 | 488.1917 | 319.9041 | 203.3108 | 503.1554 | 244.8264 | 152.671 | 126.2694 | 411.329 | 486.5077 | 349.8316 | 172.9041 | 249.0414 | 1262.9972 | 902.0596 | 399.772 | 827.6166 | 807.3988 | 744.2513 | 373.3108 | 860.1605 | 1188.5803 | 387.9819 | 240.8083 | 1344.5854
2169 | 25.8534 | 245.2332 | 366.5862 | 253.0572 | 645.6433 | 513.4691 | 344.1779 | 222.0599 | 756.5862 | 1186.6479 | 448.176 | 214.7641 | 1288.6443 | 966.4655 | 320.2922 | 206.9419 | 1154.7586 | 804.8802 | 571.4664 | 277.1171 | 1174.2931 | 553.7032 | 853.5844 | 349.823 | 647.6452 | 963.4655 | 704.4691 | 406.2931 | 1247.7595 | 373.706 | 239.7051 | 148.5281 | 482.823 | 537.8802 | 325.2323 | 176.941 | 212.1171 | 790.3484 | 534.9438 | 323.353 | 687.9973 | 565.1751 | 310.6461 | 211.1751 | 539.7622 | 133.706 | 105.177 | 63.2949 | 306.765 | 1111.1134 | 514.0599 | 375.529 | 722.7032 | 455.8811 | 263 | 165.8829 | 428.1161 | 899.412 | 377.6461 | 478.7686 | 898.6424 | 1085 | 501.9991 | 654.8203 | 1159.2922 | 352.9973 | 539.4664 | 459.1742 | 1599.5862 | 791.5889 | 390.6461 | 393.471 | 1113.176 | 755.7005 | 502.6452 | 252.1171 | 1310.5825 | 705.5834 | 472.1125 | 220.823 | 1057.9973 | 943.0599 | 630.823 | 407.7042 | 977.1733 | 1493.9963 | 1026.7014 | 511.6443 | 912.4664 | 345.471 | 257.8802 | 120.8829 | 382.647 | 1114.353 | 723.5871 | 323.3521 | 1662.8175 | 757.6452 | 381.529 | 275.529 | 721.7042 | 487.5899 | 147.9401 | 112.6461 | 487.5899 | 690.9373 | 383.059 | 283.9419 | 642.765 | 515.1152 | 346.8802 | 213.706 | 554.2903 | 243.647 | 155.9991 | 137.9991 | 440.9982 | 513.706 | 359.8829 | 190.8221 | 270.3521 | 1280.6507 | 943.2913 | 442.3502 | 842.3521 | 878.7631 | 770.3512 | 403.1161 | 846.1797 | 1208.1171 | 413.3512 | 237.824 | 1403.1742
2170 | 25.8644 | 266.9116 | 372.7477 | 285.7568 | 702.9648 | 553.3276 | 367.3055 | 229.5045 | 788.4603 | 1159.2432 | 453.6372 | 237.4603 | 1298.4292 | 1046.1507 | 337.6151 | 225.0754 | 1266.204 | 911.3014 | 594.3849 | 275.7256 | 1196.8231 | 572.9779 | 894.8231 | 378.7125 | 679.1196 | 998.9779 | 723.8894 | 407.7256 | 1357.819 | 400.6904 | 262.801 | 148.4292 | 495.8894 | 550.6814 | 373.3497 | 176.0221 | 224.5708 | 870.7658 | 553.602 | 338.8231 | 750.8321 | 608.2834 | 326.8894 | 214.0663 | 582.3718 | 131.0221 | 104.6372 | 71.8452 | 322.801 | 1136.3628 | 552.4914 | 405.0311 | 781.5356 | 478.5045 | 280.1417 | 163.9779 | 412.2875 | 945.4693 | 407.0754 | 514.3145 | 925.0442 | 1142.3587 | 547.5578 | 661.4955 | 1160.1327 | 394.0754 | 612.4914 | 495.0975 | 1663.1065 | 774.4734 | 424.0311 | 431.8763 | 1101.4955 | 803.5266 | 534.1196 | 260.6151 | 1361.1859 | 720.5176 | 503.3628 | 227.9558 | 1122.1507 | 885.8935 | 682.7879 | 428.8894 | 1021.0754 | 1513.3849 | 1104.7879 | 547.801 | 918.9025 | 344.9779 | 273.9779 | 128.8673 | 425.1949 | 1096.8583 | 785.0623 | 304.606 | 1768.5889 | 810.8763 | 419.602 | 266.792 | 735.6372 | 501.394 | 158.9558 | 103.8362 | 501.394 | 724.5708 | 405.7347 | 305.3718 | 665.394 | 548.1417 | 372.208 | 207.0663 | 588.8894 | 245.9779 | 183.0975 | 149.9337 | 467.5045 | 541.3497 | 373.801 | 193.407 | 297.7568 | 1323.5889 | 1027.2703 | 465.9337 | 809.7608 | 863.2875 | 796.5045 | 436.7347 | 778.8493 | 1282.5447 | 428.9558 | 256.0975 | 1398.2211
2171 | 25.8754 | 271 | 379.0738 | 279.5387 | 766.3395 | 590.4908 | 397.4317 | 229.6125 | 837.6236 | 1185.6089 | 491.2657 | 238.2694 | 1225.4981 | 1092.834 | 351.417 | 237.6716 | 1403.7381 | 991.2362 | 596.7011 | 275.0148 | 1211.7159 | 615.6384 | 1006.4244 | 383.941 | 674.8967 | 1086.6347 | 781.3838 | 424.476 | 1441.3247 | 409.6273 | 283.417 | 132.583 | 514.4317 | 551.3432 | 370.5387 | 171.6421 | 222.2841 | 925.8782 | 583.417 | 342.2989 | 751.8376 | 604.5092 | 349.4613 | 226.4465 | 603.3875 | 146.7897 | 116.7454 | 83.0295 | 366.9225 | 1087.9889 | 623.6827 | 417.6568 | 806.3579 | 465.8524 | 293.0295 | 183.1328 | 403.0295 | 984.1033 | 415.6421 | 562.7897 | 900.166 | 1148.4945 | 569.3432 | 677.5055 | 1161.3875 | 434.2067 | 642.0443 | 538.9078 | 1670.8672 | 774.0443 | 491.0554 | 440.2398 | 1114.8192 | 824.7749 | 579.5941 | 244.8672 | 1410.9816 | 676.4059 | 530.8782 | 223.9557 | 1107.7195 | 878.1328 | 718.7454 | 443.4022 | 988.6753 | 1463.3173 | 1127.3137 | 595.2805 | 924.1181 | 350.7011 | 269.6273 | 146.417 | 435.583 | 1005.3948 | 819.0443 | 320.2067 | 1809.0443 | 848.1181 | 413.1513 | 277.1181 | 747.7454 | 559.6679 | 152.941 | 96.6716 | 559.6679 | 720.2694 | 438.1476 | 293.4797 | 682.0295 | 565.059 | 395.4613 | 207.3579 | 600.0443 | 247 | 202.7306 | 157.7011 | 497.834 | 547.2694 | 384.3432 | 227.6089 | 267.3616 | 1340.5387 | 1030.107 | 434.7491 | 802.0886 | 838.583 | 814.7454 | 454.3727 | 787.9078 | 1368.0259 | 450.476 | 277.0738 | 1416.2067
2172 | 25.8864 | 283.9937 | 391.5227 | 269.6322 | 837.5553 | 629.039 | 442.2972 | 217.7935 | 839.2192 | 1136.2961 | 526.097 | 236.7355 | 1181.7418 | 1170.3362 | 359.7355 | 242.1033 | 1372.9863 | 1036.4068 | 660.549 | 317.0327 | 1156.1865 | 613.432 | 1059.0516 | 356.0643 | 659.7935 | 1198.0717 | 825.4647 | 440.1549 | 1553.2719 | 443.2455 | 318.4068 | 143.4131 | 511.4257 | 536.2708 | 356.529 | 185.0453 | 223.7871 | 974.1939 | 632.7682 | 340.6839 | 754.6776 | 600.8388 | 342.2708 | 229.8967 | 555.7091 | 157.7355 | 123.3678 | 95.942 | 396.9937 | 1003.451 | 610.7545 | 418.3678 | 787.4837 | 466.942 | 325.7745 | 159.1739 | 403.6322 | 986.2129 | 460.5037 | 572.3678 | 842.4963 | 1129.4773 | 624.7102 | 737.5553 | 1068.5732 | 466.4647 | 682.6133 | 545.4257 | 1628.49 | 770.8451 | 471.1802 | 438.7871 | 1090.8061 | 788.9157 | 589.4773 | 240.7871 | 1438.8388 | 661.7355 | 537.4773 | 218.9484 | 1075.3224 | 893.8388 | 722.6322 | 490.4004 | 953.3224 | 1403.0706 | 1180.0264 | 625.6259 | 917.6386 | 346.1612 | 265.6322 | 129.432 | 412.903 | 993.7229 | 846.6196 | 331.2645 | 1808.5806 | 910.0264 | 452.9747 | 280.8967 | 781.7229 | 522.5026 | 159.942 | 105.8904 | 522.5026 | 747.4584 | 455.5227 | 281 | 645.7028 | 543.0126 | 372.5416 | 217.8904 | 613.942 | 247 | 230.9357 | 146.3224 | 487.1159 | 542.3161 | 400.7292 | 222.7481 | 248.3678 | 1289.5996 | 1037.568 | 474.7619 | 762.1802 | 821.3741 | 851.4584 | 461.7871 | 786.903 | 1346.864 | 475.7292 | 251.2255 | 1461.4584
2173 | 25.8974 | 276.7728 | 418.6698 | 294.0621 | 843.9881 | 703.9591 | 471.8349 | 216.0885 | 783.6605 | 1068.103 | 595.9591 | 263.0621 | 1097.6367 | 1288.7173 | 367.2655 | 269.7583 | 1413.9709 | 1107.782 | 656.2801 | 329.7345 | 1046.1559 | 595.3421 | 1094.8085 | 368.6698 | 651.823 | 1248.1004 | 857.6698 | 439.5192 | 1668.7939 | 452.3421 | 354.9736 | 146.1268 | 509.0885 | 523.4306 | 380.4544 | 195.4808 | 251.7583 | 1058.9207 | 663.4425 | 348.6579 | 825.7437 | 628.3659 | 383.0357 | 234.9617 | 570.9855 | 175.0119 | 118.4306 | 103.7847 | 394.646 | 948.5456 | 636.4663 | 420.3923 | 760.5958 | 463.646 | 345.5694 | 176.7199 | 421.1004 | 1006.9736 | 493.8349 | 589.7081 | 820.3038 | 1064.0409 | 682.0238 | 782.2774 | 1024.6962 | 477.0885 | 721.5812 | 582.7701 | 1603.9498 | 706.7371 | 507.782 | 478.5931 | 1033.2299 | 805.897 | 623.5931 | 266.6698 | 1377.9524 | 656.7345 | 559.0621 | 232.6196 | 1048.8111 | 825.9907 | 747.0621 | 521.8349 | 910.7992 | 1367.8996 | 1122.1559 | 644.2272 | 872.0145 | 334.646 | 262.9115 | 139.1889 | 387.5958 | 944.3448 | 873.3157 | 307.0264 | 1765.926 | 956.5812 | 480.4808 | 311.7199 | 795.6962 | 511.0119 | 179.6196 | 118.354 | 511.0119 | 757.823 | 508.428 | 346.4399 | 640.531 | 499.6724 | 382.3659 | 246.3659 | 614.823 | 233.7728 | 203.7107 | 145.5694 | 488.1387 | 496.749 | 432.3659 | 248.2008 | 293.5548 | 1230.3184 | 1033.3804 | 498.9115 | 716.6341 | 767.9643 | 877.8349 | 505.7701 | 695.4597 | 1355.2391 | 486.4808 | 241.8732 | 1393.1559
2174 | 25.9084 | 275.9592 | 419.79 | 306.4169 | 898.636 | 702.6616 | 466.3731 | 248.1722 | 745.997 | 966.9848 | 635.7523 | 289.5861 | 979.4017 | 1296.5362 | 378.5015 | 281 | 1410.4516 | 1103.9531 | 705.053 | 351.0876 | 963.2416 | 582.79 | 1131.9623 | 361.997 | 618.1193 | 1355.1844 | 898.4638 | 434.4577 | 1773.9714 | 424.9124 | 334.5362 | 144.7085 | 483.787 | 463.6147 | 406.1692 | 192.0408 | 260.8746 | 1108.8369 | 697.0469 | 342.0816 | 912.8022 | 592.2416 | 429.5046 | 231.6239 | 577.3731 | 175.6239 | 117.4169 | 89.4139 | 381.79 | 918.58 | 663.5015 | 444.3792 | 714.0347 | 420.3262 | 352.2508 | 205.1692 | 376.5739 | 953.947 | 518.1284 | 609.0438 | 776.7839 | 994.4924 | 724.3384 | 891.1844 | 972.5739 | 541.7614 | 779.8838 | 581.8716 | 1536.2386 | 631.1162 | 512.4547 | 505.6269 | 968.2416 | 779.4924 | 642.1254 | 289.0438 | 1324.4954 | 632.0377 | 565.7492 | 227.6646 | 975.5301 | 751.7839 | 777.1284 | 520.6239 | 839.7401 | 1226.8095 | 1170.3068 | 671.2538 | 760.7748 | 321.79 | 279.003 | 125.3293 | 388.5015 | 894.3293 | 884.2508 | 346.5922 | 1673.6117 | 997.8807 | 464.997 | 284.9093 | 864.7206 | 499.58 | 193.0846 | 110.6646 | 499.58 | 715.6178 | 515.1223 | 365.0816 | 623.3293 | 476.79 | 407.8777 | 232.2039 | 586.787 | 240.0438 | 204.003 | 138.79 | 483.79 | 491.1692 | 442.2915 | 264.7085 | 288.2039 | 1183.2447 | 1006.4547 | 510.7523 | 697.4985 | 725.7462 | 839.0755 | 521.1662 | 638.4547 | 1305.6555 | 547.5107 | 253.21 | 1293.2386
2175 | 25.9194 | 288.0905 | 431.1358 | 281.7737 | 927.6038 | 745.0563 | 503.0828 | 241.5397 | 719.5397 | 925.5132 | 651.0905 | 315.4603 | 896.7097 | 1300.5773 | 407.2263 | 276.6755 | 1378.4227 | 1132.3621 | 729.5585 | 342.2605 | 885.6192 | 527.8267 | 1159.0188 | 365.8113 | 598.5132 | 1317.4381 | 944.1546 | 477.6866 | 1793.9095 | 394.819 | 336.415 | 152.2075 | 485.2528 | 432.0717 | 412.7208 | 194.3245 | 270.8113 | 1166.7318 | 771.7054 | 356.7395 | 933.3962 | 532.6379 | 455.1358 | 239.0905 | 563.6302 | 185.2528 | 120.1623 | 95.2528 | 382.3245 | 872.0794 | 692.2263 | 475.6225 | 669.17 | 411.9283 | 385.7131 | 180.7284 | 391.0376 | 937.3698 | 557.7131 | 623.3698 | 698.4569 | 906.3664 | 761.947 | 858.4834 | 903.4304 | 578.0905 | 797.117 | 610.0376 | 1412.6986 | 617.4868 | 571.0299 | 512.883 | 881.2494 | 716.5927 | 671.1093 | 291.8377 | 1308.9547 | 632.3698 | 561.8377 | 230.9283 | 910.8002 | 677.7814 | 755.4492 | 540.3433 | 791.7737 | 1055.2383 | 1136.5286 | 705.2263 | 697.4945 | 279.8002 | 293.9283 | 136.181 | 345.1512 | 856.17 | 866.5397 | 333.8454 | 1614.8907 | 995.8642 | 456.5585 | 280.2528 | 942.7318 | 474.702 | 187.351 | 124.0188 | 474.702 | 696.117 | 544.3168 | 401.3621 | 570.0342 | 467.234 | 419.766 | 223.4415 | 526.2682 | 221.936 | 218.2075 | 147.2528 | 432.4304 | 471.053 | 440.5585 | 276.532 | 250.6115 | 1148.6567 | 992.9547 | 488.6115 | 665.8907 | 680.2417 | 813.7925 | 524.3245 | 615.585 | 1237.4757 | 601.5508 | 264.9283 | 1236.7284
2176 | 25.9304 | 314.7243 | 467.7878 | 288.86 | 907.0087 | 734.4792 | 541.7243 | 250.86 | 703.2035 | 842.3653 | 671.86 | 304.6739 | 835.4835 | 1296.0044 | 444.8556 | 280.8643 | 1364.6696 | 1097.4201 | 708.4748 | 355.5252 | 842.4748 | 494.8052 | 1240.0373 | 381.7287 | 574.7418 | 1365.4355 | 1047.8972 | 492.0678 | 1725.4923 | 357.6783 | 336.1357 | 156.4661 | 432.8227 | 410.6739 | 423.9956 | 198.1991 | 271.3348 | 1239.5121 | 804.0635 | 354.4705 | 990.6434 | 531.6608 | 422.1488 | 234.9365 | 543.14 | 190 | 101.9409 | 88.2713 | 354.6783 | 809.3523 | 739.3852 | 500.86 | 634.5426 | 395.9409 | 386.2713 | 184.3939 | 376.4114 | 841.3062 | 579.5295 | 671.7156 | 683.0635 | 840.1488 | 798.3895 | 877.6477 | 855.6783 | 578.3348 | 830.453 | 661.3852 | 1251.8489 | 575.2844 | 621.1904 | 527.9278 | 844.0044 | 680.6739 | 681.267 | 271.2079 | 1241.7592 | 576.6236 | 565.3983 | 250.86 | 859.3479 | 628.8774 | 743.267 | 540.2035 | 732.8862 | 967.4835 | 1060.1531 | 710.6017 | 671.8731 | 289.1904 | 296.267 | 125.6739 | 310.8731 | 813.4792 | 861.9322 | 329.7965 | 1556.3523 | 981.2035 | 468.4617 | 301.86 | 989.1226 | 435.0131 | 175.2035 | 102.4114 | 435.0131 | 623.1619 | 576.3261 | 414.8009 | 566.9278 | 448.14 | 439.593 | 257.7199 | 512.8643 | 223.2626 | 221.733 | 172.5252 | 398.6061 | 433.8774 | 448.7965 | 275.1357 | 214.3435 | 1074.0262 | 920.6279 | 510.9869 | 643.2713 | 652.0722 | 831.5252 | 532.5974 | 557.221 | 1155.7592 | 659.1817 | 246.7418 | 1163.4244
2177 | 25.9414 | 334.9253 | 500.6147 | 286.056 | 883.348 | 731.1987 | 579.8133 | 232.404 | 635.9013 | 752.628 | 754.7693 | 316.6333 | 758.1373 | 1282.3107 | 473.6333 | 274.056 | 1261.38 | 1062.6027 | 708.4533 | 368.9627 | 753.0133 | 499.6893 | 1261.7827 | 391.2173 | 535.696 | 1346.2987 | 1101.8693 | 461.1867 | 1709.4533 | 341.0373 | 323.5653 | 158.4907 | 409.0187 | 406.9813 | 418.3107 | 197.5093 | 237.2053 | 1194.156 | 775.696 | 378.5773 | 1060.4467 | 517.6213 | 444.5213 | 232.7453 | 501.224 | 168.3853 | 107.6707 | 98.9067 | 341.0187 | 716.7213 | 784.304 | 498.056 | 556.684 | 359.932 | 406.596 | 191.4907 | 312.336 | 803.5093 | 632.9373 | 699.18 | 639.7893 | 772.336 | 816.1987 | 903.1987 | 763.0133 | 668.676 | 869.0867 | 701.0867 | 1187.6213 | 534.404 | 603.168 | 552.3787 | 776.6467 | 602.448 | 672.8013 | 301.2667 | 1125.5787 | 528.9133 | 540.168 | 251.0373 | 773.7027 | 583.224 | 715.4227 | 563.832 | 661.336 | 836.4733 | 1065.1053 | 732.8507 | 576.56 | 276.3853 | 301.2173 | 130.18 | 282.64 | 743.864 | 896.54 | 377.72 | 1465.9573 | 933.28 | 519.2107 | 316.1987 | 1060.1173 | 376.7893 | 183.18 | 111.3787 | 376.7893 | 532.156 | 613.5587 | 398.348 | 525.7893 | 391.3173 | 431.8387 | 284.1613 | 479.9693 | 201.6587 | 227.2173 | 180.7453 | 358.7147 | 385.2427 | 437.0933 | 288.652 | 202.0187 | 951.616 | 876.3667 | 555.0493 | 613.6587 | 611.9693 | 787.572 | 559.596 | 507.4413 | 1056.7213 | 649.9133 | 252.416 | 1079.8827
2178 | 25.9524 | 349.6067 | 533.7592 | 278.4543 | 855.2713 | 716.5746 | 625.6083 | 260.366 | 629.9101 | 672.1477 | 743.8764 | 349.5169 | 686.9968 | 1266.878 | 472.4238 | 308.8812 | 1214.1204 | 934.2648 | 710.2424 | 395.244 | 685.6035 | 465.1188 | 1294.0931 | 385.4238 | 487.8764 | 1253.4799 | 1069.6035 | 513.1252 | 1619.5698 | 317.7865 | 357.1236 | 166.5762 | 401.939 | 384.5136 | 463.488 | 201.5457 | 251.7592 | 1136.3917 | 837.7319 | 406.183 | 1117.9117 | 473.8764 | 484.2745 | 260.2745 | 445.0578 | 144.3323 | 119.5762 | 86.5746 | 336.2119 | 684.1814 | 827.5778 | 509.3949 | 506.4831 | 315.9069 | 370.3002 | 190.4848 | 290.2119 | 701.4783 | 668.183 | 697.6966 | 559.6019 | 696.4205 | 830.3644 | 921.9101 | 685.7544 | 753.7913 | 921.1846 | 674.3917 | 1084.0241 | 483.573 | 598.5457 | 570.3644 | 707.2697 | 549.2103 | 675.3034 | 320.061 | 1100.8186 | 482.8764 | 541.6067 | 260.3644 | 736.1204 | 509.6324 | 664.3307 | 579.061 | 608.9069 | 736.39 | 988.9037 | 766.0016 | 498.2697 | 282.6372 | 289.3628 | 139.3034 | 267.4238 | 669.451 | 870.8764 | 367.9984 | 1344.0241 | 901.3323 | 594.0947 | 321.2729 | 1115.8507 | 363.2729 | 176.3933 | 119.5152 | 363.2729 | 517.5762 | 561.1172 | 409.6677 | 482.7255 | 395.9711 | 460.3355 | 279.9085 | 449.817 | 225.5778 | 212.3323 | 179.4848 | 376.5474 | 351.5441 | 411.5441 | 285.6661 | 204.7881 | 886.9374 | 852.5746 | 588.7287 | 539.0867 | 524.2376 | 763.1814 | 583.1525 | 470.4831 | 1012.0594 | 679.6388 | 265.3339 | 1033.2408
2179 | 25.9634 | 421.5225 | 598.9737 | 300.0975 | 777.9377 | 697.1415 | 652.3983 | 282.7788 | 584.1856 | 583.7076 | 747.3983 | 372.6284 | 578.6101 | 1197.5572 | 492.3275 | 336.3983 | 1102.2118 | 883.4513 | 708.4602 | 395.3008 | 618.336 | 429.9025 | 1259.575 | 398.3983 | 476.7699 | 1125.8313 | 1070.0886 | 525.3008 | 1490.6809 | 334.0975 | 351.2919 | 166.6903 | 352.2652 | 369.3008 | 432.8051 | 223.7877 | 226.8936 | 1135.0089 | 909.5047 | 403.5309 | 1237.2483 | 428.1237 | 529.186 | 250.522 | 427.9203 | 162.8674 | 148.947 | 87.1593 | 305.7432 | 636.575 | 813.3627 | 556.3453 | 465.6636 | 305.7699 | 385.2568 | 193.8496 | 290.5398 | 655.9114 | 713.2657 | 695.2301 | 500.5042 | 594.8491 | 872.4958 | 890.0441 | 619.9555 | 807.186 | 915.7521 | 661.4602 | 911.3356 | 449.2123 | 622.3275 | 580.9292 | 641.9555 | 556.9381 | 708.5665 | 313.5309 | 989.0525 | 415.5661 | 541.9203 | 283.2479 | 677.336 | 433.0263 | 624.053 | 600.2479 | 531.0174 | 710.0708 | 901.9466 | 809.4161 | 486.0797 | 223.097 | 274.2212 | 141.7699 | 271.1593 | 591.9466 | 805.8758 | 394.6462 | 1236.7784 | 881.3716 | 602.9911 | 303.522 | 1204.7615 | 382.478 | 163.761 | 121.5398 | 382.478 | 475.3449 | 573.3364 | 443.4869 | 402.3979 | 365.2741 | 454.0619 | 287.0089 | 417.053 | 184.4157 | 187.7521 | 162.0619 | 339.0352 | 309.5839 | 393.4513 | 283 | 182.1326 | 829.2652 | 796.9555 | 592.9203 | 549.5665 | 505.3894 | 712.4953 | 573.3716 | 432.7432 | 973.6636 | 703.5487 | 258.761 | 966.4864
2180 | 25.9744 | 406.8007 | 603.4847 | 308.5644 | 730.405 | 652.3253 | 661.6932 | 263.316 | 512.6811 | 545.0093 | 773.1196 | 409.8527 | 495.4633 | 1106.8192 | 556.8833 | 351.1687 | 968.3132 | 861.227 | 671.2363 | 445.4078 | 579.316 | 376.8498 | 1252.2577 | 395.9601 | 437.8897 | 1076.227 | 1016.1167 | 555.0705 | 1435.9694 | 313.6228 | 338.9601 | 164.4356 | 330.1779 | 337.2762 | 454.1993 | 224.5246 | 211.5246 | 1060.5616 | 931.0399 | 393.9601 | 1344.0918 | 379.6719 | 520.4448 | 273.9416 | 377.721 | 179.3865 | 157 | 93.6932 | 293.089 | 540.0861 | 838.0705 | 561.1779 | 409.0676 | 290.3559 | 350.1964 | 209.8619 | 264.405 | 621.4939 | 765.6747 | 702.8221 | 438.1566 | 548.8804 | 882.4356 | 864.9203 | 493.1445 | 852.4171 | 916.2577 | 665.6932 | 769.6598 | 379.6413 | 579.721 | 597.8619 | 646.5552 | 493.1658 | 651.5124 | 296.9203 | 919.8007 | 409.1687 | 528.4847 | 299.9509 | 560.8776 | 390.9694 | 581.5829 | 596.6135 | 485.1872 | 589.8868 | 861.316 | 820.7822 | 433.81 | 211.8221 | 270.2178 | 154.5153 | 251.8804 | 556.7915 | 734.4633 | 408.1288 | 1191.316 | 835.5431 | 659.4477 | 296.4356 | 1269.4569 | 340.2854 | 157.8712 | 116.5246 | 340.2854 | 426.8007 | 592.3865 | 416.8007 | 363.0093 | 337.5737 | 427.8804 | 302.5153 | 345.6413 | 169 | 191.3865 | 152.3068 | 279.632 | 295.8712 | 355.5829 | 303.3374 | 171.089 | 757.1167 | 731.1566 | 619.3772 | 509.721 | 444.4235 | 609.6107 | 592.4662 | 408.3559 | 849.7979 | 747.6747 | 281.8128 | 949
2181 | 25.9854 | 423.2184 | 631.7799 | 343.9579 | 716.644 | 589.3577 | 726.5598 | 291.1634 | 459.6586 | 447.8383 | 777.411 | 398.3431 | 435.5081 | 1048.8641 | 584.123 | 338.9045 | 882.7962 | 791.6457 | 599.0292 | 413.1246 | 506.2622 | 346.5211 | 1184.8787 | 441.6699 | 419.8495 | 971.6878 | 1072.2993 | 590.2184 | 1309.0309 | 258.3301 | 322.699 | 170.356 | 324.822 | 289.2751 | 413.9467 | 236.5065 | 210.7945 | 1011.9871 | 950.4789 | 427.7524 | 1329.2476 | 372.9725 | 541.2184 | 301.8624 | 367.178 | 168.4935 | 123.6311 | 81.4935 | 264.1926 | 469.3301 | 819.9871 | 562.178 | 384.0825 | 288.3835 | 387.2588 | 224.3285 | 272.89 | 557.3852 | 777.7945 | 725.6569 | 417.644 | 479.4402 | 867.699 | 864.9725 | 461.2055 | 902.5194 | 902.11 | 717.8478 | 718.9596 | 351.6715 | 581.89 | 605.9725 | 570.1668 | 437.4531 | 599.6311 | 337.4919 | 840.0842 | 401.6715 | 502.7541 | 303.7945 | 548.0955 | 369.699 | 521.9467 | 597.178 | 405.7007 | 537.1376 | 790.6457 | 758.2347 | 356.2347 | 184.6036 | 263.644 | 167.534 | 241.233 | 541.6715 | 681.6586 | 364.5081 | 1094.4273 | 769.9742 | 707.7799 | 323.0129 | 1296.0955 | 288.8641 | 167.3285 | 123.7395 | 288.8641 | 379.6586 | 528.2622 | 420.5065 | 353.233 | 271.8237 | 400.5486 | 318.712 | 324.0275 | 162.644 | 205.123 | 164.5065 | 261.4385 | 288.644 | 347.589 | 320.123 | 148.5486 | 694.1246 | 646.2897 | 637.3285 | 487.2605 | 392.0421 | 560.3706 | 551.3301 | 427.8349 | 827.534 | 725.6311 | 275.4935 | 876.7007
2182 | 25.9964 | 466.3055 | 657.7493 | 326.376 | 670.6266 | 557.4439 | 717.1828 | 294.3525 | 451.6136 | 400.6031 | 735.8538 | 453.5091 | 388.5013 | 907.4935 | 664.4517 | 340.6475 | 830.8877 | 708.0365 | 553.9556 | 406.6475 | 468.0235 | 333.1253 | 1187.9765 | 437.0574 | 369.5927 | 906.4674 | 1032.1044 | 612.329 | 1284.4543 | 259.7154 | 320.6136 | 168.7728 | 310.2846 | 273.3525 | 416.329 | 231.1253 | 190.0235 | 950.1384 | 1025.9974 | 427.3185 | 1345.8172 | 373.1932 | 550.4204 | 278.7624 | 377.6815 | 155.3185 | 145.658 | 90.1018 | 265.0679 | 461.0339 | 817.8407 | 553.3185 | 339.047 | 284.9661 | 394.3525 | 230.2272 | 254.4099 | 484.1044 | 788.4883 | 701.1488 | 424.8747 | 392.8094 | 822.047 | 851.4778 | 436.7963 | 1036.4387 | 828.6162 | 727.7728 | 697.671 | 319.1488 | 598.9086 | 637.658 | 506.2402 | 418.3185 | 616.8172 | 340.9321 | 786.5013 | 367.5352 | 432.4569 | 295.9321 | 528.6031 | 330.5013 | 513.8407 | 613.329 | 384.5796 | 456.9687 | 660.2428 | 775.8851 | 323.8642 | 194.7493 | 236.1828 | 165.9661 | 218.2167 | 531.7389 | 709.9191 | 353.8068 | 1012.1044 | 707.5927 | 695.0235 | 322.7389 | 1379.8721 | 242.6945 | 154.671 | 121.1593 | 242.6945 | 319.9791 | 484.376 | 454.658 | 310.8538 | 246.3185 | 365.1488 | 301.0235 | 287.5013 | 158.5796 | 184.6031 | 163.9661 | 223.6945 | 232.9452 | 303.6266 | 300.4099 | 134.1253 | 660.2167 | 645.7493 | 627.0914 | 471.2846 | 354.7624 | 519.3081 | 558.3629 | 386.3995 | 762.2298 | 686.342 | 299.4308 | 774.094
2183 | 26.0074 | 520.0494 | 697.2204 | 347.0296 | 583.8256 | 530.8848 | 711.8191 | 300.3717 | 392.2072 | 382.7138 | 687.5033 | 490.9342 | 348.0559 | 871.0855 | 707.5724 | 373.9441 | 758.1118 | 651.046 | 497.0361 | 430.1152 | 420.4079 | 317.0756 | 1095.5295 | 453.477 | 353.9046 | 837.2072 | 970.4934 | 572.5888 | 1153.3091 | 249.0756 | 280.046 | 211.4112 | 296.352 | 261.352 | 415.0855 | 251.1152 | 167.7993 | 936.1809 | 949.6249 | 435.648 | 1380.6678 | 305.0164 | 498.5789 | 299.0296 | 308.7401 | 132.523 | 144.8092 | 93.8191 | 261.2664 | 385.8256 | 773.9506 | 603.4211 | 319.171 | 227.4835 | 395.4572 | 231 | 224.4276 | 445.4276 | 817.2105 | 725.125 | 368.0263 | 358.7993 | 790.7039 | 835.7138 | 338.444 | 1097.9441 | 809.5428 | 702.4276 | 593.2532 | 304.6283 | 582.3421 | 604.8651 | 458.3223 | 400.4375 | 609.0756 | 342.2763 | 765.7138 | 352.6283 | 459.5921 | 271.0658 | 459.1118 | 311.352 | 436.3684 | 567.0361 | 344.6842 | 408.6941 | 588.7697 | 834.7829 | 283.1414 | 208.0099 | 221.8092 | 166.6382 | 219.5526 | 482.4934 | 667.398 | 384.3059 | 962.7796 | 646.0361 | 733.4112 | 304.6184 | 1409.648 | 238.0954 | 168.2007 | 130.648 | 238.0954 | 289.1612 | 476.3618 | 451.352 | 308.3717 | 233.352 | 345.7138 | 302.5526 | 274.9046 | 126.0559 | 177.3618 | 159.7237 | 209.2664 | 194.6085 | 295.4572 | 333.4967 | 112.3421 | 599.3026 | 572.1874 | 657.5823 | 449.1612 | 352.9145 | 496.2566 | 570.3717 | 386.4671 | 620.6808 | 698.6579 | 283.0658 | 709.7697
2184 | 26.0184 | 537.4823 | 739.0862 | 333.0471 | 567.8314 | 494.4079 | 724.4706 | 315.3019 | 363.204 | 342.5883 | 633.2746 | 489.349 | 342.6627 | 775.2237 | 764.545 | 395.9647 | 680.8158 | 598.7687 | 451.0824 | 437.4941 | 399.3608 | 314.8314 | 1048.2275 | 444.6981 | 325.5648 | 770.7923 | 921.7569 | 589.6039 | 1077.4432 | 260.9647 | 285.9647 | 200.2157 | 280.6981 | 227.4079 | 399.8667 | 270.9647 | 188.9411 | 886.1178 | 934.8196 | 437.1686 | 1367.0471 | 301.6392 | 484.5059 | 304.1686 | 312.9529 | 123.8431 | 162.9529 | 94 | 224.2511 | 410.5685 | 755.6863 | 578.4197 | 292.0589 | 236.6156 | 372.7216 | 231 | 208.1922 | 382.6354 | 734.8747 | 726.1804 | 335.0471 | 352.5059 | 766.8785 | 788.106 | 313.0118 | 1149.894 | 777.4079 | 721.1098 | 525.2746 | 286.3726 | 518.9727 | 636.7372 | 442.5177 | 387.0236 | 576.902 | 343.8314 | 685.6825 | 319.4079 | 400.8276 | 281.796 | 507.3528 | 299.8549 | 393.2275 | 560.1569 | 315.2157 | 405.3255 | 574.8431 | 810.2511 | 267.5177 | 165.9374 | 210.0236 | 139.5648 | 241.7843 | 488.6274 | 640.204 | 376.0353 | 913.4315 | 567.659 | 744.6627 | 292.6863 | 1368.7687 | 214.0589 | 143.7334 | 138.8196 | 214.0589 | 263.3844 | 451.8902 | 432.3726 | 268.4315 | 216.0353 | 343 | 309.8196 | 217.4668 | 114.8431 | 169.5177 | 166.4823 | 181.3961 | 188.1686 | 263.5765 | 311.2393 | 111.3255 | 547.0942 | 530.8902 | 649.204 | 440.0118 | 291.6472 | 464.7334 | 598.6039 | 406.4588 | 609.6274 | 700.5059 | 276.3373 | 719.9529
2185 | 26.0294 | 625.0535 | 786.4958 | 350.0915 | 526.6606 | 404.4775 | 663.7254 | 307.0324 | 342.2831 | 275.1 | 543.0085 | 532.7141 | 336.2507 | 712.1296 | 758.9704 | 407.4366 | 549.8874 | 487.9169 | 402.6606 | 461.6225 | 395.3127 | 309.938 | 945.1352 | 450.4366 | 325.9056 | 693.3507 | 873.5042 | 667.5549 | 928.7958 | 210.8493 | 295.7493 | 188.407 | 225.693 | 266.5577 | 368.3155 | 243.6282 | 175.4394 | 795.8521 | 901.4099 | 461.4662 | 1298.038 | 287.1268 | 463.7521 | 341.1211 | 274.8169 | 149.9972 | 149.2831 | 95.6873 | 219.531 | 415.6254 | 689.8817 | 649.4605 | 313.1535 | 201.3479 | 363.7817 | 218.3451 | 211.062 | 356.6577 | 740.6225 | 688.7225 | 331.8437 | 316.5662 | 718.2859 | 761.2831 | 310.3127 | 1137.9085 | 733.0352 | 682.9732 | 515.1563 | 307.4662 | 481.6901 | 657.6549 | 434.2507 | 357.1592 | 530.5042 | 318.6901 | 657.3451 | 289.3775 | 361.6901 | 314.5282 | 454.1944 | 270.1592 | 355.9409 | 527.2535 | 328.7169 | 346.1 | 471.9169 | 812.2803 | 237.3155 | 188.2155 | 201.2507 | 172.8084 | 195.3803 | 475.1268 | 546.7282 | 373.8437 | 926.9352 | 581.8408 | 720.5338 | 283.407 | 1341.4394 | 189.5958 | 132.938 | 135.7817 | 189.5958 | 259 | 418.3155 | 407.0648 | 232.1592 | 189.3775 | 305.0352 | 313.531 | 208.531 | 108.938 | 165.469 | 169.6873 | 188.6549 | 180.407 | 246.8761 | 305 | 114.531 | 541.531 | 518.407 | 671.3099 | 430.5634 | 323.7141 | 419.3479 | 583.7521 | 407.469 | 598.6577 | 660.3479 | 292.8732 | 714.7197
2186 | 26.0404 | 633.2964 | 799.9917 | 365.9833 | 476.2024 | 392.4238 | 631.4571 | 306.7119 | 324.4488 | 287.1107 | 523.7202 | 550.4155 | 307.6095 | 635.5238 | 789.3822 | 410.7119 | 532.2798 | 457.0167 | 360.7619 | 470.2798 | 358.194 | 303.0083 | 903.8893 | 476.8226 | 328.7119 | 649.3297 | 815.4988 | 686.9917 | 871.3297 | 210.4155 | 255.0583 | 208.3988 | 230.5512 | 218.5071 | 362.144 | 265.3905 | 154.025 | 722.7952 | 896.7119 | 466.856 | 1205.5404 | 278.8643 | 451.4405 | 314.6178 | 239.6095 | 135.3131 | 138.2964 | 96.856 | 208.8726 | 376.4821 | 712.2381 | 602.3714 | 274.3464 | 239.3655 | 353.5845 | 227.1274 | 207.7202 | 348.8643 | 713.3297 | 685.4238 | 311.4571 | 310.856 | 624.4047 | 668.125 | 280.8976 | 1110.8893 | 648.9642 | 752.8917 | 472.2024 | 294.025 | 513.806 | 648.8726 | 406.4655 | 354.5679 | 463.0833 | 349.95 | 632.7452 | 326.9417 | 375.831 | 306.3048 | 442.856 | 239.3214 | 343.1524 | 568.0774 | 290.0583 | 322.1607 | 483.8143 | 825.9833 | 209.7452 | 168.3214 | 192.2964 | 170.5845 | 216.8143 | 460.0167 | 535.9917 | 402.2464 | 835.1333 | 531.3631 | 783.6203 | 294.8393 | 1297.7702 | 202.2631 | 140.5595 | 126.4405 | 202.2631 | 207.6428 | 385.6095 | 438.0941 | 205.6012 | 168.7369 | 288.5845 | 350.806 | 185.8893 | 122.5512 | 152.1607 | 170.856 | 165.3214 | 177.2881 | 195.3547 | 310.1357 | 126.9833 | 532.5845 | 475.0583 | 662.3048 | 415.3048 | 313.169 | 375.194 | 571.4405 | 377.0417 | 601.9917 | 657.2798 | 313.1191 | 726.6952
2187 | 26.0514 | 659.7838 | 862.6453 | 394.0473 | 440.348 | 371.2939 | 598.4797 | 338.2568 | 302.8986 | 298.8142 | 469.1689 | 558.0777 | 284.7669 | 568.4324 | 752.4561 | 441.3885 | 561.652 | 423.7432 | 350.6588 | 486.6284 | 365.0237 | 280.2939 | 862.6655 | 513.9933 | 310.7669 | 584.9594 | 710.6249 | 687.1318 | 820.8513 | 212.8682 | 226.2939 | 218.0777 | 236.473 | 215.0777 | 342.0304 | 259.5811 | 144.054 | 679.1385 | 875.2939 | 500.8615 | 1129.4864 | 273.6588 | 441.3176 | 312.473 | 267.9933 | 117.2399 | 146.5507 | 77.8986 | 187.8986 | 350.8986 | 622.4932 | 637.1487 | 261.7905 | 236.5811 | 318.1385 | 229 | 186.1621 | 346.2635 | 718.4189 | 638.2466 | 312.3412 | 245.8817 | 582.2162 | 684.2568 | 295.9696 | 1113.946 | 619.5034 | 677.4391 | 475.4189 | 265.8209 | 434.9121 | 636.5811 | 394.1858 | 360.2095 | 443.4493 | 381.1791 | 610.7669 | 299.2702 | 379 | 319.6284 | 418.6892 | 233.2635 | 312.4797 | 490.0439 | 325.5439 | 322.6047 | 409.9898 | 794.1385 | 224.2331 | 170.0777 | 176.2399 | 163.7905 | 219.3953 | 471.0237 | 473.6182 | 403.527 | 786.875 | 472.5101 | 825.3885 | 311.7601 | 1184.206 | 182.4257 | 146.3412 | 140.6284 | 182.4257 | 218.1014 | 349.7432 | 409.2702 | 196.7905 | 197.2568 | 248.7973 | 302.3006 | 183.7365 | 106.7669 | 137.8446 | 177.946 | 153.1858 | 161.3716 | 187.8682 | 269.3243 | 117.7128 | 463.277 | 466.2635 | 666.0777 | 394.7669 | 301.3176 | 374.2095 | 535.2702 | 407.598 | 594.3176 | 618.929 | 309.9223 | 667.3547
2188 | 26.0624 | 708.0267 | 865.8363 | 402.4027 | 415.7477 | 329.2565 | 548.2122 | 342.1195 | 299.1195 | 300 | 415.2122 | 603.0267 | 270.553 | 512.4511 | 772.416 | 423.9866 | 521.9733 | 392.584 | 323.584 | 454.6591 | 337.9423 | 272.5973 | 793.6015 | 480.2565 | 308.8805 | 490.5881 | 716.2523 | 683.4779 | 713.619 | 216.5221 | 221.2389 | 180.2565 | 269.5798 | 231.8496 | 319.6282 | 290.5798 | 136.8363 | 617.6458 | 831.4954 | 541.2214 | 1075.9733 | 243.9423 | 438.2389 | 302.4336 | 258.0309 | 118.5221 | 110.137 | 89.9691 | 186.7611 | 325.1061 | 618.5664 | 661.6107 | 242.5088 | 219.1504 | 292.7477 | 235.1637 | 192.6859 | 308.137 | 653.0793 | 640.6859 | 283.9423 | 269.4603 | 554.2256 | 656.4202 | 264.6591 | 1056.0042 | 559.7652 | 691.2966 | 457.6282 | 234.7034 | 414.0752 | 647.3275 | 396.5221 | 318.7343 | 419.9866 | 370.0309 | 541.9599 | 286.0752 | 327.929 | 333.447 | 414.1195 | 228.5973 | 275.4202 | 550.0844 | 288.8538 | 298.345 | 361.8981 | 780.1947 | 254.2966 | 172.7611 | 187.208 | 163.8805 | 192.584 | 437.7786 | 479.8496 | 396.8363 | 796.9691 | 450.9114 | 766.6015 | 331.6107 | 1110.7652 | 163.1504 | 119.7034 | 147.4027 | 163.1504 | 205.1504 | 307.137 | 432.1771 | 219.7744 | 165.0175 | 239.4779 | 324.8187 | 182.2389 | 93.4336 | 131.5973 | 186.0443 | 144.0752 | 154.5973 | 189.7611 | 296.4603 | 127.447 | 519.0401 | 421.9733 | 638.8229 | 411.3718 | 263.8981 | 349.4645 | 539.6859 | 390.9866 | 556.0175 | 576.8981 | 308.1195 | 655.3584
2189 | 26.0734 | 713.1072 | 881.9637 | 401.2143 | 393.3578 | 307.9291 | 548.2498 | 325.9291 | 323.9991 | 277.6794 | 375.0726 | 605.4287 | 281.4996 | 508.6785 | 776.8928 | 451.3561 | 471.3587 | 400.6067 | 295.0009 | 454.4641 | 349.178 | 230.0372 | 749.287 | 501.7848 | 304.5359 | 483.4641 | 656.5022 | 724.07 | 662.5013 | 192.0009 | 229.9283 | 208.9274 | 224.8946 | 234.8928 | 299.1435 | 265.5367 | 135.1072 | 577.8583 | 745.6457 | 514.7511 | 986.9672 | 234.643 | 388.8946 | 295.643 | 256.8928 | 119.8928 | 134.4633 | 89.3215 | 180.7502 | 350.5704 | 536.9672 | 642.5722 | 276.6059 | 212.5359 | 270.3578 | 272.6059 | 199.357 | 319.9637 | 625.2507 | 607.1798 | 278.2143 | 247.2152 | 459.9317 | 643.9646 | 243.9291 | 1076.5704 | 504.6802 | 656.6085 | 453.2143 | 264.9274 | 371.0372 | 651.6785 | 364.8583 | 345.1417 | 384.8583 | 382.2852 | 594.605 | 292.1426 | 308.5004 | 337.6785 | 427.3924 | 238.7139 | 265.643 | 451.9681 | 291.0354 | 291.4287 | 339.1435 | 727.2161 | 228.5367 | 179.2498 | 177.3933 | 165.7857 | 177.3933 | 470.4987 | 445.3941 | 453.1409 | 778.465 | 449.8928 | 814.248 | 293.8228 | 1048.5376 | 194.0346 | 107.9646 | 139.9646 | 194.0346 | 176.2152 | 304.6785 | 397.6085 | 180.1444 | 174.2852 | 228.2861 | 276.3233 | 200.7493 | 104.4996 | 117.6076 | 165.5722 | 143 | 144.1789 | 170.3578 | 304.5713 | 116.5004 | 441.3959 | 408.8574 | 621.6076 | 422.0354 | 269.7139 | 340.643 | 584.7485 | 362.108 | 568.8565 | 514.8591 | 324.0709 | 671.0709
2190 | 26.0844 | 808.9424 | 863.1823 | 414.5768 | 384.6642 | 310.5256 | 506.4595 | 357.4893 | 312.5181 | 248.7516 | 322.1237 | 582.3721 | 263.0874 | 472.7953 | 736.2697 | 432.3721 | 446.0874 | 348.5981 | 277.5181 | 492.1098 | 336.5181 | 290.1684 | 669.8753 | 492.3284 | 277.7516 | 480.3795 | 563.0139 | 723.5693 | 635.3721 | 179.0437 | 217.4232 | 190.3721 | 247.9637 | 225.0437 | 272.5618 | 288.2484 | 169.3944 | 593.9126 | 694.3646 | 560.7814 | 1008.6791 | 235.8102 | 384.8102 | 350.2121 | 238.8977 | 92.8465 | 147.0512 | 84.4744 | 182.7154 | 340.4232 | 479.0288 | 608.3209 | 264.7079 | 198.4232 | 248.9925 | 280.6205 | 197.2846 | 294.8465 | 578.6493 | 632.8689 | 301.533 | 229.5181 | 433.613 | 626.7079 | 225.7079 | 973.1962 | 421.1599 | 657.4307 | 431.2772 | 289.8177 | 342.467 | 709.9275 | 374.5768 | 330.8977 | 374.6642 | 391.2409 | 535.9265 | 289.3795 | 288.8977 | 348.8614 | 412.7079 | 244.5256 | 246.8977 | 440.8102 | 282.9488 | 278.3284 | 324.3284 | 744.533 | 196.9414 | 175.4744 | 172.3795 | 193.1535 | 201.3433 | 428.839 | 452.7665 | 440.0874 | 806.774 | 402.0288 | 763.9776 | 327.92 | 952.3934 | 159.9851 | 117.8614 | 161.6279 | 159.9851 | 192.9126 | 294.1386 | 415.6279 | 189.4819 | 158.8028 | 196.226 | 295.3433 | 172.226 | 113.2409 | 104.2335 | 143.0874 | 141.1898 | 170.1535 | 146.2772 | 286.8977 | 112.2846 | 471.7303 | 403.4744 | 639.0075 | 419.3795 | 269.1898 | 308.3209 | 571.8977 | 384.3433 | 523.9339 | 518.8614 | 323.2846 | 651.2772
2191 | 26.0954 | 776.7992 | 851.8259 | 447.1919 | 344.5515 | 280.7255 | 441.4511 | 345.4041 | 245.864 | 265.2656 | 334.4308 | 615.7789 | 280.2656 | 449.7344 | 754.0178 | 480.4574 | 438.4956 | 326.4867 | 248.4778 | 477.6518 | 310.23 | 280.4867 | 650.9911 | 518.5222 | 275 | 462.5692 | 552.1652 | 693.6429 | 625.6607 | 186.2567 | 223.3393 | 237.54 | 214.3037 | 219.413 | 276.4219 | 286.413 | 123.46 | 598.7522 | 633.1207 | 523.7992 | 852.371 | 233.2478 | 373.0737 | 343.1563 | 237.9174 | 98.2567 | 137.9085 | 91.3393 | 191.2567 | 339 | 446.4778 | 649.953 | 245.5692 | 228.1919 | 227.7344 | 270.9085 | 170.3952 | 328.6963 | 536.3863 | 555.2681 | 293.9085 | 228 | 420.0737 | 541.5159 | 227.6696 | 831.7281 | 419.4219 | 675.4308 | 448.2656 | 249.7992 | 342.7522 | 680.2211 | 364.9911 | 348.2656 | 352.8996 | 387.413 | 540.9263 | 291.7522 | 272.3215 | 361.0089 | 377.0559 | 242.2478 | 228.4867 | 437.3304 | 246.2211 | 273.3304 | 315.6607 | 695.6251 | 181.1563 | 164.9085 | 160.0737 | 177.6518 | 207.6696 | 443.2656 | 479.6874 | 488.4574 | 768.7166 | 438.2834 | 709.3774 | 313.6518 | 917.3126 | 158.7522 | 109.8259 | 140.1474 | 158.7522 | 193.1652 | 245.2948 | 358.3685 | 179.9911 | 162.5044 | 171.8996 | 275.0648 | 179.0915 | 102.9911 | 110.3393 | 160.2656 | 133.6607 | 143.6429 | 150.4219 | 242.7992 | 108.3304 | 429.2122 | 355.2948 | 650.1741 | 414.413 | 243.3126 | 309.587 | 567.2478 | 383.3304 | 533.6785 | 457.6162 | 334.0089 | 682.0267
2192 | 26.1064 | 778.5782 | 845.4218 | 436.9842 | 355.8752 | 278 | 442.5079 | 373.7504 | 268.8207 | 259.5624 | 332.2812 | 652.4692 | 278.2812 | 387.5696 | 650.0143 | 522.188 | 424.9842 | 342.6643 | 235.7733 | 452.7575 | 306.1406 | 280.8594 | 640.703 | 498.6872 | 261.9842 | 465.6485 | 486.9211 | 732.8365 | 614.7733 | 188.8594 | 229.5782 | 189.0071 | 229.594 | 221.7891 | 264.9139 | 298.0861 | 142.2425 | 549.7259 | 648.4534 | 572.9929 | 794.3041 | 275.7662 | 336.6714 | 316.8981 | 211.9684 | 111.0861 | 135.1406 | 97.5782 | 191.0703 | 360.3831 | 409.6011 | 626.109 | 242.1406 | 210.5466 | 247.3831 | 227.2338 | 184.7346 | 312.4763 | 513.4763 | 546.1406 | 289.2812 | 212.1951 | 325.1004 | 546.0861 | 197.3199 | 759.5696 | 379.0932 | 627.7259 | 379.3429 | 296.2038 | 305.812 | 677.9297 | 345.406 | 321.1793 | 338.9139 | 426.0474 | 565.2425 | 274.3357 | 283.0861 | 348.9842 | 408.3989 | 217.8278 | 221.4218 | 441.6485 | 258.8049 | 284.1564 | 269.4447 | 671.4763 | 201.3831 | 184.4534 | 162.7188 | 176.9297 | 217.297 | 463.594 | 435.515 | 498.5782 | 848.6729 | 389.0071 | 681.7575 | 310.1406 | 876.8823 | 174.8049 | 80.1793 | 136.1406 | 174.8049 | 178.1248 | 213.109 | 334.406 | 178.0703 | 159.2812 | 188.594 | 275.7891 | 174.4218 | 86.1951 | 104.4921 | 141.5466 | 140.4376 | 157.7346 | 147.2812 | 233.4218 | 80.109 | 463.1177 | 375.1722 | 634.2654 | 425.1564 | 283.7662 | 319.297 | 516.7962 | 383 | 545.2267 | 425.9684 | 328.4921 | 690.5782
2193 | 26.1174 | 811.9697 | 795.0745 | 451.0718 | 339.1022 | 231.8425 | 394.0165 | 376 | 264.4061 | 249.5801 | 249.1049 | 606.0165 | 229.9585 | 394.3039 | 684.3896 | 527.826 | 430.5939 | 337.4061 | 223.6961 | 467.9558 | 292.8121 | 285.71 | 622.1022 | 521.4917 | 292.0857 | 441.5083 | 498.0138 | 671.9447 | 587.6243 | 177.6961 | 180.0745 | 196.3039 | 230.058 | 243.6658 | 248.9282 | 302.768 | 133.6381 | 544.116 | 576.5248 | 632.5774 | 757.0884 | 266.7541 | 273.7126 | 331.0138 | 232.6078 | 83.7403 | 125.5801 | 90.4641 | 167.4502 | 328.0884 | 398.5221 | 682.4034 | 210.9143 | 221.2459 | 225.4502 | 256.0277 | 187.884 | 292.1602 | 435.6988 | 538.4641 | 285.232 | 211.942 | 314.232 | 496.1325 | 227.9697 | 722.0303 | 365.6381 | 668.2735 | 413.5636 | 263.2624 | 315.2459 | 654.4502 | 374.1437 | 332.1879 | 335.174 | 424.29 | 554.7541 | 301.2597 | 286.826 | 343.29 | 417.5939 | 246.1437 | 208.7541 | 379.8286 | 266.5939 | 318.9116 | 249.9862 | 636.0884 | 199.232 | 174.6961 | 167.71 | 171.3481 | 206.6961 | 420.7265 | 435.768 | 487.6961 | 800.3646 | 449.0553 | 714.8536 | 307.174 | 852.3342 | 176 | 115.6796 | 125.6381 | 176 | 195.8398 | 217.5939 | 327.3481 | 178.942 | 139.2182 | 191.884 | 263.7541 | 153.2762 | 104.7818 | 98.3481 | 133.4061 | 121.2182 | 141.1022 | 144.174 | 223.5801 | 118.5056 | 500.8536 | 379.826 | 597.2044 | 421.29 | 234.2485 | 315.29 | 503.5801 | 362.2762 | 517.7403 | 384.4364 | 383.5774 | 616.5828
2194 | 26.1284 | 748.1546 | 709.9318 | 461.5428 | 332.2743 | 267.1713 | 372.8687 | 418.9427 | 245.8687 | 274.7656 | 266.9028 | 544.7888 | 259.4456 | 408.3599 | 585.8462 | 444.0232 | 376.606 | 359.9028 | 244.9485 | 513.8512 | 329.217 | 274.5486 | 575.1945 | 543.0399 | 312.1313 | 424.7315 | 453.1945 | 692.8113 | 577.4115 | 209.4456 | 197.0399 | 215.1313 | 244.3142 | 235.4572 | 232.7315 | 282.0058 | 134.9086 | 546.8628 | 534.783 | 572.0632 | 703.4688 | 230.6916 | 280.4971 | 323.4115 | 215.8687 | 106.8113 | 108.7772 | 108.1313 | 171.7257 | 326.9543 | 343.606 | 636.3774 | 246.217 | 218.1829 | 205.8687 | 242.7315 | 182.2743 | 301.4971 | 406.1887 | 517.9601 | 266.8687 | 256.8512 | 297.7772 | 494.9086 | 225.2286 | 622.6633 | 324.9202 | 630.9202 | 414.0914 | 240.9601 | 271.1488 | 667.3142 | 383.6343 | 305.3258 | 335.9543 | 463.1255 | 561.6343 | 261.9659 | 283.1829 | 358.2685 | 372.1945 | 235.5943 | 200.3657 | 381.7257 | 260.32 | 279.9659 | 265.2228 | 653.0856 | 209.4971 | 170.1829 | 163.2286 | 191.0399 | 208.8628 | 458.0798 | 373.9717 | 550.9368 | 779.8229 | 429.143 | 674.0573 | 343.2627 | 734.5777 | 166.4572 | 113.2286 | 147.9028 | 166.4572 | 187.4572 | 178.8745 | 331.7714 | 190.4514 | 156.1313 | 158.6002 | 277.3142 | 168.2228 | 104.0914 | 84.6401 | 133 | 127.6343 | 147.6343 | 147.8171 | 270.7141 | 122.9086 | 503 | 387.6343 | 657.9826 | 419.0914 | 272.9884 | 321.68 | 507.7714 | 348.5943 | 557.0341 | 404.9028 | 336.4231 | 658.7598
2195 | 26.1394 | 746.9331 | 692.468 | 486.1643 | 306.8691 | 286.3983 | 392.298 | 396.8357 | 243.0669 | 237.3371 | 224.5042 | 586.4623 | 229.1031 | 435.0975 | 547.17 | 519.2589 | 408.7966 | 351.3343 | 233.4346 | 456.0725 | 317.468 | 289.4652 | 596.1977 | 537.234 | 293.6686 | 431.7326 | 431.6686 | 722.0306 | 569.2674 | 182.0028 | 183.5014 | 242.0975 | 216.0028 | 216.6351 | 249.3983 | 288.7326 | 136.9331 | 490.9388 | 490.4708 | 602.83 | 661.3705 | 238.6657 | 285.8329 | 308.5014 | 216.9331 | 104.1337 | 108 | 89.6686 | 163.3008 | 329.8997 | 386.4289 | 671.6963 | 228.6686 | 237.3314 | 203.0669 | 290.3286 | 186.8329 | 334.8635 | 367.3037 | 495.7354 | 283.3983 | 199.0725 | 277.6686 | 474.702 | 205.6686 | 556.1394 | 327.8329 | 639.6323 | 412.0669 | 214.8691 | 249.6686 | 644.8023 | 388.8329 | 325.2646 | 303.1365 | 481.4317 | 523.3371 | 283.1977 | 290.7326 | 389.9303 | 376.766 | 222.4346 | 184.5348 | 346.2368 | 240.6686 | 280.8997 | 281.4652 | 637.5683 | 214.8329 | 155.5014 | 156.234 | 206.4986 | 180.0028 | 430.0363 | 423.1949 | 525.0028 | 778.0334 | 407.702 | 653.6351 | 332.4346 | 720.3008 | 162.1337 | 94.6351 | 124.8357 | 162.1337 | 153.17 | 189.5654 | 296.2368 | 179.4011 | 173.4317 | 153.1337 | 241.2703 | 148.702 | 104 | 105.2646 | 117.5348 | 146.3649 | 145.1003 | 135.4346 | 260.4346 | 98.8357 | 484.6351 | 364.8023 | 643.6017 | 399.6686 | 258.5683 | 277.5377 | 468.3705 | 398.2618 | 578.3314 | 350.9054 | 327.234 | 674.532
2196 | 26.1504 | 671.6275 | 625.4373 | 508.535 | 329.4927 | 248.8243 | 286.3039 | 408.7252 | 219.5073 | 233.0634 | 214.1902 | 588 | 237.7886 | 391.9511 | 477.4796 | 554.3025 | 410 | 275.6275 | 208.5284 | 431.4861 | 294.4861 | 287.0634 | 546.0991 | 592.7952 | 265.5918 | 439.8309 | 434.9155 | 752.3659 | 560.1902 | 187.852 | 206.4927 | 224.4016 | 207.1691 | 241.4504 | 267.6195 | 289 | 137 | 482.148 | 520.3236 | 626.5139 | 658.0423 | 230.1902 | 290.8943 | 350.0911 | 227.7675 | 93.2325 | 111.9155 | 84.1057 | 154.1902 | 355.4504 | 364.5073 | 661.2536 | 239.7464 | 229.1902 | 202.0211 | 273.4016 | 172.317 | 310.5496 | 361.1057 | 525.3448 | 265.4016 | 218.535 | 279.9366 | 465.1902 | 235.3448 | 502.1202 | 317.2325 | 612.5918 | 412.9789 | 279.5838 | 245.0845 | 643.0211 | 325.3739 | 326.9789 | 310.8098 | 451.6552 | 498.5073 | 277.148 | 265.5496 | 412.535 | 440.6261 | 202.4227 | 204.5561 | 357.7252 | 260.5561 | 248.6975 | 265.3593 | 615.465 | 207.1691 | 159.8943 | 175.5773 | 213.852 | 199.5561 | 463.2602 | 415.2114 | 552.387 | 779.9577 | 378.613 | 605.0357 | 349.6195 | 711.1902 | 159.0634 | 96.9366 | 137.7041 | 159.0634 | 184.3025 | 178.2536 | 299.8943 | 175.0845 | 127.0146 | 178.4504 | 241.9577 | 161.7041 | 101.0634 | 106 | 157.1334 | 125.465 | 143.0423 | 142.8309 | 236.5073 | 102.8943 | 478.1268 | 367.9155 | 622.4439 | 388.2325 | 273.6618 | 308.3025 | 507.1334 | 331.4796 | 579.9789 | 365.6407 | 349.5139 | 674.0211
2197 | 26.1614 | 629.3627 | 600.2123 | 531.7966 | 299.2742 | 281.6993 | 284.9912 | 428.8231 | 219.9912 | 233 | 197.1504 | 555.2919 | 222.1415 | 413.7966 | 464.1061 | 519.3184 | 367.3804 | 344.372 | 222.8673 | 402.2565 | 298.9558 | 302.8585 | 543.0177 | 576.1592 | 292.7523 | 448.9204 | 408.2388 | 724.2565 | 558.0177 | 181.0619 | 179.2477 | 235.8939 | 209.9735 | 243.9823 | 212.4954 | 317.7435 | 137 | 482 | 469.46 | 661.6904 | 607.4511 | 240.9027 | 285.0531 | 317.3007 | 228.9912 | 98.9469 | 122.9027 | 82.0177 | 161.9292 | 343.115 | 330.3007 | 654.0619 | 255.8585 | 190.345 | 204.9735 | 288.8585 | 195.7877 | 284.23 | 344.1504 | 504.1946 | 277.885 | 212.0619 | 236.3892 | 464.0088 | 200.3184 | 457.3892 | 320.9646 | 557.4865 | 408.0442 | 278.0265 | 235.0885 | 669.7612 | 361.6639 | 341.8673 | 291.1769 | 482.7169 | 561.4339 | 276.0088 | 269.9558 | 430.8408 | 423.1681 | 219.8408 | 199.0531 | 349.0796 | 261 | 276.7435 | 266.9823 | 605.0885 | 171.3184 | 171.8939 | 178.9735 | 207.0619 | 200 | 452.1061 | 403.1061 | 553.9912 | 810.7258 | 406.7435 | 570.3007 | 294.4954 | 626.7519 | 155.0354 | 77.1769 | 130.0708 | 155.0354 | 185 | 195.8408 | 291.0796 | 191.8496 | 143.8408 | 165.1238 | 242.9912 | 168.9381 | 85.1415 | 93.115 | 154.0354 | 142.8408 | 136.0619 | 149.9381 | 243.9292 | 98.0442 | 461.1504 | 312.4954 | 585.3273 | 412.7789 | 221.4688 | 277.2831 | 461.4157 | 368.655 | 603.7877 | 372.9381 | 365.8585 | 717.6108
2198 | 26.1724 | 613.6348 | 536.9414 | 481.0689 | 325.9035 | 259.9552 | 308.9104 | 424.0724 | 224.0379 | 227.0241 | 180.093 | 521.9587 | 240.062 | 367.9897 | 429.8036 | 529.9966 | 402.9552 | 306.0103 | 209.0448 | 402.0448 | 276.9931 | 279.0999 | 530.0482 | 618.0551 | 282.0103 | 420.9724 | 369.1137 | 688.0345 | 537.845 | 198.9759 | 199.9483 | 221.093 | 215.9966 | 230.9897 | 239.9552 | 277.1998 | 121.9931 | 537.9139 | 413.9208 | 645.1103 | 606.9587 | 202.1378 | 270.9242 | 308.0172 | 226.9276 | 103.9966 | 106.9483 | 92.0689 | 172.9897 | 379.8587 | 320.9449 | 683.9139 | 227.0689 | 234.8966 | 209.0448 | 295.0172 | 184.0138 | 323.8346 | 319.9242 | 509.9931 | 268.9655 | 230.9724 | 202.9966 | 408.9035 | 223.9552 | 432.9242 | 305.9862 | 595.8346 | 444.8208 | 261.9897 | 255.9828 | 644.9414 | 339.8691 | 323.969 | 311.1378 | 486.031 | 512.0827 | 263.9828 | 248.1103 | 425.9483 | 410.0276 | 230.9621 | 240.8553 | 354.9311 | 225.062 | 297.0034 | 272.8829 | 568.9173 | 200.9724 | 165.0448 | 178.9793 | 192.0482 | 197.9276 | 453.1826 | 412.9793 | 567.124 | 768.9483 | 389.2377 | 551.7623 | 333.031 | 677.8897 | 155.0103 | 87.969 | 123.0965 | 155.0103 | 142.0861 | 174.9931 | 292.9966 | 180.0655 | 148.9828 | 163.9414 | 221.0138 | 145.0448 | 103.0034 | 99.9414 | 126.062 | 111.0551 | 149.9552 | 143.969 | 202.0034 | 97.9897 | 455.2102 | 361.0172 | 592.9518 | 367.1964 | 276.0069 | 311.9242 | 497.062 | 337.093 | 613.938 | 353.0345 | 360.0586 | 660.1619
2199 | 26.1834 | 507.701 | 519.7797 | 502.0543 | 298.5508 | 247.0472 | 282.9056 | 445.1102 | 235.3619 | 233.8898 | 206.4807 | 508.9299 | 257.6223 | 364.4335 | 372.0401 | 528.402 | 389.9056 | 308.5122 | 221.8112 | 414.6853 | 275.4721 | 307.4177 | 543.4177 | 635.0229 | 284.7797 | 413.0944 | 401.764 | 697.9213 | 493.7396 | 191.8426 | 184.764 | 248.3934 | 214.6223 | 228.1731 | 226.9056 | 334.3548 | 119.8426 | 512.8269 | 391.0787 | 676.7482 | 595 | 242.299 | 248.8269 | 312.7325 | 205.9371 | 103.3462 | 92.0472 | 112.236 | 170.1416 | 339.0944 | 305.2518 | 659.1888 | 246.8898 | 205.2518 | 221.1502 | 299.5122 | 188.0315 | 276.299 | 297.4807 | 507.4492 | 259.2518 | 222.8898 | 201.9213 | 381.4564 | 210.7325 | 410.5594 | 301.5436 | 548.1102 | 393.1888 | 258.7167 | 250.6381 | 627.0715 | 302.2675 | 315.4091 | 350.2446 | 494.5122 | 536.6452 | 259.3462 | 279.6538 | 411.2046 | 417.4807 | 219.8584 | 199.2203 | 334.3863 | 242.9371 | 298.1888 | 239.1259 | 544.6381 | 193.0629 | 178.299 | 172.8112 | 206.1416 | 177.4406 | 506.1259 | 406.8269 | 603.1259 | 754.7081 | 457.6066 | 483.7711 | 341.5122 | 645.7797 | 158.7711 | 79.2518 | 150.5436 | 158.7711 | 166.8741 | 172.8426 | 290.8827 | 198.7167 | 144.236 | 147.1731 | 225.299 | 157.8269 | 104.0629 | 83.0629 | 143.5751 | 126.6066 | 136.8741 | 134.9843 | 202.8426 | 94.7167 | 515.1188 | 365.9528 | 578.6381 | 423.2132 | 277.9843 | 290.0944 | 515 | 364.4091 | 596.2203 | 362.1502 | 377.4091 | 707.3777
2200 | 26.1944 | 487.5986 | 504.1782 | 567.8038 | 332.5796 | 249.5516 | 275.6266 | 451.8599 | 258.1401 | 227.2522 | 174.4204 | 442.8128 | 234.2242 | 330.1772 | 311.3834 | 491.8689 | 383.6637 | 277.4955 | 209.7478 | 394.4955 | 304.5516 | 270.7478 | 507.2522 | 698.5796 | 270.3273 | 418.7758 | 386.5796 | 693.5325 | 540.2522 | 181.9439 | 170.3924 | 272.3273 | 191.9529 | 239.5886 | 220.7197 | 294.8689 | 109.972 | 501.972 | 396.8969 | 661.5325 | 594.6356 | 260.019 | 237.1311 | 297.3453 | 201.8879 | 124.1872 | 95.1682 | 126.3834 | 178.8318 | 345.1962 | 321.1121 | 670.0471 | 240.981 | 220.1311 | 169.0931 | 270.2332 | 190.4765 | 295.028 | 265.5886 | 472.2993 | 274.8038 | 215.2433 | 197.6446 | 408.8228 | 194.028 | 382.4394 | 273 | 554.9159 | 405.1121 | 241.1401 | 227.5235 | 569.0561 | 320.037 | 341.2522 | 302.972 | 464.1962 | 576.4114 | 280.2993 | 258.7567 | 422.7948 | 385.5325 | 211.5045 | 213.1401 | 296 | 239.0841 | 308.8789 | 246.5516 | 521.7197 | 198.2332 | 196.019 | 160.6917 | 215.5325 | 204.7758 | 512.2062 | 395.4114 | 612.2612 | 799.4765 | 432.9159 | 530.4865 | 312.4574 | 632.6166 | 205.2062 | 94.3834 | 122.4484 | 205.2062 | 159.5325 | 163.4765 | 221.5606 | 180.7197 | 158.4955 | 158.0841 | 244.028 | 146.7758 | 107.972 | 87.2242 | 116.9159 | 102.9249 | 128.8318 | 134.2242 | 193.4765 | 77.5045 | 461.6817 | 362.7197 | 556.1682 | 375.065 | 276.4394 | 296.1401 | 514.2713 | 389.3554 | 610.1121 | 309.6446 | 403.4765 | 731.4484
2201 | 26.2054 | 437.3066 | 440.3146 | 562.2499 | 318.6048 | 233.9194 | 229.5724 | 447.887 | 261.992 | 236.4032 | 187.4276 | 467.8955 | 241.7178 | 371.2419 | 291.2096 | 520.7098 | 371.8387 | 262.4594 | 201.0403 | 375.3066 | 288.4759 | 262.0806 | 513.9034 | 685.0886 | 246.9597 | 410.9597 | 372.9273 | 712.3225 | 549.2419 | 180.4435 | 182.6292 | 250.0886 | 224.0727 | 260.2419 | 211.6451 | 325.2822 | 110.0886 | 500.7984 | 428.4435 | 681.7337 | 581.3146 | 226.8064 | 207.3225 | 342.226 | 198.9273 | 96.2822 | 101.2016 | 105.4435 | 173.2419 | 352.3225 | 324.6371 | 638.1289 | 273.7501 | 191.5724 | 206.1936 | 311.5485 | 206.7178 | 295.7581 | 285.1533 | 447.3952 | 269.0483 | 190.6934 | 218.9517 | 369.4112 | 195.9676 | 361.0647 | 271.7501 | 551.879 | 408.113 | 244.992 | 212.2096 | 571.3225 | 355.9194 | 348.992 | 303.8547 | 471.1613 | 555.3146 | 258.3788 | 283.1857 | 383.0966 | 403.6775 | 230.1289 | 218.4435 | 295.2339 | 240.8711 | 269.6775 | 232.0483 | 512.121 | 238.1777 | 162.0806 | 150.5645 | 233.6775 | 196.0324 | 451.2096 | 376.008 | 656.5645 | 817.5321 | 429.7984 | 477.7984 | 363 | 653.3146 | 143.9273 | 73.4838 | 137.7984 | 143.9273 | 177.0324 | 179.0324 | 241.121 | 170.3952 | 142.6127 | 160.2339 | 243.9114 | 140.0886 | 106.7178 | 95.6854 | 113.8387 | 134.8387 | 123.6451 | 141.7178 | 209.3549 | 95.2419 | 517.1374 | 352.2339 | 562.121 | 412.5645 | 257.8064 | 302.2096 | 487.8308 | 367.4032 | 613.5565 | 333.0886 | 419.7984 | 745.5485
2202 | 26.2164 | 398.8939 | 423.3161 | 591.1056 | 332.737 | 232.6313 | 267.1052 | 468.3161 | 238.1578 | 245.1056 | 151.2626 | 392.6321 | 234.9474 | 376.3161 | 319.1582 | 490.7891 | 367.737 | 320.053 | 201.8422 | 336.2626 | 277.5783 | 267.104 | 467.0514 | 711 | 247.4205 | 411.6835 | 393.1587 | 719.1056 | 550.8964 | 189.3691 | 151.3152 | 274.4743 | 202.5261 | 263.2643 | 227.2104 | 331.053 | 135.4213 | 496.7365 | 438.8948 | 722.0004 | 563.5269 | 245.8948 | 216.6835 | 300.6839 | 219.7374 | 102.7896 | 106.0526 | 115.1056 | 179.1578 | 358.6848 | 316.4735 | 663.9478 | 242.6848 | 227.9478 | 187.1052 | 279.2096 | 199.5265 | 290.263 | 265.6313 | 432.2635 | 294.7891 | 229.1587 | 195.21 | 399.2121 | 219.4209 | 315.0526 | 244.6305 | 548.4739 | 388.3152 | 222.4731 | 240.8422 | 579.5261 | 353.6317 | 324.6843 | 344.7382 | 473.3165 | 536.4221 | 312.9482 | 241.3679 | 432.6317 | 397.1574 | 254.1065 | 228.737 | 274.9482 | 216.3148 | 260.5269 | 258.3679 | 513.79 | 173.0518 | 164.4209 | 164.5787 | 226.947 | 175.5779 | 480.263 | 401.3152 | 669.6317 | 850.2647 | 424.1582 | 473.1578 | 362.8422 | 636.5791 | 166.1578 | 87.2622 | 132.7896 | 166.1578 | 153.0004 | 157.21 | 244.1578 | 156.1052 | 179.79 | 141.2635 | 216.0534 | 166.9996 | 99.5265 | 111.6843 | 109.6317 | 131.263 | 138.8422 | 135 | 194 | 101.3157 | 449.1052 | 336.0518 | 561.4225 | 423.58 | 275.4217 | 328.7904 | 461.8939 | 379.6305 | 601.4743 | 357.4217 | 413.8952 | 710.6317
2203 | 26.2274 | 428.7815 | 408.8583 | 576.2331 | 327.351 | 244.3894 | 268.7404 | 452.1709 | 242.5576 | 229.5841 | 173.4808 | 366.5722 | 236.9854 | 362.6371 | 301.3391 | 504.6106 | 360.9881 | 303.2596 | 199.8437 | 358.351 | 305.2861 | 320.6636 | 524.1417 | 714.3099 | 270.7934 | 441.6106 | 360.8437 | 696.2503 | 480.245 | 161.4278 | 176.6874 | 249.9854 | 210.8318 | 215.6874 | 231 | 316.9205 | 125.1298 | 508.1828 | 435.1828 | 704.9086 | 538.3629 | 243.5457 | 246.7404 | 313.8437 | 197.0649 | 99.3245 | 106.351 | 99.2596 | 181.7404 | 336.6874 | 322.2093 | 643.4424 | 220.5576 | 209.3894 | 188.2861 | 321.2715 | 189.7669 | 294.4808 | 273.2358 | 417.4424 | 304.5867 | 198.1417 | 217.5841 | 331.5192 | 225.8318 | 315.6755 | 290.1179 | 538.1563 | 411.8702 | 248.0914 | 239.6225 | 588.6106 | 346.8702 | 319.8437 | 289.245 | 446.7642 | 492.2715 | 276.9735 | 263.9232 | 427.7523 | 415.1444 | 203.8172 | 222.8967 | 240.4662 | 255.3391 | 236.8556 | 282.8053 | 491.2861 | 210.4543 | 170.5722 | 175.9735 | 242.5073 | 221.5457 | 485.2596 | 423.5073 | 664.1033 | 782.2861 | 412.3099 | 475.6106 | 361.1033 | 629.3894 | 168.0914 | 126.8318 | 127.6371 | 168.0914 | 135.7788 | 177.5073 | 246.351 | 158.4543 | 159.1682 | 129.1033 | 183.0768 | 183.2477 | 91.3894 | 104.8967 | 104.1033 | 135.6755 | 138.2066 | 135.3894 | 194.5192 | 106.4159 | 455.0238 | 372.8053 | 498.4927 | 381.4278 | 248.5576 | 290.8821 | 491.0675 | 423.9497 | 578.3483 | 327.9616 | 396.7907 | 701.9881
2204 | 26.2384 | 364.8351 | 380.3929 | 591.993 | 314.8351 | 252.1614 | 265 | 398.2351 | 261.2948 | 236.9965 | 165.7684 | 342.9895 | 275.3685 | 344.2351 | 260.7614 | 498.614 | 333.3895 | 306.0737 | 210.9965 | 348.2281 | 295.6175 | 287.3088 | 552.2983 | 763.7754 | 241.3929 | 435.614 | 372.4597 | 606.3332 | 527.2316 | 182.7719 | 201.386 | 294.0035 | 194.6947 | 239.3053 | 229.9193 | 356.8386 | 126.3053 | 481.6912 | 410.7754 | 720.5473 | 553.7545 | 234.2211 | 241.9193 | 323.8386 | 199.621 | 103.9228 | 96.7684 | 103.6175 | 177.9228 | 360.2281 | 285.8596 | 619.3018 | 243.4632 | 215.386 | 176.1474 | 347.2843 | 171.2281 | 288.9298 | 222.4736 | 396.0035 | 244.7052 | 228.1509 | 222.9895 | 340.621 | 211.1614 | 312.6982 | 263.6947 | 522.8351 | 407.8386 | 236.8526 | 260.6071 | 585.2386 | 345.0772 | 331.8456 | 335.4597 | 502.0035 | 525.2316 | 290.4597 | 236.6245 | 449.8351 | 374.9263 | 229.3053 | 209.8561 | 275.614 | 218.8526 | 278.386 | 279.0737 | 478.9158 | 217.6175 | 151.1579 | 187.6842 | 225.4771 | 212.6071 | 487.3018 | 405.9368 | 678.9193 | 773.6982 | 457.0667 | 467.2211 | 378.7754 | 637.2386 | 155.8491 | 109.1509 | 108.9263 | 155.8491 | 149.7017 | 157.6982 | 236.2281 | 167.007 | 175.9228 | 146.621 | 213 | 157.8456 | 97.386 | 89.0772 | 118.3018 | 129.6877 | 167.8386 | 141.7719 | 199.993 | 97.6912 | 509.9123 | 368.6877 | 518.4562 | 401.614 | 271.2316 | 316.2246 | 421.7754 | 379.8456 | 638.0667 | 314.3123 | 441.9403 | 681.4911
2205 | 26.2494 | 326.7316 | 405.2417 | 571.1319 | 279.1476 | 275.405 | 262.0471 | 414.7897 | 222.6374 | 225.5369 | 165.2371 | 303.6891 | 223.3109 | 360.1633 | 236.2371 | 490.2417 | 348.4788 | 298.4898 | 200.7897 | 335.4207 | 303 | 291.7159 | 517.3155 | 784.5526 | 267.584 | 429.7472 | 363.4945 | 685.4991 | 527.1365 | 189.405 | 204.7472 | 302.2417 | 203.7159 | 233.416 | 216.7316 | 332.0738 | 118.7159 | 477.9105 | 429.3155 | 745.81 | 504.4317 | 204.4898 | 229.8053 | 296.5683 | 216.4945 | 104.2528 | 95.7002 | 111.7159 | 175.2998 | 352.7002 | 332.4474 | 600.3422 | 252.8478 | 219.7316 | 157.5793 | 267.5369 | 163.1633 | 309.4945 | 262.5683 | 410.3266 | 286.7629 | 218.0895 | 188.7002 | 356.2417 | 237.2684 | 334.9686 | 273.8791 | 486.5212 | 386.1162 | 256.584 | 230.6578 | 606.7205 | 346 | 328.5683 | 324.8838 | 514 | 526.9262 | 284 | 267.1052 | 410.1209 | 385.4631 | 219.5683 | 241.9419 | 266.7048 | 242.3422 | 282.5526 | 269.2528 | 451.9419 | 228.4003 | 166.1633 | 160.2684 | 275.8681 | 186.595 | 472.9954 | 452.1052 | 668.6845 | 796.8635 | 422.6264 | 432.1209 | 401.8791 | 662.6578 | 166.179 | 98.1052 | 120.7159 | 166.179 | 178.8681 | 177.7316 | 227.6264 | 190.1947 | 173.1209 | 164.1209 | 211.2103 | 158.6845 | 101.4631 | 90.3579 | 99.7002 | 113.9105 | 142.2684 | 150.6421 | 177.5212 | 94.8053 | 473.8948 | 353.9843 | 497.6734 | 394.3155 | 274.6264 | 298.4003 | 442.821 | 378.9843 | 602.9105 | 330.3579 | 501 | 780.6734
2206 | 26.2604 | 321.0486 | 365.0663 | 603 | 304.2568 | 249 | 229.8627 | 431.9823 | 276.2344 | 230.084 | 186.1681 | 269.6947 | 250.7918 | 367.8273 | 258.6947 | 451.1858 | 335.1373 | 331.2521 | 215.9468 | 298.4911 | 300.6592 | 298.3894 | 492.7479 | 786.31 | 245.4426 | 416.5929 | 341.2568 | 600.8319 | 500.6461 | 166.4603 | 192.5089 | 263.8982 | 212.6284 | 256.3716 | 215.7302 | 340.0486 | 126.4071 | 481.4781 | 407.394 | 697.4781 | 517.2698 | 238.5752 | 237.2876 | 283.9337 | 190.7965 | 115.6769 | 113.1018 | 117.5752 | 159.8319 | 366.0308 | 335.8805 | 613.7965 | 291.4911 | 219.6461 | 193.5929 | 272.5929 | 174.5929 | 284.8142 | 249.3231 | 426.2652 | 266.8497 | 219.5089 | 206.1018 | 318 | 240.3053 | 354.6947 | 292.5929 | 498.1326 | 427.084 | 236.3763 | 219.239 | 549.6415 | 351.4958 | 313.5929 | 286.5621 | 511.761 | 519.0532 | 284.1018 | 260.8497 | 393.9514 | 381.3231 | 205 | 219.5752 | 226.7656 | 254.1681 | 281.4603 | 280.6769 | 428.0486 | 260.1503 | 175.6592 | 164.4248 | 246.1373 | 211.9823 | 531.3539 | 442.7965 | 697.5266 | 821.7255 | 429.6106 | 415.8497 | 419.0663 | 650.4248 | 169.3231 | 90.5266 | 129.1195 | 169.3231 | 148.4248 | 174.6947 | 228.4024 | 181.084 | 154.5089 | 146.0177 | 194.3231 | 183.5574 | 94.9823 | 93.8982 | 114.5574 | 113.3053 | 146.7302 | 149.1373 | 190.0486 | 103.7124 | 482.6947 | 362.8805 | 472.5752 | 369.7479 | 280.3894 | 332.8982 | 441.8142 | 386.5574 | 604.0355 | 334.7124 | 501.2035 | 749.5705
2207 | 26.2714 | 295.1141 | 344.464 | 603.4563 | 330.9468 | 250.825 | 217.0454 | 420.9468 | 242.5437 | 220.7453 | 169.2015 | 268.711 | 201.3611 | 334.0798 | 256.7985 | 448.5626 | 350.464 | 309.5094 | 194.247 | 291.9468 | 282.8516 | 293.1141 | 519.9657 | 847.5892 | 264.8328 | 414.1406 | 370.2281 | 611.2968 | 522.175 | 197.9125 | 194.1484 | 264.3688 | 226.0875 | 242.7985 | 231.7985 | 314.7985 | 129.4297 | 515.6389 | 451.0077 | 729.1295 | 511.5815 | 230.5626 | 230.6578 | 298.7187 | 192.1938 | 104.4563 | 112.8594 | 107.2813 | 172.6921 | 336.9314 | 328.0532 | 600.9357 | 287.7985 | 238.5515 | 190.6844 | 270.711 | 171.1141 | 293.2547 | 257.2359 | 361.9923 | 290.7453 | 223.5437 | 208.3688 | 319.0266 | 238.3233 | 349.6046 | 283.753 | 466.8516 | 420.1406 | 274.6578 | 236.2624 | 520.3422 | 397.3765 | 314.1063 | 315.6046 | 490.175 | 536.6467 | 286.711 | 280.5249 | 418.4031 | 393.3422 | 206.4828 | 211.2204 | 255.3233 | 241.8782 | 312 | 267.8594 | 399.1484 | 234.0875 | 159.3345 | 178.2547 | 261.5781 | 204.1406 | 508.9125 | 436.5515 | 710.7719 | 783.0875 | 432.4906 | 436.4374 | 400.4031 | 661.4031 | 182.0266 | 104.6844 | 135.6921 | 182.0266 | 161.7985 | 168.9202 | 180.9391 | 175.1672 | 158.3156 | 156.9391 | 206.5703 | 164.9657 | 87.9391 | 92.3156 | 93.9125 | 117.825 | 158.1218 | 163.2093 | 165.1406 | 110.4563 | 479.8859 | 357.5626 | 460.1141 | 392.5172 | 275.1141 | 332.1141 | 442.7264 | 366.4828 | 622 | 342.0266 | 508.7032 | 695.2924
2208 | 26.2824 | 297.5163 | 322.686 | 607.3791 | 314.4946 | 260.4512 | 260.083 | 403.231 | 239.2527 | 211.8845 | 162.4946 | 276.314 | 245.4008 | 361.538 | 260.8519 | 475.6605 | 324.3899 | 326.2202 | 233.3574 | 278.2744 | 305 | 293.2419 | 538.1985 | 889.9892 | 249.6426 | 422.6209 | 371.6209 | 568.5163 | 503.7039 | 200.9566 | 176.2961 | 273.9892 | 218.2419 | 250.2636 | 240.6535 | 316.8303 | 126.1372 | 476.7798 | 417.0938 | 670.9062 | 577.9132 | 264.7364 | 230.6535 | 281.7581 | 213.462 | 107.1155 | 104.1264 | 126.6426 | 150.917 | 384.1372 | 337.3031 | 516.0108 | 289.8303 | 203.2419 | 193.0938 | 275.5341 | 175.917 | 301.231 | 232.4225 | 390.9783 | 273.7977 | 219.6209 | 217.4837 | 322.7039 | 202.8845 | 331.8845 | 249.8232 | 483.314 | 428.2419 | 271.4946 | 213.231 | 514.6605 | 372.4512 | 350.0434 | 300.6643 | 479.4116 | 475.2457 | 298.9892 | 241.6643 | 400.1876 | 399.538 | 217.2419 | 240.2093 | 220.8953 | 273.769 | 308.462 | 261.4008 | 375.9892 | 222.7039 | 190.2202 | 183.3248 | 238.4116 | 210.3465 | 519.1589 | 408.3178 | 706.4729 | 779.2853 | 422.224 | 404.7907 | 395.2023 | 650.769 | 188.3574 | 109.3682 | 111.8953 | 188.3574 | 167.4946 | 145.2527 | 193.5992 | 188.3357 | 154.2636 | 173.8953 | 207.3357 | 184.2093 | 101.9892 | 88.3899 | 99.231 | 134.4008 | 133.1806 | 127.3899 | 173.6209 | 113.8736 | 480.5163 | 388.9566 | 461.3791 | 384.2853 | 277.1372 | 332.3682 | 394.3791 | 383.686 | 622.1264 | 351.5163 | 554.8953 | 776.3031
2209 | 26.2934 | 304.2565 | 358.6725 | 608.3362 | 310.7227 | 235.1005 | 234.3865 | 400.0121 | 241.4159 | 218.4159 | 161.6343 | 236.4454 | 257.286 | 383.0884 | 241.1474 | 421.1889 | 336.4956 | 306.7227 | 220.6135 | 290.6725 | 302.9203 | 288.8319 | 511.1005 | 887.1594 | 262.5251 | 421.3865 | 371.357 | 572.3449 | 481.6256 | 174.2478 | 202.8111 | 267.9705 | 208.7019 | 261.0797 | 254.5633 | 292.5754 | 133.7227 | 498.9116 | 401.2981 | 694.8821 | 520.3362 | 256 | 245.1179 | 280.0676 | 192.6048 | 101.6932 | 104.5841 | 137.8614 | 179.3865 | 394.2184 | 293.0797 | 531.4575 | 265.8526 | 200.9116 | 177.9913 | 224.3483 | 202.0295 | 291.0797 | 256.6222 | 379.2184 | 227.0502 | 217.4159 | 203.8111 | 292.8614 | 211.2184 | 337.8614 | 294.0884 | 449.4072 | 431.3188 | 265.5044 | 203.2184 | 460.7435 | 345.2981 | 373.8406 | 323.3657 | 497.7314 | 547.4541 | 289.3657 | 264.3657 | 366.7937 | 425.8024 | 218.7937 | 213.1681 | 234 | 278.5131 | 284 | 276.7522 | 368.5841 | 195.2184 | 178.9029 | 155.1594 | 255.0676 | 193.1092 | 537.6725 | 452.1386 | 688.4454 | 804.4956 | 485.3068 | 434.1889 | 448.3275 | 662.4454 | 177.9705 | 105.8319 | 125.2773 | 177.9705 | 163.0295 | 147.1386 | 177.5546 | 168.5459 | 160.643 | 184.5044 | 181.1681 | 157.8614 | 95.1386 | 99.2478 | 88.3865 | 148.2271 | 166.3657 | 137 | 170.1681 | 110.9203 | 490.4454 | 361.5546 | 457.8995 | 402.8406 | 282.9203 | 328.2773 | 399.4956 | 418.8406 | 622.0295 | 364.7729 | 575.3483 | 730.5546
2210 | 26.3044 | 280.3962 | 334.9812 | 594.3775 | 310.2075 | 268.3209 | 242.5472 | 449.8115 | 239.4719 | 224.4716 | 177.8491 | 225.4531 | 235.5469 | 365.7547 | 214.9244 | 429.9815 | 351.5472 | 302.4341 | 214.7169 | 265.17 | 292.2641 | 291.585 | 544.17 | 924.7732 | 275.6225 | 425.6228 | 385.6981 | 529.1887 | 525.9812 | 181.6981 | 182.8491 | 266.3022 | 179.1322 | 269.019 | 226.8491 | 314.2078 | 133.0566 | 513.9812 | 430.8678 | 709.189 | 513.7735 | 252.3775 | 223.8678 | 238.3584 | 213.3397 | 105.3962 | 101.0944 | 133.9812 | 183.4719 | 405.7359 | 310.83 | 570.5665 | 292.3775 | 221.1131 | 207.6416 | 262.4531 | 190.7172 | 303.3962 | 217.4528 | 386.0568 | 252.0375 | 216.0756 | 184.1509 | 289.7359 | 225.4528 | 329.9059 | 277.8113 | 471.8306 | 473.7928 | 248.9434 | 219.2641 | 481.7359 | 371.8491 | 345.4341 | 308.2075 | 464.3962 | 516.4716 | 276.4716 | 249.2075 | 405.3775 | 435.9434 | 258.585 | 211.7735 | 235.5094 | 235.7735 | 285.6603 | 269.1509 | 361.7737 | 207.3397 | 220.9059 | 179.1887 | 216.981 | 201.2075 | 519.7169 | 445.604 | 688.1697 | 820.4528 | 481.4528 | 451.585 | 473.0753 | 662.7735 | 180.5284 | 112.0566 | 125.9434 | 180.5284 | 160.7735 | 149.9622 | 177.0756 | 196.9625 | 149.7735 | 167.9434 | 176.4528 | 160.0188 | 95.0944 | 106.3962 | 96.0944 | 132.8113 | 150.6038 | 137.7547 | 163.4906 | 99.8113 | 484.7359 | 368.9244 | 420 | 375.3397 | 274.5281 | 326.8303 | 410.6228 | 398.1319 | 613.7359 | 378.9812 | 621 | 739.1319
2211 | 26.3154 | 278.7955 | 316.3677 | 570.409 | 314.3883 | 244.4484 | 250.9364 | 415.3265 | 219.3866 | 242.6942 | 177 | 186.4278 | 278.1632 | 369.5103 | 237.1632 | 366.8574 | 348.6735 | 289.7955 | 226.5722 | 237.8368 | 311.8556 | 278.6529 | 516.8368 | 980.2269 | 289.6341 | 402.3058 | 384.8161 | 520.3058 | 518.1408 | 180.3265 | 183.7955 | 221.5309 | 169.0395 | 246.5498 | 222.8986 | 278.7749 | 135.5722 | 494.0619 | 443.2457 | 653.1839 | 532.2251 | 231.0206 | 243.2646 | 247.7955 | 206.2852 | 103.3058 | 96.4897 | 118.7955 | 160.6117 | 392.1839 | 338.1632 | 491.5722 | 267.5928 | 231.3677 | 203.2646 | 227.3453 | 163.122 | 300 | 226.2027 | 346.8368 | 284.0412 | 200.5292 | 187.7749 | 276.3471 | 228.9794 | 291.9588 | 283.2664 | 407.8574 | 430.1615 | 246.1014 | 229.2251 | 468.1839 | 365.6134 | 337.2027 | 315.4897 | 460.5103 | 535.6735 | 294.3677 | 260.5704 | 379.9399 | 424.7767 | 244.1839 | 226.6341 | 234.6959 | 259.9381 | 298.7543 | 265.756 | 332.4502 | 206.3247 | 182.3058 | 167.5309 | 241.409 | 209.3058 | 533.8574 | 419.67 | 700.9399 | 825.4484 | 481.5516 | 434.8986 | 494.0412 | 679.9193 | 162.4691 | 113.9193 | 117.3883 | 162.4691 | 183.7955 | 162.7955 | 159.4072 | 164.6323 | 169.2045 | 161.8368 | 177.0412 | 171.9399 | 94.0807 | 103.3265 | 92.7955 | 142.0206 | 152.5309 | 145.4278 | 155.8161 | 110.3265 | 472.1632 | 392.3058 | 413.9605 | 371.2233 | 294.6134 | 316.0189 | 386.8161 | 419.0018 | 610.4673 | 369.3453 | 621.1632 | 775.1821
2212 | 26.3264 | 296.0739 | 316.2495 | 556.5635 | 301.5266 | 255.2448 | 291.5542 | 411.3834 | 251.388 | 240.2125 | 174.1917 | 200.6651 | 273.9099 | 365.9469 | 241.8614 | 360.2979 | 352.0901 | 297.9469 | 211.4573 | 236.2979 | 336.9953 | 287.2656 | 514.4203 | 927.1062 | 259.4573 | 408.2979 | 394.6166 | 523.8406 | 567.1062 | 190.9515 | 188.261 | 217.6859 | 204.2125 | 270.612 | 210.5104 | 291.7714 | 120.6328 | 481.9838 | 418.2078 | 642.1547 | 512.3834 | 227.4042 | 260.1755 | 259.2818 | 217.649 | 109.8245 | 96.5427 | 129.7552 | 173.649 | 383.7875 | 334.963 | 475.9307 | 252.1963 | 220.1917 | 216.4896 | 258.3302 | 175.2286 | 302.9838 | 251.5058 | 343.5427 | 278.037 | 228.2078 | 198.6651 | 275.4573 | 228.9099 | 304.4573 | 267.6005 | 401.4734 | 461.4688 | 256.2078 | 216.4042 | 464.8776 | 346.9515 | 366.1917 | 321.6859 | 454.4896 | 538.388 | 287.0531 | 274.5219 | 356.2818 | 403.4411 | 241.2286 | 199.2656 | 191.9146 | 272.8938 | 321.8614 | 253.4781 | 308.9146 | 242.3302 | 186.3672 | 162.2818 | 224.2286 | 216.8776 | 531.3349 | 492.4526 | 672.1917 | 840.9838 | 477.425 | 423.037 | 489.6166 | 665.4942 | 171.5797 | 92.8245 | 105.0531 | 171.5797 | 189.1386 | 170.0693 | 181.4734 | 175.8083 | 158.5958 | 160.649 | 164.8938 | 148.6328 | 109.2078 | 108.1594 | 101.4734 | 132.963 | 144.2979 | 162.649 | 160.1755 | 113.5797 | 476.1594 | 395.1386 | 381.2448 | 391.3302 | 275.425 | 354.4573 | 386.963 | 383.4411 | 653.0323 | 396.4688 | 631.1271 | 803.7505
2213 | 26.3374 | 325.1729 | 356.3147 | 579.6903 | 308.1318 | 254.3246 | 229.3296 | 379.9291 | 273.1269 | 257.9341 | 159.6853 | 187.2537 | 242.4875 | 365.5075 | 259.6244 | 350.2388 | 374.1219 | 295.6294 | 227.5075 | 229.6194 | 297.1368 | 299.4166 | 510.5684 | 937 | 274.005 | 404.6244 | 412.7363 | 507.4975 | 572.1169 | 224.1729 | 174.0759 | 236.5025 | 218.7413 | 249.2488 | 222.4925 | 283.5585 | 129.6194 | 503.1368 | 391.3756 | 610.1878 | 487.6903 | 232.3097 | 258.3706 | 265.4925 | 216.3756 | 113.8831 | 89.6953 | 134.995 | 173.1269 | 365.6903 | 318.8172 | 494.2028 | 285.2488 | 210.5684 | 201.5025 | 227.9391 | 183.6294 | 309.8632 | 232.8321 | 338.1978 | 291.3657 | 206.0709 | 187.6953 | 278.7363 | 211.5734 | 310.9291 | 299.3607 | 396.9341 | 415.5684 | 235.7612 | 227.1318 | 469 | 387.1219 | 351.1219 | 335.2438 | 431.2388 | 552.051 | 289.9341 | 236.7612 | 359.2997 | 428.4366 | 249.0659 | 224.5634 | 215.9391 | 263.9391 | 333.2388 | 295.9341 | 333.3147 | 215.6953 | 169.4925 | 177.6953 | 224.9291 | 220.2488 | 538.5485 | 460.5684 | 659 | 861.5734 | 513.8072 | 445.005 | 518.2537 | 694.6853 | 177.6853 | 94.2537 | 111.5025 | 177.6853 | 171 | 162.3147 | 173.3657 | 185.4316 | 156.6953 | 157.8731 | 156.8781 | 162.8781 | 88.7612 | 120.5585 | 98.8122 | 114.5634 | 144.1928 | 163.0659 | 162.6903 | 118.1828 | 489.8731 | 373.056 | 379.7612 | 369.7662 | 312.5585 | 364.1219 | 371.005 | 410.3147 | 630.2587 | 351.5075 | 670.4316 | 774.954
2214 | 26.3484 | 283.5983 | 358.5983 | 573.393 | 319.3966 | 280.7973 | 264.5952 | 357.7999 | 278.6003 | 256.0036 | 164.6054 | 191.197 | 220.2052 | 388.2016 | 251.5967 | 320.7999 | 370.2021 | 310.8019 | 241.197 | 214.7999 | 326.3992 | 258.2068 | 530.6024 | 930.7968 | 285.5993 | 407.4023 | 380.0015 | 489.194 | 546.5977 | 186 | 210.0005 | 236.5967 | 203.2037 | 254.0041 | 210.6039 | 271.7994 | 119.4023 | 524.5952 | 407.6075 | 603.5962 | 492.3987 | 228.8019 | 251.4023 | 243.7989 | 218.6003 | 128.5967 | 110.4002 | 121.7999 | 175.5988 | 377.002 | 332.8009 | 526.196 | 284.2011 | 226.6003 | 206.3992 | 226.7973 | 192.999 | 277.9995 | 272.194 | 359.7973 | 269.2011 | 228 | 204.7984 | 249.0031 | 234.5972 | 285.7984 | 270.4043 | 386.7994 | 439.8045 | 257.3961 | 250.803 | 463.197 | 382.4017 | 349.8035 | 321.4012 | 399.5988 | 527.6029 | 285.4023 | 267.2006 | 329.6018 | 424.5993 | 256.3992 | 228.2006 | 214.7973 | 268.8004 | 308.002 | 289.2011 | 331.5962 | 238.4012 | 159.2047 | 192.5972 | 205.8014 | 218.4007 | 513.806 | 480.2021 | 665.0031 | 884.7973 | 496.9985 | 453.998 | 534.0031 | 698.805 | 176.4023 | 101.5988 | 118 | 176.4023 | 171.8004 | 165.999 | 148.1996 | 168.5993 | 177.2001 | 157.8025 | 164.1996 | 176.0026 | 112.9974 | 107.1986 | 96.3992 | 122.803 | 158.2001 | 172.6003 | 166.3982 | 103.7989 | 482.999 | 360.4023 | 405.3982 | 411.1996 | 306.4023 | 366.2052 | 379.1975 | 426.0051 | 661.2042 | 378.4038 | 663.2042 | 821.2016
2215 | 26.3594 | 273.1859 | 338.2034 | 516.0265 | 300.2214 | 268.0709 | 237.6197 | 357.4248 | 275.6902 | 278.4778 | 203.8141 | 170.3363 | 261.2124 | 384.0085 | 225.1504 | 320.8496 | 385.5132 | 321.7521 | 217.5752 | 216.5487 | 317.8761 | 303.5662 | 541.1414 | 903.4513 | 284.885 | 423.938 | 392 | 438.6017 | 536.4338 | 185.1504 | 212.7256 | 205.4778 | 233.2744 | 280.4778 | 238.2389 | 269.1239 | 134.0265 | 492.947 | 450.4953 | 570.3893 | 486.2479 | 244.4248 | 263.6902 | 236.2744 | 213.7786 | 106.3983 | 108.8141 | 121 | 166 | 391.938 | 344.6727 | 490.5397 | 290.4513 | 231.3363 | 198.0885 | 207.2744 | 192.009 | 287.5932 | 231.7966 | 341.9735 | 276.7256 | 231.6107 | 193.4868 | 275.3363 | 215.5487 | 274.4868 | 297.4158 | 379.6637 | 470.115 | 235.4957 | 265.9291 | 435.115 | 388.5662 | 368.7167 | 329.938 | 382.7786 | 544.9026 | 303.8496 | 269.6637 | 342.7256 | 425.7966 | 247.0265 | 227.6902 | 196.1239 | 275.6107 | 332.2833 | 306.0709 | 301.5752 | 247.5752 | 190.4778 | 175.0355 | 207.0175 | 225.4868 | 543.3713 | 494.8761 | 672.6457 | 861.2389 | 499.2304 | 438.4248 | 563.7346 | 729.5043 | 190.3893 | 93.6991 | 114.3893 | 190.3893 | 179.2479 | 156.938 | 145.4248 | 166.8231 | 178 | 176.7876 | 161.2124 | 191.9645 | 94.9115 | 97.6991 | 90.8496 | 143.6637 | 157.3009 | 172.2389 | 145.8406 | 95.2124 | 474.1504 | 379.6991 | 386.5397 | 408.8496 | 318.053 | 399.1414 | 364.2479 | 463.2389 | 690.6017 | 405.2389 | 705.5577 | 829.7521
2216 | 26.3704 | 283.8766 | 319.7405 | 511.5158 | 325.6519 | 282.4968 | 270.2595 | 355.405 | 266.1108 | 257.5506 | 187.0569 | 179 | 266.7184 | 331.3355 | 220.4272 | 315.9114 | 368.5411 | 311.8418 | 218.2468 | 231.1677 | 317.6393 | 281.3924 | 519.864 | 882.0917 | 296.0348 | 416.405 | 396.0443 | 432.5158 | 555.6867 | 186.7184 | 201.4842 | 181.1804 | 235.0791 | 251.2373 | 239.2342 | 275.4272 | 121.8766 | 515.8987 | 402.2121 | 561.4937 | 505.3703 | 249.595 | 243.3259 | 238.9778 | 193.7405 | 116.0791 | 97.6741 | 117.405 | 166.4494 | 393.1677 | 351.4398 | 470.405 | 281.2247 | 235.731 | 191 | 208.1804 | 219.7975 | 326.7658 | 247.5633 | 351.4272 | 278.7405 | 240.9557 | 205.0665 | 281.5285 | 237.5823 | 275.5063 | 275.0222 | 371.6741 | 460.538 | 267.4494 | 241.3924 | 414.5285 | 367.5158 | 347.2121 | 328.6962 | 360.4937 | 523.5285 | 304.5285 | 262.3481 | 328.3386 | 441.5633 | 232.8544 | 219.0095 | 202.4272 | 282.2595 | 348.6202 | 338.0222 | 308.0886 | 253.6393 | 178.7627 | 185.5063 | 176.7405 | 226.0569 | 488.7849 | 484.0791 | 612.2247 | 867.402 | 541.4398 | 450.5602 | 590.1677 | 704.5158 | 183.0665 | 99.5506 | 105.0443 | 183.0665 | 189.6076 | 154.5728 | 150.8196 | 182.1234 | 176.8766 | 176.6741 | 166.269 | 175.4272 | 103.1234 | 104 | 95.1234 | 138.595 | 145.1582 | 160.6519 | 129.2912 | 100.0443 | 469.8766 | 381.9557 | 367.9778 | 419.1899 | 291.0569 | 369.3259 | 373.2595 | 450.0791 | 687.4367 | 382.867 | 729.3164 | 820.2912
2217 | 26.3814 | 279.763 | 342.526 | 519.0705 | 321.7109 | 236.8202 | 243.9193 | 351.5311 | 290.3361 | 252.6823 | 169.763 | 175.2084 | 276.6017 | 368.8958 | 214.0521 | 283.4975 | 337.5311 | 317.5445 | 225.0521 | 238.3361 | 334.521 | 300.237 | 543.9193 | 839.3412 | 276.2656 | 406.1328 | 403.6017 | 448.6017 | 535.7681 | 196.3647 | 178.763 | 174.3983 | 226.1849 | 237.5883 | 267.8101 | 262.4168 | 120.8437 | 513.5496 | 435.1849 | 559.9378 | 513.4454 | 258.4454 | 240.526 | 230.5781 | 211.3126 | 104.8151 | 98.8151 | 109.2656 | 165.6302 | 410.526 | 317.1328 | 455.6302 | 290.2941 | 224.5546 | 190.526 | 201.6353 | 212.7294 | 301.9816 | 227.0858 | 340.0756 | 299.3933 | 231.9765 | 228.1328 | 277.9765 | 270.8101 | 274.9295 | 292.4269 | 365.9428 | 486.2084 | 274.6874 | 247.6773 | 406.474 | 390.9479 | 373.7865 | 333.9714 | 364.3882 | 525.1899 | 301.4504 | 272.9244 | 295 | 420.1378 | 227.2135 | 250.3647 | 199.8437 | 263.5025 | 302.3177 | 336.9428 | 338.6067 | 277.6302 | 208.763 | 177.5832 | 202.8437 | 218.479 | 528.1328 | 485.8488 | 620.8202 | 908.3596 | 503.8151 | 483.9193 | 622.9295 | 714.2033 | 195.7059 | 96.3412 | 116.3933 | 195.7059 | 169.8151 | 159.526 | 168.9765 | 184.1042 | 173.9479 | 176.6302 | 180.0521 | 168.3412 | 111.2656 | 102.1042 | 95.2084 | 141.9949 | 134.479 | 155.2891 | 150.2891 | 112.1042 | 471.2135 | 366.1042 | 365.2656 | 433.5731 | 280.6353 | 369.3698 | 353.0807 | 438.5781 | 723.0807 | 349.2941 | 689 | 832.047
2218 | 26.3924 | 283.9853 | 339.2551 | 472.9941 | 324.9971 | 263.4956 | 237.4897 | 368.2815 | 252.5103 | 251.9707 | 167.0059 | 165.7419 | 257.9941 | 363.783 | 212.4956 | 277.2258 | 365.2493 | 274.9883 | 227.2346 | 205.7449 | 313.4721 | 297.0117 | 539.4839 | 851.6979 | 288.5044 | 420.478 | 378.5132 | 421.7683 | 565.7449 | 173.5044 | 181.9883 | 190.5132 | 224.7654 | 286.0147 | 237.4927 | 239.7449 | 129.0029 | 492.5103 | 429.5279 | 506.2463 | 513.4663 | 252.9824 | 238.2522 | 219.0205 | 203.2199 | 101.4985 | 95.4985 | 122.2522 | 159.9941 | 413.2375 | 333.2229 | 455.2287 | 318.4956 | 235.0029 | 199.9677 | 219.5103 | 177.478 | 346.0029 | 258.522 | 321.7478 | 289.5044 | 243.7625 | 232.7566 | 295.9941 | 238.2493 | 314.7742 | 318.7713 | 347.7273 | 483.2229 | 286.2698 | 215.7273 | 410.9912 | 388.2669 | 357.9971 | 319.2551 | 336.4721 | 545.0381 | 315.5103 | 289.2581 | 293.0059 | 457.4956 | 249.4839 | 242.2111 | 209.2493 | 277.0029 | 321.9736 | 314.2405 | 332.044 | 268.5044 | 203.7625 | 202.4985 | 207.5132 | 248.5015 | 537.2434 | 516.2727 | 623.8153 | 877.6921 | 501.4956 | 474 | 678.478 | 686.7097 | 176.2287 | 93.7419 | 114.9795 | 176.2287 | 165.5015 | 161.9883 | 181.0117 | 178.2493 | 176.0029 | 165.261 | 172.0147 | 174.9648 | 117.522 | 95.0029 | 88.7331 | 119.7302 | 161.5103 | 157.2375 | 142.0176 | 112.2317 | 491.739 | 360.4985 | 371.7713 | 396.0117 | 300.2551 | 381.4868 | 369.2317 | 442.2258 | 724.5249 | 393.9472 | 709.6891 | 820.6891
2219 | 26.4034 | 289.3224 | 340.8005 | 475.8613 | 323.292 | 273.7543 | 256.1083 | 337.5693 | 241.4769 | 271.0146 | 169.893 | 179.7543 | 267.1387 | 338.893 | 219.0925 | 294.6764 | 378.0316 | 289.354 | 235.9379 | 208.1229 | 338.8613 | 285 | 547.8917 | 897.4148 | 280.8613 | 445.354 | 362.9075 | 415.5548 | 558.1837 | 171.354 | 188.7688 | 177.7847 | 208.7384 | 267.8613 | 241.5998 | 239.7688 | 128.8771 | 486.708 | 405.1849 | 499.0146 | 535.4452 | 270.2153 | 243.8467 | 198.7847 | 232.1691 | 108.2312 | 94.3844 | 124.708 | 173.5852 | 429.6618 | 362.3078 | 465.753 | 318.8151 | 241.1545 | 224.3686 | 200.3686 | 197.9075 | 347.708 | 238.354 | 323.7384 | 281.0766 | 233.0925 | 230.1849 | 301.4769 | 238.2153 | 296.0158 | 297.2616 | 375.708 | 492.9525 | 266.5231 | 240.8309 | 420.5231 | 385.9088 | 362.8309 | 316.354 | 358.4611 | 501.5535 | 300.5535 | 271.8455 | 284.1229 | 458.3382 | 267.8309 | 273.107 | 210.5231 | 272.1691 | 362.8626 | 328.708 | 290.1387 | 270.8005 | 188.1229 | 204.2616 | 195.8309 | 243.5998 | 546.8771 | 496.4002 | 559.8771 | 929.4148 | 520.6472 | 478.9696 | 718.0012 | 730.354 | 199.5693 | 104.354 | 125.5377 | 199.5693 | 170.5389 | 175.3078 | 164.5535 | 176.646 | 178.4927 | 156.354 | 159.6156 | 206.0304 | 96.5693 | 96.9696 | 97.8917 | 138.1691 | 155.1849 | 172.8771 | 134.2007 | 126.292 | 502.2153 | 381.8784 | 348.9538 | 391.8467 | 301.8005 | 400.2312 | 387.2153 | 456.7993 | 705.5389 | 428.3832 | 764.1533 | 894.247
2220 | 26.4144 | 264.7385 | 357.5354 | 479.3232 | 304.523 | 289.1785 | 274.3446 | 331.5938 | 238.3569 | 250.4062 | 189.523 | 191.0707 | 268.8799 | 373.5847 | 224.4523 | 271.2615 | 395.0214 | 302.0247 | 221.7508 | 199.4523 | 326.9877 | 289.1077 | 530.2031 | 869.4276 | 278.8462 | 436.3906 | 365.4893 | 433.1998 | 532.9539 | 180.1908 | 182.7632 | 180.2738 | 200.8799 | 258.4523 | 239.6678 | 222.8092 | 133.1661 | 504.6554 | 418.0954 | 470.4646 | 502.4646 | 269.6431 | 263.8092 | 207.5724 | 223.7138 | 118.0707 | 97.9539 | 136.9046 | 186.7845 | 431.2368 | 370.3816 | 438.1077 | 319.8709 | 255.2245 | 198.0831 | 181.2031 | 193.9169 | 348.9507 | 244.7262 | 323.5477 | 274.1077 | 232.4276 | 234.3322 | 303.2862 | 245.8586 | 312.7262 | 303.7508 | 384.0707 | 443.3692 | 259.5107 | 252.8462 | 426.3816 | 424.44 | 365.8092 | 330.3939 | 336.9877 | 505.4309 | 289.0954 | 266.2648 | 273.5354 | 454.8462 | 270.8092 | 252.1077 | 216.6554 | 271.9293 | 395.4276 | 341.4523 | 306.1201 | 293.0123 | 181.9169 | 203.0831 | 191.1414 | 231.5354 | 558.8339 | 507.3692 | 554.0337 | 904.9877 | 563.0954 | 493 | 779.4186 | 738.6431 | 201.2615 | 109.9046 | 107.2278 | 201.2615 | 189 | 175.1661 | 152 | 174.3816 | 185.2368 | 169.5724 | 169.4646 | 186.2492 | 107.2985 | 108.8092 | 81.5724 | 133 | 165.0831 | 172.5107 | 161.6308 | 112.1785 | 503.2862 | 427.0461 | 350.9293 | 413.7262 | 322.3692 | 409.7969 | 391.0247 | 424.44 | 722.0831 | 377.6554 | 742 | 918.6061
2221 | 26.4254 | 271.7139 | 340.1267 | 499.4099 | 293.4337 | 279.8465 | 266.5782 | 309.1505 | 250.0238 | 278.5723 | 176.4307 | 173.7079 | 265.4515 | 407.8703 | 227.2921 | 265.4307 | 370.0298 | 321.4337 | 238.4307 | 203.7228 | 331.0179 | 300 | 544 | 840.5544 | 293.5663 | 405.301 | 387.7139 | 400.5931 | 553.4158 | 192.2951 | 203.4247 | 176.994 | 190.8703 | 251.2772 | 259.1416 | 220.4337 | 126.4337 | 508.9881 | 429.2981 | 474.1386 | 517.2981 | 261.9821 | 246.8406 | 222.1386 | 215.5723 | 109.5782 | 119.5604 | 129.4218 | 167.003 | 418.8733 | 368.2653 | 458.7288 | 343.4128 | 230.8674 | 199.2951 | 185.8435 | 199.2861 | 324.4396 | 245.4307 | 333.8703 | 290.1505 | 230.7377 | 219.7228 | 315.4337 | 273.1446 | 319.4396 | 325.5812 | 378.1535 | 452.7228 | 246.4426 | 268.997 | 431.9911 | 415.2861 | 370.0149 | 347.997 | 331.003 | 557.5842 | 297.4367 | 318.8554 | 271.2921 | 471.8554 | 267.2891 | 269.0089 | 224.1356 | 285.5723 | 382.5782 | 350.5872 | 325 | 301.5723 | 193.5812 | 195.7109 | 167.003 | 227.003 | 569.5723 | 513.2891 | 521.2921 | 907.5872 | 550.2623 | 497.006 | 812.7109 | 745.8614 | 199.4367 | 105.2832 | 130.7049 | 199.4367 | 191.8614 | 164.4277 | 161.4426 | 185.4277 | 174.304 | 187 | 165.9851 | 181.301 | 115.7019 | 109.8674 | 100.7168 | 135.5753 | 159.1416 | 163.4486 | 157.4277 | 113.1475 | 517.7228 | 398.2861 | 364.2861 | 423.0149 | 321.994 | 420.608 | 409.003 | 437.0328 | 715.2832 | 386.2802 | 713.386 | 899.2861
2222 | 26.4364 | 275.178 | 327.3247 | 461.733 | 305.5812 | 261.1047 | 269.6439 | 319.3141 | 275.6753 | 270.1518 | 185.0733 | 162.7016 | 303.4031 | 422.5392 | 231.1361 | 265.4188 | 421.7906 | 336.5655 | 241.1047 | 221.5812 | 350.5549 | 299.1047 | 544.8953 | 814.6859 | 289.7749 | 429.2094 | 392.3718 | 420.8847 | 529.6126 | 198.4502 | 192.822 | 177.1466 | 216.5812 | 248.9372 | 249.8377 | 234.0733 | 129.9267 | 499.7435 | 435.3769 | 467.5969 | 535.0157 | 227.2251 | 223.7749 | 217.6859 | 213.4188 | 118.3141 | 103.6859 | 118 | 172 | 447.4922 | 349.8168 | 489.267 | 325.0576 | 245.3141 | 219.7749 | 173.4451 | 199.7016 | 335.5392 | 251.9843 | 351.2251 | 315.2357 | 269.4031 | 233.1047 | 325.4922 | 284.3561 | 333.5235 | 337.822 | 387.0471 | 467.5969 | 275.6702 | 262.8063 | 417 | 419.8796 | 381.5706 | 343.5969 | 329.1361 | 562.5863 | 306.822 | 314.5078 | 286.178 | 461.2408 | 278.9686 | 288.4765 | 212.8953 | 291.178 | 388.3298 | 372.822 | 324.1047 | 308.9686 | 210.8953 | 194.7749 | 175.5812 | 232.2984 | 568.6126 | 517.8063 | 531.1047 | 941.1623 | 508.911 | 499.2408 | 838.3352 | 753.2984 | 216.2827 | 109.9529 | 116.1937 | 216.2827 | 193.9267 | 168.9686 | 180.822 | 187.2827 | 211.5655 | 187 | 150.2514 | 199.2408 | 91.4188 | 120.733 | 98.733 | 145.8796 | 163.267 | 195.5078 | 152.7173 | 122.1937 | 528.4188 | 404.6702 | 374.5498 | 433.0785 | 328.7121 | 459.3193 | 406.2408 | 482.4502 | 722.9372 | 400.2776 | 636.6282 | 916.7121
2223 | 26.4474 | 285.9322 | 355.7177 | 460.6718 | 316.4858 | 263.6608 | 258.7002 | 326.0503 | 249.768 | 271.2363 | 190.7856 | 164.4858 | 322.1969 | 405 | 218.1072 | 259.1751 | 409.8425 | 346.407 | 238.6893 | 231.5536 | 329.4179 | 292.6499 | 547 | 803.5996 | 294.7856 | 430.2823 | 404.6893 | 386.1072 | 521.8249 | 174.4464 | 191.0788 | 194.4748 | 222.5142 | 276.0678 | 236.1072 | 237.9212 | 121.4179 | 524.3786 | 432.3829 | 463.7177 | 527.0678 | 224.3895 | 230.3392 | 210.9497 | 213.3895 | 116.0394 | 102.232 | 115.5142 | 175.4179 | 457.8359 | 380.7856 | 491.2604 | 350.8359 | 246.1466 | 219.814 | 199.2429 | 201.1751 | 317.6893 | 271.7396 | 353.361 | 339.9606 | 270.1751 | 231.9322 | 347.3326 | 292.5033 | 321.1357 | 337.6324 | 360.7177 | 471.1751 | 291.4858 | 261.8643 | 409.8534 | 424.0285 | 355.2823 | 347.1751 | 316.7287 | 541.4004 | 295.7571 | 307.2713 | 293.5142 | 433.6783 | 302.0109 | 292.1641 | 214.3786 | 299.7571 | 390.2254 | 367.6608 | 320.4464 | 322.6893 | 213.6214 | 205.3786 | 183.6893 | 243.2538 | 565.7965 | 513.7571 | 521.5427 | 979.1291 | 516.1182 | 501.8184 | 901.1247 | 758.9715 | 224.6214 | 124.1641 | 117.4464 | 224.6214 | 189.4573 | 179.2713 | 171.3107 | 196.8643 | 210.8425 | 191.3501 | 171.0678 | 188.1466 | 97.9146 | 112.5252 | 91.4573 | 153.7571 | 177.6893 | 191.3786 | 156.4967 | 123.7571 | 510.3676 | 428.2538 | 394.8249 | 392.7177 | 369.5536 | 425.8862 | 391.1072 | 464.6608 | 763.1182 | 441.639 | 607.221 | 948.8534
2224 | 26.4584 | 302.2126 | 339.4778 | 475.4954 | 314.8937 | 266.2477 | 285.1063 | 347.2477 | 263.4469 | 312.7255 | 187.3364 | 172.9071 | 356.9866 | 393.0485 | 222.4159 | 253.6636 | 397.5799 | 332.7079 | 258.6728 | 233.0619 | 343.1372 | 291.0753 | 557.6594 | 774.6904 | 284.8762 | 431.3406 | 404.646 | 392.677 | 523.1372 | 175.1992 | 205.4469 | 179.677 | 219.5841 | 280.7833 | 237.8318 | 218.7079 | 128.677 | 510.4025 | 492.3096 | 441.9866 | 535.3364 | 236.5397 | 231.6285 | 207.518 | 221.3406 | 122.0929 | 94.3674 | 102.5707 | 183 | 470.7389 | 379.5975 | 484.3188 | 366.323 | 264.2611 | 212.1197 | 206.8452 | 203.4159 | 316.031 | 289.0176 | 379.7699 | 323.2477 | 271.1238 | 233.031 | 379.2787 | 282.1816 | 312.8008 | 342.2033 | 361.5975 | 485.6904 | 298.2921 | 263.7389 | 427.2703 | 413.969 | 366.031 | 351.031 | 323.7079 | 569.9866 | 296.8762 | 292.8628 | 279.6017 | 436.5799 | 314.8937 | 282.8142 | 222.3674 | 306.6904 | 420.1238 | 383.4912 | 319.2611 | 317.8008 | 219.5222 | 201.4159 | 181.7079 | 264.7079 | 581.6904 | 522.6285 | 519.8586 | 1013.7255 | 547.8452 | 527.9732 | 885.0753 | 757.4025 | 220.5088 | 109.323 | 112.3849 | 220.5088 | 206.646 | 173.5841 | 173.9381 | 204.8762 | 190.5046 | 196.1548 | 174.1682 | 206.5841 | 128.4159 | 121.5088 | 109.9381 | 155.5222 | 176.031 | 182.5707 | 171.0176 | 124.8762 | 488.354 | 442.9247 | 402.5975 | 395.8586 | 374.6151 | 478.7213 | 419.6419 | 466.2787 | 777.4025 | 414.9071 | 579.6904 | 936.5531
2225 | 26.4694 | 323.9527 | 335.0296 | 454.4003 | 308.7179 | 264.0828 | 288.2644 | 328.9882 | 274.4417 | 299.7415 | 201.6233 | 175.9823 | 331.9764 | 378.7297 | 225.0473 | 221.6056 | 437.2762 | 338.3826 | 295.2939 | 238.3885 | 367.065 | 317.0532 | 598.777 | 784.6056 | 287.9704 | 415.0355 | 408.6824 | 406.4181 | 552.4299 | 187.353 | 195.6529 | 177.6588 | 227.0118 | 294.6647 | 245.1064 | 213.9882 | 128.3353 | 488.6943 | 467.641 | 430.0532 | 547.9468 | 225.3471 | 246.6115 | 250.9409 | 207.7179 | 115.6647 | 108.6351 | 91.359 | 185.3471 | 488.8007 | 408.0118 | 517.2644 | 376.7238 | 271.3471 | 245.3176 | 215.3235 | 202.6943 | 308.2998 | 264.9941 | 376.6883 | 317.3944 | 263.3353 | 226.9764 | 357.3117 | 305.9527 | 320.0946 | 323.0946 | 385.9882 | 500.2998 | 296.3057 | 266.712 | 470.8463 | 424.7179 | 366.3471 | 355.3708 | 315.6351 | 559.0591 | 312.712 | 292.0709 | 264.0118 | 487.3412 | 301.0059 | 291.9823 | 238.3117 | 320.9645 | 419.3767 | 407.7415 | 334.0591 | 318.3885 | 227.6588 | 199.353 | 184.0296 | 268.0355 | 587.9172 | 541.9704 | 533.2289 | 973.5819 | 567.3885 | 504.4417 | 913.4003 | 746.424 | 214.7002 | 108.9941 | 110.3412 | 214.7002 | 209.3412 | 178.6647 | 175.9882 | 217.359 | 214.2821 | 184.6588 | 189.0236 | 205.2939 | 119.647 | 112.6824 | 111.3176 | 169.0236 | 181.712 | 179.0709 | 160.7415 | 133 | 497.7297 | 427.0059 | 420.9527 | 425.9941 | 381.353 | 485.935 | 447.8995 | 455.712 | 768.7711 | 421.3353 | 592.9586 | 944
2226 | 26.4804 | 313.9091 | 350.5615 | 470.3957 | 328.4332 | 290.9572 | 274.0375 | 325 | 299.877 | 340.2942 | 185.6096 | 167.2192 | 327.1284 | 414.9466 | 238.5668 | 207.9519 | 414.1765 | 350.8289 | 274.7005 | 251.4385 | 386.5668 | 320.401 | 613.0587 | 775.123 | 279.0428 | 449.9413 | 424.0803 | 432.5668 | 590.5615 | 196.7808 | 195.1711 | 174.6524 | 230.6524 | 301.2995 | 261.1871 | 214.8663 | 126.2192 | 507.3851 | 461.7808 | 447.6524 | 532.4332 | 243.9038 | 228 | 221.9625 | 220.8289 | 114.6524 | 97.3048 | 104.1711 | 185.4813 | 531.9144 | 419.2995 | 502.6899 | 392.1765 | 265.5722 | 234.4385 | 212.3476 | 216.5187 | 309.1658 | 264.3904 | 384 | 349.123 | 269.5615 | 229.0803 | 363.9038 | 304.9466 | 347.8289 | 340.401 | 388.9519 | 503.2514 | 282.1337 | 286.5187 | 426.6471 | 432.9625 | 367.5241 | 363.8289 | 309.1711 | 572.3957 | 328 | 318.0856 | 272.1711 | 500.0803 | 313.7754 | 296.0803 | 243.5134 | 305.5241 | 430.8716 | 438.9091 | 360.6043 | 337 | 226.3904 | 227.2033 | 192.9572 | 263.3154 | 573.2086 | 564.3263 | 506.3423 | 962.0322 | 579.3957 | 539.262 | 928.3529 | 765.877 | 226 | 107.3476 | 113.3476 | 226 | 205.3957 | 175.9144 | 179.2995 | 213.8342 | 203.9519 | 205.9841 | 205.6899 | 201.3851 | 121.6899 | 119.0428 | 108.0856 | 171.0909 | 203.2567 | 199.1765 | 183.5668 | 124.6577 | 528.3851 | 436.2995 | 395.9625 | 424.3476 | 384.5241 | 475.8182 | 425.4706 | 473.4332 | 816.5134 | 438.337 | 562.3154 | 946.0856
2227 | 26.4914 | 322.4808 | 359.5605 | 453.6814 | 332.2802 | 288.6401 | 294.239 | 332.1977 | 294.1152 | 360.0028 | 195.2361 | 162.7198 | 346.6757 | 441.5605 | 223.9232 | 225.6785 | 407.3599 | 344.7994 | 271.0768 | 241.3599 | 379.4808 | 303.7966 | 585.1152 | 772.2475 | 295.3954 | 475.3656 | 440.8407 | 424.4011 | 595.9616 | 201.6401 | 209.4779 | 180.8378 | 229.6401 | 312.8407 | 227.239 | 214.643 | 127.8378 | 517.0825 | 471.3186 | 445.9203 | 539.5192 | 271.7994 | 232.3186 | 211.1181 | 207.9616 | 107.8819 | 94.5605 | 112.3599 | 182.3983 | 517.9232 | 444.1565 | 516.8407 | 389.6785 | 247.0797 | 223.2802 | 215.1593 | 223.2006 | 333.4395 | 272.3983 | 401.6344 | 375.0384 | 281.4395 | 257.357 | 398.9972 | 334.7994 | 329.2034 | 321.6373 | 408.1181 | 519.2034 | 274.7994 | 281.6842 | 439.5605 | 435.4338 | 369.2774 | 349.8819 | 314.8407 | 549.9232 | 336.9972 | 323.0797 | 275.3215 | 516.8407 | 348.0356 | 311.4011 | 263.0413 | 326.3513 | 423.5605 | 446.761 | 356.8051 | 347.4367 | 235.1181 | 250.4452 | 187.4011 | 245.6757 | 604.4779 | 607.7255 | 535.6757 | 1010.5548 | 575.6373 | 539.239 | 887.3684 | 765.1536 | 226 | 115.1977 | 108.9616 | 226 | 199.8378 | 177.3983 | 198.0384 | 201.0768 | 227.7966 | 242.3627 | 212.643 | 221.8791 | 134.4011 | 117.4011 | 103.7226 | 154.2418 | 212.1209 | 204.596 | 174.6814 | 111.1593 | 531.9644 | 472.3129 | 397.3542 | 439.7553 | 381.5989 | 502.3186 | 446.2802 | 490.2361 | 834.9616 | 456.404 | 532.0797 | 966.9146
2228 | 26.5024 | 304.9738 | 343.6016 | 453.4435 | 339.9323 | 277.9512 | 281.019 | 344.2556 | 297.1807 | 333.579 | 211.0452 | 169.9548 | 363.1843 | 424.4851 | 217.5375 | 242.5601 | 406.8835 | 334.2294 | 284.5339 | 235.3008 | 378.3496 | 311.0903 | 606.4173 | 725.9061 | 313.5565 | 436.5565 | 441.8383 | 413.1617 | 597.0488 | 199.5113 | 207.9738 | 191.1391 | 231.6052 | 301.5565 | 244.5375 | 199.1165 | 135.9061 | 499.2782 | 470.8121 | 441.0226 | 539.1617 | 272.7669 | 256.0036 | 218.6504 | 209.7931 | 104.4435 | 96.4661 | 115.9774 | 189.3948 | 507.0714 | 468.8383 | 519.327 | 402.4661 | 258.6766 | 237.6314 | 218.6278 | 218.8835 | 334.1391 | 284.9774 | 407.692 | 371.7182 | 299.2593 | 271.3948 | 416.4887 | 346.9323 | 319.6766 | 348.9097 | 419.3722 | 520.8419 | 282.4661 | 275.4923 | 440.7218 | 453.252 | 381.3948 | 356.1201 | 303.9286 | 554.7028 | 338.4851 | 336.5375 | 272.5827 | 489.5528 | 365.9286 | 306.8609 | 262.6992 | 356.7633 | 429.9323 | 453.654 | 347.3532 | 363.7669 | 236.3234 | 238.4923 | 176.9061 | 272.9774 | 619.3496 | 600.075 | 561.4887 | 1026.9738 | 599.9323 | 550.9548 | 819.1581 | 812.1653 | 232.6992 | 127.2556 | 108.9323 | 232.6992 | 209.7669 | 187.7444 | 202.9061 | 220.8609 | 242.5339 | 241.1427 | 197.4887 | 237.7931 | 129.1165 | 116.5827 | 99.0488 | 154.7931 | 217.7931 | 222.9061 | 183.3758 | 120.9548 | 512.2105 | 483.7145 | 418.8347 | 458.9286 | 386.8835 | 522.654 | 447.6052 | 512.3722 | 828.6052 | 447.0262 | 541.4435 | 997
2229 | 26.5134 | 283.3845 | 330.2272 | 472.536 | 342.697 | 271.7651 | 269.7613 | 347.9981 | 271.9943 | 317.5379 | 207.1515 | 169.6194 | 349.6155 | 406.536 | 239.6913 | 240.8523 | 410.767 | 321.7651 | 274.6932 | 227.8447 | 392.4583 | 297.1515 | 625.0662 | 715.9243 | 319.4545 | 448.9904 | 452.6913 | 400.3864 | 617.8447 | 193.9243 | 184.8466 | 190.6913 | 245.2272 | 284.0019 | 251.3126 | 204.8447 | 120.9262 | 501.4583 | 453.9243 | 440.2291 | 532.1534 | 257.0815 | 266.4679 | 215.3068 | 224.5398 | 114.6932 | 99.3864 | 119.8466 | 190.3826 | 530.1496 | 454.3164 | 532.6913 | 399.2349 | 277.3068 | 252.466 | 218.3845 | 208.5417 | 349.4545 | 288.0777 | 384.2234 | 368.6855 | 307.6989 | 279.6874 | 430.9185 | 362 | 338.3068 | 368.6155 | 425.3826 | 548.3106 | 294.9981 | 300.3864 | 434.697 | 438.0719 | 381.9981 | 373.8504 | 295.0757 | 596.0757 | 315.1534 | 348.6951 | 269.697 | 483.0643 | 359.3826 | 319.9962 | 270.5398 | 318.6951 | 453.0738 | 472.3087 | 371.9262 | 365.767 | 235.3787 | 255.3126 | 168.8466 | 276.4621 | 630.767 | 636.6932 | 552.8504 | 1011.1515 | 623.8428 | 552.9262 | 794.3806 | 840.2272 | 258.2253 | 127.5379 | 127.8447 | 258.2253 | 222.1477 | 188.6155 | 186.3883 | 224.3845 | 261.9128 | 266.9223 | 203.4602 | 239.4679 | 137.1515 | 127.5379 | 117.5379 | 175.3068 | 231.3864 | 226.7651 | 203.6951 | 127.5398 | 513.6951 | 444.8428 | 401.1534 | 443.5398 | 389.9981 | 561.6855 | 461.9962 | 514.9223 | 831.0777 | 493.9904 | 533.2387 | 1015.4545
2230 | 26.5244 | 295.506 | 356.9028 | 478.2389 | 326.1741 | 295.7449 | 314.506 | 352.0324 | 301.1579 | 337.4574 | 212.6356 | 154.9838 | 358.5222 | 417.3968 | 234.081 | 236.8218 | 405.3199 | 346.1417 | 284.0931 | 246.6963 | 390.1296 | 306.6032 | 648.1782 | 722.506 | 327.9718 | 478.2875 | 455.413 | 392.2065 | 633.919 | 196.1417 | 179.8259 | 178.7329 | 247.3037 | 276.7935 | 244.0931 | 208.6194 | 131.8056 | 504.6843 | 468.0444 | 445.413 | 544.3157 | 253.0768 | 253.8866 | 219.3968 | 236.8542 | 115.7611 | 93.9838 | 107.6843 | 191.8583 | 530.5426 | 450.1375 | 549.2995 | 385.9676 | 287.7449 | 237.5708 | 224.5546 | 202.5384 | 364.7167 | 291.7449 | 408.2551 | 392.6356 | 289.5222 | 284.5264 | 470.6315 | 358.4292 | 354.2995 | 356.494 | 426.8583 | 535.2065 | 309.7449 | 296.5708 | 439.5991 | 461.3806 | 390 | 358.7773 | 297.2227 | 597.8259 | 339.6153 | 342.9514 | 255.1579 | 509.8139 | 375.9352 | 366.1537 | 258.6519 | 328.8218 | 473.5384 | 470.3806 | 371.6963 | 370.2389 | 254.4454 | 244.1255 | 173.7449 | 277.5708 | 638.8097 | 623.8745 | 525.081 | 1011.8745 | 634.8259 | 533.6519 | 809.0162 | 844.6843 | 274.2551 | 129.6032 | 120.1134 | 274.2551 | 230.1458 | 207.4412 | 178.1417 | 220.6356 | 281.988 | 254.1296 | 201.0648 | 233.6315 | 145.8097 | 137.9352 | 114.0486 | 176.6194 | 233.919 | 226.5426 | 192 | 119.2227 | 513.506 | 453.8421 | 405.7773 | 435.6194 | 405.5384 | 581.2713 | 485.9352 | 506.0972 | 839.1093 | 520.5102 | 501.587 | 1061.6639
2231 | 26.5354 | 314.2272 | 363.1375 | 460.7737 | 349.8163 | 299.8642 | 324.228 | 352.1358 | 300.0009 | 366.4543 | 218.6806 | 163.3177 | 364.0009 | 437.6344 | 236.1801 | 265.0914 | 414.4969 | 359.0453 | 304.8181 | 265.5909 | 366.8181 | 313.3629 | 627.6806 | 720.3656 | 296.8181 | 457.5004 | 455.8634 | 398.772 | 625.2751 | 201.6823 | 177.4543 | 155.6362 | 232.4996 | 279.6362 | 265.2272 | 227.9068 | 152.8651 | 508.2254 | 492.0914 | 429.5013 | 564.4996 | 273.1366 | 282.772 | 213.4552 | 262.7728 | 134.4534 | 98.2272 | 98.9539 | 195.3168 | 534.2697 | 486.3194 | 566.7728 | 403.4534 | 286.1375 | 243 | 229.7276 | 211.5457 | 362.2254 | 297.9095 | 418.0435 | 390.0905 | 306.0453 | 270.4082 | 499.1366 | 362.4082 | 350.9113 | 358.2254 | 427.8625 | 543.8172 | 322.4543 | 303.6362 | 478.6362 | 455.1828 | 384.2733 | 378.953 | 291.3638 | 600.3629 | 371.1366 | 354.4543 | 285.4978 | 485.1349 | 394.9539 | 391.1384 | 253.9086 | 365.2724 | 505.8616 | 497.7249 | 409.4073 | 380.1801 | 254.5901 | 265.3629 | 184 | 280.5457 | 645.181 | 610 | 523.0896 | 1008.6353 | 614.8651 | 529.3177 | 778.5936 | 825.3185 | 277.0896 | 135.5448 | 111.6806 | 277.0896 | 206.2724 | 225.9556 | 191.0453 | 224.2263 | 256.4082 | 240.6353 | 189.8181 | 263.7728 | 135.819 | 145.5004 | 109.0905 | 179.5448 | 244.0914 | 205.3177 | 197.3177 | 124.8172 | 513.4108 | 434.5466 | 415.3177 | 429.1366 | 427.2263 | 571.2724 | 516.4073 | 520.315 | 863.3629 | 485.6832 | 521.9982 | 1078.819
2232 | 26.5464 | 302.0957 | 363.275 | 451.8988 | 384.1628 | 299.0055 | 307.5226 | 337.4049 | 291.6348 | 345.9835 | 232.7866 | 175.2134 | 364.0616 | 446.197 | 252.5732 | 254.2079 | 445.2531 | 361.1573 | 297.5732 | 267.528 | 369.264 | 319.7866 | 647.264 | 671.3543 | 291.6799 | 453.528 | 444.4104 | 412.1067 | 592.6854 | 190.2079 | 192.7976 | 157.5787 | 238.1573 | 297.1683 | 262.7866 | 252.5732 | 133.1067 | 547.3762 | 496.3762 | 418.4268 | 566.7866 | 271.528 | 297.3707 | 203.1573 | 253.8372 | 143.8372 | 99.1573 | 110.6854 | 207.9945 | 572.9439 | 499.1793 | 592.8085 | 423.3707 | 253.8317 | 258.1683 | 235.1122 | 216.4268 | 383.2585 | 308.4323 | 442.3652 | 405.6908 | 326.6963 | 289.4774 | 515.6457 | 385.2695 | 305.5171 | 402.8537 | 439.1012 | 554.2079 | 336.9549 | 311.4774 | 475.9439 | 451.5335 | 376.8427 | 381.3543 | 282.2585 | 608.472 | 356.0451 | 355.472 | 318.4213 | 500.2585 | 399.5226 | 386.7018 | 281.2244 | 368.736 | 548.6402 | 540.4268 | 433.0506 | 407.1067 | 273.7415 | 277.6854 | 179.3652 | 292.1683 | 659.528 | 630.6457 | 533.5732 | 1033.6908 | 612.803 | 555.539 | 744.4774 | 840.6457 | 298.528 | 147.9494 | 121.9945 | 298.528 | 212.264 | 206.2079 | 180.9384 | 241.9494 | 270.4213 | 277.4884 | 191.8427 | 250.2024 | 123.1573 | 135.2079 | 116.6854 | 180.1518 | 245.4268 | 235.3311 | 204.5787 | 139 | 517.0725 | 436.1177 | 415.8372 | 431.3201 | 447.0561 | 584.0055 | 542.1573 | 553.5171 | 890.011 | 485.6457 | 527.3543 | 1081.7469
2233 | 26.5574 | 310.7525 | 364.8847 | 449.5928 | 377.522 | 304.7964 | 297.8318 | 321.478 | 311.8762 | 339.4516 | 230.0354 | 179.2655 | 387.2036 | 438.5135 | 243.6018 | 258.9735 | 418.4955 | 373.009 | 278.6282 | 252.3538 | 385.7435 | 319.2036 | 636.8583 | 668.1507 | 294.7081 | 450.4955 | 432.9735 | 417.6018 | 623.1861 | 195.4072 | 209.4336 | 158.7345 | 238.7345 | 311.0619 | 253.5309 | 226.6901 | 138.6018 | 567.77 | 502.4601 | 435.2036 | 568.3718 | 286.2744 | 292.0265 | 199.8318 | 235.761 | 122.7256 | 102.77 | 112.1327 | 189.1946 | 616.735 | 540.6991 | 634.4691 | 437.9735 | 246.8762 | 259.9202 | 223.4157 | 231.0354 | 389.1771 | 319.2919 | 455.0798 | 417.6991 | 354.9027 | 288.3274 | 557.8762 | 402.3363 | 319.5224 | 405.6193 | 464.3897 | 560.2744 | 338.991 | 308.1592 | 480.2744 | 478.5753 | 378.4336 | 367.7435 | 284.2744 | 615.8408 | 352.7081 | 356.3363 | 312.9291 | 515.2834 | 420.6197 | 413.761 | 321.1417 | 377.4072 | 556.4601 | 541.1592 | 444.1771 | 428.2126 | 277.7964 | 271.761 | 178.6372 | 320.8054 | 677.7435 | 689.7879 | 543.248 | 1068.248 | 653.8762 | 588.3718 | 751.1327 | 878.1063 | 305.0354 | 142.0265 | 116.6372 | 305.0354 | 217.0354 | 206.2036 | 167.3363 | 240.7964 | 277.0709 | 295.7256 | 199.9381 | 248.3099 | 131.1063 | 133.4691 | 115.9646 | 179.7081 | 250.0619 | 265.1327 | 204.8673 | 134.23 | 553.9646 | 472.7081 | 399.0619 | 439.5664 | 441.8583 | 614.0798 | 546.2036 | 553.2126 | 921.5399 | 503.1592 | 505.5045 | 1121.5135
2234 | 26.5684 | 317.1231 | 351.6073 | 448.0399 | 351.9867 | 287.2978 | 303.9185 | 352.3644 | 305.9201 | 363.7438 | 230.4326 | 186.8103 | 385.5275 | 445.7171 | 243.8785 | 265.3111 | 416.2429 | 390 | 292.188 | 226.6622 | 412.5807 | 329.5674 | 633.4859 | 685.3927 | 299.2978 | 470.5391 | 470.9717 | 441.0666 | 664.0133 | 204.9867 | 199.2978 | 167.2429 | 244.1215 | 306.9052 | 254.797 | 211.5674 | 144.2296 | 556.609 | 512.7421 | 442 | 601.188 | 306.797 | 263.8386 | 203.6889 | 233.9185 | 115.5807 | 112.1215 | 108.3245 | 181.2296 | 667.8386 | 544.3511 | 635.8785 | 442.0815 | 267.6755 | 239.9052 | 209.1747 | 236.3378 | 400.0815 | 323.6206 | 476.8785 | 428.0399 | 384.5258 | 284.5807 | 576.8918 | 450.3627 | 345.5008 | 399.9584 | 488.8253 | 580.3511 | 324.2296 | 310.4992 | 502.5807 | 495.8918 | 400.4043 | 387.8918 | 290.9734 | 624.6489 | 367.5541 | 348.6223 | 305.8918 | 550.1215 | 451.7571 | 419.4992 | 318.2712 | 366.4742 | 585.9168 | 560.4443 | 446.1631 | 447.5408 | 280.8103 | 281.5125 | 194.0266 | 349.7288 | 699.6755 | 702.7987 | 561.1897 | 1089.3111 | 677.797 | 587.7438 | 737.5141 | 901.1481 | 313.0133 | 141.9318 | 127.5674 | 313.0133 | 225.4592 | 213.4459 | 177.9052 | 225.9734 | 307.2962 | 292.594 | 213.9052 | 279.9318 | 138.0948 | 138.8918 | 111.1082 | 188.7571 | 273.5524 | 280.4992 | 204.6622 | 125.7704 | 565.1614 | 455.8935 | 407.391 | 437.2163 | 445.1747 | 651.0399 | 553 | 585.9185 | 950.0532 | 491.2296 | 545.6606 | 1140.3111
2235 | 26.5794 | 289.1643 | 371.9164 | 446.9625 | 347 | 281.3314 | 310.4179 | 337.4236 | 291.1211 | 358.7896 | 236.0807 | 198.0836 | 396.5361 | 432.1643 | 256.9539 | 247.8357 | 430.8732 | 378.5446 | 330.1643 | 239.2046 | 433.4554 | 347.7464 | 655.8732 | 666.9164 | 299.2882 | 487.6714 | 500.1671 | 456.3775 | 651.5879 | 210.0807 | 207.9943 | 173.6254 | 238.3775 | 301.5446 | 263.8329 | 209.585 | 133.2536 | 518.9625 | 544.6657 | 444.2911 | 610.755 | 307.585 | 259.6629 | 223.4554 | 252.3314 | 122.2507 | 117.8329 | 108.6657 | 200.4957 | 632.9625 | 577.1182 | 622.8357 | 452.9943 | 272.8329 | 256.9971 | 229.7925 | 234.7925 | 400.4554 | 329.6714 | 462.0029 | 446.6657 | 407.6686 | 310.0375 | 599.5361 | 514.9136 | 334.6225 | 420.9596 | 470.0404 | 576.9193 | 347.1614 | 313.1296 | 507.4179 | 499.7493 | 410.9625 | 383.9596 | 294.2882 | 605.6282 | 368.8329 | 355.3689 | 295.0864 | 545.2939 | 477.9079 | 444.1239 | 300.9568 | 371.2825 | 623.3371 | 603.8329 | 442.3286 | 441.0836 | 294.8329 | 293.7118 | 201.2507 | 350.5475 | 690.6282 | 675.8732 | 549.9164 | 1080.0836 | 698.2882 | 604.7839 | 698.1728 | 911.7925 | 315.2507 | 153.8761 | 122.8357 | 315.2507 | 236.9568 | 220.415 | 204.6197 | 231.5793 | 328.585 | 299.6686 | 215.4179 | 290.0432 | 134.7493 | 150.0807 | 113.2075 | 186.4986 | 310.3343 | 295.0432 | 200.2507 | 140.4121 | 575.2939 | 440.7436 | 414.7954 | 425.8357 | 504.7407 | 667.3746 | 555.7493 | 607.0807 | 992.9107 | 502.7061 | 558.8818 | 1146.6629
2236 | 26.5904 | 295.2949 | 367.8244 | 413.4143 | 348.882 | 308.3497 | 313.9396 | 303.177 | 305.236 | 371.764 | 246.5885 | 188.0604 | 430.4691 | 435.0014 | 267.5899 | 235.5885 | 420.1798 | 365.941 | 318.8862 | 256.941 | 427.7093 | 345.1798 | 654.1194 | 689.1727 | 297.1208 | 487.1727 | 498.941 | 449.7051 | 629.5885 | 214.4719 | 224.1194 | 168.236 | 242.0562 | 298.8216 | 258.0014 | 212.8216 | 130.1742 | 527.2893 | 533.9438 | 437.1194 | 595.5281 | 309.4101 | 286.8806 | 229.0014 | 254.8834 | 116.177 | 112.9424 | 118.6461 | 212.0014 | 625.2921 | 601.8834 | 636.9368 | 486.0576 | 282.9986 | 256.3553 | 228.3525 | 242.7626 | 413.059 | 328.7022 | 451.823 | 466.0562 | 404.6461 | 323.3539 | 665.934 | 528.059 | 355.4101 | 429.1152 | 483.0576 | 603.5815 | 363.9424 | 302.8216 | 507.1756 | 528.8778 | 386.3539 | 372.8244 | 302.9424 | 608.8792 | 385.1152 | 376.4129 | 299.3497 | 528.4129 | 549.5239 | 465.9382 | 306.5885 | 422.8806 | 603.177 | 609.764 | 469.6461 | 462.1138 | 316.7612 | 299.6433 | 194.7065 | 360.6994 | 704.7008 | 710.8188 | 559.9396 | 1070.5309 | 723.4101 | 615.5337 | 682.2331 | 918.3511 | 318.1166 | 155.5871 | 124.7036 | 318.1166 | 237.4129 | 214.8258 | 217.2374 | 260.3511 | 333.2331 | 312.6433 | 211.4115 | 292.4101 | 145.9986 | 149.7669 | 112.2949 | 194.4115 | 314.9396 | 316.2303 | 194.6475 | 157.7065 | 545.2388 | 473.764 | 412.5857 | 425.8216 | 520.1812 | 698.1713 | 578.7612 | 613.3539 | 996.9452 | 539.8188 | 512.9438 | 1184.2317
2237 | 26.6014 | 287.1032 | 363.9312 | 410.0355 | 345.6892 | 321.7581 | 340.4151 | 318.3463 | 313.0699 | 377.9312 | 237.2753 | 181.3796 | 428.4129 | 431.3108 | 248.3097 | 228.2064 | 421.2086 | 376.1731 | 308.0011 | 258 | 418.0699 | 329.7936 | 650.2419 | 706.2408 | 291.4839 | 525.9667 | 492.7581 | 456.4484 | 643.4495 | 201.172 | 226.0355 | 164 | 246.9301 | 312.9312 | 261.0699 | 227.414 | 148.3796 | 546.8957 | 520.8624 | 428.414 | 606.414 | 323.3796 | 295.586 | 217.1032 | 243.6548 | 120.7247 | 109.828 | 117.7581 | 203.9656 | 646.6903 | 591.1376 | 659.2064 | 502.6903 | 319.0366 | 268.5194 | 221.8269 | 243.4473 | 417.3108 | 353.4484 | 462.2419 | 491.2075 | 417.7592 | 332.4495 | 700.8957 | 565.6237 | 352.9645 | 431.4817 | 499.2075 | 622.8602 | 353.586 | 326.5871 | 517.828 | 533.3774 | 391.1043 | 372.7936 | 289.6892 | 624.172 | 415.9667 | 398.9334 | 301.6537 | 555.7614 | 585.9656 | 505.7603 | 331.071 | 426.2753 | 617.3807 | 618.828 | 489.0355 | 501.4839 | 324.9989 | 344.2775 | 194.8624 | 386.0677 | 722.0333 | 738.586 | 580.6215 | 1083.2086 | 734.4828 | 599.1043 | 695.5172 | 944.4839 | 372.3141 | 164.4484 | 135.6204 | 372.3141 | 224.6892 | 229.4506 | 194.3441 | 278.2764 | 364.3807 | 354.3807 | 214.4484 | 306.8624 | 139.5505 | 146.6215 | 121.9667 | 209.5183 | 335.1387 | 351.3796 | 192.4828 | 153.0344 | 547.6226 | 478 | 435.7592 | 434.6204 | 507.3807 | 758.0366 | 582.1709 | 616.1731 | 1003.3474 | 544.4118 | 520.6226 | 1203.1376
2238 | 26.6124 | 291.9214 | 387.2888 | 440.3379 | 352.8723 | 331.8232 | 346.1179 | 330.5737 | 319.0688 | 377.0295 | 243.8526 | 205.2986 | 414.5147 | 449.3772 | 244.446 | 253.6856 | 436.0884 | 408.2789 | 329.9312 | 265.4263 | 420.6032 | 330.554 | 638.1965 | 684.1474 | 317.3968 | 513.2849 | 476.5835 | 453.5442 | 676.3379 | 200.9509 | 219.2063 | 170.9312 | 226.5737 | 320.446 | 269.5442 | 226.5344 | 153.9804 | 539.9509 | 539.3575 | 429.5147 | 608.0098 | 301.7505 | 280.6228 | 228.3575 | 240.9804 | 132.4656 | 112.0295 | 124.8526 | 216.387 | 686.2003 | 608.7839 | 667.8526 | 521.9018 | 339.5835 | 291.5639 | 227.3477 | 244.8526 | 424.9804 | 350.0491 | 468.5147 | 498.5147 | 438.8625 | 350.9804 | 724.6463 | 620.3674 | 336.5246 | 423.9312 | 529.2888 | 623.7347 | 357.4361 | 339 | 520.0295 | 535.7151 | 420.3281 | 381.9705 | 294.3968 | 628.9018 | 430.5147 | 434.9607 | 292.3968 | 568.7407 | 593.4361 | 520.0786 | 359.4951 | 443.2495 | 659.2691 | 642.3182 | 495.0786 | 529.3772 | 329.3084 | 360.0589 | 210.8821 | 386.3281 | 738.1807 | 730.554 | 593.4853 | 1108.9804 | 716.6818 | 612.4853 | 665.7898 | 972.8723 | 406.167 | 171.446 | 127.5442 | 406.167 | 233.8526 | 247.6326 | 187.4656 | 285.9902 | 396.3674 | 370.0295 | 216.9902 | 340.2102 | 143.7937 | 164.9312 | 123.6425 | 201.7211 | 369.7642 | 375.7937 | 217.7544 | 142.0982 | 584.3968 | 478 | 448.446 | 445.3575 | 532.4361 | 790.9607 | 599.1316 | 637.387 | 1027.6032 | 525.9116 | 556.9018 | 1219.7937
2239 | 26.6234 | 292.3893 | 398.8524 | 445.8377 | 384.2656 | 336.8082 | 323.8524 | 336.6992 | 335.3394 | 400.8911 | 236.168 | 218.3893 | 419.0885 | 432.0647 | 243.4041 | 282.4483 | 423.4483 | 407.1532 | 323.8082 | 275.0295 | 428.2656 | 319.9262 | 619.5221 | 696.3984 | 318.8672 | 508.8616 | 474.5959 | 434.7934 | 704.1623 | 217.6697 | 192.4188 | 179.0148 | 224.0738 | 343.2509 | 275.655 | 236.6992 | 149.9115 | 546.5221 | 563.6402 | 440.6845 | 605.4779 | 299.3394 | 275.1033 | 237.9262 | 270.3984 | 125.345 | 113.5664 | 122.2417 | 229.5221 | 721.059 | 652.3099 | 689.5959 | 557.3689 | 345.2066 | 294.1328 | 235.3893 | 270.1476 | 445.773 | 351.0885 | 488.3099 | 528.4575 | 451.4779 | 356.0443 | 769.6254 | 623.345 | 341.6107 | 427.9557 | 516.4982 | 643.3893 | 357.9115 | 336.9705 | 535.773 | 576.1623 | 439.5221 | 384.4926 | 311.0885 | 636.4631 | 436.6107 | 441.5369 | 312.1328 | 565.8173 | 603.0443 | 548.5313 | 371.6697 | 460.3746 | 710.3984 | 654.4041 | 507.2951 | 537.4336 | 356.6107 | 352.4779 | 212.8524 | 414.1476 | 770.4926 | 727.0295 | 592.4631 | 1111 | 712.714 | 644.4427 | 627.3746 | 1000.714 | 391.5369 | 203.3837 | 127.5664 | 391.5369 | 270.3099 | 228.9262 | 182.3746 | 280.9115 | 396.8082 | 392.8763 | 231.6992 | 353.7639 | 164.4926 | 173.0148 | 120.6697 | 191.1771 | 393.4926 | 393.9557 | 236.9115 | 140.6254 | 551.8728 | 468.3598 | 440.3008 | 466.0885 | 557.2656 | 831.0239 | 643.0885 | 662.1918 | 1046.428 | 580.6641 | 547.7196 | 1258.2509
2240 | 26.6344 | 299.0311 | 400.4327 | 448.5114 | 398.882 | 329.1967 | 337.9048 | 341.6853 | 341.3706 | 408.8116 | 239.4638 | 213.0787 | 422.4017 | 409.7557 | 252.5901 | 270.6853 | 421.0393 | 394.9131 | 331.2671 | 289.4721 | 421.0559 | 313.441 | 644.9048 | 696.4968 | 321.4327 | 500.2608 | 499.1884 | 435.5508 | 711.882 | 222.764 | 201.0311 | 168.5673 | 224.323 | 342.8509 | 297.9524 | 240.1263 | 154.3934 | 563.0704 | 543.8985 | 471.7081 | 585.2919 | 327.7164 | 293.9524 | 230.4017 | 275.6935 | 126.9524 | 116.9607 | 105 | 228.4017 | 707.8509 | 662.0869 | 719.9048 | 555.4968 | 318.9854 | 280.1345 | 239.4327 | 290.9131 | 454.1263 | 363.236 | 502.2443 | 546.5279 | 455.1967 | 384.4638 | 835.1637 | 629.1098 | 341.7246 | 434.8737 | 527.1719 | 630.8033 | 352.5197 | 360.9835 | 532.1739 | 568.8116 | 419.6935 | 390.7557 | 310.764 | 655.8261 | 435.6853 | 431.5279 | 331.9131 | 585.9607 | 606 | 598.0311 | 374.6853 | 457.7164 | 727.1263 | 676.0622 | 525.9607 | 526.2443 | 363.4803 | 361.9131 | 217.5508 | 433.8737 | 757.0083 | 720.6853 | 593.6377 | 1114.6377 | 769.0932 | 688.5508 | 619.5197 | 1018.118 | 413.7474 | 224.323 | 130.9607 | 413.7474 | 289.441 | 227.6377 | 198.4245 | 299.9441 | 401.6688 | 403.9689 | 240.323 | 344.7557 | 169.1967 | 175.559 | 127.8426 | 195.7246 | 411.7081 | 394.6377 | 229.441 | 163.0311 | 533.9048 | 463.1574 | 421.2836 | 479.2754 | 587.9917 | 841.5755 | 649 | 662.0083 | 1093.7081 | 621.3623 | 554.9835 | 1253.1739
2241 | 26.6454 | 307.2124 | 397.701 | 446.1051 | 367.2741 | 318.5731 | 355.7443 | 332.9361 | 328.1278 | 404.2351 | 247.5731 | 203.8288 | 419.468 | 416 | 263.8082 | 263.532 | 427.3196 | 397.5525 | 364.3629 | 304.7237 | 415.0227 | 306.6804 | 680.299 | 664.0206 | 327.2124 | 496.0661 | 522.3835 | 443.1484 | 688.7855 | 238.8104 | 209.2124 | 170.2351 | 221.5959 | 335.5114 | 291.4453 | 228.3402 | 148.5731 | 575.4041 | 535.6186 | 467.7216 | 579.7031 | 328.2124 | 301.2763 | 233.3196 | 277.8104 | 115.1257 | 112.2763 | 110.8516 | 228.1278 | 718.0661 | 658.3835 | 768.0661 | 555.4702 | 289.9794 | 287.1939 | 234.5731 | 282.9133 | 466.2784 | 379.1073 | 510.8949 | 541.9155 | 490.3218 | 431.4063 | 886.9361 | 639.2124 | 358.9382 | 461.0227 | 557.8927 | 669.6208 | 347.1484 | 387.1278 | 523.7031 | 572.2145 | 428.7259 | 400.7237 | 316.7031 | 663.7649 | 433.3196 | 414.1484 | 332.9567 | 587.6598 | 633.13 | 609.4041 | 386.6825 | 463 | 721.1918 | 708.9155 | 530.8516 | 541.2784 | 378.4269 | 370.4041 | 248.5547 | 416.9339 | 745 | 772.5796 | 596.468 | 1128.6392 | 803.2124 | 712.6392 | 620.532 | 1026.3196 | 449.0227 | 208.2969 | 127.8722 | 449.0227 | 299.1712 | 251.7465 | 204.5092 | 308.6371 | 445.662 | 409.6186 | 252.4908 | 377.598 | 171.8722 | 180.7237 | 129.3196 | 196.4475 | 434.8516 | 420.8743 | 239.7031 | 178.6598 | 563.4475 | 485.6186 | 422.9155 | 461.5298 | 596.8082 | 800.9133 | 654.8516 | 628.7216 | 1111.532 | 629.7031 | 561.9773 | 1228.2124
2242 | 26.6564 | 334.5666 | 409.9471 | 455.3143 | 331.6592 | 319.1505 | 356.354 | 360.3011 | 338.708 | 411.7345 | 230.7345 | 198.8983 | 428.7965 | 414.3672 | 275.1505 | 236.2432 | 452.1373 | 409.5046 | 365.0265 | 295.4822 | 420.7477 | 323.7701 | 692.5443 | 649.3805 | 309.3937 | 514.646 | 537.7965 | 443.4425 | 678.6195 | 237.3184 | 223.5046 | 173.3805 | 234.4293 | 328.8495 | 285.7965 | 229.2655 | 142.0753 | 574 | 543.8362 | 457.708 | 611.2258 | 334.3408 | 308.3408 | 230.3805 | 289.9247 | 109.1505 | 113.8983 | 120.8983 | 204.4202 | 739.3672 | 686.3143 | 786.646 | 564.4822 | 306.2523 | 314.7213 | 234.6063 | 261.5575 | 495.885 | 396.708 | 514.7477 | 542.0132 | 503.938 | 424.7874 | 936.6154 | 656.2258 | 370.2035 | 474.9115 | 563.1505 | 693.3052 | 351.2523 | 383.0132 | 516.0397 | 575.3052 | 458.8983 | 419.7833 | 315.5707 | 663.3408 | 456.5046 | 432.9471 | 310.8495 | 571.4954 | 648.8362 | 635.2126 | 429.5798 | 490.7568 | 733.7965 | 696.4954 | 532.7345 | 537.6857 | 393.0885 | 360.8362 | 277.6063 | 420.1241 | 712.8891 | 761.3225 | 610.6948 | 1104.2564 | 810.4293 | 734.5178 | 637.3276 | 1060.9339 | 460.1902 | 200.177 | 137.9735 | 460.1902 | 311.7213 | 267.823 | 211.2258 | 292.1902 | 468.3672 | 432.531 | 256.6592 | 409.1638 | 177.6195 | 183.4557 | 124.2035 | 178.5087 | 449.7965 | 414.8759 | 255.4425 | 179.3672 | 560.7477 | 488.938 | 446.9603 | 449.2523 | 587.469 | 807.8586 | 715.5137 | 612.177 | 1102.2035 | 637.2787 | 562.7833 | 1222.3672
2243 | 26.6674 | 363.7827 | 425.5477 | 482.9047 | 339.1397 | 340.4612 | 344.9733 | 377.3215 | 328.1906 | 419.5743 | 248.5921 | 208.5654 | 430.878 | 419.122 | 290.8958 | 238.0444 | 455.643 | 399.399 | 354.4435 | 287.7827 | 401.3126 | 329.7561 | 685.2084 | 645.7827 | 303.5654 | 524.357 | 533.1995 | 458.0178 | 715.0533 | 213.3215 | 225.643 | 172.5654 | 254.0178 | 300.8603 | 283.4257 | 238.122 | 139.0953 | 597.3748 | 552.5832 | 474.4612 | 614.5299 | 352.4612 | 331.4701 | 232.9047 | 271.7561 | 124.8958 | 116.3304 | 113.8691 | 204.7052 | 689.0242 | 716.6874 | 772.9822 | 554.5565 | 316.2262 | 322.0089 | 226.521 | 257.5565 | 456.5808 | 419.0267 | 533.1575 | 534.2173 | 486.3215 | 436.5057 | 1007.9402 | 664.5388 | 387.601 | 479.0089 | 531.0331 | 672.0953 | 368.4612 | 362.4168 | 538.8449 | 559.1042 | 434.6163 | 415.1906 | 298.7649 | 660.8691 | 456.9733 | 424.6163 | 319.5921 | 593.2884 | 599.1462 | 650.2084 | 453.2262 | 488.9556 | 711.3901 | 677.1042 | 507.1817 | 540.7052 | 386.7649 | 397.4103 | 270.6341 | 460.3659 | 689.8958 | 735.2706 | 647.4879 | 1096.5921 | 840.5921 | 749.4523 | 655.4523 | 1052.6074 | 452.5477 | 208.6785 | 161.9136 | 452.5477 | 298.9733 | 278.8005 | 226.2173 | 297.9047 | 440.2859 | 435.7738 | 254.122 | 391.5121 | 181.7738 | 185.2262 | 134.9225 | 181.9313 | 445.7561 | 399.122 | 273.9136 | 175.2173 | 566.357 | 490.8005 | 440.2948 | 453.6607 | 561.4079 | 852.601 | 737.5121 | 638.4879 | 1101.2351 | 630.5477 | 595.479 | 1221.5565
2244 | 26.6784 | 363.7247 | 451.0105 | 472.3661 | 355.413 | 344.587 | 336.9818 | 368.587 | 338.4599 | 417.1922 | 248.5038 | 220.9637 | 406.0908 | 422.8623 | 288.3117 | 246.6233 | 440.4493 | 375.0182 | 359.6883 | 295.1195 | 403.3767 | 309.1922 | 663.6415 | 664.4962 | 311.413 | 550.3585 | 538.3585 | 454.6233 | 721.935 | 209.7065 | 229.2208 | 169.6052 | 258.587 | 295.3948 | 273.2935 | 232.1922 | 144.5143 | 583.0076 | 566.7065 | 468.9168 | 637.6989 | 376.4962 | 339.0182 | 245.1195 | 267.6883 | 135.3948 | 117.2753 | 99.3117 | 228.2572 | 695.5067 | 736.239 | 748.5038 | 556.7065 | 321.9818 | 311.2103 | 225.2208 | 245.4857 | 461.8002 | 437.2572 | 563.6702 | 536.5507 | 495.7897 | 460.587 | 1027.0182 | 627.9895 | 392.0726 | 487.5507 | 528.9924 | 637.1271 | 355.522 | 348.8623 | 565.587 | 546.3298 | 409.7247 | 398.7247 | 305.5143 | 672.478 | 438.174 | 404.2753 | 339.9818 | 604.7973 | 566.1377 | 611.5765 | 455 | 458.1922 | 674.4857 | 700.1664 | 514.5975 | 529.5583 | 383.2753 | 407.5401 | 249.174 | 436.3193 | 723.1482 | 756.4312 | 674.9637 | 1152.2495 | 866.6702 | 763.239 | 685.1664 | 1053.8729 | 440.4675 | 227.0832 | 158.2103 | 440.4675 | 306.3404 | 292.9818 | 219.4493 | 304.4312 | 428.1013 | 429.4493 | 249.8987 | 366.3117 | 199.3404 | 176.761 | 144.3117 | 205.5507 | 412.6778 | 428.4599 | 295.239 | 176.9818 | 610.5612 | 458.9063 | 426.413 | 460.6702 | 547.2753 | 828.631 | 714.587 | 627.283 | 1080.8336 | 657.1482 | 610.4312 | 1208.348
2245 | 26.6894 | 374.7848 | 442.2004 | 456.0944 | 341.7285 | 333.8642 | 348.7169 | 366.5811 | 357.5811 | 397.9586 | 270.8675 | 217.1209 | 392.1623 | 424.9056 | 288.0679 | 243.2301 | 434.0944 | 388.8526 | 369.8113 | 311.7848 | 409.0944 | 324.245 | 688.6109 | 636.3212 | 313.4189 | 542.8229 | 565.6904 | 444.8377 | 692.7285 | 216.2301 | 242.9056 | 176.2036 | 243.2152 | 307.9735 | 280.7169 | 221.6755 | 144.702 | 560.9056 | 563.351 | 463.0679 | 651.7964 | 379.702 | 336.5811 | 249 | 273.1623 | 122.9851 | 115.5132 | 105.4603 | 255.0828 | 711.404 | 737.6076 | 774.9354 | 532.4305 | 325 | 300.5811 | 248.7848 | 251.6904 | 460.1325 | 428.053 | 563.4454 | 513.8494 | 483.1739 | 453.351 | 1043.7583 | 604.2566 | 397.4338 | 484.6076 | 528.4305 | 649.1241 | 352.2036 | 387.5165 | 572.298 | 560.5016 | 397.5397 | 396.4189 | 311.5132 | 675.7964 | 416.1474 | 421.6904 | 351.1358 | 581.1209 | 590.8261 | 601.5281 | 460.2301 | 430.8229 | 636.5249 | 684.5514 | 544.4603 | 506.8377 | 354.7815 | 376.7964 | 263.7583 | 419.7848 | 713.457 | 749.0265 | 612.4272 | 1144.808 | 843.2004 | 754.1474 | 680.0116 | 1112.7731 | 441.5546 | 234.6755 | 149.9056 | 441.5546 | 312.8642 | 297.1623 | 218.3245 | 315.6225 | 442.5546 | 447.5016 | 255.7848 | 391.0563 | 206.4454 | 172.4868 | 127.7964 | 188.0795 | 381.053 | 427.2417 | 286.1474 | 164.3096 | 586.0232 | 481.0446 | 444.6904 | 479.6225 | 537.3775 | 785.702 | 709.6755 | 594.702 | 1083.0828 | 700.9205 | 637.3129 | 1171.2682
2246 | 26.7004 | 365.0107 | 412.879 | 452.2198 | 385.1877 | 317.9119 | 367.4617 | 368.7802 | 326.5489 | 366.2634 | 275.8239 | 211.7802 | 379.3514 | 434.7144 | 295.1539 | 266.5827 | 446.8354 | 394.4724 | 365.6922 | 310.9119 | 385.637 | 333.2856 | 738.7366 | 656.2535 | 318.3407 | 518.9119 | 537.8345 | 455.2749 | 695.242 | 217.6263 | 235.5053 | 179.6593 | 235.7802 | 320.121 | 292.9342 | 225.3407 | 138.5934 | 599.1984 | 603.9128 | 459.4724 | 650.7144 | 354.6041 | 328.0988 | 249 | 250.8568 | 126.0222 | 113 | 107.0658 | 234.8016 | 671.8897 | 731.8132 | 784.0107 | 540.7037 | 307.7909 | 316.4288 | 268.6815 | 239.8568 | 442.7144 | 437.5827 | 550.2856 | 509.8354 | 460.0658 | 459.4947 | 1073.4288 | 612.4947 | 409.4066 | 463.9778 | 540.2642 | 643.3292 | 339.637 | 376.8345 | 581.9671 | 554.0436 | 403.0552 | 407.8683 | 297.1317 | 645.0436 | 413.4947 | 415.1975 | 343.3514 | 571.0329 | 587.8239 | 579.9555 | 452.1317 | 444.2971 | 636.7696 | 640.0107 | 561.4947 | 491.1646 | 339.5276 | 370.3737 | 275.0329 | 429.5934 | 732.1325 | 719.6699 | 595.0445 | 1115.0329 | 844.1432 | 771.6708 | 655.4724 | 1192.814 | 426.8239 | 238.3407 | 158.5276 | 426.8239 | 286.8239 | 303.9342 | 211.0988 | 315.6922 | 423.6699 | 449.9448 | 289.3301 | 357.4057 | 190.9119 | 183.3078 | 114.2527 | 199.7037 | 371.5934 | 403.5053 | 301.8906 | 157.7473 | 557.7144 | 473.7135 | 449.4724 | 453.5819 | 531.9671 | 776.6263 | 714.5276 | 580.879 | 1076.4502 | 713.0329 | 657 | 1175.1432
2247 | 26.7114 | 344.1258 | 391.6459 | 437.6744 | 398.7717 | 318.8742 | 376.6057 | 360.9144 | 321.9598 | 338.4916 | 283.8055 | 226.3256 | 365.7601 | 464.2283 | 304.0571 | 280.6057 | 449.7315 | 408.7769 | 353.9429 | 314.9027 | 393.4112 | 324.9715 | 708.3606 | 640.332 | 304.4345 | 528.3541 | 519.4799 | 452.0973 | 701.9715 | 211.1543 | 235.0571 | 183.2685 | 253.9598 | 305.6174 | 300.6342 | 222.9429 | 145.6628 | 577.9663 | 632.7886 | 458.0285 | 632.9831 | 351.6913 | 346.834 | 242.9429 | 249.537 | 123.6744 | 114.2114 | 112.6913 | 216.6628 | 643.0973 | 720.703 | 751.0117 | 522.0519 | 286.9144 | 313.6744 | 266.3087 | 264.5824 | 444.9715 | 449.1829 | 552.2685 | 507.2802 | 443.8859 | 475.0856 | 1105.1997 | 582.6576 | 419.297 | 473.594 | 540.406 | 607.2516 | 344.3826 | 351.8171 | 573.703 | 547.4799 | 397.463 | 402.6744 | 310.1997 | 625.7601 | 434.537 | 417.537 | 343.0856 | 545.9831 | 581.8742 | 579.1997 | 450.0571 | 446.0402 | 618.8975 | 614.8859 | 566.1829 | 481.6057 | 365.8055 | 395.6226 | 264.5201 | 392.4462 | 685.2866 | 693.3087 | 607.3087 | 1083.3204 | 858.7886 | 799.8742 | 666.7484 | 1200.5318 | 420.8742 | 232.9144 | 150.8859 | 420.8742 | 280.8742 | 300.1258 | 211.6628 | 306.3658 | 394.8859 | 433.7315 | 286.0117 | 353.3139 | 189.4514 | 186.5772 | 126.7484 | 208.9144 | 379.2685 | 399.4228 | 310.0973 | 159.1829 | 553.3087 | 473.4968 | 419.5603 | 431.2114 | 518.2516 | 758.6459 | 719.6057 | 559.6459 | 1037.3774 | 742.4968 | 669.7198 | 1157.6861
2248 | 26.7224 | 334.056 | 424.2599 | 454.9559 | 382.382 | 342.5399 | 377.618 | 344.494 | 339.18 | 333.472 | 291.82 | 223.876 | 348.786 | 437.3581 | 317.27 | 273.584 | 466.6299 | 424.708 | 362.36 | 325.09 | 385.7521 | 305.0781 | 710.1359 | 651.2599 | 294.236 | 518.6061 | 550.3379 | 461.068 | 702.472 | 194.404 | 226.64 | 170.6961 | 255.73 | 298.562 | 291.258 | 232.596 | 160.506 | 563.45 | 574.5262 | 448.876 | 609.966 | 349.348 | 374.124 | 237.146 | 286.8099 | 116.854 | 110.674 | 119.618 | 226.562 | 617.4601 | 712.764 | 791.7658 | 510.978 | 290.27 | 310.562 | 260 | 255.7741 | 440.528 | 431.932 | 518.6841 | 484.786 | 432.292 | 476.056 | 1086.2461 | 546.022 | 423.528 | 492.562 | 518.966 | 598.236 | 326.7181 | 372.1579 | 568.236 | 544.348 | 386.146 | 396.472 | 309.168 | 631.034 | 422.3701 | 417.73 | 329.2241 | 546.45 | 528.9082 | 567.6621 | 472.5399 | 413.9881 | 589.348 | 607 | 559.438 | 460.3701 | 334.8861 | 370.2121 | 262.09 | 379.742 | 644.652 | 687 | 636.8439 | 1067.708 | 853.056 | 841.4619 | 717.7879 | 1204.8659 | 408.0781 | 224.528 | 166.4839 | 408.0781 | 274.258 | 301.18 | 209.82 | 283.6061 | 395.034 | 436.416 | 273 | 363.876 | 168.8981 | 176.966 | 132.056 | 204.236 | 371.64 | 395.438 | 283.2241 | 160.472 | 513.6281 | 460.0101 | 397.438 | 435.09 | 483.2801 | 746.146 | 721.236 | 553.326 | 994.8081 | 740.1341 | 661.314 | 1074.8182
2249 | 26.7334 | 349.9087 | 416.3341 | 448.3096 | 379.4788 | 366.6058 | 385.5635 | 345.5635 | 322.8307 | 314.2005 | 283.588 | 218.8909 | 326.1336 | 405.9577 | 299.6793 | 272.7817 | 445.9644 | 396.7461 | 355.5033 | 329.5212 | 355.6125 | 282.0245 | 682.098 | 667.2851 | 292.4788 | 516.7149 | 554.6548 | 437.5947 | 709.6726 | 204.9087 | 207.0245 | 177.0178 | 235.5033 | 283.7216 | 287.7817 | 229.6548 | 161.3274 | 585.6058 | 598.8775 | 456.4967 | 613.824 | 335.8062 | 360.2005 | 237.8909 | 285.0913 | 118 | 119.9755 | 106.1336 | 211.7216 | 565.6615 | 734.6904 | 784.6615 | 507.6548 | 279.6125 | 325.9755 | 252.4365 | 244.7149 | 420.7216 | 445.951 | 552.9866 | 469.0668 | 398.4855 | 484.3452 | 1103.1871 | 536.1091 | 406.2428 | 497.8909 | 529.7572 | 570.0067 | 314.9332 | 343.6615 | 551.3519 | 486.0557 | 374.9154 | 386.6548 | 305.0423 | 630.3274 | 394.5033 | 402.5457 | 321.412 | 567.3452 | 472.2005 | 542.5457 | 467.6125 | 423.6904 | 514.0379 | 597.5457 | 542.1336 | 410.4432 | 330.3207 | 338.6971 | 254.5457 | 347.2918 | 660.7149 | 715.9933 | 636.9399 | 1042.2673 | 844.3274 | 858.3274 | 719.6793 | 1234.7572 | 406.4543 | 218.588 | 169.0245 | 406.4543 | 256.2851 | 270.9644 | 215.4543 | 291.1693 | 372.2673 | 426.5033 | 291.2784 | 358.2606 | 171.6481 | 171.3697 | 123.9577 | 177.2673 | 351.3942 | 376.8731 | 279.1938 | 163.8909 | 515.0601 | 452.2361 | 394 | 408.0067 | 481.2361 | 715.3764 | 715.0668 | 518.8129 | 993.7572 | 755.6236 | 646.588 | 1038.2606
2250 | 26.7444 | 352.6465 | 434.5548 | 438.8555 | 391.4942 | 331.5897 | 387.4297 | 326.8671 | 321.2813 | 317.4219 | 257.8671 | 226.4258 | 342.4181 | 401.7148 | 316.2013 | 290.4219 | 406.7226 | 396.9922 | 355.9961 | 323.9316 | 367.2051 | 289.1368 | 681.631 | 648.1484 | 289.0723 | 538.4219 | 530.0078 | 430.209 | 677.6581 | 209.5742 | 216.0645 | 176.0761 | 235.9961 | 278.7832 | 288.7148 | 223 | 158.6426 | 566.0116 | 586.4491 | 444.4374 | 632.4942 | 325.2168 | 342.2168 | 256.9922 | 271.4297 | 116.0723 | 119.289 | 110.209 | 210.6387 | 558.7071 | 735.1484 | 714.7342 | 471.4413 | 279.6387 | 316.9355 | 244.7871 | 264.4942 | 416.4258 | 431.7265 | 548.4452 | 429.6581 | 372.9316 | 467.2245 | 1128.2129 | 470.742 | 436.1974 | 502.8555 | 553.7793 | 570.3497 | 306.791 | 355.3419 | 558.3497 | 466.5664 | 368.3574 | 396.0645 | 301.5742 | 616.7187 | 398.209 | 397.6426 | 336.8516 | 547.0116 | 441.3652 | 551.7793 | 449.0039 | 417.08 | 501.9161 | 536.0955 | 496.7303 | 347.2323 | 326.6503 | 325.3613 | 270.2051 | 343.92 | 632.3007 | 720.1484 | 613.0761 | 1006.7226 | 845.4981 | 842.1484 | 686.7226 | 1288.9806 | 374.0878 | 199.2929 | 143.3652 | 374.0878 | 258.3535 | 268.3497 | 219.7148 | 273.4413 | 343.791 | 421.2129 | 280.1523 | 342.9355 | 158.8671 | 169.0723 | 129.3535 | 162.2852 | 333.7187 | 344.2245 | 281.4297 | 161.7871 | 503.5819 | 449.7187 | 377.9355 | 407.0645 | 480.6465 | 672.9394 | 692.3652 | 482.8632 | 985.0078 | 790.9922 | 654.2813 | 1037.7148
2251 | 26.7554 | 338.1769 | 465.1682 | 413.4954 | 390.4867 | 316.5133 | 366.3539 | 356.5666 | 315.8672 | 292.6016 | 268.5754 | 221.4867 | 329.1503 | 410.8231 | 328.4159 | 310.7877 | 415.6461 | 372.2918 | 361.6549 | 327.8585 | 363.2831 | 298.9646 | 650.5041 | 690.1153 | 299.7877 | 514.9113 | 526.5487 | 461.2302 | 642.2831 | 219.7523 | 210.5928 | 166.7256 | 233.1415 | 274.1415 | 295.2036 | 214.4867 | 151.7964 | 565.4072 | 561.9646 | 438.239 | 632.1415 | 306.2831 | 344.239 | 282.7169 | 245.7698 | 123.5133 | 111.7256 | 113.0708 | 244.8143 | 559.4867 | 709.6636 | 770.8589 | 425.5308 | 270.5928 | 292.938 | 233.1769 | 230.7431 | 397.0795 | 442.8497 | 539.4426 | 413.2744 | 359.1769 | 460.5487 | 1089.398 | 426.4867 | 440.9733 | 491.9026 | 527.2918 | 555.1149 | 313.0974 | 369.761 | 536.5662 | 481.5133 | 351.6282 | 374.8759 | 303.8938 | 591.3539 | 376.1857 | 382.2831 | 319.1149 | 510.3892 | 418.5574 | 512.1944 | 430.9026 | 383.7344 | 509.4867 | 521.3718 | 512.0179 | 347.5754 | 304.5574 | 295.7698 | 278.7256 | 328.1503 | 575.0175 | 672.398 | 628.6108 | 1036.602 | 802.159 | 846.7877 | 708.7436 | 1276.6723 | 341.2123 | 208.9913 | 151.3364 | 341.2123 | 229.6016 | 276.6903 | 225.5487 | 245.2123 | 348.7877 | 388.9467 | 276.5133 | 299.2918 | 155.1682 | 159.4867 | 143.1682 | 163 | 320.1415 | 308.7344 | 281.3097 | 141.0087 | 505.7169 | 429.5928 | 400.4338 | 409.4513 | 474.6903 | 641.9733 | 672.1769 | 465.1769 | 906.2385 | 764.9821 | 688.8143 | 1052.7169
2252 | 26.7664 | 349.0119 | 465.3299 | 404.3373 | 392.0045 | 312.9941 | 357.3343 | 330.2971 | 315.3373 | 305.0223 | 272.3269 | 228.3418 | 288.9807 | 428.6761 | 326.6672 | 294.3165 | 401.315 | 352.6657 | 343.3195 | 321.3254 | 342.3239 | 297.3314 | 681.3745 | 667.9643 | 301.997 | 491.9955 | 525.997 | 457.3209 | 662.688 | 204.3165 | 203 | 175.6746 | 221.6612 | 264.6627 | 294.3299 | 240.6895 | 137.9926 | 575.6686 | 563.6672 | 441 | 670.6984 | 305.3388 | 364.3462 | 256.3076 | 227.6627 | 123.997 | 112.0015 | 111.6672 | 219.3031 | 548.3284 | 683.9881 | 761.2942 | 392.6538 | 265.6686 | 302.3462 | 240.0089 | 212.3358 | 364.318 | 462.3373 | 515.9777 | 389.6478 | 360.0045 | 448.6553 | 1057.9926 | 419.9985 | 454.6835 | 458.9807 | 508.3328 | 533.3269 | 313.9955 | 340.9807 | 533.3432 | 472.6568 | 378.3596 | 373.0104 | 327.0149 | 595.0104 | 391.0223 | 355.3195 | 336.0223 | 497.6657 | 393.9866 | 462.315 | 423.3328 | 361.9926 | 463.6359 | 513.9881 | 494.9732 | 333.3135 | 250.6315 | 292.3403 | 258.9866 | 311.9985 | 573.6805 | 663.6761 | 629.3254 | 991.9509 | 754.6493 | 893.6969 | 735.3403 | 1288.3566 | 352.3462 | 210.9941 | 154.3284 | 352.3462 | 226.0104 | 250.6478 | 196.309 | 237.6657 | 313.6359 | 364.9955 | 282.3343 | 266.9896 | 155.997 | 167.0089 | 138.6597 | 167.003 | 296.6523 | 295.6657 | 282.6672 | 124.9955 | 493.318 | 428.0045 | 427.6746 | 381.9822 | 439.3076 | 604.312 | 659.6612 | 434.6478 | 888.6805 | 776.6895 | 683.318 | 1004.2912
2253 | 26.7774 | 371.2685 | 442.6164 | 409.0384 | 349.4766 | 311.7178 | 383.8192 | 318.9096 | 304.4109 | 281.9506 | 263.9233 | 231.9616 | 299.1014 | 424.8082 | 322.9233 | 272.1287 | 395.1151 | 362.1918 | 309.5397 | 289.5013 | 327.8466 | 280.3726 | 638.3368 | 622.2575 | 292.526 | 469.9753 | 509.052 | 429.2959 | 683.7945 | 181.4493 | 209.7945 | 168.7698 | 207.8082 | 268.7945 | 288.6027 | 235.6164 | 132.3205 | 534.8739 | 567.3973 | 440.3205 | 681.8082 | 309.6795 | 356.6931 | 270.2549 | 227.0384 | 113.1671 | 115.7178 | 109.9616 | 217.3452 | 509.6684 | 668.526 | 728.2055 | 396.9096 | 264.2822 | 285.8602 | 247.3589 | 199.7315 | 373.0247 | 467.0384 | 525.4603 | 362.7315 | 307.2848 | 418.578 | 1035.3342 | 402.6931 | 481.6274 | 442.6027 | 529.0631 | 532.3973 | 304.2055 | 339.5507 | 524.3726 | 444.9369 | 366.104 | 383.3973 | 304.3862 | 565.989 | 378.8219 | 356.8713 | 319.0657 | 509.9096 | 377.526 | 437.7698 | 438.6274 | 350.2055 | 439.6027 | 484.9369 | 491.948 | 324.0767 | 269.1261 | 276.6164 | 263.5891 | 304.8849 | 582.3205 | 614.9643 | 583.9122 | 1007.9206 | 717.8602 | 935.7425 | 689.041 | 1325.0631 | 330.4246 | 183.2191 | 161.8713 | 330.4246 | 216.6931 | 252.2685 | 188.1534 | 244.474 | 281.4493 | 334.8219 | 281.6411 | 249.8082 | 126.1424 | 158.052 | 143.5123 | 163.5644 | 276.8082 | 286.8466 | 266.6931 | 126.0767 | 455.8219 | 435.7562 | 441.1534 | 380.1918 | 409.7698 | 570.9753 | 650.5644 | 437.6274 | 853.1561 | 753.2712 | 649.8986 | 955.6164
2254 | 26.7884 | 378 | 453.7523 | 427.9854 | 357.0533 | 310.9248 | 398.767 | 354.745 | 289.0073 | 277.1431 | 248.1651 | 232.3835 | 294.7816 | 381.2623 | 325.8422 | 278.0679 | 400.0752 | 346.2477 | 292.466 | 297.7523 | 340.6019 | 257.0899 | 649.9634 | 632.7523 | 278.6238 | 469.9927 | 506.1505 | 399.2477 | 682.8495 | 212.6625 | 216.4587 | 171.9927 | 213.3762 | 249.8642 | 269.0146 | 189.3375 | 138.2257 | 517.767 | 570.3835 | 439.3083 | 669.3908 | 294.7816 | 355.2257 | 256.6385 | 214.8569 | 112.4587 | 106.6238 | 108.3083 | 221.8495 | 490.9248 | 649.7816 | 693.8715 | 367.0293 | 258.8495 | 285.0679 | 257.6844 | 190.233 | 341.1871 | 472.8422 | 481.6604 | 344.2403 | 319.7377 | 397.6238 | 992.4128 | 372.8642 | 477.2403 | 444.4587 | 482.277 | 491.8036 | 296.8495 | 363.6771 | 498.3229 | 422.5486 | 331.2477 | 392.6092 | 293.8422 | 553.8422 | 339.7137 | 357.1578 | 324.0606 | 500.7816 | 356.0146 | 396.0293 | 425.9394 | 331.0899 | 408.9467 | 449.4055 | 509.3762 | 315.6238 | 246.1118 | 271.1505 | 247.1724 | 278.4807 | 607.5945 | 636.0386 | 564.3083 | 961.1338 | 692.1651 | 921.0972 | 722.4148 | 1343.9927 | 303.5486 | 176.1505 | 163.5413 | 303.5486 | 195.8569 | 243.0899 | 214.1358 | 221.7137 | 269.0826 | 299.1724 | 294.1431 | 234.6238 | 129.6019 | 165.5266 | 134.8569 | 142.3229 | 251.2477 | 278.8495 | 262.4587 | 125.233 | 461.6771 | 426.2403 | 418.022 | 394.6844 | 380.4807 | 533.6385 | 628.6312 | 449.1505 | 822.3156 | 743.3009 | 634.1578 | 912.1045
2255 | 26.7994 | 335.7579 | 446.8072 | 422.6233 | 347.4709 | 299.4395 | 376.7668 | 339.8386 | 298.3767 | 247.0942 | 226.5112 | 226.6637 | 266.1749 | 377.3767 | 334.3363 | 285.1121 | 362.9821 | 348.9686 | 294.9282 | 290.1031 | 320.5426 | 278.8655 | 644.6547 | 633.5516 | 275.4081 | 460.6233 | 469.278 | 406.1928 | 650.7265 | 195.2063 | 213.7758 | 181.6323 | 211.6637 | 267.4574 | 276.4888 | 207.4977 | 151.5605 | 504.2153 | 578.0404 | 439 | 654.7354 | 273.2153 | 368.5605 | 227.8072 | 228.009 | 125.9686 | 116.0807 | 105.1839 | 203.1031 | 452.6861 | 621.8789 | 685.6323 | 366.4888 | 236.583 | 280.1435 | 240.8789 | 177.7354 | 330.7444 | 469.3677 | 488.6816 | 320.6951 | 320.1031 | 422.5695 | 934.7579 | 363 | 490.3049 | 440.3677 | 482.3453 | 463.8475 | 280.2153 | 337.5023 | 530.13 | 407.1435 | 363.5381 | 362.2063 | 310.7847 | 538.3991 | 304.7668 | 348.6637 | 322.4395 | 476.3991 | 326.8789 | 393.3767 | 414.1839 | 316.2556 | 395.296 | 445.7444 | 497.1031 | 289.1749 | 222.7758 | 261.7354 | 270.7937 | 281.3767 | 550.7086 | 620.3902 | 557.6637 | 945.4888 | 703.6009 | 857.9014 | 725.4709 | 1290.9731 | 257.1659 | 157.583 | 162.704 | 257.1659 | 222.3856 | 233.1839 | 218.3677 | 205.7758 | 233.6144 | 281.2556 | 254.9417 | 232.8161 | 122.9193 | 129.0538 | 140.9686 | 142.4484 | 234.9596 | 236.87 | 268.9282 | 124.704 | 479.8565 | 408.3274 | 381.3588 | 385.6233 | 399.5695 | 510.4395 | 610.1435 | 418.6144 | 800.3991 | 682.2287 | 684.0986 | 857.574
2256 | 26.8104 | 353.8163 | 458.6245 | 439.2061 | 346.5959 | 303.5959 | 373.4469 | 318.6878 | 283.2265 | 254.9225 | 225.0143 | 251.9367 | 262.0143 | 350.049 | 330.5531 | 291.7306 | 367.7735 | 344.6878 | 290.5531 | 310.2204 | 305.4184 | 271.6592 | 598.1837 | 637.1633 | 301.7878 | 472.9082 | 469.4755 | 415.149 | 613.6449 | 187.4469 | 196.2408 | 182.5673 | 199.6878 | 253.9286 | 253.6306 | 228.4469 | 142.3898 | 490.8367 | 578.1347 | 458.3408 | 633.5245 | 262.702 | 336.4673 | 218.851 | 231.702 | 106.6449 | 115.5531 | 129.7878 | 201.7306 | 449.8939 | 585.0633 | 665.7939 | 354.3755 | 230.8653 | 299.6388 | 223.4041 | 183.7449 | 346.8939 | 435.4816 | 479.7939 | 321.5959 | 278.6163 | 390.5878 | 917.7163 | 346.5245 | 460.751 | 449.4612 | 500.8796 | 465.0143 | 264.6878 | 318.702 | 499.7224 | 397.9857 | 355.9286 | 349.4327 | 328.4612 | 553.2061 | 312.1918 | 339.5531 | 331.6102 | 445.3612 | 321.5816 | 372.4959 | 397.2408 | 301.5388 | 386.4041 | 418.198 | 422.6653 | 288.5959 | 236.7592 | 263.4469 | 277.8367 | 271.9571 | 564.9796 | 617.8939 | 587.2347 | 904.0061 | 672.3184 | 883.2776 | 700.2408 | 1257.6878 | 243.5816 | 143.9857 | 148.6735 | 243.5816 | 180.1265 | 201.198 | 223.1633 | 201.851 | 234.3265 | 271.5531 | 241.7306 | 221.1061 | 125.5959 | 119.5959 | 133.1061 | 148.149 | 214.0918 | 221.4327 | 255.2408 | 109.2408 | 463.9429 | 436.6674 | 373.8653 | 383.5816 | 384.7796 | 521.0429 | 617.4612 | 427.9225 | 815.2061 | 639.2408 | 641.5306 | 827.3898
2257 | 26.8214 | 375.2862 | 474.0148 | 437.7994 | 336.8849 | 277.8553 | 362.8849 | 318.6431 | 266.9852 | 259.6283 | 221.8997 | 250.6135 | 233.3981 | 387.2746 | 338.2006 | 310.0296 | 365.7994 | 339.5428 | 272.699 | 294.7698 | 288.699 | 273.472 | 614.6019 | 621.0559 | 277.0264 | 445.0264 | 471.6283 | 426.2006 | 601.0855 | 184.1711 | 173.9704 | 184.9145 | 208.3865 | 228.699 | 260.2154 | 236.1003 | 128.9852 | 466.1415 | 536.1974 | 469.6431 | 643.0296 | 271.2006 | 327.1003 | 242.0444 | 212.5132 | 102.8289 | 107.1711 | 124.699 | 183.5987 | 467.3865 | 541.9408 | 670.8437 | 358.6579 | 220.3421 | 285.6842 | 233.8437 | 200.3865 | 308.2829 | 390.2122 | 452.7846 | 314.0707 | 286.8585 | 367.8997 | 900.5132 | 337.0855 | 505.7466 | 407.3685 | 474.8553 | 435.6694 | 261 | 311.1711 | 475.8997 | 401.1003 | 339.4424 | 366.0296 | 322.7994 | 572.9293 | 302.2418 | 319.5132 | 326.7994 | 456.4013 | 311.3421 | 318.2829 | 392.4572 | 288.2566 | 348.0264 | 400.8997 | 433.8881 | 259.9408 | 240.0855 | 221.5543 | 242.9408 | 249.2418 | 520.7961 | 649.2302 | 609.472 | 906.8585 | 628.5839 | 921.6727 | 682.3421 | 1244.528 | 236.9852 | 144.1859 | 161.9441 | 236.9852 | 190.0592 | 207.9441 | 217.9852 | 176.227 | 244.3717 | 275.5576 | 262.2154 | 213.8141 | 123.1711 | 123.7286 | 130.9145 | 167.9441 | 193.1563 | 210.3421 | 267.2154 | 112.472 | 440.699 | 417.9408 | 384.472 | 369.699 | 358.699 | 493.4836 | 608.8849 | 423.1563 | 825.4572 | 624.2566 | 665.546 | 803.0559
2258 | 26.8324 | 352.8091 | 485.15 | 375.4682 | 334.2227 | 275.15 | 311.3228 | 314.0727 | 245.4773 | 239.4773 | 191.8455 | 238.5909 | 255.3409 | 363.0273 | 316.8091 | 308.5909 | 353.1091 | 312.3273 | 275.15 | 274.6273 | 295.5954 | 261.4409 | 601.55 | 652.5273 | 280.3 | 403.1046 | 458.1455 | 419.6273 | 543.6909 | 175.3318 | 190.2273 | 160.0682 | 207.8136 | 217.0727 | 259.5909 | 243.1864 | 123.0364 | 481.7091 | 559.5273 | 472.4818 | 631.9591 | 277.2227 | 300.1046 | 265.0773 | 204.5182 | 104.2591 | 116.8545 | 117.5182 | 189.3364 | 436.4364 | 545.5591 | 661.1818 | 332.6227 | 215.2591 | 252.55 | 233.0727 | 185.7364 | 299.4091 | 378 | 476.3773 | 290.7364 | 277.2182 | 351.1818 | 842.1364 | 299.6955 | 487.25 | 395.9273 | 466.2227 | 439.5591 | 247.6636 | 316.4091 | 472.5182 | 398.5545 | 344.4091 | 363.85 | 310.85 | 534.0273 | 302.6682 | 311.5182 | 309.6636 | 445.4773 | 292.9227 | 325.7091 | 377.4409 | 275.3682 | 333.5182 | 364.1773 | 459.8545 | 241.3318 | 219.7364 | 225.7454 | 242.8545 | 225.2182 | 547.6409 | 581.2046 | 599.6636 | 883.1409 | 633.5591 | 898.6227 | 740.9772 | 1200.9909 | 212.5136 | 145.7409 | 146.0318 | 212.5136 | 196.0727 | 232.7818 | 212.0364 | 184.7818 | 216.3637 | 249.1046 | 280.8545 | 199.6636 | 117.2955 | 129.9273 | 129.7773 | 140.9182 | 191.7045 | 197.1091 | 262.1455 | 121.1864 | 437.2227 | 382.2909 | 386.5182 | 366.9636 | 324.8455 | 475.8136 | 562.5091 | 401.7 | 785.25 | 624.7045 | 643.25 | 849.3454
2259 | 26.8434 | 378.6472 | 452.5996 | 395.6728 | 338.0128 | 251.638 | 314.5832 | 319.532 | 231.468 | 229.234 | 204.596 | 252.5704 | 241.3912 | 365.3236 | 335.8684 | 287.9232 | 368.83 | 319.5704 | 265.1956 | 270.2468 | 277.404 | 236.4168 | 580.9488 | 614.7824 | 246.0804 | 417.8556 | 474.3364 | 440.1024 | 575.2176 | 169.9872 | 191.9744 | 159.2852 | 189.4296 | 223.2852 | 244.9488 | 226.17 | 151.3748 | 477.9488 | 569.9872 | 480.532 | 617.7148 | 262.1828 | 305.064 | 267.7404 | 213.7916 | 104 | 117.234 | 107.9616 | 176.6764 | 424.7532 | 534.4296 | 610.0548 | 336.3108 | 223.2852 | 244 | 233.2596 | 168.4552 | 308.0256 | 374.234 | 466.6636 | 293.0384 | 252.9232 | 333.1956 | 787.0932 | 287.7532 | 482.0384 | 406.2852 | 465.4936 | 454.7916 | 255.0512 | 322.0128 | 478.0256 | 386.4552 | 338.7148 | 361.7532 | 302.7276 | 497.1572 | 286.17 | 323.8044 | 315.5448 | 411.1316 | 272.936 | 291.8208 | 363.7148 | 276.5192 | 326.9744 | 339.1956 | 420.3144 | 252.5576 | 203.9616 | 217.1828 | 268.8428 | 231.8044 | 537.8848 | 572.0768 | 566.3784 | 844.1316 | 600.5868 | 835.276 | 717.0548 | 1170.4296 | 195.2084 | 136.9616 | 150.8044 | 195.2084 | 182.702 | 196.8208 | 205.7276 | 185.7276 | 206 | 239.7532 | 269.936 | 190.4808 | 110.7276 | 124.468 | 126.7404 | 158.3748 | 192.2468 | 178.936 | 257.2468 | 129.0256 | 414.6508 | 387.5576 | 381.4808 | 389.8428 | 358.6856 | 453.6636 | 561.5576 | 390.7276 | 734.0932 | 603.404 | 662.1408 | 808.4972
2260 | 26.8544 | 376.2212 | 454.7788 | 388.3317 | 308.3804 | 254.4824 | 311.0486 | 307.4557 | 255.0267 | 253.2612 | 202.0486 | 256.469 | 218.4557 | 365.6416 | 329.6902 | 288.8894 | 345.1459 | 308.1592 | 252.5796 | 290.6683 | 282.2479 | 252.1992 | 588.4824 | 596.469 | 242.9514 | 404.5663 | 435.0706 | 413.5529 | 569.0353 | 192.7302 | 187.7035 | 172.7169 | 188.5929 | 236.7169 | 241 | 218.469 | 142.6282 | 507.6816 | 546.0353 | 440.898 | 608.8761 | 235.5663 | 294.6902 | 259.3451 | 199.3937 | 112.4204 | 118.2965 | 95.0486 | 177.1239 | 397.4423 | 508.3317 | 625.6196 | 292.7741 | 221.4071 | 254.7169 | 237.8275 | 178.7788 | 283.2078 | 363.8141 | 462.2965 | 279.9247 | 217.1459 | 335.1239 | 764.2831 | 282.6416 | 471.9867 | 378.3804 | 477.2479 | 413.6015 | 256.7035 | 303.8627 | 462.3937 | 380.7035 | 329.1106 | 358.1725 | 284.9247 | 508.3718 | 282.2965 | 311.1592 | 304.4557 | 396.6416 | 273.3584 | 294.8408 | 341.0972 | 254.2698 | 318.1106 | 343.4204 | 416.7169 | 252.4071 | 177.2698 | 231.1373 | 254.5663 | 203.8494 | 538.9514 | 568.8141 | 592.2126 | 822.7521 | 623.9781 | 782.0219 | 694.3451 | 1166.531 | 181.2831 | 118.6902 | 157.2965 | 181.2831 | 181.2965 | 193.7169 | 219.3098 | 172.5176 | 201.4071 | 212.4423 | 264.2345 | 162.2078 | 115.1239 | 122.7655 | 131.3584 | 138.1459 | 168.2698 | 170.938 | 257.7655 | 108.8008 | 431.4957 | 383.5796 | 387.6549 | 381.6902 | 337.0219 | 441.6416 | 559.8761 | 383.6416 | 776.3498 | 589.1106 | 620.1811 | 809.6683
2261 | 26.8654 | 375.1111 | 446.5555 | 396.7779 | 329.3335 | 247.1111 | 311.8889 | 326.3335 | 232.6665 | 232.2221 | 203.6667 | 236.5555 | 222.5556 | 369.4445 | 314.1111 | 294.8889 | 346.8889 | 296 | 243.7777 | 271.3332 | 262.6665 | 232.5554 | 581.1111 | 572.6665 | 267.7779 | 412.7779 | 435.0001 | 403.7778 | 574.4445 | 199.2222 | 187.7778 | 166.6666 | 207.889 | 218.9999 | 226.9999 | 228.1112 | 124.5555 | 513.3333 | 494.6664 | 434.2223 | 616.3334 | 221.2222 | 293.1111 | 244.5555 | 196.3333 | 108.7777 | 99.5555 | 112.2223 | 179.7778 | 355.5554 | 507.4445 | 643.5556 | 296.6668 | 202.8888 | 231.5554 | 224.2221 | 194.6667 | 278.1111 | 372.6667 | 481.6668 | 285.8889 | 218.1112 | 305.8887 | 749.3333 | 281.7778 | 434.5554 | 373.6667 | 481 | 416.3334 | 231.111 | 317.4445 | 456.2222 | 370.6666 | 320 | 347.6666 | 289.3334 | 538.889 | 294.6667 | 288.8888 | 310.8889 | 375.5555 | 250.8888 | 279.7777 | 342.7778 | 239.2222 | 336.2223 | 346.7778 | 451.1113 | 234.6666 | 168.4444 | 222.9999 | 255.7778 | 230.5557 | 525.6666 | 563.6667 | 564.8887 | 815.8889 | 571.9997 | 813.0002 | 725.4446 | 1074.444 | 195.1112 | 121 | 153.3333 | 195.1112 | 183.5556 | 175.2221 | 211.5555 | 171.3333 | 177.4443 | 211 | 237.5554 | 181.2224 | 102.9999 | 110.5555 | 131.4444 | 153.1112 | 182.0001 | 166.8889 | 236.9999 | 108.2223 | 435.8889 | 375.5555 | 393.1111 | 381.6667 | 307.3332 | 428.3333 | 568.8889 | 374.2222 | 721.4441 | 548.8887 | 608.6667 | 766.222
2262 | 26.8764 | 366.5191 | 439.8397 | 389.1489 | 309.5574 | 263.7518 | 320.9007 | 307.7177 | 234.2709 | 208.9886 | 204.2099 | 231.7901 | 216.3092 | 355.1985 | 315.7404 | 291.2596 | 323.9276 | 296.3702 | 259.3816 | 290.0724 | 244.939 | 251.0724 | 596.9618 | 550.1985 | 253.4581 | 399.6184 | 408.1872 | 428.4923 | 546.3971 | 196.6298 | 165.0879 | 160.8397 | 216.9504 | 248.5533 | 256.1831 | 231 | 121.7901 | 478.8169 | 461.4581 | 428.0993 | 567.6453 | 227.6908 | 255.2864 | 232.3092 | 186.7291 | 104.6298 | 104.2709 | 122.7404 | 162.6184 | 377.6029 | 490.0382 | 612.0269 | 308.3206 | 196.4199 | 215.3092 | 218.4199 | 175.878 | 279.7901 | 379.9504 | 453.0269 | 284.2596 | 194.1375 | 275.668 | 688.3247 | 251.1872 | 444.7518 | 383.6908 | 447.0269 | 432.061 | 240.5915 | 294.5574 | 440.1985 | 353.7787 | 322.7404 | 356.8511 | 293.5801 | 519.1375 | 282.9886 | 294.2709 | 298.9886 | 408.7136 | 249.5305 | 247.9276 | 372.6526 | 242.5305 | 305.6566 | 325.668 | 411.8056 | 230.7901 | 190.122 | 200.8283 | 240.6184 | 222.8283 | 496.5077 | 551.1489 | 602.9845 | 808.6794 | 559.5305 | 768.1147 | 739.7404 | 1085.9235 | 219.7518 | 123 | 133.8283 | 219.7518 | 154.7673 | 169 | 177.9772 | 161.7291 | 174.9504 | 211.4199 | 263.9731 | 175.5688 | 109.2709 | 110.9504 | 114.4085 | 145.7787 | 148.4964 | 164.4199 | 227.8397 | 93.4085 | 414.4581 | 375.5801 | 384.5191 | 357.7177 | 282.2482 | 402.878 | 553.0382 | 379.1107 | 711.2709 | 534.0496 | 591.0382 | 770.9618
2263 | 26.8874 | 346.348 | 458.2567 | 396.4307 | 298.7905 | 267.3953 | 270.044 | 311.4307 | 239.4071 | 236.2922 | 192.7669 | 231.1976 | 217.2095 | 345.3835 | 323.4189 | 290.8024 | 321.8142 | 288.9764 | 234.3125 | 294.5929 | 278.1064 | 252.3835 | 573.1149 | 616.608 | 231.9527 | 386.174 | 415.2449 | 444.6284 | 521.9527 | 181.5574 | 184.6756 | 172.0355 | 189.1149 | 246.7669 | 234.5102 | 208.5338 | 134.0355 | 454.7551 | 509.7584 | 424.3953 | 614.1773 | 214.7551 | 239.3835 | 258.0827 | 178.3835 | 94.3716 | 112.6165 | 128.0118 | 172.4426 | 396.4307 | 452.9054 | 601.3953 | 300.3716 | 195.1976 | 237.8733 | 216.3953 | 182.0355 | 289.6284 | 364.9527 | 403.0794 | 284.6047 | 211.8733 | 270.8024 | 666.5811 | 230.9645 | 476.478 | 351.4983 | 432.3835 | 422.9645 | 241.7905 | 265.3362 | 451.2449 | 346.7905 | 325.6047 | 361.6047 | 287.5811 | 466.2653 | 296.652 | 284.9645 | 295 | 394.9291 | 237.3598 | 241 | 331.0558 | 223.9409 | 303.2213 | 348.0827 | 402.2095 | 218.9645 | 173.5338 | 200.8142 | 246.4307 | 214.7905 | 532.5253 | 539.9764 | 576.6842 | 811.0118 | 594.6993 | 736.1504 | 721.7433 | 1043.044 | 201.7314 | 123 | 134.6165 | 201.7314 | 166.2567 | 181.0355 | 186.8496 | 167.0236 | 171.9882 | 176.4983 | 269.7905 | 184.0355 | 99.1622 | 101.5693 | 134.0709 | 133.174 | 144.4189 | 155.9764 | 215.7669 | 117.8851 | 402.5811 | 366.3716 | 383.6047 | 369.4544 | 322.7347 | 411.4426 | 564.0473 | 360.1385 | 712.3953 | 546.6402 | 577.9764 | 809.6993
2264 | 26.8984 | 355.0345 | 444.2629 | 377.819 | 273.5603 | 249.8922 | 314.0259 | 294.4483 | 232.6681 | 218.7457 | 204.6638 | 231.8147 | 236.7371 | 327.7069 | 305.4483 | 308.1078 | 342.5517 | 271.5216 | 223.7414 | 285.0388 | 282.1121 | 259.1466 | 544.0043 | 608.7457 | 236.9612 | 365.2629 | 408.4095 | 390.7888 | 479.7112 | 181.2586 | 175.5216 | 158.7069 | 175.4828 | 239.1121 | 232.7026 | 220.9224 | 120.7069 | 442.0388 | 479.8233 | 435.4052 | 598.0431 | 200.4095 | 258.3664 | 238.1164 | 175.3707 | 103.4052 | 101.7802 | 113.5216 | 142.5991 | 382.7069 | 462.9224 | 574.1164 | 322.4397 | 206.4052 | 234.2241 | 202.9655 | 172.7802 | 287.9267 | 348.7802 | 436.1767 | 267.0776 | 237.5517 | 253.8922 | 645.4483 | 221.4828 | 399.9052 | 337.2974 | 452.9957 | 402.0776 | 246.7026 | 263.2586 | 438.7069 | 344.3707 | 297.4871 | 365.2586 | 277.8534 | 515.2845 | 296.1121 | 295.0345 | 319.4397 | 425.6595 | 229.1121 | 258.9224 | 327.9612 | 226.3319 | 344.9181 | 297.1595 | 419.2931 | 239.625 | 188.3664 | 207.7026 | 212.3405 | 193.6336 | 526.7069 | 520.8922 | 596.3276 | 851.1034 | 570.4138 | 718.1509 | 721.8879 | 1038.9612 | 174.819 | 113.2241 | 129.4828 | 174.819 | 188.9224 | 170.1509 | 186.9267 | 166.556 | 184.8491 | 184.2931 | 264.1121 | 160.1164 | 98.444 | 95.7414 | 118.0043 | 129.3707 | 159.8491 | 153.1853 | 222.7759 | 108.7069 | 405.8879 | 380.2931 | 436.9526 | 380.5172 | 315.2629 | 412.556 | 497.125 | 376.181 | 740.5129 | 519.0431 | 564.5948 | 811.4828
2265 | 26.9094 | 370.4042 | 401.1335 | 351.4994 | 265.5192 | 217.0569 | 283.9988 | 264.3647 | 229.3461 | 201.4228 | 166.6527 | 217.115 | 209.576 | 297.5377 | 304.3078 | 285.5377 | 335.4228 | 281.2311 | 261.8665 | 260.6725 | 276.0383 | 239.4994 | 533.2114 | 574.8838 | 246.4425 | 384.1545 | 396.9036 | 426.7898 | 475.9617 | 184.4808 | 190.1928 | 181.4623 | 186.4042 | 209.0569 | 221.5958 | 220.0383 | 131.0581 | 471.424 | 472.4808 | 469.424 | 547.1719 | 224.4623 | 286.1545 | 252.6737 | 164.2497 | 102.6922 | 108.0964 | 117.4425 | 173.8665 | 393.885 | 427.3066 | 562.2114 | 262.6713 | 197.4228 | 220.4228 | 228.1162 | 162.5575 | 253.9222 | 364.1928 | 454.2695 | 259.6922 | 218.0186 | 245.0383 | 610.403 | 220 | 408.2892 | 326.9036 | 434.8455 | 400.4808 | 282.7317 | 287.9814 | 426.7305 | 328.288 | 325.7317 | 345.3263 | 302.4623 | 523.2114 | 271.8455 | 294.6922 | 336.5772 | 395.9605 | 208.1533 | 264.6539 | 321.7305 | 208.1533 | 298.5946 | 294.7503 | 354.3635 | 223.4994 | 188.0383 | 185.8455 | 225.5006 | 199.7503 | 508.9419 | 548.424 | 550.9024 | 876.5389 | 560.5192 | 692.6725 | 696.5377 | 994.691 | 180.7503 | 120.9234 | 139.5772 | 180.7503 | 178.9419 | 178.5772 | 185.1731 | 173.4425 | 158.2299 | 185.5192 | 252.2497 | 172.1928 | 90.7305 | 87.5575 | 124.5772 | 126.5192 | 158.0383 | 167.885 | 237.4042 | 89.288 | 411.1347 | 366.6341 | 404.3449 | 389.4425 | 319.2695 | 418.6156 | 532.2707 | 368.5958 | 680.8443 | 509.5192 | 579.3659 | 743.8443
2266 | 26.9204 | 366.2861 | 399.7139 | 362.1063 | 268.3569 | 243.7302 | 259.0545 | 269.9101 | 227.3215 | 246.8365 | 159.6785 | 228.267 | 203.8392 | 333.9619 | 273.109 | 268.2507 | 335.5177 | 268.8937 | 266.6431 | 261.0354 | 270.8038 | 237.5177 | 573.1226 | 568.1608 | 254.7139 | 368.8583 | 424.3733 | 377.4142 | 503.8556 | 200.1063 | 188.9646 | 184.4823 | 213.3379 | 230.6948 | 212.2861 | 222.3569 | 131.4823 | 443.5913 | 457.0545 | 445.7875 | 540.5177 | 209.0191 | 250.7166 | 246.9292 | 211.515 | 106.1962 | 100.7684 | 103.0545 | 161.0191 | 352.5204 | 429.0708 | 555.1253 | 273.3379 | 189.1253 | 223.0354 | 260.8556 | 147.5722 | 269.6594 | 323.5204 | 410.6812 | 259.8392 | 218.0354 | 233.0899 | 591.4114 | 231.7493 | 375.3951 | 309.0545 | 409.0191 | 383.376 | 223.7003 | 252.7166 | 434.2316 | 351.0164 | 328.8038 | 357.7847 | 298.7684 | 526.1962 | 263.6431 | 309.9455 | 326.4114 | 419.0518 | 228.3379 | 258.2861 | 338.4632 | 226.6594 | 339.8719 | 276.0191 | 371.891 | 243.3379 | 171.8937 | 176.8038 | 224.1253 | 193.6076 | 484.1798 | 544.9292 | 594.5504 | 845.5913 | 522.2343 | 681.2861 | 686.8038 | 938.842 | 209.8556 | 112.9292 | 137.8038 | 209.8556 | 164.2507 | 181 | 177.4469 | 147.3052 | 149.4823 | 178.2861 | 240.7684 | 174.3215 | 91.5177 | 89.3569 | 113.5722 | 141.1063 | 162.8747 | 160.0899 | 196.3597 | 100.267 | 411.1608 | 369.7139 | 386.6076 | 375.0545 | 310.0899 | 427.5531 | 485.0926 | 382.7847 | 679.5886 | 542.5695 | 591.3924 | 772.8011
2267 | 26.9314 | 370.1092 | 393.3362 | 353.0786 | 276.6638 | 227.8603 | 302.8341 | 294.1397 | 204.8603 | 216.8297 | 167.6638 | 206.3057 | 225.2882 | 318.1572 | 257.6332 | 283.0306 | 321.524 | 273.6638 | 212.3537 | 260.2969 | 312.5764 | 242.2576 | 555.4542 | 583.3275 | 271.3275 | 367.5546 | 422.3362 | 360.8908 | 495.3755 | 164.6812 | 177.7817 | 171.2271 | 186.4935 | 213.8603 | 230.5851 | 242.5851 | 133.5546 | 468.5065 | 494.8734 | 424.6725 | 541 | 205.8515 | 248.9607 | 224.5633 | 180.9782 | 99.3362 | 105.8122 | 107.6638 | 179.9913 | 371.2489 | 412.2664 | 603.3886 | 286.5153 | 210.1397 | 235.9214 | 225.1266 | 166.2882 | 282.5153 | 325.2183 | 399.4454 | 246.3755 | 203.6725 | 256.5458 | 566.8603 | 253.5851 | 377.3668 | 322.179 | 434.8035 | 379.1485 | 234.8428 | 245 | 454.7336 | 352.5939 | 300.7511 | 334.6026 | 306.3668 | 492.9389 | 247.6725 | 248.2839 | 307.821 | 380.7205 | 215.1179 | 266.3668 | 303.6812 | 255.6943 | 261.4411 | 312.0218 | 390.7729 | 206.2751 | 175.8122 | 181.9607 | 220.4454 | 196.2576 | 479.1485 | 547.2576 | 556.4629 | 833.8908 | 558.4279 | 638.2751 | 682.5939 | 899.345 | 190.3057 | 116.9607 | 142.1092 | 190.3057 | 168.8122 | 164.821 | 187.9214 | 166.6943 | 174.5458 | 177 | 238.1485 | 181.6638 | 111.5851 | 87.4454 | 126.3275 | 118.4542 | 153.7817 | 158.8515 | 193.1092 | 103 | 402.4847 | 366.7424 | 376.4847 | 377.9607 | 307.1485 | 396.6419 | 488.476 | 362.1572 | 676.8908 | 508.9782 | 609.179 | 744.3843
2268 | 26.9424 | 335.5834 | 411.8678 | 357.0467 | 316.8719 | 236.9573 | 293.7236 | 272.0854 | 226.9146 | 224.6849 | 173.3191 | 234.8251 | 217.7703 | 318.3191 | 239.5874 | 284.2724 | 333.6849 | 274.1362 | 195.2256 | 216.809 | 269.0347 | 223.1322 | 545.8171 | 584.2724 | 251.5407 | 392.1869 | 425.3191 | 415.2844 | 503.3658 | 161.4553 | 178.5915 | 170.7276 | 171.498 | 213.4553 | 239.1829 | 231.3151 | 127.0894 | 436.8558 | 497.6809 | 421.1362 | 560.004 | 212.9106 | 246.5447 | 278.8759 | 193.004 | 117.004 | 94.0427 | 119.3658 | 158.0854 | 369.9533 | 395.1789 | 523.8904 | 299.2297 | 202.7703 | 224.1789 | 219.7276 | 177.7744 | 274.498 | 319.2256 | 412.8211 | 247.4553 | 202.7276 | 229.0387 | 563.8638 | 257 | 345.311 | 311.1789 | 425.3151 | 402.3231 | 237.2724 | 254.502 | 452.8171 | 334.7236 | 309.8211 | 342.9573 | 287.2683 | 506.004 | 237.2256 | 239.5915 | 298.0894 | 353.996 | 254.3271 | 280.9573 | 331.5528 | 221.992 | 301.2844 | 310.3618 | 383.498 | 205.0467 | 170.0894 | 198.5487 | 207.0427 | 192.6809 | 496.2764 | 525.5407 | 572.1869 | 809.6769 | 557.3618 | 650.8678 | 719.1442 | 873.2683 | 192.9106 | 111.9533 | 155.9573 | 192.9106 | 158.7703 | 181.8678 | 185.6809 | 158.9065 | 166.9065 | 165.7703 | 237.1362 | 176.0894 | 95.996 | 106.004 | 112.5874 | 109.6809 | 145.9533 | 151.2256 | 213.004 | 107.3191 | 397.5447 | 376.3658 | 382.7744 | 348.7663 | 320.8211 | 397.9106 | 490.1362 | 352.8171 | 715.7357 | 444.1241 | 582.6301 | 702.5834
2269 | 26.9534 | 333.5044 | 374.6989 | 379.9028 | 279.1945 | 223.23 | 268.2211 | 290.7789 | 216.1061 | 228.7522 | 154.7256 | 224.23 | 206.3628 | 305.8583 | 239.6283 | 273.4867 | 321.9822 | 263.4867 | 202.7611 | 255.5577 | 250.4867 | 215.6194 | 607.2038 | 583.1239 | 248.8761 | 362.7078 | 410.23 | 388.0795 | 490.9822 | 175.1417 | 175.4956 | 191.1505 | 204.1683 | 217.5044 | 246.1328 | 222.8672 | 118.9911 | 433.6283 | 527.6639 | 451.6639 | 492.0351 | 226.2655 | 246.8761 | 230.1767 | 170.5928 | 98.9734 | 75.3539 | 105.23 | 164.5133 | 368.1239 | 391.2478 | 526.6461 | 290.4867 | 208.0089 | 199.2211 | 210.3628 | 187.7611 | 260.7345 | 305.7345 | 449.1683 | 268.1505 | 198.6194 | 259.9205 | 571.0089 | 267.5133 | 364.5311 | 314.2567 | 432.6372 | 392.8583 | 220.3539 | 274.3983 | 480.0355 | 336.3806 | 299.7345 | 317.8406 | 284 | 502.8672 | 248.2655 | 281.1772 | 298.7522 | 372.0266 | 220.6989 | 260.2211 | 300.2034 | 201.1061 | 286.345 | 315.1328 | 388.1328 | 206.8761 | 157.6106 | 208.0089 | 207.6283 | 206.8939 | 455.1945 | 529.0089 | 587.3894 | 776.2122 | 532.345 | 606.69 | 725 | 834.9556 | 186.9911 | 97.8583 | 128.2122 | 186.9911 | 172.77 | 147.3273 | 184.1239 | 140.3539 | 165.8761 | 174.5133 | 200.2034 | 167.9911 | 99.1328 | 117.7611 | 102.115 | 117.7611 | 152.885 | 155.2567 | 208.115 | 118.5133 | 414.5222 | 359.6017 | 359.4689 | 346.6283 | 281.8228 | 379.7256 | 496.1328 | 363.3894 | 704.4778 | 477.0533 | 587.6372 | 720.655
2270 | 26.9644 | 338.442 | 360.116 | 383 | 299.652 | 224.5536 | 241.9016 | 264.6828 | 219.3304 | 222.7812 | 164.4376 | 202.4552 | 193.4508 | 341.3129 | 262.21 | 274.6652 | 324.442 | 275.326 | 213.7724 | 240.6784 | 249 | 222.1072 | 560.9191 | 595.4376 | 214.3523 | 367.7724 | 393.7856 | 417.6477 | 478.3392 | 161.8972 | 182.1072 | 165.5712 | 176.1291 | 210.8928 | 235.4508 | 215.7812 | 124.2188 | 435.7768 | 468.9235 | 487.094 | 506.8752 | 208.4552 | 236.3392 | 252.2056 | 191.8752 | 89.7812 | 100.5404 | 97.6696 | 142.9016 | 358.2276 | 401.6608 | 555.652 | 284.558 | 203.6696 | 225.3172 | 200.116 | 193.442 | 259.8884 | 279.1248 | 414.0219 | 241.6828 | 200.6652 | 245.4552 | 552.4552 | 245.0132 | 360.8928 | 325.6608 | 436.6652 | 414.9868 | 229.5492 | 234.3567 | 447.5755 | 330.7812 | 290.8928 | 321.9956 | 316.8709 | 497.558 | 220.6828 | 276.3392 | 322.9868 | 365.2276 | 211.4464 | 228.5712 | 302.1072 | 218.5448 | 288.3304 | 300.0088 | 371.232 | 199.8928 | 179.9868 | 184.1248 | 207.1116 | 219.6608 | 456.9956 | 507.79 | 570.3436 | 773.7768 | 536.1072 | 618.6564 | 678.8031 | 873.5317 | 169.1204 | 107.5492 | 129.3304 | 169.1204 | 174.1116 | 142 | 184 | 165.5404 | 146.4552 | 173.3348 | 211.8796 | 180.326 | 92.0044 | 91.4596 | 116.9912 | 118.1116 | 150.4464 | 165.7724 | 202.558 | 118.2232 | 376.1335 | 358.7768 | 392.4245 | 371.8752 | 247.5712 | 371.6696 | 455.2451 | 376.5492 | 737.5361 | 496.326 | 534.8075 | 770.1969
2271 | 26.9754 | 312.8799 | 382.418 | 377.5958 | 260.6675 | 213.291 | 259.7159 | 276.3118 | 232.6097 | 200.3834 | 161.4965 | 205.4042 | 215.418 | 328.8868 | 281.2125 | 275.9007 | 307.8868 | 258.9862 | 195.1848 | 261.418 | 236.3903 | 227.5035 | 544.0924 | 587.0924 | 217.2055 | 331.1709 | 404.6097 | 389.5751 | 503.1201 | 168.1062 | 195.6097 | 169.2055 | 172 | 227.1132 | 206.0785 | 220.4042 | 118.6952 | 438.7021 | 455.5958 | 447.7668 | 491.0855 | 222.2125 | 219.6882 | 237.9862 | 185.0924 | 108.8152 | 107.6028 | 116.8152 | 161.6166 | 320.0716 | 410.2055 | 565.3048 | 254.2772 | 181.3834 | 240.709 | 210.709 | 189.4965 | 265.4042 | 260.6882 | 393.6882 | 256.0138 | 198.2979 | 209.6744 | 532.8868 | 223.0855 | 322.1709 | 329.7021 | 424.3903 | 423.4042 | 222.8938 | 232.6028 | 433.0924 | 304.7806 | 285.4965 | 324.8014 | 342.6166 | 483.4896 | 242.2194 | 265.0924 | 280.9654 | 376.6097 | 211 | 225.9007 | 307.5035 | 220.0993 | 284.4965 | 285.3903 | 338.3765 | 207.1062 | 178.4965 | 187.3048 | 212.4042 | 215.5958 | 425.5751 | 520.3118 | 533.7737 | 772.1986 | 533.3972 | 635.4111 | 689.2125 | 891.6097 | 184.1132 | 106.2979 | 121.8938 | 184.1132 | 167.6952 | 156.4111 | 198.4111 | 151.8868 | 172.8222 | 163.9931 | 238.3187 | 142.3695 | 87.3972 | 97.9076 | 119 | 114.3972 | 146.3972 | 154.3903 | 189.3903 | 101.7875 | 386.3118 | 362.6028 | 395.1986 | 383.1062 | 302.545 | 356.5889 | 468.9145 | 357.2841 | 681.6536 | 500.7021 | 557.7228 | 756.1848
2272 | 26.9864 | 325.5207 | 385 | 360.5663 | 229.5235 | 231.1727 | 277.5207 | 235.0898 | 244.0428 | 220.8246 | 156.4351 | 188.6533 | 227.1298 | 305.0884 | 230.047 | 275.087 | 285.0014 | 273.4337 | 196.6519 | 236.6105 | 245.9558 | 215.2182 | 589.5622 | 616.1285 | 228.9558 | 386.3439 | 395.9572 | 423.4323 | 489.5663 | 150.7403 | 154.0898 | 152.6533 | 180.2169 | 179.6989 | 213.9558 | 186.3066 | 119.826 | 393.3508 | 475.0856 | 471.3025 | 531.9102 | 182.9157 | 232.6077 | 229.6091 | 190.3909 | 88.1754 | 122.6077 | 99.8273 | 159.4351 | 367.1271 | 420.1298 | 493.8743 | 238.2182 | 182.6519 | 206.3936 | 212 | 180.7831 | 247.7403 | 314.692 | 387.4351 | 262.5649 | 178.8273 | 207.826 | 500.8715 | 225.5649 | 317.087 | 340.0428 | 423 | 387.4807 | 219.261 | 245.7818 | 473.9972 | 333.0414 | 305.0856 | 334.1298 | 329.4793 | 523.0843 | 244.087 | 271.3039 | 294.2597 | 338.7417 | 220.1298 | 234.2169 | 291.5663 | 223.6519 | 277.6091 | 295.8688 | 375.1713 | 202.5221 | 171.6091 | 209.9116 | 208.4351 | 194.9144 | 480.4309 | 528.3909 | 558.3025 | 787.5207 | 535.739 | 561.2224 | 711.9986 | 820.8743 | 177.7831 | 132.4765 | 110.9572 | 177.7831 | 184.3467 | 161.6519 | 166.2196 | 134.4793 | 157.7403 | 149.3052 | 227.3052 | 164.4765 | 107.9986 | 91.6961 | 115.3481 | 108.5221 | 145.087 | 146.6091 | 183.4351 | 119.1727 | 371.5663 | 358.4351 | 395.913 | 371.2182 | 319.0428 | 332.1754 | 468.261 | 353.174 | 643.0456 | 543.9102 | 544.5663 | 747.6091
2273 | 26.9974 | 326.0747 | 360.9429 | 409.8901 | 284.367 | 225.5978 | 268.822 | 268.9363 | 234.822 | 212.822 | 173.5802 | 212.9077 | 207.644 | 282.644 | 224.0747 | 274.0747 | 299.6549 | 237.0637 | 199.7758 | 223.822 | 246.0747 | 244.5341 | 593.0747 | 562.5583 | 238.3275 | 376.2703 | 384.822 | 397.3912 | 429.7077 | 143.4484 | 163.8791 | 153.7758 | 181.9253 | 212.011 | 220.5516 | 208.9077 | 119.0747 | 391.7758 | 484.4022 | 460.1209 | 502.6901 | 211.3846 | 223.822 | 214.1956 | 173.4198 | 104.5055 | 99.9429 | 108.178 | 170.1033 | 343.3165 | 392.3165 | 516.6088 | 228.6725 | 201.5055 | 216.8791 | 215.7011 | 176.2989 | 283.011 | 302.4198 | 417.5341 | 229.6901 | 197.356 | 223.7297 | 471.167 | 237.1033 | 325.3275 | 328.0462 | 411.8967 | 381.2242 | 203.2703 | 243.2989 | 440.0637 | 339.7011 | 294.9714 | 322.0462 | 352.9824 | 466.8572 | 233.822 | 269.2242 | 274.7187 | 357.2066 | 225.6264 | 258.1319 | 282.5978 | 222.1495 | 289.9538 | 301.6264 | 390.1033 | 199.2242 | 188.5802 | 197.1956 | 202.4484 | 192.0747 | 468.4198 | 487.3626 | 586.9077 | 756.6154 | 523.0462 | 554.9253 | 652.9319 | 767.7363 | 165.8967 | 99.8396 | 114.6264 | 165.8967 | 177.6725 | 165.7011 | 218.5165 | 172.7868 | 160.6264 | 167.4308 | 225.0747 | 144.7934 | 92.4198 | 91.9253 | 133.5055 | 111.7011 | 134.822 | 146 | 192.2527 | 121.9253 | 386.6549 | 369.1033 | 371.9429 | 352.4198 | 265.4088 | 372.5626 | 449.4945 | 346.5231 | 668.6835 | 497.1099 | 523.5692 | 667.4264
2274 | 27.0084 | 304.436 | 373.0635 | 379.3101 | 266.4984 | 238.1259 | 268.9376 | 321.691 | 226.4995 | 212 | 169.3746 | 184.0603 | 224.7513 | 284.7503 | 240.8762 | 237.4349 | 310.3757 | 247.1259 | 207.5005 | 237.0635 | 259.1259 | 225.436 | 582.6868 | 592.6899 | 242.7503 | 428.4412 | 384.9376 | 429.6899 | 476.5661 | 173.0021 | 173.4381 | 185.8772 | 190.4381 | 200.9365 | 208.8116 | 239.127 | 128.3757 | 420.127 | 505.6264 | 469.3132 | 508.4381 | 202.7492 | 224.8751 | 243.9397 | 172 | 102.2497 | 95.1873 | 128.6889 | 158.8116 | 383.1905 | 412.5016 | 497.436 | 229.8751 | 180.4984 | 237.6889 | 221.6254 | 158.1863 | 250.3725 | 281.3111 | 406.8741 | 230.7503 | 199 | 214.6868 | 459.6243 | 199.5598 | 341.9386 | 287.6222 | 437.2518 | 365.0614 | 215.1259 | 265.5016 | 451.0635 | 326.8741 | 305.2508 | 311.6243 | 337.1863 | 492.0021 | 262.0645 | 264.3122 | 288.001 | 341.1863 | 216.6243 | 243.1238 | 279.1873 | 220.1249 | 311.6264 | 284.1863 | 363.8106 | 191.4995 | 166.5609 | 168.8106 | 237.6275 | 210.7513 | 484.8137 | 535.5661 | 573.0614 | 722.1228 | 548.2518 | 588.7524 | 703.3164 | 838.0677 | 210.0031 | 101.6878 | 131.8762 | 210.0031 | 161.999 | 136.9355 | 160.1831 | 178.8127 | 164.7503 | 160.5619 | 218.437 | 165.5016 | 90.0624 | 98.563 | 125.6243 | 119.5005 | 147.1259 | 132.8741 | 197.6878 | 106.0614 | 365.4984 | 360.6243 | 390.6264 | 365.0635 | 314.4412 | 367.5619 | 442.3746 | 354.4381 | 630.6847 | 510.8137 | 533.2508 | 736.9428
2275 | 27.0194 | 306.7994 | 383.4986 | 383.649 | 250.7522 | 220.0029 | 256.6519 | 268.9585 | 218.4011 | 189.2034 | 161.4011 | 231.3925 | 223.1504 | 267.9026 | 248.649 | 245.4484 | 298.6519 | 247.0501 | 211.7994 | 239.8997 | 231.5043 | 203.1032 | 573.4513 | 547.5072 | 228.7522 | 376.9083 | 368.8524 | 415.8524 | 459.1032 | 176.8997 | 172.1003 | 186.1003 | 173.9026 | 237.0444 | 209.8997 | 201.106 | 122.351 | 445.7464 | 427.212 | 436.755 | 506.1504 | 199.1504 | 224.0501 | 246.9499 | 162.5014 | 107.6991 | 92.1504 | 111.0029 | 160.8496 | 315.7106 | 388.3539 | 499.7994 | 223.351 | 188.4986 | 225.702 | 207.7522 | 139.9026 | 248 | 265.7522 | 362.3066 | 222.4513 | 186.6519 | 256.7435 | 440.9527 | 196.0501 | 294.5573 | 322.9943 | 405.755 | 391.5458 | 215.0501 | 237.5544 | 451.0501 | 285.1562 | 301.2507 | 301.5014 | 335.0501 | 474.053 | 228.8553 | 244.053 | 313.6963 | 366.596 | 203.6519 | 228.702 | 268.5516 | 219.0501 | 305.4011 | 281.1003 | 397.1447 | 186.2507 | 180.1977 | 193.596 | 215.3037 | 188.2536 | 443.2565 | 534.2507 | 549.2034 | 690.5544 | 550.9499 | 600.4986 | 724.0974 | 771.7608 | 173.106 | 114.3481 | 118.7522 | 173.106 | 135.3539 | 145.4484 | 199.6934 | 139.106 | 179.2478 | 158.1003 | 191.404 | 153.702 | 97.5989 | 95.2006 | 109.8023 | 114.3009 | 129.9527 | 148.1476 | 205.5989 | 104.0501 | 376.3481 | 345.7522 | 364.4542 | 357.4513 | 260.0587 | 345.1533 | 450.5487 | 336.0029 | 670.7435 | 486.3539 | 565.3453 | 728.702
2276 | 27.0304 | 304.1136 | 371.4921 | 364.757 | 266.3565 | 211.3028 | 247.3407 | 282.3565 | 215.1136 | 174.5299 | 156.1893 | 180.1197 | 202.7949 | 318.956 | 239.3785 | 257.5458 | 312.4322 | 263.3565 | 230.2808 | 229.4164 | 227.1136 | 221.243 | 539.3248 | 558.4701 | 219.3407 | 357.6435 | 373.7729 | 427.5079 | 482.0537 | 169.3028 | 171.0379 | 165.7949 | 199.9401 | 189.9305 | 235.0158 | 240.3724 | 123.9243 | 397.9305 | 470.1452 | 472.5238 | 488.6813 | 202.8486 | 213.4164 | 223.9085 | 175.4701 | 82.9842 | 116.0537 | 115.7729 | 162.9243 | 339.9023 | 394.6972 | 480.757 | 236.4701 | 199.5836 | 255.7887 | 209.8864 | 163.0537 | 227.7949 | 298.6752 | 401.3724 | 243.1672 | 170.6056 | 197.4225 | 490.0317 | 199.8486 | 271.7949 | 315.3785 | 379.9463 | 353.5519 | 258.2966 | 244.6593 | 399.044 | 297.4322 | 305.8107 | 319.2808 | 314.7949 | 510.5238 | 247.2051 | 275.713 | 312.1136 | 340.0977 | 222.243 | 233.7729 | 313.2209 | 211.3028 | 259.7791 | 262.7192 | 375.9085 | 175.4164 | 175.2271 | 205.5836 | 204.3785 | 186.0379 | 425.6056 | 504.1734 | 573.0158 | 738.0695 | 543.3028 | 550.9683 | 665.3468 | 817.0695 | 183.5079 | 89.022 | 126.6593 | 183.5079 | 137.8486 | 160.4322 | 165.4384 | 128.3407 | 157.8706 | 155.1136 | 207.3187 | 150.1136 | 99.9243 | 104.6215 | 111.8864 | 119.7729 | 151.1294 | 137.4542 | 162.7034 | 95.3407 | 379.8864 | 330.5678 | 385.1294 | 337.757 | 270.4701 | 365.1672 | 438.4921 | 360.0158 | 648.9463 | 468.6435 | 580.4701 | 636.596
2277 | 27.0414 | 284.5112 | 337.8691 | 349.3834 | 268.9489 | 242.182 | 238.2301 | 263.5112 | 220.8466 | 182.7699 | 143.3323 | 217.952 | 191.2812 | 274.2269 | 224.3834 | 258 | 293.5112 | 261.0767 | 220.2812 | 261.1565 | 238.6933 | 217.1278 | 555.5399 | 542.4345 | 229.7188 | 357 | 372.0511 | 433.8466 | 453.7668 | 144.639 | 179.7699 | 181.5655 | 190.2812 | 204.5655 | 204.818 | 226.409 | 117.1789 | 435.952 | 419.3803 | 471.0767 | 504.5655 | 194.2301 | 226.6421 | 256.1309 | 171.1278 | 108.3098 | 104.3323 | 108.2045 | 139.6135 | 326.3834 | 357.9713 | 529.6964 | 241.8722 | 209.7444 | 204.3803 | 233.3865 | 163.0256 | 241.6166 | 282.4601 | 352.3292 | 245.9489 | 196.3098 | 200.8466 | 425.7382 | 220.4632 | 327.5174 | 317.9233 | 415.0542 | 377.3354 | 213.2269 | 249.8722 | 411.6166 | 329.182 | 299.1789 | 322.9233 | 316.9233 | 499.3323 | 256.7699 | 266.2812 | 298.3579 | 375.0542 | 205.4601 | 225.2301 | 281.8691 | 222.6933 | 244.3579 | 294.1565 | 357.4601 | 197.4121 | 175.9744 | 175.7924 | 227.3865 | 180.1534 | 439.6166 | 531.2587 | 614.9264 | 719.5368 | 555.6677 | 570.4377 | 657.1534 | 830.6933 | 177.1789 | 85.0767 | 136.7444 | 177.1789 | 133.1278 | 142.4857 | 179.591 | 154.3098 | 161.8722 | 148.1789 | 196.3067 | 148.0511 | 101.9489 | 106.9489 | 101.2812 | 113.1789 | 178.3098 | 148.6933 | 181.4632 | 92.0767 | 361.4857 | 341.6933 | 346.048 | 356.4888 | 261.2556 | 327.0225 | 469.182 | 351.2556 | 655.7955 | 505.0287 | 556.639 | 688.543
2278 | 27.0524 | 286.9602 | 354.7611 | 356.8938 | 271.9602 | 216.3583 | 279.4427 | 271.8806 | 199.2919 | 166.2256 | 160.7611 | 173.6104 | 194.9469 | 283.854 | 259.5223 | 262.9336 | 230.836 | 229.4246 | 208.1592 | 233.3848 | 226.1725 | 205.1592 | 543.1725 | 570.6152 | 212.2389 | 355.0265 | 330.5573 | 408.345 | 461.8806 | 143.0133 | 185.9204 | 157.3318 | 164.345 | 192.1725 | 210.9071 | 218.1062 | 123.9071 | 395.5573 | 444.6417 | 466.0664 | 479.345 | 196.9602 | 209.2389 | 216.5441 | 185.8009 | 105.0531 | 96.1062 | 128.7213 | 171.5621 | 324.0265 | 381.6682 | 505.345 | 277.5223 | 192.2389 | 201.0265 | 224.1327 | 187.6682 | 232.1327 | 270.1592 | 370.7346 | 257.8408 | 200.9469 | 175.345 | 463.4692 | 196.3318 | 283.6104 | 275.5706 | 397.2521 | 366.1592 | 236.6682 | 239.146 | 423.8408 | 299.4114 | 312.8142 | 333.854 | 320.9469 | 449.6635 | 246.146 | 260.0796 | 269.3848 | 350.345 | 204.0133 | 200.3318 | 280.0133 | 209.1858 | 271.6284 | 306.8408 | 329.3716 | 197.0133 | 175.0133 | 159.2123 | 201.3583 | 189.8673 | 425.1991 | 504.3716 | 553.836 | 681.5043 | 573.7611 | 552.2521 | 615.5573 | 769.8228 | 145.4246 | 102.7611 | 122.1991 | 145.4246 | 144.8408 | 157.7877 | 175.0664 | 153.0265 | 152.1327 | 172.6682 | 192.0531 | 161.8142 | 90.1592 | 109.9602 | 123.6948 | 118.9204 | 131.637 | 149 | 170.1592 | 95.9469 | 380.7346 | 377.5223 | 358.8142 | 324.4379 | 257.0531 | 362.509 | 422.637 | 350.0133 | 643.1725 | 493.1725 | 552.0531 | 635.7299
2279 | 27.0634 | 297.9892 | 327.0274 | 362.9941 | 260.0117 | 225.9902 | 219.0597 | 274.9971 | 193.0059 | 204.9618 | 159.002 | 181.9912 | 169.0255 | 276.0078 | 239.0206 | 243.0196 | 318.9129 | 233.9951 | 198.0098 | 233.999 | 245.9804 | 201.0039 | 530.0127 | 580.9902 | 243.9687 | 332.0225 | 380.9501 | 401.0069 | 439.0225 | 158.9843 | 184.002 | 158.998 | 178.9853 | 186.0059 | 206.0049 | 223.9941 | 124.999 | 379.0157 | 442.0029 | 467.998 | 510.9687 | 203.9931 | 210.998 | 242.9736 | 171.0147 | 110.9941 | 89.0069 | 100.0284 | 135.0362 | 326.9971 | 382.999 | 470.0343 | 236.0411 | 189.0029 | 206.9941 | 216.0078 | 148.0392 | 227.0049 | 291.9785 | 372.998 | 209.048 | 191.0098 | 202.9726 | 451.0127 | 204.9912 | 298.9843 | 316.9589 | 386.0108 | 329.0362 | 218.0186 | 243.9951 | 421.0029 | 302.9961 | 276.0362 | 307.0264 | 356.9648 | 428.0206 | 239.0069 | 240.0196 | 291.9775 | 320.0294 | 188.0157 | 231.9687 | 290.9892 | 210.998 | 257.0147 | 286.0206 | 359.9696 | 191.0059 | 169.0059 | 207.952 | 201.999 | 190.999 | 472.953 | 508.9951 | 588.9648 | 685.9951 | 450.1214 | 536.0157 | 653.9618 | 755.0137 | 164.9804 | 88.0147 | 120.002 | 164.9804 | 160.9843 | 160.9971 | 167.0078 | 138.0147 | 159.9922 | 175.9971 | 194.9971 | 157.0049 | 93.9961 | 93.0166 | 111.0127 | 113.0059 | 144.9863 | 147.002 | 165.0049 | 104.9912 | 357.0235 | 304.0725 | 397.9618 | 340.9834 | 272.9843 | 350.0127 | 430.9912 | 366.9834 | 634.0088 | 469.0235 | 576.9755 | 685.9501
2280 | 27.0744 | 281.0452 | 336.328 | 374.8529 | 221.2715 | 219.7511 | 231.3959 | 262.362 | 217.0566 | 211.672 | 156.0792 | 200.8529 | 191.8529 | 323.8303 | 230.0452 | 214.7013 | 230.7126 | 233.509 | 181.0226 | 221.1018 | 220.9208 | 214.8529 | 548.8416 | 582.1923 | 216.1357 | 353.6606 | 364.9774 | 344.6108 | 439.5362 | 168.2036 | 174.7851 | 178.0113 | 170.9548 | 202.3054 | 185.0113 | 219.8643 | 108.1018 | 364.3167 | 451.7625 | 483.2987 | 488.9548 | 192.7398 | 196.0339 | 238.8869 | 158.9321 | 100.0226 | 114.5588 | 103.1923 | 157.9208 | 307.2149 | 322.509 | 473.1584 | 256.819 | 182.9321 | 209.1018 | 214.1244 | 168.7172 | 216.1357 | 279.7964 | 410.2987 | 204.0679 | 187.1584 | 206.4412 | 411.2149 | 186.0339 | 280.2602 | 323.2375 | 415.8756 | 308.2828 | 209.7964 | 280.3552 | 457.0679 | 284.8982 | 288.7964 | 293.6222 | 305.0679 | 402.9887 | 231.1244 | 261.0679 | 318.6267 | 371.638 | 188.3167 | 207.2036 | 270.7398 | 206.181 | 269.9661 | 258.0566 | 350.0679 | 189.7625 | 192.638 | 188.7625 | 219.0679 | 195.7738 | 450.7511 | 448.7013 | 577.7625 | 665.69 | 578.6652 | 590.457 | 646.1584 | 735.3733 | 158.0679 | 111.8756 | 128.8869 | 158.0679 | 138 | 135.4525 | 195.8756 | 157.8643 | 185.6154 | 141.2262 | 187.9548 | 137.8869 | 88.0113 | 120.5588 | 114.9548 | 120.1131 | 135.181 | 147.0226 | 201.6267 | 98.9321 | 337.362 | 328.9208 | 372.0792 | 325.181 | 227.4864 | 333.0905 | 392.4185 | 338.4072 | 616.0566 | 496.6154 | 520.4977 | 602.4299
2281 | 27.0854 | 284.0559 | 365.0944 | 362.6137 | 244.4807 | 198.1416 | 264.9379 | 292.9379 | 221.6696 | 183.3068 | 162.6932 | 187.1739 | 178.9056 | 308.3863 | 234.1652 | 275.882 | 292.8348 | 276.8907 | 183.8025 | 229.5279 | 213.8348 | 202.3776 | 534.1975 | 598.5515 | 227.4099 | 324.4249 | 362.9056 | 398.236 | 400.7466 | 184.9851 | 156.3068 | 179.3068 | 167.354 | 228.0559 | 186.0472 | 208.9441 | 116.9764 | 392.118 | 431.6137 | 422.6609 | 483.7726 | 170.0472 | 199.236 | 229.0708 | 153.236 | 102.2124 | 76.2596 | 119.5987 | 151.4485 | 326.0236 | 366.764 | 486.3627 | 240.6932 | 176.5751 | 217.5751 | 225 | 144 | 228.0944 | 262.6137 | 349.3304 | 210.6845 | 201.0472 | 244.2211 | 429.764 | 189.4013 | 302.1975 | 343.9292 | 404.2447 | 332.9528 | 192.3304 | 224.0944 | 461.9615 | 275.5279 | 271.6137 | 346.3006 | 311.1652 | 402.3776 | 242 | 266.5043 | 286.6609 | 340.0708 | 215.4335 | 224.9764 | 248.9913 | 221.5043 | 266.3627 | 263.4485 | 356.1888 | 169 | 160.882 | 168.5429 | 224.1739 | 176.236 | 428.5987 | 508.3714 | 557.354 | 723.8049 | 462.8969 | 543.8733 | 661.369 | 766.9615 | 164.118 | 100.9764 | 118.764 | 164.118 | 137.9056 | 173.7254 | 184.8584 | 146.2124 | 151.9528 | 160.7876 | 184.7081 | 128.6609 | 88.882 | 82.4721 | 110.9292 | 130 | 150.6932 | 148.9292 | 168.9764 | 92.7404 | 368.5279 | 321.5751 | 379.1888 | 341.7789 | 269.882 | 340.3863 | 428.8584 | 374.2124 | 619.3242 | 463.6373 | 563.4807 | 639.764
2282 | 27.0964 | 246.7588 | 367.5283 | 386.7796 | 224.5434 | 204 | 221.8256 | 250.005 | 207.4257 | 195.3898 | 151.7947 | 154.364 | 175.7538 | 282.8564 | 240.8923 | 270.4975 | 284.9591 | 231.1795 | 216.5334 | 210.5384 | 207.0718 | 218.4666 | 500.2821 | 579.7487 | 202.9641 | 341.6052 | 358.9641 | 407.5334 | 471.313 | 142.7846 | 168.5693 | 191.5334 | 180.8514 | 190.3281 | 187.7487 | 247.7128 | 115.7128 | 396.5334 | 456.5693 | 450.3589 | 433.2154 | 172.6102 | 209.8256 | 232.2154 | 163.2872 | 111.0718 | 87.1436 | 103.2154 | 168.6719 | 327.3948 | 356.7128 | 460.6102 | 227.7846 | 158.5334 | 200.3589 | 225.1795 | 144.8974 | 232.3589 | 287.2462 | 364.2204 | 238.7487 | 201.995 | 212.5384 | 420.1436 | 207.364 | 270.2563 | 339.7796 | 374.6152 | 330.2821 | 205.6411 | 227.8205 | 419.7897 | 256.9691 | 296.4616 | 276.8256 | 318.2872 | 418.7179 | 240.5283 | 244.5642 | 313.4257 | 341.313 | 193.4717 | 224.2154 | 289.3898 | 200.677 | 242.1536 | 282.2872 | 362.2771 | 169.2154 | 156.2154 | 191.8974 | 190.7538 | 185.9282 | 412.323 | 440.7487 | 570.636 | 635.0101 | 499.0309 | 577.9181 | 715.5233 | 723.7487 | 168.2821 | 100.0359 | 109.5743 | 168.2821 | 134.1077 | 121.2872 | 178.6052 | 153.8873 | 149.8564 | 151.8923 | 212.8155 | 155.2103 | 84 | 101.677 | 107.3539 | 129.5693 | 138.0359 | 145.6052 | 168.2154 | 82.3948 | 349 | 304.8974 | 385.9591 | 372.2412 | 263.7078 | 315.8256 | 424.4717 | 381.0617 | 550.4666 | 489.1026 | 540.9591 | 628.7437
2283 | 27.1074 | 292.3135 | 327.5181 | 355.2528 | 264.6508 | 205.301 | 243.3736 | 277.4819 | 192.4337 | 178.1809 | 199.1327 | 189.8917 | 194.9399 | 278.0845 | 236.6508 | 257.3373 | 257.8191 | 235.8073 | 203.7591 | 225.53 | 209.0482 | 231.4337 | 484.3248 | 574.0119 | 202.1446 | 331.3855 | 357.6145 | 395.4337 | 424.8917 | 137.6746 | 157.4819 | 178.1809 | 150.2891 | 224.169 | 182.0601 | 239.9036 | 108.2409 | 383.7109 | 443.8917 | 459.229 | 440.301 | 190.301 | 231.47 | 238.53 | 169.7954 | 112.0363 | 91.3855 | 109.9155 | 133.0482 | 336.4581 | 347.2172 | 475.6508 | 222.1927 | 148.1683 | 209.47 | 228.2653 | 168.3254 | 240.9881 | 265.699 | 396.1809 | 231.8554 | 175.7228 | 225.6026 | 424.4819 | 241.5907 | 303.5663 | 307.6264 | 417.0845 | 312.0601 | 196.6264 | 224.2046 | 439.9399 | 282.2772 | 281.47 | 298.229 | 327.7347 | 437.47 | 201.1446 | 207.2046 | 297.3254 | 298.072 | 230.8317 | 228.0726 | 274.1564 | 192.0964 | 297.5426 | 289.4218 | 317.9756 | 175.5782 | 161.8073 | 214.3135 | 210.9518 | 184.2409 | 419.9399 | 436.7947 | 536.2647 | 687.0726 | 471.6026 | 523.1564 | 649.0964 | 715.8917 | 149.0482 | 101.1446 | 124.3254 | 149.0482 | 138.2046 | 130.0601 | 168 | 125.0601 | 145.7109 | 149.1927 | 181.8191 | 134.5782 | 84.3373 | 93.6264 | 90.9155 | 118 | 139.3855 | 135.3373 | 174.8673 | 92.9036 | 349.1927 | 328.9518 | 359.3974 | 327.0238 | 231.3129 | 336.8917 | 461.8317 | 328.8554 | 564.0601 | 463.8436 | 514.2528 | 596.4937
2284 | 27.1184 | 258.3952 | 318.6652 | 377.9438 | 238.6048 | 227.19 | 229.4276 | 285.1252 | 199.6695 | 164.1771 | 181.7581 | 170.2376 | 175.8748 | 258.1857 | 211.9957 | 262.6695 | 274.2419 | 231.3348 | 200.6329 | 235.4557 | 210.4838 | 237.3995 | 548.0648 | 591.2181 | 204.8186 | 337.4276 | 348.2462 | 404.5443 | 402.9395 | 150.5767 | 165.609 | 163.27 | 157.149 | 187.2138 | 201.4881 | 237.2138 | 113.4233 | 377.6371 | 422.4838 | 444.2419 | 463.5205 | 215.2138 | 218.8833 | 247.2462 | 147.9957 | 93.7257 | 98.4557 | 126.3067 | 134.6652 | 307.0281 | 314.1167 | 449.2419 | 224.6695 | 190.9395 | 199.0605 | 195.2095 | 153.6695 | 221.9071 | 241.2095 | 379.4557 | 228.5767 | 188.7905 | 199.27 | 436.7819 | 195.2095 | 296.0281 | 315.4643 | 400.2095 | 331.1252 | 209.7257 | 249.149 | 418.0324 | 270.0562 | 273.2376 | 283.8467 | 359.8229 | 426.8186 | 205.27 | 230.6371 | 286.6005 | 336.581 | 189.9957 | 192.419 | 295.609 | 195.2095 | 250.3305 | 278.6048 | 356.3348 | 187 | 159.0886 | 181.6652 | 210.9071 | 189.7862 | 400.2095 | 489.0367 | 580.2138 | 649.3629 | 443.0929 | 548.54 | 654.0843 | 693.5767 | 150.7257 | 104 | 112.2095 | 150.7257 | 161.6976 | 150.9395 | 166.7905 | 145.2138 | 142.2376 | 152.879 | 197.3348 | 145.879 | 90.5162 | 105.9395 | 107.9719 | 117.6976 | 147 | 143.0281 | 191.879 | 90.5162 | 353.4838 | 328.0605 | 386.6371 | 363.0691 | 275.1857 | 313.9719 | 422.7495 | 327.0886 | 585.3024 | 440.5767 | 538.1533 | 625.0929
2285 | 27.1294 | 248.8545 | 330.3826 | 344.8357 | 249.3826 | 169.5281 | 206.6736 | 256.6549 | 180.0187 | 197.054 | 177.054 | 203.1081 | 202.0165 | 233.2018 | 240.3076 | 242.4366 | 276.9085 | 222.237 | 223.7442 | 228.237 | 217.4179 | 196.4179 | 516.3991 | 548 | 203.8192 | 312.5634 | 323.1102 | 409.5799 | 403.6736 | 144.2183 | 145.0187 | 183.2183 | 174.9272 | 177.2558 | 179.6009 | 223.6902 | 118.3991 | 373.1643 | 430.946 | 451.7111 | 427.183 | 202.5634 | 190.4388 | 219.5634 | 176.2348 | 103.8357 | 91.0187 | 100.2183 | 143.9813 | 320.817 | 347.7277 | 455.1102 | 206.9835 | 191.0187 | 200.946 | 214.8732 | 134.8192 | 234.3991 | 258.0353 | 368.817 | 219.8898 | 171.7464 | 217.4719 | 475.5612 | 214.1455 | 308.2889 | 247.9107 | 419.5821 | 303.8919 | 219.1081 | 265.6174 | 403.2911 | 300.6902 | 302.9063 | 297.6549 | 326.3638 | 421.6715 | 224.2723 | 226.6736 | 324.1621 | 299.6549 | 219.5447 | 227.3076 | 275.4553 | 213.7089 | 267.5799 | 287.4906 | 346.2911 | 186.2723 | 174.0353 | 192.2183 | 224.9272 | 201.1268 | 419.2183 | 444.6736 | 568.1455 | 649.9063 | 429.8004 | 585.3076 | 698.2889 | 688.4553 | 162.5094 | 101.6715 | 130.4179 | 162.5094 | 155.2536 | 148.3264 | 149.4553 | 134.8919 | 175.8357 | 152.1643 | 186.3451 | 144.4366 | 83.8732 | 104.7817 | 91.3451 | 111.763 | 148.8919 | 156.6715 | 187.8898 | 83.5821 | 360.5634 | 328.1268 | 380.4179 | 303.6009 | 248.0187 | 300.4013 | 477.2161 | 343.2723 | 589.8545 | 437.8567 | 525.2183 | 607.2889
2286 | 27.1404 | 246.1494 | 349.7862 | 334.8483 | 265.554 | 189.5103 | 229.1058 | 264.4897 | 191.8092 | 184.4046 | 165.1701 | 177.3839 | 166.1265 | 273.4046 | 210.423 | 245.7034 | 265.977 | 236.8092 | 193.8092 | 242.8092 | 212.7219 | 190.3609 | 495.8299 | 549.0207 | 224.7885 | 304.5333 | 350.9356 | 368.2115 | 423.2138 | 148.0207 | 156.0437 | 186.5103 | 173.7448 | 204.554 | 198.6391 | 209.4667 | 99.2759 | 385.9586 | 444.0207 | 497.3839 | 454.1494 | 195.3839 | 242.9172 | 217.9127 | 141.8506 | 90.6161 | 103.6598 | 104.1058 | 132.3609 | 294.9563 | 354.2781 | 482.3402 | 242.7885 | 204.2552 | 211.8942 | 225.5747 | 159.3609 | 215.8713 | 231.7241 | 343.977 | 195.3816 | 194.2552 | 196.7241 | 416.7885 | 217.446 | 266.1886 | 286.7862 | 426.6598 | 327.5747 | 193.7241 | 246.3839 | 410.8276 | 283.4046 | 240.3816 | 303.7034 | 331.0851 | 392.9356 | 215.6805 | 250.2115 | 290.6368 | 311.0621 | 203.3816 | 193.5103 | 291.2781 | 212.977 | 223.9149 | 279.3839 | 349.9149 | 178.3609 | 150.6161 | 193.554 | 225.7862 | 189.0644 | 422.9356 | 469.3173 | 569.1494 | 588.8276 | 441.1908 | 551.9356 | 650.4046 | 705.2138 | 169.1701 | 74.6368 | 123.4253 | 169.1701 | 133.2552 | 127.7655 | 171.1472 | 158.3195 | 162.0207 | 166.4897 | 179.1058 | 150.9356 | 95.7655 | 102.5954 | 84.1908 | 98.6368 | 171.3839 | 131.6782 | 162.9563 | 92.0207 | 352.6184 | 317.5103 | 374.6161 | 323.1494 | 258.7034 | 331.5954 | 429.5103 | 334.4253 | 588.8506 | 481.0459 | 528.0851 | 564.2529
2287 | 27.1514 | 239.4337 | 365.7875 | 384.761 | 246.3007 | 194.3186 | 240.8053 | 259.1947 | 180.1682 | 178.8761 | 165.8318 | 163.6993 | 188.3186 | 267.292 | 254.3007 | 223.8231 | 298 | 228.0887 | 180.708 | 229.8053 | 245.0178 | 204.2212 | 496.3364 | 558.5398 | 202.5576 | 277.9027 | 358.3716 | 397.4516 | 401.8583 | 158.8053 | 139.9648 | 191.5133 | 172.2655 | 188.708 | 186.7258 | 235.3186 | 112.708 | 365.6549 | 449.6458 | 480.8761 | 442.2742 | 179.4602 | 199.1947 | 255.2299 | 151.3894 | 105.2742 | 100.4867 | 114.6371 | 146.2212 | 316.8761 | 329.3807 | 488.4337 | 220.8496 | 204.0795 | 196.4954 | 222.7523 | 169.1325 | 235.1239 | 220.8231 | 376.3894 | 221.4867 | 193.6901 | 184.6549 | 394.1682 | 231.053 | 316.4511 | 298.5041 | 421.9292 | 325.239 | 180.7788 | 230.7523 | 446.7431 | 279.9205 | 263.4689 | 278.2212 | 332.0973 | 404.6549 | 220.1769 | 271.1769 | 316.5663 | 340.9822 | 228.1239 | 202.0178 | 264.7258 | 240.9113 | 225.4337 | 260.7345 | 349.5841 | 191.4424 | 170.3364 | 177.9027 | 241.3007 | 179.4867 | 432.4159 | 500.4246 | 562.5311 | 631.4602 | 457.4072 | 557.5219 | 644.292 | 683.5663 | 171.7788 | 102.5133 | 126.6371 | 171.7788 | 137.1682 | 134.8318 | 201.2034 | 151.6814 | 172.1239 | 161.7788 | 189.9292 | 159.053 | 103.708 | 108.4159 | 97.5841 | 125.5398 | 153.6106 | 175.0882 | 184.3894 | 103.9735 | 328.6549 | 323.5841 | 391.5133 | 315.7523 | 234.9735 | 338.7788 | 435.9735 | 337.9292 | 593.2299 | 417.0973 | 533.186 | 610
2288 | 27.1624 | 261.8904 | 334.9036 | 365.535 | 215.9471 | 186.3554 | 238.3422 | 263.8506 | 191.3289 | 186.3422 | 153.0265 | 198 | 184.3024 | 272.0832 | 222.741 | 248.465 | 293.9433 | 260.0794 | 178.7675 | 226.5747 | 271.794 | 196.5614 | 516.465 | 547.4121 | 216.6711 | 282.8109 | 334.2193 | 443.9433 | 385.5085 | 157.8771 | 182.0529 | 189.1928 | 180.9301 | 185.4518 | 208.8204 | 226.259 | 111.4253 | 375.8469 | 401.535 | 484.2855 | 415.2458 | 192.5614 | 201.3289 | 211.6578 | 154.6711 | 82.1928 | 105.1096 | 102.4386 | 139.3289 | 322.5879 | 363.5312 | 505.4083 | 238 | 176.8904 | 166.9603 | 235.9301 | 126.6314 | 227.1096 | 250.9471 | 377.0397 | 224.6843 | 166.6181 | 198.465 | 406.6446 | 211.9036 | 263.0265 | 233.9338 | 411.3422 | 342.2722 | 192.9338 | 244.8072 | 398 | 304.3686 | 243.9471 | 271 | 332.0132 | 421.9735 | 195.535 | 240.2855 | 295.4253 | 316.7278 | 220 | 229.6711 | 287.3686 | 204.0964 | 246.2458 | 250.9735 | 352.259 | 177 | 191.3422 | 178.4253 | 209.5217 | 181.6976 | 427.3289 | 454.9301 | 585.6711 | 642.1494 | 482.5312 | 514.9036 | 646.6711 | 665.8242 | 176.4783 | 99.8639 | 112.794 | 176.4783 | 147.1229 | 128.7145 | 167.206 | 155.9168 | 168.0567 | 166.8072 | 179.5614 | 143.3024 | 99.2458 | 102.2325 | 105.5217 | 112 | 148.9036 | 115.8242 | 185.1494 | 109.4915 | 339.9433 | 328.3422 | 387.1096 | 331.4518 | 242.1361 | 348.5217 | 442.259 | 327.2325 | 549.6578 | 421.0699 | 570.9036 | 609.8904
2289 | 27.1734 | 255.635 | 324.293 | 376.293 | 250.0728 | 175.4386 | 236.0483 | 280.0499 | 195.829 | 182.9509 | 138.3412 | 193.0008 | 201.0491 | 290.2193 | 242.5614 | 242.2185 | 259.6588 | 216.1457 | 184.1465 | 214.8781 | 259.0736 | 191.4149 | 499.1227 | 563.3421 | 212.0491 | 329.2684 | 340.1457 | 412.829 | 413.2684 | 165.2439 | 154.6334 | 202.2447 | 158.5606 | 180.7561 | 184.707 | 205.9509 | 120.4394 | 334.9755 | 412.9026 | 452.7062 | 400.0245 | 187.0491 | 203.3903 | 218.829 | 154.0728 | 97.5614 | 106.8535 | 105.3903 | 140.9026 | 305.8044 | 310.8781 | 461.2193 | 236.293 | 179.1702 | 191.1219 | 209.171 | 146.171 | 231.7799 | 277.5131 | 359.3404 | 215.0974 | 195.3903 | 186.3658 | 427.3159 | 202.0245 | 245.0491 | 275.829 | 409.7799 | 310.9026 | 227.635 | 225.7807 | 400.1948 | 281.9026 | 271.1227 | 276.9746 | 322.171 | 433.2447 | 209.5851 | 209.072 | 304.5614 | 341.757 | 227.9255 | 226.6342 | 267.707 | 212.1465 | 235.6579 | 269.9264 | 336.8282 | 177 | 186.6097 | 190 | 229.5123 | 167.1457 | 429.6342 | 437.8036 | 586.2921 | 611.806 | 432.8044 | 501.3666 | 642.4149 | 698.9034 | 158.5606 | 111.3421 | 106.0483 | 158.5606 | 142.5606 | 164.9272 | 178.4632 | 141.3167 | 197.5859 | 151.1948 | 177.829 | 160.9026 | 83.7807 | 96 | 125.3903 | 113.5851 | 142.1948 | 151.8298 | 163.5851 | 81.1219 | 309.4386 | 320.9272 | 384.2201 | 328.829 | 225.4149 | 366.5614 | 419.5123 | 321.7316 | 555.7316 | 444.0491 | 555.9034 | 605.5859
2290 | 27.1844 | 220.0871 | 308.5102 | 364.5369 | 261.5102 | 196.7582 | 254.2418 | 236.4427 | 208.3289 | 196.8658 | 148.6711 | 158.7449 | 185.792 | 296.1813 | 222.1813 | 280.4231 | 255.3689 | 243.3156 | 179.4764 | 213.0604 | 236.5236 | 185.3689 | 469.6373 | 538.208 | 198.8053 | 319.9733 | 361.6782 | 421.5707 | 408 | 164.1813 | 180.6178 | 167.5973 | 178.208 | 181.5502 | 196.1947 | 221.8791 | 100.6578 | 342 | 402.0471 | 495.3627 | 396.0871 | 175.5502 | 193.6311 | 232.6711 | 168.3289 | 80.1342 | 109.1075 | 103.1476 | 135.0133 | 314.9938 | 311.8791 | 467.584 | 225.7449 | 195.96 | 188.9129 | 195.8658 | 137.0267 | 252.96 | 238.6711 | 403.9262 | 225.8187 | 190.8658 | 188.0604 | 480.5369 | 201.9796 | 233.416 | 290.4764 | 432.0338 | 303.8053 | 190.208 | 215.792 | 417.208 | 270.7787 | 239.0871 | 317.0471 | 311.1476 | 406.9191 | 218.7182 | 259.7316 | 292.1005 | 306.8187 | 277.0809 | 231.9191 | 279.7316 | 202.7787 | 267.2347 | 291.6711 | 382.376 | 179.416 | 188.9867 | 190.5369 | 224.3893 | 196.8658 | 425.6578 | 492.2551 | 612.2818 | 566.4427 | 454.0738 | 468.2684 | 636.6373 | 652.1813 | 173.9129 | 87.5502 | 126.1209 | 173.9129 | 157.1075 | 144.8995 | 188.4631 | 157.5973 | 171.6578 | 168.4764 | 189.5236 | 153.8053 | 76.6107 | 98.9529 | 120.1947 | 122.584 | 159.7449 | 131.5502 | 176.8791 | 83.6107 | 328.0738 | 306.2684 | 363.1742 | 338.376 | 220.4427 | 347.6578 | 416.5369 | 325.5236 | 567.9796 | 430.9396 | 513.3422 | 578.4498
2291 | 27.1954 | 244.3954 | 291.128 | 362.4325 | 240.4651 | 215.872 | 231.0256 | 272.4325 | 204.1465 | 195.5605 | 150.9488 | 170.8791 | 179.4651 | 276.6814 | 204.1465 | 242.1024 | 283.5534 | 225.2674 | 194.128 | 205.6814 | 227 | 216.4837 | 503.3628 | 548.4907 | 203.707 | 295.6116 | 318.1094 | 398.714 | 406.8279 | 145.5605 | 137.8349 | 167.4907 | 179.5534 | 195.3628 | 191.5861 | 235.7582 | 93.3186 | 343.0256 | 386.0768 | 463.5419 | 413.5093 | 191.8535 | 166.4722 | 238.0256 | 162.2418 | 82.3186 | 84.8791 | 113.9232 | 148.1465 | 250.4466 | 328.9815 | 455.8094 | 234.5093 | 160.9559 | 172.8094 | 196.0256 | 162.4139 | 213.8535 | 244.0256 | 389.8208 | 239.3116 | 192.3442 | 183.9047 | 489.2745 | 240.8464 | 255.1536 | 299.4139 | 396.0441 | 311.7838 | 195.5093 | 215.3257 | 424.1209 | 255.5675 | 262.2233 | 297.4139 | 331.3001 | 452.1395 | 209.2745 | 255.2162 | 324.707 | 324.4139 | 228.3442 | 279.3372 | 278.5861 | 189.7652 | 297.3372 | 297.3186 | 354.8976 | 193.5349 | 175.5605 | 194.8791 | 217.2233 | 194.5349 | 420.6628 | 477.9232 | 622.4581 | 608 | 462.7582 | 475.128 | 672.4139 | 637.8094 | 157.2233 | 105.9047 | 118.5419 | 157.2233 | 135.8094 | 118.8861 | 184.1209 | 153.2418 | 163.4395 | 176.9744 | 183.6628 | 158.2674 | 89.0512 | 115.2162 | 117.0512 | 109.9303 | 170.4139 | 149.3186 | 193.8349 | 92.8279 | 331.7767 | 308.1465 | 400.3628 | 313.0953 | 265.3698 | 337.0953 | 424.1024 | 319.5163 | 610.8023 | 437 | 530.4978 | 569.3257
2292 | 27.2064 | 266.3475 | 317.1887 | 333.6525 | 250.1116 | 187.588 | 236.8412 | 238.412 | 203.412 | 194.7468 | 141.6997 | 177.588 | 190.2233 | 268.8412 | 206.2704 | 258.695 | 248.3821 | 220.8412 | 218.7595 | 205.9402 | 230.0173 | 197.588 | 488.794 | 567.4465 | 202.4764 | 302.9355 | 366.6777 | 431.5236 | 399.3648 | 144.7468 | 163.1415 | 181.6824 | 146.1289 | 182.3821 | 198.3349 | 245.6824 | 101.4764 | 351.0645 | 401.6179 | 486.1887 | 398.5236 | 192.588 | 208.5708 | 243.8412 | 155.9701 | 93.8113 | 93.0173 | 96.4764 | 147.5708 | 287.8412 | 360.5409 | 470.3774 | 223.8113 | 183.6824 | 200.717 | 207.2405 | 159.794 | 220.305 | 251.2704 | 356.8931 | 213.1116 | 204.412 | 199.2877 | 512.3774 | 209.1289 | 285.6006 | 286.6305 | 378.2405 | 329.7468 | 187.511 | 250.695 | 417.7296 | 288.3176 | 280.0472 | 286.3774 | 365.7767 | 417.5236 | 243.0173 | 238.6824 | 322.6824 | 322.7468 | 239.4591 | 251.6478 | 278.1887 | 227.8538 | 278.6997 | 307.0645 | 335.2877 | 190.2405 | 175.0645 | 200.3176 | 236.6352 | 185.5236 | 439.7767 | 461.588 | 593.1415 | 606.412 | 468.2359 | 504.3648 | 667.5708 | 659.5236 | 176.4764 | 113.6651 | 145.1588 | 176.4764 | 152.9183 | 158.206 | 176.3003 | 142.6824 | 168.6997 | 170.8412 | 198.3302 | 152.412 | 100.5236 | 99.4764 | 122.3302 | 127.3176 | 165.2532 | 161.7641 | 215.9654 | 87.9057 | 313.3176 | 310.588 | 386.4292 | 303.7468 | 281.6651 | 330.7641 | 438.4717 | 340.9528 | 596.8884 | 432.8711 | 573.489 | 606.6006
2293 | 27.2174 | 244.0914 | 294.4336 | 351.289 | 250.5249 | 189.6695 | 234.8023 | 232.9086 | 195.6844 | 203.3156 | 156.6844 | 181.5515 | 198.5781 | 266.4601 | 211.6312 | 221.3688 | 262.1977 | 222.2242 | 194.593 | 253.304 | 240.8671 | 202.9203 | 488.5515 | 587.7492 | 206.5399 | 288.9203 | 317.1329 | 430.3688 | 397.3688 | 150.0648 | 150.0531 | 177.4336 | 165.9469 | 197.3953 | 212.0648 | 239.5515 | 105.1977 | 359.7758 | 385.2359 | 467.711 | 392.2359 | 198.9469 | 200.8289 | 240.4336 | 173.7492 | 109.0382 | 107.8023 | 102.593 | 141.8821 | 293.6727 | 365.9585 | 439.7524 | 237.4984 | 180.6312 | 227.407 | 231.407 | 161.2625 | 253.8671 | 255.6047 | 406.3953 | 220.8821 | 198.7375 | 214.6429 | 476.7907 | 229.9734 | 255.8023 | 246.7524 | 406.1711 | 335.2359 | 219.2625 | 217.304 | 413.7375 | 289.1445 | 278.0797 | 255.5814 | 352.1179 | 412.0914 | 252.1562 | 241.6196 | 317.0648 | 331.8289 | 230.2774 | 222.1063 | 263.1329 | 191.3688 | 291.289 | 307.7343 | 358 | 214.407 | 183.2625 | 196.696 | 243.593 | 191.3837 | 421.4984 | 468.6312 | 587.9235 | 600.7375 | 450.8704 | 496.7226 | 661.5399 | 654.0914 | 184.1329 | 97.0266 | 141.8937 | 184.1329 | 132.3305 | 150.2359 | 164.3953 | 140.3156 | 184.0266 | 170.5133 | 172.2392 | 145.5399 | 100.2242 | 104.7375 | 94.0149 | 126.7758 | 158.5664 | 184.2625 | 183.1445 | 98 | 310.8937 | 318.4867 | 395.8821 | 313 | 265.711 | 354.4601 | 395.6993 | 347.0266 | 593.0531 | 407.9203 | 537.9734 | 579.1977
2294 | 27.2284 | 227.9827 | 292.0863 | 342.8166 | 222.8997 | 169.0863 | 233.9514 | 260.754 | 195.699 | 202.9342 | 160.3668 | 164.7681 | 183.6505 | 260.4671 | 206.6505 | 231.4843 | 266.8997 | 237.4013 | 214.9342 | 279.5502 | 217.6505 | 189.2837 | 466.0831 | 568.1489 | 212.5329 | 285.37 | 337.9655 | 435.5329 | 407.4843 | 134.301 | 158.7163 | 169.9514 | 157.2837 | 206.0658 | 188.4153 | 212.8652 | 110.8166 | 350.6505 | 373.2351 | 486.8198 | 378.5847 | 189.9169 | 198.5329 | 229.1003 | 162.5847 | 87.5674 | 101.0831 | 115.9655 | 150.8166 | 328.0173 | 328.2837 | 478.3808 | 210.7681 | 174 | 210.3668 | 214 | 150.9342 | 224.9655 | 244.9169 | 411.5815 | 226.5157 | 210.7163 | 181.7508 | 510.8198 | 222.9827 | 253.1176 | 290.699 | 407.1834 | 320.8511 | 204.3495 | 238.5987 | 425.7163 | 283.1661 | 289.3323 | 300.8997 | 344.4671 | 392.315 | 220.6505 | 264.3668 | 305.5189 | 330.8166 | 249.7163 | 242.9169 | 297.5361 | 198.1834 | 282.8166 | 266.2524 | 357.2665 | 197.1834 | 175.5016 | 178.5189 | 256.4153 | 225.3981 | 400.453 | 464.0173 | 633.2492 | 604.4639 | 481.8479 | 475.0831 | 673.9514 | 643.4843 | 207.1661 | 105.301 | 126.2179 | 207.1661 | 159.699 | 132 | 177.4671 | 137.7335 | 191.5674 | 172.2665 | 217 | 155.9342 | 108.0658 | 114.5157 | 122.4984 | 120.0345 | 171.5502 | 173.5674 | 185.6019 | 97.4498 | 296.1348 | 314.5329 | 398.5815 | 311.1661 | 274.5502 | 354.1521 | 428.7821 | 352.1834 | 600.0658 | 400.1521 | 523.2805 | 585.9169
2295 | 27.2394 | 221.9352 | 332.4994 | 341.2173 | 217.413 | 210.2821 | 249.7383 | 307.9112 | 181.9784 | 212.6735 | 160.4346 | 174.2593 | 194.5222 | 266.8043 | 211.0648 | 242.1081 | 257.4994 | 251.9124 | 222.1297 | 270.2593 | 231.653 | 196.7611 | 468.8487 | 553.9352 | 207.9568 | 326.8031 | 322.5438 | 433.1092 | 411.0637 | 148.413 | 157.3049 | 188.8692 | 162.2173 | 191.6303 | 173.0876 | 199.7839 | 112.5438 | 350.7599 | 401.9784 | 526.9772 | 399.3914 | 195.1741 | 195.9136 | 234 | 173.9988 | 97.4346 | 95.2389 | 98.587 | 145.8908 | 331.521 | 331.2605 | 449.3925 | 234.9352 | 173.2173 | 208.2821 | 216.1525 | 163.8043 | 213.8487 | 249.1957 | 385.5438 | 216.8703 | 208.9136 | 198.6963 | 552.347 | 225.7407 | 263.4778 | 276.1957 | 408.9784 | 340.1525 | 204.4358 | 224.6747 | 417.8475 | 284.7839 | 276.9136 | 295.6086 | 342.9772 | 374.6314 | 224.8692 | 265.2173 | 338.6086 | 325.8908 | 242.2389 | 249.1525 | 336.8692 | 195.6735 | 287.8703 | 305.3481 | 357.9136 | 197.6086 | 201.1741 | 214.9352 | 238.1525 | 200.5006 | 440.347 | 472.413 | 602.9977 | 555.934 | 451.1741 | 470.2173 | 684.455 | 646.868 | 197.2389 | 117.0648 | 144.8043 | 197.2389 | 148.1308 | 132.3914 | 187.5222 | 141 | 201.6303 | 176.0444 | 216.413 | 173.6963 | 94.0216 | 100.7611 | 100.9352 | 134.8692 | 178.3049 | 183.0432 | 215.9784 | 95.7827 | 320.5654 | 306.8259 | 379.7839 | 305.7395 | 267.8031 | 387.5654 | 413.1308 | 349.4778 | 582.8908 | 426.3253 | 488.2185 | 593.5222
2296 | 27.2504 | 229.0087 | 311.0884 | 340.2877 | 215 | 194.3761 | 225.2478 | 262.4781 | 183.0884 | 199.6239 | 152.9602 | 135.3586 | 216.2789 | 266.208 | 197.5442 | 231.7745 | 254.5977 | 229.1283 | 199.2478 | 231.1506 | 256.0884 | 203.1681 | 498.6638 | 567.0398 | 214.3362 | 298.5666 | 330.2964 | 445.4247 | 370.1108 | 146 | 164.3936 | 176 | 161.2877 | 191.6152 | 192.1283 | 237.5044 | 122.5841 | 328.4869 | 398.0972 | 494.8319 | 394.345 | 199.3761 | 207.2167 | 233.792 | 141.5753 | 92.6638 | 90.9116 | 99.7522 | 132.5355 | 315.863 | 324.3673 | 489.8717 | 251.7833 | 166.0486 | 195.2877 | 233.7347 | 160.0884 | 247.4072 | 248.7522 | 402.655 | 242.8717 | 222.9203 | 221.5355 | 538.518 | 268.3848 | 253.7833 | 276.5841 | 411.5442 | 347.1283 | 223.5617 | 244.2566 | 409.6239 | 321.2566 | 299.863 | 293.792 | 320.1905 | 402.2566 | 219.4869 | 261.792 | 339.0797 | 307.3362 | 248.9339 | 257.792 | 318.3848 | 186.5753 | 314.0797 | 316.0398 | 374.8319 | 204.1108 | 204.3362 | 228.0398 | 244.2964 | 215.0486 | 410.9203 | 471.4558 | 545.2216 | 537.863 | 458.7036 | 472.4072 | 642.4072 | 586.6016 | 198.3188 | 110.1992 | 141.9203 | 198.3188 | 159.1283 | 137.5355 | 198.4645 | 140.3761 | 198.7036 | 200.6327 | 215.0398 | 194.4558 | 94.5753 | 111.1195 | 114.8717 | 121.3761 | 199.9514 | 183.7347 | 216.4645 | 102.1195 | 327.4159 | 304.4956 | 421.0398 | 325.7347 | 239.7745 | 385.0573 | 428.4956 | 342.2789 | 580.5578 | 409.0311 | 511.3936 | 605.9203
2297 | 27.2614 | 212.8634 | 304.8457 | 354.5065 | 209.053 | 195.3039 | 232.2026 | 300.0883 | 175.3039 | 203.5418 | 154.7268 | 171.6608 | 236.7314 | 262.3745 | 197.947 | 261.9517 | 305.8504 | 227.066 | 202.2379 | 257.1896 | 247.8634 | 203.9647 | 504.7797 | 568.7974 | 206.2379 | 333.2379 | 329.6914 | 432.1849 | 401.881 | 144.8987 | 132.3699 | 185.0307 | 170.4405 | 208.8151 | 182.7974 | 236.8104 | 119.8987 | 360.5242 | 370.4229 | 497.119 | 372.5242 | 196.7797 | 191.8634 | 239.1673 | 157.0177 | 99.2203 | 100.0177 | 100.5065 | 143.3569 | 344.8634 | 336.8151 | 497.4405 | 259.4275 | 157.1673 | 201.1366 | 261.934 | 156.0483 | 269.9823 | 247.3039 | 417.4275 | 248.2379 | 216.1013 | 234.5595 | 606.96 | 250.0836 | 266.4935 | 276.1013 | 412.6078 | 336.4758 | 189.6914 | 231.7621 | 420.8104 | 303.2556 | 313.0047 | 295.6431 | 358.3392 | 401.8764 | 243.1543 | 261.6608 | 343.0353 | 303.4229 | 289.8857 | 262.7268 | 311.7574 | 209.947 | 322.4405 | 322.4229 | 389.013 | 233.8987 | 194.9164 | 240.1496 | 239.2862 | 205.066 | 406.3039 | 476.3392 | 602.3745 | 574.132 | 459.9693 | 486.1719 | 663.881 | 563.7397 | 231.9823 | 132.4229 | 134.2203 | 231.9823 | 156.6255 | 152.1013 | 183.4582 | 138.2203 | 211.4229 | 186.2026 | 220.5418 | 195.8151 | 113.5418 | 116.8327 | 123.3216 | 115.9164 | 217.6431 | 213.0353 | 207.8457 | 115.3216 | 334.7268 | 321.7091 | 413.987 | 356.3569 | 276.5595 | 353.8634 | 437.119 | 367.7974 | 628.7314 | 422.3086 | 476.947 | 588.7491
2298 | 27.2724 | 225.5581 | 317.1164 | 373.0698 | 193.8138 | 202.5815 | 234.6513 | 278.2554 | 187 | 206.2094 | 166.3956 | 172.1396 | 209.8372 | 250.1859 | 217.6047 | 241.651 | 267.0003 | 235.4654 | 201.1161 | 229.7208 | 260.093 | 214.9768 | 492.3725 | 559.3721 | 205.1161 | 327.6977 | 363.0698 | 444.6745 | 396.1631 | 141 | 159.2094 | 213.0466 | 171.7674 | 195.8604 | 179.4185 | 256.1399 | 115.5349 | 373.6977 | 381.3255 | 495.3953 | 380.349 | 194.1396 | 201.0698 | 267.0463 | 153.7208 | 106.5114 | 93.6977 | 112.3258 | 136.093 | 352.4419 | 324.5581 | 502.0232 | 220.3721 | 181.5581 | 187.0466 | 242.3721 | 169.605 | 267.0003 | 254.5815 | 386.2789 | 247.1161 | 233.9531 | 230.2094 | 678.395 | 264.4883 | 237.9305 | 280.9302 | 437.6279 | 321.907 | 222.6047 | 237.5349 | 445.0235 | 284.4188 | 274.6742 | 303.8372 | 351.8141 | 436.8144 | 231.3487 | 264 | 333.8836 | 321.302 | 263.2786 | 283.4651 | 361.907 | 229.1396 | 325.8604 | 328.2094 | 429.8604 | 231.1628 | 191.4416 | 262.9537 | 263.3487 | 217.1862 | 421.2557 | 472.6047 | 584.3721 | 615.1161 | 434.0463 | 475.488 | 658.8607 | 622.7674 | 242.9534 | 136.349 | 138.9534 | 242.9534 | 173 | 156 | 184.5114 | 137.1396 | 220 | 201.6742 | 230.4185 | 178.6745 | 113.6513 | 101.9765 | 124.9768 | 100.5815 | 222.5815 | 205.9765 | 234.302 | 119.5349 | 354.5349 | 347.8372 | 380.3487 | 342.1164 | 278.7208 | 384.2319 | 443.302 | 366.9765 | 599.2792 | 391.4883 | 501.0232 | 558.7906
2299 | 27.2834 | 223.9615 | 297.8171 | 367.3067 | 224.7524 | 194.3659 | 217.7345 | 309.7579 | 189.2036 | 205.5695 | 143.1417 | 168.6933 | 205.0206 | 268.1623 | 214.4691 | 253.5722 | 241.3453 | 208.4072 | 218.9794 | 241.5103 | 256.9615 | 202.7964 | 466.4869 | 554 | 210.0027 | 332.6933 | 349.7166 | 438.9587 | 367.9794 | 144.6726 | 153.6726 | 197.7964 | 173.4484 | 195.879 | 193.5103 | 222.8762 | 115.4691 | 368.8996 | 384.348 | 497.3067 | 377.7318 | 188.4897 | 198 | 275.4924 | 172.3659 | 119.6754 | 96.2448 | 88.9587 | 139.3274 | 346.2036 | 328.5928 | 523.0179 | 229.9354 | 190.2448 | 180.121 | 238.9587 | 140.0798 | 234.7345 | 241.7139 | 415.3246 | 265.9587 | 263.8404 | 217.3274 | 731.9147 | 281.8762 | 220.2421 | 288.4072 | 438.5928 | 323.0798 | 213.5928 | 227.9202 | 426.8968 | 275.0592 | 291.1857 | 304.4072 | 349.2627 | 397.8349 | 242.879 | 264.2448 | 345.8377 | 343.6754 | 320.6726 | 295.0385 | 352.5516 | 224.4691 | 330.7758 | 324.3865 | 430.3686 | 235.4897 | 218.2036 | 246.9175 | 281.7345 | 206.5695 | 431.2655 | 469.2242 | 581.6933 | 630.0413 | 444.4512 | 509.227 | 638.6107 | 615.3893 | 258.2036 | 126.6314 | 152 | 258.2036 | 172.5103 | 155.7552 | 204.7758 | 132.9587 | 233.7111 | 236.5488 | 250.8762 | 176.6314 | 96.4897 | 122.3067 | 118.9175 | 95.7937 | 224.4044 | 231.6933 | 267.2036 | 113.8377 | 355.9381 | 337.1444 | 398.9794 | 323.7964 | 295.4072 | 442.9202 | 475.4691 | 395.8762 | 576.227 | 412.304 | 506.8377 | 577.5488
2300 | 27.2944 | 200.9711 | 292.6284 | 354.1145 | 255.6002 | 206.0289 | 216.143 | 275.5999 | 199.0856 | 230.2001 | 162.3427 | 169.5431 | 203.2857 | 282.0285 | 212.5717 | 238.6287 | 263.2286 | 219.9429 | 222 | 238.4572 | 229.3427 | 201.6569 | 507.3142 | 549.3716 | 183.2853 | 329.429 | 321.5999 | 445.7714 | 375.1141 | 153.6858 | 157.5431 | 194.857 | 176.3716 | 179.5713 | 188.143 | 240.4857 | 115.8859 | 353.1712 | 373.857 | 490.7999 | 405.429 | 187.4287 | 189.5146 | 241.5428 | 181.7147 | 100.1712 | 102.6569 | 99.3713 | 128 | 346.3145 | 312.9429 | 552.6287 | 277.2857 | 188.4287 | 204.3998 | 238.5717 | 167.8285 | 239.0571 | 245.9715 | 434.5717 | 275.3427 | 217.5999 | 204.7143 | 774.8288 | 307.1997 | 248.2289 | 296.6002 | 430.9141 | 343.8859 | 205.1427 | 213.2853 | 438.6573 | 300.9429 | 263.2286 | 310.286 | 376.0289 | 419.1715 | 230.4283 | 266.5428 | 323.2289 | 323.9141 | 333.2857 | 324.7143 | 348.3427 | 236.1997 | 336.7711 | 338.6858 | 408.1145 | 237.2571 | 229.6284 | 264.1427 | 292.4854 | 241.2283 | 427.2001 | 477.6284 | 602.5995 | 615.0003 | 433.7707 | 478.8285 | 661.3142 | 599.3713 | 264.2286 | 145.0285 | 162.2853 | 264.2286 | 177.9426 | 156.0285 | 201 | 142.3427 | 274.9715 | 259.0574 | 269.5143 | 200.4283 | 106.2283 | 112.4572 | 133.5713 | 127.2857 | 278.0856 | 245.1427 | 264.7432 | 107.1712 | 352.6576 | 323.1997 | 400.9715 | 320.3427 | 301.8003 | 423.3998 | 477.1715 | 402.9433 | 553.0282 | 435.6002 | 503.7711 | 622.6851
2301 | 27.3054 | 221.0363 | 304.6529 | 346.7356 | 236.2281 | 186.7719 | 212.005 | 286 | 203.9586 | 223.114 | 173.3471 | 154.7356 | 211.3108 | 286.886 | 202.3521 | 224.114 | 265.3108 | 218.8496 | 216.8809 | 245.3158 | 244.9273 | 210.995 | 509.9586 | 541.6579 | 218.9273 | 307.8547 | 336.3108 | 456.0827 | 394.005 | 147.8083 | 142.4662 | 198.995 | 160.5752 | 189.9637 | 175.9223 | 243.5802 | 109.1191 | 365.3421 | 379.8809 | 499.7719 | 374.9637 | 184.0414 | 173.3521 | 242.7669 | 149.5752 | 115.0363 | 109.0313 | 108.7669 | 128.8083 | 339.1241 | 311.3108 | 530.8396 | 284.7719 | 181.8083 | 211.6479 | 227.8133 | 175.7306 | 245 | 237.0727 | 417.886 | 283.9223 | 235.5439 | 217.7043 | 791.4863 | 333.9273 | 220 | 280.7306 | 447.4612 | 332.8083 | 227.005 | 230.3371 | 443.5489 | 306.3158 | 264.2331 | 288 | 345.7256 | 409.6529 | 247.6529 | 274.2331 | 307.6629 | 337.2281 | 335.3835 | 326.9273 | 378.4762 | 257.5388 | 364.4612 | 343.8547 | 398.8496 | 239.886 | 244.6165 | 269.5702 | 312.6892 | 267.4248 | 428.4662 | 494.7719 | 629.5702 | 601.6629 | 489.886 | 478.3784 | 658.3007 | 610.3835 | 276.01 | 149.0777 | 192.8083 | 276.01 | 199.3471 | 156.8446 | 199.6529 | 153.886 | 285.2018 | 244.9273 | 276.3885 | 222.2331 | 133.2331 | 112.5802 | 136.9586 | 131.5388 | 287.5388 | 263.7719 | 247.9687 | 123.3835 | 333.1654 | 334.5702 | 399.886 | 334.3108 | 288.2744 | 436.3058 | 471.4248 | 382.8547 | 593.5852 | 425.9273 | 515.0263 | 658.0313
2302 | 27.3164 | 212.7183 | 301.5634 | 363.8731 | 221.7743 | 200.944 | 224.5206 | 303.1846 | 195.2537 | 217.1549 | 170.2388 | 173.5634 | 222.7183 | 283.2668 | 219.084 | 222.0989 | 282.0709 | 236.2257 | 206.0989 | 271.4926 | 244.9422 | 204.9589 | 486.8863 | 554.7463 | 224.5765 | 332.6754 | 352.7892 | 469.2668 | 402.8583 | 149.7183 | 159.1549 | 190.7052 | 165.2686 | 195.4646 | 179.7612 | 254.3097 | 122.7183 | 352.5354 | 378.1137 | 509.1549 | 381.8731 | 195.8451 | 196.2817 | 237.3806 | 145.5354 | 104.7463 | 85.944 | 105.916 | 135.7892 | 372.2817 | 326.3806 | 516.1417 | 291.3377 | 186.5354 | 194.3526 | 244.7743 | 170.2108 | 252.3246 | 228.0429 | 435.9589 | 285.5074 | 264.7332 | 255.9571 | 923.595 | 352.2537 | 227.888 | 280.2817 | 456.4235 | 345.9869 | 233.8863 | 202.4794 | 479.3377 | 322.9142 | 276.3806 | 290.8172 | 332.7612 | 413.0429 | 252.1697 | 277.9291 | 342.2817 | 320.5206 | 344.5223 | 351.7332 | 423.5765 | 253.9291 | 366.944 | 381.3526 | 416.7892 | 252.3246 | 247.3097 | 256.3955 | 297.6194 | 281.6063 | 455.0149 | 501.3377 | 611.888 | 609.8005 | 486.2668 | 465.1417 | 642.4235 | 590.224 | 311.1828 | 156.5074 | 189.3657 | 311.1828 | 200.1828 | 158.3246 | 198.2537 | 160.972 | 313.6754 | 276.776 | 292.4086 | 235.2257 | 137.2108 | 128.3806 | 128.2537 | 140.0429 | 289.8451 | 276.2537 | 280.916 | 123.2257 | 373.6194 | 315.7612 | 398.8023 | 340.084 | 310.2537 | 439.9869 | 489.5503 | 408.2388 | 625.2817 | 444.2537 | 488.1566 | 636.0709
2303 | 27.3274 | 217.2921 | 297.1199 | 359.236 | 215.1161 | 207.7659 | 215.3502 | 327.8896 | 206.8801 | 223.8202 | 159.7621 | 170.8839 | 234.9363 | 275.4063 | 205.764 | 232.9401 | 292.648 | 246.706 | 212.824 | 256.1779 | 210.4681 | 224.1817 | 478.3445 | 547.824 | 200.352 | 315.47 | 346.4776 | 454.6442 | 392.2883 | 148.706 | 156.118 | 209.7678 | 191.648 | 195.1741 | 205.2322 | 247.3539 | 117.2959 | 341.6536 | 406.648 | 508.764 | 385.1741 | 206.8202 | 196.706 | 253.9382 | 151.118 | 95.1779 | 96.294 | 116.5899 | 148.882 | 367.4139 | 348.2303 | 535.2379 | 275.8259 | 198.5861 | 207.706 | 234.588 | 162.1161 | 267.7659 | 253.3502 | 462.5319 | 293.7659 | 282.5899 | 244.1741 | 1038.4643 | 359.764 | 247.118 | 288.9382 | 476.8259 | 356.9476 | 208.5861 | 227.2322 | 460.8858 | 303.7583 | 284.706 | 296.236 | 351.47 | 427.1779 | 272.706 | 266.764 | 349.4681 | 321.9344 | 371.0618 | 366.3558 | 404.6442 | 243.058 | 378.47 | 395.882 | 443.7003 | 263.3558 | 253.2902 | 291.5281 | 293.1161 | 300.706 | 462.4177 | 486.1199 | 641.9962 | 537.7079 | 466.352 | 488.3539 | 657.2398 | 547.0599 | 323.4044 | 173.2921 | 185.2902 | 323.4044 | 185.6498 | 193.7583 | 209.5861 | 167.47 | 312.0524 | 318.176 | 304.2322 | 251.2921 | 130.5861 | 146.4082 | 133.118 | 149.4738 | 291.118 | 294.9363 | 302.7621 | 134.294 | 355.2978 | 332.706 | 386.8221 | 328.5281 | 318.9401 | 474.1741 | 522.2341 | 401.5842 | 629.5281 | 436.7697 | 526.7098 | 636.3577
2304 | 27.3384 | 232.1441 | 292.8015 | 352.2433 | 236.6393 | 197.2433 | 236.3693 | 285.6756 | 217.0181 | 238.8741 | 168.9552 | 164.6756 | 263.5496 | 298.9733 | 204.4867 | 242.4504 | 277.5048 | 249.063 | 210 | 246.6937 | 230.5945 | 191 | 517.604 | 544.6937 | 208.7567 | 319.6126 | 310.8826 | 465.0363 | 419.8111 | 139.1174 | 155.5315 | 195.8196 | 179.2615 | 199.8922 | 207.5411 | 238.2433 | 107.5315 | 320.2156 | 394.5678 | 518.8196 | 392.9552 | 218.1985 | 203.6574 | 264.1174 | 151.1441 | 87.8378 | 91.7929 | 99.2615 | 153.1441 | 350.937 | 388.27 | 540.4408 | 281.5219 | 205.2615 | 222.0086 | 248.8645 | 167.0992 | 260 | 272.5315 | 447.9189 | 302.54 | 275.9819 | 264.5133 | 1097.2967 | 373.8015 | 238.8741 | 312.5945 | 451.8922 | 335.603 | 222.9189 | 245.1622 | 451.2519 | 341.5496 | 281.5496 | 304.2519 | 362.0448 | 417.6937 | 274.1441 | 280.4952 | 354.8922 | 357.0363 | 384.4675 | 366.0715 | 428.8196 | 257.8015 | 385.3693 | 407.4952 | 478.0811 | 261.54 | 288.1078 | 317.1526 | 314.333 | 317.7652 | 435.1078 | 486.396 | 645.4781 | 551.0353 | 483.333 | 493.333 | 646.1526 | 556.54 | 354.6851 | 174.667 | 205.7118 | 354.6851 | 171.3693 | 212.5592 | 221.7748 | 180.1889 | 362.8015 | 319.7748 | 323.3874 | 265.5315 | 146.1441 | 158.2796 | 134.9819 | 133.2882 | 303.0897 | 328.7567 | 320.8378 | 138.9819 | 335.063 | 340.27 | 393.8559 | 344.0448 | 328.7567 | 490.5315 | 538.27 | 430.063 | 628.8111 | 410.2071 | 502.9189 | 616.7204
2305 | 27.3494 | 234.1327 | 310.2212 | 357.1506 | 250.2565 | 192.2744 | 234.4424 | 301.0089 | 201.6282 | 231.0532 | 160.6461 | 167.9026 | 257.6814 | 276.5487 | 198.3718 | 246.0885 | 265.5576 | 245.4866 | 215.0974 | 252.3718 | 248.6461 | 184.3096 | 493.9205 | 553.2391 | 213.4071 | 325.7788 | 331.3186 | 447.3186 | 403.1416 | 137.4781 | 155.1769 | 200.0795 | 170.3718 | 182.7346 | 173.416 | 239.9647 | 115.7788 | 346.5661 | 386.3718 | 510.336 | 367.1238 | 201.416 | 216.8584 | 251.0089 | 166.1948 | 92.9558 | 89.5576 | 88.7788 | 156.0885 | 373.9473 | 408.9115 | 581.0353 | 320.0353 | 202.1064 | 243.2565 | 266.5308 | 161.416 | 269.5576 | 285.2391 | 464.0179 | 337.7699 | 302.2038 | 281.7788 | 1190.0626 | 399.8231 | 225.3186 | 317.2654 | 429 | 353.2386 | 219.9026 | 232.3981 | 477.4071 | 335.3628 | 277.5929 | 317.3449 | 374.1769 | 411.2654 | 271.6725 | 301.3628 | 333.5929 | 339.3186 | 469.1506 | 428.664 | 440.0885 | 275.5398 | 407.6551 | 430.9115 | 475.0442 | 295.4955 | 312.2744 | 359.8231 | 353.6993 | 342.3892 | 466.9205 | 512.3096 | 613.3365 | 599.8052 | 497.2122 | 483.9553 | 694.5576 | 587.3096 | 336.248 | 159.5134 | 201.7346 | 336.248 | 185.3718 | 161.4513 | 224.416 | 199.7256 | 385.3186 | 324.6461 | 332.5134 | 261.3539 | 152.2744 | 132.7256 | 148.4602 | 136.6282 | 336.9115 | 342.3275 | 339.0179 | 144.8141 | 353.4692 | 342.1148 | 389.9558 | 349.4866 | 340.7346 | 516.3012 | 544.2565 | 445.2833 | 618.5134 | 430.9205 | 512.3275 | 636.5398
2306 | 27.3604 | 239.5808 | 309.2646 | 380.0074 | 228.2941 | 200.2941 | 225.5808 | 301.7722 | 192.3014 | 245.6544 | 181.6323 | 158.9853 | 254.0221 | 283.3235 | 205.7133 | 259.8675 | 285.6691 | 234.9117 | 229.3088 | 267.3235 | 258.3824 | 185.8822 | 497.1178 | 572.0074 | 214.3014 | 332.6912 | 320.0883 | 444.6912 | 434.1689 | 164.0957 | 145.3456 | 190.2646 | 177.3824 | 195.9926 | 189.9632 | 242.7722 | 120.3751 | 321.2131 | 393.0515 | 474.2573 | 353.9411 | 213.6618 | 217.5955 | 276.5881 | 179.0663 | 98.9706 | 105.6986 | 103.9632 | 152 | 434.0294 | 404.0663 | 577.5661 | 301.3456 | 218.4045 | 215.6691 | 260.8601 | 172.3382 | 294.6323 | 302.6838 | 501.6102 | 337.9632 | 335.6912 | 294.3162 | 1318.4924 | 431.4924 | 229.6397 | 308.6397 | 430.6544 | 325.7791 | 218.2646 | 230.2131 | 480.2867 | 331.353 | 285.3014 | 291.2941 | 364.3456 | 426.1321 | 271.5513 | 321.5072 | 344.2793 | 358.1984 | 503.2573 | 485.5881 | 450.5587 | 275.2646 | 433.0442 | 441.9485 | 461.0883 | 297.2646 | 315.6618 | 386.5292 | 387.3014 | 335.4703 | 480.3751 | 503.9558 | 654.2573 | 583.6249 | 481.9264 | 437.1247 | 717.647 | 579.9485 | 385.8896 | 183.9853 | 206.3898 | 385.8896 | 207.272 | 161.9558 | 238.3162 | 191.3751 | 402.8748 | 360.8528 | 372.8675 | 261.8748 | 155.6618 | 136.6176 | 158.9926 | 127.3014 | 362.838 | 363.9558 | 368.0074 | 140.3162 | 357.4855 | 323.2057 | 392.6618 | 334.9411 | 362.2646 | 564.6618 | 560.3452 | 463.9926 | 650.2646 | 458.9337 | 519.728 | 657.11
2307 | 27.3714 | 241.7234 | 307.4098 | 371.8229 | 230.7639 | 203.1071 | 247.6273 | 271.6273 | 212.3246 | 272.3356 | 188.5978 | 164.3727 | 246.6273 | 292.1771 | 193 | 277.1256 | 279.5093 | 249.0481 | 248.3541 | 270.3432 | 246.3727 | 213.882 | 470.8339 | 578.9224 | 210.989 | 326.6863 | 301.4907 | 444.8634 | 444.683 | 143.4907 | 140.6273 | 202.1366 | 168.118 | 198.6863 | 205.8339 | 229.4427 | 114.1476 | 364.6383 | 380.0295 | 499.118 | 361.4207 | 221.5202 | 230.4798 | 281.5278 | 166.1476 | 102.1956 | 98.3727 | 117.7749 | 146.8524 | 423.941 | 384.6273 | 615.6383 | 307.9519 | 202.7749 | 217.7454 | 274.9781 | 178.8929 | 303.6568 | 302.4612 | 529.3541 | 359.6678 | 332.4317 | 295.5683 | 1413.3575 | 503.0219 | 234.5978 | 307.7639 | 439.8339 | 373.6459 | 229.1071 | 269.1771 | 497.4612 | 348.6492 | 285.4207 | 300.6273 | 348.6459 | 462.1256 | 307.8044 | 370.1476 | 365.8634 | 381.9595 | 553.5126 | 513.5202 | 496.8153 | 303.9519 | 433.011 | 464.2361 | 456.2175 | 310.8524 | 345.4831 | 412.3322 | 383.3027 | 403.7563 | 476.5497 | 519.4317 | 665.0481 | 587.1071 | 498.2251 | 486.7934 | 703.0961 | 604.9815 | 418.7454 | 196.5793 | 198.8929 | 418.7454 | 223.3137 | 172.9705 | 239.2251 | 180.6863 | 450.7268 | 400.6568 | 391.8415 | 306.6383 | 162.1476 | 151.3727 | 151.7344 | 140.118 | 437.5422 | 393.5017 | 382.8153 | 145 | 322.8634 | 351.4798 | 387.8229 | 349.6273 | 380.6568 | 576.9815 | 648.3432 | 482.129 | 675.1771 | 459.4983 | 513.2951 | 709.0776
2308 | 27.3824 | 221.6168 | 282.2185 | 377.0621 | 220.3319 | 195.8202 | 250.3319 | 275.0428 | 230.4454 | 314.1563 | 178.6681 | 168.0664 | 251.1092 | 295.8244 | 203.6638 | 249.1756 | 278.9529 | 243.5075 | 263.8244 | 278.1092 | 262.5075 | 209.1992 | 498.7067 | 590.3126 | 200.0857 | 314.137 | 324.0857 | 461.9101 | 370.8244 | 157.9101 | 150.0857 | 193.621 | 183.4647 | 200.3555 | 207.0471 | 268.9529 | 117.6017 | 368.6488 | 375.9572 | 516.242 | 362.0621 | 228.3126 | 209.8244 | 270.6168 | 178.1328 | 102.5975 | 101.3555 | 115.758 | 135.2227 | 442.0386 | 392.3084 | 638.4882 | 336.9529 | 208.2227 | 226.9101 | 242.5075 | 199.6874 | 295.8908 | 295.3597 | 554.4218 | 386.621 | 350.0193 | 298.242 | 1575.4477 | 555.0428 | 232.8437 | 306.2185 | 458.1092 | 372.0386 | 230.3512 | 259.6723 | 517.3748 | 390.2227 | 274.6874 | 309.0193 | 340.9958 | 422.4454 | 332.2848 | 366.4925 | 371.8908 | 362.6874 | 631.3277 | 563.6789 | 544.2848 | 329.7538 | 482.3041 | 472.1798 | 506.1521 | 326.5932 | 407.8202 | 415.2378 | 376.0621 | 441.242 | 477.8479 | 517.1135 | 653.4647 | 569.8672 | 497.0428 | 475.8437 | 671.3983 | 637.0193 | 444.2613 | 231.8395 | 223.5975 | 444.2613 | 220.9336 | 172.4218 | 228.8008 | 186.9765 | 502.5932 | 406.0428 | 383.5268 | 325.5782 | 172.3555 | 156.1328 | 144.3084 | 148.7109 | 487.1135 | 446.9058 | 434.1949 | 145.3555 | 344.5311 | 334.7345 | 383.1092 | 345.5782 | 377.5118 | 604.7538 | 664.2848 | 520.4647 | 695.1756 | 440.5739 | 530.1563 | 694.8437
2309 | 27.3934 | 240.6124 | 310.942 | 381.1541 | 229.2645 | 207.7935 | 248.9674 | 289.3098 | 229.1485 | 300.5417 | 177.4583 | 177.355 | 262.1612 | 285.0453 | 221.529 | 253.7029 | 303.9873 | 258.4384 | 246.7935 | 293.942 | 293.2518 | 203.2319 | 512.8189 | 578.529 | 227.5163 | 311.0651 | 340.4837 | 457.5997 | 360.7808 | 166.471 | 172 | 194.645 | 183.0254 | 204.6775 | 191.5743 | 270.4511 | 111.1485 | 342.6775 | 373.0906 | 533 | 381.5997 | 219.8388 | 204.5616 | 297.3352 | 179.058 | 112.645 | 96.8515 | 107.0906 | 141.5616 | 459.5671 | 404.3805 | 637.8515 | 343.2319 | 224.4909 | 223.7029 | 277.6648 | 205.2192 | 303.5941 | 279.8713 | 540.8189 | 404.9293 | 378.8261 | 311.6902 | 1748.027 | 603.8785 | 240.5743 | 321.7029 | 473.5743 | 405.3805 | 254.4837 | 220.2645 | 544.9873 | 381.116 | 293.0906 | 319.4384 | 355.2899 | 412.3225 | 355.9547 | 347.5163 | 385.8459 | 385.5036 | 663.7482 | 649.2844 | 580.8261 | 356.3423 | 532.942 | 493.672 | 533.3098 | 372.6322 | 433.4003 | 466.884 | 416.5616 | 475.2844 | 473.1231 | 516.4456 | 661.8515 | 564.1612 | 532.2717 | 493.1358 | 680.4257 | 642.6577 | 501.9619 | 254.5997 | 249.4583 | 501.9619 | 220.471 | 177.942 | 230.7228 | 205.1938 | 565.1811 | 440.9039 | 439.0326 | 356.4655 | 176.3098 | 162.5743 | 158.2192 | 153.6775 | 517.3368 | 495.5036 | 471.2065 | 155.5616 | 363.6775 | 331.413 | 405.9293 | 338.9547 | 407.9365 | 630.2391 | 697.884 | 528.4837 | 717.355 | 475.8261 | 544.8586 | 705.884
2310 | 27.4044 | 212.0403 | 318.2803 | 355.5223 | 244.3419 | 210.8408 | 229.62 | 290.0594 | 231.1592 | 286 | 185.1592 | 197.2208 | 272.3015 | 288.1996 | 238.3822 | 248.1614 | 302.6985 | 247.7007 | 262.3822 | 311.1614 | 292.5987 | 189.0807 | 490.3206 | 569.1592 | 226.7388 | 340.639 | 320.3588 | 440.6603 | 378.5011 | 173.1805 | 167.8195 | 186.9406 | 170.7813 | 222.0213 | 199.8004 | 244.6603 | 118.0998 | 336.8386 | 387.38 | 517.0382 | 361.2399 | 222.62 | 219.0998 | 284.3206 | 183.5011 | 99.62 | 92.5605 | 118.7197 | 148.4989 | 440.7623 | 392.2399 | 667.7644 | 348.0829 | 253.0594 | 212.4607 | 286.1189 | 204.2208 | 337.0998 | 308.0807 | 536.9427 | 440.2421 | 395.1401 | 310.1805 | 1888.5076 | 683.8217 | 239.2993 | 308.9406 | 478 | 390.9597 | 244.62 | 229.6412 | 540.6581 | 393.0425 | 299.3991 | 308.7007 | 339.8811 | 423.482 | 343.6794 | 372.9618 | 422.1189 | 408.5202 | 676.1827 | 695.3419 | 621.843 | 381.7601 | 536.0998 | 559.1805 | 536.7197 | 386.4417 | 475.4226 | 478.9002 | 448.2017 | 507.6603 | 510.5605 | 542.62 | 659.4607 | 567.4607 | 562.5987 | 512.0998 | 683.9809 | 618.7601 | 566.1232 | 262.7432 | 268.5605 | 566.1232 | 240.862 | 197.0616 | 256.9215 | 216.0594 | 624.1636 | 499.862 | 472.8217 | 413.2399 | 178.5796 | 172.3206 | 165.2017 | 167.6009 | 601.9406 | 551.964 | 478.8004 | 165.9193 | 354.4182 | 356.482 | 414.259 | 341.5011 | 474.0807 | 688.7241 | 719.7813 | 519 | 735.7007 | 484.9193 | 580.9002 | 726.6412
2311 | 27.4154 | 221.708 | 318.8613 | 380.761 | 267.0767 | 211.0146 | 244.301 | 296.5164 | 218.0383 | 277.3686 | 185.3777 | 201.1533 | 293.6077 | 283.9617 | 241.3596 | 255.053 | 286.1003 | 246.407 | 286.938 | 320.0146 | 278.3157 | 206.8467 | 497.4307 | 538.3833 | 210.3923 | 319.1533 | 327.3247 | 446.9617 | 394.7847 | 174.115 | 170.8084 | 204.0237 | 174.699 | 225.876 | 198.938 | 251.7463 | 114.6077 | 322.4927 | 382.5073 | 501.593 | 357.646 | 240.0474 | 224.239 | 290.2537 | 181.3449 | 97.4307 | 98.823 | 113.0383 | 138.885 | 477.3157 | 422.3868 | 738.4544 | 405.8704 | 244.2153 | 227.3157 | 293.5401 | 216 | 328.8997 | 310.5456 | 571.6843 | 460.1916 | 402.1003 | 313.8613 | 1997.3923 | 712.708 | 250.7937 | 317.3923 | 489.77 | 371.876 | 251.4544 | 258.3777 | 537.5401 | 416.053 | 291.708 | 310.938 | 369.3394 | 446.115 | 360.8941 | 387.6223 | 421.6934 | 410.6077 | 728.5164 | 724.354 | 665.5693 | 390.4544 | 542.0237 | 560.5073 | 556.1477 | 408.823 | 534.3631 | 505.9324 | 494.1861 | 544.5638 | 509.3686 | 549.062 | 670 | 582.708 | 580.4871 | 505.469 | 684.053 | 614.8997 | 615.9233 | 322.5164 | 277.177 | 615.9233 | 275.4927 | 223.9233 | 272.115 | 215.062 | 670.6843 | 570.5875 | 499.354 | 425.3394 | 183.593 | 191.2007 | 186.469 | 170.1916 | 645.7024 | 624.6934 | 509.3247 | 161.4927 | 339.9854 | 366.9526 | 384.699 | 353.0767 | 490.2774 | 738.3686 | 736.646 | 526.4544 | 767.3631 | 469.8997 | 578.8997 | 769.5017
2312 | 27.4264 | 220.9075 | 320.8786 | 363.9537 | 244.0057 | 194.4046 | 257.8844 | 313.3352 | 220.6648 | 283.0174 | 197.3352 | 186.5722 | 297.4509 | 284.1676 | 199.8324 | 239.0463 | 294.4046 | 230.8093 | 280.4509 | 307.0463 | 293.9018 | 221.0231 | 499.0463 | 533.0868 | 199.2659 | 300.9306 | 355.289 | 449.1907 | 396 | 171.8786 | 184.3815 | 207.0751 | 162.7341 | 213.711 | 196.4971 | 253.5722 | 112.7861 | 331.4046 | 376.4278 | 528.9249 | 363.711 | 270.8324 | 220.0057 | 283.9769 | 162.4971 | 97.0231 | 96.5954 | 104.7399 | 148.8093 | 486.0231 | 439.4392 | 718.4392 | 432 | 255.1388 | 230.763 | 321.0982 | 217.6185 | 327.474 | 312.7573 | 581.185 | 457.5434 | 424.1619 | 328.8267 | 2039.2718 | 746.7053 | 257.9132 | 308.6936 | 513.6648 | 382.7747 | 260.9769 | 276 | 569.5491 | 394.3815 | 308.3064 | 304.4509 | 386.763 | 443.0694 | 396.5491 | 393.4683 | 441.5491 | 412.4278 | 784.1793 | 743.3758 | 693.4914 | 408.8786 | 566.5203 | 588.0117 | 580.2659 | 426.8267 | 594.4914 | 562.8267 | 524.1676 | 577.6011 | 508.5434 | 562.8324 | 663.9306 | 576.2427 | 641.4509 | 511.3295 | 672.9018 | 604.1676 | 652.7515 | 352.2833 | 317.0348 | 652.7515 | 273.4797 | 230.4046 | 272.3064 | 225.5954 | 691.5145 | 632.1676 | 510.6879 | 477.1562 | 212.5434 | 215.0925 | 186.2601 | 155.8093 | 684.815 | 654.6648 | 575.7284 | 178.0925 | 366.711 | 342.237 | 380.8267 | 340.1214 | 512.948 | 745.6359 | 749.5897 | 550.5434 | 814.948 | 476.5665 | 574.237 | 806.9769
2313 | 27.4374 | 208.1662 | 328.9154 | 362.1692 | 231.5075 | 197.9184 | 241.3323 | 309.5861 | 232.7523 | 309.3323 | 193.1692 | 192.1722 | 300.6647 | 277.0816 | 214.8398 | 245.8338 | 274.5711 | 231.5831 | 302.009 | 305.4955 | 295.3263 | 219.4139 | 499.997 | 561.6617 | 188.2538 | 304.5075 | 341.4169 | 441.7462 | 397.6677 | 172.8278 | 184.9184 | 188.4985 | 188.3383 | 223.4139 | 212.7523 | 261.2568 | 117.6707 | 343.2568 | 387.9214 | 508.3113 | 383.4199 | 268.3293 | 220.4319 | 300.1782 | 177.9184 | 99.1662 | 99.3353 | 99.0846 | 139.994 | 498.5891 | 438.027 | 682.4229 | 450.7613 | 270.0816 | 223.9154 | 310.5015 | 236.6767 | 349.0906 | 331.3323 | 629.1812 | 467.2447 | 446.2508 | 348.8278 | 2135.4379 | 797.6737 | 240.1782 | 318.7643 | 518.6647 | 422.6617 | 261.3353 | 265.1602 | 551.3263 | 418.6827 | 301.988 | 313.5015 | 389.0876 | 452.4169 | 391.2508 | 433.003 | 437.0846 | 427.2568 | 861.3503 | 776.4229 | 731 | 425.2538 | 577.4139 | 638.9154 | 585.9304 | 440.1572 | 607.4019 | 596.1692 | 521.2508 | 583.2688 | 529.5015 | 559.0786 | 667.5075 | 571.4259 | 653.003 | 525.4985 | 703.0936 | 638.3563 | 725.5165 | 365.0846 | 368.4139 | 725.5165 | 273.009 | 236.003 | 277.2447 | 224.1662 | 740.3383 | 650.9304 | 560.7733 | 510.9094 | 226.4139 | 227.4169 | 190.2477 | 154.4985 | 701.024 | 688.015 | 613.4139 | 200.8398 | 378.9154 | 347.8338 | 387.4865 | 335.003 | 525.3293 | 776.8338 | 779.4199 | 581.0906 | 813.1542 | 488.8278 | 585.2538 | 826.0966
2314 | 27.4484 | 203.5663 | 316.1282 | 354.6946 | 246.4248 | 202 | 257.4559 | 316.8584 | 231.9911 | 298.8451 | 205.0089 | 186.1149 | 294.2832 | 284.5886 | 220.1282 | 256.0134 | 268.748 | 253.3188 | 317.4203 | 308.3054 | 280.7168 | 211.2832 | 506.3054 | 550 | 218.0312 | 337.8807 | 351.4426 | 439.4381 | 406.0089 | 153.9911 | 182.5619 | 181.9955 | 190.5574 | 213.1371 | 209.8451 | 271.4203 | 117.5574 | 344.4069 | 394.7079 | 487.4737 | 372.1149 | 268.4426 | 243.1238 | 318.8451 | 179.854 | 101.4337 | 106.5752 | 105.8629 | 139.3054 | 534.3188 | 491.5574 | 735.1995 | 461.5485 | 272.8673 | 241.602 | 324.301 | 260.8584 | 356.9866 | 319.1282 | 636.8183 | 480.3233 | 453.5797 | 338.5752 | 2188.8629 | 829.0223 | 273.4381 | 354.146 | 532.4559 | 406.2787 | 269.0045 | 261.5886 | 545.301 | 440.5485 | 280.8673 | 325.1549 | 408.8718 | 449.9955 | 406.3054 | 430.9866 | 422.4025 | 441.7124 | 910.4426 | 839.9297 | 748.5975 | 445.1594 | 598.8985 | 634.2832 | 599.5441 | 445.8985 | 594.0312 | 628.6109 | 528.1505 | 622.7079 | 510.6812 | 568.1728 | 694.8718 | 610.1683 | 685.3233 | 544.3188 | 715.4114 | 663.6723 | 764.8584 | 371.4337 | 364.5752 | 764.8584 | 296.5752 | 243 | 277.0178 | 224.2876 | 752.8584 | 694.5886 | 621.0089 | 501.7257 | 240.1728 | 243.8807 | 194.4426 | 156.1505 | 749.2743 | 761.6287 | 613.4381 | 204.8406 | 372.5663 | 341.2743 | 375.3143 | 341.5708 | 551.195 | 797.7436 | 780.9911 | 579.1149 | 794.301 | 488.0178 | 618.0356 | 826.097
2315 | 27.4594 | 208.2715 | 308.298 | 337 | 243 | 208.181 | 272.7285 | 325.064 | 217.3775 | 281.819 | 202.404 | 161.883 | 307.011 | 288.521 | 229.9579 | 257.8455 | 295.0265 | 277.2605 | 308.234 | 331.7395 | 284.3245 | 213.415 | 515.6115 | 572.5164 | 235.404 | 347.9625 | 358.2605 | 454.181 | 409.1435 | 147.766 | 167.3775 | 181.532 | 181.9735 | 206.6755 | 207.83 | 255.17 | 114.2715 | 332.947 | 387.234 | 522.2605 | 358.0375 | 273.4945 | 241.894 | 302.2605 | 174.7925 | 109.0905 | 108.234 | 113.649 | 172.011 | 535.7441 | 500.6334 | 785.117 | 444.532 | 272.287 | 276.064 | 331.819 | 245.223 | 344.5585 | 302.9095 | 611.713 | 534.7504 | 464.9735 | 392.3309 | 2211.66 | 835.3135 | 275.3775 | 345.521 | 541.989 | 389.404 | 267.702 | 276.5585 | 568.713 | 421.766 | 295.298 | 321.4305 | 390.2121 | 440.2605 | 415.17 | 403.5475 | 424.1654 | 437.351 | 914.17 | 890.872 | 796.2824 | 447.1965 | 637.3245 | 635.532 | 594.4259 | 463.1326 | 643.6334 | 647.223 | 556.6334 | 642.1654 | 500.8674 | 585.287 | 696.9625 | 626.819 | 701.7815 | 559.872 | 706.6225 | 644.1014 | 780.128 | 383.5055 | 376.83 | 780.128 | 291.17 | 253.1545 | 285.053 | 247.1919 | 778.2824 | 729.4259 | 642.6865 | 514.2075 | 248.8986 | 261.9095 | 209.2075 | 157.702 | 771.2294 | 795.7815 | 628.181 | 201.479 | 376.83 | 345.5695 | 385.5475 | 335.2605 | 580.34 | 820.468 | 791.1014 | 546.053 | 817.713 | 503.117 | 618.6536 | 799.128
2316 | 27.4704 | 212.2856 | 301.8401 | 359.6895 | 241.1848 | 206.0166 | 264.6303 | 319.4787 | 215.8068 | 288.0675 | 194.4455 | 181.6054 | 312.3863 | 267.6469 | 227.6885 | 250.4455 | 299.9834 | 255.0249 | 326.8743 | 343.9917 | 282.3697 | 225.3531 | 508.1765 | 576.4953 | 221.5462 | 353.513 | 367.3364 | 460.6386 | 410.2606 | 166.3364 | 159.4538 | 182.8235 | 179.2856 | 202.2773 | 216.7311 | 244.9076 | 121.0083 | 353.4372 | 392.2606 | 521.3531 | 383.3447 | 270.3531 | 251.5628 | 300.8993 | 181.9834 | 99.8401 | 103.2773 | 111.1007 | 177.8401 | 511.7985 | 529.8152 | 764.9408 | 443.1007 | 259.5462 | 278.1932 | 314.0166 | 239.8826 | 339.4621 | 307.1682 | 634.4372 | 530.1423 | 478.1682 | 448.7394 | 2279.8328 | 811.6469 | 267.9076 | 324.1932 | 535.2523 | 398.2357 | 268.7144 | 277.3614 | 577.3697 | 419.0083 | 305.6303 | 310.7144 | 374.0509 | 443.4372 | 410.3531 | 415.3115 | 447.378 | 419.9325 | 886.6636 | 867.5545 | 826.0924 | 439.622 | 629.4704 | 627.7477 | 624.8993 | 436.6386 | 667.8235 | 648.6895 | 577.1932 | 671.7311 | 523.2773 | 574.3614 | 699.7902 | 606.2939 | 725.379 | 566.5047 | 719.5379 | 667.4538 | 821.5971 | 404.3447 | 370.756 | 821.5971 | 291.7985 | 256.4704 | 293.6054 | 265.8318 | 797.2015 | 755.3614 | 659.0924 | 541.5047 | 253.6812 | 254.3697 | 218.8068 | 152.3614 | 808.8235 | 767.6469 | 664.5888 | 209.1932 | 366.6719 | 357.1932 | 375.9834 | 326.6386 | 558 | 793.7643 | 809.9159 | 550.5213 | 808.6719 | 509.7144 | 594.4289 | 822.8993
2317 | 27.4814 | 209.0591 | 284.6018 | 382.8053 | 238.5339 | 200.9912 | 273.1947 | 311.7877 | 223.0678 | 301.7374 | 208.9234 | 209.1269 | 290.8731 | 269.3216 | 212.838 | 257.9322 | 295.6872 | 245.8555 | 345.4749 | 337.8819 | 271.0088 | 226.9409 | 512 | 576.8292 | 213.5427 | 360.2211 | 356.2387 | 444.5515 | 405.9497 | 177.1445 | 158.6018 | 193.4485 | 181.1859 | 201.3304 | 205.6105 | 264.6432 | 115.9322 | 367.8643 | 389.348 | 538.7877 | 398.8643 | 267.7462 | 257.1772 | 310.3304 | 194.9322 | 98.9146 | 94.8731 | 104.4661 | 172.2538 | 521.8053 | 543.1859 | 751.7877 | 453.7111 | 260.8643 | 281.1859 | 321.1093 | 251.5427 | 370.3543 | 321.1947 | 641.407 | 504.652 | 494.0591 | 455.3807 | 2276.2738 | 822.1772 | 256.8819 | 336.0415 | 524.3656 | 427.7965 | 259.8228 | 276.2035 | 556.2211 | 425.5842 | 292.755 | 307.4158 | 409.5251 | 445.2801 | 413.8053 | 437.1532 | 445.8555 | 423.5076 | 912.2098 | 874.0503 | 836.1859 | 463.5163 | 633.8468 | 602.7462 | 625.0088 | 455.5076 | 643.4924 | 643.3743 | 598.3631 | 684.8468 | 517.2035 | 581.1269 | 712.5427 | 594.7286 | 771.1445 | 584.348 | 722.2035 | 688.0415 | 850.4661 | 429.652 | 357.9234 | 850.4661 | 305.5339 | 254.7877 | 306.6784 | 246.6784 | 799.3128 | 775.1772 | 670.1181 | 562.6105 | 274.8819 | 255.1269 | 223.7374 | 163.3216 | 807.3304 | 775.3807 | 693.8731 | 203.7965 | 347.5251 | 332.6872 | 381.4749 | 324.0678 | 531.4334 | 752.0854 | 795.2123 | 549.8228 | 791.8555 | 528.3216 | 629.3128 | 820.6784
2318 | 27.4924 | 212.173 | 304.0484 | 385.654 | 230.8244 | 220.9593 | 289.4809 | 308.1705 | 224.3919 | 299.4351 | 212.7379 | 204.4758 | 296.0026 | 274.7379 | 224.6488 | 254.6947 | 284.1756 | 277.0484 | 326.0407 | 312.9974 | 269.2188 | 210.9109 | 521.0891 | 583.7328 | 213.6107 | 347.6972 | 328.1756 | 426.5649 | 393.8702 | 164.1298 | 157.9567 | 209.5216 | 192.0865 | 207.827 | 191.2163 | 275.4758 | 125.1323 | 358.5191 | 381.2621 | 540.4326 | 391.4326 | 261.2621 | 270.346 | 318.7405 | 187.3893 | 122.6972 | 98.5674 | 103.5649 | 181.1298 | 503.6056 | 552.173 | 764.9135 | 469.6947 | 264.9135 | 297.827 | 339.6514 | 254.0026 | 406.7379 | 337.4809 | 617.7786 | 510.8244 | 492.3893 | 461.5191 | 2183.6463 | 836.7812 | 275.0509 | 338.9084 | 514.6132 | 422.8677 | 241.8702 | 258.6488 | 551.3512 | 428.9974 | 291.1781 | 309.2646 | 395.9516 | 428.2163 | 413.3053 | 429.5242 | 466.5242 | 441.2595 | 922.86 | 863.1679 | 843.7379 | 477.0865 | 626.1679 | 582.8677 | 611.2595 | 455.5598 | 630.0458 | 629.9186 | 634.3919 | 675.2544 | 514.9567 | 583.173 | 707.827 | 608.0891 | 733.2544 | 568.1298 | 772.0993 | 721.0458 | 852.9135 | 437.7379 | 359.346 | 852.9135 | 304.5216 | 277.4809 | 297.4351 | 253.6998 | 813.6947 | 823.2672 | 696.4377 | 554.4351 | 250.9542 | 265.7837 | 240.57 | 174.9567 | 802.346 | 764.7761 | 687.5216 | 204.1298 | 357.3919 | 313.0433 | 382.6107 | 309.1272 | 510.5674 | 734.0433 | 792.1323 | 529 | 814.916 | 555.743 | 656.1323 | 796.1272
2319 | 27.5034 | 220.7572 | 322.584 | 398.416 | 228.888 | 222.7148 | 289.7334 | 310.1545 | 227 | 279.4108 | 233.0051 | 208.7012 | 284.4482 | 274.888 | 222.2479 | 249.5093 | 293.472 | 303.9253 | 333.6078 | 300.5093 | 288.7572 | 202.9627 | 542.2852 | 568.7199 | 227.9066 | 343.2241 | 349.2479 | 437.7572 | 434.2904 | 162.472 | 162.4346 | 209.4907 | 184.112 | 204.1494 | 208.5705 | 265.9627 | 124.1867 | 337.6587 | 388.9627 | 538.3413 | 390.8133 | 257.1867 | 268.3786 | 318.6214 | 180.4533 | 122.6401 | 109.4907 | 98.5654 | 166.2241 | 515.0051 | 540.6401 | 750.3175 | 448.3175 | 266.5093 | 305.9253 | 335.5093 | 235.0882 | 381.8641 | 341.1681 | 583.2988 | 507.416 | 459.9388 | 459.7946 | 2179.5518 | 810.8641 | 284.4855 | 348.8931 | 512.9388 | 415.472 | 239.112 | 251.3413 | 560.6214 | 404.2428 | 308.056 | 316.6961 | 374.9627 | 437.7199 | 405.7334 | 439.528 | 462.2988 | 451.1919 | 881.9253 | 861.9865 | 826.7148 | 463.7148 | 565.1256 | 597.0612 | 615.3039 | 421.1681 | 614.9949 | 626.5602 | 611.9762 | 611.6349 | 510.112 | 562.3175 | 721.3226 | 610.6401 | 721.1545 | 600.8184 | 780.9337 | 750.1545 | 813.2936 | 419.7334 | 379.4585 | 813.2936 | 308.9066 | 290.9813 | 294.0934 | 253.8641 | 776.6162 | 853.9813 | 695.7708 | 535.3922 | 242.944 | 264.1307 | 245.3361 | 167.1681 | 775.8455 | 743.8133 | 685.0373 | 194.6214 | 349.2054 | 325.2479 | 395.4346 | 307.2479 | 538.6639 | 718.2801 | 771.5602 | 498.5789 | 811.3361 | 579.888 | 655.1867 | 758.4295
2320 | 27.5144 | 223.438 | 322.5266 | 403 | 226.4498 | 221.562 | 262.4143 | 313.4084 | 218.4498 | 261.0236 | 249.438 | 218.4616 | 270.0591 | 284.5207 | 228.9586 | 265.5975 | 292.9882 | 298.9468 | 337.8996 | 328.1654 | 283.3907 | 213.5502 | 549.9764 | 574.9468 | 231.9941 | 345.5857 | 361.4911 | 467.5975 | 434.7755 | 166.0118 | 154.4321 | 214.5266 | 180.0118 | 189.4616 | 221.438 | 266.9941 | 113.497 | 362.195 | 387.9823 | 551.5266 | 384.9055 | 252.5325 | 277.0118 | 318.0354 | 195.0709 | 99.9114 | 102.9586 | 95.503 | 164.0591 | 508.3021 | 538.5325 | 729.9823 | 445.5857 | 256.9468 | 293.4025 | 330.4734 | 217.0827 | 382.1182 | 364.692 | 580.5798 | 489.87 | 429.9646 | 486.5916 | 2130.5687 | 761.8287 | 271.0768 | 371.9409 | 514.6507 | 398.8937 | 231.9941 | 268.0473 | 554.503 | 405.0827 | 287.3966 | 308.0118 | 371.4675 | 466.0768 | 383.4439 | 422.4084 | 461.5916 | 451.9055 | 832.1839 | 866.8641 | 821.0354 | 433.3907 | 518.4261 | 583.2784 | 641.0886 | 404.4557 | 566.373 | 592.4793 | 569.4025 | 580.5207 | 482.373 | 556.5679 | 697.3021 | 592.438 | 785.7688 | 635.9882 | 760.1772 | 786.1004 | 765.4734 | 418.5679 | 386.3671 | 765.4734 | 305.9527 | 266.3493 | 283.4675 | 246.5739 | 727.9114 | 858.5207 | 697.0945 | 526.5739 | 248.0118 | 267.0591 | 232.0295 | 160.0118 | 751.4675 | 722.8168 | 690.5443 | 189.0059 | 340.0118 | 325.4202 | 392.9646 | 327.5384 | 521.7932 | 711.0473 | 745.5266 | 448.8937 | 771.373 | 597.065 | 673.6684 | 750.5916
2321 | 27.5254 | 224.8506 | 331.6372 | 375.1768 | 215.4238 | 234.5762 | 274.2775 | 303.1525 | 221.8506 | 264.4848 | 235.9085 | 218.1829 | 262.9969 | 282.8475 | 230.2439 | 260.875 | 286.3628 | 285.878 | 331.8201 | 325.6005 | 266.0305 | 225.6067 | 544.4543 | 571.6677 | 227.3933 | 334.2377 | 363.0915 | 474.2103 | 412.4574 | 168.5152 | 146.6067 | 219 | 175.3018 | 199.4878 | 215.122 | 262.9085 | 111.9695 | 380.0579 | 376.7561 | 541.0579 | 377.2439 | 274.4878 | 264.0579 | 315.7561 | 198.2408 | 99.4269 | 101.6677 | 103.2134 | 172.9695 | 478.6067 | 560.4878 | 745.0336 | 406.9298 | 229.4512 | 286.7897 | 305.3902 | 216.5731 | 376.2377 | 383.6036 | 587.8171 | 472.122 | 450.7927 | 485.5122 | 2009.567 | 741.7592 | 276.2713 | 365.0915 | 533.3018 | 367.0884 | 246.4574 | 268.7866 | 528.7225 | 399.9359 | 271.5457 | 298.1494 | 360.3323 | 447.0548 | 383.2744 | 400.8171 | 461.0274 | 421.5731 | 754.7835 | 807.9328 | 824.4238 | 418.6067 | 502.3933 | 530.0274 | 590.5639 | 382.5731 | 541.3933 | 564.7835 | 555.061 | 565.4512 | 448.6341 | 555.1189 | 690.7927 | 560.3597 | 790.8109 | 663.369 | 797.7287 | 825.1555 | 759.4543 | 434.122 | 351.6341 | 759.4543 | 283.5731 | 263.6708 | 293.9726 | 237.3597 | 716.0915 | 829.0243 | 696.5122 | 525.0884 | 245.878 | 253.814 | 228.2408 | 160.9695 | 695.506 | 699.2134 | 678.4207 | 186.9085 | 329.1189 | 302.2103 | 393.6982 | 326.7866 | 489.061 | 709.2103 | 745.878 | 426.3628 | 746.3933 | 625.5183 | 679.3292 | 748.9969
2322 | 27.5364 | 233.3623 | 331.7797 | 340.5594 | 230.9362 | 228.0319 | 286.3391 | 290.0638 | 225.087 | 251.8667 | 210.8435 | 212.3942 | 238.0319 | 279.0551 | 235.3623 | 260.5188 | 290.9681 | 284.1101 | 313.513 | 298.5826 | 283.3536 | 214.7565 | 523.4812 | 574.3623 | 212.3942 | 311.0551 | 363.8899 | 479.771 | 447.0464 | 170.0551 | 146.8348 | 214.7797 | 155.8116 | 206.5594 | 203.1739 | 280.5739 | 106.2522 | 345.9536 | 353.1739 | 506.4261 | 388.1652 | 265.2058 | 237.3391 | 326.9913 | 184.1971 | 108.7797 | 108.0551 | 106.3072 | 167.2522 | 454.0406 | 559.6464 | 700.2783 | 369.1333 | 227.8261 | 307.6058 | 273.8667 | 209.3304 | 341.4493 | 372.0551 | 581.4725 | 441.7101 | 446.458 | 483.1884 | 1880.6493 | 682.4754 | 286.4087 | 373.2754 | 500.6232 | 346.6145 | 271.5507 | 259.8899 | 524.0464 | 388.3304 | 288.2985 | 286.4725 | 379.2667 | 429.1333 | 389.8899 | 386.5594 | 451.2754 | 388.4812 | 701.9304 | 760.2841 | 790.3478 | 396.6783 | 494.2522 | 481.2377 | 551.1565 | 352.1188 | 527.9768 | 523.0087 | 563.3304 | 506.8522 | 455.9913 | 525.5913 | 660.6087 | 561.629 | 742.4493 | 740.0609 | 805.5275 | 862.8812 | 730.5681 | 423.229 | 358.9913 | 730.5681 | 256.2841 | 275.5043 | 297.3942 | 224.3855 | 678.9073 | 763.1826 | 668.8667 | 480.8754 | 230.9217 | 234.6377 | 222.1101 | 168.4406 | 644.8348 | 657.4667 | 648.3942 | 183.4725 | 326.4957 | 320.9594 | 408.9681 | 316.8348 | 461.458 | 667.2928 | 730.9217 | 425.6928 | 738.7797 | 626.1739 | 666.9681 | 728.2522
2323 | 27.5474 | 233.6991 | 347.4338 | 361.6108 | 242.1416 | 225.9381 | 270.6814 | 290.1948 | 204.5043 | 237.2212 | 206.1152 | 216.5753 | 228.3805 | 259.4866 | 242.1771 | 275.8407 | 284.4247 | 263.8671 | 272.345 | 316.7346 | 263.9909 | 204.1593 | 491.4247 | 554.7255 | 198.2212 | 312 | 370.0974 | 470.3273 | 440.1679 | 153.1857 | 158.0355 | 205.0619 | 146.1593 | 200.0797 | 205.1948 | 292.5221 | 116.5753 | 356.6108 | 362.2126 | 506.894 | 374.8052 | 244.0797 | 248.1329 | 323.4866 | 171.4424 | 106.0797 | 107.3805 | 104.0797 | 186.2126 | 426.8407 | 566.0533 | 640.1416 | 352.4688 | 227.9645 | 289.929 | 278.1152 | 206.9026 | 334.1593 | 377.3186 | 549.1502 | 422.3364 | 412.4424 | 475.5662 | 1763.8935 | 627.3983 | 286.8848 | 355.8671 | 463.5043 | 356.1948 | 241.5926 | 247.7433 | 518.7874 | 365.3892 | 285.2654 | 307.0533 | 379.4069 | 422.1857 | 364.2476 | 345.5307 | 437.106 | 379.0974 | 621.5571 | 731.8052 | 736.0883 | 377.7788 | 454.9112 | 447.2212 | 561.2831 | 345.0974 | 490.3892 | 509.239 | 542.5485 | 491.0533 | 451.4069 | 511.0797 | 646.6286 | 558.3273 | 718.9645 | 746.2121 | 858.9031 | 896.3541 | 646.0792 | 405.6814 | 372.761 | 646.0792 | 242.3805 | 247.5662 | 276.2035 | 228.3009 | 652.1771 | 741.1771 | 660.6991 | 427.0883 | 210.7433 | 223.5043 | 227.7788 | 154.4069 | 591.2562 | 597.8052 | 631.5221 | 161.4069 | 322.584 | 321.7078 | 412.1416 | 312.3805 | 417.7255 | 607.7697 | 700.4866 | 418.7433 | 716.106 | 614.1593 | 679.3186 | 664.0792
2324 | 27.5584 | 224.9576 | 343.5716 | 399.803 | 249.5945 | 210.5716 | 282.4593 | 297.9267 | 204.5945 | 222.4055 | 214.5831 | 233.9691 | 255.158 | 277.3356 | 260.4902 | 285.5945 | 268.6873 | 240.5831 | 272.803 | 311.1237 | 230.5831 | 184.4674 | 511.575 | 557.1157 | 200.5212 | 325.2508 | 349.3941 | 470.3241 | 391.9462 | 155.6678 | 162.2703 | 198.3436 | 156.2818 | 191.614 | 201.8843 | 271.5407 | 127.3436 | 364.9886 | 386.4169 | 511.9576 | 368.7297 | 219.0504 | 287.4284 | 306 | 172.1775 | 106.4479 | 106.5521 | 105.5521 | 189.9886 | 388.5602 | 529.8924 | 644.2818 | 330.6564 | 214.8958 | 277.5945 | 273.8843 | 201.1042 | 317.228 | 403.6368 | 523.1042 | 398.3208 | 388.3322 | 461.7606 | 1617.5763 | 590.5911 | 297.189 | 327.614 | 438.7492 | 365.0309 | 215.1775 | 263.2932 | 519.4284 | 333.2703 | 285.3551 | 304.0195 | 365.9691 | 403.6873 | 344.5521 | 328.0424 | 410.5098 | 379.9267 | 568.8225 | 667.2051 | 726.1466 | 363.8843 | 408.9886 | 426.8843 | 556.3127 | 338.197 | 459.9267 | 481.0814 | 519.4479 | 455.4445 | 435.2085 | 489.3632 | 670.9267 | 547.2818 | 693.7492 | 732.772 | 893.5098 | 951.7525 | 585.0619 | 367.2165 | 349.5716 | 585.0619 | 231.614 | 246.4593 | 279.1466 | 218.2703 | 587.1204 | 672.8077 | 626.0081 | 425.9805 | 193.7182 | 221.386 | 228.7915 | 141.5212 | 529.228 | 550.8729 | 619.9267 | 141.8958 | 323.5945 | 298.2394 | 401.9267 | 325.3551 | 388.6449 | 562.6759 | 657.6026 | 410 | 714.3551 | 619.8648 | 670.8534 | 628.4593
2325 | 27.5694 | 221.0797 | 319.0983 | 398.0186 | 244.3254 | 196.2576 | 307.417 | 300.3373 | 190.8592 | 222.0305 | 213.822 | 228.9695 | 253.7304 | 282.8592 | 257.1474 | 275.8101 | 254.1474 | 257.3187 | 269.3254 | 282.583 | 226.4356 | 199.1713 | 494.5405 | 570.4356 | 203.8712 | 337.1288 | 338.8526 | 478.6627 | 398.2205 | 157.0186 | 156.5644 | 211.1102 | 160.178 | 188.5153 | 189.6932 | 263.5153 | 115.8406 | 342.9695 | 394.5153 | 513.7238 | 369.9203 | 210.417 | 286.5033 | 308.2576 | 171.7915 | 100.3559 | 115.4661 | 102.6135 | 173.0491 | 365.7729 | 510.2696 | 617.6999 | 324.6627 | 226.9814 | 281.3559 | 290.4781 | 199.178 | 312.1594 | 392.6507 | 504.2457 | 351.2457 | 342.3931 | 462.3068 | 1537.595 | 537.8101 | 309.2271 | 330.7238 | 448.8831 | 362.6932 | 209.1474 | 284.1288 | 529.7915 | 315.1474 | 282.3254 | 293.7729 | 366.0491 | 415.6746 | 315.0863 | 318.9389 | 405.3865 | 374.4356 | 543.8101 | 588.1846 | 690.8473 | 317.8287 | 408.9814 | 396.0677 | 537.5524 | 297.0863 | 414.9269 | 424.9575 | 519 | 406.3559 | 418.3745 | 493.012 | 667.1288 | 554.5644 | 680.7424 | 759.8526 | 927.3307 | 972.7118 | 546.5219 | 355.3679 | 324.5339 | 546.5219 | 207.6321 | 265.2085 | 274.3254 | 218.2085 | 490.1846 | 625.1102 | 564.3931 | 397.8473 | 181.9203 | 216.583 | 207.2457 | 151.0797 | 499.8898 | 502.5524 | 569.2828 | 130.8406 | 331.3068 | 298.822 | 378.9389 | 350.5458 | 377 | 518.7795 | 635.3068 | 391.9389 | 661.6573 | 606.0677 | 672.2882 | 626.8898
2326 | 27.5804 | 234.2272 | 326.8408 | 361.4718 | 234.7301 | 214.7243 | 300.8524 | 289.2757 | 187.9942 | 210.5456 | 223.4971 | 221.4233 | 207.8951 | 268.4602 | 236.466 | 243.1223 | 285.369 | 263.1961 | 249.3495 | 301.835 | 236.9573 | 208.4291 | 490.9884 | 594.2214 | 207.6136 | 323.0058 | 339.3495 | 481.6505 | 409.0427 | 138.3495 | 151.8097 | 224.8835 | 149.3495 | 181.6194 | 200.8039 | 260.6563 | 105.6932 | 331.3864 | 393.9631 | 510.466 | 357.3495 | 213.6563 | 283.3806 | 302.5029 | 179.6874 | 108.6874 | 100.6621 | 109.2272 | 152.699 | 366.3864 | 484.0174 | 564.8951 | 303.4291 | 204.1281 | 249.8951 | 299.0058 | 190.0796 | 292.3184 | 374.6136 | 487.2699 | 315.8155 | 331.8408 | 432.7417 | 1441.1939 | 486.0912 | 332.835 | 326.3126 | 466.1534 | 347.2757 | 204.6136 | 283.2699 | 495.0116 | 293.3126 | 276.767 | 293.233 | 367.0369 | 422.3864 | 310.4544 | 309.0369 | 397.0427 | 369.9631 | 469.5999 | 527.2815 | 630.0485 | 295.8835 | 383.2446 | 365.699 | 514.6563 | 288.9942 | 358.0485 | 355.4407 | 487.8582 | 362.7844 | 380.3553 | 477.7048 | 670.8835 | 550.3864 | 657.6621 | 763.8097 | 944.4291 | 970.9204 | 511.466 | 327.0543 | 311.6563 | 511.466 | 187.5029 | 255.0058 | 250.8893 | 203.3922 | 457.5398 | 589.864 | 563.068 | 353.7728 | 160.1223 | 182.2019 | 198.3437 | 143.466 | 456.1281 | 469.8524 | 542.6505 | 136.2641 | 316.7359 | 306.1903 | 365 | 351.7728 | 358.5456 | 488.9262 | 614.9689 | 363.5825 | 631.6078 | 584.3495 | 675.8097 | 595.8155
2327 | 27.5914 | 236.8771 | 345.2459 | 357.3138 | 230.6991 | 212.8642 | 259.5504 | 282.8899 | 178.3853 | 198.767 | 213.6862 | 220.411 | 192.1229 | 280.4239 | 230.945 | 254.7376 | 299.0972 | 246.5761 | 234.1651 | 292.3724 | 234.9871 | 188.1651 | 491.7963 | 617.3009 | 208.644 | 321.4239 | 319.4532 | 453.8385 | 370.9614 | 138.4789 | 142.6991 | 217.5761 | 145.356 | 180.4789 | 200.3431 | 258.945 | 106.356 | 358.0386 | 378.6312 | 519.0808 | 362.1908 | 204.2881 | 270.9743 | 291.6991 | 167.2073 | 96.7963 | 97.3688 | 106.5761 | 135.8771 | 350.0422 | 439.2202 | 567.3817 | 286.1101 | 195.9028 | 241.7798 | 274.4532 | 174.9871 | 266.6991 | 395.0808 | 481.2881 | 323.9707 | 316.0844 | 391.3853 | 1263.6935 | 425.4275 | 323.9614 | 319.945 | 438.1394 | 317.3303 | 210.356 | 274.9321 | 475.945 | 281.644 | 268.0422 | 283.1101 | 352.9192 | 407.2202 | 279.2367 | 301.3982 | 387.233 | 339.3174 | 409.5082 | 476.3174 | 605.1101 | 277.9743 | 335.7963 | 327.0844 | 513.5339 | 265.8385 | 321.9192 | 344.9707 | 470.8899 | 318.0972 | 387.7376 | 432.7963 | 640.6055 | 531.6862 | 628.4532 | 745.8642 | 908.8514 | 1005.4496 | 459.4146 | 279.9192 | 296.3982 | 459.4146 | 186.7119 | 226.9192 | 273.2165 | 180.1651 | 418.4825 | 537.1523 | 539.3596 | 314.6055 | 146.411 | 162.822 | 208.8899 | 123.8092 | 396.0715 | 396.1559 | 508.9486 | 130.0972 | 314.6569 | 289.3853 | 361.4661 | 332.6312 | 334.9871 | 460.0293 | 570.9614 | 348.2881 | 624.7963 | 579.178 | 640.1945 | 563.9192
2328 | 27.6024 | 237.6077 | 331.759 | 371.0064 | 211.3667 | 195.1833 | 259.4436 | 276.1769 | 174.2218 | 198.7974 | 175.341 | 219.3987 | 207.6269 | 287 | 227.5885 | 272 | 264.55 | 243.0064 | 245.4372 | 275.209 | 234.0128 | 174.7846 | 505.2603 | 552.4538 | 203.3923 | 315.3731 | 332.8487 | 429.3987 | 371.0449 | 140.5949 | 144.4115 | 203.1833 | 151.209 | 165.1577 | 194.3987 | 261 | 112.209 | 349.5372 | 378.418 | 531 | 375.4115 | 203.4051 | 258.2026 | 306.6398 | 158.0128 | 83.2026 | 96.3795 | 110.2218 | 140.8167 | 342.8038 | 425.5885 | 549.9295 | 285.6077 | 191.3667 | 237.9744 | 272.8167 | 156.5756 | 278.6333 | 386.5564 | 452.1449 | 316.55 | 284.1705 | 383.0128 | 1132.2705 | 379.1513 | 344.4885 | 316.5885 | 443.2539 | 297.7846 | 218.0128 | 234.5243 | 455.1513 | 278.7974 | 263.8103 | 282.0064 | 342.2026 | 374.9487 | 262.4308 | 287.3859 | 383.5949 | 310.9808 | 341.4794 | 431.7461 | 604.0064 | 250.7718 | 299.3538 | 322.2282 | 476.3154 | 222.759 | 308.7974 | 320.1128 | 414.8718 | 297.3795 | 394.1769 | 413.791 | 654.0769 | 497.55 | 573.9038 | 696.9166 | 861.9359 | 993.5307 | 421.7846 | 257.1769 | 258.3346 | 421.7846 | 190.6013 | 220.4115 | 264.9359 | 178.8103 | 356.9231 | 483.1128 | 492.9615 | 276.3667 | 138.7846 | 165.0064 | 187.7461 | 128.6269 | 355.3667 | 373.4436 | 447.9295 | 120.8038 | 312.1833 | 266.582 | 351.1833 | 299.9487 | 298.5372 | 409.7269 | 534.9679 | 331.1705 | 600.9808 | 554.7461 | 586.7333 | 533.3602
2329 | 27.6134 | 224.0471 | 361.9501 | 346.6163 | 216.3393 | 177.7285 | 265.6371 | 286.1801 | 203.7022 | 190.0235 | 176.6344 | 209.1828 | 204.9557 | 252.0263 | 228.9086 | 280.59 | 245.9321 | 238.8643 | 228.7756 | 250.3892 | 224.5014 | 169.3864 | 481.0499 | 539.5194 | 196.705 | 314.3629 | 311.1856 | 399.5485 | 380.705 | 140.8407 | 127.1385 | 177.1385 | 146.6371 | 166.1122 | 174.9792 | 246.2742 | 102.7285 | 343.4986 | 401.1801 | 516.2742 | 388.2036 | 198.8643 | 247.9557 | 317.1593 | 156.4778 | 90.7493 | 93.6814 | 109.0235 | 151.5222 | 342.7729 | 395.0028 | 518.7964 | 248.7313 | 185.295 | 229.3864 | 260.8199 | 153.9086 | 272.4335 | 366.2507 | 435.6136 | 273.3892 | 244.3213 | 365.5249 | 1024.5069 | 338.2299 | 319.8462 | 309.9321 | 431.162 | 280.7285 | 205.4335 | 219.59 | 447.9765 | 274.9321 | 262.0914 | 301.1801 | 350.9765 | 376.7936 | 257.2742 | 284.2271 | 369.7285 | 287.2064 | 338.3601 | 367.1413 | 537.2798 | 229.7285 | 288.6814 | 318.115 | 454.295 | 217.9765 | 263.8227 | 275.8199 | 406.2923 | 265.2299 | 385.7729 | 407.3186 | 635.0499 | 474.0235 | 543.3864 | 724.3809 | 864.7022 | 979.3629 | 379.572 | 226.6842 | 243.5222 | 379.572 | 187.3186 | 216.6371 | 235.1828 | 167.8878 | 313.3656 | 402.0055 | 476.0914 | 263.5457 | 137.0679 | 163.3186 | 169.7729 | 122.8878 | 310.0263 | 331.7784 | 426.6136 | 99.2978 | 303.3186 | 260.3864 | 345.3864 | 306.0886 | 267.3421 | 376.5692 | 519.4778 | 317.705 | 574.1385 | 530.6371 | 604.6551 | 522.2271
2330 | 27.6244 | 222.1362 | 355.8292 | 330.6259 | 207.8499 | 182.5173 | 249.7344 | 266.3326 | 181.6999 | 177.4896 | 180.8638 | 204.2517 | 201.7552 | 227.4965 | 243.2656 | 270.3534 | 238.9931 | 238.1293 | 204.1155 | 247.3949 | 200.9792 | 190.9053 | 469.6466 | 571.8914 | 177.7275 | 317.7483 | 311.7829 | 413.545 | 355.4689 | 140.1224 | 142.7898 | 180.8983 | 148.8845 | 181.8845 | 181.7759 | 231.9931 | 106.8914 | 335.1017 | 400.7344 | 496.9862 | 369.5913 | 194.3741 | 230.9862 | 302.8568 | 162.388 | 86.8638 | 92.2448 | 104 | 153.1224 | 290.6791 | 359.8499 | 516.7552 | 229.6328 | 174.8568 | 212.7275 | 233.7275 | 154.4965 | 243.8499 | 339.4689 | 402.8292 | 240.3534 | 214.9862 | 340.1086 | 945.6999 | 313.2379 | 318.2932 | 297.9862 | 411.7483 | 287.3949 | 179.3534 | 219.3672 | 444.8638 | 262.9862 | 240.8499 | 302.612 | 366.6397 | 366.0947 | 232.3534 | 291.2586 | 338.843 | 293.5242 | 313.686 | 348.2656 | 479.6051 | 211.3603 | 278.4827 | 289.3465 | 454.4965 | 221.7483 | 234.1224 | 253.7344 | 399.7137 | 263.3949 | 366.2241 | 423.7759 | 623.6466 | 463.3672 | 488.5498 | 717.9447 | 862.1017 | 962.0947 | 338.4758 | 198.231 | 242.6189 | 338.4758 | 178.7414 | 188.843 | 239.0138 | 162.1293 | 274.7137 | 366.1224 | 447.3396 | 233.8361 | 130.8638 | 154.1155 | 170.2517 | 124.6397 | 267.9723 | 266.3326 | 397.5844 | 114.4156 | 306.0069 | 286.2863 | 354.388 | 328.1362 | 271.2656 | 354.3534 | 441.5221 | 319.3811 | 541.5982 | 512.8568 | 613.9792 | 525.5035
2331 | 27.6354 | 214.2369 | 330.2566 | 328.4474 | 214.8684 | 167.0264 | 252.5723 | 272.2499 | 165.7434 | 166.6184 | 177.9145 | 227.3354 | 191.875 | 229.8289 | 255.1052 | 261.6382 | 264.0788 | 217.6646 | 206.3816 | 251.5329 | 210.5065 | 183.579 | 464.9606 | 549.6448 | 170.5526 | 300.408 | 306.8553 | 423.4277 | 362.421 | 144.7434 | 143.408 | 187.1711 | 149.1711 | 166.2172 | 186.6184 | 248.7828 | 93.579 | 349.2499 | 384.4277 | 515.2499 | 377.9736 | 183.1514 | 216.704 | 286.0658 | 159.704 | 87.1052 | 99.5723 | 102.7237 | 144.9803 | 270.2105 | 373.8881 | 488.3685 | 207.4738 | 182.9539 | 201.0855 | 247.6118 | 154.2763 | 218.9606 | 320.8948 | 384.2763 | 236.1052 | 200.0658 | 315.1316 | 881.8817 | 306.0855 | 313.3882 | 288.1711 | 405.2566 | 287.0658 | 184.0394 | 206.4277 | 418.3027 | 260.8289 | 233.8289 | 270.8356 | 341.7304 | 389.3749 | 208.9606 | 260.5396 | 313.5132 | 250.3949 | 284.7237 | 332.0264 | 469.7237 | 210.1052 | 244.8356 | 282.1052 | 422.3685 | 202.4935 | 204.283 | 232.3224 | 350.1778 | 227.9673 | 357.5526 | 429.2566 | 598.5396 | 486.8025 | 464.9342 | 712.5065 | 846.8948 | 899.2239 | 282.0528 | 195.1052 | 241.2763 | 282.0528 | 168.6184 | 166.704 | 218.8356 | 167.1908 | 273.6776 | 348.408 | 423.704 | 205.5132 | 128.5526 | 127.0264 | 148.0264 | 111.2172 | 215.7962 | 228.2172 | 375.5329 | 121.7895 | 298.1514 | 262.4344 | 361.9145 | 315.1316 | 245.4541 | 340.5329 | 428.9078 | 320.0855 | 496.1778 | 450.1186 | 570.5396 | 512.9606
2332 | 27.6464 | 223.2612 | 338.708 | 312.6902 | 210.6902 | 175.4647 | 251.4955 | 259.3317 | 165.7478 | 172.7567 | 193.3098 | 240 | 173.292 | 254.1152 | 243.6902 | 261.3496 | 240.9732 | 222.5621 | 204.7964 | 245.7478 | 224.9027 | 166.146 | 476.5134 | 538.854 | 180.4558 | 283.1946 | 299.2522 | 419.9513 | 376.6018 | 149.3009 | 150.9116 | 212.9692 | 148.3009 | 181.0179 | 183.6504 | 241.7875 | 85.2522 | 345.4379 | 349.0795 | 482.1678 | 383.8451 | 169.8451 | 217.2036 | 292.708 | 155 | 91.9513 | 106.3009 | 100.0487 | 140.3496 | 274.0487 | 334.6634 | 445.6326 | 219.0179 | 171.8362 | 201.9513 | 246.6902 | 140.6902 | 220.1063 | 310.1946 | 388.2522 | 225.3406 | 190.4469 | 284.8362 | 764.3357 | 264.0219 | 333.8232 | 273.6415 | 397.4469 | 285.9027 | 178.9911 | 229.6197 | 423.5621 | 264.3009 | 230.7964 | 237.7875 | 338.3098 | 374.8759 | 217.2612 | 230.5929 | 325.863 | 268.1331 | 270.3406 | 279.3228 | 441.6812 | 200.6415 | 246.9116 | 270.0397 | 377.0308 | 196.5531 | 189.3009 | 218.146 | 336.854 | 207.5531 | 349.2433 | 401.283 | 582.2522 | 492.2433 | 451.7875 | 715.1946 | 840.0973 | 873.6018 | 235.8362 | 160.2741 | 215.9821 | 235.8362 | 160.4469 | 166.5531 | 226.7656 | 154.0397 | 246.8759 | 296.0705 | 394.9335 | 187.9424 | 116.9911 | 122.4558 | 159.0665 | 106.5045 | 190.3982 | 243.6683 | 336.575 | 106.8937 | 273.7875 | 262.1152 | 372.1063 | 323.2125 | 220.4955 | 315.2344 | 431.3893 | 299.4866 | 475.6991 | 428.1063 | 548.3982 | 480.9335
2333 | 27.6574 | 224.7099 | 316.5023 | 335.4859 | 228.8468 | 176.3962 | 240.7099 | 248.3137 | 178.5802 | 178.6627 | 177.8042 | 212.8278 | 163.0236 | 234.8513 | 227.3844 | 261.6627 | 237.2312 | 226.0354 | 199.349 | 230.7452 | 215.0707 | 166.3137 | 477.7217 | 506.2122 | 169.757 | 290.9293 | 282.4433 | 430.2783 | 366.0707 | 148.6745 | 151.3844 | 206.7924 | 143.0354 | 173.7924 | 167.4316 | 227.3608 | 92.3019 | 328.0471 | 338.3019 | 489.4741 | 379.3373 | 164.6745 | 212.7099 | 285.0825 | 158.3137 | 96.9764 | 108.9882 | 97.6745 | 141.3255 | 267.6981 | 327.2076 | 451.2194 | 219.0825 | 157.6863 | 227.5213 | 217.1297 | 134.3255 | 211.0825 | 327.8703 | 384.0354 | 212.6981 | 193.9646 | 267.3726 | 694.7806 | 252.6038 | 318.2004 | 286.8821 | 385.7217 | 280.0471 | 175.3137 | 215.1768 | 404.5023 | 249.757 | 224.6863 | 247.8821 | 362.5684 | 359.9764 | 210.0943 | 233.6156 | 341.9646 | 284.0589 | 231.8513 | 246.0707 | 401.816 | 189.3608 | 240.0943 | 260.0118 | 381.2194 | 194.3608 | 186.6863 | 204.3962 | 335.3608 | 201.3844 | 335.3844 | 370.7688 | 580.0236 | 473.7452 | 458.5684 | 695.757 | 822.757 | 849.1532 | 215.0589 | 136.6863 | 210.6156 | 215.0589 | 156.6745 | 163.6981 | 207.1886 | 148.651 | 206.1297 | 271.0118 | 366.757 | 170.7335 | 102.7099 | 117.0589 | 155.4316 | 110 | 180.3844 | 210.2829 | 311.0354 | 101.3373 | 266.651 | 268.6981 | 368.3844 | 321.7335 | 231.5802 | 308.3019 | 410.408 | 285.6863 | 478.3137 | 434.9882 | 518.5023 | 493.184
2334 | 27.6684 | 222.3501 | 325.626 | 344.2497 | 214.3238 | 179.1003 | 243.7503 | 266.2007 | 165.0741 | 181.0251 | 163.9749 | 198.325 | 165.0502 | 250.0513 | 221.6499 | 243.7742 | 232.4994 | 219.6249 | 179.0992 | 233.4504 | 198.1995 | 170.7001 | 453.4242 | 481.2497 | 198.4516 | 282.1494 | 283.1254 | 406.6488 | 351.2246 | 123.699 | 142.2748 | 183.1745 | 150.8005 | 157.5998 | 152.9247 | 252.0012 | 102.4253 | 330.7503 | 357.8757 | 496.5496 | 342.5484 | 168.7252 | 206.2999 | 265.8495 | 138.3991 | 96.9749 | 102.5747 | 107.8005 | 152.1254 | 265 | 320.4493 | 421.8483 | 205.9247 | 160.7252 | 240 | 231.6511 | 126.8997 | 211.4253 | 299.8483 | 353.3238 | 199.1995 | 197.675 | 260.6249 | 631.6727 | 225.5986 | 303.0251 | 292.2748 | 389.7754 | 258.4493 | 180.3751 | 202.7001 | 408.2258 | 233.2246 | 225.7001 | 240.4493 | 355.4743 | 364.0251 | 201.325 | 219.7742 | 335.5496 | 252.6739 | 211.9498 | 247.4253 | 380.2246 | 193.0753 | 243.4755 | 243.4743 | 347.7982 | 181.8746 | 185 | 180.7742 | 296.5484 | 204.4253 | 328.9749 | 413.3022 | 540.8734 | 467 | 446.7491 | 683.9498 | 800.1494 | 777.9475 | 210.675 | 141.7503 | 185.9738 | 210.675 | 147.8997 | 157.6249 | 203.8255 | 138.5245 | 172.724 | 238.9487 | 323.2235 | 160.2748 | 103.7252 | 109.9749 | 154.4253 | 103.2497 | 177.3501 | 169.1243 | 297.1995 | 101 | 284.2007 | 278.1505 | 353.8746 | 320.7252 | 227.5245 | 296.1494 | 413.4504 | 300.8757 | 467.8495 | 425.1995 | 492.5245 | 475.6237
2335 | 27.6794 | 236.7464 | 340.4366 | 361.6196 | 213.3718 | 171.3156 | 259.3718 | 245.1326 | 185.2421 | 179.938 | 152.6902 | 201.4366 | 163.5634 | 238.8818 | 222.3746 | 243.9352 | 229.4366 | 198.755 | 194.7435 | 235.2507 | 199.1859 | 176.8112 | 427.5043 | 490.3718 | 192.6311 | 290.8084 | 294.8732 | 423.9294 | 324.3184 | 139.4928 | 142.7493 | 187.3098 | 140.879 | 160.4986 | 158.2478 | 252.6282 | 101.8761 | 307.8818 | 343.3184 | 498.8732 | 339.4337 | 169.6254 | 200.8761 | 279.245 | 123.1888 | 96.6873 | 98.3127 | 109.5634 | 130.8818 | 251.9409 | 290.6931 | 420.8703 | 185.817 | 169.8732 | 208.3833 | 221.3184 | 138.121 | 218.4366 | 265.1916 | 357.8703 | 195.3746 | 199.3746 | 259.6873 | 581.2565 | 212.062 | 279.2565 | 261.1326 | 397.4366 | 276.8055 | 168.9409 | 199.8761 | 379.8847 | 238.6225 | 233.8732 | 251.9323 | 340.4395 | 373.9352 | 200.3127 | 215.8112 | 360.5548 | 271.6167 | 203.8141 | 246.1888 | 363.6282 | 182.2536 | 233.1916 | 233.938 | 343.4957 | 185.2478 | 158.8818 | 172.6873 | 311.304 | 181.8818 | 350.6816 | 404.3213 | 588.9827 | 467.6873 | 398.5101 | 591.2738 | 760.696 | 756.1859 | 204.8141 | 147.7493 | 187.121 | 204.8141 | 159.121 | 150.5014 | 212.062 | 144.6844 | 193.6167 | 201.3184 | 310.1239 | 149.7522 | 91.5663 | 113.8112 | 148.3775 | 97.2507 | 162.1916 | 169.2478 | 267.2565 | 96.1888 | 285.1268 | 284.6873 | 326.3184 | 327.8112 | 224.062 | 289.6873 | 404.2536 | 299.3775 | 442.0677 | 405.5663 | 503.4957 | 445.1916
2336 | 27.6904 | 233.9051 | 330.8062 | 333.221 | 197.3121 | 170.8973 | 208.3315 | 252.9883 | 176.6125 | 164.3082 | 156.3953 | 198.8023 | 162.6996 | 202.5136 | 238.3914 | 259.8934 | 234.4981 | 190 | 182.9128 | 221.407 | 190.8062 | 161.5097 | 427.6957 | 489.0039 | 177.1027 | 270.0156 | 290.3043 | 427.5058 | 318.8973 | 129.6125 | 155.8934 | 210.1899 | 139.5969 | 158.8023 | 175.9922 | 248.3992 | 95.8023 | 301.5969 | 344.8934 | 497.1027 | 330.6086 | 171.0988 | 193.4031 | 277.5058 | 141.2887 | 84.407 | 94.5019 | 108.6996 | 124.5969 | 240.4031 | 304.7868 | 414.3082 | 182.8973 | 149.9128 | 189.1027 | 198.407 | 140.8023 | 223.4981 | 277.5891 | 362.5019 | 198.7984 | 182.5097 | 238.3121 | 527.3237 | 204.6047 | 281.2926 | 271.0872 | 364.7191 | 278.5058 | 186.7868 | 206.3953 | 381.1899 | 256.2926 | 209.0156 | 243.5097 | 321.6086 | 355.6125 | 186.0078 | 227.0949 | 339.0195 | 270.6086 | 203.0988 | 218.1144 | 399.9767 | 177.3992 | 232.2965 | 228.8023 | 353.7984 | 179.2054 | 151.8973 | 182.0949 | 289.6202 | 159.5058 | 335.1144 | 398.7946 | 563.0311 | 454.0078 | 374.0039 | 569.5891 | 755.7946 | 683.442 | 213.8934 | 123.1144 | 178.6086 | 213.8934 | 135.2171 | 150.7984 | 172.4226 | 127.6125 | 192.6086 | 195.1977 | 292.4109 | 153.6957 | 101.7907 | 94.3121 | 137.7035 | 96 | 161.2965 | 154.1105 | 223.1183 | 106.593 | 274.3043 | 291.2965 | 343.2848 | 311.1105 | 225 | 254.3198 | 368.6163 | 306.1938 | 473.5774 | 368.9167 | 452.9323 | 439.3992
2337 | 27.7014 | 239.9666 | 330.2714 | 334.7975 | 179.4572 | 162.3215 | 202.5095 | 241.3548 | 159.1691 | 164.4071 | 147.8977 | 177.7786 | 174.3904 | 221.188 | 225.7786 | 238.6596 | 253.0856 | 194.9833 | 180.8309 | 197.4905 | 194.5428 | 170.3737 | 406.7953 | 520.1713 | 175.7119 | 251.5929 | 300.5261 | 386.693 | 338.7975 | 121.4238 | 141.7786 | 168.1669 | 154.1023 | 159.1357 | 180.5762 | 230.4905 | 98.2714 | 284.0667 | 344.3048 | 483.6096 | 318.3048 | 147.0834 | 184.5929 | 245.9477 | 129.3548 | 86.119 | 92.2881 | 99.0334 | 122.0167 | 220.2025 | 309.3048 | 400.881 | 192.8309 | 130.7453 | 173.4739 | 193.7119 | 136.8643 | 220.7286 | 257.5239 | 355.3048 | 199.2881 | 173.5762 | 214.762 | 501.881 | 207.4071 | 292.6952 | 248.9644 | 356.4071 | 301.0523 | 192.0167 | 199.3215 | 396.8309 | 282.6452 | 203.4071 | 253.0856 | 329.9499 | 331.762 | 195.6618 | 225.3048 | 310.4739 | 243.3548 | 193.3215 | 189.9144 | 336.1313 | 178.7119 | 211.5072 | 227 | 330.7953 | 152.2191 | 162.5428 | 162.5072 | 271.8643 | 167.8142 | 314.7453 | 363.8454 | 569.6118 | 443.0167 | 363.1691 | 526.7431 | 753.5929 | 632.4905 | 200.4905 | 111.2881 | 154.2025 | 200.4905 | 127.6952 | 142.0334 | 184.188 | 127.9666 | 156.1001 | 174.931 | 264.0667 | 138.4905 | 102.5929 | 97.8142 | 129.3048 | 105.2547 | 145.4905 | 165.9333 | 226.7975 | 99.8977 | 271.7119 | 260.5406 | 400.5617 | 292.3215 | 205.7786 | 251.1023 | 365.2547 | 296.762 | 446.8621 | 375.2214 | 442.238 | 416.5072
2338 | 27.7124 | 226.6194 | 302.3081 | 336.9306 | 203.5193 | 162.3452 | 224.4839 | 250.3806 | 174.829 | 150.3436 | 151.6903 | 193.1742 | 163.7919 | 237.3452 | 231.0355 | 242.4839 | 244.7919 | 194.8274 | 193.4145 | 208.829 | 182.7919 | 173.379 | 423.7951 | 502.1355 | 157.8952 | 270.0016 | 279.2049 | 395.1742 | 327.171 | 129.9661 | 128.2065 | 182.0371 | 146.6887 | 146.9645 | 153.9291 | 236.0355 | 100 | 278.8968 | 318.1016 | 486.9661 | 344.9677 | 173.2097 | 199.3806 | 258.8984 | 131.1387 | 95.5177 | 89.8274 | 91.379 | 114.2065 | 233.2774 | 288.8952 | 390.0339 | 192.379 | 151.6226 | 162.9306 | 223.6919 | 137.4484 | 192.9291 | 253.2419 | 337.7919 | 183.7919 | 164.3097 | 206.8274 | 457.721 | 210 | 279.0678 | 244.6903 | 354.6548 | 271.4452 | 183.4823 | 198.621 | 407.2419 | 246.8242 | 198.7581 | 260.7242 | 328.0339 | 356.4161 | 171.5839 | 209.2403 | 299.2065 | 237.1726 | 204.2081 | 193.1387 | 302.2758 | 188.4145 | 241.8306 | 237.1387 | 337.6564 | 175.5887 | 151.5161 | 155.8968 | 271 | 173 | 319.6903 | 353.7935 | 543.6178 | 451.8629 | 362.1726 | 495.1371 | 707.5484 | 580.8242 | 181.4129 | 107.379 | 133.2403 | 181.4129 | 120.5855 | 134.379 | 177.171 | 137.7935 | 149.7935 | 187.7258 | 266.8629 | 131 | 105.0694 | 105.1726 | 128.4484 | 119.1387 | 130.0339 | 174.7242 | 212.9984 | 98.621 | 274.8968 | 254.2419 | 378.4452 | 298.1387 | 210.3113 | 263.9661 | 379.8629 | 314.1742 | 423.9306 | 352.5839 | 457.4145 | 402.6548
2339 | 27.7234 | 225.8919 | 301.3108 | 309.9595 | 211.7905 | 152.9527 | 228.527 | 245.2163 | 181.2635 | 171.2499 | 143.2163 | 202 | 144.0068 | 230.8987 | 191.811 | 239.1622 | 225.7433 | 215.358 | 188.8987 | 213.0541 | 210.8783 | 186.7297 | 430.3176 | 485.5811 | 162.0473 | 271.3716 | 277.8378 | 423.8851 | 319.2635 | 137.4189 | 136.3108 | 206.7837 | 156.7297 | 158.2026 | 147.6824 | 237.3176 | 119.1486 | 286.6284 | 332.5675 | 492.9459 | 308.865 | 176.6892 | 186.1149 | 243.696 | 130.5811 | 90.6352 | 99.3108 | 86.3176 | 117.1554 | 254.9932 | 267.2703 | 415.7229 | 195.4189 | 143.3244 | 169.1013 | 229.1081 | 121.7974 | 169.7433 | 228.0136 | 334.9459 | 205.2499 | 163.2095 | 194.2163 | 451.5743 | 198.2163 | 270.7905 | 235.4798 | 352.2635 | 267.0473 | 183.9459 | 211.7837 | 381.2771 | 223.3716 | 203.3648 | 242.5879 | 330.8919 | 353.2703 | 195.3512 | 200.3176 | 294.7905 | 288.8174 | 179.0677 | 191.1081 | 293.8987 | 172.8514 | 240.0609 | 246.1554 | 344.7365 | 172.5338 | 165.8851 | 176.1486 | 254.7974 | 166.3716 | 324.473 | 384.7229 | 540.7837 | 410.338 | 382.8851 | 448.2839 | 689.527 | 562.527 | 171.8446 | 118.5202 | 140.5202 | 171.8446 | 111.8446 | 136.6824 | 189.1486 | 129.6892 | 143.1622 | 168.0136 | 254.7974 | 142.0473 | 94.4798 | 106.7365 | 108.3784 | 106.7974 | 149.0945 | 173.527 | 205.7365 | 86.0068 | 254.642 | 269.5202 | 358.5811 | 280.642 | 192.9055 | 260.3716 | 356.7501 | 281.7569 | 462.5066 | 339.527 | 474.9932 | 443.7161
2340 | 27.7344 | 217.5171 | 279.5416 | 312.2316 | 195.2757 | 171.4633 | 218.2659 | 261.2169 | 168.2708 | 172.7635 | 143.4927 | 195.261 | 156.2218 | 219.0147 | 202.2071 | 228.5122 | 223.2463 | 190.0539 | 182.2561 | 207.5073 | 199.788 | 208.4731 | 425.2561 | 476.5122 | 185.4682 | 285.4731 | 299.7194 | 408.5367 | 337.7194 | 148.7341 | 142.9951 | 187.788 | 155.261 | 173.7341 | 152.7439 | 233.7537 | 110.2757 | 294.9902 | 343.9951 | 470.788 | 313.7145 | 167.7586 | 177.5024 | 243.9853 | 140.2316 | 79.7635 | 85.0294 | 95.4829 | 127.2365 | 236.0392 | 264.2463 | 387.8125 | 189.5122 | 141.4927 | 172 | 226.2512 | 138.4633 | 201.69 | 232.2267 | 320.2757 | 194.0343 | 163.2512 | 177.2708 | 447.5122 | 176.0294 | 243.7929 | 221.2659 | 381.951 | 283.7292 | 185 | 210.7586 | 387.4731 | 241.2169 | 204.5024 | 233.0049 | 333 | 371.9608 | 178.049 | 199 | 290.2561 | 278.5465 | 183.7243 | 185.2561 | 300.7341 | 174.2365 | 201.0539 | 264.4731 | 304.5661 | 164.0049 | 170.7537 | 175.5122 | 231.7782 | 158.7586 | 327.9951 | 381.522 | 542.7537 | 417.212 | 349.5661 | 453.2169 | 677.0196 | 550.0196 | 161.0147 | 105.7782 | 144.2512 | 161.0147 | 117.4878 | 138.7488 | 175.7831 | 116.2659 | 143.7439 | 160.9951 | 260.2316 | 150.4927 | 98.2365 | 102.5073 | 124.9608 | 106.9902 | 148.7635 | 144.5465 | 208.2463 | 96.7243 | 261.9755 | 247.7929 | 370.478 | 255.0294 | 195.4829 | 257.2512 | 347 | 281.9755 | 432.8223 | 350.2316 | 432.0784 | 441.7782
2341 | 27.7454 | 211.7168 | 286.2213 | 322.0885 | 209.0267 | 165.3008 | 196.7344 | 246.6901 | 186.832 | 181.416 | 138.1504 | 201.3717 | 176.4603 | 237.3099 | 177.5131 | 215.3451 | 235.416 | 201.0709 | 164.2565 | 209.8053 | 209.5045 | 192.6901 | 421.7168 | 506.7258 | 188.1947 | 269.6901 | 258.0528 | 393.3893 | 322.6901 | 153.5221 | 137.1504 | 193.6992 | 153 | 167.1061 | 156.2832 | 236.0443 | 111.0885 | 291.6725 | 351.8496 | 503.3365 | 319.4779 | 166.7611 | 158.7344 | 269.0709 | 149.3275 | 83.8496 | 90.416 | 110.416 | 147.5045 | 250.0709 | 278.6992 | 377.2832 | 206.7877 | 163.5488 | 161.3451 | 187.1856 | 136.8672 | 206.3893 | 237.7611 | 334.0267 | 192.3275 | 173.6549 | 179.0885 | 448.0443 | 171.5221 | 272.2922 | 237.7877 | 362.3184 | 256.7963 | 192.6107 | 205.1947 | 381.584 | 238.8672 | 222.2656 | 233.5221 | 336.0443 | 390.6549 | 184.7435 | 201.2832 | 276.0619 | 259.1061 | 186.7168 | 174.1061 | 297.9115 | 175.4779 | 202.9381 | 219.0086 | 310.7877 | 172.8939 | 146.4069 | 160.8229 | 234.3717 | 160.8053 | 315.3008 | 355.6901 | 542 | 405.9733 | 320.7787 | 461.5221 | 659.3008 | 521.6459 | 168.6549 | 87.0619 | 143.2389 | 168.6549 | 129.8939 | 148.1328 | 155.3008 | 129.7435 | 154.8939 | 145.2565 | 241.9291 | 153.5221 | 96.4336 | 84.2565 | 123.1061 | 121.1771 | 139.1504 | 135.7611 | 203.6725 | 93.6283 | 270.8053 | 267.1594 | 359.0176 | 296.9472 | 199.7611 | 285.1594 | 337.1061 | 267.2123 | 417.2389 | 357.8053 | 444.9205 | 432.1947
2342 | 27.7564 | 187.0259 | 268.5726 | 339.4671 | 208.0398 | 160.2266 | 193.3201 | 240 | 168.4792 | 169.5329 | 129.0398 | 168.4256 | 163.9861 | 226.2128 | 177.0536 | 235.9741 | 235.1332 | 211.8668 | 167.5069 | 187.7993 | 215.7734 | 179.8131 | 449.6142 | 535.5606 | 175.3996 | 258.3599 | 283.9879 | 412.654 | 358.5346 | 136.2128 | 141.9602 | 202.6401 | 156.8668 | 153.9464 | 164.7336 | 238.5467 | 104.4931 | 296.9602 | 348.5865 | 503.6263 | 322.8668 | 155.3996 | 183.1609 | 276 | 147.9066 | 103.7872 | 88.5865 | 107.0398 | 125.9325 | 269.3737 | 248.9723 | 381.0934 | 186.7059 | 152.9861 | 176.5606 | 197.4274 | 142.5069 | 193.173 | 233.3599 | 302.1055 | 184.7197 | 155.346 | 167.8529 | 414.9723 | 192.8806 | 237.5986 | 245.5467 | 340.6263 | 237.7197 | 181.8529 | 207.8668 | 374.1332 | 244.5069 | 213.3062 | 265.7076 | 321.5329 | 348.3719 | 199.2803 | 191.173 | 275.8668 | 276.8806 | 181.3599 | 186.4671 | 275.1194 | 171.1332 | 247.2146 | 227.7474 | 313.1332 | 161.3062 | 141.3201 | 161.6401 | 230.8131 | 146.5329 | 297.8529 | 371.4274 | 539.6799 | 433.2544 | 347.7076 | 395.4913 | 607.0519 | 500.0796 | 175.0934 | 96.1471 | 130.6263 | 175.0934 | 134.5467 | 133.2128 | 165.6938 | 121.0796 | 143.3062 | 165.5208 | 226.4931 | 140.8529 | 84.9464 | 88.2803 | 119.2266 | 111.8529 | 133.9066 | 145.2803 | 187.3062 | 100.2803 | 284.3737 | 262.0796 | 360.1869 | 308.1332 | 205.4135 | 235.9982 | 346.3737 | 286.5208 | 407.7197 | 312.5986 | 401.4117 | 435.6401
2343 | 27.7674 | 213.7829 | 259.6431 | 329.8583 | 192.644 | 185.9264 | 191.6431 | 246.2852 | 158.6431 | 156.3579 | 147.4269 | 184.712 | 180.9982 | 199.0018 | 178.4287 | 207.6458 | 226.1435 | 198.0727 | 171.5713 | 197.4986 | 196.3588 | 170.9292 | 432.0736 | 481.2907 | 175.9282 | 264.8565 | 275.5731 | 393.0736 | 315.219 | 120.0009 | 144.7856 | 188.287 | 160.3569 | 166.713 | 162.2861 | 217.0018 | 98.0718 | 317.8556 | 328.9301 | 478.0018 | 343.6412 | 146.5005 | 176.287 | 249.288 | 142.2861 | 98.7866 | 81.5005 | 96.3579 | 121.9282 | 280.8565 | 231.1435 | 381.2144 | 175.8574 | 134.644 | 163.1444 | 215.7847 | 115.1454 | 180.5722 | 213.9301 | 328.7111 | 207.9264 | 164.713 | 169.4995 | 426.2125 | 189.5722 | 242.0699 | 247.5713 | 352.713 | 230.2861 | 173.2861 | 186.2162 | 369.8574 | 207.7176 | 222.356 | 245.931 | 295.0018 | 339.7139 | 209.8565 | 207.6412 | 276.2144 | 249.2171 | 182.3569 | 180.7866 | 254.0727 | 175.4995 | 218.1463 | 213.7875 | 311.2144 | 167.9991 | 143.5713 | 155.1435 | 214.8583 | 144.3569 | 315.2125 | 387.4278 | 514.6449 | 430.4296 | 328.7167 | 402.712 | 587.5005 | 512.4986 | 176.7856 | 107.0708 | 142.713 | 176.7856 | 132.6431 | 136.6421 | 143.288 | 117.7856 | 132.7148 | 163.5722 | 207.5014 | 136.2144 | 96.9273 | 87.0718 | 131.5704 | 109.5713 | 133.7856 | 144.8574 | 180.6431 | 94.3579 | 273.0727 | 260.3569 | 365.1426 | 304.6431 | 191.287 | 269.2815 | 331.9301 | 312.0699 | 452.9245 | 319.4269 | 388.3569 | 432.2861
2344 | 27.7784 | 203.0515 | 264.5856 | 324.4041 | 204.1608 | 178.6371 | 195.7876 | 240.8185 | 180.3423 | 182.7258 | 140.233 | 176.0412 | 194.9794 | 182.6268 | 195.5547 | 205.9794 | 224 | 204.3732 | 163.2227 | 219.5547 | 184.6165 | 170.5959 | 437.767 | 485.7464 | 189.767 | 272.5856 | 297.9279 | 398.767 | 303.1918 | 116.2021 | 148.1918 | 170.435 | 149.0309 | 170.2021 | 155.4144 | 211 | 95.4041 | 321.4041 | 344.7464 | 491.1505 | 343.4144 | 143.4041 | 178.3835 | 227.6371 | 135.4144 | 84.0309 | 85.5856 | 85.2227 | 126.9897 | 243.103 | 257.7258 | 365.0412 | 203.7258 | 146.9588 | 163.5856 | 223.7876 | 140.5135 | 185.9794 | 207.4041 | 327.8288 | 201.435 | 176.9794 | 171.7979 | 392.901 | 195.7773 | 207.103 | 220.8701 | 325.8804 | 250.5444 | 159.233 | 198.3526 | 348.2536 | 221.7361 | 212.435 | 236.4041 | 323.3114 | 339.4041 | 189.6577 | 215.3938 | 279.9897 | 262.3423 | 173.4247 | 171.6165 | 255.5856 | 184.9794 | 226.9485 | 213.5856 | 299.8288 | 176.5856 | 136.0206 | 141.0309 | 250.099 | 145 | 303.4453 | 360.4762 | 484.8598 | 438.9691 | 322.5959 | 383.668 | 620.3114 | 444.3875 | 177.7979 | 96.233 | 128.6474 | 177.7979 | 138.3835 | 136.6062 | 158.3423 | 125.9794 | 138.1815 | 148.233 | 227.7361 | 137.5959 | 77.0618 | 89.1918 | 103.8804 | 103.6165 | 130.8082 | 145.5959 | 180 | 92.7979 | 294.5341 | 272.9691 | 342.8598 | 291.233 | 179.8185 | 268.6371 | 330.9897 | 310.6165 | 463.6062 | 307.4453 | 389 | 421.4247
2345 | 27.7894 | 210.9637 | 269.2409 | 313.467 | 175.5908 | 166.8977 | 187.2772 | 212.2624 | 178.7079 | 175.4159 | 133.7591 | 186.5841 | 167.8317 | 224.5627 | 182.9852 | 216.9125 | 224.8102 | 201.3284 | 178.0148 | 202.1238 | 204.8762 | 155.6056 | 430.467 | 511.2558 | 163.8317 | 259.4159 | 281.5033 | 370.401 | 287.7954 | 134.6353 | 127.934 | 178.3432 | 135.467 | 145.6931 | 153.1898 | 226.3944 | 103.1023 | 318.5693 | 355.6716 | 478.1749 | 329.0363 | 155.9637 | 162.1749 | 240.2046 | 138.8614 | 85.8614 | 88.6205 | 111.3581 | 143.3944 | 243.533 | 261.7591 | 356.9488 | 189.934 | 172.8762 | 156.0875 | 212.8465 | 141.7079 | 189.6205 | 181.8829 | 306.3647 | 180.9852 | 155.5033 | 184.9637 | 392.7228 | 183.4159 | 219.6865 | 242.3581 | 326.9125 | 228.4522 | 177.8762 | 206.2409 | 378.6501 | 211.7954 | 226.0148 | 236.8102 | 327.9488 | 350.3432 | 200.2046 | 220.8614 | 260.7442 | 253.4159 | 185.5841 | 187.0148 | 266.7228 | 172.4159 | 228.7591 | 221.4818 | 301.8614 | 176.3795 | 125.8977 | 158.2558 | 228.401 | 153.1023 | 328.1683 | 365.9637 | 519.5115 | 441.1898 | 311.6568 | 368.0875 | 604.6931 | 439.7739 | 185.2921 | 103.533 | 123.1898 | 185.2921 | 131.0875 | 137.6205 | 183.4455 | 139.3432 | 151.3432 | 145 | 218.6056 | 142.8614 | 84.7739 | 80.2772 | 105.7228 | 106.8614 | 138.1023 | 119.2624 | 189.7228 | 102.7228 | 300.1898 | 268.7079 | 343.4818 | 288.8102 | 203.1171 | 252.8465 | 328.7591 | 279.8317 | 454.8977 | 320.8251 | 408.4455 | 419
2346 | 27.8004 | 200.8397 | 272.4676 | 303.5973 | 195.1432 | 160.8874 | 201.4504 | 219.1603 | 177 | 156.3721 | 137.9351 | 167.7919 | 152.7748 | 230.8874 | 161.727 | 215.7099 | 216.7748 | 179.437 | 183.645 | 207.6928 | 196.0171 | 184.0783 | 438.6928 | 514.355 | 146.3072 | 265.0477 | 262.0171 | 377.8054 | 315.2558 | 144.7577 | 143.563 | 166.1946 | 141.2252 | 153.1603 | 165.3378 | 220.9523 | 95.1297 | 295.7919 | 338.9045 | 447.6793 | 310.3721 | 169.6928 | 161.2901 | 223.437 | 141.645 | 102.6279 | 81.5973 | 116.355 | 123.9694 | 225.437 | 262.645 | 378.2081 | 175.9523 | 137.6964 | 186.0783 | 191.082 | 140 | 201.5153 | 220.4162 | 321.7405 | 173.7099 | 145.8874 | 179.7748 | 381.0171 | 182.4676 | 201.9694 | 229.2595 | 331.4676 | 250.7883 | 178.8874 | 185.6144 | 390.2901 | 231.8531 | 238.2252 | 252.6279 | 310.5496 | 345.5973 | 183.437 | 211.3072 | 260.9351 | 238.4847 | 192.2901 | 212.3856 | 264.8874 | 164.0649 | 209.082 | 215.5973 | 294.7748 | 174.355 | 151.9657 | 144.9045 | 223.4676 | 154.1775 | 321.0171 | 364.8874 | 495.2766 | 420.437 | 318.0477 | 383.273 | 584.1946 | 441.355 | 164.7919 | 96.1297 | 119.7099 | 164.7919 | 119.1297 | 136.355 | 174.0171 | 137.0649 | 151.5324 | 154.0477 | 218.2901 | 141.5324 | 77.3072 | 84.5802 | 109.645 | 88.2595 | 127.6622 | 136.0306 | 205.9829 | 75.3892 | 292.5973 | 280.9829 | 357.3378 | 289 | 209 | 248.355 | 314.0171 | 276.2901 | 452.1775 | 325 | 401.4847 | 456.0135
2347 | 27.8114 | 213.8614 | 254.6341 | 299.4956 | 181.7993 | 139.9645 | 199.9911 | 226.1741 | 167.8171 | 177.2096 | 138.1652 | 178.8614 | 161.0178 | 258.3836 | 178.0355 | 206.6519 | 220.0089 | 194.2007 | 176.4867 | 187.4601 | 200.5133 | 193.5044 | 387.572 | 541.5488 | 179.0621 | 249.4689 | 296.5666 | 373.4867 | 323.6696 | 131.8082 | 134.643 | 186.3748 | 148.8437 | 162.6785 | 162.1563 | 239.8703 | 89.6607 | 319.3836 | 305.7816 | 459.5311 | 278.6164 | 158.643 | 167.0089 | 221.5044 | 116.9556 | 109.3393 | 95.8614 | 105.9822 | 127.3481 | 221 | 275.5222 | 357.1208 | 185.6874 | 144.0266 | 177.9734 | 193.6785 | 130.8171 | 196.4867 | 213.3037 | 319.3215 | 184.6874 | 164.2007 | 171.3215 | 341.2682 | 177.9911 | 224.5488 | 245.8703 | 332 | 258.6696 | 172.9911 | 196.8614 | 360.1119 | 187.7461 | 232.4867 | 245.9822 | 286.9645 | 359.8614 | 167.3126 | 208.1652 | 269.5133 | 242.6785 | 179.643 | 198.6341 | 260.6607 | 187.2096 | 215.8526 | 183.9467 | 303.0178 | 176.5044 | 123.7727 | 151.8526 | 222.3304 | 147.3215 | 332.1918 | 370.6785 | 514.714 | 416 | 346.714 | 356.1119 | 565.9734 | 425.9734 | 170.0178 | 86.4867 | 108.9822 | 170.0178 | 142.0444 | 139.3393 | 166.8259 | 129.3215 | 146.8259 | 143.4778 | 203.9734 | 134.3215 | 84.1829 | 75.9822 | 115.8437 | 84 | 102.4601 | 128.4778 | 178.9467 | 96.5488 | 242.5809 | 263.9645 | 338.2949 | 268.9645 | 183.9556 | 242.1563 | 333.5399 | 292.8614 | 437.8082 | 298.286 | 384.8082 | 450.643
2348 | 27.8224 | 224.6239 | 264.5537 | 314.2479 | 183.9297 | 161.4131 | 185.4463 | 215.9877 | 155.8348 | 159.9753 | 154.9421 | 155.7398 | 156.2232 | 233.5043 | 178.6116 | 221.095 | 252.3428 | 185.2934 | 165.6819 | 177.9174 | 186.7521 | 166.8927 | 426.132 | 528.3637 | 153.8101 | 274.8015 | 281.1282 | 374.5413 | 313.8348 | 110.3637 | 138.7768 | 181.6819 | 150.8471 | 164.8471 | 161 | 244 | 111.0247 | 336.8594 | 328.8015 | 463.8471 | 278.9297 | 143.2934 | 158.6819 | 208.4463 | 132.3305 | 94.7521 | 87.9877 | 100.6116 | 128.1529 | 226.9297 | 294.095 | 384.1899 | 204.095 | 176.6486 | 144.5043 | 201.7768 | 118.8348 | 209.4008 | 204.9174 | 293.434 | 190.3884 | 167.1529 | 144.5869 | 305.1985 | 187.1652 | 212.2108 | 237.2934 | 319.2934 | 281.8718 | 186.4008 | 194.0703 | 371.7892 | 226.2849 | 215.7521 | 225.3637 | 302.4834 | 355.3761 | 192.9544 | 199.529 | 285.4008 | 265.1776 | 165.1406 | 183.9877 | 244.7521 | 149.6448 | 205.2934 | 189.0123 | 309.2355 | 137.1861 | 141.566 | 183.6486 | 225.3884 | 153.6239 | 322.2934 | 363.529 | 486.963 | 424.471 | 328.2811 | 361.8594 | 534.1985 | 441.6363 | 163.529 | 85 | 123.095 | 163.529 | 128.3637 | 118.8224 | 190.566 | 134.7768 | 163.7892 | 146.0826 | 195.0703 | 121.1406 | 91.0826 | 93.4834 | 100.905 | 100.095 | 136.1197 | 125.1529 | 161.9877 | 106.2355 | 299.0741 | 280.3305 | 311.9753 | 261.6116 | 193.4008 | 224.905 | 324.4463 | 279.0579 | 436.6942 | 295.5413 | 417.5783 | 454.7768
2349 | 27.8334 | 197.64 | 260.9842 | 306.6873 | 173.8279 | 148.8121 | 202.7661 | 243.2194 | 166.8909 | 151.703 | 155.4218 | 165.6097 | 184.2194 | 222.8436 | 186.5939 | 196.4994 | 261.4376 | 195.8909 | 181.1879 | 189.8909 | 196.0315 | 168.8752 | 394.6085 | 493.2024 | 173.7818 | 267.9685 | 282.1564 | 401.6412 | 298.2497 | 139.657 | 123.6715 | 188.5939 | 132.9527 | 156.4061 | 163.5782 | 219.937 | 117.5782 | 302.9054 | 312.5151 | 461.4218 | 268.8279 | 164.2037 | 168.1721 | 231.7818 | 133.4218 | 117.7818 | 96.3127 | 118.0473 | 116.8279 | 223.703 | 239.4206 | 383.9842 | 182.937 | 138.1709 | 157.9067 | 196.1248 | 112.703 | 193.0933 | 216.0315 | 329.3915 | 174.6715 | 153.2497 | 155.4691 | 310.3127 | 173.5309 | 213.297 | 230.703 | 344.5006 | 285.1406 | 176.1091 | 188.703 | 349.2182 | 234.1406 | 242.2194 | 233.1721 | 314.5939 | 372.0473 | 189.4061 | 181.6715 | 269.0933 | 230.3272 | 180.1879 | 180.2812 | 254.0315 | 173.7976 | 198.703 | 190.1406 | 293.6715 | 159.2194 | 127.9527 | 158.0618 | 220.8436 | 156.7188 | 296.7963 | 349.9685 | 525.6885 | 420.8436 | 300.7963 | 338.2182 | 519.5467 | 398.5927 | 152.5467 | 72.9685 | 103.6557 | 152.5467 | 121.5624 | 114.1406 | 164.9212 | 146.3127 | 142.937 | 149.5782 | 204.3127 | 125.8752 | 94.5782 | 79.8121 | 107.4533 | 94.4061 | 141.2812 | 133.5939 | 169.4533 | 101.8436 | 273.1866 | 265.9527 | 339.7976 | 261 | 195.1406 | 238.3285 | 331.4533 | 276.8594 | 454.1879 | 310.6097 | 371.5769 | 413.0303
2350 | 27.8444 | 192.1283 | 258.2566 | 312.8452 | 171.1283 | 173.0223 | 218.2036 | 207.0308 | 148.0795 | 137.9247 | 155.8717 | 155.7964 | 158.491 | 213.2831 | 180.1548 | 198.1018 | 234.9777 | 191.8982 | 188.3584 | 202.4602 | 199.7434 | 166.5133 | 403.6903 | 481.0265 | 171.8982 | 252.9247 | 270.7964 | 398.1548 | 329.9957 | 131.9247 | 143.6638 | 172.5663 | 131.7434 | 166.3319 | 146.5663 | 188.9777 | 105.7964 | 293.5133 | 349.9692 | 480.1771 | 267.8717 | 147.9512 | 172.6151 | 221.1813 | 123.4115 | 107.1813 | 84.053 | 96.5928 | 110.6416 | 220.3849 | 260.509 | 340.2874 | 185.9735 | 147.562 | 191.509 | 169.7211 | 120.7169 | 182.1548 | 229.3319 | 309.8494 | 161.5398 | 160.5885 | 156.2566 | 273.6458 | 162.2831 | 218.3584 | 244.8187 | 316.7476 | 253.6193 | 175.7434 | 201.0753 | 351.1018 | 227.8982 | 226.9512 | 225.4115 | 305.5398 | 355.8229 | 193.2301 | 181.6151 | 247.6946 | 237.0753 | 176.8982 | 204.4072 | 256.8717 | 192.947 | 210.2036 | 211.7921 | 279.6681 | 153.5398 | 124.1283 | 141.6681 | 192.1061 | 168.3319 | 289.5133 | 367.1771 | 526.8982 | 398.2079 | 317.0488 | 332.2566 | 542.4072 | 440.6861 | 174.5355 | 92.7921 | 114.8187 | 174.5355 | 121.8717 | 109.6416 | 153.8982 | 126.2079 | 142.4867 | 135.1813 | 170.2609 | 140.0753 | 87.1548 | 81.3584 | 98.5398 | 99.1018 | 148.8452 | 136.7434 | 165.7699 | 88.7964 | 249.5663 | 302.2259 | 345 | 301.0975 | 219.4072 | 235.7699 | 318.1813 | 260.438 | 422.1326 | 299.9247 | 408.3277 | 417.3319
2351 | 27.8554 | 191.116 | 262.4199 | 314.884 | 203.7071 | 172.5801 | 181.989 | 223.9834 | 152.9558 | 176.6629 | 135.6685 | 145.1602 | 157.5359 | 212 | 180.768 | 194.5801 | 208.9006 | 205.1436 | 158.0608 | 196.9282 | 184.9724 | 159.8122 | 465.2264 | 472.9282 | 162.1602 | 260.7237 | 284.9116 | 395.232 | 312.0166 | 123.8122 | 136.3923 | 164.6961 | 132.884 | 169.768 | 191.7347 | 220.3591 | 95.1602 | 308.9116 | 292.3537 | 509.5193 | 302.4751 | 138.8122 | 157.9724 | 213.6961 | 111.3923 | 91.7403 | 88.1878 | 94.768 | 103.8122 | 252.7071 | 257.9282 | 349.0276 | 184.3481 | 134.0884 | 165.9448 | 178.3757 | 125.5359 | 193.3757 | 222.1602 | 263.5691 | 179.4475 | 156.6961 | 172.7956 | 300.7071 | 168.9558 | 181.873 | 216.9448 | 320.8398 | 264.0276 | 163.6243 | 195.9282 | 358.1878 | 211.9724 | 212.5083 | 230.1878 | 311.9558 | 336.2044 | 178.0884 | 194.3757 | 245 | 246.0718 | 177.768 | 185.0166 | 247.2763 | 163.1768 | 221.7237 | 227.3757 | 248.8288 | 144.0442 | 128.4199 | 143.5359 | 210.9834 | 203.5911 | 293.4199 | 370.884 | 494.1768 | 387.9282 | 325.3039 | 376.1988 | 511.5249 | 418.8288 | 177.116 | 89.8122 | 104.6243 | 177.116 | 135.2597 | 122.2597 | 153.884 | 120.3481 | 138.5801 | 165.7071 | 173.8398 | 132.2763 | 91.3039 | 82 | 97 | 99.116 | 130.5525 | 110.4807 | 146.4365 | 85.232 | 262.0276 | 284.1326 | 306.989 | 270.757 | 195.5967 | 257.0994 | 297.4365 | 276.5635 | 386.0608 | 320.9834 | 383.1768 | 407.5083
2352 | 27.8664 | 207.1328 | 256.7261 | 327.5477 | 191.8672 | 162.1411 | 184.1701 | 227.8963 | 150.4149 | 163.1784 | 131.2075 | 159.2366 | 148.1411 | 207.5187 | 165.7634 | 201.1701 | 223.0291 | 212.3776 | 153.1037 | 199.5851 | 179.4149 | 176.0291 | 421.0165 | 504.2656 | 165.4813 | 232.4232 | 245.7717 | 335.8463 | 321.5477 | 131.9627 | 160.0955 | 184.6142 | 145.5477 | 179.8589 | 156.7717 | 200.8008 | 96.6888 | 272.4605 | 312.6806 | 441.2986 | 284.5933 | 157.7179 | 162.2739 | 199.556 | 129.7179 | 88.2075 | 88.1037 | 97.6888 | 129.888 | 214.8754 | 239.0746 | 359.0664 | 161.5933 | 144.5477 | 166.4813 | 195.2366 | 127.7925 | 214.7179 | 200.3858 | 321.635 | 181.1037 | 154.2075 | 172.3112 | 271.8381 | 184.3403 | 189.5477 | 251.5395 | 334.5477 | 230.1493 | 170.9627 | 186.9336 | 330.3194 | 229.7179 | 215.4813 | 240.8589 | 321.0664 | 374.332 | 202.888 | 193.3112 | 260.2366 | 231.7634 | 177.1037 | 177.5187 | 213.7344 | 165.2739 | 196.112 | 223.6224 | 255.7552 | 151.0664 | 124.5187 | 141.3112 | 205.0373 | 233.0955 | 329.8507 | 377.2739 | 488.2075 | 394.1701 | 302.697 | 339.8754 | 483.697 | 453.5395 | 156.3858 | 95.2739 | 102.1037 | 156.3858 | 139.6888 | 135.6515 | 160.2739 | 117.3112 | 137.1037 | 135.0456 | 177.6888 | 129.2075 | 85.7261 | 96.3403 | 96.1037 | 81.9709 | 106.4896 | 106.1037 | 163.7179 | 81.4149 | 279.2366 | 263.9709 | 311.8589 | 299.1619 | 183.0373 | 248.3485 | 285.1411 | 244.0456 | 398.1328 | 286.3568 | 422.0209 | 420.3403
2353 | 27.8774 | 175.3833 | 279.6225 | 337.177 | 217.2568 | 161 | 191.3599 | 241.6284 | 159.9941 | 163.7257 | 113.7374 | 152.823 | 122.4689 | 215.177 | 196.7081 | 200.1829 | 256.7996 | 171.2063 | 173.8969 | 190.0059 | 191.7198 | 163.4631 | 400.4631 | 461.6635 | 167.8171 | 248.0797 | 232.823 | 368.9766 | 317.5486 | 113.9203 | 124.8405 | 162.4689 | 146.0914 | 144.6576 | 150.1829 | 211.6284 | 94.2743 | 262.5486 | 286.0176 | 435.7257 | 308.3482 | 159.0914 | 159.3658 | 213.4455 | 138.3599 | 88.9086 | 94.3599 | 103.4514 | 136.6342 | 224.5369 | 244.2685 | 325.4748 | 159.9086 | 142.3658 | 148.8288 | 193.3658 | 124.3658 | 201.5545 | 208.9027 | 287.2063 | 187.3599 | 163.9941 | 163.823 | 303.4338 | 171.4631 | 196.4514 | 261.4514 | 325.0973 | 231.4514 | 168.3658 | 184.1829 | 340.6284 | 235.6342 | 199.6459 | 226.5545 | 320.1829 | 315.4009 | 197.823 | 185.7315 | 256.5486 | 243.6284 | 178.8171 | 178.8171 | 208.1829 | 152.3716 | 219.3482 | 185.7491 | 261.5428 | 129.286 | 143.0797 | 150.0856 | 184.9203 | 226.9144 | 324.0059 | 346.2004 | 478.9144 | 414.9883 | 303.6342 | 343.177 | 536.4221 | 386.2239 | 124.0176 | 76.9203 | 85.6459 | 124.0176 | 130.9144 | 133.3658 | 150.0973 | 121.5428 | 138.8171 | 125.5486 | 173.4572 | 129 | 98.6284 | 73.4689 | 85.0973 | 107.2568 | 124.8969 | 97.823 | 134.2004 | 88.2685 | 270.0973 | 264.7257 | 292.1031 | 292.0973 | 171.0973 | 212.4748 | 296.7198 | 256.3541 | 375.4689 | 299.2626 | 421.5486 | 402.0117
2354 | 27.8884 | 196.8629 | 235.0367 | 345.3668 | 194.2162 | 160.0792 | 198.4459 | 222.7413 | 161 | 159.3958 | 142.388 | 151.0792 | 143.0212 | 216 | 167.7703 | 226.7046 | 247.1081 | 216.7259 | 186.1293 | 195.4459 | 192.0792 | 166.6042 | 405.4459 | 504.8842 | 158.7915 | 244.4749 | 273.4382 | 399.7046 | 339.1004 | 115.6834 | 124.6834 | 167.3668 | 146 | 164.0212 | 141.7124 | 210.2375 | 99.5251 | 291.4672 | 296.8127 | 492.1718 | 307.3166 | 159 | 148.8707 | 223.2876 | 117.8205 | 77.9498 | 95.9208 | 113.2085 | 107.5328 | 231.5251 | 265.2587 | 339.4961 | 180.2587 | 161.3378 | 170.9421 | 183.7915 | 148.8629 | 168.6911 | 266.1718 | 288.5251 | 156.6911 | 163.1583 | 163.9208 | 266.4826 | 178.2876 | 196.0792 | 240.8205 | 347.0212 | 225.5541 | 159.7124 | 183.0792 | 355.8127 | 218.5039 | 194.3166 | 264.5965 | 301.583 | 309.9208 | 169.3745 | 173.029 | 249.5541 | 239.4749 | 172.5541 | 165.1873 | 230.1004 | 163.8919 | 189.7703 | 170.9498 | 231.612 | 132.5251 | 131.1873 | 151 | 166.4247 | 167.0656 | 305.5039 | 352.2085 | 490.8919 | 386.612 | 322.417 | 364.2587 | 524.5039 | 369.7915 | 139.417 | 84.2085 | 119.9131 | 139.417 | 142.8919 | 152.3378 | 161.8919 | 127.5251 | 135.3166 | 136.0502 | 196.9421 | 111.5039 | 80.6622 | 80.2085 | 102.417 | 85.1371 | 103.0579 | 122.7838 | 157.7046 | 89.9208 | 284.6544 | 206.9865 | 300.1293 | 304.8127 | 203.1506 | 245.834 | 275.8996 | 271.8127 | 383.1293 | 282.583 | 397.0579 | 390.7915
2355 | 27.8994 | 177.5378 | 239.3983 | 331.0697 | 184.5349 | 160 | 190.6017 | 203.2703 | 161 | 151.5349 | 140.3343 | 127.6715 | 151.532 | 216.9331 | 172.4651 | 198.2063 | 224.5378 | 192.0726 | 172.0697 | 184.8023 | 162.1395 | 205.2588 | 382.6715 | 447.4127 | 165.4651 | 245.8663 | 278.8663 | 357.2092 | 306.4738 | 130.9303 | 124.0669 | 169.8663 | 137.6017 | 186.5291 | 168.0611 | 214.6657 | 110.2646 | 267.8721 | 304.532 | 457.8081 | 278.0726 | 143.1366 | 123.7383 | 210.936 | 109.468 | 93.7965 | 90.4012 | 93.4709 | 121.7965 | 209.6046 | 242.7383 | 332.6017 | 168.936 | 143.404 | 170.2006 | 180.2006 | 138.8692 | 171.5988 | 211.279 | 289 | 178.2617 | 167.6657 | 151.8692 | 265.7994 | 142.6075 | 192.2674 | 219.404 | 354.5988 | 225 | 170.1977 | 187.6657 | 367.2646 | 200.2035 | 221.0611 | 229.7412 | 295.3343 | 278.2732 | 177.2646 | 210.2588 | 235.936 | 219.404 | 185.9971 | 158.4012 | 193.7412 | 164.0669 | 192.5988 | 205.4594 | 249.5291 | 152.596 | 120.6686 | 132.3372 | 186.4622 | 157.3343 | 288.1366 | 372.596 | 469.6046 | 406.3954 | 340.7965 | 339.8721 | 520.2006 | 389.5291 | 132.6017 | 98.9971 | 99.6715 | 132.6017 | 144 | 123.2063 | 180.7297 | 125.2006 | 138.7326 | 132.3343 | 176.6046 | 133.3285 | 86.4651 | 87.532 | 110.532 | 81.1337 | 114.9971 | 125.9331 | 130.1395 | 66.6715 | 243.0755 | 267.3199 | 295.4012 | 274.2732 | 185.4709 | 221.0058 | 266.5349 | 261.8023 | 377.468 | 308.0611 | 370.7383 | 379.7354
2356 | 27.9104 | 193.0178 | 264.5812 | 330.9454 | 179.2728 | 154.3274 | 185.2728 | 217.1269 | 147.764 | 181.2538 | 127.7094 | 159.0901 | 161.4543 | 217 | 186.236 | 196 | 235.2906 | 201.3452 | 171 | 201.9632 | 179.8541 | 195.7094 | 411.2538 | 464.7449 | 138.5825 | 256.3998 | 259.1459 | 339.8185 | 290.764 | 120.6548 | 138.1815 | 176.618 | 128.4911 | 182.3274 | 158.6548 | 193.2551 | 97.764 | 297.1993 | 350.3807 | 446.4911 | 264.6548 | 149.5635 | 179.6714 | 197.7094 | 115.618 | 83.6548 | 86.2183 | 88.2183 | 100.3096 | 200.4365 | 246.6726 | 326.3274 | 158.5457 | 146.7272 | 192.6904 | 185.6726 | 119.0914 | 185.236 | 207 | 272.9277 | 167.7094 | 150.0368 | 172.7449 | 248.0368 | 168.363 | 199.5635 | 221.7817 | 331.3642 | 229.7272 | 154.9277 | 167.2005 | 313.1649 | 211.2906 | 226.7817 | 218.4911 | 277.0368 | 317.599 | 174.2183 | 213 | 259.5812 | 222.7272 | 177.5457 | 154.2183 | 208.9632 | 159.2728 | 199.618 | 180.5825 | 264.236 | 136.9822 | 125.6726 | 148.0178 | 205.0178 | 157 | 289.8363 | 381.5635 | 482.1815 | 405.1637 | 310.8007 | 355.9632 | 503.9277 | 410.8541 | 153.7449 | 92.4365 | 105.5635 | 153.7449 | 122.2551 | 122.8909 | 136.6193 | 141.0723 | 116.3096 | 141.4543 | 164.6002 | 130.2728 | 101.1815 | 87.0546 | 93.0368 | 87.618 | 117.8909 | 110.8731 | 128 | 75.3998 | 266.4721 | 244.5825 | 327.1447 | 270.1091 | 191.5635 | 235.0723 | 279.236 | 261 | 388.3452 | 292.9822 | 359.5457 | 401.6904
2357 | 27.9214 | 178.6764 | 267.9154 | 329.0846 | 179.9577 | 168.3658 | 168.7187 | 232.3658 | 160.4081 | 144.6911 | 119.3382 | 138.9724 | 159.1268 | 195.9301 | 169.761 | 204.6195 | 210.1415 | 173.2683 | 177.7041 | 181.9301 | 154.1838 | 161.4797 | 387.1415 | 469.8309 | 161.9008 | 261.7886 | 249.3805 | 347.6195 | 303.4081 | 135.3236 | 114.0992 | 173.1691 | 135.6618 | 181.0423 | 156.0846 | 197.7463 | 97.9577 | 281.761 | 302.2407 | 412.4797 | 305.1821 | 157.6618 | 151.3951 | 195.0846 | 124.6195 | 91.6195 | 100.3658 | 107.1545 | 100.9154 | 234.478 | 267.1122 | 309.7187 | 145.5496 | 133.5919 | 148.987 | 209.9431 | 128.535 | 203.239 | 192.6342 | 306.478 | 167.9577 | 148.0423 | 172.0846 | 235.5073 | 146.0569 | 197.1268 | 239.239 | 296.4797 | 225.2114 | 145.3805 | 154.5073 | 322.4504 | 190.9301 | 215.5073 | 267.8016 | 304.7317 | 323.8309 | 181.6618 | 188.0992 | 256.2114 | 244.0699 | 161.6764 | 175.0699 | 205.2114 | 157.0846 | 179.8878 | 212.5203 | 248.7187 | 121.6342 | 104.9301 | 127.9301 | 197.3805 | 175.1967 | 297.6618 | 348.4797 | 497.3658 | 416.4927 | 311.8732 | 304.3252 | 464.6911 | 379.4374 | 126.2683 | 94.8732 | 109.8309 | 126.2683 | 130.5772 | 142.1545 | 152.1967 | 101.7756 | 120.7463 | 125.7187 | 137.1838 | 122.3382 | 86.6764 | 95.6195 | 90.0846 | 97.5772 | 95.9724 | 126.2813 | 135.6618 | 88.4504 | 275.6618 | 227.6764 | 309.8455 | 298.7317 | 185.2959 | 225.465 | 268.5073 | 271.535 | 412.9431 | 311.1545 | 363.7886 | 390.5496
2358 | 27.9324 | 206.1304 | 259.2699 | 320.2699 | 167.3898 | 170.94 | 192.2504 | 223.2999 | 155.1799 | 143 | 128.7001 | 172.9205 | 163.8501 | 192.09 | 187.4303 | 209.8501 | 214.8201 | 168.1199 | 190.6102 | 187.7901 | 177.2504 | 182.3103 | 379.2099 | 491.3403 | 157.1799 | 278.4902 | 287.8006 | 372.2504 | 311.7601 | 152.4902 | 139.1904 | 165.2399 | 142.7901 | 145.1095 | 136.5997 | 209.6402 | 101.8801 | 323.6806 | 289.3298 | 455.6206 | 314.7601 | 171.5802 | 141.2699 | 193.06 | 122.09 | 97.8201 | 86.4498 | 106.06 | 115.5502 | 173.9191 | 234.0495 | 338.1004 | 202.2308 | 162.1004 | 132.4498 | 196.4498 | 123.1799 | 186.5397 | 192 | 303.1499 | 175.7601 | 142.1799 | 169.09 | 235.97 | 147.91 | 183.4198 | 223.5098 | 304.7001 | 227.91 | 168.2803 | 216.0809 | 353.0704 | 198.7301 | 227.6102 | 234.1095 | 301.1499 | 320.1199 | 182 | 200.5802 | 279.2803 | 258.5802 | 153.2399 | 163.3898 | 194.3298 | 187.0704 | 167.3598 | 195.5697 | 212.1095 | 139.4303 | 121.4603 | 145.4303 | 180.5098 | 173.09 | 330.0105 | 329.5397 | 434.9491 | 399.5397 | 317.8201 | 340.8006 | 476.5802 | 401.2803 | 157.0105 | 65.8996 | 82.8396 | 157.0105 | 99.9595 | 107.1095 | 114.1994 | 111.6402 | 131.6702 | 117.2399 | 146.6702 | 110.3598 | 93.7601 | 87.2699 | 98.7301 | 96.06 | 105.6702 | 99.8396 | 127.2699 | 87.06 | 311.8905 | 244.4603 | 341.9805 | 267.0195 | 217.0105 | 230.8201 | 271.8801 | 257.4498 | 382.9595 | 291.6297 | 358.1799 | 423.9505
2359 | 27.9434 | 186.3716 | 266.8584 | 299.3716 | 191.5576 | 150.3716 | 188.0885 | 211.2123 | 152.0531 | 142.0177 | 129.9823 | 166.1416 | 147.3008 | 212.6284 | 164.4247 | 180.5308 | 218.9292 | 182.7346 | 163.4955 | 167.3716 | 177.0177 | 168.2654 | 379.9823 | 500.8407 | 143.2477 | 265.2477 | 263.4601 | 346.4778 | 321.8231 | 132.3716 | 126.2477 | 173.8407 | 146.9292 | 166.593 | 131.0885 | 204.1062 | 99.0531 | 269.0086 | 309.6284 | 480.5753 | 294.3716 | 138.6016 | 152.7877 | 199.8761 | 158.3453 | 98 | 76.1769 | 88.3185 | 124.8407 | 240.7614 | 228.0885 | 280.0617 | 173.5485 | 129.6016 | 176.2037 | 175.3716 | 137.7346 | 179.1239 | 186.1062 | 287.2831 | 149.4778 | 130.2123 | 158.1946 | 225.1946 | 133.2654 | 193.8054 | 224.9646 | 307.9469 | 201.4778 | 157.2123 | 169.867 | 304.8847 | 200.9646 | 202.4601 | 214.3362 | 297.0708 | 316.0708 | 174.1416 | 187.2477 | 266.2477 | 238.3716 | 191.3099 | 175.77 | 221.5045 | 146.7432 | 192.5399 | 155.7078 | 227.6992 | 139.0177 | 138.6992 | 140.1062 | 212.4161 | 167.1062 | 314.3008 | 343.7346 | 436.9292 | 399 | 290.4955 | 360.6638 | 447.5308 | 383.3362 | 144.2477 | 76.7877 | 90.8407 | 144.2477 | 115.6992 | 108.9469 | 161.133 | 130.6638 | 120.2123 | 120.9292 | 175.4868 | 130.6284 | 74.3539 | 81.1062 | 73.4601 | 97.9646 | 130.5576 | 119.6284 | 141.7346 | 85.0354 | 271.7432 | 267.593 | 294.867 | 309.2214 | 161.0263 | 216.2654 | 246.4601 | 231.4601 | 365.3008 | 291.9823 | 334.4247 | 379.814
2360 | 27.9544 | 175.0594 | 243.1297 | 294.027 | 159.1783 | 145.027 | 171.0919 | 194.0919 | 158.9622 | 158.9081 | 153.8703 | 160.0324 | 140.0378 | 236.8703 | 175.9352 | 199.8919 | 216.0162 | 155.1513 | 169.9622 | 181.9189 | 164.0702 | 159.0486 | 382.9838 | 471.1621 | 153.9406 | 237.1513 | 261.0108 | 386.7784 | 319.0162 | 109.1243 | 105.1135 | 150.1297 | 133.0757 | 180.9243 | 158.8487 | 192.0648 | 86.0702 | 295.8487 | 290.1081 | 418.3404 | 294.9946 | 161.8703 | 153.9946 | 214.9189 | 111.2594 | 70.1513 | 79.9784 | 88 | 115.054 | 220.1189 | 211.0919 | 314.8055 | 174.9892 | 146.9027 | 173.0216 | 183.9514 | 117.1135 | 187.9514 | 233.7406 | 304.9027 | 171.8757 | 145.9135 | 142.0865 | 248.8703 | 153.8865 | 198.973 | 214.0594 | 288.1081 | 240.7838 | 165.9514 | 198.8379 | 346.7676 | 197.0216 | 227.8595 | 199.0811 | 268.1567 | 277.2107 | 187.9243 | 177.054 | 220.2486 | 213.1351 | 144.2594 | 163.0702 | 188.1837 | 142.0216 | 169.1297 | 169.9189 | 194.1837 | 120.1027 | 134.027 | 145.9676 | 181.1729 | 155.0648 | 308.0324 | 340.0216 | 486.7298 | 369.1621 | 305.9135 | 310.2756 | 446.0054 | 371.0648 | 149.9676 | 68.0486 | 93.9838 | 149.9676 | 112.0216 | 113.973 | 158.0216 | 118.0702 | 129.946 | 111.054 | 139.1999 | 105.1405 | 90.9081 | 75.0324 | 102.8379 | 93.027 | 107.1297 | 124.973 | 116.1405 | 90.9676 | 265.0324 | 220.2594 | 283.0594 | 283.1459 | 161.9892 | 190.1405 | 283.7947 | 270.7838 | 392.8487 | 303.9352 | 373.7838 | 359.1081
2361 | 27.9654 | 200.8003 | 263.1515 | 328.124 | 189.8209 | 159.938 | 165.9862 | 231.9725 | 169.876 | 144.0413 | 123.0895 | 135.0895 | 159.9036 | 199.1309 | 187.9242 | 197.938 | 187.1584 | 165.0207 | 149.0826 | 180 | 188.9656 | 169.9862 | 412.5386 | 441.186 | 174.8072 | 240.8967 | 278.0344 | 330.1584 | 273.9725 | 138.8209 | 128.9311 | 167.9725 | 126.1515 | 146.0344 | 179.8898 | 193.9449 | 96.0207 | 285.9036 | 300.8967 | 428.9174 | 310.8485 | 148.0482 | 138.8898 | 198.9931 | 135.062 | 84.0207 | 78.9862 | 92.0344 | 117.9587 | 206.1515 | 226.9587 | 358.7107 | 169.0207 | 149.9725 | 155.8416 | 202.7865 | 107.0344 | 185.0689 | 201.0482 | 274.1791 | 169.0689 | 142.0275 | 173.8623 | 214.2066 | 166.9174 | 210.0207 | 193.1446 | 330.7727 | 231.2479 | 160.9105 | 158.1997 | 324.9725 | 226.7934 | 205.8829 | 238.0413 | 307.8829 | 356.7934 | 185.938 | 171.124 | 258.7796 | 218.1584 | 144.9725 | 135.1653 | 182.0413 | 174.7658 | 181.0758 | 189.9105 | 256.9174 | 137.9036 | 157.876 | 142.0069 | 177.1033 | 135.0689 | 322.1102 | 354.814 | 470.9587 | 352.303 | 278.1377 | 316.2342 | 479.8485 | 350.3444 | 144.9518 | 90.9725 | 103.8898 | 144.9518 | 130.9105 | 111.0895 | 138.0207 | 128.8347 | 114.0344 | 124.9036 | 159.8829 | 125.0482 | 89.9311 | 79.0895 | 80.1171 | 108.9174 | 112.0275 | 104.062 | 130.0275 | 91.9862 | 243.3995 | 245.0344 | 322.7383 | 265.1309 | 178.9656 | 186.1928 | 268.8485 | 274.7452 | 377.365 | 327.8347 | 360.8278 | 363.4614
2362 | 27.9764 | 172.1151 | 284.9041 | 346.1342 | 164.4027 | 150.5397 | 164.6521 | 228.2685 | 151.7123 | 149.8658 | 136.0192 | 148.0192 | 145.5589 | 217.9041 | 176.674 | 189.0192 | 209.7123 | 168.2877 | 160.9808 | 179.4822 | 184.5562 | 167.8082 | 346.4986 | 467.9233 | 147.6712 | 226.326 | 282.5397 | 353.4027 | 270.9973 | 113.2685 | 119.1918 | 163.7699 | 147.9041 | 151.1534 | 163.6164 | 186.7863 | 98.7507 | 272.2493 | 286.4986 | 417.0575 | 289.3644 | 155.211 | 123.8055 | 198.1342 | 143.5397 | 87 | 77.4603 | 96.789 | 112.1534 | 227.6548 | 221.0192 | 317.2685 | 171.6548 | 145.5589 | 133.211 | 172.1342 | 111.9041 | 194.674 | 208.1151 | 299.0027 | 178.674 | 146.2301 | 153.6931 | 243.8658 | 154.7123 | 213.2301 | 214.3069 | 298.211 | 266.5014 | 148.3644 | 186.9808 | 320.8082 | 197.0575 | 189.2685 | 243.674 | 290.7315 | 326.6931 | 176.9041 | 188.6931 | 227.7671 | 241 | 141.1918 | 158.4247 | 187.7123 | 141.3069 | 191.7507 | 177.3069 | 244.789 | 124.1918 | 140.1534 | 142.5589 | 191.6548 | 144.7123 | 337.5205 | 328.1918 | 464.3096 | 396.1726 | 298.0959 | 349.6931 | 457.5781 | 399.4438 | 137.9425 | 86.8658 | 88.4411 | 137.9425 | 118.537 | 124.3069 | 141.2301 | 105.4986 | 118.7699 | 111.0384 | 142.7699 | 132.1342 | 80.0384 | 91.8274 | 97.1918 | 97.2685 | 115.5781 | 112.8466 | 133.8274 | 90.0384 | 299.8685 | 249.9616 | 285.7671 | 284.3069 | 174.211 | 214.1534 | 246.9616 | 238.1151 | 428.5616 | 303.3863 | 336.9973 | 429.1178
2363 | 27.9874 | 177.8112 | 280.535 | 351.7727 | 185.3776 | 127.4091 | 197.4965 | 240.3321 | 137.0629 | 143.1573 | 136.7797 | 149.1259 | 123.2832 | 212.9371 | 160.1259 | 189.5594 | 195.2832 | 183.1259 | 159.7482 | 153 | 211.7098 | 157.7168 | 372.0315 | 464.8811 | 181.2133 | 243.1573 | 257.8356 | 372.5209 | 320.8356 | 126.2447 | 129.535 | 152.7553 | 142.9056 | 159.2203 | 144.1259 | 226.1503 | 85.8112 | 284.7482 | 310.8671 | 420.2832 | 307.4336 | 165.3706 | 163.5839 | 204.6853 | 120.1888 | 87.1259 | 100.3077 | 86.1259 | 120.3462 | 210.9126 | 222.3462 | 329.9615 | 154.472 | 123 | 144.5664 | 179.2518 | 107.9126 | 178.535 | 213.4336 | 249.3532 | 161.528 | 157.8741 | 138.3462 | 237.1259 | 139.9371 | 224.1189 | 229.4021 | 309.5035 | 241.1888 | 167.3147 | 184.93 | 311.7238 | 199.8741 | 203.3776 | 226.3391 | 277.4406 | 311.2518 | 172.8182 | 172.6224 | 266.6853 | 240.4965 | 151.9441 | 129.8497 | 174.1958 | 156.4021 | 178.7168 | 193.3147 | 234.2518 | 133.9371 | 147.0874 | 121.2902 | 173.9371 | 130.1573 | 312.9056 | 336.2692 | 430.4791 | 404.6853 | 303.2518 | 333.5909 | 436.3776 | 370.8427 | 134.4021 | 80.3776 | 109.6783 | 134.4021 | 144.993 | 139.8112 | 152.2133 | 130.6224 | 107.0629 | 113.1888 | 131.5664 | 137.93 | 81.7797 | 82.6224 | 106.6853 | 110.3077 | 94.6923 | 105.1888 | 125.2518 | 91.4021 | 242.6923 | 248.2518 | 324.7797 | 298.993 | 184.7168 | 221.465 | 244.8741 | 244 | 356.3847 | 272.8497 | 387.3391 | 383.5909
2364 | 27.9984 | 172.8752 | 298.4441 | 314 | 195.9497 | 140.9627 | 183.094 | 191.5819 | 139.919 | 148.2626 | 130.2626 | 153.0438 | 132.3063 | 212.1378 | 164.0438 | 176.6126 | 203.2998 | 187.2188 | 151.8687 | 153.5251 | 171.6499 | 149.7877 | 372.3436 | 490.2058 | 157.4814 | 246.2058 | 223.1815 | 326.7812 | 285.7002 | 103.7439 | 144.6434 | 176.4814 | 140.175 | 165.6062 | 147.825 | 199.5186 | 80.3501 | 277.3063 | 277.3566 | 427.5559 | 288.8622 | 145.3436 | 121.094 | 194.3874 | 125.2123 | 90.2998 | 79.4814 | 89.5186 | 129.9497 | 236.4619 | 232.8687 | 297.9125 | 169.1313 | 123.175 | 161.7374 | 186.7374 | 134.7309 | 194.8687 | 195.5624 | 291.2188 | 146.8687 | 153.3874 | 148.7374 | 239.2058 | 138.5251 | 197.3501 | 210.6062 | 324.7374 | 246.7374 | 176.4749 | 153.9692 | 333.2123 | 196.0875 | 214.8687 | 206.5689 | 289.906 | 318.6062 | 196.7747 | 161.4376 | 256.4749 | 224.2998 | 179.8622 | 155.2998 | 210.6062 | 138.6126 | 170.4376 | 202.0373 | 241.4749 | 132.3938 | 118.7812 | 159.9935 | 172.4814 | 134.6937 | 307.987 | 285.1005 | 476.1313 | 393.7309 | 311.6126 | 320.7812 | 449.1815 | 365.6937 | 116.2626 | 91.9562 | 69.7002 | 116.2626 | 114 | 133.6499 | 128.2626 | 119.5689 | 109.0438 | 118.8687 | 149.3063 | 105.0875 | 74.6062 | 70.825 | 97.0438 | 89.3501 | 116.1313 | 111 | 132.2561 | 73.7877 | 264.5251 | 254.906 | 317.3874 | 269.094 | 176.7877 | 205.2188 | 242.1378 | 244.3063 | 397.6872 | 299.3501 | 366.1686 | 371.8314
2365 | 28.0094 | 192.0561 | 331.9618 | 312.0942 | 172.5516 | 162.1121 | 205.4306 | 245.7019 | 158.5987 | 155.121 | 136.2242 | 152.3745 | 139.6726 | 234.3095 | 166.121 | 190.3924 | 190.5223 | 191.7197 | 148.4955 | 165.2803 | 163.5516 | 167.1682 | 359.5134 | 452.2331 | 166.7669 | 208.7376 | 248.2713 | 322.0561 | 301.0561 | 121.8497 | 117.3542 | 186.5516 | 143.5516 | 156.2153 | 144.6726 | 189.7847 | 88.5045 | 284.2242 | 307.3924 | 397.9058 | 265.065 | 131.5045 | 144.4955 | 181.5045 | 109.121 | 75.3924 | 90.1682 | 79 | 109.4663 | 182.6815 | 229.1032 | 296.1121 | 171.4955 | 127.3924 | 156 | 180.1592 | 107.4484 | 191.0471 | 186.3924 | 294.6548 | 144.8968 | 140.3924 | 143.2242 | 201.2331 | 149.0471 | 203.7108 | 202.8408 | 318.3274 | 241.2803 | 163.879 | 195.5427 | 314.8229 | 199.6815 | 212.5045 | 221.5223 | 267.8497 | 310.5605 | 171.1771 | 171.6726 | 242.87 | 209 | 155.4484 | 140.7847 | 200.2064 | 151.7197 | 180.6726 | 179.2064 | 229.2153 | 140.2153 | 114.8408 | 137.5516 | 183.0561 | 127.6637 | 265.121 | 328.87 | 475.9172 | 366.4484 | 323.4866 | 315.4395 | 474.7197 | 360.5694 | 122.2242 | 90.0471 | 85.3924 | 122.2242 | 114.1682 | 125.6637 | 133.8318 | 130.9911 | 109.4395 | 115.8318 | 155.4395 | 108.1771 | 66.7847 | 68.5694 | 98.1121 | 97.5045 | 117.935 | 109.3745 | 115.5516 | 91 | 275.8318 | 233.4102 | 304.5045 | 292.6076 | 193.4395 | 211.7376 | 267.9529 | 251.6166 | 369.4484 | 305.3185 | 347.5516 | 390.3363
2366 | 28.0204 | 192.1799 | 365.6834 | 282.187 | 165.8885 | 162.5648 | 177.9065 | 208.8452 | 136.7086 | 173.0432 | 138.3598 | 124.0432 | 149.7015 | 191.7337 | 182.813 | 197.1367 | 230.2663 | 187.3417 | 141.3669 | 169.5216 | 155.9317 | 169.0432 | 383.0864 | 475.187 | 147.4352 | 237.9317 | 235.0432 | 323.6151 | 302 | 151.9065 | 154.813 | 178.1115 | 136.3417 | 144.0935 | 155.5216 | 202.5899 | 96.6583 | 284.2412 | 313.6583 | 429.5899 | 281.2914 | 141.4352 | 136.6151 | 190.5467 | 127.2482 | 81.9317 | 92.3166 | 78.795 | 148.5397 | 211.9568 | 214.3417 | 296.1548 | 163.2734 | 132.7698 | 155.1115 | 167.1618 | 114.5216 | 175.0683 | 192.2482 | 272.7518 | 162.1186 | 146.795 | 146.3166 | 222.8885 | 132.5216 | 183.6402 | 218.9136 | 307.3417 | 245.1799 | 145.4784 | 171.8885 | 297.392 | 226.7698 | 219.7015 | 261.7447 | 296.813 | 318.3598 | 190.3166 | 182.5899 | 208.2302 | 211.392 | 163 | 155.7769 | 171.9819 | 146.5899 | 191.7266 | 150.7769 | 216.4784 | 127.1367 | 128.8633 | 131.5719 | 182.7015 | 124.187 | 284.5467 | 295.187 | 423.6583 | 371.8814 | 297.1115 | 304.9065 | 467.5397 | 386.2482 | 126.2734 | 74.0683 | 93.2985 | 126.2734 | 117.7518 | 120.2734 | 131.205 | 113.5899 | 100.8201 | 113.8885 | 144.5648 | 127.4533 | 80.3417 | 94.0432 | 98.8382 | 104.7698 | 99.7266 | 82.0683 | 110.1186 | 91.7518 | 272.0432 | 271.608 | 313.4784 | 284.5648 | 183.8633 | 237.1045 | 281.198 | 260.3598 | 375.9749 | 274.198 | 339.3849 | 394.3598
2367 | 28.0314 | 180.8387 | 367.5011 | 310.7099 | 176.7099 | 145.1771 | 166.3866 | 233.2743 | 160 | 160.887 | 118.1771 | 111 | 132.9676 | 228.4356 | 155.9834 | 196.5004 | 195.4996 | 190.8711 | 158.0648 | 164.1289 | 157.1771 | 155.5968 | 358.532 | 499.7588 | 165.1454 | 234.1778 | 222.8063 | 330.8711 | 300.1454 | 136.9194 | 126.3708 | 166.3225 | 140.3549 | 158.5968 | 147.3873 | 195.1454 | 91.8387 | 235.419 | 308.4356 | 422.3873 | 261.4672 | 159.3067 | 144.5968 | 198.9676 | 119.5803 | 81.4032 | 83.1613 | 77.8546 | 114.3549 | 221.8553 | 218.4356 | 273.5803 | 166.5968 | 117.532 | 143.5644 | 181.3873 | 109.6127 | 177.9352 | 182.0806 | 281.7905 | 187.2102 | 142.7099 | 137.8063 | 234.6775 | 127.1289 | 202.049 | 242.4197 | 308.8553 | 234.3225 | 155.5479 | 182.3067 | 328.7099 | 209.5968 | 201.2743 | 229.7099 | 270.951 | 297.3708 | 181.9676 | 176.5162 | 227.4996 | 241.9842 | 161.2261 | 176.1296 | 195.694 | 140.8387 | 186.468 | 176.2095 | 222.5162 | 127.6292 | 125.7099 | 151.468 | 162.1778 | 151.9842 | 292.8063 | 327.7415 | 420.1289 | 346.306 | 284.9194 | 292.6609 | 434.7257 | 375.2743 | 129.5968 | 76.9352 | 110.5162 | 129.5968 | 128 | 123.0324 | 133.5968 | 108.3225 | 113.0482 | 125.3549 | 127.4996 | 119.5968 | 84.1937 | 81.4032 | 83.5644 | 89.1289 | 97.3708 | 84.6127 | 138.2743 | 100.1454 | 259.4032 | 240.3708 | 317.0166 | 266.2095 | 183.532 | 188.3542 | 244.7739 | 237.6775 | 362.3225 | 241.4025 | 333.7415 | 375.3866
2368 | 28.0424 | 179.3717 | 281.8142 | 296.4646 | 163.0222 | 167.7698 | 201.7476 | 223.5132 | 159.7212 | 168.3052 | 141.6991 | 113.6019 | 144.3717 | 221.1637 | 186.6547 | 170.0443 | 223.0708 | 177.5354 | 137.8585 | 176.7212 | 180.0486 | 151.9292 | 372.8406 | 444.4604 | 145.6726 | 208.5976 | 234.2302 | 321.9958 | 281.3231 | 134.3274 | 141.9071 | 167.9557 | 134.8585 | 153.6283 | 133.2745 | 175.3009 | 89.8142 | 260.584 | 304.8807 | 403.4425 | 298.584 | 140 | 140.8363 | 208.792 | 146.9557 | 85.9071 | 86.8585 | 98.2566 | 106.5354 | 187.7656 | 211.1637 | 304.7656 | 159.8627 | 135.1858 | 152.2302 | 165.509 | 124.9335 | 197.6769 | 182.8142 | 269.0222 | 144.9292 | 128.8363 | 147 | 229.8849 | 141.208 | 158.3939 | 213.1858 | 270.1193 | 235.9557 | 192.5618 | 164.6726 | 312.792 | 202.1193 | 194.5797 | 216.9514 | 312.5618 | 313.0929 | 192.2566 | 170.1637 | 255.4425 | 219.5575 | 143.1373 | 134.7877 | 161.717 | 139.8363 | 169.9292 | 189.5354 | 213.8406 | 114.6019 | 112.6726 | 132.9778 | 132.8807 | 127.8849 | 301.2566 | 355.0972 | 433 | 380.4689 | 286.7877 | 320.9557 | 440.1193 | 366.5354 | 125.3717 | 97.8849 | 104.0708 | 125.3717 | 125.9557 | 113.0222 | 130.5797 | 111.6283 | 124.0708 | 119.8585 | 153.0264 | 115.5575 | 76.0222 | 85.6283 | 90.4646 | 102.2788 | 113 | 103.7434 | 129.2566 | 79 | 260.8406 | 256.1858 | 287.2745 | 281.0222 | 200.2566 | 234.8406 | 265.9292 | 239.2038 | 367.9514 | 298.9113 | 365.1858 | 409.6325
2369 | 28.0534 | 180.3696 | 280.3156 | 303.2095 | 173.6313 | 150.6295 | 170.4209 | 208.4739 | 155.4218 | 142.8417 | 132.5773 | 139.9469 | 146.527 | 211.2113 | 156.1043 | 189.4739 | 214.4209 | 174.0522 | 157.0522 | 171.4748 | 165.3138 | 160.5791 | 343.7887 | 484.4227 | 160.9478 | 266.2662 | 256.1052 | 358.6853 | 301.58 | 117.5791 | 138.4748 | 147.7896 | 151.8426 | 152.1043 | 171.7383 | 189.9991 | 86.4218 | 241.7347 | 329.7383 | 402.6295 | 273.9478 | 141.9991 | 151.3147 | 194.3696 | 127.6313 | 86.8939 | 103.7374 | 90.3687 | 104.7356 | 207.1043 | 203.7365 | 319.8957 | 139.7365 | 134.8957 | 171.58 | 191.6835 | 98.4721 | 173.6322 | 183.7356 | 278.0531 | 151.3696 | 138.5782 | 145.1061 | 216.1592 | 150.7905 | 173.5782 | 215.4209 | 247.9973 | 214.8426 | 157.9478 | 178.58 | 303.7356 | 178.6295 | 207.58 | 231.8966 | 280.7887 | 315.8939 | 186.1574 | 162.5261 | 246.9487 | 225.4209 | 163.4739 | 160.527 | 192.9496 | 146.6322 | 177.4218 | 183.527 | 189.2077 | 137.2644 | 129.5261 | 121.8426 | 161.3156 | 116.6844 | 294.9469 | 316.1583 | 431.1061 | 347.9991 | 312.1061 | 301.6313 | 414.3147 | 361.4739 | 128.3687 | 87.4209 | 96.7887 | 128.3687 | 106.4209 | 122.3687 | 144.9478 | 108.1052 | 114.0531 | 136.4218 | 124 | 120.4739 | 86.8417 | 82.1052 | 93.9478 | 103.2113 | 111.6322 | 108.4748 | 120.3165 | 78.7896 | 235.5764 | 253.4757 | 326.6853 | 289.1061 | 192.4739 | 205.8939 | 268.4766 | 288.4218 | 387.8417 | 261.2095 | 364.8957 | 367.4209
2370 | 28.0644 | 156.1199 | 281.9424 | 315.3021 | 181.9377 | 184.4772 | 173.765 | 204.8225 | 145.4077 | 147.8849 | 160.1798 | 145.6499 | 134.9401 | 199.0552 | 172.0024 | 182.6499 | 214.0048 | 183.705 | 164.7074 | 151.5276 | 199.1798 | 156.4125 | 359.8225 | 458.5828 | 153.5276 | 213.2302 | 257.1175 | 337.9424 | 282.1751 | 115.1751 | 116.8825 | 148.4676 | 140.4724 | 167.2973 | 145.6451 | 206.0599 | 78.1703 | 283.4101 | 298.4748 | 437.6523 | 269.9353 | 159 | 168.8849 | 175.6427 | 130.5324 | 101.1199 | 91.2374 | 86.0576 | 125.4149 | 222.0623 | 209.4125 | 301.8825 | 148.8202 | 119.8202 | 151.235 | 201.1822 | 126.8249 | 163.6451 | 207.235 | 266.8273 | 134.4053 | 148.5923 | 130.7626 | 194.4676 | 132.3525 | 185.3549 | 216.1798 | 298.6475 | 206.1751 | 149.8225 | 155.4149 | 321.9472 | 215.8849 | 187.5875 | 200.5252 | 297.0576 | 330.5899 | 192.1223 | 165.3549 | 222.2926 | 228.6475 | 158.8825 | 148.235 | 153.6427 | 135.3525 | 163.4125 | 172.1751 | 240.3549 | 118.1127 | 133.295 | 116.2302 | 166.5851 | 114.9401 | 304.2926 | 311.2326 | 414.295 | 365.9401 | 295.53 | 305.1199 | 432.5899 | 353.2398 | 123.1175 | 92.0576 | 108.5923 | 123.1175 | 111.7626 | 116.4125 | 138.9377 | 110.0576 | 107.5276 | 122.295 | 127.7602 | 114.9424 | 92.5899 | 84.5276 | 86.2926 | 89.1751 | 100.235 | 90.9952 | 108.2326 | 76.765 | 278.2997 | 215.5875 | 307.5875 | 277.9353 | 186.8249 | 223.4101 | 219.5804 | 275.235 | 387.127 | 278.4724 | 345.2374 | 367.4748
2371 | 28.0754 | 148.3601 | 275.0384 | 264.351 | 199.4571 | 146.1151 | 174.8556 | 210.8702 | 169.1828 | 132.7788 | 138.0914 | 132.1535 | 142 | 221.6106 | 158.4278 | 164.6106 | 188.5429 | 187.9616 | 153.702 | 165.0768 | 177.2212 | 153.298 | 368.3364 | 490.2742 | 163.4424 | 243.3126 | 257.0914 | 326.1061 | 287.6253 | 121.7934 | 118.7258 | 163.8849 | 127.3126 | 147.3894 | 157.351 | 200.947 | 114.2742 | 293.351 | 275.5045 | 407.7258 | 307.2596 | 156.5338 | 151.5722 | 200.0147 | 114.0768 | 85.4424 | 78.649 | 92.3126 | 105.0914 | 204.0237 | 204.7404 | 299.4424 | 168.1828 | 141.5192 | 152.351 | 170.3217 | 116.1828 | 176.7788 | 208.7404 | 242.351 | 173.9616 | 122.4808 | 142.702 | 211.4424 | 133.1828 | 175.9853 | 198.6343 | 296.3894 | 215.7788 | 153.9232 | 139.2449 | 284 | 202.9853 | 189.9232 | 226.7404 | 300.1975 | 316.4955 | 165.9853 | 156.2449 | 232.9086 | 227.8172 | 156.4424 | 150.649 | 180.7404 | 139.0384 | 160.6874 | 182.1682 | 225.7934 | 146.404 | 124.8849 | 146.9616 | 182.2359 | 121.6106 | 311.7934 | 328.1298 | 411.8556 | 369.6253 | 289.5338 | 295.2833 | 425.8939 | 330.3217 | 125.1682 | 98.8318 | 83.1298 | 125.1682 | 126.1682 | 112.1298 | 158.5338 | 119.5576 | 117.4424 | 116.8702 | 152.2359 | 111.4131 | 83.8172 | 95.2212 | 96.5192 | 96.8318 | 104.8556 | 119.702 | 124.2212 | 77.0768 | 241.5722 | 224.9323 | 312.1298 | 310.9763 | 176.053 | 229.3273 | 266.0147 | 275.3126 | 331.3747 | 267.3894 | 335.7641 | 346.7111
2372 | 28.0864 | 189.1528 | 284.8418 | 264.831 | 172.126 | 163.874 | 192.2949 | 207.3003 | 154.3003 | 140.2734 | 137.9677 | 155.0269 | 142.5684 | 219.7104 | 166.2949 | 166.2627 | 216.0161 | 179.7212 | 143.1528 | 177.0108 | 172.4263 | 159.3003 | 381.1582 | 469.8687 | 153.1421 | 232.7104 | 248.3003 | 300 | 268.5522 | 108.2788 | 138.2788 | 148.2573 | 116.8525 | 151.2788 | 153.7104 | 208.0323 | 94.0108 | 287.5791 | 298.2896 | 424.2949 | 304.8794 | 142.6997 | 143.8472 | 179.4316 | 126.7212 | 77.6997 | 83.1421 | 82.4209 | 101.5576 | 226.0269 | 205.2734 | 289.7104 | 159.268 | 146.2788 | 148.5684 | 200.8848 | 105.7051 | 185.126 | 205.0108 | 240.9839 | 166.2896 | 123.1206 | 135.1367 | 199.2949 | 122.4209 | 194.5791 | 217.0323 | 297.7212 | 221 | 142.1313 | 163.6997 | 287.4101 | 219.732 | 181.2573 | 225.8525 | 272.5468 | 295.9839 | 186.2842 | 180.9839 | 236.0161 | 236.874 | 146.1421 | 151.5899 | 175.4478 | 144.2949 | 172.2842 | 189.4316 | 211.4263 | 128.4316 | 107.4101 | 139.4316 | 155.5791 | 119.9946 | 296.4316 | 327.437 | 427.874 | 355.3832 | 276.268 | 321.8687 | 447.169 | 358.3272 | 135.1313 | 92.563 | 83.2896 | 135.1313 | 133.1475 | 113.4263 | 141.5737 | 126.0215 | 105.5791 | 115.4316 | 130.6943 | 139.1528 | 93.1582 | 89.4316 | 96.874 | 89.4263 | 122.1528 | 110.0054 | 118.7158 | 89.5791 | 233.1367 | 252.6006 | 307.6006 | 286.1313 | 162.4316 | 200.4155 | 243.8687 | 262.8633 | 356.8418 | 272.5576 | 358.3003 | 379.0323
2373 | 28.0974 | 179.5438 | 295.9792 | 289.9861 | 189.9193 | 147.7788 | 179.9931 | 192.6982 | 141.0876 | 149.8317 | 168.4423 | 127.1613 | 146.6175 | 224.1544 | 159.5507 | 194.6245 | 204.3296 | 174.4701 | 135.242 | 164.5369 | 174.2281 | 145.9332 | 367.8595 | 461.6314 | 156.7788 | 234.0668 | 236.477 | 302.4701 | 289.8456 | 116 | 144.3018 | 176.6036 | 119.5299 | 157.9193 | 157.2281 | 170.9193 | 85.5507 | 278.3825 | 294.3088 | 412.9193 | 279.6912 | 151.7441 | 151.1405 | 173.0668 | 121.3157 | 94.1544 | 82.6106 | 88.0668 | 114.8387 | 200.9402 | 218.8456 | 293.8456 | 175.1405 | 150.1405 | 150.6106 | 167.8387 | 116.9263 | 197.0529 | 194.6982 | 259.7719 | 163.3894 | 142.4423 | 141.8456 | 191.1613 | 130.5369 | 186.1544 | 183.1613 | 292.1613 | 221.4631 | 152.8387 | 178.9193 | 308.3088 | 203.242 | 210.0668 | 233.3157 | 296.4493 | 314.4631 | 186.1474 | 195.9124 | 225.2559 | 219.3894 | 148.0807 | 133.6245 | 158.1683 | 137.242 | 178.9402 | 186.4631 | 217.8595 | 125.6175 | 131.3964 | 137.5438 | 149.0069 | 126.6175 | 293.1544 | 323.0945 | 408.6912 | 398.1266 | 295.0737 | 308.8456 | 422.477 | 313.5438 | 144.2949 | 99.2212 | 81.0069 | 144.2949 | 129.8526 | 115.3825 | 138.6912 | 106.0139 | 95.4562 | 114.6245 | 150.0668 | 126.1474 | 77.0807 | 83.8387 | 79.6982 | 92.6175 | 111.4631 | 107.7051 | 117.7719 | 80.3825 | 240.7719 | 220.1544 | 281.9471 | 293.2879 | 159.1544 | 211.1405 | 233.1613 | 254.1474 | 373.6912 | 284.758 | 339.6843 | 339.6036
2374 | 28.1084 | 186.8333 | 264.3334 | 273.0001 | 180.6666 | 159.9999 | 166.4999 | 208.3335 | 161.3334 | 128.1667 | 131.9998 | 139.6667 | 150.6667 | 217.8331 | 178 | 200.8331 | 234.6666 | 185.4999 | 155.8334 | 161.6667 | 171.5 | 161.8333 | 384.9999 | 486.1666 | 168.8333 | 224.5002 | 257.9999 | 318.8334 | 290.5 | 115.6667 | 129.3332 | 162.1668 | 112.5002 | 148.1666 | 154 | 173.0002 | 101.8333 | 278.1667 | 293.6666 | 403.8333 | 275.4999 | 121.0002 | 134.1668 | 162.8335 | 133.6667 | 94.8333 | 77.8334 | 72.6667 | 104.1667 | 222.3332 | 218.5 | 287.9999 | 154.3333 | 128.8333 | 145.3334 | 163.8335 | 120.3333 | 160.5 | 202.5 | 259.3332 | 168.4999 | 119.5002 | 139.3333 | 202.1666 | 128 | 186.5 | 191.9999 | 301.4999 | 223.8333 | 144.1667 | 171.8333 | 312.5001 | 214.9998 | 198.0001 | 241.1665 | 279.8334 | 319.3334 | 177 | 176.8333 | 258.3332 | 224.8333 | 151.1666 | 149.5 | 183.6667 | 155.1667 | 205 | 185.9999 | 235.9999 | 126.3333 | 142.9998 | 143.3333 | 161.3334 | 129.8333 | 288.8332 | 345.9999 | 402.8332 | 352.8333 | 288.6666 | 307.1666 | 445.8333 | 327.0001 | 121.9999 | 83.6667 | 89.4999 | 121.9999 | 136.1666 | 111.6667 | 136.3333 | 124.8333 | 92.3335 | 127.1666 | 131.8333 | 119.3334 | 83.1667 | 75.0001 | 89.3333 | 96 | 111.9999 | 123.1665 | 119.4999 | 80.0001 | 243.5 | 218.3333 | 320.3333 | 266.1668 | 162.3334 | 195.5001 | 241.1666 | 248.1668 | 373 | 268 | 335.6669 | 326.0001
2375 | 28.1194 | 180.6421 | 271.5369 | 287.7992 | 165.0738 | 147.8634 | 158.1161 | 229.7992 | 172.0943 | 137.0424 | 106.4741 | 143.3579 | 154.3579 | 188.2637 | 171.7473 | 163.4003 | 213.3579 | 174.9686 | 170.5369 | 159.4631 | 176.3156 | 162.9686 | 365.5369 | 477.5259 | 153.8948 | 250.0205 | 237.4522 | 330.3156 | 294.7787 | 115.4317 | 104.0424 | 181.2104 | 143.7473 | 129.3579 | 158.6421 | 204.4098 | 87.4317 | 290.5997 | 285.5792 | 389.6107 | 262.821 | 158.8415 | 150.347 | 190.2104 | 133.2527 | 95.2527 | 88.2418 | 77.2527 | 107.8525 | 197.7582 | 219.7473 | 270.6954 | 156.7158 | 131.4003 | 157.5369 | 199.2418 | 112.179 | 172.832 | 190.168 | 239.5792 | 154.0314 | 155.5683 | 131.179 | 196.5683 | 132.832 | 185.2527 | 179.3265 | 286.8525 | 221.5683 | 162.5055 | 164.5683 | 321.5997 | 185.3579 | 219.0738 | 218.2527 | 285.0519 | 326.168 | 177.8839 | 171.1899 | 226.9686 | 210.9576 | 137.8948 | 156.6421 | 190.7363 | 150.2732 | 192.4836 | 174.0424 | 217.0628 | 116.9372 | 110.8634 | 133.2213 | 166.9262 | 129.5369 | 264.4317 | 318.4945 | 380.7473 | 346.832 | 286.2951 | 301.7473 | 437.4945 | 353.4317 | 114.5055 | 81.821 | 79.5055 | 114.5055 | 116.821 | 111.7896 | 132.1052 | 111.1475 | 120.0738 | 121.4003 | 123.3265 | 128.6626 | 83.1052 | 89.4631 | 90.1052 | 96.8948 | 100.7473 | 98.1052 | 107.7158 | 92.1366 | 240.4741 | 210.7268 | 319.1366 | 285.2104 | 173.4631 | 220.4945 | 225.5792 | 263.6311 | 377.1052 | 270.4945 | 367.3675 | 353.8634
2376 | 28.1304 | 177.0875 | 265.0111 | 251.1364 | 169.6175 | 157.5574 | 180.06 | 189.3114 | 142.1312 | 152.3388 | 126.8087 | 141.7487 | 156.3825 | 214.5137 | 169.2513 | 182.4426 | 214.235 | 180.7487 | 167.0712 | 156.235 | 167.6338 | 170.2787 | 363.4099 | 457 | 156.8525 | 217.6175 | 211.4262 | 321.4426 | 279.6175 | 123.1475 | 118.1913 | 167.6449 | 139.53 | 135.2076 | 155.2787 | 171.3114 | 93.8087 | 276.1475 | 291.2898 | 394.8962 | 262.5738 | 129.8251 | 133.7813 | 179.705 | 135.7487 | 98.5137 | 72.3388 | 80.3225 | 101.06 | 209.8362 | 213.8087 | 301.4099 | 161.9125 | 145.8525 | 158.8525 | 179.8962 | 114.7213 | 190.9837 | 172.7813 | 254.0875 | 145.5738 | 147.8525 | 136.3988 | 191.7213 | 152.1312 | 186.9837 | 193.4426 | 284.6501 | 214.235 | 175.1475 | 160.8688 | 305.4262 | 185.6612 | 222.8525 | 225.9125 | 255.4752 | 308.5901 | 164.9288 | 192.705 | 235.8087 | 198.3388 | 143.3388 | 154.8087 | 166.2076 | 124.3825 | 163.7376 | 178.0548 | 196.3225 | 133.0875 | 122.8525 | 145.1312 | 161.235 | 132.765 | 279.9889 | 307.9563 | 382.8414 | 367.47 | 300.2624 | 297.9125 | 425.8087 | 354.2624 | 129.06 | 80.4262 | 91.765 | 129.06 | 117.7213 | 123.06 | 128.9563 | 122.1475 | 122.1312 | 133.5574 | 131.3225 | 98.8251 | 77.6612 | 86.235 | 85.8087 | 105.2076 | 97.4863 | 96.4863 | 108.1312 | 79.5738 | 260.6175 | 241.2076 | 304.8525 | 276.8087 | 172.1475 | 205.94 | 239.3225 | 243.0163 | 369.3662 | 254.7924 | 317.4426 | 369.4863
2377 | 28.1414 | 169.8142 | 223.7433 | 283.0086 | 171.4603 | 143.6106 | 189.5397 | 214.7787 | 132.8496 | 155.5575 | 125.7965 | 127.1858 | 153.1149 | 205.4248 | 174.2476 | 191.5043 | 210.1858 | 171.2744 | 148.1063 | 157.6815 | 188.2921 | 163.8142 | 347.4603 | 448.8582 | 154.0354 | 224.1418 | 218.9737 | 329.8937 | 272.0972 | 127.7965 | 122.2567 | 148.6638 | 145.3894 | 150.9646 | 142.7079 | 196.3716 | 93.2035 | 276.1149 | 257.8056 | 388.4425 | 267.8496 | 142.7433 | 148.7079 | 172.0354 | 124.4425 | 86.3716 | 80.0354 | 72.2567 | 112.9823 | 187.5929 | 209.1327 | 277.7256 | 164.9114 | 145.478 | 150.3362 | 169.7787 | 124.0354 | 175.0177 | 197.0709 | 241.522 | 153.9028 | 146.2567 | 152.761 | 197.7787 | 140.8142 | 172.8496 | 206.3716 | 310.0795 | 217.717 | 175.522 | 170.3539 | 301.3716 | 181.8319 | 207.8228 | 235.6284 | 296.0618 | 322.5043 | 186.3362 | 182.1858 | 228.6901 | 215.3985 | 147.5752 | 152.3716 | 177.283 | 127.8319 | 185.1504 | 146.9914 | 190.4957 | 127.239 | 117.1858 | 128.1149 | 159.8319 | 132.9468 | 311.283 | 311.7965 | 412.3362 | 365.478 | 275.3894 | 303.5575 | 430.2921 | 325.9291 | 143.6284 | 79.2213 | 93.9823 | 143.6284 | 125.4071 | 132.5397 | 131.7787 | 124.3539 | 110.8142 | 114.7256 | 123.8673 | 113.5752 | 73.0177 | 78.3185 | 87.0354 | 116.283 | 107.1858 | 106.3894 | 96.8142 | 83.4248 | 260.6284 | 255.3362 | 306.3099 | 276.2035 | 174.3539 | 193 | 228.6106 | 252.2653 | 363.1595 | 244.7347 | 325.4866 | 382.8496
2378 | 28.1524 | 174.0792 | 214.2375 | 237.41 | 184.1367 | 144.9208 | 181.18 | 209.7842 | 143.7842 | 147.5108 | 128.2375 | 125.9425 | 139.3959 | 210.5683 | 144.7484 | 172.3742 | 206.7842 | 181.5249 | 165.3091 | 174.4892 | 201.6616 | 167.8633 | 361.6475 | 422.3959 | 163.9425 | 248.4458 | 251.0933 | 308.1584 | 259.1443 | 123.9783 | 135.2158 | 160.1941 | 141.0575 | 147.1009 | 128.9067 | 194.1009 | 97.4534 | 261.1009 | 294.5033 | 395.41 | 272.0933 | 128.4892 | 132.9642 | 181.0792 | 130.3308 | 78.4892 | 88.8633 | 83.2733 | 111.1584 | 195.4967 | 194.8633 | 261.4317 | 146.7267 | 154.3666 | 136.6475 | 167.3742 | 130.2733 | 183.7484 | 208.2516 | 231.7701 | 177.2158 | 157.2733 | 145.0792 | 189.3308 | 143.1367 | 187.0217 | 200.6475 | 286.6692 | 232.9924 | 150.8774 | 163.0217 | 296.5108 | 190.0792 | 180.0868 | 235.5249 | 268.2375 | 302.295 | 175.2375 | 177.0575 | 211.5542 | 249.8557 | 141.1367 | 152.4751 | 152.0792 | 136.5108 | 172.9208 | 188.0575 | 206.1725 | 138.8059 | 116.3742 | 115.0434 | 161.1725 | 125.3167 | 282.6258 | 314.4534 | 394.7625 | 373.5033 | 282.2809 | 294 | 442.3666 | 322.5184 | 148.0575 | 81.41 | 85.8991 | 148.0575 | 124.8416 | 117.9208 | 126.3525 | 111.41 | 111.1941 | 105.9859 | 139.8633 | 108 | 76.5683 | 66.0434 | 97.1584 | 89.3525 | 106.59 | 105.0792 | 102.8059 | 86.8416 | 260.5249 | 241.4317 | 317.1865 | 275.0575 | 161.41 | 193 | 240.2809 | 229.3384 | 389.9349 | 275.6475 | 309.0651 | 390.5466
2379 | 28.1634 | 175.9469 | 234.2125 | 239.4063 | 179.6312 | 142.1406 | 188.3656 | 209.4563 | 142.5437 | 150.0344 | 137.7187 | 138.9313 | 158.0875 | 203.7531 | 163.5969 | 178.7187 | 209.4219 | 164.1594 | 148.3 | 169.2281 | 172.0344 | 168.0344 | 360.35 | 439.9469 | 170.3156 | 220.5437 | 234.4407 | 314.4031 | 300.7907 | 115.0531 | 134.4031 | 149.5437 | 136.0531 | 157.8937 | 147.3843 | 197.3656 | 104.2968 | 273.9469 | 267.0719 | 389.1938 | 250.65 | 126.4219 | 149.6312 | 181.35 | 115.8594 | 79.3875 | 88.35 | 78.5969 | 113.9812 | 222.4718 | 198 | 265.2813 | 156.4219 | 125.5094 | 141.9656 | 164.8218 | 124.2281 | 193.9281 | 186.8937 | 256.1219 | 174.8062 | 153.2813 | 145.35 | 184.4407 | 136.8062 | 187.2781 | 202.5437 | 292.4718 | 195.9656 | 137.0907 | 170.35 | 307.2469 | 191.0344 | 221.0344 | 208.1219 | 266.9968 | 303.5781 | 187.2281 | 175.2469 | 236.2625 | 211.175 | 140.7375 | 169.1562 | 158.0531 | 141.5437 | 168.7719 | 183.2813 | 181.8062 | 143.9812 | 125.4563 | 135.8594 | 143.1938 | 135.2968 | 281.5282 | 330.1938 | 393.2657 | 337.6312 | 307.5593 | 303.5813 | 414.8781 | 363.2093 | 142.5969 | 71.3156 | 75.1063 | 142.5969 | 115.8594 | 111.7187 | 125.825 | 104.0531 | 106.475 | 127.4718 | 144.825 | 107.5437 | 75.9125 | 85.7187 | 101.1219 | 89.2813 | 112.8062 | 110.5969 | 109.5781 | 78.5437 | 242.9313 | 240.7187 | 271.7907 | 268.2281 | 158.8438 | 195.9656 | 271.9469 | 260.7907 | 358.5969 | 275.4906 | 330.5937 | 382.1063
2380 | 28.1744 | 168.0383 | 247.575 | 263.8154 | 175 | 143.5958 | 146.3554 | 205.7108 | 141.7213 | 146.9721 | 127.5958 | 151.7875 | 152.8362 | 192.9721 | 167.5575 | 170.2787 | 209.9408 | 184.7596 | 166.2683 | 166.3937 | 162.2404 | 157.2787 | 351.1254 | 435.1638 | 164.6342 | 222.8467 | 256.0279 | 308.5192 | 312.1917 | 122.7213 | 126.3554 | 144.6342 | 140.8362 | 145.4425 | 132.2125 | 162.3275 | 82.5958 | 261.23 | 287.1533 | 399.7596 | 259.1533 | 132.9512 | 145.2404 | 170.2021 | 116.0871 | 98.5958 | 77.2021 | 85.6829 | 96.2787 | 204.9338 | 204.0104 | 277.5679 | 165.115 | 133.5192 | 147.6725 | 138.54 | 126.9233 | 175.77 | 183.8188 | 247.5679 | 162.0767 | 151.6237 | 137.5679 | 207.9512 | 131.77 | 158.9338 | 200.7596 | 273.4913 | 209.1254 | 174.6342 | 157.5192 | 320.6725 | 176.9129 | 212.6829 | 193.0767 | 234.0592 | 300.4146 | 186.7979 | 192.2787 | 223.8746 | 207.885 | 151.4425 | 142.8746 | 166.2021 | 139.7596 | 164.8746 | 185.4704 | 170.7596 | 126.5192 | 128.6829 | 136.5679 | 155.6829 | 116 | 306.7596 | 330.9025 | 433.5296 | 340.2125 | 291.2613 | 324.7004 | 416.46 | 337.3833 | 139.1046 | 69.7213 | 92.6063 | 139.1046 | 113.9233 | 105.6829 | 131.5192 | 111 | 122.5958 | 107.77 | 141.3833 | 106.9617 | 73.9512 | 78.9617 | 87.5192 | 93.1533 | 102.2404 | 116.7213 | 96.3171 | 76.7596 | 259.3937 | 232.0383 | 287.7596 | 271.6446 | 180.0767 | 206.2404 | 256.5854 | 265.9408 | 366.4042 | 276.6167 | 309.5505 | 393.115
2381 | 28.1854 | 165.5054 | 209.834 | 224.8123 | 170.4512 | 144.3394 | 154.426 | 193.5488 | 143.4946 | 159.1444 | 121.769 | 130.0108 | 146.2527 | 211.7148 | 169.8231 | 173.8123 | 189.1517 | 177.935 | 143.9892 | 154.7473 | 166.5379 | 153.7365 | 359.7365 | 435.935 | 148.4404 | 229.6931 | 235.5163 | 305.9892 | 292.704 | 126.2635 | 124.3177 | 140.3177 | 126.9242 | 151.2635 | 150.1985 | 174.231 | 80.0542 | 247.296 | 274.4946 | 401.2744 | 245.231 | 134.4369 | 141.4512 | 175.2635 | 125.2094 | 92.2635 | 80.4946 | 94.0325 | 94.2527 | 218.3754 | 217.4404 | 286.1877 | 156.7581 | 137.0542 | 140.3177 | 171.4621 | 126.935 | 187.4296 | 204.101 | 252.9025 | 157.2635 | 127.5596 | 156.8015 | 186.6931 | 146.9675 | 177.1769 | 197.7256 | 287.4296 | 213.9458 | 160.9675 | 163.3285 | 309.2744 | 168.296 | 213.704 | 192.3069 | 271.2744 | 311.8375 | 180.9783 | 190 | 215.769 | 214.4729 | 155.4946 | 141.3394 | 165.9567 | 138.2419 | 155.7581 | 175.834 | 180.1083 | 125.7581 | 133.2419 | 141.6498 | 159.231 | 118.5271 | 301.9567 | 302.7906 | 434.8267 | 361.9892 | 317.3646 | 283.1877 | 389.0975 | 314.834 | 115.296 | 75.7906 | 95.6606 | 115.296 | 118.7365 | 108.9783 | 124.1877 | 109.9892 | 116.769 | 122.2094 | 120.8556 | 113.0325 | 64.065 | 81.7473 | 83.7256 | 87.5704 | 102.7473 | 111.4187 | 100.3502 | 78.7798 | 238.3971 | 225.4621 | 290.5379 | 277.2202 | 173.2419 | 207 | 228.7365 | 240.8556 | 368.9458 | 302.9567 | 342.0035 | 377.6823
2382 | 28.1964 | 173.36 | 233.53 | 249.42 | 167.335 | 162.11 | 193.195 | 203.025 | 143.06 | 131.325 | 127 | 131.145 | 147.795 | 203.855 | 179.665 | 190.735 | 222.785 | 165.035 | 140.735 | 155.855 | 174.35 | 150.795 | 359.975 | 419.06 | 128.025 | 228.84 | 241.06 | 300.085 | 280 | 131.06 | 143 | 152.11 | 113.95 | 150.23 | 132.025 | 171.12 | 95.795 | 255.495 | 276.975 | 397.93 | 242.385 | 175.135 | 131.71 | 178.205 | 117.445 | 93.615 | 80.855 | 97.965 | 96.325 | 188.24 | 203.915 | 269.53 | 158.735 | 145.375 | 151.845 | 163.12 | 110.06 | 173.3 | 167.505 | 228.12 | 153.315 | 141.085 | 171.265 | 178.155 | 137.47 | 168.395 | 199.745 | 273.035 | 211.19 | 158.625 | 181.025 | 310.7 | 187.89 | 199.94 | 206.145 | 278.795 | 278.13 | 168.64 | 187.085 | 210.4 | 209.915 | 148.7 | 158.315 | 152.94 | 136.265 | 158.265 | 190.255 | 205.23 | 121.64 | 126.495 | 116.65 | 151.35 | 127.855 | 295.83 | 317.18 | 392.83 | 360.325 | 286.095 | 269.71 | 416 | 330.58 | 126.94 | 83.29 | 74.445 | 126.94 | 112.88 | 104.59 | 113.625 | 106.735 | 117.23 | 110.735 | 147.88 | 118.025 | 81.735 | 79.41 | 78.06 | 108.59 | 99.88 | 91.385 | 121.64 | 83.555 | 213.71 | 214.47 | 298.88 | 276.42 | 178.42 | 204.35 | 228.975 | 260.725 | 357.71 | 285.435 | 295.73 | 359.265
2383 | 28.2074 | 182.404 | 231.9498 | 274.7139 | 197.1092 | 145.2183 | 163.2271 | 192.2773 | 148.2183 | 139.1594 | 126.7227 | 129.6048 | 147.2271 | 208.1681 | 149.059 | 183.8996 | 200 | 173.7315 | 140 | 164.6048 | 174.4868 | 151.3363 | 359.0765 | 433.9235 | 131.1418 | 261.404 | 239.5634 | 295.6048 | 278.6135 | 129.5634 | 133.5721 | 134.6637 | 133.2271 | 137.8319 | 125.7139 | 180.0502 | 95.2271 | 243.5546 | 279.9586 | 373.8821 | 253.7139 | 144.1681 | 142 | 169.6223 | 128.3865 | 83.9498 | 87.3865 | 84 | 104.4366 | 196.9498 | 218.3777 | 277.3777 | 157.4454 | 133.655 | 132.8319 | 167.3363 | 110.5044 | 180.7904 | 181.941 | 247.0327 | 136.9912 | 129.1179 | 167.0088 | 193.8406 | 135.7227 | 196.6725 | 217.6223 | 279.5132 | 230.5809 | 169.5132 | 168.6135 | 292.6725 | 200.345 | 200.3275 | 207.3777 | 271.0175 | 313.1594 | 156.2685 | 179 | 193.4781 | 215.2271 | 132.6135 | 137 | 150.8319 | 133.1179 | 163.9912 | 167.3865 | 196.7139 | 110.3777 | 117.8821 | 119.8406 | 150.3777 | 129.9498 | 316.7139 | 320.7315 | 405.9498 | 365.1092 | 287.1944 | 287.2096 | 421.8231 | 316.655 | 123.4454 | 72.7227 | 87.0502 | 123.4454 | 108.3865 | 107.8908 | 136.7139 | 111.5458 | 107.6048 | 110.2773 | 138.1179 | 108.3865 | 79.8908 | 73.6135 | 81 | 104.6813 | 89.5634 | 94.1179 | 104 | 72.059 | 228.9912 | 211.6135 | 296.0502 | 297.8319 | 187.0765 | 199.7729 | 236.6725 | 246.3689 | 360.5132 | 270.8319 | 349.0677 | 358.6135
2384 | 28.2184 | 164.6333 | 224.2896 | 293.3708 | 191.3126 | 148.3938 | 164.3979 | 190.1042 | 146.4709 | 153.475 | 118.4709 | 130.023 | 141.1583 | 215.5562 | 152.3979 | 168.2896 | 209.5562 | 164.9229 | 138.5521 | 166.4709 | 188.7875 | 151.9229 | 336.6874 | 451.6793 | 165.1042 | 240.7375 | 224.2354 | 300.3667 | 277.6062 | 117.7104 | 117.6874 | 141.3167 | 128.8958 | 146.3708 | 140.6062 | 180.7605 | 89.7375 | 254.2666 | 261.6062 | 383.7104 | 275.5562 | 139.9729 | 146.9229 | 151.5521 | 128.8146 | 76.2896 | 94.1854 | 88.3437 | 112.8146 | 191.0271 | 222.8376 | 280.6793 | 156.5792 | 144.3396 | 141.95 | 160.6833 | 108.0541 | 171.2666 | 187.8687 | 267.6793 | 149.7104 | 126.2395 | 149.3667 | 189.3708 | 139.0541 | 186.8417 | 197.2354 | 262.8958 | 194.4479 | 146.8146 | 171.95 | 278.498 | 190.2395 | 203.7875 | 219.9459 | 259.7687 | 320.2354 | 167.9729 | 178.7104 | 228.6062 | 207.1313 | 135.0812 | 149.1624 | 160.8187 | 130.5291 | 185.1083 | 171.2896 | 205.8146 | 131.9229 | 128.8687 | 109.5792 | 165.2625 | 115.05 | 305.8335 | 318.0041 | 404.3708 | 361.9188 | 259.8999 | 301.6563 | 432.0771 | 331.1042 | 123.1583 | 75.475 | 97.6062 | 123.1583 | 111.4208 | 111.3708 | 141.7605 | 129.1854 | 120.475 | 116.7916 | 128.8687 | 114.0271 | 74.3938 | 70.8687 | 85.9229 | 88.2125 | 84.6604 | 91.2395 | 101.6833 | 72.2125 | 237.6563 | 211.475 | 292.9959 | 291.8917 | 153.9729 | 206.7104 | 230.8958 | 275.7104 | 351.1354 | 269.2354 | 324.9188 | 352.9729
2385 | 28.2294 | 180.2262 | 224.0944 | 302.0569 | 167.2831 | 140.7925 | 177.4532 | 188.7356 | 136.7543 | 158.6794 | 108.1506 | 112.83 | 143.3963 | 217.567 | 170.2831 | 165.9813 | 229.6794 | 180.0187 | 137.415 | 153.7356 | 173.8869 | 165.2075 | 351.0569 | 420.2075 | 154.0757 | 246.2075 | 218.0187 | 293.2262 | 293.9618 | 122.7356 | 136.8869 | 149.7169 | 137.5094 | 157.17 | 149.1131 | 168.4337 | 92.7925 | 272.4719 | 268 | 371.5288 | 286.3214 | 134.0944 | 154.7738 | 160.0749 | 121.3019 | 75.1888 | 103.5094 | 89.9438 | 107.1131 | 192.6794 | 201.1506 | 254.3393 | 152.2644 | 124.3775 | 149.642 | 163.4524 | 115.2831 | 191.2831 | 184.8682 | 236.8112 | 155.0375 | 143.6981 | 141.9243 | 199.2644 | 145.9813 | 196.0749 | 191.3206 | 274.2262 | 201.9243 | 142.3206 | 177.227 | 264.3963 | 209.8112 | 199.1506 | 216.6412 | 277.5858 | 314.9243 | 166.0187 | 176.1888 | 222.3214 | 211.9431 | 150.3019 | 172.6607 | 179.0944 | 143.2644 | 194.4345 | 173.5094 | 197.3963 | 136.1513 | 121.0382 | 114.9243 | 159.7925 | 107.0562 | 260.1693 | 336.2457 | 407.3214 | 347.9056 | 295.8112 | 291.3019 | 419.3963 | 323.6981 | 116.3401 | 82.7925 | 94.0382 | 116.3401 | 108.4906 | 120.3588 | 126.1131 | 131.2644 | 125.0757 | 128.585 | 129.4906 | 116.585 | 72.2262 | 79.3393 | 94.0757 | 100.3963 | 102.5094 | 102.9625 | 97.8112 | 83.7925 | 232.1318 | 220.6037 | 274.1506 | 282.2637 | 162.2824 | 202.9813 | 244.0375 | 259.0008 | 386.2637 | 250.3401 | 313.9243 | 341.6607
2386 | 28.2404 | 189.0575 | 223.2567 | 280.9734 | 164.1416 | 136.4292 | 159.0575 | 201.6858 | 156.6858 | 151 | 120.0884 | 129.9691 | 139.8009 | 187.7923 | 163.3717 | 159.6283 | 212.2609 | 177.5751 | 147.7125 | 165.115 | 176.8541 | 168 | 340.0266 | 435.5665 | 144.0575 | 240.8317 | 232.5399 | 300.1725 | 309.4601 | 132.23 | 125.7167 | 154.4292 | 140.0575 | 141.6017 | 138.9201 | 176.2875 | 96.9116 | 267.9691 | 272.3983 | 344.3717 | 254.8009 | 133.885 | 151.5974 | 190.1991 | 129.2258 | 78.8541 | 103.5442 | 76.5399 | 115.4558 | 187.8275 | 220 | 275.5974 | 140.2567 | 124.1683 | 132.3983 | 176.0884 | 118.4249 | 188.4558 | 178.0266 | 235.0309 | 168.6858 | 149.9116 | 139.6901 | 188.1991 | 137.7433 | 201.0661 | 197.7433 | 268.9201 | 216.9691 | 156.5974 | 151.885 | 267.7125 | 205.8317 | 217.3717 | 228.23 | 250.394 | 312.6901 | 173.3142 | 174.5133 | 195.1991 | 219.2034 | 146.9159 | 150.1459 | 171.9734 | 136.9116 | 174.0841 | 177.6283 | 189.7167 | 125.2258 | 100.8275 | 125.5708 | 150.0884 | 130.1991 | 291.3408 | 316.6283 | 380.5133 | 350.6283 | 287.1193 | 284.4601 | 409.8317 | 336.8232 | 104.9957 | 80 | 81.0532 | 104.9957 | 106.2567 | 105.3142 | 130.6858 | 109.5176 | 111.9159 | 118.9159 | 128.1991 | 118.5399 | 80.7433 | 89.9159 | 85.6283 | 98.6858 | 107.5708 | 91.8275 | 104.8275 | 80.6858 | 233.9468 | 233.6241 | 292.3717 | 293.7476 | 189.5442 | 195.6858 | 247.9468 | 229.2258 | 390.836 | 241.509 | 316.0884 | 337.9957
2387 | 28.2514 | 182.7561 | 226.6529 | 269.3471 | 174.2645 | 147.364 | 172.0169 | 200.6735 | 154.3677 | 145.1238 | 113.8143 | 147.2439 | 138.2645 | 224.1445 | 158.0825 | 171.4465 | 190.0206 | 161.2439 | 154.7355 | 158.0619 | 196.2852 | 163.1032 | 355.2439 | 449.9418 | 144.2852 | 234.7524 | 242.4709 | 295.9587 | 293.6529 | 125.9587 | 117.2233 | 149.0413 | 134.409 | 134.9381 | 111.2233 | 164.6942 | 104.818 | 270.5535 | 285.2645 | 353.1201 | 250 | 143.8762 | 157.5122 | 177.3714 | 140.1032 | 95.6735 | 93.3884 | 86.1445 | 120.7148 | 208.6905 | 220 | 290 | 148.2233 | 138.7355 | 142.3265 | 158.0619 | 135.409 | 185.227 | 192.9175 | 235.7111 | 170.2852 | 154.8799 | 122.5085 | 190.7148 | 133.3677 | 158.5085 | 193.0413 | 251.0169 | 227.7148 | 163.1651 | 156.3265 | 285.182 | 200.0788 | 200.6566 | 217.0619 | 278.7767 | 297.7936 | 169.1032 | 181.6323 | 202.6116 | 204.2439 | 144.2026 | 140.182 | 166.8762 | 152.6529 | 172.8799 | 178.3471 | 180.2439 | 131.2064 | 113.2026 | 129.6529 | 142.5085 | 130.7561 | 296.4709 | 316.0413 | 395.4672 | 346.7767 | 275.3227 | 268 | 401.7936 | 362.1032 | 125.7355 | 82.2852 | 103 | 125.7355 | 107.6942 | 103.7148 | 128.6942 | 94.591 | 103 | 113.591 | 122.5535 | 118.6773 | 76.6942 | 84.2645 | 88.3058 | 95.3884 | 116.8762 | 102.5704 | 106.4296 | 81.9587 | 215.591 | 245.9418 | 289.0413 | 266.0206 | 180.0413 | 205.4465 | 226.9794 | 250.8762 | 351.9381 | 266.1032 | 296.4296 | 350.5741
2388 | 28.2624 | 171.6288 | 220.8863 | 269.0162 | 178.6288 | 167.1764 | 184.7077 | 197.6288 | 146.2576 | 134.6937 | 124.5476 | 148.2088 | 139.5801 | 212.71 | 150.6775 | 192.3225 | 193.4199 | 168.3063 | 147.3225 | 156.0649 | 186.0952 | 148.5963 | 352.4826 | 411.8214 | 146.9675 | 244.4339 | 227.645 | 298.645 | 283.4826 | 121.1926 | 121.9026 | 143.645 | 123.6125 | 138.6288 | 128.7749 | 165.0487 | 88.645 | 244.5963 | 279.8052 | 356.724 | 264.5661 | 154.3063 | 143.0487 | 151.1624 | 124.2413 | 91.9513 | 90.3712 | 77.0812 | 107.5314 | 227.4826 | 221.6937 | 267.9815 | 160.6937 | 136.4037 | 141.645 | 169.276 | 131.3875 | 168.4199 | 192.8701 | 259.5963 | 167.2088 | 130.9513 | 140.71 | 188.0325 | 133.7262 | 175.0162 | 201.1948 | 269.4664 | 228.4199 | 149.3712 | 162.4199 | 295.9351 | 214.1601 | 176.4524 | 214.3875 | 259.1926 | 315.9675 | 155.2738 | 178.2251 | 205.9675 | 216.3875 | 156.6613 | 150.9351 | 165.1114 | 146.8863 | 160.4686 | 181.7424 | 198.1461 | 117.4362 | 123.6288 | 137.7749 | 156.645 | 124.3712 | 286.3875 | 306.5801 | 409.2413 | 338.0325 | 294.4339 | 275.4524 | 409.4664 | 341.4989 | 120.0162 | 91.0649 | 94.5314 | 120.0162 | 103.645 | 102.3875 | 119.5638 | 114.5336 | 105.0325 | 115.9188 | 108.1137 | 91.4362 | 74.6775 | 86.5963 | 83.8863 | 87.9675 | 118.1601 | 112.3387 | 100.0487 | 83.29 | 224.0162 | 221.71 | 268.4015 | 269.4199 | 174.645 | 210.7402 | 220.8701 | 268.4199 | 361.0487 | 257.355 | 288.0162 | 334.7912
2389 | 28.2734 | 164.1917 | 214.7228 | 267.8407 | 178.9145 | 146.5959 | 156.7228 | 193 | 141.5635 | 145.7228 | 115.9676 | 127.0324 | 136.3718 | 211.3303 | 149.1917 | 187.4896 | 204.351 | 167.1062 | 146.7021 | 157.3303 | 165.0738 | 147.0207 | 334.2124 | 398.4249 | 144.0531 | 217.0531 | 223.9469 | 294.2448 | 267.6697 | 105.0531 | 115.2656 | 141.7021 | 119.1062 | 143.127 | 151.127 | 171.7021 | 90.9145 | 255.307 | 268.2863 | 341.2539 | 288.4365 | 155.2124 | 152.5104 | 175.2124 | 114.0531 | 88.4573 | 92.1917 | 89.4573 | 95.2124 | 207.4573 | 225 | 237.2863 | 162.9469 | 138.3303 | 137.2448 | 193.171 | 130.9793 | 173.7344 | 174.8938 | 244.3303 | 149.8938 | 128.1593 | 144.9469 | 173.7461 | 145.2124 | 182.9676 | 210.9585 | 252.3718 | 231.6282 | 164.5311 | 171.9469 | 282.7344 | 187.7344 | 198.3718 | 213.9793 | 242 | 317.2656 | 158.1801 | 161.8407 | 204.1062 | 215.6282 | 159.1593 | 135.6282 | 140.4573 | 146.3394 | 173.1386 | 181.522 | 212.1386 | 127.1917 | 119.351 | 139.4158 | 158.9145 | 130.4041 | 291.5959 | 303.0207 | 397.649 | 339.1917 | 272.6697 | 275.6075 | 371.6606 | 319.2448 | 118.8407 | 93.0324 | 75.8938 | 118.8407 | 103.8083 | 107.2448 | 114.3718 | 126.7137 | 104.4365 | 103.8938 | 120.9469 | 98.7344 | 77.7552 | 76.9469 | 73.8614 | 85.4041 | 104.3718 | 106.6813 | 109.1593 | 78.351 | 233.0207 | 232.9676 | 259.5195 | 268.7668 | 176.9145 | 184.2656 | 209.9145 | 260.044 | 361.0324 | 260 | 303.6904 | 350.7021
2390 | 28.2844 | 159 | 224.2767 | 259.82 | 180.3701 | 141.4567 | 167.3667 | 198.8132 | 149.6367 | 158.1834 | 121.1868 | 114.1834 | 146.73 | 201.9066 | 147.6333 | 168.6469 | 199.9134 | 165.1868 | 138.9168 | 146.4533 | 178.91 | 157.82 | 352.5365 | 406.3701 | 147.8166 | 216.82 | 220.5467 | 299.3565 | 280.7267 | 106.2733 | 121.0034 | 143.7267 | 118.2767 | 149.4635 | 160.3701 | 179.5399 | 92.0068 | 289.5467 | 289.5467 | 372.0934 | 271.2801 | 160.4567 | 152.0968 | 166.2869 | 120.7233 | 96.27 | 87 | 95.4533 | 104.4533 | 201.8268 | 213.0102 | 253.0968 | 154.4601 | 131.0866 | 126.0068 | 168.4601 | 108.9168 | 162.5467 | 174.9966 | 234.18 | 140.5501 | 133.6367 | 140.82 | 158.893 | 152.2733 | 174.017 | 178.4601 | 257.2801 | 220.18 | 173.4635 | 164.9134 | 282.4465 | 181.9966 | 201.4635 | 207.5399 | 237.6401 | 310.6502 | 173.0034 | 167.9898 | 212.7233 | 207.4499 | 162.4567 | 133.2632 | 149.3599 | 160.1002 | 164.1834 | 159.6401 | 187.5603 | 128.5399 | 114.9134 | 124.1766 | 168 | 133 | 283.1868 | 301.8302 | 390.0968 | 344.1732 | 290.4499 | 252.9966 | 347.3633 | 318.5433 | 127.8964 | 76.1868 | 80.3565 | 127.8964 | 111.5433 | 103.2733 | 123.2767 | 109.8132 | 101.0866 | 104.3599 | 113.1868 | 98.0832 | 82.4533 | 71.73 | 87.9032 | 90.9066 | 111.0968 | 94.2733 | 108.0968 | 74.6401 | 247.09 | 239.2767 | 302.6367 | 253.0866 | 179.4601 | 198.7233 | 221.9066 | 242.4397 | 350.7267 | 254.9134 | 311.1969 | 367.9864
2391 | 28.2954 | 165.0099 | 207.7364 | 259.3806 | 174.1368 | 141.1318 | 159.8731 | 206.7463 | 152.005 | 151.2438 | 115.5074 | 122.2686 | 135.3657 | 203.995 | 158.883 | 154.5124 | 182.7364 | 154.6244 | 137.2735 | 151.6293 | 179.2537 | 155.1269 | 340.1993 | 402.0149 | 154.3806 | 216.7562 | 223.2587 | 312.4975 | 282 | 115.8929 | 115.8781 | 153.2636 | 116.0099 | 140.7661 | 134.9802 | 183.8632 | 77.2488 | 284.7463 | 272.7265 | 353.8533 | 262.0099 | 156.7512 | 139.6244 | 138.8781 | 130.5025 | 92.9851 | 87.7512 | 87.4827 | 108.5025 | 194.8929 | 190.8731 | 262.0347 | 138.8682 | 136.995 | 122.1417 | 158.8781 | 95.6293 | 164.1318 | 186.883 | 229.7364 | 137.0099 | 128.8682 | 138.5025 | 178.8533 | 149.1219 | 151.3954 | 180.5223 | 243.1269 | 222.4975 | 158.005 | 161.6343 | 297.1219 | 189.3756 | 187.883 | 217.8731 | 243.5223 | 288.3855 | 169.3806 | 179.6095 | 210.4827 | 207.2339 | 159.1269 | 156.005 | 154.3707 | 145.1318 | 172.6442 | 158.3855 | 155.1219 | 139.6244 | 109.3806 | 125.9851 | 163.4926 | 120.6046 | 277.5074 | 292.776 | 390.3954 | 355.6145 | 304.1318 | 266.0099 | 367.5322 | 329.3855 | 146.8632 | 72.3855 | 90.4926 | 146.8632 | 125.7661 | 106.5074 | 110.8682 | 111.7364 | 113.3806 | 110.1219 | 109.3855 | 110.2438 | 91.3855 | 66.7512 | 93.9901 | 91.117 | 101.6293 | 102.3905 | 102.7611 | 71.5074 | 243.7413 | 229.8731 | 281.341 | 248.8533 | 171.7562 | 198.7364 | 221.3657 | 257.9752 | 359.888 | 253.888 | 292.3954 | 386.2339
2392 | 28.3064 | 169.1814 | 199.8186 | 263.8363 | 181.733 | 153.8186 | 158 | 196.4055 | 152.733 | 136.466 | 123.7758 | 130.9143 | 132.9824 | 212.1889 | 161.9572 | 163.5088 | 184.5164 | 176.4988 | 148.199 | 149.4055 | 183.7758 | 160.1033 | 303.7406 | 415.0605 | 165.4307 | 236.1889 | 223.466 | 297.5869 | 287.0428 | 123.5264 | 116.6725 | 149.932 | 129.8791 | 165.1461 | 120.4307 | 169.2846 | 80.267 | 280.2242 | 270.723 | 349.4559 | 270.0605 | 153.733 | 136.2846 | 145.1033 | 126.4055 | 82.267 | 86.6725 | 82.0252 | 105.9572 | 202.9143 | 197.534 | 274.1133 | 140.1461 | 133.9395 | 142.1033 | 158.1209 | 105.8186 | 175.6549 | 181.4232 | 217.9395 | 143.8967 | 125.4912 | 137.1209 | 168.2494 | 139.7934 | 184.801 | 204.1637 | 238.4055 | 230.4736 | 172.6977 | 163.2418 | 284.3023 | 186.8967 | 193.6725 | 212.8967 | 262.5088 | 291.2418 | 180.4307 | 165.5516 | 201.1461 | 201.0857 | 150.9143 | 143.932 | 154.6549 | 142.3023 | 192.7758 | 162.0176 | 162.4736 | 135.1209 | 121.2066 | 121.0857 | 156.3879 | 109.3098 | 278.7934 | 306.199 | 398.9748 | 343.8363 | 297.8111 | 269.7934 | 376.7254 | 327.9748 | 129.9572 | 85.7154 | 82.6121 | 129.9572 | 134.6297 | 117.8791 | 106.7154 | 98 | 119 | 111.6549 | 125.0428 | 109.4307 | 93.8539 | 78.0857 | 85.5516 | 87.6549 | 99.0176 | 114.5088 | 112.8363 | 79.7758 | 234.2242 | 221.0176 | 254.3703 | 254.5416 | 163.6473 | 193.1461 | 215.8791 | 237.6549 | 343.5188 | 265.4484 | 317.6549 | 362.2418
2393 | 28.3174 | 157.999 | 215.003 | 249.5938 | 170.1976 | 158.1976 | 165.6038 | 186 | 143.1986 | 140.6088 | 120.5978 | 117.8014 | 145.2006 | 218.5928 | 157.6018 | 157.5998 | 214.2026 | 201.5948 | 134.4032 | 134.1976 | 180.1976 | 162.999 | 322.5978 | 400.7944 | 161.994 | 248.996 | 212.001 | 274.3982 | 296.2006 | 115.8044 | 112.1996 | 135.8064 | 142.003 | 165.5938 | 136.6038 | 165.8004 | 83.7984 | 288.6048 | 282.9889 | 346.5847 | 252.993 | 140.997 | 131.5998 | 159.6048 | 115.5998 | 82.997 | 77.7974 | 94.1966 | 99.6008 | 187 | 209.3992 | 240.8014 | 149.7984 | 132.5978 | 143.7984 | 152.8004 | 117.4012 | 181.3992 | 160.8004 | 223.8014 | 145.003 | 131.4002 | 122.5958 | 186 | 134.8024 | 192.1936 | 200.7974 | 241.6068 | 237.7954 | 169.991 | 158.6038 | 282.6068 | 184.001 | 196.003 | 202.7974 | 244.1936 | 289.005 | 175.7934 | 165.1986 | 215.6008 | 211.3972 | 144.2046 | 124.6038 | 159.5988 | 135.004 | 183.5948 | 154.6018 | 186.2036 | 131.4012 | 126.998 | 135.3992 | 143.7934 | 131.6038 | 284.6078 | 285.5998 | 378.7994 | 348.8034 | 270.5968 | 271.6058 | 352.8064 | 325.4082 | 124.8014 | 79.995 | 82.4002 | 124.8014 | 116.1966 | 131.2036 | 115.4022 | 98 | 117.3992 | 116.999 | 126.997 | 114.7984 | 84.6018 | 90.7984 | 85.5988 | 91.7984 | 99.2056 | 106.5988 | 108.1986 | 80.1996 | 227.3972 | 226.008 | 275.2046 | 292.3972 | 147.8034 | 203.5988 | 230.4042 | 234.1946 | 325.4143 | 263.4002 | 317.3962 | 358.8044
2394 | 28.3284 | 162.4249 | 212.8627 | 245.8497 | 156.8126 | 148.525 | 173.2876 | 188.8874 | 142.2999 | 157.5127 | 110.2876 | 119.525 | 149.0625 | 190.4001 | 163.825 | 164.4249 | 204.6753 | 176.1002 | 151.4249 | 141.8497 | 171.7625 | 154.6376 | 325.4873 | 404.6246 | 151.0124 | 236.175 | 220.7749 | 273.9499 | 290.5751 | 132.2999 | 105.6376 | 141.3878 | 145.2251 | 150.7124 | 142.2251 | 174.0124 | 96.3247 | 292.6877 | 258.6623 | 334.8243 | 241.8998 | 135.2999 | 134.7124 | 162.8627 | 123.6623 | 79.7749 | 78.2498 | 80.7001 | 111.0748 | 186.5875 | 209.0625 | 256.9623 | 141.7001 | 120.2251 | 138.175 | 158.5374 | 109.8627 | 175.2876 | 176.4372 | 239.9623 | 161.0124 | 134.8874 | 120.7248 | 178.5751 | 147.7749 | 171.7625 | 195.0625 | 264.475 | 215.7502 | 141.7625 | 175.7749 | 299.2876 | 195.2498 | 197.1627 | 197.0625 | 232.4249 | 300.7001 | 161.7749 | 165.1249 | 202.7378 | 204.65 | 159.2375 | 134.7625 | 153.9375 | 143.7001 | 166.35 | 163.2999 | 187.1002 | 132.9375 | 117.7001 | 128.8751 | 117.4001 | 134.3377 | 315.0124 | 282.1126 | 360.0878 | 357.7625 | 272.5498 | 283.6376 | 363.3377 | 344.2251 | 123.6376 | 77.7873 | 85.475 | 123.6376 | 110.9499 | 137.0501 | 119.525 | 103.3624 | 109.6376 | 104.9252 | 129.9623 | 110.825 | 89.175 | 85.5875 | 83.2375 | 84.9375 | 114.7625 | 101.35 | 101.9375 | 85.1874 | 214.0501 | 234.7378 | 272.5003 | 270.3878 | 181.9245 | 196.7001 | 231.4502 | 220.8874 | 365.1126 | 274.3123 | 299.8126 | 367.0501
2395 | 28.3394 | 171.7257 | 207.6196 | 255.1061 | 150.5487 | 154.77 | 171.3983 | 184.9291 | 143.177 | 148.2478 | 116.8939 | 130.8674 | 153.6991 | 192.4691 | 161.6017 | 160.9822 | 188.0709 | 169.2213 | 155.2035 | 163 | 172.1239 | 147 | 344.0974 | 427.6283 | 152.5043 | 235.4248 | 231.1239 | 285.6726 | 277.8761 | 123.8317 | 105.2213 | 131.77 | 131.4778 | 152.177 | 136.5487 | 173.8052 | 105.283 | 269.5043 | 274.3983 | 369.4069 | 256 | 136.6017 | 134.4778 | 149.5487 | 123.6813 | 85.0265 | 82.7787 | 74.7257 | 119.7522 | 200.4426 | 208.177 | 259.9822 | 137 | 118.7965 | 147.6196 | 155.23 | 107.1683 | 170 | 179.1591 | 250.2035 | 166.7522 | 137.7257 | 134.3009 | 176.4957 | 148.3539 | 178.4957 | 192.4778 | 274.3717 | 193.8052 | 146.7965 | 177.6283 | 293.1504 | 200.6283 | 186.8496 | 203.823 | 240.4513 | 293.8761 | 170 | 175.2478 | 185.6726 | 209.1239 | 158.8761 | 137.6726 | 152.6991 | 136.8761 | 175.0444 | 165.8761 | 179.7965 | 138.0709 | 117.6726 | 115.7787 | 125.4161 | 122.8496 | 310.9822 | 287.3452 | 348.3187 | 356.4248 | 283.9026 | 283.6461 | 351.4248 | 324.9556 | 117.6991 | 81.5574 | 84.7522 | 117.6991 | 120.1239 | 123.2035 | 120.6726 | 111.8496 | 100.3009 | 92.4248 | 139.3539 | 114.1239 | 84.6017 | 82.0265 | 85 | 91.3452 | 103.6548 | 90.5043 | 102.823 | 91.0265 | 221.8674 | 218.0974 | 248.1504 | 263.3187 | 201.5574 | 192.8496 | 226.8939 | 229.6726 | 357.6017 | 277.9556 | 271.0352 | 364.6726
2396 | 28.3504 | 178.3044 | 227.682 | 249.9269 | 156.6224 | 152.7024 | 174.6888 | 180.5561 | 150.2381 | 141.3248 | 131.6888 | 137.8878 | 152.0663 | 206.2517 | 156.5629 | 139.8146 | 200.3112 | 183.3707 | 139.4439 | 158.1922 | 165.8214 | 146.5629 | 354.4966 | 409.3844 | 144.4966 | 227.6956 | 232.6888 | 293.7483 | 283.3044 | 116.0527 | 123.7415 | 144.5629 | 128.3707 | 139.1327 | 146.119 | 163.9337 | 84.4371 | 270.2381 | 281.6224 | 344.4439 | 250.7551 | 138.9932 | 133.9932 | 154.3112 | 114.4303 | 81.8741 | 80.4303 | 79.5561 | 117.0595 | 216.9405 | 214.3639 | 241.4371 | 143.5561 | 136.7415 | 158.0663 | 145.9337 | 114.9473 | 168.6888 | 163.3044 | 231.3844 | 151.3844 | 137.3112 | 135.0595 | 166.5834 | 138.8741 | 175.5766 | 185.4371 | 272.5102 | 187.8673 | 164.3044 | 177.3044 | 286.318 | 205.1122 | 196.3044 | 213.3707 | 242.2449 | 284.0068 | 159.9473 | 165.7024 | 192.8741 | 208.0663 | 145.5102 | 146.1854 | 154.5629 | 126.5697 | 179.5102 | 160.8146 | 185.9405 | 131.9541 | 127.9337 | 116.4898 | 156.4303 | 118.7551 | 296.8078 | 285.1395 | 384.2313 | 358.3112 | 287.0527 | 301.4303 | 358.5561 | 320.4898 | 125.682 | 70.3044 | 83.8078 | 125.682 | 116.0068 | 108.7551 | 120.881 | 107.318 | 99.7483 | 101.7415 | 123.318 | 114.3776 | 86.9932 | 78.8741 | 84.1259 | 92.6361 | 89.6888 | 94.7347 | 100.5697 | 89.1854 | 244.6224 | 227.1854 | 246.5629 | 272.1327 | 176.7551 | 194.4371 | 229.0136 | 256.9796 | 344.2585 | 247.2653 | 271.5357 | 365.7551
2397 | 28.3614 | 179.5633 | 229.6076 | 267.5949 | 165.3924 | 141.0886 | 165.3608 | 180.0127 | 144.4305 | 134.886 | 125.9557 | 122.3481 | 147 | 206.6962 | 161.3924 | 142.8417 | 198.8544 | 179.114 | 131.8987 | 157.8417 | 174.2214 | 151.3924 | 345.9684 | 399.2468 | 139.114 | 228.2341 | 243.5822 | 280.2722 | 288.6076 | 123.6076 | 126.9114 | 154.3354 | 130.4051 | 141.1329 | 145.0127 | 167.2025 | 81.4051 | 284.6519 | 262.9811 | 327.4621 | 239.057 | 151.1456 | 138.5063 | 152.8544 | 125.7975 | 79.4051 | 91.7975 | 93.3924 | 121.8987 | 203.114 | 217.7659 | 254.5822 | 149.3038 | 133.1709 | 147.1583 | 165.3797 | 107.8417 | 164.7532 | 181.6392 | 211.3608 | 145.7405 | 127.7659 | 135.8544 | 152.481 | 135.5063 | 162.3924 | 176.5633 | 259.8987 | 211.0316 | 164.2152 | 174.519 | 283.0443 | 201.3292 | 209.2468 | 210.4621 | 258.4367 | 256.576 | 154.6392 | 150.943 | 186.3608 | 201.6519 | 120.3165 | 146.7532 | 141.8671 | 121.4937 | 164.2025 | 154.4051 | 177.057 | 115.3924 | 126.2595 | 117.4178 | 149.3735 | 110.6519 | 279.633 | 280.8291 | 387.8798 | 355.9557 | 294.1583 | 305.6076 | 362.9557 | 325.9114 | 129.4051 | 83.2468 | 81.0127 | 129.4051 | 110.1456 | 110.0886 | 118.8417 | 105.8417 | 107.3924 | 113 | 120.0443 | 118.1898 | 87.9114 | 84.4937 | 85.2468 | 78.4494 | 98.3354 | 110.057 | 100.4367 | 79.4178 | 246.6519 | 220.5633 | 263.0759 | 264.6962 | 167.7532 | 193.6519 | 215.943 | 263.3292 | 345.1329 | 229.6962 | 311.943 | 371.1329
2398 | 28.3724 | 170.7699 | 221.1534 | 260.9202 | 168.7684 | 154.2316 | 156.9233 | 172.6933 | 124.0015 | 151.1549 | 122.8466 | 119.8451 | 138.6902 | 220.1564 | 149.9954 | 152.3083 | 186.2285 | 173.0015 | 136.2316 | 166.8466 | 182.3804 | 153.3834 | 335.0782 | 388.1503 | 136.6948 | 234.6135 | 251.1518 | 263.7699 | 271.3819 | 123.9248 | 117 | 152.227 | 133.8482 | 152.7684 | 130.7684 | 165.9233 | 91.773 | 276.9985 | 253.0813 | 330.388 | 245.9264 | 140.6887 | 140.3865 | 154.5399 | 124.7684 | 81.0015 | 95.385 | 91.6902 | 116.0752 | 203.9264 | 208.1549 | 254.3037 | 155.2331 | 118.6933 | 140.9233 | 167.3037 | 119.1549 | 167.0782 | 193.3083 | 216.773 | 152.1534 | 126.4647 | 137.0782 | 169.4617 | 137.8482 | 172.2316 | 171.4632 | 247.612 | 226.0767 | 150.5383 | 156.1534 | 281.5368 | 187.0798 | 199.9969 | 201.385 | 235.8374 | 247.388 | 164.4617 | 172.2347 | 177 | 200.9233 | 118.9264 | 142.615 | 134.3865 | 133.4632 | 168.9248 | 147.6917 | 177.0015 | 114.6135 | 122.1549 | 111.2331 | 140.0798 | 125.158 | 263.4647 | 296.2301 | 356.9218 | 359.3098 | 277.7669 | 309.1564 | 354.7684 | 317.385 | 129.6166 | 80.4601 | 76.9248 | 129.6166 | 104.3052 | 116.7684 | 124.1534 | 107.9218 | 115.8466 | 113.9233 | 116.2285 | 121.9218 | 67.8436 | 83.9985 | 84.3067 | 76.2301 | 100.3819 | 97.5368 | 112.4617 | 70.0015 | 251.4632 | 217.7715 | 267.842 | 276.7715 | 167.3083 | 185.9985 | 209.9969 | 248.6181 | 365.5399 | 237.1549 | 318 | 369.842
2399 | 28.3834 | 166.8908 | 215.6824 | 258.9008 | 162.6303 | 155.6303 | 167.4789 | 177.9479 | 129.1042 | 152.6824 | 117.4168 | 120.2134 | 136.5832 | 212.5682 | 143.2184 | 157.3697 | 191.6403 | 161.5161 | 142.8437 | 150.0421 | 160.1563 | 145.1563 | 358.5361 | 399.015 | 135.6774 | 243.5361 | 242 | 264.6303 | 265.6353 | 126.3126 | 128.3747 | 144.9529 | 127.3597 | 148.5261 | 123.2084 | 170.3176 | 95.2555 | 266.2084 | 276.3176 | 336.0471 | 253.8908 | 133.9529 | 146.1563 | 155.7866 | 125.2655 | 91.3176 | 78.5161 | 81.5261 | 92.8337 | 205.7295 | 211.5782 | 255.4318 | 152.7295 | 128.6874 | 140.1042 | 140.9429 | 115.4689 | 172.0521 | 189.8387 | 234.4739 | 151.4218 | 145.2655 | 140.1563 | 165.7345 | 149.8437 | 173.6303 | 184.6874 | 247.6403 | 219.7866 | 153.7916 | 156.8437 | 271.6303 | 182.4639 | 182.6824 | 206.7916 | 203.5832 | 290.9629 | 166.4218 | 175.3076 | 185.5311 | 203.8958 | 128.3126 | 156.1663 | 166.2234 | 128.6774 | 174.6303 | 155.4268 | 175.9429 | 119.2705 | 115.6253 | 125.3176 | 145.4168 | 122.6724 | 301.2234 | 292.5261 | 346.7345 | 364.2605 | 276.005 | 283.5582 | 348.6303 | 317.5782 | 125.0471 | 74.9479 | 84.5261 | 125.0471 | 107.2184 | 114.4218 | 113.9429 | 103.4218 | 118.4739 | 117.8437 | 101.7345 | 115.0521 | 75.9058 | 78.8958 | 80 | 85.7966 | 95.5832 | 97.5832 | 107.7866 | 75.1042 | 238.6203 | 220.8908 | 248.0521 | 278.4689 | 165.7345 | 191.7966 | 199.3176 | 249.6253 | 372 | 238.6824 | 305.2034 | 359.5311
2400 | 28.3944 | 155.5963 | 211.5138 | 267.5963 | 169.7248 | 153.4862 | 173.1376 | 181.4312 | 129.9174 | 149 | 116.7798 | 129.4037 | 142.5688 | 189.5413 | 151.7064 | 163.8899 | 205.5963 | 155.1009 | 145.0275 | 148.9358 | 169.156 | 137.1376 | 340.0734 | 404.6881 | 131.7248 | 248.4128 | 249.7798 | 253.2477 | 283.211 | 122.9174 | 133.7064 | 141.4128 | 116.3211 | 140.7064 | 124.3486 | 158.9266 | 80 | 269.7798 | 265.8991 | 320.1376 | 245.5138 | 143.5413 | 129.8716 | 147.0826 | 124.6514 | 82.844 | 71.1284 | 82.4587 | 96.9082 | 191.9725 | 194.9266 | 268.2477 | 139.4587 | 128.6789 | 130.7064 | 149.0183 | 115.7248 | 167.1101 | 177.1651 | 227.2202 | 156.4037 | 148.055 | 140.8899 | 167.8073 | 156.4037 | 192.8807 | 190.5138 | 257.2202 | 215.4587 | 150.7064 | 156.5963 | 268.5138 | 171.7523 | 187.7523 | 218.7798 | 222.6972 | 305.4679 | 160.7064 | 174.0734 | 177.4954 | 198.7064 | 132.6972 | 153.5505 | 170.6055 | 118.8899 | 169.5688 | 153.8716 | 173.8073 | 117.9541 | 101.5413 | 129 | 144.2936 | 107.7523 | 310.9541 | 290.5413 | 357.0734 | 347.8165 | 264.1193 | 257.0734 | 343.0826 | 316.9725 | 124.6972 | 78.4312 | 87.8899 | 124.6972 | 126.4037 | 111.6239 | 110.8349 | 107.4312 | 114.6239 | 112.2477 | 108.6697 | 113.5138 | 82.4404 | 81.9725 | 83.8899 | 106.3211 | 105.945 | 105.5138 | 103.945 | 72.5138 | 228.156 | 219.8073 | 247.4862 | 271.9174 | 150.789 | 189.9541 | 212.2385 | 234.5688 | 350.6055 | 250.5596 | 296.3486 | 375.8073
2401 | 28.4054 | 157.4838 | 207.5103 | 249.0441 | 180.4985 | 143.5309 | 177.4721 | 193.9706 | 139.9794 | 138.0324 | 122.0088 | 133.4985 | 138.0059 | 182.5162 | 155.9647 | 172.9853 | 201.0029 | 169.5015 | 160.4515 | 152.0529 | 183.4956 | 148.95 | 301.4956 | 392.4603 | 145.4897 | 229.5103 | 242.0471 | 250.9794 | 287.0206 | 122.0118 | 119.5191 | 129.9882 | 124.9941 | 145.4632 | 132.4926 | 153.4691 | 79.5015 | 270.0235 | 236.0441 | 306.0265 | 254.9706 | 129.5368 | 127.9647 | 146.9912 | 128.4721 | 87.9471 | 88.4897 | 82.0059 | 110.0265 | 203.9676 | 204.9235 | 252.0206 | 147.4809 | 136.4574 | 119.5103 | 154.5368 | 128.9912 | 171.475 | 161.5279 | 222.9882 | 156.0118 | 143.0088 | 143.0059 | 177.4926 | 144.5456 | 199.05 | 183.0206 | 253.9853 | 203.5397 | 156.4603 | 165.9618 | 256.0353 | 173.0235 | 196.0029 | 211.0471 | 230.5103 | 292.4868 | 156.9882 | 189.0029 | 173.4574 | 191.9971 | 135.0265 | 148.9529 | 154.4809 | 113.0294 | 164.0088 | 144.9853 | 178.0088 | 121.4397 | 99.5015 | 131.9912 | 137.5426 | 119.9912 | 290.0118 | 302.9588 | 379.9794 | 320.0471 | 259.4456 | 268.0147 | 346.4809 | 308.5279 | 135.0029 | 86.4838 | 91.5015 | 135.0029 | 128.5044 | 102.5132 | 114.0088 | 108.5044 | 109.0029 | 113.4691 | 108.0382 | 112.5015 | 77.9647 | 69.5397 | 87.0029 | 95.5515 | 102.5162 | 117.9618 | 103.0088 | 73.4956 | 236.0147 | 228.9941 | 251.9882 | 281.4809 | 157.9471 | 192.4426 | 217.0147 | 233.9941 | 326.5044 | 257.0294 | 294.0235 | 367.0471
2402 | 28.4164 | 160.4459 | 208.5974 | 239.1082 | 175.381 | 141.1731 | 168.6105 | 200.4243 | 152.1082 | 137.2164 | 112.3593 | 132.9784 | 139.0649 | 179.0433 | 158.2944 | 170.8485 | 193.8702 | 170.0216 | 161.6754 | 135.5325 | 179.8918 | 156.2944 | 319.3463 | 408.0433 | 163.8138 | 218.3377 | 225.4892 | 264.1298 | 274.8918 | 110.8485 | 114.5325 | 134 | 123.9351 | 145.7403 | 146.2381 | 170.1298 | 89.2164 | 252.8052 | 243.9869 | 300.5757 | 255.2944 | 125.1731 | 132.3377 | 149.4892 | 129.3161 | 100.8918 | 90.4675 | 76.4243 | 102.0216 | 208.3593 | 232.0216 | 245 | 160.1298 | 143.3377 | 126.7272 | 157.8354 | 117.697 | 164.1646 | 169.8787 | 236.7056 | 147.4026 | 132.3377 | 143.5541 | 168.2511 | 132.0649 | 187.1082 | 167.8269 | 247.2511 | 216.0518 | 162.3377 | 165.7187 | 253.7056 | 165 | 192.9567 | 207.2597 | 254.5843 | 290.8702 | 166.1082 | 173.1862 | 179.8269 | 198.1082 | 119.8702 | 145.078 | 163.0433 | 123.9436 | 167.6407 | 163.2813 | 172.4459 | 137.9134 | 106.6623 | 127.8485 | 134.7489 | 125.5541 | 291.619 | 302.1862 | 369.1213 | 329.0302 | 261.6537 | 272.1948 | 337.6754 | 301.0433 | 134.5108 | 82.2944 | 89.9784 | 134.5108 | 126.4892 | 103.1082 | 118.1515 | 112.1082 | 111.5757 | 112.2511 | 108.7922 | 108.9351 | 78.7619 | 64.6839 | 83.4459 | 89.2381 | 106.1948 | 116.697 | 109.1948 | 82.6623 | 234.5757 | 229.4675 | 247.3161 | 287.4892 | 174.9784 | 200.7619 | 205.8702 | 253.8787 | 336.7489 | 232.697 | 272.2078 | 358.1515
2403 | 28.4274 | 158.5231 | 197.3067 | 235.1071 | 153.2604 | 154.2311 | 155.708 | 181.3067 | 142.876 | 148.0462 | 116.9853 | 130.9538 | 138.8613 | 188.3236 | 144.8151 | 180.2165 | 200.5547 | 176.2311 | 140.7227 | 137.5231 | 170.292 | 162.6933 | 344.416 | 383.8444 | 154.9831 | 228.2627 | 217.1533 | 268.9538 | 272.6156 | 100.8613 | 112.8613 | 130.8613 | 119.4307 | 144.9853 | 142.876 | 165.5378 | 93.2458 | 260.7396 | 262.8613 | 330.6787 | 236.0609 | 135.0924 | 123.4307 | 147.4307 | 116.292 | 94.4307 | 85.3382 | 86.6009 | 103.5231 | 195.1996 | 218.876 | 241.8613 | 160.2458 | 145.416 | 129.1533 | 165.6764 | 110.2773 | 160.5085 | 170.2604 | 230.3067 | 149.2773 | 132.3236 | 147.5693 | 152.8151 | 135.5231 | 187.292 | 179.3551 | 226.0609 | 227.3991 | 164.9391 | 151.4307 | 270.3236 | 192.7249 | 183.6764 | 216.9076 | 269.4915 | 282.3844 | 171.5231 | 168.416 | 179.8467 | 198.8151 | 134.9245 | 127.5693 | 158.7227 | 134.584 | 177.1387 | 160.8298 | 168.4307 | 139.7542 | 112.4307 | 115.7689 | 140.2458 | 123.292 | 293.3382 | 298.4622 | 366.6471 | 353 | 254.8929 | 287.8004 | 352.8174 | 295.1533 | 142.3236 | 69.8613 | 87.4307 | 142.3236 | 123.9076 | 117.9391 | 128.1387 | 117 | 112.9076 | 109.3698 | 109.4453 | 108.6156 | 77.9391 | 73 | 78.3844 | 96.3382 | 112.3844 | 94.6302 | 118 | 79.5378 | 242.708 | 229.0462 | 241.0924 | 291.708 | 174.5231 | 202.5547 | 213.0778 | 251.6449 | 371.0169 | 226.8467 | 272.0778 | 353.4915
2404 | 28.4384 | 155.2522 | 188.354 | 238.7789 | 149.7789 | 153.1106 | 157.323 | 172.8894 | 133.2124 | 147.9292 | 133.4248 | 130 | 142.9602 | 188.0398 | 135.1106 | 183.2211 | 203.4336 | 180.4646 | 124.8274 | 152.9913 | 175.1018 | 170.0398 | 329.9777 | 373.9205 | 147.9205 | 229.5441 | 222.3143 | 254.0795 | 273.3938 | 118.3453 | 120.1726 | 115.1503 | 111.5752 | 152.8584 | 138.031 | 158.1416 | 96.5664 | 271.7168 | 261.6062 | 336.2609 | 230.2124 | 146.1018 | 134.8497 | 145.4646 | 114.1416 | 84.4336 | 84.677 | 92.0398 | 107.7478 | 190.0708 | 206.5354 | 236.323 | 150.7168 | 151.8584 | 119.323 | 155.7876 | 107.4336 | 160.8274 | 175.3453 | 242.7701 | 144.8274 | 132.0398 | 157.031 | 161.8497 | 132.7876 | 188.354 | 173.9777 | 232.5267 | 208.5752 | 157.4025 | 173.5577 | 269.5044 | 188.5529 | 186.1018 | 206.4336 | 238.1193 | 285.354 | 163.969 | 167.8982 | 191.2832 | 198.2124 | 144.8982 | 124.7168 | 154.0708 | 144.2037 | 165.5441 | 147 | 189.4869 | 140.7168 | 111.5354 | 120.1018 | 144.6372 | 121.1416 | 299.6372 | 287.1193 | 362.3317 | 357.2832 | 263.5354 | 281.6857 | 364.4733 | 286.3938 | 143.1106 | 68.6062 | 80.646 | 143.1106 | 117.1814 | 120.5752 | 127.7876 | 106.2919 | 113.1416 | 117.5354 | 115.1329 | 102.4336 | 82.7168 | 73.5354 | 73.8584 | 95.1416 | 102.5044 | 95.5664 | 105.1503 | 73.7478 | 229.3317 | 223.0398 | 255.8497 | 269.7653 | 184.6372 | 196.8671 | 213.7565 | 251.6683 | 368.4423 | 228.1106 | 292.031 | 351.031
2405 | 28.4494 | 158.5723 | 201.2154 | 237.4984 | 170.9539 | 148.5477 | 158.2862 | 169.7846 | 140.9292 | 143.1661 | 120.3784 | 134.9292 | 144.6185 | 198.9785 | 143.6923 | 166.4277 | 199.2862 | 180.5477 | 132.9785 | 167.6431 | 183 | 151.9507 | 324.7877 | 390.9292 | 161.6431 | 215.3569 | 230.2615 | 242 | 276.9292 | 127.2369 | 123.5231 | 122.0739 | 115.8585 | 154.8339 | 143.9046 | 158.9046 | 99.6185 | 263.6185 | 255.88 | 310.1415 | 245.597 | 153.4523 | 125.1876 | 153.2154 | 119.2862 | 78.0954 | 74.9507 | 90.3815 | 103.3323 | 188.8092 | 193.3077 | 251.5262 | 150.2862 | 136.3077 | 126.8585 | 153.5477 | 106.0246 | 162.9539 | 176.5692 | 240.2584 | 156.8124 | 118.88 | 154.1415 | 165.88 | 139.3108 | 182.0461 | 178.6462 | 248.1908 | 211.2154 | 143.5477 | 177.5692 | 275.0493 | 180.5262 | 185.2369 | 185.8554 | 246.837 | 290 | 167.9539 | 169.3108 | 191.1661 | 207.5723 | 151.24 | 134.6923 | 149.5231 | 149.0461 | 155.7385 | 154.6677 | 198.0461 | 140.2862 | 108.7138 | 121.9754 | 144.7846 | 125.1908 | 310.7385 | 273.3815 | 360.1446 | 360.4523 | 256.3323 | 283.0493 | 362.5016 | 317.8616 | 126.4984 | 68.3569 | 72.7138 | 126.4984 | 121.7631 | 113.3569 | 120.6185 | 106.3108 | 100.2123 | 116.3569 | 132.7385 | 107.0493 | 79.5477 | 69.6185 | 72.5477 | 92.0708 | 95.4523 | 98.0708 | 102.7631 | 79.1908 | 233.1693 | 225.7631 | 261.5231 | 245.9046 | 184.7846 | 186.3815 | 206.7385 | 265.3569 | 335.9507 | 228.4769 | 268.8769 | 365.6677
2406 | 28.4604 | 158.4002 | 211.3599 | 226.32 | 173.2802 | 146.2001 | 153.7202 | 169.7198 | 151.7198 | 136.6401 | 115.0797 | 129.4803 | 133.1602 | 213 | 155 | 159.32 | 197.5201 | 168.1204 | 154.2798 | 146.0406 | 183.56 | 145.3599 | 342.5998 | 376.5205 | 148.4404 | 221.2798 | 221.2802 | 250.3998 | 256.9207 | 126.7198 | 111.7202 | 132.5201 | 136.8796 | 154.6401 | 149.1598 | 163.5998 | 87.6002 | 265.0399 | 256.7198 | 315.9996 | 238.6405 | 139.0004 | 116.3998 | 161.12 | 122 | 80.68 | 75.6397 | 86.1602 | 98.68 | 192.0399 | 197.6796 | 244.7206 | 155.2399 | 136.7597 | 145.0797 | 155.68 | 119.3998 | 159.6803 | 157.4002 | 225.1598 | 156.0805 | 130.9195 | 147.12 | 156.6401 | 139.1602 | 184.7597 | 191.7202 | 237.1204 | 207.3603 | 147.9199 | 164 | 281.0801 | 192.7601 | 180.2399 | 185.6397 | 245.4806 | 294.4799 | 164.1204 | 165.2403 | 181.2802 | 202.3603 | 156.2802 | 148.2399 | 131.5604 | 139.44 | 177.5995 | 173.3197 | 191.7999 | 136.2802 | 116.6397 | 119.7999 | 145.8398 | 134.8398 | 287.2407 | 287.0797 | 354.2005 | 354.9601 | 247.7601 | 253.8011 | 365.2802 | 310.8409 | 129.3197 | 77.0797 | 78.9598 | 129.3197 | 113.8804 | 104.7202 | 118.12 | 109.5201 | 93.0399 | 113.32 | 125.4803 | 112.5201 | 81.68 | 64.32 | 70.7601 | 99.1997 | 106.1997 | 96.2399 | 123.4396 | 78.2001 | 267.4795 | 224.6002 | 253.0801 | 251.1598 | 161.7604 | 181.6002 | 212.9199 | 276.8796 | 355.6792 | 235.2399 | 284.8789 | 345.6807
2407 | 28.4714 | 166.5901 | 201.9822 | 222.7109 | 171.4337 | 143.4277 | 150.8614 | 169.2772 | 159.8614 | 153.4574 | 110.9822 | 119.1386 | 144.1683 | 198.1208 | 151.5663 | 141.9763 | 196.8614 | 164.295 | 154.2772 | 125.7109 | 184 | 147.4277 | 329.8317 | 381.4574 | 145.5841 | 225.2832 | 224.0119 | 281.6079 | 262.0356 | 130.8554 | 116.301 | 134.7228 | 138.9881 | 143.9881 | 145.9881 | 154.2654 | 89.5841 | 264.9941 | 248.8377 | 304.109 | 237.0178 | 132.0059 | 132.1564 | 165.4337 | 123.1446 | 83.1446 | 81.1386 | 88.0119 | 109.1564 | 193.1386 | 204.8495 | 226.8554 | 148.4159 | 140.2772 | 140.4099 | 151.8495 | 121.4218 | 156.295 | 163.8732 | 221.9881 | 153.4455 | 142.1386 | 138.8436 | 156.2891 | 123.8436 | 185.9881 | 186 | 237.301 | 199 | 147.5663 | 163.4277 | 282.0059 | 215.6079 | 197.4514 | 184.8436 | 226.4277 | 284.2654 | 165.4455 | 153.1386 | 187.4455 | 199.7228 | 151.5723 | 136.2654 | 143.3188 | 139 | 176.9763 | 172.1268 | 188.2772 | 129.2832 | 118.705 | 120.8554 | 140.5545 | 117.5367 | 288.1742 | 274.3981 | 336.4218 | 355.0059 | 253.4396 | 263.6316 | 345.1208 | 294.301 | 127.5545 | 87.2891 | 77.9881 | 127.5545 | 105.4337 | 109.8732 | 117.8554 | 119.1623 | 102.1505 | 106.8495 | 114.5663 | 108.4277 | 76.705 | 65.2891 | 73.0059 | 104.5663 | 109.8495 | 102.0059 | 131.1386 | 86.8732 | 239.6516 | 209.9881 | 256.295 | 263.4396 | 170.7465 | 184.7287 | 211.4218 | 253.808 | 374.4218 | 241.0059 | 290.103 | 322.1386
2408 | 28.4824 | 174.2463 | 193 | 221 | 165.2315 | 147.6765 | 167.0296 | 156.2315 | 159.6617 | 149.8784 | 111.9376 | 117.5846 | 152.908 | 186.4154 | 145.4926 | 143.445 | 203.092 | 178.3531 | 142.9852 | 130.4302 | 189.8457 | 149.3383 | 317.5846 | 396.5846 | 156.092 | 226.9228 | 237.5994 | 260.834 | 296.9524 | 118.8932 | 115.8161 | 134.908 | 123.6469 | 150.2759 | 135.3235 | 168.5518 | 97.1691 | 269.5994 | 248.9376 | 315.6438 | 239.5698 | 144.3531 | 137.2463 | 164.4926 | 128.092 | 80.4926 | 93.0296 | 91.0772 | 111.3235 | 188.6617 | 184.0476 | 245.2907 | 143.7537 | 145.3531 | 130.4154 | 136.8932 | 116.2463 | 158.0772 | 169.6617 | 199.0476 | 139.7865 | 141.7537 | 134.9228 | 165.5994 | 122.8457 | 180 | 183.6617 | 238.5698 | 200.7537 | 140.3235 | 155.9852 | 260.4482 | 191.9112 | 221.8605 | 177.4154 | 217.8161 | 288.0296 | 171.6617 | 152.1691 | 181.3858 | 206.3383 | 151.4154 | 137.6913 | 148.8932 | 140.7537 | 162.6617 | 175.1068 | 184 | 120.4006 | 119.8457 | 142.2135 | 147.7833 | 101.7537 | 294.0624 | 272.445 | 338.2611 | 336.371 | 269.5222 | 283.4778 | 314.1247 | 306.092 | 127.7685 | 81.4006 | 73.7537 | 127.7685 | 106.8309 | 113.908 | 118.1691 | 107.9556 | 101.3235 | 97.1543 | 105.5698 | 109.1691 | 77.2611 | 76.3531 | 68.4006 | 90.8932 | 105.4154 | 95.6469 | 115.555 | 89.1543 | 236.6438 | 223.8753 | 253.9852 | 256.7241 | 179.4006 | 184.4006 | 202.7389 | 219.0624 | 349.371 | 256.8605 | 289.2907 | 339.2907
2409 | 28.4934 | 160.466 | 200.1623 | 245.4712 | 157.2094 | 166.5183 | 172.2251 | 167.3089 | 144.8691 | 139.9843 | 120.1937 | 128.1466 | 153.5811 | 193.1623 | 145.9843 | 150.6126 | 205.4031 | 168.8848 | 147.5497 | 140.3717 | 177.8848 | 160.5497 | 319.7906 | 376.7068 | 149.4503 | 233.178 | 230.466 | 241.9529 | 289.3037 | 120.5497 | 109.4031 | 133.1937 | 124.7592 | 139.3037 | 137.3717 | 175.8534 | 100.3874 | 279.7749 | 251.822 | 337.1937 | 255.2932 | 155.178 | 134.2094 | 154.2408 | 116.0786 | 90.534 | 93.4503 | 84.2251 | 110.3874 | 195.9529 | 188.712 | 250.0471 | 139.6283 | 132.9005 | 138.9529 | 129 | 114.4031 | 150.6283 | 179.9372 | 203.1152 | 125.7906 | 140.0157 | 140.5811 | 173.3874 | 140.7277 | 184.178 | 180.2094 | 238.178 | 209.1623 | 156.0995 | 153.3874 | 244.7906 | 176.3246 | 222.644 | 174.0157 | 226.9214 | 305.1623 | 181.9372 | 164.3403 | 163.2408 | 202.6283 | 154.5811 | 138.2408 | 140.4031 | 159.9057 | 165.7749 | 168.0786 | 191.7592 | 110.822 | 108.4817 | 145.466 | 159 | 103.5969 | 297.7434 | 284.3874 | 322.9005 | 327.5654 | 288.3403 | 273.6126 | 316.1152 | 310.7906 | 128.0314 | 81.3717 | 75.5969 | 128.0314 | 110.7749 | 96.0786 | 113.6283 | 108.5183 | 100.9843 | 105.534 | 106.9686 | 115.9686 | 69.6597 | 77.0314 | 80.9057 | 84.7906 | 115.1466 | 93.178 | 100.0314 | 85.5969 | 225.9634 | 218.9005 | 265.1152 | 262.3246 | 171.0157 | 178.4031 | 213.3246 | 213.1937 | 337.5811 | 254.6597 | 291.6597 | 370.9057
2410 | 28.5044 | 164.183 | 196.4719 | 235.1974 | 168.183 | 158.1164 | 162.9085 | 175.9543 | 128.6902 | 157.8379 | 127.7006 | 128.2994 | 147.4719 | 188.8628 | 158.3555 | 149.6091 | 183.6798 | 156.1726 | 153.3909 | 148.8732 | 174.9189 | 169.4366 | 314.2994 | 372.1372 | 136.9085 | 230.5177 | 223.8274 | 246.1726 | 243.7151 | 112.9896 | 98.6445 | 146.7921 | 137.3098 | 143.0562 | 139.7817 | 154.3804 | 97.736 | 275.0811 | 271.5385 | 310.5883 | 253.7713 | 151.9085 | 128.1268 | 153.264 | 112.0915 | 89.2536 | 85.1726 | 79.7817 | 113.8274 | 197.1268 | 187.2079 | 236.0811 | 145.1372 | 126.7006 | 140.1268 | 127.7817 | 120.0915 | 166.4928 | 188 | 214 | 132.4823 | 138.6091 | 141.7817 | 166.4719 | 144.5177 | 183.3451 | 196.0562 | 219.6798 | 210.9543 | 154.9896 | 148.9085 | 258.183 | 189.6549 | 215.7817 | 195.7568 | 249.183 | 295.9896 | 185.1268 | 170.7817 | 176.2745 | 193.5177 | 149.6902 | 147.6196 | 143.0457 | 151.2889 | 159.8628 | 181.7568 | 189.6902 | 117.1372 | 113.8379 | 113.2432 | 148.0353 | 120.447 | 298.9085 | 299.4013 | 319.1372 | 325.2994 | 288.0811 | 284.7921 | 338.5738 | 316.8732 | 122.7817 | 85.6091 | 82.0915 | 122.7817 | 107.9085 | 105.4928 | 113.6549 | 103.1621 | 108.4823 | 109.736 | 127.447 | 127.3098 | 79.0562 | 74.2183 | 81.4262 | 83.5634 | 98.8523 | 85.6445 | 102.0915 | 90.8732 | 217.7921 | 212.7006 | 263.817 | 268.3451 | 165.3451 | 190.183 | 208.3804 | 218.8732 | 354.0104 | 240.9085 | 276.0811 | 356.1974
2411 | 28.5154 | 163.2998 | 220.2077 | 217.3786 | 178.4858 | 152.2144 | 152.7856 | 181.4573 | 145.129 | 135.6851 | 111.4925 | 122.1357 | 131.4355 | 187.3501 | 165.6214 | 149.3786 | 178.0787 | 165.5645 | 153.6214 | 153.8643 | 161.9782 | 173.4858 | 318.7002 | 379.8643 | 148.536 | 226.3786 | 228.1072 | 230.0854 | 250.9648 | 118.9146 | 99.4573 | 152.5142 | 128.9498 | 160.2144 | 134.0285 | 154.943 | 92.5142 | 267.5142 | 280.407 | 325.3149 | 261.4003 | 139.2998 | 133.7002 | 164.0787 | 112.8928 | 81.7571 | 77.7856 | 79.6214 | 93.2496 | 195.2429 | 183.9715 | 241.5645 | 155.3501 | 129.7571 | 133.8928 | 125.7571 | 114.6784 | 169.057 | 173.7069 | 204.6784 | 144.7002 | 149.5645 | 137.8928 | 156.0285 | 136.0285 | 176.6499 | 190.2211 | 219.0502 | 213.9715 | 153.4573 | 167.9933 | 267.2429 | 169.0067 | 200.7069 | 196.0854 | 229.0352 | 290.7286 | 170.1926 | 184.2931 | 171.2211 | 209.2646 | 142.5142 | 161.9715 | 166.129 | 142.3501 | 170.7789 | 181.464 | 180.6499 | 119.8928 | 122.7571 | 107.3216 | 149.0787 | 124.1642 | 286.9213 | 306.7571 | 331.8358 | 323.4858 | 268.0854 | 273.7354 | 330.464 | 301.3568 | 107.7069 | 82.8928 | 85.3786 | 107.7069 | 115.8074 | 106.1926 | 126.5645 | 104.1859 | 104.5427 | 102.6499 | 124.9498 | 110.2496 | 71.3568 | 73.1357 | 83.3216 | 92.5645 | 97.0502 | 90.1859 | 109.1072 | 94 | 232.2144 | 220.7286 | 262.2144 | 261.6499 | 159.2714 | 201.7286 | 192.7856 | 236.9146 | 340.0067 | 238.2429 | 272.8643 | 346.4573
2412 | 28.5264 | 166.8722 | 238.6337 | 227.7734 | 161.6218 | 143.9591 | 147.7325 | 167.6218 | 142.1567 | 136.9131 | 121.1806 | 127.971 | 130.0698 | 202.0409 | 171.0698 | 148.3663 | 187.4361 | 166.9302 | 151.4651 | 153.7325 | 157.3373 | 167.3952 | 324 | 382.9012 | 154.8313 | 226.6337 | 223.029 | 245.7155 | 269.9012 | 132.4361 | 122.3782 | 148.4651 | 121.6337 | 151.3254 | 133.5349 | 152.1278 | 89.0988 | 242.552 | 246.5809 | 312.6797 | 244.1158 | 132.7325 | 151.0409 | 143.017 | 123.0409 | 90.506 | 78.4361 | 81.2675 | 103.448 | 192.8313 | 197.7734 | 248.6337 | 158.6337 | 133.4361 | 135.8024 | 143.3782 | 118.506 | 166.8024 | 171.971 | 205.971 | 146.8313 | 149.029 | 126.494 | 173.2794 | 137.4361 | 189.8433 | 177.4651 | 222.5639 | 200.523 | 145.9591 | 180.7325 | 264.8313 | 177.1806 | 185.029 | 216.1516 | 229.1107 | 293.6337 | 158.5639 | 180.3254 | 169.8722 | 189.3134 | 144.1687 | 149.1567 | 162.523 | 148.8024 | 162.6218 | 161.8603 | 171.029 | 110.3952 | 127.0698 | 129.477 | 147.6627 | 120 | 292.7734 | 290.7905 | 315.7206 | 325.6337 | 261.5349 | 287.3492 | 336.2095 | 298.8722 | 107.2385 | 89.2385 | 79.2964 | 107.2385 | 114.7615 | 111.0409 | 125.3952 | 120.506 | 112.6746 | 106.971 | 125.8722 | 105.2385 | 65.7036 | 75.8024 | 75.0579 | 92.029 | 98.029 | 103.971 | 100.2266 | 105.4072 | 225.2266 | 224.2675 | 242.552 | 265.971 | 167.1397 | 193.2266 | 207.3782 | 246 | 340.5758 | 233.2964 | 280.971 | 352.2675
2413 | 28.5374 | 165.5398 | 224.7875 | 222.3716 | 169.0886 | 140.8761 | 156.6903 | 171.8585 | 133.646 | 144.4778 | 119.1415 | 137.1682 | 149.7966 | 206.4159 | 161.0796 | 166.0886 | 206.1506 | 179.5045 | 146.7699 | 152.354 | 173.9204 | 151.3716 | 323.354 | 380.7699 | 149.7699 | 230.8761 | 208.0176 | 242.5574 | 268.4159 | 128.5398 | 118.7875 | 145.0619 | 117.4778 | 138.8318 | 137.5841 | 148.292 | 93.8142 | 254.8409 | 229 | 300.4602 | 257.071 | 137.1682 | 144.4335 | 155.7789 | 133.2301 | 92.1239 | 77.7699 | 87.8142 | 121.5222 | 203.2744 | 212.3983 | 238.0176 | 144.1415 | 128.2477 | 138.646 | 150.1239 | 122.0619 | 172.8761 | 153.3892 | 220.9824 | 148.2301 | 135.9557 | 139.0886 | 185.646 | 141.292 | 182.2034 | 168.8938 | 211.6017 | 219.425 | 162.2568 | 171.6017 | 260.4159 | 169.9733 | 181 | 204.575 | 239.646 | 292.0619 | 156 | 160.7256 | 189.2125 | 200.717 | 140.8318 | 140 | 157.5222 | 140.6636 | 155.8761 | 168.2744 | 165.708 | 115.0443 | 128.0619 | 128.0176 | 134.3097 | 127.1062 | 313.8585 | 274.2477 | 333.717 | 327.292 | 283.0267 | 275.5131 | 319.8051 | 289.1415 | 130.7346 | 82.3716 | 77.9381 | 130.7346 | 119.6903 | 107.0176 | 115.1858 | 127.9381 | 108.3716 | 114.2301 | 123.2477 | 111.292 | 70.9381 | 77.354 | 79.2744 | 93.1682 | 91.4159 | 93.7875 | 116.6108 | 91.9733 | 250.0091 | 230.8142 | 247.7346 | 273.8761 | 183.9824 | 199.2744 | 208.3097 | 247.292 | 370.3187 | 232.5841 | 266.9114 | 351.708
2414 | 28.5484 | 162 | 209.7586 | 213.3667 | 168.467 | 161.4328 | 170.5581 | 164.2005 | 126.7586 | 148.5831 | 116.2665 | 141.3166 | 138.5923 | 197.1002 | 151.3667 | 160.2005 | 210.3918 | 193.2665 | 142.3667 | 146.0752 | 173.1002 | 142.3667 | 313.1253 | 367.1504 | 143.3918 | 232.3417 | 224.3826 | 257.3576 | 275.5581 | 111.1754 | 115.6082 | 128.8588 | 129.4829 | 146.533 | 136.3667 | 162.8246 | 97.6583 | 268.3417 | 242.8246 | 314.533 | 238.1094 | 132.1002 | 129.7586 | 145.9841 | 125.7836 | 78.8087 | 89.1663 | 94.9499 | 123.3417 | 181.6925 | 214.3667 | 253.066 | 143.8998 | 132.5581 | 126.4921 | 149.3166 | 110.467 | 153.934 | 160.7494 | 233.5831 | 161.1913 | 137.5831 | 154.2665 | 172.8337 | 130.8087 | 165.7586 | 176.1913 | 209.6333 | 202.0843 | 166.4419 | 155.8087 | 236.6174 | 166.8998 | 190.8747 | 195.5831 | 213.0092 | 294.9499 | 156 | 179.6742 | 189.1002 | 201.8588 | 133.3918 | 135.3918 | 140.2506 | 135.6583 | 149.4419 | 176.9749 | 169.6082 | 121.3166 | 109.8839 | 136.4829 | 129.6583 | 116.5171 | 302.2506 | 256.8337 | 346.7084 | 339.1913 | 256.4762 | 261.6583 | 315.8747 | 292.1913 | 123.2255 | 75.3417 | 84.2665 | 123.2255 | 113.1504 | 116.7995 | 113.3166 | 92.1345 | 111.2164 | 116.6583 | 107.1504 | 106.7335 | 81.2164 | 69.7586 | 84.6834 | 99.9499 | 93.9499 | 99.8246 | 109.2506 | 76.3918 | 253.1754 | 222.1504 | 251.4169 | 267.4419 | 184.7335 | 182.959 | 191.1504 | 244.0501 | 371.467 | 222.1504 | 268.8496 | 343.7586
2415 | 28.5594 | 173.4003 | 224.1063 | 218.706 | 174.4003 | 150.2113 | 159.5761 | 165.3885 | 137.7533 | 139.2585 | 125.0354 | 133.2821 | 145.4357 | 190.3176 | 147.9882 | 158.0354 | 189.8937 | 175.2113 | 162.4593 | 165.1299 | 162.9646 | 150.3885 | 335.4948 | 361.6706 | 148.3766 | 222.6115 | 206.4934 | 264.6352 | 257.1995 | 126.8005 | 117.3294 | 134.4121 | 139.6824 | 155.353 | 147.7415 | 141.164 | 92.6352 | 253.9173 | 257.3766 | 313.294 | 223.353 | 140.7415 | 136.0591 | 145.0709 | 115.6352 | 82.3885 | 85.9409 | 86.2703 | 101.5407 | 169.6824 | 204.2821 | 251.9291 | 140.6234 | 125.5997 | 126.0354 | 141.2821 | 115.0591 | 163.7887 | 176.3648 | 225.5997 | 158.9527 | 126.9173 | 148.9527 | 163.3176 | 116.9527 | 190.836 | 175.294 | 218.3766 | 209.1418 | 162 | 137.9291 | 253.1654 | 178.3766 | 214.1063 | 190.9527 | 223.1418 | 271.517 | 167.4003 | 174.2231 | 192.3766 | 198.0236 | 136.3648 | 130.3176 | 148.7769 | 139.353 | 159.0827 | 168.6115 | 184.7415 | 124.0118 | 112.4003 | 121.8701 | 132.6824 | 119.0591 | 287.9764 | 272.8005 | 318.8464 | 344.3294 | 277.9069 | 249.2585 | 325.6942 | 321.4711 | 118.6942 | 65.6115 | 75.5997 | 118.6942 | 107.6706 | 115.6115 | 110.647 | 104.5184 | 107.2821 | 102.9173 | 112.4003 | 106.0118 | 90.0236 | 82.7651 | 76.6234 | 108.0354 | 96 | 109.6824 | 105.0354 | 80.706 | 231.9173 | 213.3176 | 225.1995 | 268.3648 | 178.647 | 168.9882 | 206.4593 | 234.6234 | 373.3766 | 218.6824 | 264.9409 | 343.353
2416 | 28.5704 | 155.7816 | 220.7079 | 206.2934 | 180.6829 | 155.0237 | 160.1947 | 157.7079 | 147.0487 | 135.7316 | 115.7079 | 122.171 | 160.829 | 189.6829 | 143.5855 | 167.1461 | 198.0724 | 160.9026 | 160.7079 | 148.0013 | 149.7566 | 152.2684 | 334.6592 | 339.4645 | 162.2434 | 252.1448 | 189.9513 | 264.0487 | 269.9013 | 121.6105 | 115.6342 | 120.5131 | 131.4395 | 157 | 138.9763 | 146.1211 | 84.5368 | 264.0724 | 240.5131 | 288.8302 | 236.6092 | 156.5605 | 134.8539 | 148.9513 | 123.2434 | 88.3658 | 86.4632 | 84.4145 | 103.2921 | 191.4869 | 195.2197 | 246.6829 | 143.829 | 112.4882 | 135.1461 | 143.829 | 110.4395 | 176.0487 | 164.6592 | 228.8776 | 141.3421 | 124.7803 | 165.4869 | 154.4882 | 130.7553 | 180.4158 | 180.8776 | 213.1224 | 221 | 149.025 | 153.8526 | 262.9026 | 179.2684 | 178.6118 | 189.7316 | 224.7566 | 283.5842 | 160.7079 | 181.3895 | 178.9276 | 197.9513 | 155.3895 | 125.5855 | 138.196 | 153.975 | 159.171 | 159.2197 | 167.0987 | 134.5605 | 109.8053 | 125.3408 | 118.9763 | 113.7566 | 275.0737 | 297.6579 | 289.6105 | 351.5118 | 276.1474 | 246.4145 | 323.2197 | 309.7816 | 114.8539 | 65.7803 | 80.9263 | 114.8539 | 110.0487 | 110.3171 | 116.5118 | 104.2934 | 107.7803 | 106.2434 | 109.8053 | 100.171 | 93.3658 | 81.4395 | 81.1947 | 89.1474 | 100.0974 | 106.2197 | 97.0737 | 79.2197 | 246.8526 | 216.0974 | 231.7553 | 245.7329 | 158.5618 | 181.6579 | 200.6105 | 206.4158 | 341.4894 | 239.1211 | 246.3421 | 327.9276
2417 | 28.5814 | 170.7219 | 199.0107 | 213.5989 | 153.8877 | 150.877 | 176.5134 | 154.0856 | 149.3904 | 146.0374 | 130.8556 | 134.2941 | 147.3155 | 194.1711 | 145.4759 | 164.4385 | 172.6311 | 181.246 | 141.0963 | 147.5134 | 173.5026 | 157.2567 | 314.0963 | 327.6096 | 146.8396 | 231.8503 | 210.5508 | 267.7807 | 263.6203 | 119.5615 | 118.4759 | 119.3422 | 122.8289 | 166.0374 | 142.4278 | 146.6578 | 87.5615 | 272 | 243.5134 | 276.9144 | 233.6578 | 142.23 | 130.6096 | 147.3048 | 110.6203 | 79.9626 | 82.0481 | 99.2085 | 106.9144 | 189.877 | 201.3422 | 246.3048 | 142.8289 | 118.0374 | 135.9144 | 144.2193 | 121.2941 | 173.5241 | 155.9144 | 228.8289 | 154.4652 | 136.7326 | 140.9359 | 183.6738 | 119.5348 | 159.9626 | 187.7807 | 218.7326 | 219.6096 | 158.9893 | 154.2674 | 242.23 | 166.877 | 185.8075 | 197.2567 | 208.1818 | 274.1444 | 156.3904 | 184.1337 | 166.8289 | 199.0856 | 152.5722 | 133.7326 | 134.5615 | 138.4492 | 149.0481 | 150.0481 | 165.7326 | 118.8396 | 116.4278 | 120.1818 | 114.7807 | 116.647 | 273.4759 | 282.4492 | 318.1497 | 338.3155 | 245.8396 | 256.3422 | 319.6096 | 305.2567 | 112 | 75.647 | 88.0856 | 112 | 111 | 103.0481 | 120 | 114.3797 | 115.5615 | 104.0481 | 97.6578 | 97 | 82.1818 | 75.6096 | 77.353 | 83.8663 | 104.7807 | 100.5241 | 106.5989 | 91.5989 | 261.1711 | 211.7433 | 245.5615 | 259.7219 | 161.123 | 174.7915 | 210.3797 | 225.5882 | 334.0374 | 247.3048 | 262.246 | 361.016
2418 | 28.5924 | 162.1906 | 194.1224 | 210.8028 | 151.1972 | 148.8299 | 179.1701 | 169.1495 | 154.2449 | 153.5374 | 116.6532 | 132.5103 | 137.8776 | 208.7346 | 156.9047 | 145.728 | 193.3739 | 176.8505 | 142.7823 | 136.0205 | 165.4832 | 150.8028 | 311.5374 | 341.1495 | 135.8776 | 229.4421 | 205.8505 | 261.9252 | 251.0477 | 122.7075 | 107.9729 | 129.3673 | 107.5579 | 150.8981 | 142.0477 | 147.2449 | 93.5374 | 245.8233 | 242.6327 | 275.2925 | 241.3196 | 146.0271 | 138.4897 | 146.2925 | 99.4626 | 83.7823 | 69.8028 | 102.1701 | 101.7551 | 197.0271 | 202.8776 | 248.1224 | 151.6122 | 143.9318 | 140.6598 | 133.0953 | 130.8299 | 165.6327 | 167.7346 | 214.9729 | 144.6056 | 136.0477 | 130.9523 | 181.728 | 120.1972 | 157.415 | 183.3402 | 227.2449 | 193.5308 | 155.3878 | 152.8299 | 260.8841 | 199.4963 | 188.0953 | 209.1972 | 205.8299 | 272.7823 | 173.9795 | 192.6122 | 158.6327 | 202.1224 | 160.7346 | 127.3878 | 140.5374 | 125.4626 | 160.1495 | 161.1495 | 177.7823 | 119.1972 | 115.3402 | 116.415 | 120.9523 | 116.4626 | 284.9047 | 285.0271 | 344.6122 | 345.1495 | 266.0066 | 257.8776 | 299.8981 | 307.2925 | 104.9252 | 77.585 | 80.5103 | 104.9252 | 117.3673 | 103.5374 | 115.0477 | 117.0477 | 115.1701 | 101 | 89.0477 | 94.8776 | 71.3402 | 66.5103 | 80.3673 | 95.1972 | 113.7823 | 107.4897 | 108.0477 | 87.3878 | 241.0682 | 218.9047 | 226.0682 | 268.8776 | 156.0953 | 190.9318 | 198.1906 | 229.8028 | 363.4692 | 238.5103 | 288.9795 | 362.728
2419 | 28.6034 | 160.4779 | 218.0326 | 209.1593 | 167.9558 | 140.643 | 162.8849 | 173.5605 | 137.286 | 156.4395 | 127.1535 | 119.643 | 132.6814 | 188.8081 | 165.3186 | 148.357 | 204.4012 | 156.6046 | 139.5221 | 137.6372 | 162.0384 | 150.5988 | 323.0768 | 360.6756 | 159.4721 | 225.643 | 194.9616 | 235.9674 | 248.2802 | 123.7198 | 106.5988 | 128.4012 | 121.4337 | 161.714 | 152.2361 | 152.5988 | 96.4395 | 255.8733 | 231.3628 | 315.307 | 246.7198 | 133.7256 | 154.9558 | 137.3628 | 99.4395 | 83.4012 | 73.9175 | 88.7639 | 95.6814 | 204.1593 | 174.6488 | 233.1651 | 151.6814 | 144.3628 | 156.6756 | 135.4779 | 124.0825 | 181.714 | 166.5221 | 199.2035 | 148.5163 | 141.1977 | 136.5988 | 172.1209 | 139.8349 | 180.3128 | 165.1651 | 223.2419 | 185.1593 | 148.8407 | 149.6814 | 249.6872 | 167.4953 | 185.4395 | 189.9674 | 207 | 272.4012 | 171.643 | 174.6872 | 169.6756 | 199.4012 | 162.4012 | 142.4337 | 140.5605 | 139.1151 | 177.5163 | 164.8407 | 178.1209 | 108.6046 | 110.8407 | 119.8791 | 130.1977 | 112.8407 | 285.4012 | 270.5663 | 325.2476 | 350.2802 | 249.2093 | 267.0768 | 301.357 | 299.0825 | 117.1151 | 79.1593 | 73.4012 | 117.1151 | 113.5221 | 107.1593 | 120.9175 | 114.2802 | 111.1209 | 92.3628 | 104.2744 | 100.4779 | 64.6814 | 74.5163 | 82.2802 | 97.5605 | 107.643 | 102.3628 | 97.3628 | 78.6814 | 233.4395 | 240.2744 | 222.7581 | 240.6488 | 152.7198 | 188.4837 | 217.3512 | 215.2035 | 375.4395 | 227.0825 | 298.1593 | 342.3244
2420 | 28.6144 | 163.7321 | 201.3779 | 224.6412 | 155.4306 | 139.9282 | 157.7321 | 162.0191 | 135.1244 | 140.8947 | 128.4115 | 108.6794 | 138.3206 | 194.3732 | 147.2344 | 169.5694 | 217.6412 | 167.8374 | 145.0527 | 143.9282 | 169.1244 | 139.555 | 339.445 | 378.445 | 141.1817 | 213.2153 | 200.3206 | 254.8183 | 240.6794 | 126.1962 | 110.1962 | 142.3732 | 116.555 | 148.5023 | 147.4832 | 156.1962 | 94.8038 | 252.287 | 237.5168 | 333.1962 | 226.5023 | 148.9618 | 140.2344 | 145.713 | 121.2297 | 84.9282 | 63.0909 | 82.5359 | 96.9282 | 197.6794 | 183.7656 | 221.1435 | 147.8038 | 133.6794 | 155.4115 | 162.89 | 110.0191 | 180.9473 | 173.5168 | 201.5885 | 146.4115 | 132.287 | 145.3206 | 153.4306 | 122.5741 | 182.4115 | 165.5885 | 209.287 | 202.8374 | 151.6603 | 141.4115 | 229.287 | 163.6412 | 197.713 | 203.6938 | 196.0191 | 291.4977 | 168 | 146.2344 | 185.9809 | 213.3732 | 141.9665 | 152.1962 | 150.2488 | 142.0718 | 177.6076 | 166.1962 | 190.9091 | 124.9618 | 104.8756 | 121.7321 | 109.5741 | 108.3397 | 273.7512 | 263 | 319.6603 | 327.3062 | 262.158 | 251.9665 | 317.445 | 308.445 | 126.6603 | 78.5359 | 79.3206 | 126.6603 | 108.0718 | 97.0191 | 125.4641 | 119.8565 | 112.1962 | 105.8374 | 113.9282 | 105.1962 | 73.2488 | 81.9282 | 65.8947 | 97 | 115.713 | 104.1244 | 88.1435 | 90.1771 | 253.7656 | 215.5214 | 240.3732 | 262.9427 | 161.7847 | 203.7656 | 190.665 | 251.9953 | 344.5214 | 259.1389 | 277.7703 | 349.445
2421 | 28.6254 | 166.2331 | 202.4209 | 215.113 | 172.0749 | 149.9322 | 143.113 | 164.6992 | 137.7444 | 138.7218 | 123.0226 | 111.2105 | 135.7895 | 205.2105 | 128.8347 | 142.5042 | 208.113 | 173.2556 | 147.2556 | 160.6314 | 188.1201 | 136.4887 | 323.9025 | 345.0381 | 131.7444 | 247.9619 | 205.9774 | 248.6243 | 258.0975 | 105.4138 | 126.6314 | 126.4138 | 110.5113 | 143.9774 | 146.2331 | 141.3686 | 88.7895 | 236.0904 | 252.1653 | 299.0155 | 225.6992 | 151.7895 | 138.2105 | 135.8573 | 122.3008 | 89.7218 | 74.3757 | 74.5565 | 125.541 | 182.346 | 183.5565 | 221.2331 | 128.3912 | 122.8121 | 145.5565 | 164.5565 | 101.5339 | 182.4661 | 161.3686 | 217.3983 | 151.4435 | 128 | 144.2782 | 173.7966 | 134.8418 | 183.7218 | 166.5113 | 245.1949 | 209.7444 | 143.3234 | 145.6992 | 222.7669 | 164.5339 | 182.6469 | 214.9774 | 180.0904 | 292.3008 | 158.3234 | 171.7514 | 172.8799 | 223.4661 | 132.7669 | 147.7895 | 156.9774 | 138.0226 | 164.8347 | 167 | 179.6243 | 119.6017 | 129.7966 | 120.5113 | 135.2401 | 119.654 | 282.654 | 280.1201 | 339.6088 | 327.9322 | 272.4887 | 265.0975 | 330.1879 | 318.2105 | 117.5791 | 74.2782 | 85.7218 | 117.5791 | 117.4209 | 107.887 | 126 | 115.3008 | 122.6766 | 101.5791 | 117.9774 | 114.1879 | 71.7895 | 71.8347 | 71.1653 | 87.3234 | 106.6017 | 106.7444 | 94.9322 | 81.6017 | 220.0607 | 236.7514 | 228.1356 | 273.5113 | 145.6469 | 222.1653 | 195.6088 | 236.4816 | 344.9096 | 239.2486 | 260.3234 | 322.7373
2422 | 28.6364 | 185.916 | 224.916 | 219.8363 | 148.2211 | 154.5133 | 149.3496 | 174.5664 | 134.2168 | 145.2965 | 115.9469 | 125.1063 | 148.3762 | 202.4602 | 143.1594 | 149.1594 | 212.8363 | 169.2168 | 135.6504 | 154.6504 | 155.4113 | 158.9426 | 317 | 325.9469 | 146.3762 | 217.8714 | 213.8098 | 239.9734 | 265.7566 | 108.593 | 119.8937 | 131.8629 | 115.2965 | 154.8363 | 137.9203 | 151.1328 | 95.3231 | 254.6992 | 235.5707 | 285.4867 | 230.2699 | 155.2965 | 123.3539 | 149.916 | 107.8937 | 94.7832 | 70.9203 | 77.2965 | 109.2742 | 187.8363 | 190.8363 | 225.0266 | 141.6727 | 120 | 131.6504 | 146.1105 | 119.6727 | 171.8937 | 155.2434 | 195.5175 | 155.5133 | 142.3762 | 146.0266 | 178.4602 | 123.8406 | 163.814 | 170.5398 | 231.0043 | 211.5133 | 157.4028 | 149.5133 | 232.593 | 169.8098 | 183.5664 | 217.5133 | 216.102 | 307.4028 | 163.3231 | 162.5707 | 175.3231 | 205.3273 | 145.6195 | 137.6769 | 150.4336 | 152.1328 | 168.5664 | 157.1637 | 173.2434 | 136.186 | 126.8937 | 113.1902 | 131.8672 | 127.7832 | 265.814 | 250.4378 | 316.491 | 343.1063 | 279.0531 | 278.0531 | 333.7566 | 303.3539 | 117.7832 | 73 | 87.7566 | 117.7832 | 111.1637 | 108.4602 | 124.4867 | 120.5664 | 122.9734 | 107.8363 | 130.3496 | 111.7035 | 78.3231 | 69.5133 | 81.8098 | 74.1637 | 105.0266 | 110.7832 | 97.2434 | 89.1063 | 259.7216 | 205.6281 | 223.5133 | 275.2699 | 149.593 | 189.6812 | 192.1637 | 236.593 | 348.2434 | 233.2965 | 247.9203 | 339.9957
2423 | 28.6474 | 168.1631 | 224.0796 | 219.9243 | 161.8369 | 150.3864 | 171.4544 | 181.6136 | 126.8485 | 141.6175 | 110.9243 | 137.4583 | 153 | 193.3107 | 148.2311 | 150.3864 | 194.4699 | 157.2349 | 145.8408 | 160.2272 | 165.299 | 161.3864 | 341.6058 | 308.6213 | 153.3068 | 215.534 | 220.6136 | 220.5456 | 259.0796 | 122.7651 | 110.6175 | 132.1553 | 139.299 | 146.466 | 154.9922 | 152.1553 | 100.3068 | 238.1631 | 261.2952 | 285.7689 | 245.6097 | 135.4699 | 119.5379 | 149.8485 | 115.534 | 80.6213 | 73.3825 | 84.3825 | 110.9961 | 192.5379 | 187.0796 | 251.3748 | 154.1515 | 122.3068 | 129.5379 | 131.7728 | 110.6213 | 169.5379 | 176.5301 | 188.5379 | 144.466 | 141.6175 | 135.466 | 171.6175 | 134.1476 | 160.8447 | 174.3068 | 233.534 | 212 | 157.6175 | 153.8447 | 212.1631 | 177.3825 | 186.7689 | 195.701 | 224.3864 | 259.9437 | 160.6175 | 148.3107 | 195.6854 | 205.1515 | 134.6213 | 150.3787 | 148 | 153.1553 | 157.9281 | 168.6097 | 176.8447 | 126.8524 | 106.0835 | 107.1553 | 133.1515 | 109.0078 | 295.9088 | 257.4544 | 302.3864 | 340.0796 | 261.7767 | 247.7048 | 336.3068 | 281.0835 | 121.3068 | 74.5379 | 85.6932 | 121.3068 | 105.6932 | 122.3787 | 118.6175 | 118.3864 | 112.7728 | 114.0757 | 121.6971 | 89.2388 | 77.1553 | 81.534 | 84.7689 | 87.9165 | 95.2349 | 113.5379 | 100.0757 | 80.6971 | 246.2427 | 199.6893 | 224.7689 | 269.8485 | 148.3864 | 190.3029 | 207.4544 | 239.2311 | 374.1437 | 228.0796 | 270.3748 | 351.3068
2424 | 28.6584 | 172.7184 | 197.0008 | 215.0939 | 170.5624 | 149.7812 | 180.1249 | 170.5004 | 145.3118 | 150.9371 | 120.1559 | 125.1568 | 140.5004 | 204.2808 | 141.7502 | 170.8743 | 180.969 | 171.9681 | 140.6253 | 134.0947 | 156.3755 | 172.4996 | 340.4065 | 329.7804 | 158.6873 | 228.3747 | 230.5935 | 256.4049 | 257.7812 | 128.3437 | 127.7494 | 118.5004 | 135.0629 | 162.5306 | 146.938 | 145.5314 | 83.0319 | 237.2498 | 263.1878 | 269.5943 | 232.0319 | 133.6873 | 128.5935 | 137.0629 | 135.4057 | 85.3747 | 82.8122 | 93.031 | 107.7502 | 192.2188 | 195.9371 | 238.6882 | 144.2816 | 129.2498 | 140.1559 | 132.9061 | 113.031 | 187.1869 | 164.2506 | 181.1878 | 128.5004 | 128.2816 | 135.9061 | 173.9061 | 134.3127 | 151.8441 | 153.1257 | 219.8131 | 215.9061 | 153.6563 | 147.969 | 244.8424 | 179.7812 | 201.062 | 189.7812 | 230.031 | 280.7175 | 157.4376 | 146.7812 | 161.9388 | 207 | 155.7804 | 145.6253 | 145.6563 | 149.6563 | 158.6873 | 156.5943 | 199.093 | 132.1559 | 115.062 | 116.1559 | 132.6563 | 100.6563 | 289.8131 | 267.6873 | 289.2816 | 352.062 | 267.7184 | 254.4057 | 349.4996 | 318.1861 | 121.2188 | 87.4996 | 88.1249 | 121.2188 | 124.5306 | 112.938 | 117.7812 | 138.8743 | 106.0939 | 117.3437 | 110.1878 | 95.4996 | 74.4376 | 75.6253 | 75.6253 | 102.3747 | 86.5314 | 92.1257 | 108.031 | 83.2498 | 244.031 | 219.9681 | 232.031 | 247.6882 | 150.9061 | 195.5624 | 216.1249 | 246.031 | 353.8759 | 225.2188 | 238.1576 | 351.2188
2425 | 28.6694 | 163.3038 | 205.8703 | 217.9676 | 175.7611 | 146.8259 | 158.7816 | 170.1741 | 139.8908 | 131.7816 | 117.4454 | 121.7935 | 128.2713 | 207.2065 | 160.6314 | 157.1621 | 179 | 164.3038 | 137.2065 | 138.6962 | 182.1536 | 162.5102 | 318.9556 | 326.6843 | 157.6195 | 236.5546 | 202.8464 | 244.9881 | 223.8789 | 132.9676 | 126.6519 | 125.3157 | 132.7935 | 163.2389 | 143 | 142.413 | 89.9027 | 245.3481 | 237.1946 | 288.8054 | 230.9676 | 137.3805 | 143.6962 | 156.2184 | 139.2065 | 79.2713 | 83.413 | 94.2065 | 101.2389 | 201.5222 | 199 | 225.8584 | 160.8379 | 135.7611 | 120.7816 | 122.0973 | 115 | 172.9556 | 156.6195 | 202.0119 | 128.1741 | 125 | 166.36 | 158.3362 | 125.0649 | 175.186 | 187.4693 | 199.9232 | 217 | 164.1092 | 160.2833 | 245.6843 | 185.5546 | 186.7492 | 207.4573 | 234.3805 | 280.6843 | 149.8584 | 163.6638 | 165.2833 | 197.4778 | 150.3038 | 135.0649 | 157.6962 | 149.7935 | 162.3805 | 156.7611 | 188.3362 | 122.3038 | 115.0324 | 111.8584 | 144.6962 | 123.8054 | 279.4454 | 261.8584 | 315.36 | 361.5546 | 267.8259 | 243.1297 | 279.9966 | 317.3038 | 118.6195 | 74.3362 | 85.0324 | 118.6195 | 113.3362 | 114.5546 | 110.0649 | 131.5102 | 108.9676 | 110.0649 | 122.2833 | 101.3805 | 79.5546 | 65.8584 | 84.9027 | 81.9881 | 102.4573 | 109.0119 | 97.3038 | 73.0973 | 254.7287 | 220.2389 | 253.8379 | 281.6757 | 154.3805 | 197.587 | 212.2389 | 247.2065 | 342.8259 | 234.5222 | 256.36 | 352.587
2426 | 28.6804 | 159.1942 | 199.5245 | 216.5826 | 165.7187 | 134.7187 | 161.8639 | 161.3303 | 153.1161 | 137.2813 | 112.7768 | 122.8058 | 127.6116 | 207.8058 | 160.3594 | 147.971 | 169.3303 | 163.4174 | 128.1361 | 150.8639 | 180.3303 | 159.8058 | 345.4264 | 323.1942 | 155.3884 | 218.6607 | 204.6697 | 224.4955 | 282.6877 | 128.3594 | 104.0491 | 131.2232 | 130.5826 | 149.9129 | 137.3594 | 146.029 | 80.1071 | 269.5626 | 244.6987 | 269.2142 | 240.8639 | 147.6697 | 113.1562 | 133.7968 | 135.7768 | 85.0581 | 78.971 | 89.1652 | 104.029 | 190.3013 | 199.8058 | 240.9219 | 161.1652 | 133.7768 | 112.5826 | 135.1161 | 116.6116 | 160.7477 | 174.5335 | 219.2813 | 148.3393 | 114.5245 | 167.5536 | 160.4464 | 114.1361 | 179.5826 | 183.4955 | 212.1161 | 221.029 | 151.6897 | 162.3884 | 232.5245 | 178.9419 | 206.98 | 201.5245 | 214.8549 | 318.2903 | 168.9509 | 151.8839 | 192.3684 | 201.4464 | 146.1942 | 142.6697 | 146.4955 | 136.3013 | 154.9419 | 146.7187 | 181.5826 | 113.3594 | 115.6116 | 123.6987 | 127.8549 | 119.5245 | 278.8058 | 275.3103 | 293.991 | 325.1271 | 265.3884 | 254.3103 | 292.4264 | 290.6316 | 138.9509 | 74.029 | 85.6116 | 138.9509 | 109 | 110.3594 | 113.6406 | 108.6607 | 108.3884 | 115.2523 | 113.9129 | 91.5245 | 74.5536 | 79.3103 | 78.3303 | 97.7568 | 109.4174 | 132.7277 | 102.0581 | 93.3684 | 270.6987 | 194.8258 | 260.6116 | 272.6607 | 164.6697 | 214.1161 | 186.02 | 231.6897 | 333.1361 | 241.029 | 273.6697 | 342.5245
2427 | 28.6914 | 161.4543 | 204.3629 | 218.4543 | 168.7267 | 152.4524 | 137.0028 | 167.181 | 141.4562 | 136.7276 | 122.6353 | 118.0914 | 150.0886 | 201.4552 | 150.0009 | 140.4552 | 180.0895 | 165.6362 | 149.7248 | 142.3647 | 172.2733 | 160 | 305.5504 | 342.6343 | 160.7267 | 231.9981 | 236.4515 | 212.819 | 257.2771 | 117.1828 | 113.7257 | 132 | 134.0905 | 140.4552 | 107.3666 | 138.0009 | 94.18 | 248.8209 | 241.4552 | 285.0886 | 237.2733 | 142.6371 | 118.9076 | 139.2715 | 134.1819 | 86.1819 | 87.8172 | 87.1819 | 107.4543 | 191.9086 | 199.1819 | 240.9095 | 151.819 | 130.5457 | 122.6353 | 147.9991 | 99.82 | 159 | 170.819 | 219.5457 | 146.4552 | 132.4524 | 132.458 | 149.7285 | 134.0886 | 169.1828 | 182.4543 | 229.9076 | 225.2724 | 169.2705 | 182.4524 | 226.7276 | 195.8162 | 185.1847 | 222.7248 | 214.0905 | 294.458 | 143.7304 | 169.2705 | 193.0914 | 200.5457 | 141.0914 | 143.3638 | 140.5457 | 128.0914 | 154.6362 | 170.9972 | 186.7267 | 128.3619 | 120.9086 | 129.4543 | 123 | 119.4543 | 279 | 266.7285 | 299.2715 | 338.9067 | 274.8172 | 267.8172 | 313.0895 | 311.1791 | 136.6371 | 78.2724 | 78.6371 | 136.6371 | 103.2733 | 113.0905 | 107.6371 | 124.4524 | 112.0905 | 122.7267 | 115.0905 | 110.2705 | 72.1819 | 67.4562 | 84.181 | 103.8181 | 110.8181 | 96.0952 | 108.9086 | 86.7285 | 236.3675 | 230.7229 | 262.6362 | 272.0905 | 155.5466 | 177.9133 | 195.5438 | 210.82 | 337.5448 | 216.639 | 247.3666 | 321.1838
2428 | 28.7024 | 175.2862 | 206.8304 | 212.3569 | 142.5972 | 142.8834 | 146.7774 | 196.4028 | 137.1696 | 136.8304 | 128.3216 | 124.4735 | 156.6608 | 196.6784 | 151.3216 | 176.3675 | 188.8127 | 180.1166 | 140.8834 | 152.4558 | 160.205 | 174.947 | 308.2862 | 344.5088 | 135.4276 | 228.5265 | 218.0884 | 230.9293 | 272.0812 | 119.9823 | 117.8304 | 123.6961 | 135.8304 | 143.9823 | 121.7597 | 126.0353 | 83.053 | 264.5901 | 238.3392 | 285.0177 | 226.0353 | 143.4912 | 118.6784 | 146.9823 | 124.0353 | 79.3569 | 82.5265 | 96.1343 | 103.8481 | 183.0353 | 212.2862 | 215.9188 | 129.2403 | 115.053 | 118.3569 | 136.7138 | 107.6254 | 173.947 | 162.3569 | 199.9011 | 163.2685 | 132.0177 | 129.1519 | 161.1166 | 146.4735 | 176.9647 | 187.9823 | 184.0071 | 228.4912 | 149.0884 | 193.6431 | 245.0989 | 165.9541 | 203.0812 | 175.6856 | 195.9011 | 302.7774 | 176.0282 | 169.0177 | 196.9823 | 217.4381 | 163.2509 | 142.1696 | 126.7138 | 133.6431 | 155.8304 | 167.0353 | 178.0353 | 117.053 | 107.8834 | 103.4276 | 123 | 120.8304 | 281.4912 | 280.6078 | 307.8127 | 338.1873 | 249.5972 | 263.3569 | 279.4629 | 293.7491 | 125.0353 | 80.6608 | 78.6608 | 125.0353 | 115.2862 | 99.053 | 124.2685 | 116.5442 | 99.7138 | 108.2226 | 113.5088 | 121.6431 | 80.3039 | 83.9293 | 68.5619 | 81.5795 | 118.4735 | 106.0989 | 112.4912 | 86.4912 | 230.4912 | 186.8552 | 230.6149 | 263.0353 | 149.6784 | 193.9116 | 216.4381 | 241.0459 | 336.5088 | 215.9823 | 257.6078 | 348.5547
2429 | 28.7134 | 162.8318 | 196.0452 | 212.6854 | 152.1682 | 155.1682 | 142.4159 | 176.7196 | 134.472 | 159.7523 | 118.0452 | 122.6293 | 142.6745 | 210.3255 | 139.3598 | 150.2928 | 188.3146 | 182.1573 | 159.067 | 148.2586 | 171.4829 | 169.5732 | 303.4159 | 322.933 | 165.3925 | 243.0109 | 223.9548 | 252.6963 | 262.6745 | 112.5732 | 112.9439 | 117.7866 | 128.4159 | 172.8084 | 125.1573 | 156.0218 | 88.4268 | 297.6511 | 242.2134 | 271.3598 | 253.4938 | 145.6854 | 140.7523 | 141.2586 | 106.8318 | 97.3816 | 91.1121 | 81.9891 | 115.6402 | 203.7523 | 196.4611 | 230.3816 | 134.2695 | 115.3707 | 123.7414 | 147.4829 | 120.9548 | 163.5171 | 162.6854 | 200.2134 | 146.7757 | 152.9097 | 146.0109 | 166.528 | 126.933 | 162.9891 | 170.4611 | 208.5498 | 202.0343 | 144.8427 | 172.2477 | 235.5171 | 177.5389 | 185.2477 | 198.7072 | 219.8084 | 273.9782 | 183.1573 | 167.1573 | 196.3146 | 204.1464 | 155.3598 | 147.8988 | 144.2243 | 126.5732 | 160.2134 | 171.7414 | 185.2695 | 134.2243 | 135.3364 | 121.595 | 134.7975 | 131.1121 | 278.6293 | 305.9097 | 284.5623 | 312.5623 | 248.2134 | 270.4268 | 307.3925 | 307.5389 | 125.528 | 76.7866 | 81.528 | 125.528 | 134.0109 | 100.2134 | 120.4159 | 109.7866 | 107.9548 | 83.0903 | 131.5389 | 109.5171 | 78.6293 | 78.7305 | 77.6402 | 102.2804 | 114.1012 | 99.8879 | 96.1464 | 67.6184 | 237.7414 | 223.19 | 219.7866 | 250.8879 | 164.1682 | 182.1464 | 225.0561 | 234.5171 | 340.2134 | 242.2804 | 267.7414 | 356.6854
2430 | 28.7244 | 174.5345 | 199.9848 | 205.3053 | 154.145 | 143.4655 | 150.4047 | 168.5801 | 137.4199 | 147.7556 | 117.7099 | 131.4047 | 140 | 207.0152 | 139.5649 | 166.2292 | 207.6643 | 174.3053 | 138.2059 | 134.1755 | 145.786 | 162.8702 | 318.2444 | 312.1602 | 166.0152 | 237.4503 | 183.2515 | 247.4503 | 281.3742 | 116.1298 | 109.4351 | 111.8702 | 108.1907 | 138.6714 | 131.8398 | 146.6105 | 102.8245 | 243.1521 | 246.4199 | 252.7556 | 227.3661 | 150.2748 | 139.0152 | 143.4199 | 129.6491 | 103.5649 | 93 | 75.5801 | 117.145 | 201.1602 | 204.9696 | 226.3053 | 144.5497 | 134.8093 | 145.5193 | 134.6105 | 126.4199 | 138.7708 | 151.8854 | 216.3895 | 166.9392 | 137.3357 | 138.7404 | 167.855 | 138.3895 | 155.7252 | 182.3895 | 209.8702 | 209.8245 | 141.5801 | 168 | 235.5649 | 207.5041 | 196.3895 | 198.1602 | 214.7404 | 278.2596 | 175.3053 | 147.3357 | 201.1298 | 201.855 | 153.855 | 143.8702 | 140.3053 | 120.7252 | 156.7252 | 157.6105 | 175.8854 | 119.1907 | 101.6714 | 106.3357 | 132.7252 | 119.3205 | 287.4047 | 249.2971 | 280 | 312.2748 | 262.6795 | 248.9158 | 300.3205 | 311 | 113.1755 | 81.1298 | 87.9848 | 113.1755 | 105.3661 | 89.8854 | 126.6947 | 113.2748 | 115.1298 | 100.3742 | 116.1907 | 119.8245 | 76.2901 | 77 | 74.0152 | 95.0304 | 107.8702 | 92.0152 | 109.2444 | 84.5193 | 249.2596 | 210.6258 | 228.4047 | 268.6643 | 156.7404 | 179 | 190.0913 | 234.5649 | 328.1755 | 235.0304 | 235.6562 | 355.2901
2431 | 28.7354 | 141.4423 | 221.8143 | 198.7964 | 160.9381 | 171.3541 | 144.1947 | 175.8053 | 131.9292 | 152.8053 | 116.2655 | 120.8583 | 148.6726 | 204.2655 | 134.7964 | 142.2476 | 217.9381 | 187.7434 | 132.2655 | 147.6107 | 163.5488 | 146.3893 | 313.1947 | 315.3363 | 136.3804 | 222.9911 | 217.6285 | 228.6548 | 245.9733 | 118.7345 | 102.0619 | 134.416 | 118.0089 | 129.3982 | 130.3982 | 139.6637 | 104.1327 | 243.4071 | 261.7434 | 255.2036 | 240.2125 | 168.3452 | 138 | 155.2744 | 108.8494 | 89.2566 | 83.4601 | 83.6726 | 97.053 | 183.522 | 190.522 | 242.3452 | 136.4601 | 125.8583 | 144.6637 | 122.4601 | 122.6637 | 161.0178 | 176.8851 | 225.0708 | 130.2387 | 135.7345 | 128.3274 | 173.2036 | 147.9381 | 164.5399 | 178.0619 | 220.2744 | 198.1238 | 169.6196 | 165.3982 | 263.7524 | 173.8405 | 191.1947 | 187.4601 | 229.478 | 303.416 | 147.1149 | 172.6196 | 199.3982 | 196.7964 | 152.2655 | 123.9203 | 132.0619 | 149.4869 | 148.1947 | 161.9381 | 165.3274 | 106.4601 | 126.2214 | 123.8143 | 100.7786 | 120.469 | 294.2036 | 257.2125 | 285.2036 | 335.5488 | 276.2744 | 241.531 | 310.1417 | 257.2298 | 114.469 | 76.7964 | 72.522 | 114.469 | 97.3982 | 108.8143 | 114.1238 | 111.3982 | 123.8053 | 83.1857 | 106.0619 | 102.9203 | 63.8583 | 87.4071 | 73.8673 | 103.4071 | 104.3982 | 104.8762 | 129.3452 | 81.9292 | 234.522 | 198.3274 | 216.9911 | 271.1327 | 155.8673 | 183.3363 | 184 | 232.3982 | 341.6107 | 239.9381 | 255.1506 | 331.584
2432 | 28.7464 | 162.3866 | 196.8543 | 210.3137 | 146.1681 | 139.0588 | 161.4706 | 162.0476 | 145.0728 | 136.409 | 106.3249 | 118.1204 | 120.0952 | 211.916 | 133.1204 | 158.2297 | 221.6387 | 172.409 | 144.3137 | 153.5182 | 159.084 | 159.8319 | 320.7955 | 331.8319 | 146.0728 | 233.3137 | 202.8907 | 231.2773 | 247.0364 | 127.7955 | 115.9524 | 118.6499 | 121.7591 | 142.1933 | 126.4818 | 144.2773 | 82.8907 | 207.1793 | 265.7591 | 256.8796 | 270.2661 | 148.1316 | 146.7955 | 134.1316 | 101.4818 | 81.7227 | 78.4818 | 70.0476 | 101.916 | 198.591 | 193.2773 | 215.9748 | 160.507 | 132.7955 | 146.6387 | 130.6751 | 127.2773 | 165 | 181.8796 | 197.8543 | 142.4706 | 144.7955 | 134.916 | 169.6022 | 137.5658 | 177.4342 | 204.2661 | 221.1204 | 200.3978 | 164.3249 | 152.6863 | 242.493 | 182.0728 | 181.2045 | 186.8796 | 190.6611 | 263.9019 | 171.1457 | 174.3613 | 166.4566 | 210.9524 | 156.3978 | 148.2661 | 135.3978 | 149.6022 | 171.6275 | 140.1316 | 167.5182 | 121.7115 | 107.2521 | 119.084 | 128.5434 | 128.0364 | 251.0224 | 245.0476 | 272.8067 | 343.3978 | 275.3613 | 241 | 312.8796 | 302.6527 | 107.9636 | 87.4342 | 73.5182 | 107.9636 | 111.0728 | 114.6387 | 122.5546 | 125.0728 | 111.8067 | 96.7115 | 107.6387 | 100 | 79.591 | 87.2409 | 77.5182 | 89.1681 | 97.8431 | 92.0476 | 117.0476 | 79.2409 | 220.5658 | 197 | 222.916 | 269.2409 | 175.3501 | 176.9636 | 210.3866 | 246.0728 | 329.9272 | 231.3249 | 258.1204 | 368.4594
2433 | 28.7574 | 160.6489 | 205.4858 | 229.8369 | 172.539 | 138.4592 | 155.0816 | 163.5674 | 125.5958 | 142.0266 | 118.3776 | 121.5674 | 149.89 | 197.8387 | 146.3776 | 153.8653 | 222 | 144.1365 | 131.7305 | 145.0816 | 164.2429 | 163.7837 | 285.4344 | 327.7571 | 150.6755 | 230.5408 | 212.4858 | 217.7305 | 232.8387 | 108.4876 | 127.8103 | 121.3511 | 116.6489 | 141.3245 | 142.945 | 155.7021 | 93.3776 | 230.539 | 262.4326 | 243.6224 | 241.8936 | 145 | 137.2979 | 158.6471 | 120.6205 | 89.9184 | 70.8653 | 63.5408 | 98.5408 | 194.7571 | 207.3776 | 208.4326 | 150.6224 | 135.7837 | 151.4592 | 133.7837 | 119.0816 | 154.2979 | 171.2979 | 217.1879 | 153.9184 | 138.8653 | 124.406 | 196.6471 | 140.4592 | 172.7571 | 188.3795 | 200.4876 | 200.1082 | 146.055 | 147.4326 | 244.3511 | 201.8369 | 173.7571 | 156.6773 | 195.7021 | 282.9716 | 199.0798 | 137.4344 | 173.594 | 223.7021 | 157 | 146.6489 | 133.3245 | 137.406 | 174.1082 | 145.9184 | 188.5124 | 103.4876 | 107.5674 | 135.8369 | 150.8369 | 114.7305 | 279.7819 | 274.2145 | 287.0532 | 302.9752 | 251.8121 | 259.7287 | 290.7039 | 284.1365 | 108.7837 | 83.6489 | 84.7021 | 108.7837 | 94.2713 | 100.7305 | 108.8387 | 112.7305 | 105.5408 | 95.4326 | 108.8918 | 112.4858 | 61.4876 | 86.1082 | 73.5408 | 99.4858 | 121.9716 | 101.594 | 105.1897 | 84.3511 | 226.1347 | 206.8103 | 232.9184 | 264.5408 | 169.0816 | 180.4592 | 192.5958 | 215.8936 | 354.7553 | 216.6224 | 248.1897 | 341.8936
2434 | 28.7684 | 160 | 179.876 | 213.9173 | 154.3008 | 164.3158 | 152.1917 | 164 | 129.3289 | 144.8083 | 106.438 | 140.0827 | 150.3835 | 205.0413 | 141.6711 | 151.1917 | 241.891 | 157.2744 | 145.3703 | 132.2463 | 158.6711 | 144.109 | 293.6579 | 359.5488 | 152.8083 | 208.3008 | 197.7256 | 266.6315 | 256.3158 | 115.0413 | 126.2876 | 109.3421 | 124.1372 | 133.7669 | 124.2049 | 146.1504 | 88.6711 | 257.5075 | 249.3421 | 281.7819 | 217.2049 | 145 | 142.3289 | 136.6842 | 118.4793 | 89.1917 | 80.8496 | 68.4248 | 106.1372 | 204.8496 | 202.6711 | 218.8496 | 160.7537 | 144.1372 | 111.314 | 121.3421 | 121.6165 | 167.4661 | 166.3835 | 196.4925 | 145.0545 | 123.5339 | 136.562 | 163.8346 | 130.1504 | 194.6034 | 176.0545 | 222.4116 | 213.562 | 150.3289 | 145.1917 | 233.2463 | 165.1222 | 188.3703 | 186.453 | 189.7669 | 314.9323 | 169.4512 | 179.1108 | 170.4793 | 189.7388 | 152.4793 | 134.2463 | 132.0959 | 132.3835 | 167.6711 | 150.6165 | 167.4925 | 112.7537 | 101.6711 | 111.7801 | 125.876 | 111.1917 | 269.5339 | 282.5207 | 298.0413 | 364.9059 | 249.9041 | 276.4661 | 309.6992 | 328.9191 | 111.7124 | 64.0132 | 76.9587 | 111.7124 | 99.2331 | 115.2744 | 103.3835 | 122.7537 | 120.3703 | 92.2876 | 97.2463 | 104.9587 | 84.3158 | 67.0132 | 69.3835 | 103.7124 | 96.9718 | 95.7669 | 91.3421 | 81.3835 | 241.4661 | 234.2199 | 229.4793 | 235.0677 | 169.8083 | 180.0959 | 207.1785 | 227.3703 | 338.109 | 246.6447 | 221.6842 | 309.0677
2435 | 28.7794 | 160.9164 | 203.5763 | 231.2449 | 152.9164 | 126.6773 | 161.1643 | 145.6715 | 145.5792 | 144.0836 | 129.7435 | 119.0894 | 151.8329 | 220.6628 | 161.1614 | 138.17 | 214.6744 | 159.9164 | 129.5879 | 151.1614 | 151.585 | 145.6657 | 303.2478 | 309.8473 | 157.5821 | 191.3372 | 194.1671 | 218.8473 | 263.5821 | 114.1671 | 121.4179 | 129.9942 | 121.3343 | 144.9135 | 123.8329 | 133.0865 | 88 | 244.4208 | 247.0836 | 251.1758 | 208.585 | 130.3372 | 140.2507 | 133.0836 | 104.2536 | 89.9164 | 87.4986 | 80.9135 | 100.585 | 195.0029 | 175.4237 | 248.4092 | 150.0865 | 130.3372 | 117.0807 | 125.4986 | 118.3343 | 179.9971 | 156.8357 | 196.7493 | 150.415 | 143.0778 | 140.7493 | 169.1643 | 145.4957 | 168.5908 | 184.1643 | 209.4208 | 189.3401 | 144.585 | 149.5821 | 228.3343 | 178.4121 | 193.6657 | 209.2449 | 199.9971 | 318.9164 | 181.5792 | 168.4208 | 175.4986 | 203.4121 | 148.3343 | 125.6686 | 142.0807 | 141.1643 | 156.0029 | 167.4957 | 164.0836 | 124.9971 | 113.83 | 138.3256 | 130.3314 | 116.4986 | 262.5014 | 276.585 | 318.2449 | 313.3488 | 249.0836 | 263.3372 | 314.7493 | 301.9251 | 117.4986 | 72.0807 | 78.7493 | 117.4986 | 116.4957 | 110.585 | 136.9077 | 124 | 118.3343 | 115.8271 | 88.6686 | 111.3314 | 66.8386 | 60.4179 | 84.5792 | 76.5072 | 115.9942 | 107.83 | 113.8271 | 85.5821 | 246.6657 | 238.8329 | 247.3285 | 247.5792 | 161.7522 | 191.9135 | 199.8357 | 241.83 | 297.5101 | 209.6773 | 263.9048 | 321.5792
2436 | 28.7904 | 160.0713 | 180.9247 | 218.1405 | 153 | 160.1487 | 156.4277 | 148.6436 | 133.998 | 147.7149 | 128.2851 | 109.5703 | 144.5703 | 190.4236 | 149.998 | 153.7169 | 185.9959 | 162.7861 | 129.8574 | 148.3564 | 165.8595 | 154.3584 | 333.7189 | 333.7902 | 143.1405 | 233.6497 | 169.8534 | 256.721 | 241.7108 | 116.7861 | 107.0693 | 129.2139 | 138.6456 | 132.998 | 120.2851 | 146.8595 | 89.8574 | 277.3625 | 237.7128 | 284.2199 | 245.1487 | 161.5051 | 111.2098 | 138.5723 | 117.8595 | 80.7128 | 82.4277 | 91.2872 | 113.002 | 212.5743 | 176.7149 | 225.9247 | 153.6436 | 119.7128 | 126.3584 | 132.501 | 103.1405 | 151.2811 | 153.2139 | 183.0693 | 156.5723 | 145.9287 | 125.2118 | 149.5682 | 131.2118 | 158.5703 | 198.002 | 216.3584 | 185.1426 | 158.8595 | 144.4277 | 257.7189 | 174.4277 | 220.9328 | 197.0693 | 200.0713 | 285.5662 | 167.2118 | 153.998 | 169.499 | 181.7821 | 121.0672 | 124.0713 | 150.4297 | 148.501 | 154.0713 | 182.9307 | 168.6436 | 128.7861 | 96.4257 | 126.1405 | 127.2851 | 117 | 288.9328 | 250.9247 | 289.3524 | 320.0733 | 235.0693 | 239.7108 | 327.0733 | 297.1426 | 116.1426 | 64.6416 | 87.3584 | 116.1426 | 113.3564 | 112.7861 | 116.7821 | 110.0693 | 112.4277 | 104.998 | 91.7149 | 110.1426 | 85.4318 | 78.5743 | 83.2139 | 97.2179 | 108.7128 | 97.8554 | 104.8554 | 77.6416 | 224.7108 | 236.2139 | 231.3544 | 247.1426 | 151.7128 | 189.2851 | 198.0713 | 229.998 | 373.8696 | 201.3564 | 263.3564 | 337.8595
2437 | 28.8014 | 172.2331 | 202.5252 | 232.9971 | 168.0561 | 163 | 165.4101 | 169.7022 | 136.764 | 163.0561 | 97.8878 | 117.4691 | 127.0619 | 199.2921 | 184.7583 | 165.3511 | 175.5309 | 171.4691 | 126.236 | 147.059 | 153.8259 | 172.8791 | 295.5367 | 350.1151 | 142 | 213.4748 | 215.9914 | 258.118 | 244.705 | 127.3511 | 133.2892 | 138.4101 | 129.6489 | 145.1741 | 132.2331 | 167.7612 | 91.882 | 243.3007 | 244.5281 | 289.823 | 255.5281 | 146.1209 | 125.9381 | 127.7079 | 114.295 | 77.177 | 70.7079 | 100.4691 | 113.059 | 181.0647 | 180.764 | 199.5338 | 148.354 | 122.764 | 129.823 | 125.4719 | 138.6993 | 165.9381 | 173.7022 | 196.1151 | 166.4101 | 130.9439 | 129.646 | 181.8763 | 126.236 | 194.6993 | 197.118 | 178.4187 | 209.4662 | 143.0619 | 157.1741 | 232.7108 | 155.1799 | 216.4129 | 186.5899 | 173.6518 | 276.4129 | 171.646 | 160.5281 | 172.764 | 193.1741 | 147.2302 | 115.5309 | 133.1209 | 125.4748 | 178.4662 | 164.2388 | 184.9971 | 138.4101 | 116.6432 | 140.0561 | 132.646 | 125.4691 | 259.0057 | 244.295 | 329.3453 | 341.7022 | 282.9324 | 258.7022 | 296.0057 | 307.3511 | 140.4662 | 77.1741 | 73.8849 | 140.4662 | 121.4691 | 120.5281 | 108.4129 | 108.059 | 104.4719 | 85.1799 | 112.7022 | 100.5899 | 53.1237 | 69.6489 | 79.236 | 108.4101 | 94.8259 | 90.4129 | 86.1209 | 63.8259 | 245.5842 | 210.5928 | 199.8878 | 257.3511 | 166.9971 | 167.3568 | 196.118 | 225.236 | 306.6014 | 225.4662 | 260.177 | 348.4101
2438 | 28.8124 | 152.0274 | 173.4945 | 187.2884 | 161.3736 | 159.1868 | 174.5797 | 159.5604 | 129.3736 | 147.7939 | 114.1127 | 106.5604 | 135.533 | 179.0274 | 140.2884 | 153.6071 | 189.2995 | 161.5137 | 158.4121 | 133.6538 | 172.066 | 174.9533 | 326.3654 | 332.8873 | 138.1868 | 229.1594 | 193.2609 | 228.4477 | 246.9066 | 128 | 117.8406 | 109.4477 | 127.0934 | 140.2802 | 148.2528 | 142.3076 | 99.6264 | 239.0934 | 235.467 | 284.2802 | 233.1208 | 157.3929 | 140.3462 | 155.599 | 105.4203 | 79.8599 | 67.1401 | 90.5137 | 110.1401 | 188.533 | 191.4863 | 217.066 | 149.9066 | 135.3929 | 127.1401 | 118.3269 | 109.5412 | 185.1127 | 151.1675 | 207.4863 | 134.5879 | 128.0934 | 138.5797 | 158.2609 | 123.1401 | 193.1868 | 174.1208 | 209.3654 | 186.2142 | 142 | 145.6071 | 214.7939 | 185.4588 | 214.0934 | 216.5055 | 192.9726 | 306.5055 | 141.4945 | 181.9726 | 186.3462 | 207.3462 | 154.7198 | 129.2995 | 130.0934 | 147.8325 | 169.5137 | 158.2335 | 150.728 | 121.8406 | 104.6538 | 124.7939 | 112.0274 | 124.0934 | 246.5137 | 267.8325 | 309.1208 | 344.9066 | 265.0274 | 271.4396 | 284.467 | 312.7665 | 109.5879 | 78 | 66.3269 | 109.5879 | 120.0934 | 104.7939 | 115.6264 | 87.0274 | 116.3929 | 103.066 | 96.8406 | 111.4396 | 78.6457 | 99.5055 | 79.9533 | 95.6538 | 95.9066 | 106.2061 | 107.8792 | 71.5797 | 242.2335 | 218.533 | 207.533 | 258.9533 | 145.1208 | 206.9918 | 190.2802 | 217.3736 | 327.7391 | 235.5797 | 233.3076 | 316.5879
2439 | 28.8234 | 151.9656 | 193.243 | 213.0021 | 158.1032 | 159.9656 | 145.0667 | 153.2065 | 145.415 | 126.7226 | 123.6903 | 108.8968 | 133.1032 | 224.3483 | 175.658 | 171.3462 | 199.6559 | 172.5871 | 141.6538 | 132.0344 | 191.3462 | 162.4473 | 315.4473 | 354.2086 | 152.4838 | 229.0344 | 221.9333 | 258.8645 | 257.6215 | 128 | 104.4473 | 134.0709 | 128.9312 | 148.6903 | 143.2065 | 145.8279 | 111.5871 | 258.3118 | 260.1053 | 246.342 | 269.658 | 141.585 | 147.7591 | 124.17 | 127.2086 | 69.3785 | 73.7591 | 83.2409 | 103.2409 | 206.3806 | 188.1376 | 241.1741 | 159.6559 | 111.8603 | 135.6903 | 144.0709 | 106.0688 | 139.6517 | 152.8968 | 182.8947 | 131.0688 | 150.2086 | 139.9656 | 150.2409 | 142.3118 | 160.17 | 196.1741 | 222.5871 | 206.243 | 149.7247 | 142.1032 | 246.83 | 202.4494 | 185.9979 | 172.6173 | 179.5162 | 300.2753 | 161.243 | 173.3441 | 174.4473 | 183.8603 | 144.3785 | 125.1721 | 147.3806 | 125.8259 | 149.6882 | 172.4838 | 184.7268 | 147.0709 | 131.0365 | 121.1032 | 119.6903 | 120.1376 | 274.9677 | 246.7914 | 297.3785 | 342.1032 | 267.8624 | 272.9656 | 286.8968 | 312.0344 | 115.7247 | 96.3462 | 73.7247 | 115.7247 | 113.2409 | 103.0344 | 112.1376 | 99.5183 | 89.9635 | 102.0688 | 114.3462 | 117.7935 | 74.2065 | 72.9979 | 85.7935 | 92.1032 | 111.4494 | 105.0688 | 89.6882 | 85.5183 | 260.3462 | 205.4817 | 236.9677 | 266.7247 | 144.9656 | 165.5485 | 209.3118 | 231.4838 | 392.7289 | 217.6538 | 238.7591 | 332.3806
2440 | 28.8344 | 152.9779 | 199.8673 | 214 | 147.2433 | 159.0221 | 158.6682 | 147.1327 | 147.9558 | 128.9336 | 123.0221 | 123.6682 | 141.8009 | 182.9733 | 153.5309 | 162.2212 | 186.3097 | 168.1106 | 119.4866 | 134.9336 | 195.9115 | 170.8009 | 326.7346 | 324.6857 | 163.7567 | 255.4028 | 210.2876 | 224.7963 | 265.823 | 118.2212 | 132.3585 | 115.4424 | 123.1327 | 149 | 136.1548 | 132.3097 | 81.6857 | 255.0885 | 259.0442 | 282.1595 | 243.6193 | 135.1327 | 122.5751 | 117.1327 | 128 | 79.7567 | 77.9115 | 75.177 | 101.0442 | 207.9779 | 178.2212 | 228.3097 | 156.0885 | 132.5134 | 119.376 | 134.2433 | 105.0221 | 156.5797 | 154.9558 | 198.624 | 144.6903 | 112.8627 | 142.9336 | 167.6018 | 131.2654 | 159 | 207.7567 | 206.376 | 221.6682 | 147.0664 | 148.8452 | 212.7963 | 165.8405 | 190.8673 | 177.8452 | 211.2701 | 293.1548 | 169.823 | 168.1106 | 183.7788 | 201.5797 | 136.177 | 146.5134 | 141.1548 | 128.9115 | 190.071 | 147.5751 | 145.9069 | 121.5972 | 127.1106 | 110.2433 | 126.8452 | 120.9779 | 275.0221 | 258.7124 | 279.3982 | 309.7299 | 263.1106 | 245.6193 | 296.7788 | 295.376 | 128.7124 | 76.4645 | 82.8009 | 128.7124 | 113.9779 | 112.7788 | 125.6903 | 119.5576 | 115.4028 | 105.9115 | 113.0442 | 86.7078 | 75.9558 | 76.8894 | 57.6415 | 97.8673 | 81.6857 | 102.0664 | 130.071 | 81.1106 | 226.7742 | 224.5576 | 243.8673 | 262.1106 | 181.1816 | 205.071 | 187.5087 | 235.9115 | 328.5041 | 233.624 | 240.9558 | 305.6193
2441 | 28.8454 | 159.9312 | 193.0688 | 221.9214 | 171.7543 | 134.2457 | 149.0983 | 160.8624 | 135.1278 | 135.9312 | 114.0885 | 103.2064 | 123.1867 | 201.8034 | 145.0786 | 151.1081 | 194.9115 | 170.9705 | 146.7248 | 147.8722 | 159.3637 | 150.2064 | 294.3243 | 318.059 | 144.1966 | 256.9902 | 214.9509 | 223.0098 | 232.3342 | 103.1474 | 100.3243 | 121.9312 | 119.0393 | 145.0393 | 130.059 | 142.8919 | 79.0197 | 285.6953 | 260.9803 | 283.9902 | 248.941 | 134.0098 | 152.6953 | 149.6757 | 112.1573 | 69.1081 | 80.9705 | 67.0786 | 101.9902 | 215.9214 | 194.8329 | 259.6855 | 175.8034 | 130.0295 | 123.9509 | 130.0393 | 104.0098 | 152.0491 | 150.0491 | 204.941 | 145.9902 | 135.7641 | 131.1179 | 178.8919 | 134.9607 | 168.9017 | 154.5307 | 201.0491 | 204.1769 | 165.8133 | 143.059 | 262.4988 | 166.9803 | 192.9803 | 192.8526 | 197.1474 | 265.2752 | 156.1376 | 160.0786 | 172.1179 | 208.9312 | 122.1376 | 161.8526 | 141.9902 | 135.9312 | 177.1376 | 167.7936 | 162.8231 | 135.8526 | 116.1081 | 116.9312 | 129.9705 | 138.8231 | 278.9607 | 277.8133 | 265.1376 | 319.8919 | 269.9312 | 258.8624 | 268.285 | 322.7248 | 125.0393 | 63.1278 | 80.0295 | 125.0393 | 111.0295 | 103.0983 | 109.1671 | 107.1278 | 107.0885 | 118.8722 | 101.1179 | 96.8919 | 67.0885 | 86.9017 | 68.8821 | 99.9803 | 95.8526 | 117.8427 | 112.1867 | 85.9509 | 229.9607 | 197.2752 | 207.3637 | 238.2359 | 164.1769 | 184.2162 | 191.9509 | 226.0983 | 360.6658 | 223.1081 | 214.2654 | 372.3317
2442 | 28.8564 | 152.0271 | 195.9409 | 217.0049 | 156.9507 | 168.9113 | 154.9557 | 157.0517 | 145.9729 | 166.9778 | 108.0788 | 136.8916 | 143.9631 | 188.0148 | 139.9975 | 140.0493 | 182.9877 | 173.0369 | 139.0049 | 157.9261 | 155.0345 | 154.0148 | 333.8744 | 337.9655 | 138.0025 | 216.0468 | 183.1108 | 220.0246 | 242.0887 | 128.9975 | 114.9581 | 109.0345 | 122.064 | 137.0172 | 121.0148 | 143.9828 | 78.032 | 238.0246 | 239.0123 | 287.899 | 232.9261 | 146.0172 | 137.0222 | 112.0985 | 105.0025 | 86.0197 | 82.0049 | 76.0222 | 97.0739 | 180.0099 | 195.9951 | 224.9877 | 138.1157 | 155.9483 | 138.0025 | 129.9975 | 93.0123 | 175 | 160.9926 | 190.0099 | 148.9926 | 127.0566 | 164.9138 | 162.9803 | 131.0172 | 156.0788 | 161.0616 | 203.0099 | 201.0443 | 166.8793 | 144.0837 | 234.9975 | 169.0271 | 190.0739 | 201.9926 | 195.0123 | 302.8473 | 182.9754 | 165.0074 | 171.0099 | 189.9877 | 138.9926 | 118.064 | 155.9803 | 105.0837 | 173.0369 | 146.0788 | 172.9975 | 124.0443 | 126.9631 | 105.0345 | 124.9754 | 123.0123 | 285.9778 | 246.9951 | 309.9507 | 367.8375 | 273.9877 | 256 | 280.0665 | 271.1428 | 135.9089 | 71.0246 | 66.0394 | 135.9089 | 101.0443 | 107.9975 | 124.9655 | 102.0493 | 121.936 | 105 | 105.0049 | 112.9754 | 84.9729 | 93.9803 | 76.0246 | 79.0049 | 95.0123 | 116.9606 | 112.9704 | 88.9581 | 238.9754 | 218.9631 | 220.0739 | 263.0123 | 166.0074 | 191.9729 | 197.9729 | 217.0443 | 327 | 237.9335 | 234.0985 | 326.1034
2443 | 28.8674 | 163.1623 | 172.6492 | 219.1328 | 137.177 | 133.1328 | 137.3541 | 177.6607 | 135.0738 | 157.4394 | 139.3951 | 93.3098 | 129.2508 | 194 | 139.1033 | 159.8377 | 178.0738 | 187.5574 | 140.8082 | 128.1033 | 168.6312 | 159.6017 | 282.9557 | 323.9705 | 139.3393 | 234.6754 | 227.7787 | 229.8377 | 276.9525 | 127.9852 | 98.1475 | 122.8672 | 147.6164 | 143.9705 | 126.9705 | 137.1033 | 91.0885 | 248.6639 | 244.1918 | 246.9557 | 204.1213 | 153.0148 | 145.8525 | 151.941 | 106.2361 | 93.9852 | 83.9557 | 85.0148 | 126.9705 | 184.5901 | 193.8525 | 220.1033 | 184.3803 | 135.0148 | 139.0443 | 129.2508 | 98.0443 | 175.3541 | 157.8377 | 194.4279 | 145.8967 | 149.6754 | 129.882 | 154.9557 | 137.9115 | 187.9262 | 185.6902 | 206.9262 | 219.0295 | 118.4426 | 177.7492 | 233.7344 | 180.2508 | 219.6607 | 199.0443 | 199.9705 | 241.2803 | 172.882 | 167.7639 | 174.882 | 185.3541 | 136.1328 | 144.1623 | 147.7639 | 138.8377 | 187.823 | 177.8525 | 172.1918 | 141.6164 | 112.1033 | 118.9705 | 115.2361 | 127.823 | 277.2803 | 245.2066 | 290.2803 | 301.7197 | 269 | 255.0705 | 306.0263 | 328.4689 | 99.4869 | 80.8082 | 81.7934 | 99.4869 | 118.8525 | 107.3393 | 111.0295 | 121.941 | 95.9115 | 104.9557 | 106.9852 | 103.059 | 73.7639 | 85.823 | 85.8967 | 81.059 | 100.2508 | 100.9852 | 100.8377 | 72.1623 | 229.0443 | 204.2066 | 249.5869 | 267.941 | 169 | 180.9705 | 187.1033 | 234.6902 | 327.6344 | 211.4869 | 273.5279 | 367.6607
2444 | 28.8784 | 173.8107 | 215.405 | 227.973 | 149.2704 | 142.5409 | 160.6755 | 155.4598 | 139.973 | 120.3245 | 99.2704 | 113.9189 | 145.0534 | 193.8918 | 145.6214 | 148.6214 | 183.9195 | 158.1352 | 128.1893 | 135.3786 | 144.9195 | 133.0541 | 280.4327 | 321.8918 | 161.378 | 213.1082 | 212.378 | 218.8918 | 208.5145 | 126.5673 | 108.027 | 114.5409 | 122.2434 | 142.2704 | 125.4598 | 143.4321 | 96.5673 | 292.6214 | 257.1352 | 244.027 | 277.6478 | 153.7566 | 135.7025 | 148 | 121.8107 | 92.7025 | 81.1893 | 86.027 | 124.2427 | 222.9453 | 184.2975 | 226.2157 | 142.7296 | 135.7296 | 141.1075 | 145.973 | 101.3516 | 198.2968 | 147.8384 | 222.1075 | 139.0541 | 128.3516 | 122.3516 | 152.4327 | 132.1893 | 182.8648 | 165.3245 | 202.1082 | 220.5132 | 147.8918 | 160.7566 | 216.3516 | 196.405 | 196.5402 | 202.1082 | 198.3786 | 261.1359 | 165.2434 | 151.7566 | 167.3245 | 207.5937 | 144.4591 | 154.2968 | 132.1623 | 128.3245 | 176.5138 | 168.3786 | 185 | 116.2164 | 118.973 | 116.7296 | 131.2704 | 115.7836 | 295.405 | 259.4868 | 307.6748 | 284.4604 | 269.7032 | 194.2441 | 242.2441 | 292.2968 | 131.8377 | 68.9466 | 68.2704 | 131.8377 | 108.8648 | 129.2968 | 112.973 | 117.5402 | 90.7573 | 101.9189 | 106 | 107.027 | 58.4868 | 73.9459 | 79.0811 | 85.595 | 116.5402 | 99.4321 | 89.8648 | 83.2434 | 231.6214 | 218.0541 | 222.027 | 263.5132 | 169.1082 | 179.9466 | 194 | 214.7032 | 368.2421 | 242.6478 | 242.1623 | 344.2968
2445 | 28.8894 | 167.1967 | 193.3706 | 227.5507 | 159.3934 | 161.4886 | 150.2194 | 170.5052 | 139 | 133.0227 | 109.7474 | 110.1739 | 112.3768 | 189.6066 | 132.2754 | 134.6853 | 215.9379 | 163.0393 | 135.6687 | 148.6853 | 177.8033 | 135.5114 | 296.236 | 318.1573 | 139.7474 | 217.354 | 190.472 | 215.0393 | 263.4327 | 111.1967 | 109.6687 | 132.7019 | 131.118 | 151.8033 | 142.3147 | 123.4327 | 81.7474 | 278.1739 | 260.9773 | 244.7246 | 229.118 | 145.3147 | 125.9047 | 147.528 | 115.7867 | 81.9607 | 88.2754 | 86.1739 | 97.5114 | 185.59 | 194.764 | 198.118 | 133.118 | 127.3768 | 110.0227 | 144.1739 | 113.4886 | 172.8033 | 176.5446 | 190.9834 | 140.6853 | 140.4493 | 134.6066 | 167.6066 | 138.1739 | 177.9213 | 175.9773 | 206.6687 | 203.0787 | 144.708 | 152.236 | 228.528 | 175.354 | 179.9213 | 205.2526 | 211.1739 | 302.2754 | 174.0787 | 143.4327 | 179.0787 | 158.4948 | 126.0227 | 128.8427 | 138.472 | 139.882 | 194.41 | 181.0953 | 185.59 | 124.2754 | 117.4493 | 107.4327 | 139.6232 | 107.764 | 273.9607 | 276.5673 | 260.0393 | 337.2754 | 293.6232 | 238.41 | 287.4327 | 267.9834 | 125.8427 | 101.5839 | 78 | 125.8427 | 104.354 | 104.4327 | 111.528 | 101.3934 | 118.118 | 98.6066 | 106.472 | 107.9213 | 76.3147 | 72.8654 | 81.5673 | 107.0393 | 100.4327 | 79.5507 | 85.59 | 92.5114 | 219.1801 | 220 | 223.8261 | 246.3934 | 172.292 | 212.4266 | 194.6294 | 239.2133 | 307.0848 | 196.0455 | 247.4493 | 319.0393
2446 | 28.9004 | 170.9158 | 178.4646 | 240.4321 | 168.2256 | 149.3614 | 178.7093 | 135.7037 | 138.3805 | 156.761 | 126.8642 | 91.4456 | 145.3289 | 181.3423 | 139.5163 | 128.1874 | 190 | 165.0842 | 151.0707 | 141 | 174.1874 | 149.1358 | 301.2772 | 322 | 156.2963 | 226.2581 | 200.9675 | 216.826 | 273.6902 | 116.1033 | 125.6386 | 102.6521 | 133.4321 | 145.8642 | 149.5354 | 134.413 | 100.0516 | 256.761 | 237.3939 | 238.5679 | 233.4972 | 152.7419 | 147.5354 | 136.0516 | 135.1033 | 80.587 | 94.3289 | 67.4456 | 110.6195 | 200.0516 | 188.7419 | 201.7744 | 136.3614 | 159.9675 | 134.2256 | 124.7744 | 101.4646 | 169.0842 | 141.413 | 213.5028 | 133.826 | 127.3614 | 125.0516 | 158.3098 | 118.2065 | 176.7744 | 151.9809 | 222.3805 | 204.8451 | 161.0707 | 157.0707 | 216.2772 | 183.3289 | 180.1167 | 188.2907 | 192.9618 | 307.1414 | 174.1931 | 154.5679 | 180.5354 | 194.4837 | 150.4321 | 126.1874 | 148.5544 | 136.8451 | 179.8451 | 160.4972 | 199.2256 | 131.3614 | 105.3939 | 117.5354 | 107.4972 | 102.9293 | 273.9293 | 264.761 | 263.0134 | 343.8451 | 258.9675 | 224.9809 | 295.8833 | 292.0516 | 121.1223 | 67.5163 | 78.4646 | 121.1223 | 112.3805 | 113.9675 | 100.9293 | 110.5354 | 120.0707 | 90.1874 | 117.6386 | 106.5679 | 83.4321 | 93.0191 | 71.5163 | 107.0707 | 110.4837 | 93.1549 | 101.1874 | 103.2963 | 247.1223 | 219.7419 | 242.0382 | 255.587 | 155.1549 | 174.0516 | 210.2581 | 220.3614 | 356.8126 | 245.1223 | 234.4646 | 320.7744
2447 | 28.9114 | 153.1732 | 187.2557 | 228.9134 | 154.6392 | 155.0412 | 155.2557 | 165.466 | 127.1918 | 134.1918 | 107.9175 | 117.8082 | 134.2784 | 205.9361 | 148.2969 | 149.5526 | 191.0227 | 183.6577 | 134.1918 | 139.8495 | 193.8268 | 170.0639 | 286.466 | 323.9815 | 127.2371 | 229.7216 | 184.4289 | 231.6165 | 267.4247 | 117.3608 | 118.1051 | 131.5938 | 123.0639 | 125.6392 | 141.1278 | 141.9361 | 100.6165 | 235.0866 | 261.3381 | 248.3608 | 259.3381 | 146.7216 | 138.5526 | 136.7443 | 136.8082 | 74.7258 | 82.5753 | 92.5938 | 121.6165 | 202.2144 | 185.5979 | 216.4474 | 143.8309 | 140.8082 | 119.2969 | 138.7443 | 110.7031 | 186.635 | 150.2144 | 188.1093 | 149.0639 | 134.6392 | 127.2784 | 163.9361 | 122.4474 | 189.6577 | 169.3608 | 209.5299 | 203.534 | 144.9588 | 140.767 | 205.0454 | 170.6165 | 217.466 | 210.9134 | 225.4845 | 274.534 | 143.2371 | 165.4474 | 172.0639 | 186.4062 | 140.3196 | 148.6392 | 123.4701 | 133.6804 | 175.0185 | 186.8495 | 185.6392 | 138.2557 | 130.8082 | 111.0454 | 133.7856 | 119.233 | 289.2742 | 243.5979 | 298.2742 | 339.7216 | 240.6392 | 242.233 | 259.2371 | 292.9361 | 105.9588 | 76.8722 | 86.0412 | 105.9588 | 102.0866 | 95.2557 | 118.1918 | 103.0866 | 103 | 110.7856 | 111.3196 | 115.9773 | 73.4474 | 75.1278 | 81.6392 | 90 | 102.0866 | 96.0639 | 122.4887 | 72.767 | 231.5113 | 214.0412 | 205.8722 | 248.8949 | 159.534 | 174.7443 | 215.3196 | 228.0866 | 334.5526 | 227.2103 | 241.5938 | 333.2103
2448 | 28.9224 | 184.0093 | 192.9052 | 211.7807 | 163.2379 | 141.6097 | 159.9907 | 142.0093 | 130.4573 | 137.6859 | 135.5427 | 114.3903 | 152.3903 | 203.9331 | 138.8383 | 143.1524 | 204.3996 | 164.7621 | 136.9238 | 124.0669 | 162.7621 | 168.3327 | 269.5633 | 350.4851 | 157.8661 | 213.4387 | 217.7899 | 226.8383 | 260.2956 | 108.3048 | 104.8476 | 112.9052 | 124.2286 | 134.1617 | 143.1524 | 142.5242 | 94.8476 | 248.7899 | 235.0186 | 240.5242 | 235.3048 | 130.2101 | 131.0093 | 133.6097 | 132.7807 | 99.0855 | 89.8569 | 72.9145 | 114.171 | 220.1524 | 207.9331 | 221.552 | 154.552 | 136.6282 | 111.2101 | 133.0948 | 120.4665 | 152.6859 | 166.8569 | 215.723 | 149.9238 | 143.4665 | 147.7528 | 164.5242 | 129 | 171.0669 | 159.3903 | 188.8383 | 226.6859 | 157.0948 | 152.8383 | 234.9238 | 167.2863 | 195.381 | 195.9145 | 177.8104 | 296.2379 | 174.9331 | 173.2956 | 171.3996 | 205.552 | 144.0093 | 156.9331 | 143.171 | 130.829 | 146.6859 | 169.619 | 193.7807 | 141.1617 | 128.0762 | 138.8661 | 116.9145 | 107.5427 | 265.7528 | 263.8754 | 273.2286 | 319.3996 | 250.6859 | 232.9814 | 291.0093 | 292 | 120.3048 | 75.4573 | 73.6004 | 120.3048 | 117.0855 | 97.9331 | 118.7899 | 117.7807 | 102.7714 | 94.2956 | 114.4758 | 99.8569 | 78.0948 | 77.381 | 90.619 | 90.2286 | 117.1617 | 95.8569 | 114.0855 | 83.4665 | 238.2379 | 201.7528 | 206.6673 | 262.1524 | 178.8754 | 171.1524 | 219.619 | 244.9907 | 329.6004 | 202.2956 | 221.4573 | 310.5818
2449 | 28.9334 | 170.761 | 212.0175 | 197.4425 | 154.177 | 147.23 | 169.0795 | 134.8675 | 137.8585 | 146.885 | 129.3805 | 108.6815 | 144.7345 | 191 | 147.4955 | 145 | 182.345 | 174.531 | 135.0265 | 138.593 | 170.938 | 137.8585 | 334.345 | 318.584 | 132.9735 | 242.292 | 193.0355 | 235.6725 | 271.398 | 111.9115 | 104.3275 | 136 | 127.708 | 125.6815 | 141.991 | 158.9645 | 93.885 | 226.1595 | 211.177 | 259.531 | 238.5575 | 154.699 | 120.593 | 140.9115 | 119.3275 | 87.823 | 76.6195 | 84.885 | 93.77 | 219.9645 | 194.9115 | 205.416 | 138.593 | 122.9205 | 137.823 | 112.301 | 113.469 | 160.823 | 155.655 | 177.77 | 147.584 | 136.5575 | 165.106 | 180.6105 | 129.531 | 184.7965 | 155.0975 | 199.2655 | 233.23 | 135.77 | 163.0885 | 231.7875 | 195.354 | 199.115 | 205.407 | 221.2565 | 288.1505 | 162.531 | 187.761 | 152.7965 | 188.7965 | 132.4425 | 142.8495 | 120.5575 | 152.823 | 157.655 | 165.9735 | 179.2655 | 132.3275 | 128.7345 | 112.823 | 128.0885 | 102.531 | 284.522 | 226.531 | 274.407 | 301.8585 | 261.2125 | 255.4955 | 281.7435 | 292.9735 | 123.469 | 79.4955 | 92.292 | 123.469 | 106.708 | 86.239 | 95.2745 | 104.77 | 99.7345 | 109.9115 | 96.885 | 86 | 57.478 | 82.885 | 84.6725 | 94.1505 | 106.115 | 84.301 | 101.407 | 76.823 | 228.9115 | 222.646 | 236.345 | 260.106 | 141.885 | 173.9735 | 211.7255 | 255.053 | 346.9645 | 218.708 | 226.115 | 352.8495
2450 | 28.9444 | 158.8063 | 173.5198 | 202.2016 | 156.2016 | 129.4032 | 138.6976 | 181.5731 | 153.0691 | 156.8063 | 123.8063 | 123.3794 | 141.3953 | 192.4111 | 132.2016 | 146.4111 | 158.3261 | 179.2945 | 123.9921 | 152.7826 | 160.3024 | 159.7214 | 309.9229 | 303.3953 | 142.7984 | 229.6581 | 213.6976 | 226.1324 | 229.2411 | 113.0158 | 116.9921 | 132.6739 | 135.7055 | 145.9229 | 115.6364 | 137.1937 | 102.0929 | 266.1779 | 216.4269 | 265.1008 | 234.3024 | 129.5889 | 135.8913 | 139.8992 | 131.7905 | 85.7984 | 82.1937 | 93.1937 | 111.1937 | 198.7984 | 191.9842 | 222.0158 | 152.7826 | 152.7984 | 136.1008 | 134.4881 | 107.6976 | 161.3182 | 181.1858 | 196.2016 | 139.1483 | 132 | 124.3953 | 156.8063 | 136.1087 | 189.8755 | 185.5889 | 201.7984 | 214.2945 | 152.7905 | 163.2945 | 211.419 | 198.0929 | 190.504 | 189.2016 | 192.3953 | 300.1008 | 166.5889 | 172.1779 | 165.6364 | 194.8834 | 139.5198 | 130.7984 | 115.6976 | 150.6976 | 186.915 | 173.581 | 181.0929 | 143.2786 | 125.0929 | 113.2174 | 128.5968 | 109.3103 | 260.6127 | 234.9229 | 259.2095 | 319.5889 | 279.5652 | 239.8992 | 307.4802 | 302.3953 | 116.4881 | 66.0158 | 85.2016 | 116.4881 | 113.496 | 99.9071 | 128.8992 | 121.6897 | 99.2174 | 107.1858 | 106.8063 | 86.8063 | 83 | 89.7826 | 71.1008 | 104.9921 | 97.2016 | 105.6818 | 85.4032 | 75.1008 | 231.1245 | 220.4111 | 207.2866 | 223.0237 | 152.7134 | 183.1937 | 178.8063 | 234.496 | 327.1087 | 223.8834 | 218.3103 | 341.8063
2451 | 28.9554 | 166.7916 | 196.5654 | 203.8869 | 159.2439 | 131.9823 | 138.6009 | 151.3392 | 121.5055 | 160.2683 | 130.5477 | 102.3747 | 136.1131 | 204.6608 | 134.6785 | 157.643 | 187.4346 | 173.3392 | 118.2794 | 131.8692 | 162.7738 | 181.1729 | 332.3814 | 302.8625 | 139.9823 | 188.898 | 209.643 | 257.2328 | 274.1907 | 129.9823 | 106.9823 | 104.5831 | 142.1131 | 167.816 | 155.1907 | 131.6962 | 92.9823 | 241.2262 | 243.4945 | 262.9468 | 238.357 | 122.7672 | 126.9047 | 137.3038 | 120.6608 | 85.6962 | 75.3392 | 86.5654 | 104.7916 | 196.7738 | 174.7916 | 238.7561 | 135.2616 | 148.3991 | 136.5477 | 123.714 | 106.0177 | 180.643 | 164.7738 | 197.5477 | 190.6075 | 134.0355 | 121.0355 | 165.1308 | 143.5122 | 162.9047 | 170.5122 | 199.7738 | 208.1131 | 139.8514 | 157 | 229.9468 | 188.643 | 196.8093 | 192.1308 | 186.7738 | 303.0355 | 156.2616 | 147.9047 | 202.4767 | 175.7738 | 157.816 | 128.6608 | 115.6009 | 146.3038 | 197.8936 | 152.4523 | 173.1131 | 122.5055 | 116.5477 | 131.4169 | 123.9823 | 121.4523 | 274.0953 | 261.9047 | 269.8869 | 308.0177 | 244.5233 | 239.7916 | 283.5299 | 295.5299 | 104.6962 | 82.4169 | 83.9468 | 104.6962 | 110.0177 | 108.2262 | 126.7561 | 111.1308 | 118.3215 | 94.0532 | 113.4346 | 94.7916 | 84.0177 | 72.3747 | 73.357 | 95.5477 | 98.4346 | 85.9047 | 89.5654 | 76.2262 | 256.0599 | 230.5122 | 198.1064 | 252.4878 | 168 | 177.4701 | 189.3925 | 231.8093 | 340.5055 | 207.0355 | 231.4701 | 349.9047
2452 | 28.9664 | 169.1134 | 198.7433 | 202.2478 | 165.4896 | 126.7582 | 158.1224 | 155.6299 | 146.994 | 130.6239 | 126.7493 | 119.4955 | 139.2567 | 200.7463 | 140.6269 | 147.3761 | 181.8716 | 174.7463 | 143.2478 | 126.2597 | 161.5015 | 157.4955 | 306.006 | 337.991 | 133.1284 | 241.8716 | 201.6328 | 215.7612 | 256.7403 | 120.6209 | 100.5045 | 117.8776 | 140.2418 | 134.6239 | 141.2507 | 144.2478 | 85.5015 | 247.3881 | 214.997 | 239 | 247.997 | 165.1164 | 133.8746 | 128.0119 | 117.8746 | 97.4955 | 76.3701 | 89.4955 | 110.6239 | 194.6239 | 183.5075 | 227.997 | 149.991 | 126.4955 | 132.7493 | 141.8657 | 114.3761 | 170 | 160.4925 | 193.2478 | 161.3672 | 148.3701 | 134.2418 | 169.6119 | 126.8836 | 169.1224 | 151.5015 | 197.6239 | 209 | 126.5104 | 155.2448 | 207.2537 | 178.7522 | 210.4985 | 199.3701 | 185.8776 | 312.4806 | 171.1164 | 154.1224 | 164.4925 | 175.003 | 128.009 | 125.4985 | 136.3761 | 133.5015 | 154.8896 | 158.5075 | 173.8746 | 145.7373 | 113.5015 | 117.6209 | 119.5104 | 122.994 | 267.2507 | 268.6239 | 275.0179 | 315.6239 | 278.1194 | 249.1343 | 272.5015 | 286.7582 | 118.3761 | 72.1313 | 61.3791 | 118.3761 | 118.1254 | 109.1224 | 117.2507 | 119.6239 | 111.6209 | 116.8716 | 108.2478 | 100.4985 | 90.1194 | 90.8746 | 83.8746 | 92 | 94.3761 | 94.8806 | 94.6269 | 78.3761 | 236.5104 | 214.6358 | 238.4776 | 270.2448 | 166.997 | 187.1194 | 211.8627 | 239.9821 | 366.3701 | 221.4955 | 247.8896 | 353.3642
2453 | 28.9774 | 142.826 | 184.239 | 196.174 | 142.5143 | 143.6597 | 152.6883 | 163.0077 | 134.652 | 128.5507 | 125.5507 | 107.2104 | 154.3117 | 192.5507 | 145.9636 | 151.3766 | 173.174 | 169.3039 | 134.4207 | 146.7974 | 165.6883 | 148.9273 | 317.7974 | 314.2467 | 141.1377 | 238.5429 | 219.587 | 238.9714 | 241.1585 | 112.9273 | 114.0286 | 125.9273 | 122.7896 | 133.9273 | 142.3117 | 138.4493 | 89 | 273.5947 | 210.478 | 242.4416 | 240.174 | 148.7169 | 134.5143 | 150.2181 | 115.0727 | 86.4493 | 66.9273 | 79.9636 | 108.6883 | 193.1013 | 199.2104 | 217.0077 | 132.3039 | 117.9273 | 131 | 120.9636 | 115.2104 | 167.3844 | 145.7533 | 187.4493 | 143.9636 | 135.4857 | 117.478 | 142.1663 | 144.7974 | 165.2026 | 157.2026 | 195.6883 | 212.0286 | 150.4493 | 143.5507 | 215.3117 | 181.6597 | 205.348 | 187.0364 | 188.1454 | 270.4416 | 148.3844 | 148.2753 | 145.348 | 181.8623 | 149.9636 | 125.3039 | 139.6883 | 134.6597 | 188.7247 | 175.4493 | 168.8701 | 117.5143 | 116.8623 | 109.652 | 142.0727 | 109.5507 | 274.3689 | 268.3403 | 311.0804 | 315.6156 | 267.2026 | 266.5947 | 276 | 302.1454 | 120.587 | 86.4493 | 70.587 | 120.587 | 116.1091 | 104.239 | 119.826 | 117.2753 | 102.413 | 111.2026 | 103.2753 | 98.9273 | 77.413 | 88.8987 | 82.0364 | 93.3766 | 99.3403 | 104.6961 | 99.6883 | 82.3766 | 260.4493 | 235.3194 | 191.6442 | 261.4416 | 161.1013 | 174.5507 | 183.9636 | 203.7819 | 350.4571 | 209.348 | 276.6311 | 324.6961
2454 | 28.9884 | 148.15 | 191.4002 | 194.8988 | 153.4996 | 133.4987 | 158.3496 | 144.1482 | 146.1996 | 134.8492 | 128.2502 | 101.5483 | 152.0994 | 196.4499 | 151.2502 | 159.4002 | 174.1482 | 186.701 | 114.2493 | 129.4011 | 168.8004 | 159.5004 | 309.8483 | 302.0985 | 146.1987 | 262.0515 | 211.9015 | 222.0994 | 282.2511 | 122.7507 | 128.0515 | 136.0506 | 136.5492 | 144.6504 | 138.7498 | 134.7001 | 89.8997 | 244.5004 | 226.7498 | 261.9006 | 238.1491 | 168.701 | 142.5004 | 115.0497 | 102.85 | 84.6495 | 79 | 84.8004 | 112.85 | 200.7498 | 193.0985 | 192.15 | 154.3496 | 130.8997 | 131.2999 | 126.4002 | 104.7498 | 153.2493 | 162.7001 | 186.2493 | 148.5004 | 128.2493 | 136.5989 | 159.6007 | 132.6495 | 176.7507 | 172.7995 | 201.7995 | 231.2999 | 147.7498 | 148.9494 | 213.8492 | 166.2502 | 195.5998 | 183.6992 | 167.1491 | 294.9991 | 132.8997 | 150.15 | 138.7489 | 183.9991 | 156.4499 | 143.0009 | 141.6007 | 125.0985 | 186.1003 | 171.5501 | 147.9485 | 127.7498 | 117.6495 | 119.1003 | 132.9991 | 112.7001 | 304.4011 | 280.7507 | 279.3984 | 336.0488 | 278.7507 | 239.4499 | 275.5501 | 281.7489 | 121.8988 | 84.0497 | 71.1491 | 121.8988 | 99.1996 | 113.0497 | 125.2999 | 118.85 | 104.5501 | 121.701 | 107.0994 | 109.2005 | 79.7001 | 81.5501 | 75.85 | 99.1508 | 113.1003 | 85.7995 | 101.6007 | 88.9006 | 251.3017 | 209.1491 | 225.6504 | 269.2041 | 164.4011 | 179.6495 | 187.9006 | 244.7986 | 325.8988 | 201.9991 | 240.6007 | 305.0497
2455 | 28.9994 | 150.4602 | 189.1357 | 219.4337 | 162.649 | 158.5133 | 165.6755 | 175.9173 | 151.5398 | 149.2153 | 120.9173 | 132.5133 | 160.2683 | 204.0296 | 149.1092 | 151.6193 | 204.9438 | 167.351 | 130.5694 | 112.0827 | 161.5133 | 150.5133 | 336.0234 | 330.5133 | 166.7551 | 229.4571 | 186.4072 | 233.1888 | 263.9204 | 110.1622 | 103.0827 | 122.8908 | 147.5928 | 138.6225 | 143.1622 | 133.4867 | 93.2153 | 240.3807 | 226.9438 | 248.053 | 252.2683 | 149.351 | 139.6786 | 123.2714 | 106.8674 | 92.702 | 77.702 | 77.351 | 114.1092 | 200.2949 | 217.4571 | 197.0562 | 161.351 | 136.4867 | 131.5398 | 123.3245 | 111.4337 | 170.7051 | 161.3245 | 199.8112 | 139.6755 | 139.3775 | 150.6724 | 146.1622 | 137.4571 | 163.6755 | 180.8908 | 203.7254 | 229.1061 | 150.8643 | 157.0796 | 227.7286 | 167.5164 | 196.4041 | 200.947 | 183.702 | 310.053 | 142.7051 | 150.8378 | 162.5959 | 202.1357 | 155.1061 | 129.5694 | 130.9204 | 151.8908 | 177.2683 | 165.1061 | 177.947 | 131.0265 | 126.6755 | 115.1357 | 143.5101 | 111.1622 | 279.7816 | 262.3214 | 299.1295 | 356.8908 | 269.2449 | 240.2153 | 274.7847 | 298.1326 | 140.7551 | 86.5928 | 87.2153 | 140.7551 | 107.7847 | 117.702 | 127.649 | 117.1888 | 101.6755 | 103.3245 | 116.0796 | 100.1357 | 78.1622 | 76.7286 | 74.8378 | 82.351 | 105.8908 | 98.5959 | 89.4602 | 78.6225 | 221.4337 | 221.6459 | 214.4306 | 199.5991 | 146.7582 | 188.5133 | 181.8409 | 263.2387 | 348.3245 | 219.8112 | 225.2153 | 309.702
2456 | 29.0104 | 156.4292 | 194.3019 | 233.0472 | 163.3819 | 151.2875 | 165.5708 | 160.5236 | 143.6509 | 140.5236 | 109.887 | 127.3819 | 139.6037 | 222.3204 | 158.6037 | 130.7454 | 186.4435 | 162.4292 | 149 | 126.9528 | 158.3019 | 150.9671 | 302.4435 | 327.1273 | 154.1889 | 215.0759 | 206.6181 | 230.9199 | 277.2546 | 112.0472 | 117.7782 | 114.9671 | 132.6181 | 145.4292 | 141.9056 | 135.6509 | 84.5236 | 261.9528 | 208.6181 | 236.4292 | 239.1088 | 144.9528 | 164.113 | 132.2074 | 127.462 | 86.8727 | 75.7125 | 74 | 120.9856 | 178.269 | 196.269 | 217.4764 | 154.5093 | 136.3819 | 126.7926 | 124.4764 | 119.462 | 188.3676 | 162.8255 | 202.7782 | 138.1745 | 132.2218 | 119.7454 | 149.9671 | 117.8398 | 163.7454 | 171.2218 | 167.2833 | 208.3019 | 148.3162 | 143.7454 | 220.3634 | 190.9384 | 180.731 | 205.7639 | 179.6181 | 298.4292 | 166.8255 | 154.1889 | 168.844 | 204.6838 | 140.0616 | 147.3019 | 144.7782 | 141.3491 | 163.5852 | 149.3634 | 184.8583 | 125.3019 | 121.1602 | 119.4292 | 110.3491 | 113.7454 | 255.0329 | 231.4292 | 252.0759 | 342.8583 | 276.4004 | 237.1088 | 286.0944 | 273.5236 | 124 | 69.8727 | 77.1273 | 124 | 115.6037 | 112.0472 | 127.6838 | 109.8583 | 101.9199 | 105.5236 | 103.0944 | 104.9528 | 80.9199 | 65.5236 | 77.1417 | 81.269 | 94.1273 | 108.1602 | 97.5236 | 88.3963 | 231.731 | 195.9199 | 202.2833 | 239.113 | 160.8255 | 184.0801 | 206.2403 | 222.3634 | 350 | 224.3491 | 225.0759 | 312.425
2457 | 29.0214 | 151.6156 | 188.5712 | 234.6371 | 156.418 | 129 | 171.1317 | 163.9341 | 141.0659 | 144.121 | 129.8239 | 114.0659 | 127.7688 | 186.3737 | 149.1976 | 141.2419 | 198.9341 | 160.0444 | 146.5712 | 124.2419 | 156.8683 | 164.8132 | 310.4503 | 321.1317 | 171.9449 | 254.6371 | 216.7581 | 233.6478 | 272.3629 | 118.4946 | 110.6922 | 129.1868 | 142.7581 | 139.3078 | 134.0551 | 135.121 | 85.1317 | 257.5605 | 220.2527 | 236.0995 | 259.7795 | 140.5605 | 141 | 118.2527 | 107.6815 | 86.8898 | 95.3844 | 78.1102 | 110.0336 | 190.8683 | 204.3844 | 213.6922 | 139.121 | 125.6815 | 138.6922 | 125.3629 | 102.1102 | 160.1761 | 160.6922 | 195.6922 | 138.2527 | 139.4946 | 124.0766 | 160.4503 | 135.6263 | 168.4503 | 175.3185 | 196.4395 | 201.8239 | 162.879 | 151.2527 | 238.7581 | 171.6707 | 175.418 | 188.2863 | 191.4395 | 288.3844 | 165.2527 | 173.0659 | 135.9341 | 189.1868 | 162.1317 | 143.4395 | 139 | 140.1976 | 184.0215 | 170.1868 | 172.0551 | 125.5497 | 106.3629 | 111.6263 | 112.9341 | 121.6263 | 250.5283 | 226.9893 | 286.2191 | 330.4288 | 236.7917 | 264.3185 | 292.4503 | 284.0336 | 123.2527 | 77.7366 | 75.6156 | 123.2527 | 111.2419 | 117.3737 | 114.0551 | 98.3629 | 108.1976 | 108.7473 | 113.9341 | 102.9893 | 84.3951 | 68 | 88.6371 | 91.6263 | 91.8683 | 89.4395 | 99.8132 | 91.6371 | 223.8024 | 204.2527 | 226.9556 | 218.9893 | 168.0336 | 178.1761 | 177.5054 | 239.2634 | 341.0322 | 223.1976 | 259.4059 | 350.2742
2458 | 29.0324 | 164.3938 | 180.9868 | 216.8141 | 172.4335 | 133.5797 | 159.8141 | 161.2299 | 136.6018 | 149.3982 | 116.3982 | 109.0044 | 114.5577 | 189.7921 | 140.7833 | 152.3938 | 201.0088 | 180.407 | 140.1815 | 130.615 | 161.2167 | 163.0044 | 301.7701 | 313.9956 | 160.8097 | 232.6326 | 206.0044 | 208.7877 | 277.6414 | 122.8097 | 104.2035 | 127.0132 | 131.407 | 140.2211 | 145.7877 | 144.1815 | 78.3938 | 258.2079 | 221.1815 | 262.4203 | 231.1683 | 140.8097 | 139.8053 | 123.363 | 112.4114 | 74.9736 | 81.2123 | 92.0176 | 136.2387 | 197.2079 | 189.615 | 208 | 150.1727 | 133.3982 | 133 | 136.0176 | 117.0132 | 173.2035 | 157.1903 | 192.7877 | 138.1859 | 144.2079 | 114.7657 | 145.5974 | 136.7877 | 157.7789 | 170.1903 | 194.3982 | 190.9868 | 160.5886 | 149.3938 | 229.7965 | 185.8273 | 195.615 | 215.4203 | 188.2035 | 274.0132 | 162.3982 | 171.1903 | 143.1859 | 187.6106 | 152.2079 | 141 | 145.3718 | 131.385 | 163.1771 | 165.6238 | 176.6194 | 138.2123 | 120.0044 | 106.8141 | 118.593 | 118.0088 | 287.8185 | 243.385 | 251.7833 | 341.7965 | 265.385 | 251.8229 | 271.8229 | 314.6194 | 124.978 | 89.8185 | 89.3894 | 124.978 | 119.0088 | 119.5974 | 119.2167 | 111.2079 | 97.593 | 114.5886 | 115.0132 | 116.1991 | 63.5842 | 67.4026 | 74.5974 | 88.407 | 96.0176 | 88.1947 | 99.7965 | 79.3894 | 235.0044 | 213.5445 | 207.7877 | 232.5974 | 189.0616 | 190.2079 | 174.385 | 223.9956 | 307.5753 | 214.7833 | 218.2167 | 325.5797
2459 | 29.0434 | 168.0973 | 194.0571 | 201.8859 | 144.9597 | 150.0973 | 147.6342 | 140.1309 | 138.3826 | 151.2114 | 120.1141 | 107.9597 | 149.406 | 198.4799 | 151.1376 | 162.0168 | 195.5369 | 172.5201 | 155.7315 | 120.594 | 154.0336 | 161.7483 | 323.0805 | 319.6913 | 153.2685 | 207.4396 | 207.9195 | 219.3658 | 247.3994 | 117.1712 | 101.8456 | 114.7886 | 128.3826 | 125.8624 | 151.9261 | 152.1208 | 86.7483 | 254.8054 | 235.2517 | 244.8456 | 249.6409 | 130.9429 | 136.6913 | 154.2517 | 105.594 | 95.1946 | 73.8054 | 73.0403 | 102.4799 | 191.9027 | 174.7315 | 211.8054 | 168.0235 | 134.3658 | 134.9027 | 124.4396 | 110.4967 | 171.6913 | 164.3087 | 204.8456 | 153.7483 | 138.9027 | 143.8859 | 152.0168 | 146.9429 | 174.0403 | 179.8456 | 193.2517 | 198.5604 | 172.2685 | 153.5201 | 227.5033 | 167.7651 | 183.2685 | 191.2919 | 181.406 | 272.9933 | 164.4228 | 173.2349 | 157.9027 | 186.245 | 141.406 | 144.3826 | 164.6577 | 143.5201 | 174.6979 | 152.4564 | 158.7886 | 122.9832 | 113.0403 | 101.8222 | 122.6745 | 111.4799 | 271 | 252.5604 | 268.4799 | 341.1946 | 275.6174 | 234.594 | 255.8624 | 293.1946 | 140.349 | 74.6913 | 95.406 | 140.349 | 115.2282 | 119.4631 | 106.5369 | 102.9429 | 104.6342 | 117.1778 | 101.9429 | 115.0973 | 79.2685 | 67.1141 | 79.9597 | 80.9429 | 88.4564 | 93.4228 | 103.8456 | 86.0403 | 232.0571 | 254.0637 | 216.0403 | 240.4967 | 143.1477 | 183.8456 | 191.3826 | 235.6108 | 326.406 | 225.7718 | 211.6678 | 342.7315
2460 | 29.0544 | 163.237 | 191.9601 | 194.6711 | 154.6578 | 144.3422 | 146.8682 | 159.6445 | 150.7763 | 152.4474 | 125.9867 | 121.4607 | 139.3555 | 204.2237 | 137.0266 | 169.3941 | 201.4208 | 165.4341 | 148.763 | 131.0919 | 178.8549 | 165.5127 | 307.4873 | 337.1851 | 158.4474 | 232.9481 | 224.4075 | 217.2237 | 273.6578 | 124.526 | 107.4607 | 114 | 132.723 | 143.3289 | 134.3422 | 123.0266 | 94.3156 | 267.6578 | 236.6325 | 250.0133 | 217.2636 | 127.8948 | 136.5127 | 145.7896 | 123.0266 | 81 | 90.0133 | 69.3822 | 107.3289 | 193.1838 | 178.9481 | 218.6178 | 144.4607 | 130.2104 | 143.3422 | 139.6578 | 124.5127 | 172.8549 | 155.763 | 207.3289 | 162.6578 | 136.3808 | 138.4607 | 170.1318 | 141.4341 | 167.6977 | 186.5792 | 188.8162 | 183.7896 | 172.0786 | 150.0133 | 209.237 | 188.7497 | 193.8162 | 184.5406 | 171.8029 | 307.9601 | 166.8948 | 155.3822 | 163.2104 | 210.9601 | 131.1318 | 150.0653 | 136.7497 | 140.237 | 146.711 | 183.1052 | 161.3555 | 112.2503 | 103.7896 | 124.7364 | 117.8029 | 121.474 | 274.8029 | 238.9081 | 264.1571 | 322.526 | 258.0786 | 251.3555 | 274.4474 | 285.711 | 121.8815 | 81.8948 | 82.2237 | 121.8815 | 126.526 | 106.8682 | 111.9734 | 103.6977 | 104.5393 | 96.0533 | 97.5526 | 107.1052 | 83.3289 | 76.5659 | 93.0133 | 80.3555 | 117.5393 | 98.1318 | 102.0786 | 77.237 | 237.7896 | 214.0919 | 219.5406 | 255.4075 | 173.6178 | 181.4075 | 200.1971 | 248.5646 | 322.8429 | 215.8162 | 239.526 | 337.7763
2461 | 29.0654 | 164.8642 | 174.7759 | 199.5959 | 154.4839 | 143.5722 | 147.7996 | 145.6401 | 148.584 | 153.056 | 121.8319 | 127.6401 | 142.5043 | 203.3481 | 146.7521 | 147.0119 | 188.056 | 167.3158 | 138.8761 | 123.8319 | 165.0119 | 145.8319 | 329.528 | 367.7403 | 159.292 | 251.7284 | 209.2479 | 218.236 | 263.5722 | 111.1239 | 115.528 | 115.8879 | 109.4278 | 140.764 | 134.2801 | 137.472 | 83.1121 | 258.5161 | 259.5365 | 245.9083 | 234.2801 | 130.528 | 118.7199 | 152 | 135.3481 | 83.3599 | 96.528 | 89.5722 | 110.8998 | 179.3718 | 196.3124 | 194.416 | 154.6519 | 124.236 | 143.0442 | 133.1121 | 105.3039 | 158.5399 | 152.7199 | 200.9881 | 159.8879 | 110.3158 | 148.18 | 172.0323 | 136.944 | 184.944 | 189.3243 | 206.8761 | 193.0679 | 153.3481 | 147.3243 | 218.8879 | 175.6519 | 205.7403 | 224.8523 | 189.4839 | 288.18 | 166.6961 | 172.7403 | 156.292 | 192.1239 | 145.7759 | 135.6757 | 125.3039 | 145.4041 | 162.9202 | 177.8761 | 168.9881 | 123.6961 | 108.82 | 106.5399 | 126.9881 | 134.8761 | 286.82 | 252.3039 | 239.2037 | 308.416 | 246.6638 | 260.1681 | 284.4957 | 301.9202 | 126.0238 | 78.8642 | 86.0679 | 126.0238 | 112.8879 | 96 | 103.6638 | 120.708 | 98.8319 | 123.292 | 100.0119 | 108.9558 | 78.6401 | 79.1681 | 91.9762 | 90.82 | 103.1002 | 104.0442 | 86.6519 | 82.6401 | 238.1002 | 203.528 | 253.0085 | 236.236 | 156.9677 | 170.4957 | 188.18 | 201.6078 | 355.584 | 239.3039 | 224.944 | 331.1002
2462 | 29.0764 | 153.1934 | 186.0069 | 206.262 | 170.2414 | 132.7037 | 174.5789 | 139.4897 | 141.2689 | 148.5103 | 127.0275 | 120 | 122.9451 | 197.7517 | 136.9725 | 160.9931 | 188.2277 | 186.2826 | 140.6899 | 132.5034 | 181.2277 | 159.4691 | 318.5652 | 334.2963 | 155.0137 | 215.7243 | 220.7654 | 219.7449 | 250.2208 | 117.5034 | 114.2483 | 123.2414 | 131.4691 | 143.2277 | 119.7449 | 139.9931 | 75.5103 | 249.7105 | 211.9382 | 211.4897 | 229.9245 | 127.262 | 132.262 | 150.5103 | 123.5446 | 90.7517 | 92.2689 | 73.9863 | 125.2826 | 200.7449 | 157.4622 | 199.7449 | 154.0412 | 126.2414 | 126.0069 | 126.0069 | 113.0206 | 162.0549 | 162.2895 | 184.0275 | 157.5584 | 130.0275 | 149.5172 | 142.0069 | 133.2963 | 186.7586 | 159.738 | 196.524 | 209.9519 | 157.4348 | 130.6487 | 215.5652 | 180.2551 | 179.4966 | 190.7586 | 196.0549 | 295.476 | 157.2414 | 149.7243 | 151.5172 | 197.262 | 152.786 | 161.2895 | 136.9931 | 132.5309 | 145.4622 | 173.2346 | 158.7311 | 114.9863 | 104.0069 | 115.0206 | 116.4828 | 127.0069 | 281.262 | 266.2277 | 268.2551 | 309.2757 | 258.7929 | 252.2414 | 302.5652 | 278.0069 | 143.0618 | 63.4691 | 91.7792 | 143.0618 | 121.2346 | 100.2208 | 119.0206 | 118.7792 | 108.7449 | 118.7654 | 113.9931 | 121.0275 | 76.4622 | 74.2208 | 65.7517 | 84.2689 | 84.2483 | 92.4691 | 91.2551 | 75.9931 | 226.2002 | 205.476 | 204.7037 | 242.4622 | 180.786 | 191.5446 | 190.0137 | 225.9794 | 343.5515 | 238.3307 | 230.7311 | 318.4554
2463 | 29.0874 | 172.5279 | 187.4298 | 192.401 | 168.5279 | 149.8731 | 146.5144 | 145.824 | 134.824 | 142.9577 | 109.5279 | 122.0846 | 148.3029 | 193.6125 | 150.5634 | 159.5702 | 196.176 | 172.5211 | 157.8799 | 135.9509 | 183.9645 | 173 | 296.7749 | 317.648 | 151.3452 | 232.2115 | 213.2183 | 221.7394 | 264.824 | 119.648 | 113.176 | 124.3943 | 142.6548 | 148.8308 | 129.5567 | 143.7817 | 70.7394 | 265.6971 | 239.7817 | 224.8595 | 250.2318 | 126.6903 | 129.0846 | 145.2183 | 106.6057 | 87.3943 | 81.9154 | 79.7394 | 113.6057 | 204.824 | 182.599 | 195.4856 | 135.2183 | 128.9577 | 122.4789 | 125.6057 | 106.8663 | 146.599 | 144.4789 | 174.6057 | 135.648 | 134.1556 | 142.7817 | 141.6057 | 113.5211 | 182.7394 | 176.2047 | 194.599 | 226.0491 | 177.401 | 169.0068 | 198.7259 | 179.3029 | 182.0423 | 189.3452 | 177.2115 | 314.599 | 163.0846 | 162.0423 | 148.9509 | 193.5634 | 135.176 | 137.7462 | 139.4789 | 120.5634 | 164.8663 | 174.0068 | 167.7817 | 119.6971 | 104.1269 | 106 | 123.2183 | 120.0914 | 276 | 268.7039 | 269.648 | 304.7749 | 238.4366 | 251.3097 | 276.6057 | 281.2538 | 121.2115 | 78.824 | 79.5211 | 121.2115 | 119.401 | 112.2183 | 115.9932 | 108.6057 | 113.0846 | 102.3587 | 112.8308 | 110.824 | 89.6125 | 89.0846 | 69.4298 | 81.2115 | 87.6057 | 101.8308 | 96.5567 | 77.9577 | 239.1405 | 218.8663 | 220.5702 | 259.2606 | 159.7885 | 175.3875 | 181.9154 | 234.2183 | 324.3875 | 206.3875 | 245.7749 | 334.0914
2464 | 29.0984 | 158.6372 | 206.5488 | 173.7302 | 153.8186 | 145.093 | 159.5427 | 154.0015 | 138.3628 | 144.093 | 105.7332 | 123.6341 | 148.4527 | 180.7256 | 160.1845 | 155.186 | 193.7287 | 170.4527 | 131.7271 | 145.6341 | 163.1814 | 164.5412 | 322.003 | 328.8186 | 150.9055 | 234.721 | 207.1799 | 221 | 264.5427 | 125.907 | 106.3628 | 124.0945 | 131.907 | 139.7287 | 147.9985 | 141.907 | 76.003 | 256.8155 | 242 | 252.4558 | 222.4588 | 142.4558 | 138.8201 | 137.8155 | 112.9085 | 83.0015 | 77.9101 | 82.2744 | 118.2713 | 210 | 201.5427 | 178.9101 | 136.5473 | 133.0945 | 127.2759 | 132.7271 | 114.1814 | 167.9985 | 151.7317 | 183.9101 | 150.093 | 167.4482 | 144.7271 | 145.9985 | 119.6387 | 182.8186 | 201.997 | 210.2683 | 214.4558 | 155.4558 | 156.6387 | 229.2713 | 184.0915 | 182.5442 | 202.5488 | 192.8186 | 321.8095 | 170.3643 | 169.9116 | 164.6372 | 201.8201 | 133.0015 | 128.4588 | 144.003 | 132.6402 | 169.7256 | 158.6372 | 166.7256 | 123.9131 | 114.6372 | 107.0915 | 116.907 | 109.8186 | 276 | 248.5457 | 271.814 | 325.6372 | 240.2759 | 233.0884 | 279.907 | 291.3598 | 132.7256 | 81.5442 | 79.0899 | 132.7256 | 106.1875 | 112.7287 | 127.8155 | 115.7271 | 117.6357 | 78.5473 | 102.0915 | 108.9055 | 79.7271 | 94.1814 | 78.9985 | 95.1814 | 96.9101 | 89.1814 | 119.0915 | 77.4558 | 226.279 | 219.0869 | 205.8171 | 263.0015 | 140.907 | 188.0015 | 179.2744 | 233.6372 | 343.0046 | 216.8186 | 249.4466 | 325.4558
2465 | 29.1094 | 159.293 | 214.3047 | 189.2852 | 155.1445 | 146.875 | 151.8438 | 157.1523 | 140.2735 | 145.0195 | 138.9961 | 113.9961 | 139 | 177.7031 | 169.7226 | 169.0156 | 193.5859 | 165.8477 | 130.7148 | 143.4102 | 168.1289 | 144.2813 | 334.293 | 335.2774 | 133.7109 | 198.5781 | 207.5508 | 223.8516 | 246.5781 | 118.9922 | 104.8516 | 139.293 | 120.1445 | 142.4375 | 137.7187 | 128.4336 | 89.4336 | 250.6992 | 236.582 | 255.8633 | 248.9922 | 142.1562 | 143.0117 | 124.5703 | 111.1406 | 84.7265 | 78.4375 | 85.582 | 113.2813 | 203.1562 | 186.4297 | 189.1328 | 142.0078 | 146.2969 | 136.3008 | 128.293 | 115.9922 | 163.9922 | 168.4414 | 187.5742 | 156.4375 | 137.1367 | 140.293 | 137.1445 | 127.7226 | 183.2891 | 181.4375 | 187.2852 | 215.0117 | 157.4375 | 169.2852 | 218.0078 | 183.8633 | 185.6953 | 212.0156 | 188.4414 | 275.7148 | 174.5703 | 178.7226 | 168.7148 | 203.4453 | 139.0039 | 147.0078 | 151.1601 | 155.8516 | 164.707 | 157.8672 | 153.4375 | 141.8633 | 116.1484 | 113.4219 | 101.1523 | 115.1367 | 273.7187 | 252.8516 | 260.5508 | 329.1445 | 259.5664 | 224.8477 | 263.2969 | 282.6914 | 120.0078 | 74.7148 | 68.8672 | 120.0078 | 131.4375 | 117.1484 | 114.8555 | 110.7226 | 110.293 | 96.2695 | 102.4336 | 94.8477 | 76.4336 | 86.0117 | 74.1367 | 93.5703 | 99.1484 | 88.7109 | 110.5937 | 82.5742 | 256.1484 | 195.4336 | 204.2774 | 255.3164 | 143.6875 | 194.2891 | 186.0039 | 231.4414 | 356.164 | 214.4375 | 213.1328 | 325.1562
2466 | 29.1204 | 150 | 190.2974 | 192.0821 | 160.7336 | 127.5694 | 158.2919 | 151.5439 | 151.5383 | 128.3285 | 132.3995 | 119.5949 | 143.1643 | 191.1077 | 156.323 | 146.0255 | 172.323 | 172.6459 | 135.3541 | 154.2097 | 178.2664 | 153.2719 | 320.8357 | 340.9434 | 140.9745 | 202.7903 | 225.9689 | 228.6204 | 247.5183 | 126.6204 | 111.1077 | 126.7281 | 123.6515 | 137.5439 | 136.7592 | 124.677 | 90.7336 | 261.6715 | 233.4107 | 248.5949 | 246.5072 | 127.9745 | 131.2719 | 125.1077 | 113.1077 | 74.6515 | 68.1898 | 75.5694 | 113.9434 | 190.1643 | 190.3541 | 194.8612 | 132.5949 | 131.1077 | 119.677 | 121.677 | 121.2408 | 172.5128 | 159.7336 | 180.1332 | 141.4306 | 137.8612 | 136.0566 | 142.4362 | 125.0311 | 182.2719 | 170.8923 | 188.2974 | 199.4051 | 150.4617 | 165.5383 | 202.051 | 181.6515 | 208.5128 | 182.7791 | 166.3485 | 280.949 | 173.9179 | 165.9179 | 167.1077 | 196.8724 | 137.2719 | 135.2153 | 137.626 | 156.1587 | 171.6204 | 149.4617 | 147.6515 | 130.7281 | 109.8923 | 122.5949 | 104.1698 | 116.4561 | 270.677 | 251.374 | 280.1587 | 334.1388 | 254.3995 | 236.7026 | 258.8158 | 309.323 | 109.7026 | 69.5383 | 62.2464 | 109.7026 | 115.8357 | 105.8357 | 111.6459 | 97.6204 | 99.5128 | 113.4306 | 98.677 | 99.8612 | 75.6515 | 71.2974 | 81.7026 | 93.8102 | 92.8923 | 100.1388 | 89.1388 | 81.3796 | 242.1587 | 197.0311 | 211.7281 | 221.0566 | 176.5128 | 191.7847 | 180.7026 | 220.6515 | 321.2664 | 203.2974 | 229.8668 | 315.1388
2467 | 29.1314 | 145.9734 | 195.6461 | 200.7169 | 170.9468 | 142.1948 | 135.4779 | 156.1591 | 145.0266 | 142.7344 | 102.5221 | 118.5221 | 146.1857 | 186.8318 | 148.4513 | 135.2831 | 167.8584 | 160.9292 | 140.876 | 129.885 | 162.4779 | 161.9292 | 319.0532 | 328.0708 | 150.1682 | 241.6727 | 208.239 | 222.6903 | 260.7078 | 115.7344 | 112.7169 | 123.6461 | 135.1416 | 147.7344 | 139.9468 | 122.637 | 98.4689 | 242 | 249.637 | 245.9734 | 226.9116 | 135 | 142.0974 | 128.5753 | 116.5753 | 81.4955 | 75.6461 | 85.5487 | 105.0974 | 198.1416 | 203.9292 | 182.0266 | 141.124 | 125.2831 | 122.2831 | 125.5221 | 107.5221 | 163.4247 | 173.0442 | 179.6727 | 131.761 | 128.4337 | 144.2831 | 164.9558 | 142.1416 | 187.5221 | 164.637 | 186.5221 | 199.8584 | 158.2123 | 153.4513 | 197.2922 | 191.9026 | 202.5221 | 164.0974 | 158.885 | 285.0884 | 167.4513 | 157.2831 | 158.1857 | 218.3715 | 148.4071 | 132.8584 | 146.8494 | 134.115 | 168.4779 | 165.2656 | 153.2565 | 121.1416 | 114.4779 | 118.115 | 139.0708 | 108.7435 | 270.8052 | 234.0974 | 267.0974 | 344.4955 | 238.1506 | 231.0442 | 283.8584 | 300.2123 | 110.5487 | 68.2922 | 79.7877 | 110.5487 | 113.7435 | 97.8584 | 101.7877 | 96.6461 | 101.885 | 105 | 104.0708 | 88.5753 | 84.9734 | 73.5487 | 79.4513 | 90.6903 | 101.8143 | 112.0442 | 104.1416 | 81.4247 | 228.1682 | 211.3539 | 207.7169 | 235.1682 | 168.6637 | 193.8318 | 177.2123 | 234 | 306.4071 | 205.239 | 230.9468 | 332.9292
2468 | 29.1424 | 147.305 | 210.5084 | 212.2203 | 156.2203 | 159.8644 | 150.3389 | 142.305 | 152.0678 | 130.678 | 107.949 | 111.0678 | 130.678 | 182.322 | 139.8475 | 129.5762 | 173.9322 | 160.1017 | 124.678 | 140.1017 | 165.1017 | 170.4406 | 333.4577 | 327.6102 | 155.9661 | 246.5086 | 213.5762 | 221.678 | 250.627 | 113.9829 | 124.5423 | 131.1017 | 131.9661 | 136.6441 | 131.3389 | 114.627 | 89.4914 | 248.1186 | 239.3728 | 234.7458 | 246.1017 | 135.6441 | 144.8136 | 138.4238 | 128.0339 | 69.8983 | 86.322 | 83.5255 | 107.1695 | 187.8814 | 200.2035 | 207.7456 | 147.3391 | 123.1186 | 115.2881 | 120.322 | 104.5762 | 153.8983 | 158.4577 | 200.9322 | 136.7287 | 145.6441 | 130.2035 | 167.9492 | 143.4745 | 181.3559 | 158.5592 | 181.322 | 208.5084 | 153.9322 | 148.3898 | 216.3898 | 181.7797 | 201.5084 | 168.4238 | 178.4406 | 261.0678 | 173.6609 | 158.0168 | 145.2542 | 182.2035 | 151.8136 | 127.3391 | 123.2881 | 126.1525 | 176.2542 | 171.8475 | 146.305 | 119.5762 | 117.4238 | 116.5931 | 130.2374 | 123.4238 | 267.305 | 249.0508 | 274.322 | 318.4069 | 253.9832 | 222.8983 | 279.6272 | 295.9322 | 121.0847 | 88.3559 | 79.8814 | 121.0847 | 128.1017 | 108.1186 | 107.678 | 107.322 | 115 | 103.7119 | 103.9322 | 104.2203 | 73.4238 | 81.1864 | 76.322 | 90.9661 | 112.1695 | 97.7797 | 93.5594 | 72.8644 | 236.2203 | 201.6441 | 210.2035 | 262.5421 | 164.1525 | 188.0339 | 171.6441 | 231.1017 | 305.6272 | 211.2203 | 211.3898 | 335.6441
2469 | 29.1534 | 166.9941 | 215.3195 | 216.3254 | 173.364 | 159.3077 | 163.9822 | 156.9793 | 146.997 | 129.6657 | 123.6449 | 116.6953 | 135.3491 | 186.6775 | 133.0119 | 141.3521 | 183.3491 | 160.3521 | 117.9822 | 137.0089 | 165.3521 | 175.3077 | 314.2988 | 322.642 | 162.0119 | 218.0059 | 217.6627 | 223.6746 | 277.3462 | 136.003 | 127.9881 | 122.3254 | 128.9852 | 129.9763 | 148.6598 | 134.6598 | 108.6331 | 257.6568 | 243.6272 | 234.3462 | 240 | 137.3343 | 138.0059 | 129.3225 | 123.6598 | 77.6716 | 82.3195 | 73.003 | 100.6775 | 184.6953 | 188.6686 | 237.3165 | 135.6923 | 132.6568 | 125.0207 | 129.358 | 107.6598 | 167.3551 | 150.3314 | 204.3314 | 150.6598 | 139.3106 | 129.0326 | 150.3314 | 150.3195 | 171.9763 | 182.3254 | 183.003 | 209.642 | 157.3314 | 136.3077 | 220.6894 | 172.9941 | 207.9911 | 178.3788 | 185.3136 | 263.6864 | 182.2988 | 171.3165 | 153.0089 | 178.358 | 141.3284 | 113.0178 | 117.642 | 136.6657 | 173.9793 | 159.997 | 165.6598 | 119.3165 | 115.6775 | 127.2928 | 124.0237 | 120.0059 | 279.9733 | 263.3254 | 266.642 | 300.3521 | 250.0444 | 230.6716 | 266.6805 | 311.7012 | 126.3047 | 80.3136 | 74.3551 | 126.3047 | 113.9763 | 118.3254 | 110.0089 | 105.6598 | 111.3225 | 105.6805 | 105.6598 | 114.6775 | 70.6746 | 84.3165 | 72.3195 | 91.3284 | 101.3314 | 100.361 | 89.361 | 78.3254 | 239.6568 | 209.0178 | 197.3314 | 283.5886 | 175 | 184.6627 | 175.6746 | 227.6746 | 288.6686 | 218.6686 | 209.3373 | 334.9941
2470 | 29.1644 | 160.9203 | 195.255 | 213.7729 | 189.1474 | 143.0398 | 152.6932 | 143.3466 | 134.6016 | 131.7729 | 108.3068 | 130.4542 | 142.1873 | 194 | 141.6932 | 150.8805 | 199.5458 | 169.1873 | 121.9442 | 138.494 | 174.8805 | 157.3068 | 302.4382 | 314.6653 | 162.0279 | 220.9602 | 213.9602 | 221.3745 | 269.7092 | 131.4144 | 121.0398 | 117.3466 | 122.1195 | 125.4382 | 138.1076 | 131.3864 | 85.3068 | 262.4382 | 228.4382 | 243 | 235.1474 | 129.3347 | 143.3864 | 129.6255 | 108.9482 | 80.3068 | 71.9602 | 77.4263 | 103.8407 | 188.4025 | 194.1593 | 224.6136 | 150.2271 | 119.4144 | 141.7729 | 138.3745 | 106.8127 | 179.2271 | 165.984 | 192.255 | 142.1873 | 130.9322 | 151.6932 | 152.4661 | 134.4144 | 165.0119 | 182.8924 | 186.0398 | 193.6932 | 158.0797 | 125.5856 | 223.1754 | 170.7331 | 189.1754 | 199.6415 | 181.3585 | 280.1195 | 155.5339 | 160 | 160.0398 | 192.2271 | 141.8127 | 121.8805 | 116.9442 | 134.6136 | 158.2669 | 161.4661 | 162.7331 | 111.8127 | 116.761 | 101.3864 | 130.2949 | 123.3068 | 274.1314 | 266.6653 | 254.506 | 316.8127 | 272.7211 | 238.8526 | 273.2271 | 327.0279 | 104.5737 | 74.2789 | 87.9602 | 104.5737 | 99.3864 | 116.4661 | 117.0398 | 100.3068 | 103.6534 | 110.494 | 102.0398 | 116.4542 | 78.4263 | 67.1076 | 69.5856 | 89.3864 | 105.1992 | 113.4542 | 108.3466 | 73.3466 | 246.8645 | 209.9084 | 197.3864 | 232.3864 | 171.5339 | 183.7331 | 180.6534 | 233.6932 | 306.2908 | 223.8127 | 217.1992 | 336.5458
2471 | 29.1754 | 146.5918 | 176.0767 | 206.0795 | 181.4356 | 145.7178 | 154.3589 | 147.9479 | 123.6137 | 138.7945 | 112.0246 | 113.1808 | 135 | 190.7699 | 140.8466 | 147.1534 | 208.2055 | 162.3589 | 147.3343 | 132.1534 | 172.9479 | 159.2301 | 309.285 | 326.3343 | 148.0767 | 215.7699 | 213.0767 | 194.7973 | 236.4877 | 120.0767 | 114.7452 | 121.2301 | 127.2822 | 144.1288 | 130.589 | 130.7699 | 85.4356 | 264.6192 | 243.8986 | 227.2083 | 234.2548 | 126.2794 | 141.6932 | 134.6685 | 94.3068 | 81.1534 | 74.6657 | 79.1288 | 93.8466 | 172.9972 | 201.6411 | 219.4877 | 146.7945 | 108.4356 | 141.2575 | 120.7699 | 109.6932 | 174.7178 | 183.285 | 174.5123 | 136.7945 | 145.0767 | 150.1288 | 156.8466 | 125.589 | 185.589 | 194.3589 | 186.9233 | 199.3068 | 160.2055 | 138.7178 | 202.9479 | 171.1288 | 162.1288 | 185.589 | 199 | 285.2822 | 161.5616 | 158.5644 | 165.589 | 196.3315 | 140.0274 | 128.5616 | 144.4877 | 135.9479 | 149.9754 | 166.2055 | 160.6165 | 120.7945 | 105.7178 | 101.8466 | 116.3068 | 114.463 | 285.5151 | 286.589 | 261.5644 | 325.0767 | 253.2575 | 253.7425 | 264.411 | 309.4877 | 105.3835 | 83.6932 | 81.3343 | 105.3835 | 105.9479 | 122.2822 | 117.5644 | 100.4356 | 106.2301 | 98.411 | 100.411 | 113.8959 | 82.2822 | 63.537 | 84.8712 | 94.1534 | 112.4877 | 101.9233 | 98.9507 | 74 | 252.9014 | 199.0493 | 197.8466 | 252.9452 | 170.7425 | 183.7699 | 184.6657 | 228.5397 | 333.0521 | 215.5672 | 226.6411 | 339.8219
2472 | 29.1864 | 147.6733 | 185.0527 | 191.5375 | 173.6066 | 147.3712 | 159.0831 | 153.144 | 140.856 | 137.1745 | 115.4321 | 109.5375 | 139.0831 | 190.9391 | 133.2881 | 138.0082 | 201.6593 | 168.5679 | 133.0609 | 145.2799 | 174.8033 | 163.5152 | 300.4461 | 324.3103 | 155.9391 | 209.6288 | 216.856 | 187.8478 | 229.4016 | 123.4848 | 107.7951 | 125.5152 | 128.9696 | 140.1136 | 130.6897 | 123.5152 | 89.856 | 239.6511 | 228.5457 | 201.5984 | 243.8033 | 140.3489 | 135.1136 | 126.9087 | 107.1967 | 80.5457 | 86.3407 | 84.3934 | 98.5375 | 200.9696 | 196.5457 | 211.6593 | 147.0304 | 112.1136 | 125.0609 | 120.1967 | 112.0222 | 179.3407 | 161.0831 | 182.3407 | 148.1663 | 148.4848 | 142.0304 | 161.1663 | 125.6897 | 175.6675 | 181.637 | 193.9087 | 199.9473 | 154.4016 | 142.5984 | 209.6288 | 167.1136 | 168.507 | 197.1967 | 193.4321 | 289.1967 | 185.4848 | 162.3103 | 168.6593 | 212.2576 | 131.7951 | 137.637 | 140.3712 | 142.6288 | 145.0831 | 160.7728 | 156.5679 | 124.3712 | 110.3407 | 102.4543 | 119.9169 | 113.2494 | 266.856 | 281.1218 | 253.8033 | 336.6511 | 247.8255 | 248.4098 | 255.4016 | 293.1218 | 118.7119 | 78.5984 | 81.5375 | 118.7119 | 110.0304 | 120.2576 | 113.8864 | 100.7728 | 111.6288 | 96.0831 | 98.0831 | 117.9778 | 83.5984 | 75.5152 | 89.2576 | 97.2576 | 108 | 91.0913 | 86.5679 | 75.856 | 237.7869 | 206.2354 | 199.5679 | 274.9778 | 170.9778 | 178.3712 | 186.3185 | 226.2799 | 332.6815 | 205.8173 | 222.6593 | 327.7424
2473 | 29.1974 | 178.3156 | 186.6459 | 211.7523 | 157.5339 | 152.9853 | 156.0294 | 146.1651 | 147.4514 | 127.8495 | 112.1358 | 122.8495 | 140.2477 | 202.1505 | 132.3688 | 126.7376 | 187.1798 | 172.2477 | 138.7229 | 149.8789 | 161.3982 | 167.5193 | 331.5339 | 331.1651 | 164.8495 | 217.8202 | 220.3835 | 210.6991 | 244.9412 | 122.5486 | 113.7138 | 126.4514 | 136.6844 | 141.6165 | 123.4514 | 121 | 83.7964 | 259.4514 | 230.5871 | 210.9853 | 235 | 132.4367 | 135.0826 | 136.6312 | 115.6165 | 76.0679 | 81.6459 | 96.2477 | 111.8495 | 201.7817 | 189 | 210.9853 | 142.3835 | 119.3688 | 123.8202 | 127.8495 | 119.0294 | 178.4807 | 170.3009 | 182.2477 | 153.1798 | 148.6991 | 126.6459 | 169.2477 | 118.8349 | 157.9706 | 171.655 | 205.1651 | 187.6165 | 150.767 | 152.7523 | 218.2036 | 165.9321 | 189.9174 | 199.8642 | 185.1505 | 286.112 | 185.3156 | 163.4129 | 168.3688 | 209.8495 | 140.3982 | 119.6018 | 136.3982 | 147.3688 | 159.6697 | 160.0679 | 161.7817 | 126.1505 | 115.233 | 108.8495 | 120.6697 | 126.3303 | 292.6256 | 266.3688 | 249.6018 | 351.8495 | 268.2715 | 237.7229 | 240.2624 | 283.7376 | 126.1505 | 80.6991 | 87.9468 | 126.1505 | 114.2036 | 113.2477 | 109.3156 | 106.2036 | 104.0973 | 107.2183 | 105.3835 | 107.1358 | 84.9321 | 74.9174 | 81.4807 | 89.8642 | 114.1358 | 87.888 | 93.3156 | 78.233 | 269.1651 | 197.2036 | 210.1505 | 264.1358 | 165.1211 | 175.9321 | 186.888 | 230.112 | 339.3982 | 232.5486 | 214.3156 | 335.9027
2474 | 29.2084 | 179.5409 | 170.3958 | 207.5436 | 158.0211 | 153.4802 | 143.1662 | 143.562 | 150.6253 | 127.9789 | 120.0211 | 119.4169 | 127.8338 | 198.8549 | 140.2084 | 142.0211 | 179.7282 | 167.3535 | 147.7732 | 137.9366 | 153.6042 | 181.1662 | 330.8338 | 313.9183 | 161.0211 | 222.1056 | 209.9183 | 204.6253 | 253.2534 | 122.1451 | 111.4986 | 125.2718 | 139.0211 | 145.3535 | 131.7704 | 125.3535 | 75.3324 | 264.6042 | 247.5831 | 222.9578 | 229.0634 | 144.0211 | 130.0211 | 134.7282 | 109.4591 | 83.3747 | 71.7282 | 85.0211 | 105.2507 | 189.8549 | 185.0422 | 205.1478 | 137.8549 | 126.0211 | 130.0845 | 128.7704 | 104.5831 | 166.8127 | 165.2929 | 186.0606 | 144.1451 | 137.876 | 110.7916 | 160 | 121.438 | 172.4169 | 189.2296 | 210.0817 | 195.7071 | 145.2718 | 153.6887 | 225.876 | 161.3958 | 191.0211 | 188.0211 | 188.9789 | 280.0606 | 171.5014 | 162.6437 | 180.1662 | 210.7704 | 136.562 | 126.6042 | 129 | 153.6253 | 166.0634 | 167.3747 | 154.2084 | 131.1662 | 107.2718 | 109.3747 | 123.5014 | 123.8944 | 313.1689 | 275.7916 | 247.1056 | 355.5409 | 269.0027 | 257.8549 | 230.9155 | 308.5198 | 119.6887 | 69.0845 | 72.6253 | 119.6887 | 118.314 | 114.2902 | 107.3747 | 119.4169 | 99.3324 | 110.438 | 117.4775 | 117.3958 | 75.6465 | 74.438 | 74.562 | 78.8127 | 114.1056 | 107.7916 | 94.9366 | 78.5831 | 245.9816 | 199.3352 | 206.0634 | 273.2084 | 172.7098 | 177.3324 | 195.314 | 216.5409 | 335.562 | 230.9578 | 211.9789 | 341.0634
2475 | 29.2194 | 174.5338 | 183.6501 | 193.9371 | 165.6098 | 148.9371 | 151.2645 | 162.0582 | 144.5516 | 140.7936 | 126.3856 | 118.6323 | 130.7533 | 192.6323 | 133.6951 | 136.9597 | 184.2871 | 169.5113 | 126.6726 | 142.5113 | 165.6501 | 190.0807 | 326 | 309.242 | 149.4306 | 219.242 | 190.1435 | 194.879 | 228.4259 | 121.8387 | 123.5516 | 111.7355 | 129.4709 | 147.1032 | 131.9193 | 134.4484 | 91.5291 | 254.6144 | 245.9193 | 222.4709 | 225.3049 | 138.1435 | 136.7936 | 139.2871 | 113.242 | 87 | 82.8161 | 81.1839 | 104.9371 | 189.7533 | 175.7355 | 194.2017 | 143.0582 | 128.7129 | 120.4484 | 128.9193 | 112.3049 | 155.6144 | 154.0403 | 197.8387 | 145.0629 | 128.6323 | 118.5291 | 157.1435 | 120.0807 | 172.8565 | 178.4709 | 224.5516 | 203.6951 | 143.9775 | 145.4887 | 222.7533 | 162.4081 | 188.4081 | 185 | 189.9597 | 282.8612 | 172.9146 | 179.6951 | 180.1032 | 219.8968 | 142 | 128.0403 | 130.6323 | 144.6951 | 156.5919 | 159.5741 | 166.8743 | 129.4709 | 103.5291 | 114.6323 | 113.8968 | 127.9822 | 289.8161 | 283.9371 | 242.2017 | 346.4531 | 234.6951 | 242.6548 | 243 | 306.9418 | 115.9775 | 64.7533 | 66.9597 | 115.9775 | 106.4887 | 121.0225 | 104.879 | 112.1032 | 106.9597 | 105 | 121.5338 | 108.2064 | 77.9775 | 73.8968 | 80 | 81.4887 | 100.2242 | 101.2467 | 101.9597 | 82.2242 | 225.3049 | 182.4484 | 205.5694 | 262.2064 | 158.6323 | 177.2064 | 186.3452 | 227.4484 | 336.5113 | 227.6144 | 214.1839 | 334.4484
2476 | 29.2304 | 150.3761 | 189.3892 | 209.0443 | 166.7078 | 153.9557 | 149.2743 | 168.3892 | 138.8982 | 149.5353 | 135.3761 | 124.7832 | 145.7832 | 203.4072 | 127.6814 | 134.4204 | 175.5796 | 160.0575 | 134.9425 | 132.6371 | 173.9114 | 193.0575 | 321.3761 | 321.5353 | 142.0443 | 237.4204 | 201.1329 | 204.9161 | 247.1018 | 112.1018 | 118.7389 | 110.3629 | 121.6814 | 145.0443 | 130.2036 | 131.6946 | 88.3892 | 251.509 | 240.8407 | 220.9868 | 239.3054 | 136.1018 | 145.115 | 130.5796 | 128.8982 | 81.115 | 85.6814 | 77.4778 | 111.2168 | 206.0443 | 176.0443 | 210.6814 | 151.491 | 130.2743 | 127.3629 | 117.9557 | 123.3629 | 153.3497 | 158.2611 | 183.0575 | 151.394 | 135.2036 | 132.6239 | 161.8275 | 126.4204 | 177.4204 | 167.7389 | 207.5485 | 203.5665 | 159.9425 | 144.0132 | 223.491 | 181.0754 | 191.1018 | 165.2432 | 187 | 246.23 | 188.0132 | 164.8539 | 181.8275 | 218.1725 | 126.0264 | 131.8407 | 135.5222 | 137 | 156.8407 | 158.1329 | 169.3449 | 117.4778 | 115.1018 | 117.4204 | 119.3186 | 118.3054 | 285.5353 | 293.1593 | 261.6239 | 328.5665 | 236.2479 | 227.7964 | 248.885 | 288.8407 | 123.5353 | 75.5796 | 63.5796 | 123.5353 | 114.2611 | 107.5796 | 106.509 | 115.9293 | 103.1593 | 105 | 102.8407 | 99.885 | 80.0707 | 74.2743 | 77.4778 | 85.8982 | 98.2168 | 95.4647 | 101.5222 | 78.9557 | 240.1461 | 194.8275 | 215.4778 | 251.7832 | 153.8539 | 170.9868 | 192.3761 | 234.7832 | 337.9868 | 215.6814 | 219.3054 | 331.6946
2477 | 29.2414 | 154.8162 | 191.0366 | 200.8573 | 151.7591 | 140.9429 | 149.0856 | 151.4326 | 140.7591 | 132.4817 | 119.0491 | 122.645 | 143.2123 | 210.2409 | 136.4897 | 130.6735 | 184.5183 | 162.4897 | 147.2203 | 133.1918 | 172.1142 | 176.8858 | 322.7877 | 347.7467 | 151.5959 | 231.0776 | 214.2123 | 218.0205 | 259.5183 | 125.3836 | 123.4897 | 118.0285 | 125.7306 | 138.1552 | 142.0571 | 126.4612 | 71.4326 | 256.4817 | 235.0776 | 238.9224 | 262.2774 | 146.7877 | 133.5388 | 135.1918 | 114.3185 | 80.355 | 94.4897 | 77.0285 | 103.8368 | 206.0776 | 185.1632 | 210.8368 | 145.653 | 125.3265 | 125.9429 | 111 | 121.5103 | 162.9429 | 152.645 | 183.7591 | 153.2328 | 130.6164 | 137.2694 | 157.5593 | 121.3756 | 169.7797 | 170.3265 | 197.653 | 218.7591 | 155.7797 | 141.653 | 209.8653 | 199.0776 | 191.4041 | 150.5468 | 199.5468 | 254.661 | 177.8653 | 165.3836 | 183.6164 | 207.2979 | 104.8653 | 131.2694 | 129.0491 | 145.653 | 171.4121 | 176.8368 | 146.7306 | 122.653 | 116.7021 | 118.8653 | 122.1918 | 131.3265 | 285.355 | 284.2123 | 281.4121 | 349.3836 | 254.6244 | 227.6244 | 247.4817 | 292.1632 | 119.4612 | 75.8653 | 69.0571 | 119.4612 | 119.8938 | 107.8653 | 125.3265 | 126.6735 | 98.9715 | 113.653 | 101.4041 | 106.7021 | 77.5468 | 69.3265 | 77.4612 | 90.355 | 100.7877 | 98.6735 | 97.645 | 81.9509 | 233.992 | 205.2694 | 215.4612 | 263.4897 | 150.9224 | 189.355 | 188.1632 | 233.0776 | 345.9715 | 218.4326 | 217.6164 | 328.6244
2478 | 29.2524 | 171.6696 | 215.1101 | 202.8021 | 153.1057 | 135.6696 | 154.3259 | 148.8854 | 138.1012 | 131.5685 | 114.0089 | 120.119 | 149.0179 | 196.436 | 139.2158 | 136.5685 | 201.0045 | 167.4405 | 142.4271 | 136.8854 | 172.872 | 171.2336 | 346.3482 | 369.753 | 163.0089 | 218.4405 | 204.8899 | 200.1101 | 275.5595 | 127.8765 | 116.872 | 118.8854 | 132.0045 | 135.5729 | 140.2113 | 133.6696 | 84.0134 | 252.8988 | 241.5729 | 243.9955 | 252.5238 | 148.1012 | 133.0045 | 137.5506 | 104.3393 | 85.1057 | 98.1057 | 84.5595 | 107.6741 | 201.8943 | 185.3304 | 201.7708 | 141.4271 | 126.5506 | 124.2292 | 124.3482 | 115.2247 | 155 | 149.6741 | 178.8765 | 134.0089 | 124.564 | 134.5551 | 141.3393 | 128.683 | 170.5685 | 172.4405 | 193.4271 | 215.6562 | 139.6607 | 144.1012 | 202.9911 | 174.4271 | 191.1146 | 173.6741 | 196.8676 | 285.6652 | 170.1012 | 171.436 | 176.6741 | 217.0089 | 120.2381 | 134.7842 | 119.5595 | 152.5506 | 180.1012 | 158.872 | 143.6607 | 128.2158 | 114.1101 | 112.436 | 120.5461 | 132.9955 | 301.6741 | 276.2247 | 296.3304 | 354.9911 | 253.9911 | 247.0134 | 246.1235 | 294.5551 | 121.7753 | 71.2158 | 75.6652 | 121.7753 | 108.9821 | 112.5595 | 130.5551 | 124.1146 | 85.2113 | 116.1012 | 112.2381 | 106.3348 | 94 | 75.8899 | 87.3393 | 97.7753 | 99.8765 | 96.1146 | 89.7798 | 83.8765 | 235.247 | 199.4405 | 220.4449 | 265.3259 | 152.436 | 186.9866 | 194.5595 | 222.6652 | 330.8765 | 229.2336 | 203.1101 | 316.8676
2479 | 29.2634 | 161.7403 | 205.7698 | 212.168 | 165.7465 | 143.5723 | 147.8289 | 145.9145 | 138.888 | 148.2862 | 111.6547 | 125.7994 | 147.9673 | 183.3421 | 132.9145 | 154.2006 | 205.0855 | 160.7138 | 131.2302 | 133.4572 | 151.7138 | 173.9409 | 349.7403 | 338.7994 | 164.7698 | 212.1711 | 205.5428 | 202.6579 | 252.1089 | 119.944 | 107.6019 | 121.4012 | 139.7434 | 144.9409 | 144.4603 | 129.2271 | 86.1975 | 264.4572 | 247.7403 | 220.7107 | 220.6579 | 137.9145 | 127.4837 | 125.1415 | 111.4572 | 88.1446 | 86.056 | 80.7698 | 111.8849 | 198.8554 | 187.0295 | 194.4572 | 137.5459 | 117.7994 | 132.9145 | 132.3126 | 114.3421 | 166.888 | 152.5132 | 179.6314 | 139.8849 | 134 | 123.9409 | 146.1711 | 149.2862 | 182.7138 | 159.7698 | 175.3717 | 193.4837 | 143.9736 | 140.3157 | 201.6874 | 160.4868 | 188.5988 | 192.5163 | 173.4572 | 292.6874 | 175.9176 | 179.3748 | 187.2862 | 223.3421 | 126.5692 | 140.6283 | 133.1446 | 136.9409 | 172.2006 | 153.2597 | 149.8025 | 109.5692 | 114.8289 | 115.3453 | 116.4012 | 127.5428 | 321.888 | 293.6314 | 276.5397 | 371.0622 | 245.3717 | 249.1975 | 267.7729 | 294 | 118.0855 | 71.7729 | 71.2566 | 118.0855 | 91.7434 | 107.8554 | 113.0824 | 123.8849 | 85.8025 | 106.8289 | 119.0264 | 101.5988 | 86.2271 | 80.2006 | 89.8849 | 95.9145 | 96.0591 | 91.3126 | 96.1151 | 73.2006 | 266.6579 | 202.3157 | 215.9705 | 251.5132 | 141.1711 | 179.3157 | 194.4277 | 216.4277 | 334.3748 | 226.4541 | 203.3717 | 305.8025
2480 | 29.2744 | 146.939 | 194.5914 | 195.6952 | 171.0972 | 138.6705 | 156.5733 | 151.2257 | 147.3657 | 143.0181 | 118.598 | 116.8353 | 129.5733 | 185.5733 | 138.6952 | 149.5486 | 203.061 | 164.5733 | 137.7133 | 138.6952 | 150.8781 | 163.4086 | 335.8781 | 338.7562 | 155.4695 | 223.1466 | 200.3048 | 203.7133 | 243.4152 | 114.3229 | 119.4695 | 120.0791 | 137.3657 | 138.1647 | 145.2191 | 120.939 | 82.2685 | 266.4086 | 226.3657 | 220.1285 | 217.4876 | 134.7743 | 132.8419 | 125.4514 | 113.4086 | 89.14 | 87.9209 | 70.061 | 120.1466 | 188.4695 | 190.6524 | 195.939 | 152.7133 | 114.4695 | 133.5305 | 125.7562 | 117.5124 | 171.6095 | 144.5914 | 201.2257 | 141.5733 | 141.5124 | 119.0428 | 153.3905 | 142.14 | 175.3048 | 162.2076 | 179.3476 | 205.8847 | 146.6705 | 140.5486 | 219.5124 | 170.2867 | 182.8781 | 200.7743 | 180.1038 | 283.28 | 186.14 | 186.8534 | 174.5058 | 233.0857 | 120.5124 | 146.0428 | 137.4267 | 123.5914 | 175.061 | 162.8962 | 174.86 | 94.1219 | 116.061 | 123.8962 | 123.061 | 115.2619 | 325.201 | 299.7315 | 264.0428 | 381.8105 | 245.1219 | 233.061 | 262.9143 | 294.4695 | 122.6343 | 74.8962 | 65.1829 | 122.6343 | 104.3905 | 100.2867 | 108.0247 | 103.9753 | 100.061 | 106.6524 | 110.2257 | 108.5552 | 81.6952 | 83.061 | 83.1219 | 86.201 | 108 | 85.6952 | 93.0181 | 83.1038 | 250.341 | 215.2257 | 212.3476 | 251.1038 | 147.9209 | 178.14 | 183.3657 | 225.0857 | 337.6277 | 201.4267 | 212.5124 | 321.939
2481 | 29.2854 | 160.5272 | 189.6273 | 214.4908 | 164.5272 | 135.1908 | 164.9636 | 157.4454 | 139.5546 | 140.2636 | 125.1729 | 111.9636 | 137.9636 | 181.9183 | 142.5546 | 130.8455 | 206.8181 | 165.2546 | 133.5182 | 134.8455 | 163.1181 | 170.7818 | 335.5909 | 351.1908 | 162.2636 | 240.8909 | 195.9636 | 199.0364 | 257.0274 | 119.7544 | 117.5909 | 116.7454 | 132.4454 | 144.4818 | 131.709 | 120.5546 | 88.2546 | 255.9547 | 229.6363 | 234.8455 | 228.7544 | 128.5546 | 146.2546 | 138.3727 | 116.4454 | 86.1908 | 100.4091 | 80.0817 | 121.0274 | 201.0453 | 181.7364 | 192.6637 | 150.4454 | 121.2636 | 133.9636 | 116.5092 | 117.3274 | 153.291 | 149.2636 | 194.9183 | 144.6637 | 133.1365 | 132.7818 | 158.7 | 131.4818 | 166.6273 | 191.418 | 171.4001 | 213.282 | 142.2272 | 149.7907 | 219.8092 | 170.6273 | 196.0817 | 183.4728 | 186.5182 | 277.3817 | 189.9362 | 167.6637 | 164.2272 | 230.6911 | 125.1455 | 153.5182 | 126.1455 | 133.5635 | 160.991 | 152.6273 | 169.6183 | 106.4544 | 109.7 | 126.1545 | 126.8181 | 109.709 | 307.2182 | 301.9362 | 267.6637 | 364.3 | 242.0364 | 226.2182 | 234.4728 | 304.1545 | 120.3274 | 69.4454 | 69.2272 | 120.3274 | 114.0364 | 107.3727 | 112.4728 | 94.2272 | 110.0817 | 108.3363 | 101.991 | 109.1001 | 82.6637 | 79.5909 | 81 | 93.2725 | 110.4091 | 99.1908 | 92.1908 | 81.3274 | 232.7818 | 220.9636 | 203.4365 | 249.8092 | 150.7728 | 170.3727 | 191.9362 | 230.8819 | 323.3 | 208.4544 | 220.5182 | 318.1819
2482 | 29.2964 | 174 | 195.388 | 209.7299 | 167.6238 | 150.9175 | 147.224 | 159.4941 | 141.9528 | 141.0707 | 110.941 | 118.9293 | 127.1415 | 192.2347 | 134.0825 | 133.8467 | 206.0707 | 163.9293 | 141.8939 | 119.5648 | 168.5648 | 170.0825 | 343.8703 | 343.2004 | 159.1179 | 251.4116 | 190.5766 | 207.8821 | 239.0354 | 128.6474 | 116.9764 | 123.5059 | 130.5413 | 138.5766 | 140 | 127.8939 | 81.9882 | 259.3055 | 236.0472 | 240.3172 | 231.7181 | 136.8821 | 139.9882 | 129.1533 | 108.612 | 95 | 97.0707 | 89.5295 | 107.4823 | 199.1769 | 192.2937 | 201.8349 | 148.0472 | 127.5059 | 128.0825 | 118.3055 | 110.4705 | 160.8114 | 150.5648 | 182.5059 | 137.0354 | 127.8467 | 133.5648 | 157.4469 | 137.4352 | 170.9057 | 184.7771 | 187.6699 | 204.8349 | 146.0472 | 155.1769 | 220.8821 | 174.4116 | 186.7535 | 170.4941 | 184.5177 | 276.7181 | 205.0118 | 172.388 | 177.9293 | 226.3172 | 130.8821 | 143.1179 | 127.9293 | 142.0472 | 136.1297 | 156.9057 | 167.3172 | 133.8585 | 114.8703 | 125.1297 | 125.5766 | 130.3526 | 317.7996 | 287.8596 | 284.2465 | 368.5177 | 262.7406 | 222.9646 | 230.8821 | 290.7771 | 122.8585 | 74.4587 | 73.5413 | 122.8585 | 119.4234 | 109.0236 | 117.2819 | 109.4116 | 121.0118 | 112.5059 | 97.8821 | 92.4941 | 84.9175 | 79.9646 | 76.059 | 107.5766 | 101.1415 | 104.0472 | 106.4352 | 82.3762 | 232.5766 | 201.7417 | 198.4116 | 235.0707 | 168.1993 | 188.329 | 199.6002 | 213.0825 | 342.3408 | 219.059 | 220.9882 | 332.2701
2483 | 29.3074 | 176.532 | 195.3785 | 186.4552 | 161.0512 | 148.8849 | 148.7622 | 151.8977 | 141.9488 | 135 | 121.064 | 114.3657 | 117.0256 | 210.9488 | 132.5704 | 149.532 | 202.0256 | 163.9232 | 133.7825 | 120.0768 | 152.3657 | 174.1407 | 355.5064 | 335.6215 | 146.9744 | 233.6802 | 203.7494 | 208.3785 | 241.5704 | 116 | 129.1279 | 126.5448 | 121.936 | 145.1663 | 133.9232 | 136.4936 | 82.4936 | 261.3145 | 217.3145 | 242.3273 | 225.1535 | 144.468 | 135.936 | 132.7111 | 92.9232 | 84.8721 | 76.8209 | 78.8977 | 107.4808 | 188.0512 | 200.8849 | 208.4041 | 146.532 | 137.1279 | 127.0768 | 129.4296 | 122.6215 | 161.3017 | 152.5959 | 181.4936 | 142.6087 | 136.4424 | 129.5192 | 164.0256 | 142.4936 | 192.6727 | 160.4936 | 199.7953 | 204.3145 | 144.0256 | 143.5448 | 218.3401 | 185.0384 | 193.3582 | 178.0895 | 195.6599 | 263.5704 | 195.8977 | 178.4552 | 178.936 | 226.3017 | 134.9232 | 130.9744 | 138.5576 | 134.4552 | 140.6983 | 162.468 | 172.8721 | 136.3785 | 130.5576 | 116.0256 | 126.5959 | 149.0768 | 325.3785 | 270.6727 | 308.532 | 384.2175 | 264.7442 | 243.7239 | 254.1663 | 281.1791 | 129.4296 | 81.5448 | 67.9744 | 129.4296 | 118.4041 | 105.4808 | 126.8849 | 108.3913 | 124.0512 | 101.8721 | 113.5704 | 107.6855 | 92 | 75.4041 | 73.532 | 92.8977 | 92.5448 | 101.0128 | 96.3017 | 89.9616 | 224.9872 | 194.1663 | 206.5064 | 227.9872 | 184.3785 | 187.8081 | 181.3785 | 212.5831 | 349.4169 | 209.9488 | 206.3017 | 342.8849
2484 | 29.3184 | 174.8505 | 186 | 180.4066 | 154.6263 | 138.9626 | 160.701 | 145.5561 | 140.5934 | 132.9253 | 117.701 | 110.2242 | 135.598 | 205.9626 | 136.4439 | 156.6682 | 191.5515 | 169.4111 | 127.3737 | 141.5606 | 146.6682 | 192.2616 | 357.5561 | 337.2197 | 160.0419 | 215.2242 | 222.4813 | 207.8177 | 253.7429 | 131.0419 | 123.4394 | 118.0702 | 125.299 | 146.5889 | 131.6308 | 131.3318 | 89.2616 | 253.2242 | 216.4859 | 238.8551 | 226.6263 | 138.3692 | 131 | 129.8086 | 98.4111 | 81.7429 | 75.4485 | 74.1121 | 108.0747 | 198.2242 | 182.1449 | 193.7384 | 147.9626 | 137.1449 | 127.8131 | 123.4813 | 127.3318 | 155.3364 | 153.2571 | 186.1869 | 159.299 | 130.9626 | 127.8879 | 161.3318 | 141.4813 | 176.4348 | 161.0374 | 195.4111 | 202.9672 | 154.299 | 159.9672 | 224.154 | 182.2944 | 211.1449 | 180.8505 | 199.1823 | 257.0702 | 185.4066 | 182.7803 | 181.2616 | 225.0045 | 138.8551 | 144.0419 | 147.6682 | 137.2242 | 149.2571 | 156.3692 | 156.7758 | 127.5187 | 133.4439 | 120.5934 | 124.6636 | 144.6263 | 311.8505 | 289.7056 | 305.2944 | 392.1823 | 256.4111 | 236.103 | 236.3975 | 302.7803 | 122.4439 | 81.3692 | 74.8177 | 122.4439 | 108.4066 | 108.1495 | 114.3692 | 106.7429 | 116.5889 | 95.1121 | 109.6636 | 113.1823 | 82.6636 | 83.5606 | 79.1121 | 85.5187 | 92.3692 | 95.2571 | 88.2616 | 79.7384 | 218.2944 | 197.6636 | 215.299 | 232.1869 | 162.5515 | 168.3318 | 175.1121 | 212.7758 | 339.3692 | 199.7758 | 210.7101 | 328.8131
2485 | 29.3294 | 163.5663 | 179.0973 | 193.9293 | 162.5223 | 133.7522 | 147.1592 | 144.8761 | 141.4071 | 125.6902 | 117.4337 | 111.2212 | 136.6636 | 184.2919 | 142.4337 | 154.6283 | 185.8408 | 172.0353 | 138.5929 | 134.2299 | 162.6815 | 189.9734 | 347.8495 | 348.5842 | 158.6016 | 212.5044 | 219.3451 | 225.5576 | 253.0973 | 134.9114 | 114.3098 | 105.4071 | 124.5044 | 141.8408 | 125.4424 | 138.6815 | 98.1239 | 249.9734 | 229 | 239.5044 | 215.938 | 132.3451 | 133.6549 | 126.3364 | 113.7788 | 83.7522 | 82.7522 | 73.8141 | 106.8141 | 196.0353 | 187.8674 | 202.3984 | 151.2478 | 125.8761 | 119.8141 | 126.7168 | 112.9114 | 154.9734 | 152.8408 | 181.4424 | 143.637 | 131.062 | 125.531 | 154.3451 | 134.6283 | 165.4603 | 158.8141 | 199.0973 | 215 | 155.6283 | 155.0087 | 234.0973 | 177 | 201.469 | 171.6902 | 191 | 269.894 | 170.7875 | 163.4511 | 184.2832 | 222.0707 | 152.7788 | 144.7256 | 138.7256 | 133.9734 | 152.5576 | 165.7435 | 162.1505 | 113.1326 | 127.7522 | 127.2478 | 114.938 | 133.4071 | 318.0886 | 287.0353 | 304.2478 | 374.9734 | 249.4777 | 213 | 235.6109 | 283.4511 | 119.938 | 72.1592 | 75.0353 | 119.938 | 105.469 | 113.062 | 103.5663 | 113 | 101.7522 | 109.6815 | 103.1239 | 114.0266 | 79.8408 | 92.6902 | 78.2832 | 80.6902 | 93.7788 | 96.9647 | 89.1592 | 74.5929 | 209.8141 | 185.2832 | 221.4071 | 238.062 | 160.5576 | 177.2744 | 190.7435 | 218.1505 | 330.6902 | 201.9647 | 213.6016 | 331.4337
2486 | 29.3404 | 168.9519 | 182.7788 | 204.1971 | 168.2212 | 140.8654 | 138.7404 | 142.4567 | 137.2837 | 135.6683 | 122.3702 | 99.9375 | 135.4086 | 184.4711 | 166.3846 | 149 | 180.1346 | 162.2837 | 137.2837 | 128.0385 | 158.875 | 182.5433 | 334.1971 | 357.3942 | 146.0865 | 213.6923 | 212.6538 | 230.7404 | 247.5433 | 124.3702 | 118.4567 | 118.6683 | 122.4327 | 139.9375 | 136.0144 | 150.6298 | 101.6298 | 240.9135 | 221.9375 | 236.3462 | 222.0625 | 139.7788 | 127.3077 | 141.1731 | 120.1731 | 84.8894 | 89.8654 | 69.3702 | 101.2837 | 193.3462 | 201 | 232.2981 | 156.0865 | 130.5192 | 123.5192 | 122.3942 | 106.7163 | 164.3846 | 154.7404 | 184.4086 | 140.9279 | 131.4567 | 119.4808 | 157.4327 | 139.3221 | 164.7885 | 151.6538 | 207.6683 | 207.3942 | 153.8029 | 134.5673 | 221.4808 | 178.0865 | 186.3317 | 182.7548 | 189.9135 | 264.3798 | 173.5817 | 165.5336 | 177.1971 | 209.0865 | 150.4808 | 138.4327 | 124.8269 | 131.976 | 161 | 171.024 | 167.6538 | 103.8029 | 120.7404 | 126.6538 | 121.0625 | 127.6538 | 304.726 | 289.2356 | 299.3077 | 371.3462 | 249.7548 | 237.9904 | 247.5048 | 261.0865 | 119 | 76.7788 | 70.7163 | 119 | 99.5673 | 102.0481 | 113.8413 | 118.976 | 103.4327 | 111.851 | 106.0865 | 114.3942 | 83.3702 | 87.4567 | 76.6298 | 74.3702 | 101.8029 | 106.1731 | 88.8894 | 81.4327 | 227.101 | 191.7788 | 211.851 | 233.5673 | 168.4567 | 181.3942 | 206.3462 | 223.6538 | 337.9519 | 205.1971 | 202.7163 | 347.2356
2487 | 29.3514 | 185.1112 | 199.889 | 191.5555 | 166.1112 | 148.3333 | 143.2223 | 145.3334 | 137.2222 | 159.1112 | 125.4445 | 102.889 | 149.4445 | 201.6667 | 164.8888 | 138.9999 | 176.5556 | 178.3335 | 139.4445 | 145.6667 | 147 | 172.9999 | 323.7777 | 359.3334 | 144.2222 | 234.889 | 201.7777 | 232.4445 | 242.4444 | 120.2222 | 113 | 113.2221 | 121.4444 | 150.1112 | 137.5554 | 135.8888 | 109.6667 | 244.4445 | 211.5555 | 240 | 210.2221 | 127.4443 | 127.2223 | 119.6665 | 118.1111 | 92.8889 | 86.2221 | 74.6667 | 97.8889 | 200.8889 | 202.1111 | 212.6664 | 145.8888 | 136.5556 | 132.3334 | 118.7778 | 102.8888 | 163.9999 | 151.4444 | 184.5555 | 157.1111 | 131.5556 | 107.3333 | 162.5556 | 141.8888 | 169.2224 | 155.2223 | 191.1109 | 211.0001 | 152.5555 | 137.2223 | 214.8889 | 168.4444 | 195.6668 | 184.8888 | 187.8889 | 263.6668 | 187.2222 | 166.9998 | 181.7778 | 220.5556 | 142.7778 | 146.3334 | 120.7778 | 136.4444 | 167.6667 | 158.7777 | 162.3333 | 118.6668 | 120.2222 | 119.6666 | 123.6666 | 128.4445 | 307.6668 | 288.6666 | 287.5555 | 381.1112 | 275.2223 | 243.9999 | 250.889 | 279.2224 | 112.8888 | 76.111 | 75.7778 | 112.8888 | 102.7778 | 93.1111 | 115.7777 | 133.4445 | 102.4444 | 107.2222 | 95.8888 | 108.5556 | 80.3333 | 81.4444 | 73.5555 | 83.0001 | 105 | 106.3333 | 96.3334 | 87.6667 | 254.5556 | 203.3334 | 210 | 217.8888 | 165.7778 | 183.889 | 192.7776 | 207.7777 | 361.3334 | 218.6668 | 212.7778 | 346.111
2488 | 29.3624 | 191.1357 | 188.2609 | 203.8748 | 161.9144 | 160.6284 | 152.4071 | 142.8893 | 145.2464 | 158.3465 | 123.8893 | 108.2964 | 151.2964 | 187.1002 | 149.1607 | 142.357 | 184.975 | 181.643 | 147.5428 | 143.0501 | 164.6034 | 155.3465 | 332.1963 | 375.6535 | 140.2964 | 225.5574 | 203.9499 | 225.7787 | 251.6284 | 125.3821 | 110.7036 | 117.7391 | 120.4071 | 139.7181 | 139.0606 | 133.7892 | 105.3465 | 246.8643 | 226.1712 | 241.1357 | 218.1462 | 125.7641 | 134.7036 | 118.0355 | 116.7036 | 92.025 | 77.1357 | 91.357 | 99.2714 | 186.9645 | 195.6179 | 186 | 137 | 122.2359 | 125.3465 | 122.1357 | 117.8748 | 173.5783 | 164.6284 | 191.1963 | 155.7286 | 130.2964 | 118.4677 | 157.3215 | 133.5929 | 168.8037 | 169.5178 | 181.6284 | 228.6535 | 152.975 | 148.1107 | 217.975 | 161.1357 | 195.9645 | 195.3069 | 183.025 | 257.1503 | 183.3215 | 154.975 | 172.668 | 225.025 | 144.975 | 142.7536 | 125.2464 | 136 | 169.025 | 171.7391 | 168.9499 | 120.6179 | 128.8142 | 115.2964 | 121.5679 | 123.4822 | 291.4968 | 281.4321 | 275.4822 | 396.2213 | 241.8433 | 222.9144 | 252.3465 | 287.8893 | 115.3821 | 74.1107 | 79.1357 | 115.3821 | 111.2714 | 102.5178 | 120.1963 | 122.2609 | 110.4927 | 105.025 | 103.4677 | 107.2964 | 88.0856 | 84.3821 | 80.2213 | 88.7286 | 105.5679 | 114.0856 | 95.025 | 75.9394 | 266.975 | 219.0606 | 206.0501 | 227.1712 | 170.2464 | 180.7787 | 182.4071 | 216.7391 | 349.8538 | 224.0501 | 208.2108 | 335.5929
2489 | 29.3734 | 176.3362 | 170.5191 | 204.4964 | 146.878 | 163.4582 | 163.122 | 136.6794 | 141.2979 | 152.7404 | 117.0993 | 110.4809 | 152.9007 | 197.6255 | 145.2596 | 155.6411 | 178.7177 | 171.2596 | 138.2369 | 138.1603 | 165.817 | 153.2213 | 325.0156 | 345.2908 | 145.3816 | 223.9234 | 204.1986 | 227.8624 | 252.7177 | 126.939 | 122.4426 | 113.4355 | 128.8014 | 141.1447 | 139.1376 | 132.7177 | 96.8397 | 243.0993 | 230.1376 | 239.0993 | 229.1986 | 139.3206 | 130.7787 | 128.0993 | 124.9617 | 91.3206 | 75.6794 | 90.1376 | 98.0993 | 188.5036 | 195.2213 | 205.7248 | 141.061 | 104.0383 | 113.939 | 127.061 | 135.3206 | 166.2142 | 165.7177 | 196.1986 | 142.3972 | 127.8397 | 114.756 | 146.0383 | 146.6638 | 164.9617 | 171.3589 | 206.5036 | 211.0539 | 160.6411 | 149.6794 | 207.6567 | 159.6794 | 185.9007 | 196.6567 | 180 | 253.5036 | 183.061 | 157.4199 | 174.9234 | 222 | 134.6567 | 142.061 | 125.939 | 120.9163 | 164.8397 | 173.817 | 172.0993 | 129.5036 | 124.1376 | 120.9617 | 121.4199 | 116.4199 | 287.4653 | 281 | 265.5191 | 390.0766 | 237.427 | 235.7248 | 261.8241 | 274.1376 | 112.2979 | 80.9007 | 72.4582 | 112.2979 | 117.061 | 115.3816 | 110.695 | 106.2596 | 107.817 | 95.0383 | 107.2979 | 97.878 | 91.5191 | 79.5574 | 80.4582 | 87.5801 | 107.7404 | 107.6567 | 86.1986 | 73.5418 | 254.9163 | 238.8624 | 217.9844 | 257.8241 | 154.695 | 183.4426 | 191.3816 | 207.7943 | 338.8014 | 226.7021 | 207.5418 | 317.9163
2490 | 29.3844 | 168.5546 | 170.9622 | 198.7395 | 153.4411 | 153.2605 | 153.0043 | 150.4033 | 142.7016 | 143.9286 | 113.2227 | 118.9243 | 142.5589 | 204.5211 | 141.0378 | 148.3362 | 186.2184 | 174.7395 | 118.9286 | 136.0378 | 155.8151 | 162.9243 | 308.4454 | 337.5503 | 158.8865 | 225.1135 | 200.5924 | 212.4875 | 254.2941 | 126.9622 | 122.2984 | 111.2562 | 128.853 | 139.5967 | 129.0378 | 136.0714 | 84.7438 | 251.6638 | 227.7395 | 238.1849 | 236.8487 | 144.5546 | 136.4789 | 120.0757 | 125.2605 | 90.6303 | 74 | 84.7773 | 103.7016 | 198.4076 | 188.3362 | 208.1513 | 143.4076 | 112.626 | 115.147 | 134.7395 | 116.8573 | 158.4789 | 151.8908 | 206.2184 | 141.1092 | 141.8108 | 116.0335 | 143.3697 | 143.7816 | 173.5546 | 173.3319 | 207.5211 | 207.2184 | 145.6346 | 158.6638 | 205.7016 | 171.626 | 198.6638 | 176.9286 | 169.9286 | 275.8487 | 181.2605 | 147.5211 | 166.0421 | 211.3362 | 140.4033 | 145.5924 | 147.8822 | 118.8865 | 174.0714 | 147.8194 | 168.2227 | 119.8573 | 116.4076 | 135.4789 | 128.7016 | 117.7773 | 295.2984 | 276.2605 | 278.9957 | 383.5546 | 252.0757 | 253.5546 | 244.3783 | 297.2141 | 111.9243 | 72.9286 | 68.1849 | 111.9243 | 117.6303 | 114.0757 | 107.2562 | 105.5924 | 103.7395 | 94.147 | 100.4076 | 105.0335 | 80.1135 | 78.5168 | 75.5924 | 95.7016 | 106.0378 | 106.8865 | 86.147 | 92.0335 | 237.4832 | 213.4162 | 229.2227 | 263.4832 | 161.3276 | 188.0378 | 181.1892 | 210.5503 | 338.2605 | 219.9665 | 211.2227 | 322.4033
2491 | 29.3954 | 169.7906 | 174.2094 | 205.6283 | 166.8378 | 144.5575 | 156.5134 | 161.6047 | 151.6283 | 139.4189 | 112.6047 | 125.4189 | 130.3717 | 195.5811 | 132.9528 | 144.6283 | 194.7906 | 168.3244 | 110.1858 | 132.7906 | 162.8614 | 160.348 | 313.2567 | 325.6016 | 168.0236 | 207.5102 | 208.8614 | 194.5811 | 253.348 | 124.1622 | 123.0472 | 118.3953 | 123.5811 | 144.7228 | 136.6756 | 142.3953 | 88.4898 | 265.0472 | 229.1858 | 222.0677 | 245 | 149.4189 | 121.8347 | 117.8142 | 113.5339 | 80.5339 | 74.6047 | 76.3244 | 104.7669 | 193.1622 | 177.3717 | 202.4189 | 143 | 118.3717 | 112.5575 | 124.6961 | 109.652 | 163.7906 | 164.537 | 193.0205 | 132.6961 | 130.8347 | 123.1858 | 145 | 127.348 | 170.5575 | 164.3008 | 202.8142 | 220.0236 | 140.8614 | 163.5811 | 211.6047 | 182.8142 | 193.3008 | 169.3953 | 169.0472 | 291.2567 | 178 | 159.1417 | 151.7906 | 218.5134 | 138.3008 | 153.8614 | 146.2535 | 124.3953 | 166.4866 | 155.9559 | 167 | 136.2362 | 119.0236 | 135.348 | 123.115 | 108.115 | 278.5102 | 253.6488 | 281.5339 | 396.2803 | 249.8142 | 249.348 | 251.2126 | 303.4866 | 110.5575 | 80.5134 | 75.0472 | 110.5575 | 122.0472 | 109.3953 | 107.7433 | 115.0709 | 90.6724 | 101.8378 | 109.0709 | 98.8819 | 74.6047 | 70.9055 | 75.3953 | 93.1386 | 99.7669 | 100.3008 | 100.4898 | 89.5102 | 229.9764 | 195.8378 | 210.463 | 271.6992 | 154.2299 | 169.0677 | 177.652 | 229.4425 | 321.6961 | 224.9087 | 221.4898 | 331.1858
2492 | 29.4064 | 177.6382 | 175 | 191.3406 | 171.851 | 156.4254 | 149.9576 | 148.4257 | 142.2767 | 153.9573 | 118.5531 | 115.2767 | 136.6382 | 196.4681 | 145.0424 | 163.6594 | 200.7872 | 153.9788 | 128.1275 | 133.234 | 166.149 | 177.5955 | 320.4681 | 307.766 | 169.383 | 201.851 | 213.383 | 210.2764 | 246.5319 | 115.4469 | 122.0639 | 105.6597 | 118.2979 | 140.1916 | 136.8298 | 143.851 | 84.2767 | 264.6809 | 221.8298 | 218.4042 | 237.5958 | 125.0854 | 133.2976 | 130.7233 | 109.234 | 76.851 | 63.8937 | 83.5743 | 106.3191 | 210.3615 | 178.7021 | 204 | 147.3191 | 127.1063 | 122.5743 | 136.9785 | 112.766 | 175.3403 | 173.5319 | 182.9361 | 127.0851 | 119.4681 | 105.9576 | 148.0851 | 128.5531 | 167 | 163.4042 | 188.5746 | 229.4042 | 147.851 | 149.0427 | 202.1278 | 163.6385 | 183.149 | 175.7872 | 178.5531 | 287.9788 | 172.4469 | 172.234 | 169.5105 | 205.1704 | 138.6382 | 144.8086 | 140.7872 | 134.4893 | 165.6382 | 174.5531 | 163.9149 | 137.4895 | 133.3403 | 127.9149 | 124.6382 | 126.9146 | 256.8086 | 260.7445 | 271.6809 | 404.234 | 250.383 | 255.4893 | 267.149 | 265.6173 | 108.234 | 72.1066 | 75.9149 | 108.234 | 118.5958 | 110.851 | 114.7233 | 121.617 | 99.1275 | 116.1063 | 105.7448 | 93.3191 | 73.149 | 75.3403 | 76.851 | 99.7233 | 95.149 | 108.0424 | 101.8298 | 86.3191 | 231.7021 | 226.7657 | 207.0212 | 255.3194 | 154.1914 | 173.4254 | 177.0639 | 235.4681 | 314.234 | 218.5746 | 229 | 321.9788
2493 | 29.4174 | 176.0776 | 202.6895 | 182.8879 | 168.5949 | 150.2673 | 139.3707 | 159.5085 | 141.9224 | 163.8879 | 118.2242 | 123.1034 | 154.5861 | 197.3707 | 143.6725 | 165.1897 | 182.9743 | 170.3964 | 141.8879 | 142.181 | 160.5949 | 169.6036 | 318.2242 | 319.5861 | 150.1208 | 209.2931 | 210.4828 | 200.8622 | 253.181 | 118.9224 | 118.3707 | 131.3532 | 125.4396 | 133.5172 | 131.7414 | 140.5949 | 85.181 | 250.6725 | 229.3275 | 234.9568 | 225.4483 | 128.5085 | 137.0432 | 131.7069 | 118.8103 | 80.5172 | 65.8103 | 83.1897 | 104.5949 | 211.6725 | 196.1034 | 204 | 142.4483 | 142.181 | 112.1208 | 127.9743 | 110.7414 | 184.2586 | 158.7846 | 167.1208 | 138.4396 | 136.7327 | 109.2154 | 138.0432 | 123.1897 | 172.6638 | 184.362 | 187.8103 | 218.2673 | 145.8535 | 155.474 | 200.4051 | 168.6206 | 197.1034 | 180.6293 | 185.7758 | 309.4309 | 165.2242 | 179.2931 | 187.2931 | 191 | 131.4139 | 147.9568 | 131.1553 | 127.7846 | 182.3275 | 154.7157 | 165.1465 | 126.6293 | 131.5604 | 129.1465 | 118.6725 | 127.8966 | 260.5861 | 247.8277 | 257.6725 | 381.0864 | 282.0946 | 261.3707 | 247.7501 | 280.7239 | 117.181 | 72.3275 | 77.1465 | 117.181 | 119.0345 | 98.7846 | 112.5604 | 115.7069 | 110.3707 | 112.9311 | 104.4051 | 101.6638 | 73.2586 | 84.8879 | 70.4483 | 90.638 | 102.181 | 114.8535 | 97.3707 | 75.1553 | 250.362 | 227.638 | 205.0776 | 246.2931 | 166.7758 | 191.8103 | 186.5861 | 240.7758 | 335.7671 | 213.4051 | 211.3794 | 320.1465
2494 | 29.4284 | 168.1504 | 194.6193 | 194.2656 | 168.5664 | 127.5265 | 144.7744 | 162.6592 | 123.5441 | 161.792 | 119.208 | 122.376 | 162.6416 | 185.4512 | 140.8496 | 156.792 | 199.5135 | 183 | 146.208 | 139.3008 | 160.5664 | 150.584 | 318.5664 | 323.792 | 141.5664 | 232.2479 | 195.5265 | 217.2303 | 259.9248 | 124.2832 | 113.5088 | 131.7521 | 131.6416 | 131.1504 | 137.416 | 144.416 | 87.4336 | 220.9025 | 233.4336 | 242 | 224.208 | 131.0176 | 136.416 | 124.1504 | 116.9424 | 89.6992 | 72.2832 | 90.1904 | 105.8496 | 203.7168 | 201.792 | 203.3584 | 138 | 143.1504 | 114.4735 | 109.2256 | 102.9424 | 167.0353 | 151.6416 | 175.2479 | 136.9424 | 140.2256 | 125.9824 | 138.0576 | 116.7168 | 177.2832 | 164.4865 | 195.5664 | 226.3231 | 152.9824 | 162.7168 | 206.8496 | 175.1504 | 212.416 | 186.7744 | 184.792 | 301.9025 | 170.6992 | 167.6017 | 169.1857 | 195.4912 | 136.1904 | 149.2256 | 131.9824 | 120.6416 | 191.5664 | 160.2479 | 149.0353 | 131.4912 | 136.9072 | 132.2832 | 121.624 | 134.3983 | 267.3584 | 229.1504 | 268.9647 | 363.1504 | 276.6193 | 234.0529 | 251.7567 | 295.5088 | 123.2832 | 68.7344 | 91.1904 | 123.2832 | 113.0176 | 101.9072 | 104.4336 | 115.208 | 109.3584 | 91.6017 | 97.376 | 108.8496 | 75.9248 | 88.5664 | 70.4912 | 79.0752 | 96.7344 | 100.8096 | 109.1904 | 79.1904 | 259.3584 | 210.9424 | 223.4559 | 231.3937 | 173.4912 | 194.4336 | 187.584 | 238.5088 | 334.5265 | 226.2656 | 204.8496 | 330.9824
2495 | 29.4394 | 171.2311 | 197.3472 | 187.5761 | 176.5389 | 140.2705 | 153.2311 | 162.5772 | 131.2705 | 144.3067 | 119.0383 | 114.3844 | 153.8456 | 186.8467 | 136.4611 | 143.23 | 211.4228 | 167.9606 | 134.2683 | 144.1544 | 172.4622 | 165.3089 | 330.4622 | 335.0778 | 154.1161 | 229.9222 | 188.6539 | 212.7295 | 272.1161 | 120.4228 | 118.8467 | 119.0383 | 124.8072 | 149.2705 | 129.8839 | 146.0383 | 91.885 | 221.7317 | 213.6911 | 244.6156 | 210.9606 | 135.1544 | 139.3461 | 130.5006 | 114.3078 | 86.1533 | 70.3844 | 83.9222 | 106.6922 | 201.6922 | 179.7295 | 205.6156 | 124.9222 | 134.4611 | 128.5389 | 115.8083 | 100.3078 | 169.4239 | 146.115 | 189.9233 | 134.9617 | 136.3461 | 126.4228 | 144.6156 | 134.9628 | 181.2694 | 170.8861 | 206.8083 | 211.8061 | 150.8072 | 157.4228 | 185.46 | 173.6539 | 196.3833 | 182.8072 | 192.1544 | 277.8839 | 189.3855 | 167.5006 | 157.6539 | 207.8083 | 150.1928 | 131.6145 | 126.5378 | 121 | 191.0383 | 167.0767 | 178.8872 | 138.5772 | 135.1533 | 125.1533 | 127 | 137.7689 | 266.3461 | 263.6178 | 284.885 | 361.6539 | 272.1544 | 233.3855 | 245.6145 | 305.4239 | 122.0383 | 63 | 91.4611 | 122.0383 | 104.7306 | 113.2311 | 110.8467 | 114.3844 | 117.5006 | 107.1939 | 94.6156 | 112.3078 | 70.4611 | 80.8456 | 80.1928 | 90.4239 | 104.7317 | 86.1533 | 106.8456 | 93.8467 | 251.8072 | 188.6911 | 201.9978 | 246.5017 | 166.1917 | 202.8083 | 168.9606 | 223.5761 | 339.4239 | 228.0767 | 226.6167 | 322.9222
2496 | 29.4504 | 175.3324 | 194.6749 | 181.6662 | 164.6779 | 149.6676 | 138.6793 | 155.0073 | 151.9927 | 151.9883 | 116.6676 | 128.9883 | 127.6823 | 197.6618 | 134.3324 | 148.3251 | 210.9985 | 175.3221 | 136.9927 | 131.0132 | 170.672 | 168.3382 | 345.3265 | 330.6749 | 155.3367 | 221.6676 | 210.3177 | 207.9956 | 258.6808 | 112.0044 | 113.0073 | 119.9985 | 112.3397 | 147.3426 | 129.3294 | 134.3411 | 77.6793 | 256.9795 | 228.6486 | 248.6647 | 224.9839 | 131.3397 | 143.6633 | 129.6706 | 113.6676 | 95.3236 | 82.3236 | 78.3324 | 111.3294 | 218.9868 | 160.3397 | 209.6647 | 145.9795 | 111.0146 | 140.6603 | 133.6574 | 95.6706 | 168.672 | 166.3163 | 181.3411 | 143.9942 | 125.3411 | 134.6589 | 144.0015 | 124.3484 | 161.6823 | 192.3265 | 201.3411 | 225.6472 | 145.6676 | 155 | 206.9752 | 198.6486 | 194.6603 | 186.9942 | 187.6735 | 293.3177 | 183.6764 | 173.9985 | 152.0044 | 187.019 | 134.0146 | 131.9942 | 138.9868 | 136.3221 | 186.0029 | 161.6691 | 179.3484 | 132.3397 | 132.3324 | 124.3309 | 125.6676 | 136.9985 | 247.347 | 282.3338 | 282.0044 | 382.6516 | 247.0219 | 229.0088 | 237.3338 | 301.3411 | 116.3367 | 81.653 | 86.0015 | 116.3367 | 109.6618 | 104.6749 | 110.3367 | 107.6706 | 120.0015 | 99.3484 | 105.3265 | 102.3411 | 73.6618 | 83.9942 | 70.6764 | 90.3382 | 105.3382 | 97.9883 | 103.9985 | 84.6764 | 254.6618 | 194.9883 | 208.6486 | 224.0263 | 145.6779 | 198.0073 | 191.6443 | 216.3338 | 361.9883 | 194.0234 | 223.6764 | 320.6633
2497 | 29.4614 | 167.8584 | 201.2125 | 207.1034 | 146.1445 | 160.534 | 125.9292 | 158.8201 | 149.5368 | 151.1799 | 110.5722 | 129.5368 | 158.3867 | 208.4632 | 135 | 131.6105 | 188.932 | 192.4986 | 143.3569 | 140.3187 | 168.3569 | 157.5368 | 321.5042 | 311.4306 | 148.2507 | 233.8909 | 233.2125 | 210.3215 | 269.3541 | 111.7139 | 113.4278 | 104.0382 | 129.0326 | 137.6076 | 121.8229 | 129.6785 | 79.1771 | 271 | 238.9646 | 214.0411 | 214.2889 | 121.5722 | 131.7521 | 133.7847 | 119.7847 | 93.8584 | 78.1445 | 73.5722 | 105.1799 | 201.5396 | 194.6728 | 220.4986 | 149.1445 | 111.1771 | 138.2153 | 137.2861 | 107.9264 | 175.1771 | 178.6785 | 184.8201 | 137.8229 | 140.3541 | 125.7521 | 145.7139 | 126.8909 | 146.2507 | 187.5014 | 192.6076 | 224.0736 | 144.3215 | 144.1445 | 206.3598 | 196.0736 | 194.2507 | 178.1091 | 178.9292 | 276.1827 | 168.1799 | 166.1799 | 155.1062 | 190.9618 | 135.534 | 143.4632 | 152.7493 | 148.7493 | 189.4278 | 167.4632 | 165.6076 | 140.8909 | 131.6431 | 119.8938 | 122.2861 | 141.3924 | 271.2451 | 255.5396 | 252.5396 | 365.8612 | 242.1771 | 228.4278 | 254.6402 | 282.4306 | 114 | 76.0736 | 76.1799 | 114 | 107.5722 | 91.5368 | 118.8555 | 102.9646 | 115.6076 | 95.7847 | 113.3924 | 107.8555 | 70.8938 | 81.2153 | 69.4278 | 93.7847 | 99.2861 | 107.3569 | 86.6813 | 93.6048 | 220.6841 | 200.2861 | 201.3244 | 234.4958 | 151.5694 | 182.8229 | 196.1445 | 227.534 | 350.3244 | 203.1034 | 202.0736 | 352.8527
2498 | 29.4724 | 166.0723 | 181.4418 | 204.494 | 150.6707 | 159.7831 | 139.1968 | 148.494 | 132.8755 | 143.5462 | 99.0201 | 127.3815 | 149.6788 | 205.7831 | 130.8554 | 142.4136 | 183.5261 | 187.3293 | 133.6386 | 137.9478 | 173.1446 | 143.6386 | 323.1968 | 310.5261 | 144.6185 | 224.1125 | 221.731 | 212.763 | 259.6587 | 115.763 | 96.6587 | 108.4337 | 111.3695 | 146.3614 | 137.7228 | 141.7429 | 84 | 252.3494 | 232.4739 | 217.0321 | 202.6185 | 110.0201 | 122.9277 | 125.2571 | 121.6185 | 90 | 68.8554 | 85.506 | 95.4739 | 193.8353 | 179.8434 | 227.0723 | 139.8554 | 126.3614 | 134.3092 | 138.763 | 110.1647 | 166.8755 | 170.7109 | 177.2571 | 143.3614 | 150.6908 | 116.9277 | 137.3293 | 140.5984 | 163.3413 | 168.494 | 178.5663 | 222.5261 | 144 | 153.506 | 214.5783 | 177.9478 | 180.2571 | 192.7228 | 174.9277 | 242.9679 | 154.3293 | 159.9277 | 160.763 | 199.6908 | 141 | 127.6587 | 117.6989 | 146.1647 | 168.5141 | 175.1446 | 170.2169 | 138.7109 | 130.3092 | 132.8875 | 122.3815 | 144.3815 | 267.6587 | 252.6707 | 254.506 | 345.4016 | 258.0522 | 256.5582 | 244.3494 | 284.9799 | 114 | 80.7429 | 74.763 | 114 | 111.2169 | 100.4337 | 115.7109 | 104.763 | 109.8554 | 112.1245 | 110.8554 | 102.6386 | 75.5984 | 84.9076 | 82.3614 | 90.7831 | 98 | 94.8755 | 84.9076 | 83.0402 | 217.506 | 198.3092 | 206.506 | 230.731 | 173.8875 | 183.5261 | 177.1848 | 224.0201 | 365.1767 | 209.4739 | 211.5783 | 357.3293
2499 | 29.4834 | 154.3451 | 180.5458 | 209.9519 | 153.5939 | 144.3451 | 127.7206 | 178.5587 | 141.7641 | 145.5153 | 102.0305 | 117.4542 | 133.4847 | 180.5023 | 137.4366 | 146.3756 | 189.5153 | 171.0481 | 136.7336 | 128.0786 | 172.1878 | 141.8122 | 334.2183 | 282.7687 | 143.2969 | 209.9695 | 205.5634 | 224.5458 | 262.2488 | 118.4061 | 106.2794 | 104.1573 | 114.4672 | 138.3451 | 135.0481 | 120.9871 | 82.5939 | 241.8908 | 249.6855 | 230.2183 | 243.4801 | 132.716 | 128.3275 | 131.2488 | 116.0786 | 81.5634 | 94.4191 | 64.5809 | 114.7946 | 205.8427 | 189.6068 | 218.1573 | 144.3275 | 136.6244 | 129.0786 | 130.1573 | 106.5939 | 164.5153 | 158.5634 | 182.5458 | 136.7512 | 139.7512 | 123.0305 | 156.2007 | 125.0176 | 175.5153 | 171.1397 | 196.2007 | 222.8908 | 143.2969 | 144.5328 | 208.642 | 203.9343 | 182.0305 | 190.7512 | 174 | 257.2007 | 147.1878 | 142.1267 | 154.9695 | 196.4847 | 127.642 | 137.9824 | 120.6855 | 141.1878 | 147.4542 | 157.3145 | 179.3275 | 135.7031 | 137.7336 | 128.7512 | 131.4366 | 129.5328 | 265.3275 | 265.4366 | 277.1702 | 349.7336 | 248.9389 | 248.3145 | 239.5153 | 284.7817 | 133.6855 | 76.8603 | 66.8603 | 133.6855 | 122.4366 | 101.0786 | 133.0916 | 120.0611 | 107.2969 | 108.8603 | 109 | 103.6244 | 83.9214 | 92.2183 | 70.8298 | 95.0305 | 102.9214 | 99.5458 | 103.4672 | 79.9214 | 229.6244 | 192.3756 | 201.0481 | 224.4061 | 167.642 | 187.4061 | 165.3756 | 235.4672 | 380.9214 | 191.5328 | 232.358 | 346.6725
2500 | 29.4944 | 164.7376 | 205.0296 | 214.2847 | 153 | 151.8762 | 142.891 | 159.6635 | 148 | 149.8614 | 127.1757 | 115.146 | 141.2995 | 193.891 | 134.5619 | 144 | 204.5915 | 158.8466 | 141.4307 | 126 | 174.5767 | 152.2995 | 317.4011 | 320.7894 | 149.4381 | 219.1608 | 202.7153 | 203.2476 | 266.2847 | 116.1386 | 106.8466 | 115.7376 | 118.1386 | 146.5915 | 117.8392 | 137.1831 | 89.1534 | 247.4381 | 227.9556 | 229.854 | 251.7005 | 146.8614 | 138.1534 | 120.2624 | 99.6931 | 85.1534 | 91.6931 | 59.4381 | 127.5767 | 204.9926 | 206.4381 | 212.5693 | 157.7302 | 140.4307 | 127.7153 | 131.0074 | 95.9852 | 161.7079 | 142.1238 | 205.5989 | 155.6063 | 138.5767 | 131.7228 | 171.0074 | 144.1831 | 185.5841 | 180.7228 | 198.1312 | 227.0074 | 151.5841 | 148.7302 | 201.5693 | 195.5397 | 173.5545 | 180.9926 | 174 | 279.8762 | 151.7228 | 136.4307 | 134.8317 | 182.1238 | 135.5915 | 148.146 | 123.9926 | 134.9926 | 174.4751 | 171.1757 | 151.2402 | 134.5693 | 126.6931 | 122.5693 | 141.4381 | 122.2847 | 270.8614 | 262.5619 | 246.8021 | 338.6931 | 243 | 229.2698 | 251.7302 | 295.1608 | 119.109 | 77.2921 | 65.8614 | 119.109 | 115.9852 | 101.8614 | 115.5323 | 108.1164 | 96.9852 | 103.5693 | 116.8688 | 103.854 | 75.9852 | 82.5545 | 75.4455 | 93.7079 | 94.9852 | 91.8392 | 93.5471 | 77.7079 | 239.1534 | 210.0296 | 201.7228 | 254.3291 | 159.854 | 195.1534 | 168.7228 | 227.6931 | 351.5249 | 225.7746 | 231.5619 | 361.1683
2501 | 29.5054 | 154.2644 | 204.2684 | 190.7157 | 143.5407 | 169.3698 | 138.9026 | 164.1909 | 134.9026 | 145.1789 | 143.2763 | 113.8171 | 145 | 207.3658 | 166.1988 | 146.1829 | 199.0855 | 175.6461 | 136.9066 | 130.3658 | 173.8171 | 147.996 | 330.3738 | 296.6143 | 152.7276 | 221.8171 | 198.6342 | 195.9105 | 246.3539 | 119.3658 | 112.004 | 118.3618 | 121.3658 | 152.7276 | 127.5527 | 131.2644 | 83.2684 | 254.3658 | 249.4712 | 226.8171 | 229.0815 | 137.8131 | 130.0855 | 118.3658 | 117.2843 | 84.3618 | 80.1789 | 91.1054 | 106.4433 | 192.8131 | 195.1749 | 227.2803 | 188.1949 | 128.6302 | 133.0934 | 142.4593 | 103.6422 | 160 | 142.8211 | 186.8051 | 165.7276 | 146.5487 | 114.3539 | 149.7157 | 144.0855 | 177.3578 | 171.3578 | 187.7237 | 232.6382 | 157.1829 | 146.4513 | 209.7316 | 178.996 | 185.7356 | 191.3698 | 179.8211 | 283.5447 | 144.5407 | 150.8251 | 134.5487 | 181.3658 | 149.004 | 151.9105 | 130.7316 | 141.004 | 187.7276 | 185.0934 | 166.6501 | 133.2724 | 130.4593 | 111.8131 | 131.6302 | 124.1829 | 269.8171 | 244.7197 | 242.1869 | 342.4593 | 253.9145 | 222.0895 | 258.9105 | 302.9105 | 109.2724 | 77.5447 | 71.3658 | 109.2724 | 120.004 | 103 | 115.1869 | 111.1869 | 103.9145 | 103.7276 | 110.5407 | 110.2763 | 80.7316 | 80.9105 | 66.1749 | 102.9145 | 109.1909 | 91.3658 | 81.9066 | 85.4593 | 217.2604 | 205.6302 | 232.3778 | 217.9761 | 178.6461 | 194.3618 | 186.2803 | 218.3618 | 332.4513 | 219.4433 | 221.7237 | 355.6302
2502 | 29.5164 | 156.8791 | 214.3187 | 207.1575 | 179.2161 | 153.2821 | 139.9194 | 144.3627 | 148.4981 | 165.1978 | 131.2015 | 118.1795 | 144.2601 | 180.1429 | 145.7033 | 157.359 | 196.4799 | 183 | 140.1795 | 137.1795 | 158.9414 | 142.7802 | 308.403 | 299.978 | 153 | 218.0403 | 222.8974 | 209.5787 | 246.3993 | 128.3993 | 123.1392 | 110.3407 | 127.4396 | 152.2601 | 144.8388 | 125 | 89.619 | 238.2418 | 207.2455 | 229.6996 | 236.7985 | 137.6996 | 145.2381 | 130.359 | 114.9011 | 74.8608 | 89.0989 | 82.022 | 108.359 | 220.0769 | 187.7802 | 196.0037 | 182.4616 | 126.9597 | 138.6996 | 138.6007 | 110.9597 | 171.8388 | 154.619 | 179.9597 | 159.3407 | 154.9194 | 135.1172 | 159.4981 | 141.4799 | 162.641 | 164.7802 | 193.1392 | 232.5201 | 158.7399 | 150.6593 | 210.7802 | 173.0403 | 194.9597 | 190.0806 | 173.1209 | 300.0183 | 167.6373 | 148.6007 | 132.5604 | 215.5567 | 139.4213 | 146.3407 | 133.2601 | 141.0403 | 194.6593 | 184.7802 | 182.9194 | 113.7619 | 136.9597 | 100.6007 | 112.2015 | 129.4396 | 269 | 253.0586 | 239.3407 | 346 | 263.9194 | 252.0769 | 238.5421 | 295.1209 | 123.7985 | 86.619 | 69.3004 | 123.7985 | 123 | 105.9597 | 108.641 | 103.1612 | 102.0806 | 102.5201 | 93.6813 | 107.8205 | 81.0403 | 83.4799 | 81.7179 | 95.1612 | 96.7619 | 106.3187 | 97.7582 | 77.1612 | 262.7545 | 220.978 | 206.7436 | 261.4138 | 143.8242 | 201.8791 | 200.8791 | 237.7179 | 365.5164 | 210.2601 | 207.9011 | 350.2601
2503 | 29.5274 | 153.9823 | 197.1857 | 206.2212 | 171.1857 | 135.469 | 147.2655 | 149.2921 | 138.4512 | 171.4956 | 137.2833 | 121.5044 | 139.4867 | 202.3453 | 130.9911 | 145.9557 | 169.9202 | 170.2123 | 148.0177 | 131.4778 | 179.5754 | 163.8143 | 296.4956 | 328.0621 | 142.469 | 210.9823 | 218.4601 | 220.77 | 270.0621 | 114.4512 | 113.2123 | 111.7611 | 118.469 | 139.9645 | 130.9468 | 126.5044 | 86.9823 | 241.0266 | 204.5488 | 236.2655 | 218.6813 | 123.9557 | 157.2655 | 120.4601 | 111 | 90.0532 | 73.4424 | 75.7522 | 106.7345 | 199.4069 | 177.2212 | 217.602 | 174.7433 | 121.9823 | 121.1946 | 127.7256 | 118.0177 | 173.7433 | 153.4867 | 193.7877 | 162.2655 | 141.2034 | 130.7079 | 166.7522 | 163.0621 | 160.5044 | 172.2744 | 190.7345 | 241.0266 | 152.9823 | 157.5133 | 219.0266 | 183.2833 | 212.5488 | 188 | 208.3631 | 303.7433 | 172.4867 | 158.0355 | 151.3099 | 220.9823 | 146.2833 | 137.9823 | 120.2123 | 136.2389 | 188.7256 | 173.469 | 168.4512 | 116.7788 | 113.9291 | 116.8054 | 116.7788 | 122.7256 | 265.9911 | 239.9468 | 251.2921 | 331.7079 | 253.9645 | 235.9202 | 255.0709 | 308.5488 | 119.9734 | 77.2123 | 67.2478 | 119.9734 | 115.4778 | 113.0177 | 107.2567 | 101.2567 | 103.7611 | 105.0089 | 97.2744 | 113.5222 | 85.2655 | 81.7433 | 69.4424 | 99.2744 | 117.0798 | 90.6901 | 104 | 83.531 | 253.4158 | 219.7256 | 203.7788 | 243.8847 | 169.6197 | 209.5133 | 204.7522 | 237.4778 | 384.0177 | 204.7345 | 223.5576 | 362.7877
2504 | 29.5384 | 168.0547 | 191.529 | 190.0034 | 157.8839 | 160.2868 | 145.1775 | 162.4096 | 136.8225 | 164.1194 | 124.1808 | 129.6451 | 153.2902 | 177.0681 | 137.6451 | 153.2321 | 177.8192 | 149.9453 | 135.4743 | 133.587 | 176.5324 | 162.5904 | 308.9966 | 318.9453 | 137.471 | 212.8225 | 201.0034 | 204.6518 | 263.2389 | 119.7031 | 118.1741 | 113.7645 | 121.1161 | 133.7065 | 125.7645 | 133.8806 | 78.1194 | 232.5324 | 224.5257 | 226.5324 | 235.4643 | 144.2288 | 134.5357 | 120.587 | 117.8806 | 96 | 80.7611 | 87.4676 | 110.3516 | 199.7031 | 177.8225 | 213.7098 | 143.4197 | 128.4096 | 120.3516 | 114.2969 | 119.2355 | 182.1741 | 153.529 | 175.8293 | 160.942 | 137.529 | 149.6997 | 162.413 | 155.4743 | 164.8225 | 178.8225 | 207.3482 | 232.5324 | 154.058 | 129.9487 | 216.6484 | 177.8259 | 213.413 | 168.1228 | 180.4811 | 294.5904 | 175.587 | 142.8873 | 147.2969 | 207.5324 | 146.1775 | 132.1775 | 128.2321 | 149.5257 | 186 | 187.5837 | 169.8806 | 135.2902 | 106.7645 | 117.6484 | 114.6484 | 116.942 | 267.2935 | 234 | 241.4743 | 369.8124 | 266.8192 | 242.2902 | 246.1808 | 319.3516 | 123.1161 | 81.4096 | 60.1194 | 123.1161 | 119.1161 | 121.8806 | 117.1741 | 109.6451 | 104.2355 | 115.1741 | 90.0614 | 120.587 | 83.942 | 88.6451 | 79.8192 | 99.7065 | 120.6484 | 107.6997 | 102.471 | 73.2389 | 276.1093 | 213.1775 | 196.2969 | 273.8124 | 189.1161 | 214.8225 | 198.1194 | 269.4029 | 374.5324 | 204.529 | 218.5324 | 367
2505 | 29.5494 | 159.7944 | 194.3304 | 184.4464 | 162.2144 | 139.4816 | 147.884 | 171.9912 | 136.4464 | 140.2496 | 118.2232 | 113.3568 | 154.8928 | 199.4024 | 116.696 | 140.6872 | 205.5272 | 166.7504 | 146.536 | 146.652 | 183.0984 | 139.0264 | 321.5448 | 313.2232 | 136.2232 | 239.6344 | 206.3216 | 201.3304 | 269.0984 | 112.1248 | 100.8032 | 110.116 | 115.232 | 139.2144 | 116.6784 | 150.7592 | 79.884 | 242.9824 | 224.3392 | 207.464 | 224.3568 | 141.1248 | 126.2232 | 108.7944 | 121.5536 | 82.0176 | 82.6696 | 80.9016 | 110.4464 | 196.0088 | 183.6608 | 233.0808 | 161.188 | 135.6608 | 118.116 | 107.8928 | 101.9104 | 170.2408 | 155.5536 | 190.7504 | 163.884 | 114.696 | 134.4728 | 162.5536 | 133.9104 | 180.7592 | 193.9824 | 183.4816 | 230.5536 | 154.2232 | 134.2056 | 198.6872 | 171.116 | 219.768 | 172.0984 | 179.652 | 278.0176 | 163.7944 | 179.7239 | 145.5536 | 204.7768 | 172.9648 | 140.3216 | 127.3392 | 151.6696 | 186.7768 | 181.348 | 178.2144 | 123.6872 | 119.4288 | 136.9736 | 140.9648 | 124.5448 | 277.3216 | 258.8576 | 255.6432 | 321.6329 | 244.812 | 253.9912 | 242.5536 | 283.7136 | 121.116 | 79.3392 | 81.304 | 121.116 | 114.7856 | 104.58 | 106.7944 | 116.6608 | 109.4376 | 105.5712 | 104.8664 | 111.9016 | 63.58 | 75.464 | 102.0896 | 90.4552 | 109.6784 | 104.9016 | 116.7592 | 74.4376 | 251.8208 | 215.1072 | 204.652 | 245.8297 | 160.7048 | 201.2408 | 192.116 | 237.2761 | 374.1072 | 218.9824 | 227.4288 | 363.8928
2506 | 29.5604 | 157.5782 | 201.3127 | 206.0945 | 146.64 | 142.0473 | 152.9455 | 159.0073 | 162.04 | 174.2436 | 123.5236 | 108.7891 | 134.2727 | 212.9455 | 129.7273 | 144.68 | 205.6873 | 158.7964 | 151 | 149.2109 | 171.0073 | 151.1491 | 319.2655 | 298.0073 | 141.5236 | 215.4364 | 205.8436 | 227.2509 | 263.32 | 116.8909 | 98.1564 | 112.9455 | 127.2036 | 146.5236 | 129.7818 | 131.3273 | 91.2582 | 241.4764 | 255.3527 | 210.1018 | 220.5782 | 134.0545 | 126.7891 | 123.9382 | 107.0073 | 85.1018 | 69.3746 | 81.9455 | 109.2109 | 187.6873 | 195.2582 | 214.7491 | 171.3673 | 133.0545 | 132.7818 | 103.8436 | 115.1491 | 152.5855 | 150.4764 | 177.2727 | 152.3746 | 163.2363 | 120.8982 | 161.4218 | 136.1018 | 166.8509 | 175.9055 | 171.0545 | 255.4618 | 142.9527 | 141.1564 | 210.5709 | 184.2036 | 183.3346 | 193.1491 | 188.5236 | 285.0473 | 164.7345 | 166.7491 | 144.4218 | 198.6873 | 173.1091 | 148.5236 | 148.0945 | 144.6873 | 180.6873 | 177.2109 | 161.8509 | 137.1491 | 122.2109 | 130.3746 | 118.2255 | 114.3746 | 291.0473 | 243.1164 | 251.5309 | 363.1818 | 252.7818 | 240.2182 | 270.6182 | 283.2582 | 134.2036 | 81.9455 | 84.0545 | 134.2036 | 127.2036 | 102.1564 | 114.8364 | 114.8436 | 106.2655 | 104.3673 | 95.0073 | 112.9455 | 66.68 | 73.3673 | 88.8509 | 91.1564 | 118.0473 | 99.6327 | 103.64 | 80.7345 | 237.2655 | 222.3127 | 198.5309 | 245.8364 | 174.0945 | 190.6873 | 205.2036 | 237.6254 | 365.5309 | 216.6873 | 238.1018 | 396.1418
2507 | 29.5714 | 158 | 186.9724 | 203.1848 | 140.3972 | 138.589 | 159.6097 | 154.1986 | 160.9862 | 164.9655 | 115.3834 | 126.6304 | 145.829 | 200.3765 | 129.3903 | 152.6097 | 201.5959 | 167.0207 | 148.5959 | 138.582 | 177.418 | 154.3972 | 294.7599 | 289.9931 | 140.5959 | 215.0138 | 198.589 | 220.3765 | 239.3627 | 138.2331 | 115.0276 | 102.7807 | 147.829 | 158.418 | 137.2055 | 118.589 | 90.7945 | 235.1917 | 235.1503 | 232.0345 | 232.2193 | 130.5959 | 128.6028 | 115.3765 | 117.4249 | 85.3972 | 86.8359 | 75.7876 | 94.5751 | 222.8635 | 214.829 | 202.3903 | 177.6097 | 129.5959 | 126.582 | 117.4249 | 107.9793 | 160.2193 | 180.2539 | 198.4456 | 150.6028 | 158.7669 | 108.582 | 173.8221 | 155.6304 | 184.4387 | 174.8083 | 193.2401 | 227.5406 | 156.829 | 145.2055 | 214.1986 | 211.2401 | 192.2331 | 197.1986 | 181.9862 | 294.411 | 161.9931 | 151.1848 | 143.1986 | 207.418 | 133.3351 | 143.589 | 129.9586 | 138.1917 | 185.411 | 187.418 | 180.2401 | 138.7945 | 126.0069 | 130.2055 | 118.0138 | 133.4387 | 278.7738 | 258.6373 | 218.5475 | 350.9517 | 243.3765 | 256.0345 | 250.753 | 294.0138 | 138 | 66.9724 | 71.7807 | 138 | 130.1986 | 111.0138 | 109.9862 | 118.0069 | 104.1986 | 96.1848 | 95.0069 | 108.3903 | 86.6304 | 71.5959 | 82.3972 | 84.7876 | 114.589 | 106.2124 | 103.8083 | 79.5959 | 278.4732 | 207.171 | 212.829 | 246.5959 | 167.9793 | 184.9931 | 194.5751 | 273.8566 | 378.2262 | 219.8083 | 228.7807 | 360.924
2508 | 29.5824 | 182.4102 | 217.9879 | 217.2734 | 161.9691 | 136.1863 | 151.2359 | 146.677 | 135.4035 | 160.8137 | 141.4785 | 131 | 165.4598 | 194.559 | 149.9691 | 150.7453 | 214.0188 | 159.4223 | 134.9812 | 134.3727 | 171.0496 | 147.4906 | 297.1367 | 272.7266 | 144.0684 | 216.1863 | 206.7641 | 189.3351 | 260.8512 | 125.9129 | 103.5402 | 104.882 | 137.3539 | 141.4719 | 129.0496 | 143.8512 | 91.6273 | 231.559 | 226.3727 | 234.559 | 229.3043 | 135.6957 | 140.3914 | 137.2238 | 112.0496 | 81.7453 | 79.7949 | 78.0684 | 114.5965 | 205.1488 | 209.2359 | 216.4598 | 158.6582 | 137.1367 | 137.0188 | 126.6957 | 111.5094 | 177.6461 | 163.5898 | 179.7762 | 134.7266 | 162.1367 | 136.9195 | 163.9812 | 151.0496 | 172.0992 | 181.6957 | 181.9129 | 251.5469 | 160.1863 | 154.9504 | 228.6461 | 187.7078 | 194.559 | 175.8445 | 175.9316 | 280.5402 | 171.5777 | 162.0188 | 150.323 | 201.0496 | 161.4289 | 137.118 | 149.2238 | 149.2051 | 197.5777 | 190 | 157.5215 | 137.1863 | 139.2051 | 126.9316 | 125.6957 | 127.6086 | 273.3727 | 250.9812 | 232.1555 | 334.2359 | 274.1742 | 265.882 | 240.7453 | 314.7145 | 132.3043 | 76.8324 | 77.9504 | 132.3043 | 128.3727 | 117.882 | 121.8324 | 132.8324 | 104.8137 | 88.3043 | 89.4906 | 115.9504 | 79.6086 | 71 | 76.3043 | 91.1367 | 121.1367 | 95.7949 | 88.7266 | 80.6273 | 240.1797 | 217.6461 | 193.4035 | 265.5281 | 176.3914 | 184 | 204.8324 | 263.2855 | 388.5094 | 225.0684 | 217.0496 | 399.634
2509 | 29.5934 | 164.873 | 201.2211 | 210.2625 | 180.2154 | 158.301 | 144.0442 | 146.6519 | 130 | 158.5221 | 125.699 | 132.6519 | 171.4779 | 197.3039 | 137.6548 | 150.826 | 192.2211 | 166.0856 | 133.6519 | 129.0442 | 186.3452 | 164.1712 | 288.2625 | 303.6905 | 141.6961 | 228.3894 | 219.7375 | 216.8645 | 248.8288 | 98.873 | 94.2183 | 97.7404 | 134 | 152.6933 | 119.5663 | 136.7846 | 96.9558 | 235.1298 | 234.2596 | 222.4365 | 225.5221 | 130.3923 | 132.2625 | 116.5692 | 116.6077 | 95.0414 | 83.6077 | 85.6077 | 110.0884 | 232.8645 | 214.4337 | 232.3894 | 162.2596 | 137.3481 | 124.3067 | 137.0856 | 102.2625 | 190.0856 | 139.8288 | 178.9558 | 146.6933 | 158.2183 | 139.0442 | 160.174 | 168.8231 | 162.2183 | 175.5663 | 203.6049 | 243.3067 | 154.2183 | 155.3481 | 230.3481 | 193.3894 | 179.1327 | 164.3923 | 173.3481 | 291.8673 | 180.6077 | 179.0414 | 165.2154 | 194.0442 | 141.0913 | 129.3923 | 145.9144 | 129.699 | 209.9116 | 205.6933 | 187.3424 | 141.1298 | 108.9615 | 136.7375 | 123.6961 | 112.6106 | 277.1298 | 255.4337 | 232.8702 | 350.1712 | 286.9558 | 233.1355 | 253.2154 | 271.9201 | 121.9144 | 81.6519 | 83.3039 | 121.9144 | 143.6933 | 101.6548 | 109.3067 | 111.2211 | 119.0414 | 94.4337 | 96.2596 | 97.3509 | 86.9116 | 80.9116 | 80.7817 | 100.4337 | 104.0028 | 114.475 | 99.0414 | 83.4779 | 261.2126 | 230.0856 | 187.174 | 263.3923 | 202.127 | 195.5635 | 222.0414 | 245.7846 | 421.3866 | 198.7432 | 216.6519 | 389.525
2510 | 29.6044 | 133.1759 | 196 | 203.8087 | 167.0732 | 157.9705 | 148.0295 | 141.9705 | 141.7356 | 175.6033 | 130.2208 | 121.2644 | 150.2054 | 200.5148 | 139.0295 | 154.353 | 203.7651 | 177.2208 | 138.1913 | 127.1617 | 164.0141 | 162.9705 | 358.0899 | 321.0591 | 156.0886 | 224.294 | 221.1617 | 198.8524 | 248.353 | 105.7356 | 118.1476 | 111.9268 | 128.1322 | 155.1617 | 130.5738 | 134 | 94.647 | 241.8678 | 237.6765 | 229.2208 | 253.5006 | 134.8678 | 114.0732 | 129.4416 | 131.4121 | 89.6175 | 78.294 | 73.5879 | 102.1322 | 217.3671 | 179.1168 | 239.1913 | 137.1759 | 124.4262 | 131.8973 | 111.3376 | 106.706 | 176.9114 | 161.9859 | 195.9268 | 161.7356 | 158.6765 | 133.8087 | 186.8241 | 174.6765 | 157.647 | 198.3094 | 209 | 215.6906 | 152.1617 | 152.4852 | 202.3376 | 214.4416 | 193.6033 | 172.2208 | 179.706 | 284.9409 | 200.4416 | 161.8819 | 176.3825 | 217.3094 | 150.9268 | 142.2503 | 128.9114 | 149.3094 | 221.2208 | 216.5443 | 190.8087 | 130.2644 | 122.1181 | 111.3376 | 117.1322 | 111.6765 | 270.4557 | 262.8678 | 218.5879 | 339.7497 | 335.7805 | 239.4121 | 277.7946 | 297.2067 | 144.3094 | 81.1617 | 82.3235 | 144.3094 | 142.8087 | 138.2362 | 115.2208 | 108.5148 | 125.353 | 114.4416 | 98 | 108.0886 | 73.0732 | 89.706 | 75.294 | 98.647 | 92.4557 | 103.9114 | 96.9705 | 76.4557 | 259.6175 | 229.4852 | 207.9564 | 244.3967 | 174.3081 | 223.9859 | 203.2054 | 253.8973 | 400.3376 | 229.0449 | 222.8678 | 393.3825
2511 | 29.6154 | 172.2284 | 174.7364 | 216.6087 | 169.1033 | 157.8505 | 143.8967 | 151.2065 | 139.7473 | 159.4375 | 143.9076 | 126.6549 | 180.0218 | 231.6196 | 153.6087 | 138.8396 | 215.5055 | 167.0924 | 145.8044 | 142.3098 | 189.6196 | 161.1495 | 304.8069 | 299.1847 | 142.8396 | 245.1142 | 215.0462 | 205.0571 | 243.8967 | 129.2636 | 128.9538 | 107.3451 | 124.4484 | 126.932 | 144.9076 | 130.5978 | 100.8044 | 230.2418 | 211.6331 | 231.8505 | 230.0815 | 129.1956 | 122.9076 | 126.1956 | 120.3913 | 77.7935 | 80.4022 | 80.356 | 104.4022 | 228.3098 | 205.1713 | 248.5055 | 170.2745 | 127.9538 | 142.5055 | 126.4131 | 112.2527 | 172.2989 | 164.4484 | 190.4945 | 159.7473 | 164.1033 | 135.5516 | 189.4484 | 176.7011 | 178.2636 | 189.3913 | 207.2989 | 256.0789 | 182.6196 | 133.288 | 192.7473 | 183.9782 | 215.712 | 217.3778 | 184.4022 | 303.4131 | 181.0353 | 166.5055 | 160.9891 | 221.1495 | 138.6902 | 151.8044 | 124.2989 | 158.2527 | 234.9076 | 215.4484 | 194.2527 | 145.0109 | 115.7935 | 124.712 | 108.3451 | 130.712 | 271.5516 | 258.8967 | 234.712 | 322.5407 | 288.0135 | 247.1033 | 245.4266 | 312.5055 | 136.2418 | 78.4484 | 78.5978 | 136.2418 | 127.5407 | 135.7935 | 105.9429 | 103.0462 | 112.3913 | 108.644 | 107.356 | 128.0109 | 78.5055 | 75.6902 | 79.1033 | 85.2418 | 119.068 | 129.068 | 92.5978 | 79.2527 | 246.0924 | 219.644 | 224.7582 | 258.8615 | 208.8262 | 229 | 196.4484 | 273.0109 | 424.7691 | 224.0924 | 229.9538 | 423.068
2512 | 29.6264 | 142.0352 | 177.9027 | 225.0399 | 171.7257 | 142.4689 | 163.7081 | 143.5088 | 157.9824 | 152.5487 | 137.3716 | 113.3318 | 176.5088 | 186.9555 | 137.8804 | 160.1594 | 185.0751 | 190.0223 | 154.7655 | 151.0399 | 186.3716 | 171.354 | 338.7303 | 332.9648 | 164.1594 | 241.2345 | 210.5487 | 217.354 | 264.5709 | 128.6858 | 129.1372 | 112.9027 | 130.0399 | 133.2169 | 146.1372 | 117.0575 | 95.9601 | 231.4513 | 218.2169 | 227.6858 | 235.354 | 135.7655 | 141.3939 | 138.8054 | 120.5885 | 91.5311 | 62.8804 | 72.5088 | 105 | 233.5885 | 205.823 | 196.5041 | 177 | 143.6682 | 123.2919 | 122.2345 | 101.7831 | 170.2743 | 178.6682 | 190.7257 | 157.2743 | 145.1548 | 139.4513 | 160.5264 | 177.8628 | 162.1548 | 179.2345 | 192.3318 | 261.4115 | 157.8007 | 133.4513 | 206.6682 | 184.0399 | 213.823 | 193.0751 | 186.7257 | 263.8582 | 166.646 | 168 | 145.0575 | 219.2743 | 150.6682 | 157.3142 | 136.0797 | 171.0797 | 224.0575 | 186.5264 | 179.4689 | 162.6682 | 130.3939 | 132.3142 | 113.9027 | 131.4115 | 289.2567 | 273.5311 | 246.6284 | 339.8452 | 271.9601 | 245.4115 | 246.7655 | 296.7433 | 143.4912 | 72.823 | 73.6858 | 143.4912 | 144.8452 | 124.5088 | 112.6284 | 91.646 | 125.5311 | 110.4513 | 99.5088 | 112.8804 | 74.823 | 89.3939 | 73.9601 | 100.2567 | 108.4689 | 125.3716 | 94.5885 | 67.9203 | 235.3716 | 220.5885 | 233.9027 | 279.2567 | 185.8007 | 229.8628 | 192.5487 | 280.3142 | 444.6682 | 234.0797 | 244.8054 | 414.1946
2513 | 29.6374 | 150.8772 | 184.2508 | 226 | 183.3767 | 150.5015 | 182.7523 | 154.2518 | 164.5005 | 162.5015 | 144.7513 | 123.2518 | 162.7482 | 204.3787 | 134.1249 | 134.2457 | 211.5046 | 195.7503 | 158.6254 | 133.6223 | 167.4975 | 172.1249 | 371.3787 | 337.2497 | 153.998 | 229.4985 | 214.3756 | 202.3726 | 283.7523 | 104.3716 | 111.4975 | 121.001 | 127.4995 | 160.3787 | 138.999 | 121.1259 | 93.2497 | 267.8802 | 215.6244 | 220.8741 | 255.3777 | 163.2538 | 139.6244 | 148.001 | 129.7513 | 79.998 | 74.002 | 71 | 91.8731 | 236.6254 | 220.7523 | 231.7564 | 171.7492 | 146.8751 | 122.6254 | 124.5005 | 106.1259 | 169.1249 | 171.3736 | 190.1249 | 159.6254 | 163.8782 | 169.7543 | 163.8761 | 151.7462 | 173.8772 | 181.5005 | 216.2538 | 222.4944 | 159.1259 | 162.8792 | 216.8761 | 182.3746 | 193.7472 | 211.6284 | 205.3777 | 311.2579 | 165 | 158.3736 | 162.2528 | 223.3756 | 151.2497 | 168.5015 | 149.3767 | 158.998 | 228.1259 | 219.6305 | 179.6254 | 163.2497 | 118.1228 | 103.2457 | 118.5005 | 131 | 291.1249 | 276.8751 | 262.002 | 389.3818 | 295.5036 | 232.7482 | 262.002 | 285.2487 | 136.2487 | 82.5015 | 73 | 136.2487 | 135.7482 | 115.999 | 136.7533 | 128.5056 | 128.8751 | 112.7503 | 98.8751 | 117.001 | 93.2528 | 81.4985 | 82.6264 | 103 | 130.5036 | 102.9969 | 108.1269 | 96.6295 | 272.5056 | 244.6284 | 222.7482 | 259.2467 | 189.7513 | 202.871 | 206.002 | 276.6244 | 447 | 256.1279 | 225.9969 | 435.6284
2514 | 29.6484 | 145.9007 | 200.9735 | 194.9404 | 156.6026 | 152 | 167.2516 | 170.1987 | 146.3642 | 189.7351 | 134.4636 | 120.5629 | 178.7484 | 195.5761 | 142.8742 | 144.1987 | 208.0132 | 177.3642 | 128.8278 | 125.6755 | 192.51 | 169.3377 | 323.5297 | 310.3774 | 165.3113 | 220.0132 | 227.4239 | 206.2119 | 286 | 113.4239 | 114.3245 | 117.5629 | 119.0132 | 129.3907 | 128.2384 | 129.0993 | 92.1126 | 241.9404 | 226.5364 | 204.9139 | 260.6623 | 158.1258 | 144.3245 | 127.702 | 113.2516 | 85.0993 | 76.8874 | 71 | 110.4106 | 219.2516 | 188.3907 | 257.5232 | 192.298 | 148.7748 | 123 | 106.3642 | 105.2252 | 183.1987 | 175.3245 | 202.4239 | 170.649 | 162.5629 | 135.841 | 162.3377 | 166.6358 | 172.4503 | 188.2119 | 202.2516 | 261.3709 | 186.6226 | 164.3377 | 198.4768 | 182 | 187.4503 | 188.3774 | 202.6755 | 302.139 | 161.4503 | 183.6226 | 159.6755 | 230.2119 | 172.298 | 140.7152 | 145.6755 | 137.4768 | 201.49 | 229.4371 | 155.1523 | 155.0132 | 129.3113 | 120.298 | 121.6623 | 148.7484 | 289.2252 | 268.1258 | 224.9536 | 376.4768 | 331.8345 | 271.8213 | 259.5629 | 299.0861 | 135 | 82.2252 | 81.8742 | 135 | 145.5364 | 143.3974 | 123.139 | 126.9007 | 133.4371 | 113.8874 | 105.2119 | 140.1855 | 83.5761 | 84.4371 | 63.5894 | 94.1258 | 137.5497 | 130.1722 | 114.4371 | 83.2516 | 249.6026 | 215.1655 | 196.1523 | 248.0132 | 178.5761 | 214.0861 | 205.3377 | 279.5497 | 427.4768 | 263.4371 | 204.3642 | 464.7351
2515 | 29.6594 | 149.4985 | 197.6017 | 201.7965 | 175.4927 | 172.6933 | 130.811 | 135.1119 | 129.6046 | 166.0087 | 135.6991 | 122.6991 | 178.3009 | 218.2922 | 145.7994 | 163.0945 | 202.5015 | 186.6962 | 144.7936 | 133.9971 | 176.2064 | 161.8023 | 349.3895 | 298.0029 | 166.1003 | 211.8023 | 238.8968 | 232.1919 | 268.9055 | 131.1948 | 117.6991 | 108.9026 | 140.4927 | 129.4985 | 137.7965 | 156.0916 | 84.8023 | 271.2893 | 221.702 | 192.2035 | 226.811 | 177.6933 | 153.0974 | 129.4985 | 125.3954 | 80.6017 | 87.7965 | 78.1977 | 105.8023 | 252.0887 | 201.9942 | 230.3096 | 170.7078 | 155.298 | 141.8939 | 138.189 | 118.4956 | 171.5044 | 157.1061 | 212.0974 | 180.9971 | 162 | 156.1919 | 202.4869 | 165.4012 | 173.7994 | 203.3954 | 200 | 243.6076 | 162.109 | 137.9084 | 196 | 195.4956 | 219.3895 | 216.4898 | 178.6076 | 333.2893 | 165.4985 | 178.9026 | 143.7049 | 244.4956 | 163.3038 | 158.593 | 137.8023 | 156.593 | 228.5901 | 238.0974 | 197.8852 | 171.0945 | 124.702 | 136.4956 | 121.1003 | 119.5102 | 295.298 | 285.8939 | 244.2922 | 379.3983 | 338.6991 | 272.5015 | 284.1919 | 320.7936 | 143.9971 | 102.6933 | 67.7049 | 143.9971 | 118.2093 | 136.2035 | 97.6076 | 129.5988 | 121.4041 | 113.1003 | 114.9971 | 135.8023 | 82.8997 | 80.5015 | 74.4956 | 97.4985 | 120.0058 | 115.1061 | 99.7049 | 100.7936 | 272.9913 | 228.0945 | 191.2006 | 267.6933 | 182.3983 | 244.7907 | 211.298 | 243.1119 | 457.3895 | 264.8997 | 254.1832 | 459.0029
2516 | 29.6704 | 174.624 | 180.584 | 212.12 | 193.504 | 160.408 | 152.536 | 153.8 | 154.448 | 163.912 | 122.32 | 124.824 | 159.76 | 195.464 | 154.208 | 150.408 | 210.208 | 190.736 | 148.824 | 168.744 | 204.096 | 162.824 | 343.88 | 299.736 | 178.768 | 222.856 | 236.352 | 206.728 | 281.592 | 120.232 | 117.088 | 116.208 | 134.792 | 139.12 | 132.616 | 119.784 | 93.12 | 235.784 | 234.68 | 216.536 | 259.48 | 144.432 | 137.584 | 133.648 | 123.352 | 72.704 | 99.032 | 88.12 | 105 | 247.792 | 230.448 | 245.24 | 188.976 | 123.168 | 129.408 | 125.584 | 98.112 | 202.832 | 158.648 | 221.208 | 199.328 | 164.736 | 154.44 | 162.312 | 161.352 | 182.208 | 200.44 | 220.064 | 242.824 | 149.88 | 143.208 | 212.416 | 205.208 | 203.848 | 184.432 | 211.568 | 297.784 | 156.88 | 164.32 | 146.56 | 226.848 | 155.616 | 127.256 | 131.528 | 159 | 231.088 | 235.352 | 166.52 | 149.288 | 159.568 | 106.992 | 109.144 | 122.384 | 274.112 | 276.144 | 259.768 | 330.752 | 333.528 | 236.432 | 266.024 | 319.352 | 139.528 | 88.584 | 74.208 | 139.528 | 146.008 | 116.76 | 125.096 | 117.232 | 102.672 | 111.176 | 102.32 | 134.088 | 84.824 | 92.768 | 83.296 | 105.296 | 143.536 | 115.736 | 106.208 | 81.112 | 251.376 | 202.64 | 222.008 | 242.64 | 165.672 | 258.944 | 214.736 | 234.44 | 426.344 | 256.792 | 229.904 | 463.472
2517 | 29.6814 | 161.287 | 182.6972 | 213 | 181.1356 | 150.6814 | 169.7887 | 168.9401 | 174.5615 | 183.4101 | 126.5458 | 113.9085 | 149.6814 | 226.2745 | 136.5142 | 161.0158 | 218.3943 | 182.6814 | 160.0915 | 149.817 | 192.2113 | 156.53 | 336.53 | 318.4858 | 163.3627 | 220.3028 | 227.6814 | 204 | 267.287 | 125.47 | 132.713 | 79.104 | 127.53 | 138.1514 | 133.8486 | 120.6215 | 74.5899 | 271.7445 | 249.8644 | 228.2429 | 244.5142 | 135.4542 | 141.5458 | 157.1073 | 149.8045 | 88.6373 | 100 | 78.8328 | 135.5016 | 238.6814 | 223.7571 | 236.8328 | 166.9684 | 146.8045 | 140.9401 | 108.287 | 119.1073 | 193.9842 | 158.0757 | 210.9085 | 164.9526 | 169.6215 | 140.1356 | 200.5174 | 160.0757 | 195.0158 | 167.6498 | 228.47 | 240.2271 | 174.8802 | 169.8802 | 237.1073 | 206.9243 | 198.3028 | 196.713 | 182.6498 | 308.7887 | 193.896 | 185.183 | 161.7887 | 218.53 | 165.1672 | 151.7288 | 137.47 | 171.9401 | 243.0158 | 209.1198 | 196.2745 | 147.9243 | 124.1797 | 122.4858 | 110.7729 | 109.1356 | 276.6215 | 261.1356 | 263.7729 | 386.079 | 362.5773 | 245.0158 | 282.4858 | 344.8802 | 141.7729 | 91.6215 | 87.0158 | 141.7729 | 137.9085 | 134.4101 | 116.9085 | 119.6972 | 110.2429 | 135.0316 | 105.6215 | 141.3943 | 74.8328 | 94 | 91.3943 | 86.5899 | 120.1198 | 129.8644 | 115.3186 | 91.0158 | 244.3785 | 231.4259 | 211.1356 | 248.3186 | 188.0316 | 235.0441 | 190.9684 | 228.4542 | 460.896 | 290.1988 | 221.4542 | 483.4101
2518 | 29.6924 | 156.2537 | 190.4927 | 193.3318 | 165.0147 | 145.3171 | 145.7123 | 150.3318 | 147.9025 | 181.2537 | 130.7463 | 131.7317 | 179.9072 | 195.283 | 130.3171 | 141.3952 | 203.0781 | 171.6976 | 151.6342 | 134.8878 | 183.5073 | 158.8097 | 337.8732 | 290.9659 | 168.5561 | 233.1122 | 219.5073 | 193.6976 | 298.7804 | 116.6342 | 98.4099 | 111.5901 | 124.1903 | 138 | 127.4439 | 139.7317 | 95.478 | 257.2049 | 243.5073 | 241.1756 | 249.5561 | 135.9366 | 138.2537 | 137.4586 | 123.9025 | 64.7123 | 92.5073 | 91.1122 | 108.0294 | 218.3318 | 231.4292 | 203.2196 | 170.6195 | 142.4439 | 127.0147 | 108.8732 | 118.1903 | 192.0634 | 175.7951 | 214.6829 | 161.0634 | 132.5367 | 168.0341 | 165.6001 | 142.2049 | 161.3464 | 181.8585 | 188.7269 | 238.1268 | 191.0487 | 166.3805 | 175.3124 | 195.761 | 209.239 | 220.478 | 194.0487 | 359.6389 | 172.6488 | 157.0294 | 149.8878 | 203.0147 | 136.9659 | 148.3805 | 121.1415 | 147.7123 | 246.8097 | 218.239 | 181.2049 | 141.4439 | 117.2537 | 126.8097 | 112.8732 | 127.6682 | 270.4439 | 256.2537 | 256.5073 | 391 | 341.5854 | 250.6829 | 244.6635 | 312.3464 | 143.8732 | 83.5708 | 88 | 143.8732 | 117.3318 | 121.9513 | 124.4292 | 124.6829 | 125.0487 | 121.0781 | 111.6195 | 127.9513 | 75.8732 | 87.4439 | 66.7123 | 107.478 | 132.0487 | 102.9025 | 118.8097 | 87.3171 | 246.8097 | 228.3805 | 214.6829 | 267.7317 | 164.7123 | 238.6195 | 215.2243 | 260.7804 | 456.5073 | 287.3805 | 245.3512 | 448.4733
2519 | 29.7034 | 143.6646 | 175.818 | 204.3354 | 160.2045 | 156.3865 | 156.3354 | 158.4887 | 146.9489 | 193.3354 | 135.7444 | 133.9489 | 136.4541 | 204.3865 | 155.6196 | 136.2045 | 219.0797 | 208.0061 | 189.9038 | 130.2045 | 181.1023 | 157.1023 | 314.2782 | 283.3068 | 168.0511 | 221.6646 | 215.2045 | 238.5459 | 282.9714 | 123.591 | 128.2617 | 105.4601 | 132.5399 | 124.7158 | 125.1023 | 155.2331 | 98.8977 | 237.0225 | 234.4601 | 210.6872 | 221.5338 | 134.1023 | 131.3579 | 120.818 | 128.6421 | 75.3354 | 79.6646 | 85.3579 | 106 | 241.6707 | 230.1023 | 237.0572 | 201.364 | 110.6872 | 138.3354 | 140.3128 | 131.2842 | 183.4601 | 171.3068 | 203.6135 | 173.3354 | 167.0061 | 162.409 | 183.8752 | 148.591 | 178.9263 | 197.2331 | 206.8752 | 260.773 | 172.0737 | 154.6135 | 200.4151 | 187.409 | 204.3068 | 208.7158 | 193.1023 | 307.9653 | 203.2617 | 168.2842 | 149 | 213.3865 | 167.2617 | 159.3865 | 142.773 | 162.1309 | 247 | 256.955 | 187.591 | 133.409 | 108.4601 | 123.2045 | 114.8977 | 116.6646 | 248.1759 | 293.955 | 242.7158 | 374.8691 | 325.7669 | 279.4662 | 260.0286 | 338.4662 | 138.3068 | 83 | 93.6932 | 138.3068 | 127.3865 | 122.8977 | 125.9489 | 116.4601 | 133.591 | 131.3865 | 112 | 118.4601 | 76 | 83.2045 | 70.6932 | 120.3865 | 125.409 | 107.6421 | 99.0737 | 84.1534 | 258.3865 | 217.5624 | 198.8691 | 270.8977 | 189.5685 | 211.4827 | 205.6135 | 265.8466 | 443.6646 | 291.7444 | 267.8752 | 434.6135
2520 | 29.7144 | 185.2913 | 196.1456 | 185.7767 | 163.8447 | 138.7379 | 160.8447 | 153.233 | 159.4951 | 181.5049 | 136.9612 | 124.3884 | 159.9515 | 215.5728 | 150.2718 | 134.0777 | 215.1942 | 170.5922 | 154.5146 | 127.1165 | 180.0388 | 163.7282 | 321.6505 | 301.2621 | 152.6214 | 247.9126 | 234.2233 | 207.3592 | 278.1553 | 121.1165 | 141.534 | 128.068 | 148.3786 | 144.1845 | 115.3884 | 113.7087 | 86.5049 | 238.8835 | 245.534 | 208.0388 | 241.1456 | 164.7573 | 131 | 122.8835 | 125.1553 | 88.4951 | 89.5728 | 76.3495 | 110.8058 | 223.7767 | 233.8447 | 256.301 | 183.7767 | 146.4854 | 140.9223 | 135.2718 | 101.2427 | 169.5437 | 180.6116 | 190.5049 | 158.6214 | 156.5049 | 152.3884 | 166.7379 | 126.8932 | 183.8447 | 204.7282 | 187.8155 | 237.0097 | 164.2718 | 158.8058 | 216.4175 | 211.0291 | 201.1165 | 211.8447 | 177.6214 | 330.9515 | 182.8932 | 169.9612 | 155.7282 | 221.6893 | 148.8155 | 125.3981 | 146.8835 | 151.466 | 260.4563 | 212.8641 | 187.0388 | 137.8058 | 122.4175 | 108.5825 | 116.9223 | 137.1456 | 262.3786 | 238.3301 | 237.1942 | 379.767 | 323.0777 | 284.8447 | 276.3786 | 318.8544 | 136.0777 | 92.6116 | 77.6602 | 136.0777 | 134.7282 | 110.5049 | 143.301 | 134.2621 | 109.9709 | 100.2816 | 114.8835 | 128.5728 | 88.4951 | 79.1553 | 67.1553 | 111.3884 | 99.0485 | 124.3398 | 100.8835 | 89.767 | 267.6505 | 215.0777 | 194.1553 | 258.5049 | 200.6116 | 221.534 | 208.8447 | 242.932 | 400.7087 | 275.6602 | 247.8544 | 445.534
2521 | 29.7254 | 181.1593 | 220.3629 | 192.7876 | 166.9204 | 141.8938 | 174.6284 | 137.4247 | 162.9204 | 163.4778 | 144.7876 | 119.1327 | 154.1858 | 223.7876 | 122.7432 | 137.8938 | 255.8851 | 174.8407 | 147.1593 | 128.9469 | 180.9735 | 163.0265 | 304.4778 | 283.5043 | 164.6549 | 232.4512 | 247.6549 | 225.4691 | 290.6549 | 113.2124 | 109.876 | 109.5309 | 127.584 | 134.292 | 135.4426 | 124.6549 | 93.7876 | 268.2037 | 219.7167 | 239.1506 | 244.9204 | 141.6636 | 132.9469 | 125.9204 | 113.3185 | 89.9735 | 83.1858 | 68.2124 | 110.0265 | 226.8938 | 220.3716 | 219.0352 | 183.9735 | 135.3451 | 125.4247 | 130.1327 | 119.4691 | 179.708 | 186.8407 | 195.8407 | 182.3364 | 164.7611 | 160.7611 | 188.3895 | 156.1771 | 165.5043 | 172.876 | 181.1593 | 241.8407 | 157.1858 | 169.708 | 218.9469 | 185.7167 | 213.6549 | 198.3716 | 197.4426 | 331.0265 | 179.0796 | 182.6549 | 116.0883 | 222 | 181.0975 | 140.5488 | 116.8229 | 175.3364 | 216.221 | 209.0531 | 194.7876 | 171.0975 | 103.5309 | 126.4957 | 123.8142 | 111.7167 | 265.9204 | 271.0444 | 281.779 | 328.4068 | 330.7876 | 308.3629 | 286.7346 | 383.2215 | 141.8407 | 84.2389 | 87.708 | 141.8407 | 134.0265 | 138.2302 | 136.2124 | 132.0796 | 136.2568 | 110.6815 | 116.9469 | 123.1593 | 68.5574 | 72.1858 | 68.9469 | 115.8673 | 123.3099 | 119.1593 | 100.0265 | 84.1593 | 263.1327 | 215.9735 | 202.7611 | 264.8142 | 169.8494 | 200.584 | 217.7611 | 258.5488 | 442.8055 | 299.3364 | 245.0531 | 453.7876
2522 | 29.7364 | 148.4704 | 196.3563 | 193 | 180.8005 | 175.5154 | 160.2138 | 178.4014 | 160.0428 | 159.057 | 143.0285 | 120.9715 | 162.8717 | 235.829 | 159.4584 | 163.6294 | 226.4418 | 192.7434 | 133.1995 | 153.6437 | 168.1853 | 152.1568 | 302.0285 | 334.2588 | 141.3421 | 210.3136 | 246.0285 | 216.1425 | 265.3706 | 111.0285 | 108.0143 | 122.8005 | 128.9715 | 147.8005 | 130.0855 | 127.9572 | 88.0855 | 257.171 | 232.8005 | 219.2993 | 234.1568 | 140.0143 | 153.7007 | 123.0428 | 114.9715 | 71.2708 | 71.171 | 99.5439 | 97.1853 | 225.0285 | 237.7434 | 216.0285 | 180.057 | 140.9145 | 143.7292 | 119.1568 | 120 | 169.1568 | 192.9145 | 216.7007 | 145.5416 | 161.057 | 146.2138 | 191.9572 | 144.1853 | 208.3729 | 164.114 | 231.2731 | 248.9002 | 173.7577 | 175.9145 | 197.3136 | 199.7862 | 194.2851 | 162.5131 | 197.0143 | 311.2851 | 174.0713 | 160.3278 | 139.6437 | 198.3421 | 161.2993 | 160.7149 | 157.4014 | 169.0998 | 221.9002 | 206.0428 | 187.114 | 176.9287 | 116.8005 | 119.114 | 109.2138 | 116.9145 | 298.5296 | 271.0143 | 270.1853 | 365.4441 | 347.7577 | 313.9287 | 276.1568 | 383.0285 | 119.3278 | 77.0998 | 92.9287 | 119.3278 | 144.8432 | 134.0713 | 119.2423 | 136.9287 | 114.3278 | 118.886 | 99.2566 | 126.943 | 82.7862 | 77.9145 | 66.0428 | 110.0855 | 108.228 | 102.2423 | 100 | 91.886 | 236.3848 | 217.9715 | 200.0428 | 266.9715 | 186.7434 | 201.9715 | 203.2138 | 234.3563 | 397.6699 | 303.943 | 260.772 | 439.2138
2523 | 29.7474 | 149.9961 | 226.9392 | 194.9961 | 177.0078 | 149.053 | 156.0078 | 138.0804 | 132.0549 | 149.0196 | 151.9823 | 125.9902 | 155.0157 | 225.0216 | 145.0294 | 164.998 | 195.0608 | 195.9941 | 144.9765 | 131.0451 | 176.9823 | 159.9843 | 353.898 | 331.0078 | 163.9549 | 223.9725 | 203.0844 | 192.0471 | 296.9372 | 130.9608 | 125.9647 | 121.0039 | 123.0118 | 148 | 130 | 131.9922 | 97.9804 | 267.9784 | 247.9706 | 229.9784 | 222.0235 | 129.0216 | 148.0118 | 126.9922 | 135.9588 | 87.9666 | 77.9863 | 84.0314 | 108.9765 | 232.9843 | 225.0255 | 218.9941 | 168.0235 | 116.0491 | 146.9941 | 138.9608 | 114.0118 | 187.9627 | 159.0667 | 235.9627 | 146.9961 | 167.9863 | 162.9666 | 161.0608 | 150.9863 | 182.053 | 194.9392 | 225.0137 | 213.0706 | 166.0157 | 159.0334 | 207.9784 | 185.0294 | 212.9627 | 219.8862 | 183.0275 | 287.0471 | 159.0294 | 179.9608 | 176.9274 | 195.0059 | 159.0039 | 134.053 | 135.0451 | 164.0098 | 212.0196 | 223.9647 | 189.9941 | 140.0726 | 103.0275 | 120.9961 | 99.0196 | 106.0216 | 267.0628 | 273.9941 | 257.0255 | 389.9529 | 329.0373 | 330.9666 | 283.9843 | 419.9274 | 118.002 | 75.0039 | 81.0235 | 118.002 | 140.0098 | 120.0275 | 115.0078 | 129.0157 | 121.9843 | 103.0314 | 101.9941 | 108.0373 | 87.9902 | 90.9745 | 65.002 | 123.9725 | 117.9804 | 127.949 | 105.9882 | 97.9882 | 230.0118 | 221.9922 | 216.9666 | 231.0706 | 204.9647 | 193.0177 | 189.0275 | 239.9882 | 397 | 301.0059 | 280.9608 | 408.0608
2524 | 29.7584 | 187.6488 | 185.9484 | 192.9484 | 155.062 | 137.2066 | 159.2582 | 163.062 | 154.8761 | 162.7211 | 130.1653 | 111.969 | 184.6075 | 258.6695 | 165.1033 | 173.7211 | 209.0103 | 178.9587 | 161.8554 | 139.2066 | 147.2789 | 173.9897 | 316.3305 | 304.969 | 118.5165 | 219.9587 | 208.2686 | 220.1343 | 291.6798 | 100.1239 | 107.093 | 123.8141 | 121.9897 | 142.0207 | 133.1756 | 128.2789 | 86.0516 | 236.4235 | 256.0103 | 236.9587 | 227.2995 | 130.1549 | 164.8347 | 126.8657 | 122.0413 | 90.8761 | 69.1446 | 91.8244 | 97.0207 | 231.8761 | 256.4939 | 244.7728 | 166.9484 | 152.7314 | 120.0207 | 141.7831 | 116.0723 | 168.9793 | 149.3305 | 196.7005 | 153.0207 | 134.1343 | 151.1549 | 162.0103 | 148.8347 | 214.9587 | 198.9587 | 204.1033 | 243.907 | 154.0723 | 162.8451 | 210.1033 | 192.9277 | 224.7934 | 147.6094 | 187.1756 | 300.1239 | 173.8761 | 168.7831 | 158.062 | 215.8554 | 147.0516 | 144.0413 | 130.3409 | 144.969 | 199.1653 | 193.9897 | 201.7624 | 137 | 118.9897 | 120.9174 | 127.969 | 108.3615 | 282.8554 | 249.907 | 221.2066 | 356.9897 | 344.7831 | 329.1446 | 244.2789 | 375.5268 | 133.938 | 85.9174 | 84.062 | 133.938 | 108.3512 | 117.969 | 119.062 | 121.8347 | 117.907 | 118.0103 | 118.9277 | 114 | 69.0413 | 79.9793 | 73.0103 | 119.8347 | 108.1446 | 105.9793 | 120.6695 | 83.093 | 255.8451 | 229.9484 | 229.6075 | 269.6075 | 178.1136 | 231.5249 | 207.7934 | 238.969 | 371.1033 | 285.8347 | 282.9174 | 405.6695
2525 | 29.7694 | 153.7512 | 181.1357 | 188.181 | 161.4072 | 157.5881 | 183.0273 | 168.7738 | 142.8643 | 136.8596 | 145.3666 | 109.3167 | 147.4976 | 226.7964 | 174.3893 | 146.7512 | 209.7286 | 174.5702 | 147.6155 | 159.1357 | 174.0226 | 172.8869 | 347.8417 | 302.4298 | 167.525 | 216.4072 | 233.4571 | 231.8238 | 261.4976 | 112.4072 | 115.9321 | 105.9548 | 120.9548 | 143.5476 | 149.5476 | 154.6833 | 90.7738 | 276.6381 | 257.0679 | 232.3893 | 256.0679 | 145.1131 | 148.7738 | 114.4976 | 126.0679 | 78.9548 | 82.9774 | 74.7512 | 99.3167 | 220.3167 | 208.5203 | 222.7286 | 162.1131 | 127.3393 | 122.2941 | 121.1583 | 122.525 | 166.7964 | 181.0452 | 169.2893 | 155 | 146.8869 | 166.1131 | 163.0905 | 133.0452 | 211.3619 | 194.4797 | 213.525 | 234.9774 | 160.5024 | 148.0226 | 219.7964 | 185.9548 | 204.8417 | 205.4797 | 203.3666 | 310.9821 | 162.181 | 148.3167 | 163.7059 | 202.4298 | 151.5024 | 148.0452 | 162.5024 | 142.2488 | 214.7964 | 193.0905 | 179.5429 | 137.4524 | 117.7059 | 113.2262 | 125.1357 | 142.3893 | 269 | 241.3393 | 240.7512 | 355.9774 | 324.0905 | 342.525 | 271.4976 | 425.4119 | 128.5655 | 77.8417 | 89.7964 | 128.5655 | 141.3666 | 115.0452 | 124.8417 | 105.819 | 109.4524 | 119 | 112.4072 | 113.7286 | 73.0679 | 78.2036 | 73.9321 | 104.2262 | 121.6607 | 104.2262 | 89.0452 | 91.7512 | 241.181 | 224.9774 | 192.4524 | 232.2262 | 188.3666 | 186.6334 | 188.2488 | 237.0179 | 380.3214 | 270.6334 | 274.7964 | 374.3845
2526 | 29.7804 | 144.4663 | 187.5935 | 195.8604 | 178.9302 | 182.0923 | 141.5237 | 158.3367 | 137.2444 | 173.4065 | 118.9775 | 122.2669 | 168.5462 | 217.7905 | 148.4538 | 137.187 | 198.5237 | 156.1746 | 131.6982 | 163.9876 | 174.8254 | 168.2095 | 341.3142 | 320.5112 | 147.2095 | 233.8254 | 210.384 | 183.1296 | 282.1272 | 129.3367 | 113.5935 | 104.1746 | 119 | 124.6982 | 130.768 | 140.5462 | 81.3142 | 261.1047 | 259.2669 | 206.384 | 259.5237 | 149.7556 | 138.1621 | 135.4414 | 128.6509 | 76.8953 | 81.5462 | 65.3266 | 112.2669 | 232.813 | 230.3018 | 211.5586 | 166.4763 | 142.0349 | 135.5586 | 127.8953 | 102.2095 | 158.5586 | 183.1396 | 223.7432 | 154.9651 | 142.384 | 170.6858 | 166.8604 | 135.5586 | 225.6385 | 172.6633 | 194.3266 | 234.4189 | 139 | 147.8828 | 211.4888 | 184.2793 | 198.5237 | 183.4538 | 177.2219 | 267.9426 | 170.0349 | 161.5112 | 150.7905 | 219.9527 | 130.9077 | 149.4763 | 140.6858 | 152.9302 | 205.7556 | 196.616 | 202.3367 | 155.7781 | 105.2444 | 123.1746 | 130.197 | 116.4189 | 268.8604 | 255.0225 | 230.1746 | 355.5237 | 327.8254 | 322.2444 | 292.6509 | 400.0698 | 151.813 | 71.8379 | 81.5237 | 151.813 | 113.7905 | 117.4888 | 118.0349 | 98.3142 | 127.7432 | 119.7331 | 129.6858 | 102 | 76.1746 | 86.6858 | 71.4538 | 114.1396 | 106.7207 | 114.2444 | 91.0349 | 80.5811 | 248.1621 | 222.7083 | 211.616 | 242.5586 | 161.2095 | 213.7207 | 199.0349 | 279.7432 | 352.0124 | 298.2095 | 266.5237 | 391.2095
2527 | 29.7914 | 183.9143 | 203.8112 | 192.6608 | 176.5752 | 157.6608 | 156.472 | 140.1888 | 144.9912 | 156.8584 | 145.2448 | 103.0384 | 155.6224 | 212.236 | 160.1032 | 170.1888 | 213.0944 | 160.6224 | 151.2832 | 136.1888 | 170.5664 | 173.764 | 348.4895 | 304.7815 | 153.708 | 228.0088 | 220.056 | 240.4895 | 260.3129 | 110.9056 | 129.9528 | 109.3304 | 119.1416 | 143.764 | 151.9056 | 127.2448 | 91.3217 | 262.2535 | 239.3776 | 216.3864 | 273.528 | 144.3217 | 115.472 | 120.3304 | 119.7552 | 74.5664 | 69.7552 | 101.056 | 92.1888 | 200.1416 | 211 | 227.2832 | 152.3304 | 141.9143 | 150.4336 | 124.1504 | 108.472 | 173.1032 | 187 | 190.4161 | 153.7168 | 152.292 | 162.3304 | 162.9056 | 150.8112 | 189.2745 | 190.6224 | 229.4895 | 244.4895 | 139.1888 | 118.4633 | 224.5752 | 191.9056 | 212.2448 | 195.8584 | 209.0647 | 294.3776 | 169.8199 | 148.3776 | 145.944 | 191.1888 | 154.5839 | 135.0944 | 131.6224 | 151.1416 | 199.6136 | 186.236 | 184.944 | 122.6136 | 112.6136 | 127.0088 | 108.6136 | 129.1801 | 263.7727 | 228.9912 | 235.3776 | 369.0088 | 321.8671 | 329.1416 | 285.7377 | 402.236 | 119.5192 | 95.0472 | 95.3392 | 119.5192 | 108.4248 | 130.6696 | 118.3864 | 107.6136 | 93.472 | 137.9703 | 119.6783 | 103.6521 | 81.1416 | 79.1329 | 83.6224 | 117.292 | 98.4808 | 120.1032 | 92.8968 | 70.1801 | 224.9528 | 188.7465 | 201.5192 | 256.7255 | 167.708 | 204.6783 | 199.6224 | 243.8199 | 378.4895 | 302.5839 | 281.3776 | 394.9703
2528 | 29.8024 | 163.6063 | 199.2861 | 206.1785 | 166.5722 | 170.8215 | 166.2975 | 145.3683 | 164.2266 | 151.7987 | 131.0114 | 125.4278 | 148.357 | 215.5722 | 144.4987 | 173.3456 | 213.8101 | 155.9152 | 156.2266 | 140.5949 | 182.7139 | 168.2266 | 318.1785 | 263.1531 | 168.5949 | 210.9152 | 200.762 | 210.8329 | 304.2633 | 109.5354 | 126.8582 | 115.5241 | 122.238 | 139.238 | 148.9291 | 112.9519 | 116.3937 | 228.0114 | 244.7393 | 205.6658 | 263.5835 | 168.6798 | 125.2975 | 125.3937 | 132.9177 | 85.2266 | 83.9886 | 82.4165 | 97.1899 | 202.5835 | 210.7025 | 232.0481 | 159.119 | 119.5241 | 139.2975 | 107.9519 | 117.4646 | 155.6544 | 187.357 | 216.6203 | 147.762 | 138.0595 | 167.0367 | 162.3683 | 146.9405 | 213.4532 | 184.0114 | 198.3456 | 213.643 | 144.1899 | 148.0595 | 215.0481 | 189.1671 | 197 | 194.3089 | 171.0709 | 301.4051 | 146.9519 | 156.119 | 164.643 | 195.7139 | 127.3683 | 137.7734 | 124.6544 | 153.1671 | 210.8101 | 190.2266 | 202.6317 | 135 | 124.0481 | 108.0709 | 121.8924 | 150.1443 | 237.9291 | 249.4165 | 241.6317 | 348.643 | 297.2152 | 329.1443 | 337.6684 | 407.2975 | 129.643 | 95.5241 | 81.8215 | 129.643 | 117.1785 | 122.9886 | 107.0114 | 119.5835 | 103.4759 | 98.8443 | 94.2494 | 135.7506 | 85.3089 | 101.4646 | 75.1076 | 102.0481 | 88.2975 | 101.762 | 109.9291 | 93.643 | 221.7987 | 224.9519 | 210.2152 | 233.3202 | 181.2861 | 179.0114 | 193.4278 | 221.4278 | 350.4987 | 274.238 | 285.0734 | 353.0481
2529 | 29.8134 | 189.0718 | 190.0817 | 211.01 | 144.6461 | 169.2921 | 171.8614 | 166.7846 | 152.7896 | 121.0917 | 147.2871 | 145.8465 | 153.3539 | 196.3021 | 183.3539 | 162.9282 | 195.5025 | 200.5643 | 145.4357 | 149.8564 | 192.2054 | 155.7129 | 323.2253 | 312.9282 | 175.7029 | 254.1286 | 196.9282 | 221.9183 | 264.4456 | 117.0668 | 94.2921 | 108 | 126.2154 | 142.6411 | 131.4975 | 129.005 | 91.7178 | 243.7129 | 209.1436 | 232.5743 | 255.8515 | 133.5075 | 129.9282 | 102.656 | 101.5075 | 74.2203 | 69.5075 | 89.2154 | 114.4925 | 196.5025 | 206.2154 | 217.1436 | 160.7846 | 114.1535 | 142.6361 | 123.1436 | 108.7846 | 167.0768 | 191.995 | 181.0817 | 143.7129 | 136.7747 | 136.4307 | 182.4925 | 146.4307 | 190.1485 | 198.1336 | 188.4307 | 208.1436 | 162.3539 | 150.2203 | 200.6461 | 176.3589 | 195.6361 | 212.99 | 186.8515 | 291.3539 | 161.4975 | 158.8614 | 159.2871 | 205.5643 | 148.4257 | 151.0768 | 135.3589 | 139.0668 | 191.1386 | 178.6461 | 180.4257 | 135 | 108.7129 | 123.4207 | 134.8515 | 106.9382 | 222.005 | 254.9232 | 220.5743 | 342.7846 | 269.2154 | 285.1485 | 284.6411 | 409.9183 | 125.2203 | 87.2822 | 80.8664 | 125.2203 | 120.5743 | 108.3639 | 123.8614 | 114.0768 | 111.2871 | 127.9282 | 112.8515 | 115.6411 | 104.1336 | 91.1336 | 61.6461 | 88.651 | 94.3639 | 97.7846 | 94.5075 | 87.8564 | 192.1685 | 239.0668 | 183.7946 | 266.4108 | 171.005 | 195.7178 | 214.1336 | 243.8564 | 387.9183 | 278.1436 | 225.8714 | 337.3539
2530 | 29.8244 | 187.8981 | 215.7389 | 236.1592 | 154.3452 | 183.73 | 181.907 | 162.3185 | 160.4777 | 171.5618 | 150.9159 | 120.6904 | 145.1682 | 235.8892 | 173.1503 | 162 | 200.907 | 182.4204 | 163.0752 | 147.5796 | 168.4955 | 152.5885 | 348.0485 | 312.2522 | 144.1503 | 213.8892 | 198.4382 | 198.5312 | 305.1325 | 106.093 | 111.2522 | 106.8229 | 128.5796 | 137.5796 | 126.6815 | 141.1592 | 101.5045 | 238.5707 | 211.8408 | 237.5618 | 242.6815 | 153.1682 | 130.6815 | 136.3274 | 120.2433 | 90.0841 | 89.6726 | 90.9911 | 95.5885 | 204.5134 | 207.907 | 218.0752 | 157.2433 | 140.3185 | 124.6637 | 125.8408 | 105.8318 | 178.0574 | 179.3363 | 208.5045 | 139.1592 | 111.1949 | 142.2522 | 161.3185 | 150.4026 | 204.0752 | 176.0178 | 192.6548 | 209.4955 | 152.907 | 163.73 | 208.8318 | 180.6637 | 178.9248 | 186.7478 | 170.9911 | 284.2611 | 155.3363 | 168.9911 | 162.3274 | 188.3452 | 140.5796 | 162.3937 | 138.907 | 129.27 | 179.1592 | 187.5885 | 172.9159 | 133.0663 | 105.4204 | 103.5885 | 118.1503 | 130.5707 | 235.2522 | 240.7478 | 230.4382 | 340.0841 | 272.7567 | 301.2611 | 282.5223 | 382.6637 | 139.0663 | 78.6726 | 102.73 | 139.0663 | 125.8981 | 124.7389 | 133.5707 | 127.0752 | 114.0752 | 125.3185 | 97.5796 | 111.9248 | 78.3185 | 68.3452 | 71.093 | 110.2522 | 110.5707 | 95.0841 | 112.4026 | 85.8318 | 254.3759 | 225.8229 | 205.0663 | 222.0178 | 180.2166 | 206.5134 | 191.2611 | 241.2433 | 360.5796 | 276.5529 | 259.9733 | 330.0089
2531 | 29.8354 | 163.0727 | 203.7346 | 224.0126 | 169.6746 | 158.9036 | 169.3254 | 145.6382 | 135.6019 | 146.1564 | 146.6272 | 147.109 | 146.6145 | 202.7709 | 156.3728 | 160.6509 | 192.0837 | 186.1327 | 152.2291 | 143.2891 | 165.0237 | 159.9637 | 304.8073 | 314.5182 | 126.2164 | 187.8057 | 224.3254 | 233.3381 | 262.9036 | 116.8436 | 112.5545 | 94.4818 | 124.2891 | 132.0363 | 144.6145 | 131.2291 | 105.8436 | 225.0237 | 221.0964 | 211.2891 | 260.9036 | 155.5782 | 148.3254 | 129.4818 | 111.5182 | 90.2291 | 95.6509 | 80.0964 | 101.6145 | 217.2417 | 197.831 | 207.9036 | 149 | 122.1927 | 121.9637 | 131.7235 | 106.6019 | 146.3855 | 181.6509 | 208.9889 | 131.6382 | 144.4581 | 144.8073 | 145.0237 | 135.2164 | 194.6746 | 197.4218 | 178.8673 | 205.6382 | 140.5182 | 139.8673 | 207.1927 | 177 | 189 | 183.1327 | 161.7346 | 299.06 | 158.9036 | 156.6509 | 155.5782 | 200.0126 | 135.7109 | 135.9273 | 129.4092 | 152.5545 | 170.5782 | 191.9763 | 173.6382 | 113.3491 | 112.4092 | 111.2528 | 98.3855 | 114.6145 | 238 | 240.1201 | 253.7235 | 341 | 280.2291 | 312.109 | 307.4945 | 379.3855 | 118.0964 | 86.6746 | 78.6746 | 118.0964 | 104.2528 | 111.4818 | 117.7109 | 118.1564 | 103.9036 | 107.1927 | 93.9637 | 120.6509 | 61.06 | 81.5545 | 84.4455 | 111.6509 | 96.3491 | 98.6019 | 95.1927 | 82.6509 | 207.6619 | 211.3618 | 185.4455 | 245.3491 | 144.5893 | 220.7836 | 203.7472 | 230.9763 | 357.06 | 237.7472 | 225.7472 | 342.5419
2532 | 29.8464 | 146.9558 | 219.1307 | 195.4567 | 175.8913 | 156.1528 | 163.6519 | 158.5009 | 161.7385 | 155.3701 | 116.4125 | 121.7385 | 145.7164 | 208.9134 | 183.9797 | 151.1511 | 219.3701 | 177.674 | 145.5433 | 145.4567 | 183.8913 | 165.523 | 303.4346 | 306.7403 | 145.1528 | 258.2836 | 215.1749 | 200.9558 | 263.9558 | 107.5212 | 100.1952 | 102.5857 | 126.8913 | 123.6519 | 140.4567 | 123.2394 | 95 | 244.5433 | 223.7385 | 210.0883 | 260.1087 | 158.3922 | 140.2615 | 133.7827 | 109.9338 | 83.8693 | 84.4125 | 79.9134 | 98.9779 | 182.7606 | 212.0662 | 206.0221 | 148.8913 | 126.4991 | 129.0442 | 103.2818 | 128.0662 | 154.2376 | 171.3905 | 164.2818 | 145.2615 | 131.6299 | 144.1952 | 162.8048 | 153.9355 | 201.9779 | 191.0221 | 188.8472 | 222.0866 | 138.2818 | 147.7827 | 207.9134 | 178.1952 | 188.2394 | 178.6944 | 176.2615 | 310.5212 | 157.5654 | 145.3039 | 160.1307 | 174.2818 | 131.5875 | 150.6316 | 149.0662 | 139.2173 | 175.5654 | 172.1307 | 186.1749 | 125.1307 | 132.3922 | 119.523 | 104.2818 | 109.6961 | 238.9779 | 257.7182 | 227.0203 | 339.0442 | 274.1952 | 285 | 286.1731 | 383.7606 | 121.2818 | 91.3701 | 84.0221 | 121.2818 | 113.6095 | 115.2394 | 114.4567 | 126.6095 | 102.7827 | 109.8693 | 102.7827 | 110.3905 | 71.5433 | 69.1952 | 96.6961 | 98.8913 | 108.4567 | 120.0662 | 98.8472 | 70.4346 | 244.2818 | 200.8896 | 199.7606 | 257.8913 | 193.197 | 199.0221 | 191.2394 | 213.4125 | 363.6316 | 226.9779 | 215.8472 | 352.4364
2533 | 29.8574 | 164.2095 | 209.3391 | 191.8466 | 175.2419 | 141.8466 | 166.702 | 140.6609 | 170.1945 | 146.596 | 129.8142 | 136.8466 | 164.7344 | 200.6047 | 150.5162 | 176.5162 | 205.7581 | 172.2182 | 150.3628 | 140.0324 | 179.7344 | 138.6933 | 303.6134 | 282.6609 | 131.0885 | 232.2506 | 193.9352 | 216.6372 | 282.298 | 120.8466 | 109.6372 | 127.8877 | 125.5162 | 130.3304 | 134.1858 | 120.8704 | 96.9352 | 251.0561 | 236.4277 | 185.015 | 256.1621 | 136.0324 | 130.596 | 131.8791 | 134.4838 | 90.2743 | 95.3953 | 71.9676 | 105.6696 | 189.9676 | 188.1209 | 200.298 | 150.5399 | 145.5486 | 114.3304 | 122.9676 | 104 | 185.8553 | 187.3715 | 185.7818 | 131.8466 | 144.1534 | 153.6372 | 154.8142 | 136.0324 | 208.4277 | 182.7581 | 202.4277 | 228.581 | 153.3715 | 146.1209 | 200.0885 | 191.5075 | 183.0885 | 207.3067 | 162.2419 | 317.7992 | 155.2095 | 157.2095 | 150.581 | 192.2743 | 120.9676 | 123 | 126.3304 | 139.7905 | 169.4601 | 162.8229 | 165.1771 | 117.7581 | 112.4514 | 92.3304 | 120.702 | 103.2332 | 251.2332 | 239.1209 | 260.2419 | 321.7905 | 277.7107 | 288.0237 | 300.3715 | 385.8877 | 138.4277 | 78.7257 | 77.3304 | 138.4277 | 97.5075 | 110.3304 | 111.9352 | 108.4514 | 101.6047 | 115.5486 | 101.2419 | 127.4601 | 75.2743 | 80.6933 | 86.2419 | 97.5162 | 103.7581 | 99.7494 | 114.9676 | 75.6933 | 262.6372 | 234.3478 | 207.8142 | 259.419 | 162.2419 | 189.3067 | 185.3628 | 225.3628 | 334.7905 | 233.0648 | 231.8466 | 321.8941
2534 | 29.8684 | 173.1332 | 197.3282 | 195.202 | 176.0673 | 142.1375 | 151.1992 | 161.0659 | 130.2651 | 173.2048 | 142.3338 | 142.467 | 144.7307 | 204.8668 | 147.6662 | 169.669 | 202.9341 | 152.533 | 150.6017 | 131.2679 | 156.7994 | 152.9986 | 285.129 | 305.7307 | 137.2679 | 207.1318 | 210.1318 | 211.8682 | 293.9369 | 123.5358 | 105.9341 | 100.5316 | 124.5316 | 141.8654 | 121.2665 | 146.4684 | 110.0673 | 261.8696 | 229.1304 | 232.5358 | 228.9944 | 134.0631 | 158.2707 | 128.7349 | 136.8008 | 86.7321 | 87.5358 | 79.533 | 98.3983 | 194.0687 | 187.6676 | 215.4011 | 169.1332 | 135.9327 | 121.3352 | 130.6662 | 105.4656 | 162.0673 | 160.6003 | 202.0028 | 139.5989 | 137.2006 | 149.9341 | 167.6003 | 128.467 | 195.0014 | 180.3338 | 192.7321 | 210.0673 | 130.0645 | 145.0014 | 204.2693 | 215.2679 | 193.5316 | 192.8682 | 167.331 | 280.533 | 166.9313 | 165.0673 | 135.3983 | 186.7335 | 130.5316 | 122.8668 | 137.4656 | 134.5975 | 154.5975 | 150.4642 | 178.6017 | 116.6662 | 122.6003 | 103.5989 | 104.3997 | 130.4025 | 276.9369 | 238.5344 | 258.0028 | 319.5287 | 232.2665 | 307.3352 | 273.7335 | 356 | 127.533 | 83.8668 | 88.1992 | 127.533 | 125.3983 | 117.7349 | 124.4011 | 117.9341 | 106.2665 | 106.4656 | 102.6003 | 109.7994 | 72.1318 | 89.4025 | 84.1361 | 93.3338 | 101.2006 | 124.6017 | 117.7363 | 86.0014 | 253.2048 | 185.6003 | 216.8696 | 272.0701 | 165.0659 | 177.1332 | 185.202 | 228.3997 | 327.8654 | 220.7321 | 242.5301 | 370.0014
2535 | 29.8794 | 178.2204 | 237.1413 | 180.9293 | 166.2162 | 112.5863 | 160.3098 | 164.9439 | 142.0894 | 128.8545 | 137.4179 | 136.2349 | 167.8358 | 203.8545 | 152.7277 | 149.4553 | 197.7464 | 160.2204 | 137.7651 | 122.9106 | 165.9293 | 164.6902 | 330.0936 | 326.5073 | 129.0561 | 218.817 | 219.343 | 200.474 | 263.6196 | 104.2017 | 99.8732 | 113.9439 | 140.8545 | 152.1081 | 125.474 | 130.7277 | 103.8545 | 242.5676 | 252.0707 | 211.7464 | 276.343 | 161.9064 | 122.8919 | 118.8025 | 131.183 | 97.5634 | 65.4366 | 84.3098 | 113.1268 | 182.5676 | 181.474 | 201.5073 | 169.5634 | 139.6715 | 109.2391 | 131.7983 | 112.9626 | 156.4366 | 160.6196 | 173.9626 | 148.5447 | 131.5634 | 142.5634 | 163.1081 | 129.948 | 184.0374 | 174.1081 | 202.3992 | 207.6383 | 146.6528 | 131.2723 | 187.7131 | 201.817 | 205.6341 | 179.8732 | 185.3617 | 284.1455 | 182.7983 | 157.9813 | 148.2349 | 184.2723 | 146.2723 | 122.2911 | 142.343 | 156.3992 | 174.9439 | 171.8358 | 164.6008 | 119.8358 | 119.2723 | 113.7089 | 106.2723 | 109.9106 | 243.4179 | 234.2391 | 233.8919 | 356.0332 | 235.8919 | 291.8919 | 275.3472 | 356 | 130.4179 | 81.9813 | 89.8877 | 130.4179 | 138.817 | 103.4366 | 109.7796 | 110.4179 | 106.9813 | 116 | 104.657 | 116.1642 | 82.0707 | 67.7464 | 61.6196 | 87.5447 | 97.7464 | 107.1081 | 91.8919 | 74.6008 | 210.4553 | 188.8212 | 190.1455 | 238.5447 | 168.7983 | 178.2911 | 172.6757 | 231.4366 | 337.0894 | 239.5676 | 265.0145 | 353.9439
2536 | 29.8904 | 196.5293 | 194.6299 | 204.4735 | 149.0503 | 160.0014 | 164.8436 | 147.7891 | 125.9469 | 131.7877 | 131.6327 | 120.7891 | 159.264 | 201.5796 | 157.4735 | 157.5265 | 203.1075 | 180.4232 | 152.1061 | 138.264 | 186.4749 | 156.0531 | 350.8463 | 302.8436 | 149.0517 | 208.5251 | 190.5796 | 212.0014 | 272.6871 | 120.317 | 105.7891 | 95.0531 | 137.6327 | 141.6313 | 141.4204 | 133.264 | 99.5279 | 276.8994 | 227.4218 | 222.7891 | 247.264 | 133.736 | 135.1048 | 143.8436 | 138.7919 | 91.0545 | 71.4721 | 90.8952 | 109.5782 | 220.0558 | 196.1578 | 184.3143 | 163.8436 | 128.0517 | 139.3687 | 111.7374 | 101.4735 | 165.4707 | 176.7891 | 162.1578 | 140.1578 | 127.1061 | 137.0014 | 152.3156 | 154.2653 | 194.1061 | 164.5782 | 191.4218 | 223.3184 | 126.0531 | 129.0517 | 229.0545 | 185.0545 | 181.1048 | 187.0517 | 169.2081 | 283.7374 | 162.1061 | 149.6327 | 134.8408 | 187.102 | 139.317 | 124.7891 | 120.9972 | 148.1048 | 160.9455 | 164.8422 | 177.1564 | 115.5251 | 115.1578 | 112.3156 | 103.2626 | 121.4762 | 236.5279 | 261.3701 | 244.2109 | 321.6313 | 245.264 | 296.5293 | 293.7374 | 353.0014 | 125.1061 | 78.3673 | 59.2081 | 125.1061 | 122.8436 | 108.8408 | 90.5238 | 105.8952 | 98.317 | 112.3701 | 125.5293 | 119.8436 | 60.4204 | 78.4735 | 77 | 82.3143 | 109.5782 | 97.1048 | 105.3673 | 86.8408 | 217.1061 | 219.8436 | 195.7347 | 235.5238 | 153.3143 | 185.3659 | 205.3687 | 232.4218 | 318.7374 | 270.743 | 225.2626 | 333.1592
2537 | 29.9014 | 163.0207 | 214.8506 | 207.6805 | 174.9793 | 142.8092 | 150.5517 | 154.5517 | 128.1081 | 154.0414 | 121.5517 | 125.3402 | 148.1288 | 192.469 | 160.1701 | 150.2368 | 181.9333 | 154.8092 | 140.6391 | 127.9793 | 173.3402 | 152.7011 | 304.8919 | 286 | 156.938 | 223.8712 | 184.531 | 200.2528 | 242.4023 | 111.6137 | 107.2782 | 91.8712 | 120.7472 | 143.6391 | 146.6184 | 124 | 81.3402 | 225.8919 | 220.0207 | 227.8506 | 235.4483 | 138.5771 | 135.8966 | 126.3195 | 113.5517 | 71.7011 | 88.469 | 87.7218 | 116.2989 | 185.2115 | 192.7472 | 200.1288 | 146.1494 | 135.2575 | 133.1081 | 109.2528 | 105.3609 | 196.2575 | 174.5357 | 169.2942 | 141.8506 | 121.9333 | 124.4023 | 154 | 133.7632 | 182.9793 | 171.6391 | 181.1288 | 195.5517 | 127.1241 | 140.8506 | 206.9793 | 165.7011 | 186.8299 | 193.0667 | 193.7679 | 278.1908 | 149.7885 | 140.5724 | 149.6598 | 216.8346 | 123.9793 | 132.2321 | 149.4276 | 151.7885 | 180.469 | 167.2321 | 188.8966 | 129 | 115.8299 | 111.7885 | 111.0207 | 99.3148 | 225.3148 | 240.2989 | 240.6805 | 323.469 | 236.1701 | 264.0414 | 287.5103 | 334.1081 | 116.8712 | 87.7679 | 81.8966 | 116.8712 | 107.531 | 122.1288 | 117.4483 | 99.6598 | 90.4643 | 94.3609 | 94.062 | 104.8712 | 67.3195 | 84.4023 | 76.3195 | 98.2989 | 114.5977 | 102.3195 | 111.5357 | 100.1288 | 210.062 | 207.2528 | 220.4897 | 255.8139 | 170.4897 | 207.8759 | 199.2782 | 227.9126 | 313.531 | 202.8919 | 231.8299 | 315.9586
2538 | 29.9124 | 168.7296 | 220.824 | 210.088 | 168.5408 | 139.0064 | 161.176 | 164.0815 | 124.9442 | 160.1695 | 131.6288 | 127.7296 | 140.4592 | 185.3648 | 161.7296 | 130.1009 | 204.5408 | 152.8305 | 134.3648 | 128.6545 | 173.7232 | 157.5343 | 318.088 | 287.0944 | 143.2768 | 229.1695 | 196.9249 | 219.3584 | 254.4399 | 131.7038 | 99.1073 | 101.912 | 105.8369 | 136.6352 | 132.6288 | 125.2768 | 86.1009 | 237.8112 | 223.176 | 233.2768 | 227.0129 | 119.0064 | 114.8305 | 125.1888 | 121.8047 | 78.7232 | 78.2639 | 97.1695 | 109.6416 | 194.3584 | 170.5408 | 193.9185 | 144.1888 | 119.5537 | 116.0815 | 127.8112 | 112.3648 | 171.6159 | 143.7296 | 198.5408 | 149.2833 | 139.2511 | 141.912 | 155.2768 | 152.0751 | 181.2833 | 165 | 178.0193 | 213.1073 | 157.1824 | 148.2833 | 198.5343 | 176.5537 | 182.5343 | 171.3712 | 181.4657 | 284 | 138.088 | 157.2704 | 149.4592 | 178.824 | 123.0129 | 140.8863 | 134.3712 | 140.6352 | 174.0944 | 185.5537 | 164.7296 | 130.2768 | 112.2639 | 105.1953 | 114.176 | 130.8047 | 256.4399 | 231.088 | 244 | 318.1888 | 238.0944 | 271.8112 | 284.3455 | 323.8305 | 128.9185 | 73.824 | 57.912 | 128.9185 | 116.824 | 117.0129 | 107.5537 | 102.1953 | 120.3584 | 103.0064 | 111.7361 | 115.0944 | 70.0193 | 97.8991 | 71.7232 | 92.7361 | 97.4528 | 100.6416 | 84.5601 | 91.3648 | 226.6416 | 223.9871 | 218 | 213.7361 | 164.8991 | 176.3584 | 183.9936 | 244.4463 | 317.7167 | 211.8927 | 228.4463 | 308.0064
2539 | 29.9234 | 171.4159 | 217.1236 | 204.8318 | 160.6371 | 148.9735 | 151.8318 | 154.6992 | 169.3273 | 143.7788 | 119.4424 | 129.8318 | 147.3894 | 192.4515 | 157.991 | 145.3629 | 205.3984 | 169.4159 | 135.6106 | 147.9645 | 170.7257 | 143.9469 | 312.6371 | 296.0886 | 146.469 | 211.4159 | 220.5576 | 215.1151 | 241.2037 | 102.8939 | 119.469 | 109.3098 | 125.9645 | 136.5576 | 122.9735 | 132.7523 | 100.7788 | 234.6196 | 208.9645 | 232.7698 | 241.4955 | 128.7788 | 131.0265 | 134.6106 | 107.8674 | 71.6371 | 71.8408 | 78.0531 | 112.522 | 183.6902 | 171.3984 | 208.9469 | 154.1947 | 129.8318 | 103.7788 | 119.9469 | 113.6106 | 142.0265 | 146.4159 | 187.522 | 160.3008 | 115.0355 | 142.4955 | 160.8053 | 129.6106 | 198.7257 | 163.8318 | 203.8318 | 230.7432 | 155.2743 | 161.8318 | 183.9735 | 182.5485 | 170.5045 | 182 | 192.3273 | 284 | 138.0886 | 151.2743 | 154.8318 | 182.1326 | 142.3629 | 115.0355 | 144.6106 | 139 | 171.991 | 188.6281 | 171.1151 | 132.8849 | 106.8143 | 123.8053 | 105.4159 | 111.8053 | 237.9469 | 230.1151 | 239.9114 | 325.6637 | 246.6992 | 257.7167 | 262.1151 | 334.3804 | 138.1061 | 80.8318 | 66.8674 | 138.1061 | 122.2743 | 132.469 | 118.2212 | 118.0796 | 113 | 108.4955 | 117.3804 | 117.0796 | 95.6371 | 85.3629 | 65.6106 | 102.6637 | 92.2743 | 106.6371 | 104.6106 | 88.7167 | 241.9825 | 201.9114 | 215.6637 | 228.1416 | 154.1151 | 167.8318 | 174.6106 | 229.4955 | 302.4159 | 191.9204 | 215.6371 | 319.3363
2540 | 29.9344 | 164.2393 | 174.0026 | 197.5162 | 154.379 | 153.207 | 149.6908 | 164.6184 | 152.4838 | 142.0323 | 116.9327 | 124.586 | 144.6533 | 216.8629 | 135.4166 | 150.586 | 226.2717 | 170.5188 | 132.7581 | 123.8279 | 180.7581 | 148.0673 | 307.414 | 311.3441 | 134.5511 | 207.9651 | 221.2044 | 223.4463 | 268.9301 | 119.8279 | 110.6559 | 117.4114 | 106.5886 | 145.0699 | 122.2393 | 142.0699 | 100.6882 | 255.7904 | 192.4864 | 207.8279 | 233.8305 | 129.9301 | 125.9651 | 134.4489 | 132.1397 | 69.1048 | 95.379 | 72.4838 | 99.3816 | 178.0725 | 202.414 | 212.0323 | 154.379 | 125.207 | 109.4838 | 124.0673 | 107.6533 | 145.2445 | 146.2768 | 170.4489 | 147.6908 | 136.6882 | 128.207 | 154.8254 | 130.6908 | 206.4812 | 163.5537 | 196.9301 | 207.7581 | 147.9327 | 152.4463 | 194.6235 | 162.2768 | 185.414 | 189.2445 | 174.2419 | 279.2393 | 155 | 144.3467 | 146.2743 | 168.2096 | 146.9651 | 141.4489 | 140.3092 | 138.1721 | 148.7956 | 152.5511 | 181.7232 | 124.9677 | 133.5511 | 122.586 | 101.1371 | 112.6559 | 243.3092 | 243.621 | 222.1721 | 316.414 | 255.1695 | 238.5511 | 281.0026 | 304.3092 | 123.4489 | 78.6908 | 83.2743 | 123.4489 | 113.4838 | 121.1721 | 113.9651 | 103.9301 | 115.0699 | 101.8655 | 93.7257 | 109.3467 | 86.6882 | 84.5834 | 67.1048 | 92.586 | 84.7257 | 101 | 101.5511 | 72.4489 | 266.5136 | 187.2768 | 206.8279 | 239.2768 | 165.5511 | 165.8978 | 177.3467 | 217.8978 | 305.1746 | 205.6559 | 214.9677 | 334.1746
2541 | 29.9454 | 146.8771 | 194.8072 | 184.4916 | 151.7807 | 154.6578 | 158.9036 | 149.7012 | 162.4386 | 126.0699 | 126.2988 | 124.7542 | 131.7277 | 216.7639 | 162.1928 | 150.5349 | 188.1229 | 175.8868 | 127.3422 | 127.8771 | 177.7542 | 139.5783 | 309.4386 | 302.1494 | 133.1663 | 210.7976 | 198.5614 | 203.8072 | 257.4916 | 119.4916 | 115.2458 | 100.6048 | 126.6843 | 150.3687 | 112.9904 | 144.0795 | 90.8506 | 231.3422 | 224.5614 | 209.9036 | 253.8602 | 127.2626 | 121.5614 | 138.4651 | 137.4747 | 80.1229 | 91.6843 | 82.6578 | 120.4386 | 206.8072 | 201.1494 | 201.5518 | 151.5614 | 130.3855 | 122.2892 | 109.8771 | 98.4554 | 168.1759 | 160.0434 | 177.0964 | 158.4386 | 129.4819 | 133.8241 | 140.9205 | 135.5181 | 196.0699 | 181.2193 | 188.3422 | 203.8771 | 157.5181 | 135.4386 | 215.1759 | 174.5084 | 184.1494 | 212.3952 | 178.7807 | 268.2361 | 152.588 | 157.4916 | 135.7277 | 185.4217 | 141.2458 | 147.2193 | 130 | 135.8771 | 173.6313 | 153.7976 | 169.7108 | 142.4651 | 129.3157 | 120.7807 | 97.947 | 117.6843 | 232.7807 | 245.5614 | 217.2458 | 318.6578 | 228.3422 | 243.7446 | 295.4482 | 296.1928 | 126.3687 | 82.4217 | 70.9735 | 126.3687 | 118.6145 | 117.5614 | 112.1928 | 97.0964 | 117.7374 | 129.1494 | 107.7807 | 122.053 | 76.8506 | 64.3855 | 81.6313 | 90.5614 | 96.588 | 101.2193 | 94.9036 | 76.0265 | 231.8072 | 196.8771 | 199.4747 | 244.053 | 158.0265 | 175.0265 | 191.3687 | 220.4482 | 326.0964 | 211.3422 | 228.0795 | 343.4747
2542 | 29.9564 | 151.621 | 192.0945 | 178.0208 | 152.8525 | 151.2108 | 155.4631 | 167.4528 | 154.9688 | 137 | 109.9367 | 129.8422 | 145.2316 | 188.4528 | 167.9159 | 153.6843 | 184.5369 | 152.5576 | 127.7788 | 129.1475 | 184.4631 | 157.4424 | 303.5898 | 293.8525 | 146.8422 | 232.6843 | 203.0208 | 197.8525 | 252.1786 | 109.0841 | 109.9263 | 112.3686 | 122.9263 | 134.2845 | 137.1267 | 121.6106 | 97.9896 | 230.8525 | 217.2108 | 206.6947 | 243.5784 | 139.2845 | 123.242 | 134.621 | 102.7788 | 80.0104 | 83.0633 | 79.4424 | 118.0633 | 194.8318 | 191.9263 | 226.4424 | 144.6739 | 142.5265 | 131.2212 | 112.7684 | 120.1371 | 152.8525 | 172.2108 | 184.4735 | 156.2949 | 143.9055 | 150.3053 | 161.7684 | 120.8629 | 200.7477 | 178.5265 | 194.1049 | 203.9896 | 138 | 133.9896 | 194.0633 | 181.4424 | 174.6947 | 187.6636 | 176.6106 | 294.9263 | 150.4839 | 145.6947 | 148.3053 | 170.3364 | 133.379 | 144.0633 | 127.4528 | 135.2949 | 180.9159 | 170.8214 | 164.7155 | 133.758 | 127.7477 | 115.8318 | 112.9263 | 114.8525 | 230.6106 | 233.811 | 214.242 | 311.9688 | 223.9159 | 270.8847 | 271.4839 | 298.2108 | 121.1682 | 65.7155 | 75.4528 | 121.1682 | 109.4839 | 113.9159 | 118.8422 | 103.5472 | 102.7051 | 120.621 | 105.6106 | 109.8525 | 85.1475 | 78.379 | 90.5369 | 89.4631 | 93.7892 | 102.6947 | 90.3053 | 69.9263 | 227.0104 | 201.621 | 173.3468 | 226.0633 | 145.6739 | 174.947 | 179.9159 | 197.6417 | 321.8951 | 210.1578 | 215.568 | 309.242
2543 | 29.9674 | 156.2684 | 208.0246 | 197.0246 | 160.4632 | 130.7807 | 159.9263 | 160.2193 | 130.3649 | 139.9263 | 121.7807 | 129.6579 | 144.0491 | 186.0965 | 155.879 | 145.7561 | 183.9754 | 170.2684 | 134.5614 | 127.8772 | 189.414 | 139 | 284.8526 | 291.7088 | 143.2439 | 224.0246 | 221.7807 | 202.293 | 279.6579 | 114.7807 | 113.4877 | 98.9509 | 132.3403 | 113.2193 | 115.1456 | 120.9017 | 87.5123 | 234.8053 | 208.2421 | 207.0491 | 264.1983 | 124.3158 | 133.5123 | 138.7807 | 112.4877 | 83.9035 | 73.414 | 66.6088 | 107.4386 | 182.4632 | 194.7561 | 210.6824 | 129.9263 | 131.4877 | 116.879 | 115.9017 | 103.0246 | 162.1702 | 155.1947 | 192.3421 | 146.4386 | 133.8281 | 148.9754 | 158.0737 | 138.707 | 174.3912 | 167.4877 | 218.0737 | 202.0473 | 143.121 | 129.6088 | 189.5351 | 174.9491 | 180.1702 | 160.9754 | 177.8526 | 290.6842 | 170.293 | 147.7561 | 143.3176 | 201.414 | 130.6824 | 132.2193 | 118.0246 | 127.1456 | 172.293 | 152.3649 | 186.2684 | 123.2439 | 110.414 | 100.7807 | 114.7807 | 118.3176 | 227.9509 | 209.2666 | 225.2439 | 296.6315 | 219.4386 | 240.3403 | 286.1719 | 283.8772 | 131.3421 | 81.9035 | 70.6579 | 131.3421 | 126.8544 | 110.414 | 114.5123 | 110.293 | 113.5368 | 124.5368 | 105.6333 | 115.7561 | 78.0246 | 73.4877 | 87.5368 | 87.586 | 109.0983 | 104.7561 | 93.8526 | 75.9263 | 236.2684 | 208.4632 | 208.5614 | 212.5123 | 134.0965 | 192.1474 | 173.7316 | 216.6597 | 301.0965 | 213.7561 | 237.9035 | 325.121
2544 | 29.9784 | 152.4631 | 201.4139 | 193.4877 | 157.5716 | 130.5861 | 167.9754 | 160.67 | 144.4385 | 148.4877 | 115.9508 | 139.4631 | 143.8915 | 204.2315 | 137.4385 | 146.5369 | 190.8423 | 168 | 126.4877 | 140.6946 | 197.1823 | 144.8915 | 310.9407 | 279.3198 | 143.4877 | 219.2069 | 214.67 | 194.1823 | 282.547 | 104.8524 | 114.4877 | 105.2562 | 148.1577 | 120.0738 | 124.9262 | 132.2315 | 85.7438 | 216.2171 | 228.7294 | 211.246 | 221.8075 | 145.2069 | 130.7192 | 135.5123 | 119.123 | 67.8177 | 82.9754 | 81.4385 | 111.1577 | 189.3054 | 192.7192 | 215.9262 | 146.4284 | 132.2315 | 107.1476 | 126.4631 | 107.6845 | 173.5369 | 160.2069 | 184.0738 | 154.5123 | 147.3155 | 148.4139 | 154.3792 | 139.1331 | 157.7685 | 172.5861 | 205.4139 | 224.9508 | 153.877 | 144.1823 | 204.4732 | 204.5716 | 188.9754 | 167.5861 | 191.9016 | 272.0984 | 161.1577 | 146.2315 | 134.4877 | 205.5962 | 142.8423 | 137.0246 | 129.6006 | 142.5615 | 172.3792 | 167.2069 | 176.5716 | 123.7438 | 125.8669 | 99.8177 | 108.9508 | 112.5615 | 239.8915 | 234.1577 | 229.8423 | 333.2909 | 225.9754 | 248.7294 | 263.2069 | 301.5615 | 126.4038 | 66.5861 | 80.7192 | 126.4038 | 113.4877 | 115.3647 | 116.0738 | 107.0492 | 110.5369 | 117.1823 | 117.6946 | 113.7192 | 73.2069 | 75 | 82.2315 | 82.3792 | 100.5861 | 104 | 109.6946 | 86.0246 | 236.3054 | 215.3054 | 204.33 | 218.4284 | 151.7192 | 173.7438 | 175.7438 | 207.246 | 313.8524 | 213.2562 | 216.6946 | 338.6946
2545 | 29.9894 | 156.3217 | 191.2684 | 194.9529 | 142.4426 | 139.6578 | 161.5103 | 154.7582 | 134.3156 | 153.711 | 108.1209 | 126.9468 | 161.2684 | 202 | 138.8586 | 149.1209 | 202.8053 | 167.1947 | 129.0266 | 134.5369 | 182.5574 | 159.584 | 326.7521 | 303.7726 | 140.1209 | 214.8053 | 203.3894 | 192.1741 | 265.2479 | 93.4426 | 113 | 107.0738 | 136.4631 | 133.2951 | 117.8791 | 130.2684 | 91.1741 | 195.4221 | 201.6373 | 231.6046 | 254.1681 | 135.7049 | 125.3894 | 130.289 | 127.5574 | 80.5635 | 80.8053 | 74.8053 | 102.9529 | 192.2418 | 193.5635 | 204.0471 | 156.7255 | 124.3627 | 127.416 | 124.9529 | 125.4365 | 173.4365 | 157.416 | 190.0471 | 159.7582 | 124.7316 | 137.7316 | 161.9468 | 130.6373 | 168.8586 | 177.8996 | 195.8053 | 222.4897 | 136.3156 | 136.8053 | 176.3688 | 181.9262 | 186.5369 | 176.6578 | 176.2418 | 281.3156 | 152.416 | 150.9795 | 135.6844 | 183.5103 | 147.0205 | 138.3894 | 158.5779 | 144.6312 | 185.8525 | 162 | 162.2479 | 126.2213 | 135.4365 | 109.3422 | 105.1475 | 118.9262 | 248.1414 | 220.0471 | 235.8996 | 314.4426 | 234.0061 | 229.9589 | 260.416 | 306.3156 | 138.8259 | 77.2684 | 73.7787 | 138.8259 | 111.4631 | 100.0738 | 126.0738 | 111.3894 | 108.0205 | 109.2684 | 106.7049 | 112.416 | 68.2684 | 73.3894 | 78.9262 | 102.2951 | 108.3156 | 102.6578 | 101.6578 | 85.2418 | 249.4426 | 221.7316 | 208.8996 | 245.1004 | 146.3894 | 171.6578 | 172.3156 | 226.5308 | 301.1004 | 212.1209 | 211.6106 | 329.3156

## QA QC
Number of Samples | 30
Number of Replicates | 3
Mineralogy
SQ Mineralogy | Sample Name | Mineralogy | Analcime | Anthrophylite | Carlfrancisite | Ganophyllite | Loughlinite | Muscovite | Nontronite | Palygorskite | Sanidine | Sepiolite | Vermiculite | TOTAL
Sample | 1 Etosha 18 RIP 0-2 cm | Regular | 39 |  | 10 |  | 9 | 21 |  | 4 |  | 18 |  | 101
|  | 400C | 39 | 7 | 11 |  | 5 | 24 |  | 6 |  | 7 |  | 99
|  | 500C | 13 | 12 | 16 |  |  | 51 |  | 7 |  |  |  | 99
|  | Glycolated | 36 |  |  |  | 18 | 34 | TR | 6 |  | 7 |  | 101
Replicate | 1R Etosha 18 RIP 0-2 cm | Regular | 7 |  | 20 | 16 | 15 | 27 |  | 4 |  | 8 | 2 | 99
|  | 400C | 8 |  | 18 | 6 | 7 | 51 |  | 7 |  | 3 |  | 100
|  | 500C | 31 |  | 48 |  |  |  |  | 21 |  | TR |  | 100
|  | Glycolated | 9 |  | 56 | 8 | 16 |  |  | 7 |  | 3 |  | 99
Mineralogy
SQ Mineralogy | Sample Name | Mineralogy | Analcime | Biotite | Carlfrancisite | Ganophyllite | Microcline | Palygorskite | Phologopite | Rutile | Sanidine | Sepiolite | TOTAL
Sample | 13 Etosha 22 sub 0.1-10 cm S2 | Regular | 4 | 12 | 25 | 7 | 9 | 12 | 26 |  |  | 4 | 99
|  | 400C | 19 | 20 | 26 |  |  |  | 29 | 5 |  | 1 | 100
|  | 500C | 12 | 25 | 14 |  | 16 |  | 28 | 4 |  |  | 99
|  | Glycolated | 6 | 12 | 24 | 3 |  | 11 | 39 | 5 |  |  | 100
Replicate | 13R Etosha 22 sub 0.1-10 cm S2 | Regular | 4 | 7 | 21 | 3 | 21 | 6 | 24 | 4 | 10 |  | 100
|  | 400C | 6 | 19 | 18 | 1 | 21 | 7 | 16 | 5 | 7 |  | 100
|  | 500C | 7 | 16 | 19 |  | 17 | 10 | 21 | 5 | 4 |  | 99
|  | Glycolated | 11 | 13 | 30 | 3 |  | 7 | 33 | 4 |  |  | 101
Mineralogy
SQ Mineralogy | Sample Name | Mineralogy | Analcime | Biotite | Carlfrancisite | Ganophyllite | Hexahydrite | Illite | Loughlinite | Montmorillonite | Microcline | Orthoclase | Palygorskite | Parnauite | Phologopite | Quartz | Rutile | Sanidine | Sepiolite | Vermiculite | Zeolite | TOTAL
Sample | 25 Etosha 22 sub 5-15 cm S5 | Regular | 8 | 7 | 24 | 7 |  |  |  |  | 13 |  |  | 3 | 21 | 5 | 4 |  | 5 | 2 |  | 99
|  | 400C | 13 | 18 | 15 | 5 |  |  |  |  | 17 |  |  |  |  | 5 | 4 |  | 8 |  | 15 | 100
|  | 500C | TR | 26 | TR |  | TR |  |  |  | 14 |  | 9 |  | 41 |  |  | 10 |  |  |  | 100
|  | Glycolated | 10 |  | 18 | 2 |  | 37 |  |  |  | 10 |  |  | 18 | 4 | [5 empty] | 99
Replicate | 25 Etosha 22 sub 5-15 cm S5 REPLICATE | Regular | 16 | 8 | 15 | 5 |  |  | 8 | TR | 14 |  |  | 3 | 10 |  | 3 | 10 | 4 | 3 |  | 99
|  | 400C | 13 |  | 19 | 3 |  |  | 3 |  | 20 |  |  |  | 25 |  |  | 15 |  |  |  | 98
|  | 500C | 2 | 14 | 19 |  |  |  | 15 |  | 17 |  |  |  | 20 | 4 |  | 10 |  |  |  | 101
|  | Glycolated | 5 | 13 | 16 |  |  |  | 3 |  | 25 |  | 4 |  | 11 | 6 | 4 | 13 |  |  |  | 100
Mineralogy
Sample Name
regular
replicate
Peak Lists
samploe name regular |  | sample name replicate
Mineralogy
Sample Name
regular
replicate
Peak Lists
samploe name regular |  | sample name replicate

## Lab Verification
Measurement Details
Operator Name | J. Leidelmeijer
Sample Name | Si Reference Disc
SOP File Name | Silica SOP
Analysis
Measurement Program | Silicon.xrdmp
Particle Name | Silicon
Start Position | 25.0004
End Position | 119.9854
Step | 0.011
Date(s) Analyzed | 12/24/23; 1/25/24; 2/7/24; 2/26/24; 4/29/24
Result
Reference Code | 01-077-2107
Mineral Name | Silicon
Chemical Formula | Si
Semi Qunat % | 100
Peak List |  | 28.443 2-theta (+ or - 0.2° 2-theta) | 47.283 2-theta (+ or - 0.2° 2-theta) | 56.125 2-theta (+ or - 0.2° 2-theta).
120423 Si Disc
No. | Pos. [°2θ] | FWHM Left [°2θ] | Area calc. | Phase Assignment | h | k | l | Multiplicity | F obs. | F calc. | Area [cts*°2θ] | Derivation | Backgr.[cts] | d-spacing [Å] | Height [cts] | K-A2 / K-A1 Ratio | Matched | Matched by | Rel. Int. [%] | Source | Status | Integral Breadth [°2θ]
1 | 25.6243 | 0.1082 | 0 | [5 empty] | 0 | 0 | 17.46 | K-α1 + K-α2 | 814.77 | 3.4737 | 163.5 |  | FALSE |  | 0.34 | Search Peaks | Included | 0.10677
2 | 28.4221 | 0.0792 | 0 | [5 empty] | 0 | 0 | 5132.29 | Pure K-α1 | 860.92 | 3.1377 | 48601.25 |  | TRUE | 01-077-2107 | 100 | Search Peaks | Included | 0.1056
3 | 28.5062 | 0.0396 | 0 | [5 empty] | 0 | 0 | 1107.04 | Pure K-α2 | 861.67 | 3.1287 | 20966.73 |  | FALSE |  | 43.14 | Search Peaks | Included | 0.0528
4 | 35.5891 | 0.3168 | 0 | [5 empty] | 0 | 0 | 24.97 | Pure K-α1 | 580.51 | 2.5206 | 59.12 |  | FALSE |  | 0.12 | Search Peaks | Included | 0.4224
5 | 42.475 | 0.3168 | 0 | [5 empty] | 0 | 0 | 32.59 | Pure K-α1 | 477.36 | 2.1265 | 77.16 |  | FALSE |  | 0.16 | Search Peaks | Included | 0.4224
6 | 43.5714 | 0.528 | 0 | [5 empty] | 0 | 0 | 32.64 | Pure K-α1 | 465.91 | 2.0755 | 46.36 |  | FALSE |  | 0.1 | Search Peaks | Included | 0.704
7 | 47.2814 | 0.0924 | 0 | [5 empty] | 0 | 0 | 3069.11 | Pure K-α1 | 494.6 | 1.9209 | 24911.62 |  | TRUE | 01-077-2107 | 51.26 | Search Peaks | Included | 0.1232
8 | 47.4083 | 0.066 | 0 | [5 empty] | 0 | 0 | 1104.61 | Pure K-α2 | 493.97 | 1.9161 | 12552.37 |  | FALSE |  | 25.83 | Search Peaks | Included | 0.088
9 | 50.2837 | 0.1584 | 0 | [5 empty] | 0 | 0 | 12.05 | Pure K-α1 | 400.61 | 1.8131 | 57.07 |  | FALSE |  | 0.12 | Search Peaks | Included | 0.2112
10 | 56.1006 | 0.0924 | 0 | [5 empty] | 0 | 0 | 1630.24 | Pure K-α1 | 396.73 | 1.6381 | 13232.45 |  | TRUE | 01-077-2107 | 27.23 | Search Peaks | Included | 0.1232
11 | 56.2501 | 0.0792 | 0 | [5 empty] | 0 | 0 | 736.01 | Pure K-α2 | 394.09 | 1.6341 | 6969.78 |  | FALSE |  | 14.34 | Search Peaks | Included | 0.1056
12 | 69.107 | 0.1056 | 0 | [5 empty] | 0 | 0 | 407.5 | Pure K-α1 | 281.4 | 1.3581 | 2894.21 |  | TRUE | 01-077-2107 | 5.96 | Search Peaks | Included | 0.1408
13 | 69.2992 | 0.0924 | 0 | [5 empty] | 0 | 0 | 180.28 | Pure K-α2 | 279.4 | 1.3548 | 1463.29 |  | FALSE |  | 3.01 | Search Peaks | Included | 0.1232
14 | 76.3492 | 0.0924 | 0 | [5 empty] | 0 | 0 | 583.47 | Pure K-α1 | 289.38 | 1.2463 | 4735.99 |  | TRUE | 01-077-2107 | 9.74 | Search Peaks | Included | 0.1232
15 | 76.5754 | 0.0924 | 0 | [5 empty] | 0 | 0 | 295.77 | Pure K-α2 | 288.64 | 1.2432 | 2400.73 |  | FALSE |  | 4.94 | Search Peaks | Included | 0.1232
16 | 88.0065 | 0.0924 | 0 | [5 empty] | 0 | 0 | 718.71 | Pure K-α1 | 289.24 | 1.1088 | 5833.66 |  | TRUE | 01-077-2107 | 12 | Search Peaks | Included | 0.1232
17 | 88.2838 | 0.1188 | 0 | [5 empty] | 0 | 0 | 467.44 | Pure K-α2 | 289.95 | 1.1061 | 2951 |  | FALSE |  | 6.07 | Search Peaks | Included | 0.1584
18 | 94.9297 | 0.1056 | 0 | [5 empty] | 0 | 0 | 425.94 | Pure K-α1 | 270.04 | 1.0454 | 3025.16 |  | TRUE | 01-077-2107 | 6.22 | Search Peaks | Included | 0.1408
19 | 95.2514 | 0.1188 | 0 | [5 empty] | 0 | 0 | 233.58 | Pure K-α2 | 266.43 | 1.0427 | 1474.59 |  | FALSE |  | 3.03 | Search Peaks | Included | 0.1584
20 | 106.6805 | 0.1056 | 0 | [5 empty] | 0 | 0 | 243.38 | Pure K-α1 | 268.15 | 0.96024 | 1728.58 |  | TRUE | 01-077-2107 | 3.56 | Search Peaks | Included | 0.1408
21 | 107.0912 | 0.1452 | 0 | [5 empty] | 0 | 0 | 152 | Pure K-α2 | 263.31 | 0.95769 | 785.14 |  | FALSE |  | 1.62 | Search Peaks | Included | 0.1936
22 | 114.0803 | 0.1056 | 0 | [5 empty] | 0 | 0 | 427.37 | Pure K-α1 | 287.61 | 0.91806 | 3035.27 |  | TRUE | 01-077-2107 | 6.25 | Search Peaks | Included | 0.1408
23 | 114.4902 | 0.1188 | 0 | [5 empty] | 0 | 0 | 234.09 | Pure K-α2 | 290.34 | 0.91594 | 1477.82 |  | FALSE |  | 3.04 | Search Peaks | Included | 0.1584
012524 Si Disc
No. | Pos. [°2θ] | FWHM Left [°2θ] | Area calc. | Phase Assignment | h | k | l | Multiplicity | F obs. | F calc. | Area [cts*°2θ] | Derivation | Backgr.[cts] | d-spacing [Å] | Height [cts] | K-A2 / K-A1 Ratio | Matched | Matched by | Rel. Int. [%] | Source | Status | Integral Breadth [°2θ]
1 | 25.6095 | 0.1299 | 0 | [5 empty] | 0 | 0 | 16.58 | K-α1 + K-α2 | 811.73 | 3.4756 | 129.37 |  | FALSE |  | 0.27 | Search Peaks | Included | 0.12813
2 | 28.4285 | 0.0792 | 0 | [5 empty] | 0 | 0 | 5027.06 | Pure K-α1 | 847.05 | 3.1371 | 47604.77 |  | TRUE | 01-077-2107 | 100 | Search Peaks | Included | 0.1056
3 | 28.5118 | 0.0396 | 0 | [5 empty] | 0 | 0 | 1126.63 | Pure K-α2 | 848.79 | 3.1281 | 21337.61 |  | FALSE |  | 44.82 | Search Peaks | Included | 0.0528
4 | 42.4864 | 0.1056 | 0 | [5 empty] | 0 | 0 | 14.97 | Pure K-α1 | 469.79 | 2.126 | 106.34 |  | FALSE |  | 0.22 | Search Peaks | Included | 0.1408
5 | 43.8296 | 0.6336 | 0 | [5 empty] | 0 | 0 | 37.15 | Pure K-α1 | 465.52 | 2.0639 | 43.98 |  | FALSE |  | 0.09 | Search Peaks | Included | 0.8448
6 | 47.2861 | 0.0792 | 0 | [5 empty] | 0 | 0 | 2610.92 | Pure K-α1 | 479.93 | 1.9208 | 24724.59 |  | TRUE | 01-077-2107 | 51.94 | Search Peaks | Included | 0.1056
7 | 47.4138 | 0.066 | 0 | [5 empty] | 0 | 0 | 1084.42 | Pure K-α2 | 480.52 | 1.9159 | 12322.98 |  | FALSE |  | 25.89 | Search Peaks | Included | 0.088
8 | 50.2767 | 0.264 | 0 | [5 empty] | 0 | 0 | 11.66 | Pure K-α1 | 405.23 | 1.8133 | 33.12 |  | FALSE |  | 0.07 | Search Peaks | Included | 0.352
9 | 56.105 | 0.0792 | 0 | [5 empty] | 0 | 0 | 1403.67 | Pure K-α1 | 384.6 | 1.638 | 13292.34 |  | TRUE | 01-077-2107 | 27.92 | Search Peaks | Included | 0.1056
10 | 56.2606 | 0.066 | 0 | [5 empty] | 0 | 0 | 606.33 | Pure K-α2 | 383.1 | 1.6338 | 6890.06 |  | FALSE |  | 14.47 | Search Peaks | Included | 0.088
11 | 69.1081 | 0.0924 | 0 | [5 empty] | 0 | 0 | 347.53 | Pure K-α1 | 270.63 | 1.3581 | 2820.84 |  | TRUE | 01-077-2107 | 5.93 | Search Peaks | Included | 0.1232
12 | 69.3064 | 0.0924 | 0 | [5 empty] | 0 | 0 | 190.68 | Pure K-α2 | 270.58 | 1.3547 | 1547.75 |  | FALSE |  | 3.25 | Search Peaks | Included | 0.1232
13 | 76.3601 | 0.0924 | 0 | [5 empty] | 0 | 0 | 563.96 | Pure K-α1 | 277.1 | 1.2462 | 4577.58 |  | TRUE | 01-077-2107 | 9.62 | Search Peaks | Included | 0.1232
14 | 76.5867 | 0.0924 | 0 | [5 empty] | 0 | 0 | 283.51 | Pure K-α2 | 278.18 | 1.243 | 2301.18 |  | FALSE |  | 4.83 | Search Peaks | Included | 0.1232
15 | 88.0122 | 0.1188 | 0 | [5 empty] | 0 | 0 | 906.18 | Pure K-α1 | 277.62 | 1.1088 | 5720.81 |  | TRUE | 01-077-2107 | 12.02 | Search Peaks | Included | 0.1584
16 | 88.2875 | 0.0924 | 0 | [5 empty] | 0 | 0 | 352.79 | Pure K-α2 | 276.65 | 1.106 | 2863.57 |  | FALSE |  | 6.02 | Search Peaks | Included | 0.1232
17 | 94.9386 | 0.1188 | 0 | [5 empty] | 0 | 0 | 455.42 | Pure K-α1 | 264.35 | 1.0453 | 2875.14 |  | TRUE | 01-077-2107 | 6.04 | Search Peaks | Included | 0.1584
18 | 95.2594 | 0.1056 | 0 | [5 empty] | 0 | 0 | 199.96 | Pure K-α2 | 266.42 | 1.0426 | 1420.21 |  | FALSE |  | 2.98 | Search Peaks | Included | 0.1408
19 | 106.6832 | 0.1188 | 0 | [5 empty] | 0 | 0 | 274.17 | Pure K-α1 | 265.59 | 0.96022 | 1730.88 |  | TRUE | 01-077-2107 | 3.64 | Search Peaks | Included | 0.1584
20 | 107.0708 | 0.1056 | 0 | [5 empty] | 0 | 0 | 119.52 | Pure K-α2 | 264.25 | 0.95782 | 848.85 |  | FALSE |  | 1.78 | Search Peaks | Included | 0.1408
21 | 114.0677 | 0.1188 | 0 | [5 empty] | 0 | 0 | 478.26 | Pure K-α1 | 277.48 | 0.91812 | 3019.29 |  | TRUE | 01-077-2107 | 6.34 | Search Peaks | Included | 0.1584
22 | 114.5201 | 0.1056 | 0 | [5 empty] | 0 | 0 | 215.8 | Pure K-α2 | 278.06 | 0.91579 | 1532.64 |  | FALSE |  | 3.22 | Search Peaks | Included | 0.1408
020724 Si Disc
No. | Pos. [°2θ] | FWHM Left [°2θ] | Area calc. | Phase Assignment | h | k | l | Multiplicity | F obs. | F calc. | Area [cts*°2θ] | Derivation | Backgr.[cts] | d-spacing [Å] | Height [cts] | K-A2 / K-A1 Ratio | Matched | Matched by | Rel. Int. [%] | Source | Status | Integral Breadth [°2θ]
1 | 25.6292 | 0.0866 | 0 | [5 empty] | 0 | 0 | 10.83 | K-α1 + K-α2 | 829.75 | 3.473 | 126.73 |  | FALSE |  | 0.27 | Search Peaks | Included | 0.085419
2 | 28.4334 | 0.0792 | 0 | [5 empty] | 0 | 0 | 5031.96 | Pure K-α1 | 845.4 | 3.1365 | 47651.09 |  | TRUE | 01-077-2107 | 100 | Search Peaks | Included | 0.1056
3 | 28.5171 | 0.0396 | 0 | [5 empty] | 0 | 0 | 1090.94 | Pure K-α2 | 845.19 | 3.1275 | 20661.74 |  | FALSE |  | 43.36 | Search Peaks | Included | 0.0528
4 | 35.4643 | 0.6336 | 0 | [5 empty] | 0 | 0 | 23.34 | Pure K-α1 | 591.73 | 2.5292 | 27.63 |  | FALSE |  | 0.06 | Search Peaks | Included | 0.8448
5 | 42.4973 | 0.1584 | 0 | [5 empty] | 0 | 0 | 19.29 | Pure K-α1 | 459.14 | 2.1254 | 91.36 |  | FALSE |  | 0.19 | Search Peaks | Included | 0.2112
6 | 43.7214 | 0.6336 | 0 | [5 empty] | 0 | 0 | 33.58 | Pure K-α1 | 463.77 | 2.0687 | 39.75 |  | FALSE |  | 0.08 | Search Peaks | Included | 0.8448
7 | 47.2899 | 0.0792 | 0 | [5 empty] | 0 | 0 | 2609.87 | Pure K-α1 | 494.97 | 1.9206 | 24714.67 |  | TRUE | 01-077-2107 | 51.87 | Search Peaks | Included | 0.1056
8 | 47.4171 | 0.066 | 0 | [5 empty] | 0 | 0 | 1102.26 | Pure K-α2 | 495.18 | 1.9158 | 12525.63 |  | FALSE |  | 26.29 | Search Peaks | Included | 0.088
9 | 50.2882 | 0.3168 | 0 | [5 empty] | 0 | 0 | 16.61 | Pure K-α1 | 394.26 | 1.8129 | 39.33 |  | FALSE |  | 0.08 | Search Peaks | Included | 0.4224
10 | 56.1096 | 0.0792 | 0 | [5 empty] | 0 | 0 | 1406.3 | Pure K-α1 | 373.38 | 1.6378 | 13317.25 |  | TRUE | 01-077-2107 | 27.95 | Search Peaks | Included | 0.1056
11 | 56.2608 | 0.066 | 0 | [5 empty] | 0 | 0 | 602.03 | Pure K-α2 | 372.56 | 1.6338 | 6841.25 |  | FALSE |  | 14.36 | Search Peaks | Included | 0.088
12 | 69.1169 | 0.0924 | 0 | [5 empty] | 0 | 0 | 352.98 | Pure K-α1 | 290.56 | 1.358 | 2865.08 |  | TRUE | 01-077-2107 | 6.01 | Search Peaks | Included | 0.1232
13 | 69.3104 | 0.0792 | 0 | [5 empty] | 0 | 0 | 158.45 | Pure K-α2 | 290.35 | 1.3546 | 1500.44 |  | FALSE |  | 3.15 | Search Peaks | Included | 0.1056
14 | 76.3606 | 0.0924 | 0 | [5 empty] | 0 | 0 | 572.95 | Pure K-α1 | 276.08 | 1.2462 | 4650.61 |  | TRUE | 01-077-2107 | 9.76 | Search Peaks | Included | 0.1232
15 | 76.5826 | 0.0924 | 0 | [5 empty] | 0 | 0 | 289.97 | Pure K-α2 | 273.18 | 1.2431 | 2353.63 |  | FALSE |  | 4.94 | Search Peaks | Included | 0.1232
16 | 88.0129 | 0.1188 | 0 | [5 empty] | 0 | 0 | 911.18 | Pure K-α1 | 284.89 | 1.1088 | 5752.38 |  | TRUE | 01-077-2107 | 12.07 | Search Peaks | Included | 0.1584
17 | 88.2932 | 0.1188 | 0 | [5 empty] | 0 | 0 | 442.83 | Pure K-α2 | 285.11 | 1.106 | 2795.64 |  | FALSE |  | 5.87 | Search Peaks | Included | 0.1584
18 | 94.9387 | 0.0924 | 0 | [5 empty] | 0 | 0 | 363.54 | Pure K-α1 | 254.06 | 1.0453 | 2950.79 |  | TRUE | 01-077-2107 | 6.19 | Search Peaks | Included | 0.1232
19 | 95.2542 | 0.0924 | 0 | [5 empty] | 0 | 0 | 178.76 | Pure K-α2 | 256.94 | 1.0427 | 1451.01 |  | FALSE |  | 3.05 | Search Peaks | Included | 0.1232
20 | 106.7069 | 0.132 | 0 | [5 empty] | 0 | 0 | 285.19 | Pure K-α1 | 251.69 | 0.96007 | 1620.38 |  | TRUE | 01-077-2107 | 3.4 | Search Peaks | Included | 0.176
21 | 107.0514 | 0.1056 | 0 | [5 empty] | 0 | 0 | 116.51 | Pure K-α2 | 250.3 | 0.95794 | 827.52 |  | FALSE |  | 1.74 | Search Peaks | Included | 0.1408
22 | 114.0732 | 0.1452 | 0 | [5 empty] | 0 | 0 | 596.99 | Pure K-α1 | 288.53 | 0.9181 | 3083.64 |  | TRUE | 01-077-2107 | 6.47 | Search Peaks | Included | 0.1936
23 | 114.5086 | 0.1188 | 0 | [5 empty] | 0 | 0 | 241.24 | Pure K-α2 | 286.31 | 0.91585 | 1522.98 |  | FALSE |  | 3.2 | Search Peaks | Included | 0.1584
022624 Si Disc
No. | Pos. [°2θ] | FWHM Left [°2θ] | Area calc. | Phase Assignment | h | k | l | Multiplicity | F obs. | F calc. | Area [cts*°2θ] | Derivation | Backgr.[cts] | d-spacing [Å] | Height [cts] | K-A2 / K-A1 Ratio | Matched | Matched by | Rel. Int. [%] | Source | Status | Integral Breadth [°2θ]
1 | 25.6348 | 0.0866 | 0 | [5 empty] | 0 | 0 | 9.94 | K-α1 + K-α2 | 854 | 3.4722 | 116.38 |  | FALSE |  | 0.24 | Search Peaks | Included | 0.085419
2 | 28.4332 | 0.0792 | 0 | [5 empty] | 0 | 0 | 5069.92 | Pure K-α1 | 838.69 | 3.1365 | 48010.59 |  | TRUE | 01-077-2107 | 100 | Search Peaks | Included | 0.1056
3 | 28.5153 | 0.0396 | 0 | [5 empty] | 0 | 0 | 1108.97 | Pure K-α2 | 838.29 | 3.1277 | 21003.21 |  | FALSE |  | 43.75 | Search Peaks | Included | 0.0528
4 | 42.4837 | 0.0792 | 0 | [5 empty] | 0 | 0 | 11.44 | Pure K-α1 | 472.77 | 2.1261 | 108.35 |  | FALSE |  | 0.23 | Search Peaks | Included | 0.1056
5 | 43.6177 | 0.6336 | 0 | [5 empty] | 0 | 0 | 25.78 | Pure K-α1 | 473.07 | 2.0734 | 30.52 |  | FALSE |  | 0.06 | Search Peaks | Included | 0.8448
6 | 47.2903 | 0.0792 | 0 | [5 empty] | 0 | 0 | 2658.09 | Pure K-α1 | 500.43 | 1.9206 | 25171.3 |  | TRUE | 01-077-2107 | 52.43 | Search Peaks | Included | 0.1056
7 | 47.4176 | 0.066 | 0 | [5 empty] | 0 | 0 | 1088.92 | Pure K-α2 | 497.71 | 1.9157 | 12374.1 |  | FALSE |  | 25.77 | Search Peaks | Included | 0.088
8 | 56.1094 | 0.0792 | 0 | [5 empty] | 0 | 0 | 1395.59 | Pure K-α1 | 377.1 | 1.6378 | 13215.78 |  | TRUE | 01-077-2107 | 27.53 | Search Peaks | Included | 0.1056
9 | 56.264 | 0.0792 | 0 | [5 empty] | 0 | 0 | 703.04 | Pure K-α2 | 377.37 | 1.6337 | 6657.54 |  | FALSE |  | 13.87 | Search Peaks | Included | 0.1056
10 | 69.1141 | 0.1056 | 0 | [5 empty] | 0 | 0 | 403.17 | Pure K-α1 | 281.16 | 1.358 | 2863.43 |  | TRUE | 01-077-2107 | 5.96 | Search Peaks | Included | 0.1408
11 | 69.3081 | 0.0792 | 0 | [5 empty] | 0 | 0 | 153.78 | Pure K-α2 | 281.35 | 1.3547 | 1456.26 |  | FALSE |  | 3.03 | Search Peaks | Included | 0.1056
12 | 76.3631 | 0.1056 | 0 | [5 empty] | 0 | 0 | 647.23 | Pure K-α1 | 274.94 | 1.2461 | 4596.8 |  | TRUE | 01-077-2107 | 9.57 | Search Peaks | Included | 0.1408
13 | 76.5884 | 0.0924 | 0 | [5 empty] | 0 | 0 | 291.4 | Pure K-α2 | 272.53 | 1.243 | 2365.27 |  | FALSE |  | 4.93 | Search Peaks | Included | 0.1232
14 | 88.0166 | 0.1056 | 0 | [5 empty] | 0 | 0 | 800.47 | Pure K-α1 | 274.63 | 1.1087 | 5685.12 |  | TRUE | 01-077-2107 | 11.84 | Search Peaks | Included | 0.1408
15 | 88.2907 | 0.1056 | 0 | [5 empty] | 0 | 0 | 413.84 | Pure K-α2 | 274.06 | 1.106 | 2939.18 |  | FALSE |  | 6.12 | Search Peaks | Included | 0.1408
16 | 94.9253 | 0.1056 | 0 | [5 empty] | 0 | 0 | 396.44 | Pure K-α1 | 252.58 | 1.0454 | 2815.64 |  | TRUE | 01-077-2107 | 5.86 | Search Peaks | Included | 0.1408
17 | 95.2417 | 0.1056 | 0 | [5 empty] | 0 | 0 | 205.5 | Pure K-α2 | 250.87 | 1.0428 | 1459.5 |  | FALSE |  | 3.04 | Search Peaks | Included | 0.1408
18 | 106.6964 | 0.1452 | 0 | [5 empty] | 0 | 0 | 312.5 | Pure K-α1 | 271.93 | 0.96014 | 1614.13 |  | TRUE | 01-077-2107 | 3.36 | Search Peaks | Included | 0.1936
19 | 107.0742 | 0.132 | 0 | [5 empty] | 0 | 0 | 144.64 | Pure K-α2 | 267.64 | 0.9578 | 821.82 |  | FALSE |  | 1.71 | Search Peaks | Included | 0.176
20 | 114.0679 | 0.132 | 0 | [5 empty] | 0 | 0 | 532.41 | Pure K-α1 | 303.46 | 0.91812 | 3025.05 |  | TRUE | 01-077-2107 | 6.3 | Search Peaks | Included | 0.176
21 | 114.5172 | 0.1188 | 0 | [5 empty] | 0 | 0 | 240.28 | Pure K-α2 | 300.62 | 0.9158 | 1516.95 |  | FALSE |  | 3.16 | Search Peaks | Included | 0.1584
042924 Si Disc
No. | Pos. [°2θ] | FWHM Left [°2θ] | Area calc. | Phase Assignment | h | k | l | Multiplicity | F obs. | F calc. | Area [cts*°2θ] | Derivation | Backgr.[cts] | d-spacing [Å] | Height [cts] | K-A2 / K-A1 Ratio | Matched | Matched by | Rel. Int. [%] | Source | Status | Integral Breadth [°2θ]
1 | 25.6109 | 0.1732 | 0 | [5 empty] | 0 | 0 | 21.72 | K-α1 + K-α2 | 799.25 | 3.4754 | 127.15 |  | FALSE |  | 0.27 | Search Peaks | Included | 0.17084
2 | 28.4234 | 0.066 | 0 | [5 empty] | 0 | 0 | 4124.61 | Pure K-α1 | 818.18 | 3.1376 | 46870.55 |  | TRUE | 01-077-2107 | 100 | Search Peaks | Included | 0.088
3 | 28.5066 | 0.0396 | 0 | [5 empty] | 0 | 0 | 1093.96 | Pure K-α2 | 817.53 | 3.1286 | 20718.85 |  | FALSE |  | 44.2 | Search Peaks | Included | 0.0528
4 | 35.5762 | 0.6336 | 0 | [5 empty] | 0 | 0 | 33.14 | Pure K-α1 | 576.47 | 2.5215 | 39.23 |  | FALSE |  | 0.08 | Search Peaks | Included | 0.8448
5 | 40.5238 | 1.056 | 0 | [5 empty] | 0 | 0 | 32.45 | Pure K-α1 | 486.93 | 2.2243 | 23.05 |  | FALSE |  | 0.05 | Search Peaks | Included | 1.408
6 | 42.4747 | 0.1584 | 0 | [5 empty] | 0 | 0 | 17.3 | Pure K-α1 | 464.86 | 2.1265 | 81.92 |  | FALSE |  | 0.17 | Search Peaks | Included | 0.2112
7 | 43.5657 | 0.528 | 0 | [5 empty] | 0 | 0 | 46.77 | Pure K-α1 | 454.28 | 2.0758 | 66.44 |  | FALSE |  | 0.14 | Search Peaks | Included | 0.704
8 | 45.7236 | 0.528 | 0 | [5 empty] | 0 | 0 | 30.58 | Pure K-α1 | 442.48 | 1.9827 | 43.43 |  | FALSE |  | 0.09 | Search Peaks | Included | 0.704
9 | 47.2796 | 0.0792 | 0 | [5 empty] | 0 | 0 | 2580.15 | Pure K-α1 | 471.59 | 1.921 | 24433.22 |  | TRUE | 01-077-2107 | 52.13 | Search Peaks | Included | 0.1056
10 | 47.4076 | 0.066 | 0 | [5 empty] | 0 | 0 | 1095.15 | Pure K-α2 | 471.26 | 1.9161 | 12444.87 |  | FALSE |  | 26.55 | Search Peaks | Included | 0.088
11 | 56.0993 | 0.0924 | 0 | [5 empty] | 0 | 0 | 1615.11 | Pure K-α1 | 379.96 | 1.6381 | 13109.67 |  | TRUE | 01-077-2107 | 27.97 | Search Peaks | Included | 0.1232
12 | 56.254 | 0.0792 | 0 | [5 empty] | 0 | 0 | 687.78 | Pure K-α2 | 378.21 | 1.634 | 6513.05 |  | FALSE |  | 13.9 | Search Peaks | Included | 0.1056
13 | 69.1062 | 0.0792 | 0 | [5 empty] | 0 | 0 | 292.29 | Pure K-α1 | 276.12 | 1.3581 | 2767.91 |  | TRUE | 01-077-2107 | 5.91 | Search Peaks | Included | 0.1056
14 | 69.3064 | 0.0924 | 0 | [5 empty] | 0 | 0 | 174.41 | Pure K-α2 | 275.48 | 1.3547 | 1415.66 |  | FALSE |  | 3.02 | Search Peaks | Included | 0.1232
15 | 76.3555 | 0.0924 | 0 | [5 empty] | 0 | 0 | 558.13 | Pure K-α1 | 274.78 | 1.2462 | 4530.24 |  | TRUE | 01-077-2107 | 9.67 | Search Peaks | Included | 0.1232
16 | 76.5768 | 0.0924 | 0 | [5 empty] | 0 | 0 | 278.16 | Pure K-α2 | 274.03 | 1.2432 | 2257.78 |  | FALSE |  | 4.82 | Search Peaks | Included | 0.1232
17 | 88.0054 | 0.1188 | 0 | [5 empty] | 0 | 0 | 895.15 | Pure K-α1 | 270.16 | 1.1088 | 5651.22 |  | TRUE | 01-077-2107 | 12.06 | Search Peaks | Included | 0.1584
18 | 88.2859 | 0.1188 | 0 | [5 empty] | 0 | 0 | 447.96 | Pure K-α2 | 269.42 | 1.106 | 2828.02 |  | FALSE |  | 6.03 | Search Peaks | Included | 0.1584
19 | 94.9277 | 0.1188 | 0 | [5 empty] | 0 | 0 | 453.73 | Pure K-α1 | 264.65 | 1.0454 | 2864.46 |  | TRUE | 01-077-2107 | 6.11 | Search Peaks | Included | 0.1584
20 | 95.2401 | 0.1056 | 0 | [5 empty] | 0 | 0 | 208.2 | Pure K-α2 | 264.04 | 1.0428 | 1478.68 |  | FALSE |  | 3.15 | Search Peaks | Included | 0.1408
21 | 106.6882 | 0.1188 | 0 | [5 empty] | 0 | 0 | 271.79 | Pure K-α1 | 253.23 | 0.96019 | 1715.82 |  | TRUE | 01-077-2107 | 3.66 | Search Peaks | Included | 0.1584
22 | 107.0783 | 0.1188 | 0 | [5 empty] | 0 | 0 | 129.95 | Pure K-α2 | 254.75 | 0.95777 | 820.42 |  | FALSE |  | 1.75 | Search Peaks | Included | 0.1584
23 | 114.0666 | 0.1452 | 0 | [5 empty] | 0 | 0 | 581.38 | Pure K-α1 | 287.74 | 0.91813 | 3003.02 |  | TRUE | 01-077-2107 | 6.41 | Search Peaks | Included | 0.1936
24 | 114.4969 | 0.1188 | 0 | [5 empty] | 0 | 0 | 229.38 | Pure K-α2 | 287.32 | 0.91591 | 1448.12 |  | FALSE |  | 3.09 | Search Peaks | Included | 0.1584

## Data Review
Mineralogy and Microscopy Laboratory
Geosciences and Environmental Change Science Center |  |  | Analyst |  |  | Jenifer Leidelmeijer |  | Date | 2024-08-08
United States Geological Survey
X-ray Diffraction
Data Review |  |  | Reviewer |  |  | Michelle Leung |  | Date | 2024-08-08
Project | Namibia Clays
# of Samples | 30
Date of Analysis | 10/10/23 - 4/25/24
Analyst Name | Jenifer Leidelmeijer
Date of Review | 2024-08-08
Reviewer Name | Michelle Leung
|  | Analyst (First Level) |  | Analyst Comments (explain issues found) | Reviewer (Second Level) |  | Reviewer Comments (Note any issues found during review) | Actions Taken (explain resolution taken for first and second level review)
|  | YES | NO |  | YES | NO
From the Lab Verification Tab: Was a Manufacturer Reference Material analyzed with the sample batch? (i.e. Silicon Disc) |  | YES |  |  | YES
From the Lab Verification Tab: Are the Panalytical Aeris Silicon Disc (Manufacturer Reference Material) results within the accepted criteria? |  | YES |  |  | YES
Peak position at: | 28.443 2-theta (+ or - 0.2° 2-theta) | YES |  |  | YES
Peak position at: | 47.283 2-theta (+ or - 0.2° 2-theta) | YES |  |  | YES
Peak position at: | 56.125 2-theta (+ or - 0.2° 2-theta). | YES |  |  | YES
From the QA QC Tab: Are replicates required for this work (more than 6 samples)? (if no, skip the next two question). |  | YES |  |  | YES
From the QA QC Tab: Were sufficient replicates analyzed (at least 10% for batches > 6 samples)? |  | YES |  |  | YES
From the QA QC Tab: Were there sufficient sample to make and analyze replicates? |  | YES |  |  | YES
From the QA QC Tab: Do replicate scans contain similar mineralogy when compared? (Natural Variability is expected) |  | YES |  |  | YES
Scan Data Tab: Is anchor scan data included in the data package? |  | YES |  |  | YES
Scan Data Tab: Is anchor scan data provided and complete for ALL samples |  | YES |  |  | YES
MM Lab + Aeris Scan Paramters Tab: Is the analyst name provided in the data package? |  | YES |  |  | YES
MM Lab + Aeris Scan Paramters Tab: Are the measurement parameters included in the data package? |  | YES |  |  | YES